## Cal Specs and Observations
Run #1 of the USA Celani Cell #1
Conditions:
| Gas: | Helium | Percent: | 100
| Starting Pressure: | 3.4567 | bar
| Start Date: | 2012-10-30 22:12
| End Date: | 2012-10-31 10:56
| Glass Type: | Quartz
| Vacuum before Gas Loading: | 0.08 | Torr | 0.00010666 | bar
Observations:
| Max T_Mica: | 335.4332 | Celcius
| Max Pressure: | 5.9182 | bar
| Max T_Mica Rise: | 315.8578 | Celcius

## 10_30_2012
Date | T_Board | T_Mica | T_GlassIn | T_Well | T_GlassOut | Pressure | Current Blue | Voltage Blue | Power Blue | Impedance Blue | Current Red | Voltage Red | Power Red | Impedance Red | T_Ambient | Pressure | Total Power | T_Mica Rise | T_GlassIn Rise | T_Well Rise | T_GlassOut Rise
2012-10-30 | 26.1458 | 271.8524 | 213.926 | 267.0601 | 192.715 | 81.6714 | 1.3376 | 33.6966 | 45.0734 | 25.1914 | 2.109 | 33.599 | 70.861 | 15.931 | 21.634 | 5.6304 | 115.9344 | 250.2184 | 192.292 | 245.4261 | 171.081
2012-10-30 | 26.1421 | 271.8604 | 213.9381 | 267.0681 | 192.7034 | 81.6795 | 1.3376 | 33.696 | 45.0718 | 25.1914 | 2.109 | 33.5983 | 70.8592 | 15.9308 | 21.6263 | 5.631 | 115.931 | 250.2341 | 192.3118 | 245.4418 | 171.0771
2012-10-30 | 26.1466 | 271.8687 | 213.9426 | 267.0726 | 192.814 | 81.6668 | 1.3376 | 33.6955 | 45.0705 | 25.1913 | 2.109 | 33.5977 | 70.8568 | 15.9308 | 21.6388 | 5.6301 | 115.9273 | 250.2299 | 192.3038 | 245.4338 | 171.1752
2012-10-30 | 26.1443 | 271.8703 | 213.9599 | 267.0781 | 192.7567 | 81.6651 | 1.3375 | 33.6942 | 45.0674 | 25.1912 | 2.1089 | 33.5966 | 70.8521 | 15.9308 | 21.6084 | 5.63 | 115.9195 | 250.2619 | 192.3515 | 245.4697 | 171.1483
2012-10-30 | 26.1451 | 271.8828 | 213.9292 | 267.0322 | 192.7143 | 81.6645 | 1.3375 | 33.6923 | 45.0626 | 25.191 | 2.1088 | 33.5947 | 70.8443 | 15.9307 | 21.6173 | 5.63 | 115.907 | 250.2655 | 192.3119 | 245.4149 | 171.097
2012-10-30 | 26.1458 | 271.9261 | 213.9772 | 267.0873 | 192.8525 | 81.6789 | 1.3374 | 33.6916 | 45.06 | 25.1914 | 2.1087 | 33.5939 | 70.8408 | 15.9308 | 21.5939 | 5.6309 | 115.9008 | 250.3322 | 192.3833 | 245.4934 | 171.2586
2012-10-30 | 26.1496 | 271.8523 | 213.9771 | 267.1106 | 192.9978 | 81.6755 | 1.3374 | 33.6913 | 45.0591 | 25.1914 | 2.1087 | 33.5936 | 70.8396 | 15.9308 | 21.6218 | 5.6307 | 115.8987 | 250.2305 | 192.3553 | 245.4888 | 171.376
2012-10-30 | 26.1473 | 271.8578 | 214.0181 | 267.0811 | 192.9522 | 81.6737 | 1.3374 | 33.6902 | 45.0568 | 25.191 | 2.1087 | 33.5926 | 70.8352 | 15.9308 | 21.6114 | 5.6306 | 115.892 | 250.2464 | 192.4067 | 245.4697 | 171.3408
2012-10-30 | 26.1466 | 271.8183 | 214.0134 | 267.0376 | 192.8376 | 81.6842 | 1.3373 | 33.6893 | 45.0535 | 25.1916 | 2.1086 | 33.5917 | 70.8318 | 15.9308 | 21.6509 | 5.6313 | 115.8853 | 250.1674 | 192.3625 | 245.3867 | 171.1867
2012-10-30 | 26.1473 | 271.8694 | 214.0535 | 267.0966 | 192.8579 | 81.687 | 1.3373 | 33.6881 | 45.0507 | 25.1913 | 2.1086 | 33.5904 | 70.8272 | 15.9305 | 21.6355 | 5.6315 | 115.8779 | 250.2339 | 192.418 | 245.4611 | 171.2224
2012-10-30 | 26.1556 | 271.8661 | 214.0854 | 267.0738 | 192.9016 | 81.6784 | 1.3373 | 33.6879 | 45.0505 | 25.1911 | 2.1085 | 33.5903 | 70.8259 | 15.9307 | 21.6318 | 5.6309 | 115.8764 | 250.2343 | 192.4536 | 245.442 | 171.2698
2012-10-30 | 26.1526 | 271.8864 | 214.0588 | 267.0825 | 192.9811 | 81.6743 | 1.3373 | 33.6872 | 45.0488 | 25.1911 | 2.1085 | 33.5895 | 70.8231 | 15.9306 | 21.5967 | 5.6306 | 115.8719 | 250.2897 | 192.4621 | 245.4858 | 171.3844
2012-10-30 | 26.1578 | 271.8411 | 214.0758 | 267.0954 | 193.0924 | 81.6853 | 1.3372 | 33.6858 | 45.0452 | 25.1911 | 2.1084 | 33.5882 | 70.8183 | 15.9304 | 21.638 | 5.6314 | 115.8634 | 250.2031 | 192.4378 | 245.4574 | 171.4544
2012-10-30 00:01 | 26.1578 | 271.8567 | 214.0916 | 267.111 | 192.9863 | 81.6784 | 1.3372 | 33.6855 | 45.0443 | 25.191 | 2.1084 | 33.5878 | 70.8165 | 15.9305 | 21.65 | 5.6309 | 115.8609 | 250.2067 | 192.4416 | 245.461 | 171.3363
2012-10-30 00:01 | 26.1556 | 271.8661 | 214.1563 | 267.101 | 192.8662 | 81.6818 | 1.3372 | 33.6842 | 45.0411 | 25.1909 | 2.1083 | 33.5867 | 70.8123 | 15.9303 | 21.6398 | 5.6311 | 115.8534 | 250.2263 | 192.5165 | 245.4612 | 171.2264
2012-10-30 00:01 | 26.1503 | 271.9074 | 214.1353 | 267.0841 | 192.8492 | 81.6749 | 1.3371 | 33.6837 | 45.0396 | 25.191 | 2.1083 | 33.586 | 70.8096 | 15.9303 | 21.6225 | 5.6307 | 115.8492 | 250.2849 | 192.5128 | 245.4616 | 171.2267
2012-10-30 00:01 | 26.1601 | 271.921 | 214.149 | 267.1016 | 192.8747 | 81.6824 | 1.3371 | 33.6827 | 45.0371 | 25.1909 | 2.1082 | 33.585 | 70.8054 | 15.9303 | 21.6724 | 5.6312 | 115.8425 | 250.2486 | 192.4766 | 245.4292 | 171.2023
2012-10-30 00:01 | 26.1548 | 271.8808 | 214.1792 | 267.0924 | 192.9047 | 81.6685 | 1.3371 | 33.6818 | 45.0347 | 25.1909 | 2.1082 | 33.5843 | 70.8022 | 15.9303 | 21.6872 | 5.6302 | 115.8369 | 250.1936 | 192.492 | 245.4052 | 171.2175
2012-10-30 00:01 | 26.1563 | 271.9017 | 214.1334 | 267.1289 | 192.8434 | 81.6818 | 1.337 | 33.6804 | 45.0308 | 25.191 | 2.1081 | 33.5827 | 70.7956 | 15.9304 | 21.6445 | 5.6311 | 115.8265 | 250.2572 | 192.4889 | 245.4844 | 171.1989
2012-10-30 00:01 | 26.1488 | 271.902 | 214.1456 | 267.1331 | 192.6591 | 81.6818 | 1.337 | 33.6804 | 45.0308 | 25.191 | 2.1081 | 33.5828 | 70.7965 | 15.9302 | 21.6531 | 5.6311 | 115.8273 | 250.2489 | 192.4925 | 245.48 | 171.006
2012-10-30 00:01 | 26.1458 | 271.9145 | 214.0993 | 267.1184 | 192.4714 | 81.6749 | 1.337 | 33.6795 | 45.0286 | 25.1908 | 2.108 | 33.5818 | 70.7917 | 15.9304 | 21.6701 | 5.6307 | 115.8203 | 250.2444 | 192.4292 | 245.4483 | 170.8013
2012-10-30 00:01 | 26.1518 | 271.9321 | 214.0738 | 267.1244 | 192.3635 | 81.6761 | 1.337 | 33.6794 | 45.0281 | 25.191 | 2.1081 | 33.5818 | 70.7924 | 15.9302 | 21.6521 | 5.6308 | 115.8205 | 250.28 | 192.4217 | 245.4723 | 170.7114
2012-10-30 00:01 | 26.1481 | 271.9362 | 214.074 | 267.1518 | 192.3912 | 81.8286 | 1.443 | 36.3994 | 52.5251 | 25.2244 | 2.2708 | 36.1871 | 82.1751 | 15.9355 | 21.6403 | 5.6413 | 134.7003 | 250.2959 | 192.4337 | 245.5115 | 170.7509
2012-10-30 00:01 | 26.1496 | 272.2636 | 214.3551 | 267.1417 | 192.7424 | 82.054 | 1.4471 | 36.5173 | 52.8428 | 25.2354 | 2.2846 | 36.4107 | 83.1857 | 15.9371 | 21.694 | 5.6568 | 136.0286 | 250.5696 | 192.6611 | 245.4477 | 171.0484
2012-10-30 00:01 | 26.1488 | 272.6468 | 214.6221 | 267.2536 | 192.8398 | 82.1078 | 1.4468 | 36.5139 | 52.8299 | 25.2369 | 2.2844 | 36.4074 | 83.1701 | 15.9372 | 21.6812 | 5.6605 | 136 | 250.9656 | 192.9409 | 245.5724 | 171.1586
2012-10-30 00:01 | 26.1481 | 273.0262 | 214.8143 | 267.3617 | 193.0198 | 82.139 | 1.4467 | 36.5121 | 52.8221 | 25.2382 | 2.2843 | 36.4057 | 83.1611 | 15.9374 | 21.6203 | 5.6627 | 135.9832 | 251.4059 | 193.194 | 245.7414 | 171.3995
2012-10-30 00:01 | 26.1473 | 273.3589 | 214.9789 | 267.4736 | 193.1566 | 82.1644 | 1.4466 | 36.5111 | 52.8187 | 25.2385 | 2.2842 | 36.4046 | 83.1564 | 15.9374 | 21.5954 | 5.6644 | 135.9751 | 251.7635 | 193.3835 | 245.8782 | 171.5612
2012-10-30 00:01 | 26.1496 | 273.6752 | 215.1584 | 267.6158 | 193.2728 | 82.191 | 1.4466 | 36.5106 | 52.8152 | 25.2395 | 2.2842 | 36.4042 | 83.1537 | 15.9375 | 21.6418 | 5.6662 | 135.9689 | 252.0334 | 193.5166 | 245.974 | 171.631
2012-10-30 00:01 | 26.1443 | 274.0266 | 215.3303 | 267.7971 | 193.2715 | 82.217 | 1.4465 | 36.5097 | 52.8107 | 25.2404 | 2.2841 | 36.4033 | 83.1489 | 15.9377 | 21.6325 | 5.668 | 135.9596 | 252.3941 | 193.6978 | 246.1646 | 171.639
2012-10-30 00:01 | 26.1541 | 274.4163 | 215.5607 | 268.0323 | 193.4817 | 82.2551 | 1.4464 | 36.5091 | 52.8072 | 25.2411 | 2.284 | 36.4025 | 83.1448 | 15.9378 | 21.6343 | 5.6707 | 135.952 | 252.782 | 193.9264 | 246.398 | 171.8474
2012-10-30 00:01 | 26.1488 | 274.6474 | 215.7207 | 268.2407 | 193.6886 | 82.2788 | 1.4463 | 36.5084 | 52.8034 | 25.242 | 2.284 | 36.4021 | 83.1418 | 15.938 | 21.6089 | 5.6723 | 135.9453 | 253.0385 | 194.1118 | 246.6318 | 172.0797
2012-10-30 00:01 | 26.1503 | 274.8892 | 215.8364 | 268.4559 | 193.8158 | 82.2996 | 1.4425 | 36.4114 | 52.525 | 25.2412 | 2.2782 | 36.3102 | 82.7235 | 15.9378 | 21.6385 | 5.6737 | 135.2486 | 253.2507 | 194.1979 | 246.8174 | 172.1773
2012-10-30 00:01 | 26.1481 | 275.1466 | 215.9878 | 268.6634 | 193.8136 | 82.3355 | 1.4421 | 36.402 | 52.4969 | 25.2416 | 2.2773 | 36.296 | 82.6583 | 15.9379 | 21.6243 | 5.6762 | 135.1553 | 253.5223 | 194.3635 | 247.0391 | 172.1893
2012-10-30 00:01 | 26.1601 | 275.3601 | 216.1494 | 268.9124 | 194.2816 | 82.3522 | 1.4417 | 36.3908 | 52.4636 | 25.242 | 2.2766 | 36.2847 | 82.6072 | 15.9379 | 21.6363 | 5.6774 | 135.0708 | 253.7238 | 194.5131 | 247.2761 | 172.6453
2012-10-30 00:01 | 26.1548 | 275.6144 | 216.341 | 269.1207 | 194.5042 | 82.3713 | 1.4416 | 36.3888 | 52.4565 | 25.2427 | 2.2765 | 36.2827 | 82.5968 | 15.938 | 21.639 | 5.6787 | 135.0533 | 253.9754 | 194.702 | 247.4817 | 172.8652
2012-10-30 00:01 | 26.1533 | 275.8531 | 216.5207 | 269.3523 | 193.7717 | 82.4152 | 1.4416 | 36.389 | 52.4568 | 25.2429 | 2.2765 | 36.2829 | 82.5974 | 15.9381 | 21.6134 | 5.6817 | 135.0543 | 254.2397 | 194.9073 | 247.7389 | 172.1583
2012-10-30 00:02 | 26.1458 | 276.0857 | 216.6116 | 269.5584 | 194.3262 | 82.4048 | 1.4416 | 36.3922 | 52.4642 | 25.2438 | 2.2767 | 36.286 | 82.6115 | 15.9382 | 21.638 | 5.681 | 135.0756 | 254.4477 | 194.9736 | 247.9204 | 172.6882
2012-10-30 00:02 | 26.1556 | 276.2737 | 216.7828 | 269.7817 | 195.0868 | 82.4216 | 1.4418 | 36.3981 | 52.4799 | 25.2443 | 2.277 | 36.2919 | 82.6368 | 15.9384 | 21.6559 | 5.6821 | 135.1167 | 254.6178 | 195.1269 | 248.1258 | 173.4309
2012-10-30 00:02 | 26.1481 | 276.5179 | 216.988 | 270.0228 | 194.7294 | 82.4464 | 1.4421 | 36.4047 | 52.4974 | 25.2451 | 2.2774 | 36.2986 | 82.6672 | 15.9384 | 21.6162 | 5.6839 | 135.1647 | 254.9017 | 195.3718 | 248.4066 | 173.1132
2012-10-30 00:02 | 26.1421 | 276.783 | 217.1828 | 270.3002 | 195.3288 | 82.4771 | 1.4422 | 36.4103 | 52.5124 | 25.2457 | 2.2778 | 36.3042 | 82.6927 | 15.9385 | 21.6383 | 5.686 | 135.2051 | 255.1447 | 195.5445 | 248.6619 | 173.6905
2012-10-30 00:02 | 26.1526 | 276.96 | 217.3508 | 270.4931 | 195.3983 | 82.5083 | 1.4435 | 36.4426 | 52.6039 | 25.2465 | 2.2798 | 36.3369 | 82.8405 | 15.9387 | 21.6127 | 5.6881 | 135.4444 | 255.3473 | 195.7381 | 248.8804 | 173.7856
2012-10-30 00:02 | 26.1451 | 277.1887 | 217.5402 | 270.7574 | 195.3279 | 82.5337 | 1.4455 | 36.4952 | 52.7534 | 25.2477 | 2.283 | 36.3887 | 83.0755 | 15.9389 | 21.7055 | 5.6899 | 135.8289 | 255.4832 | 195.8347 | 249.0519 | 173.6224
2012-10-30 00:02 | 26.1354 | 277.4848 | 217.7076 | 270.9922 | 195.3457 | 82.562 | 1.4461 | 36.5117 | 52.7992 | 25.2485 | 2.284 | 36.4049 | 83.149 | 15.9391 | 21.6597 | 5.6918 | 135.9482 | 255.8251 | 196.0479 | 249.3325 | 173.686
2012-10-30 00:02 | 26.1436 | 277.675 | 217.8497 | 271.2604 | 195.4286 | 82.588 | 1.4465 | 36.5232 | 52.8301 | 25.2496 | 2.2847 | 36.4164 | 83.2013 | 15.9391 | 21.676 | 5.6936 | 136.0314 | 255.999 | 196.1737 | 249.5844 | 173.7526
2012-10-30 00:02 | 26.1369 | 277.858 | 217.9769 | 271.4555 | 195.5556 | 82.6175 | 1.4467 | 36.5306 | 52.8504 | 25.2502 | 2.2852 | 36.4238 | 83.235 | 15.9391 | 21.6693 | 5.6957 | 136.0855 | 256.1887 | 196.3076 | 249.7862 | 173.8863
2012-10-30 00:02 | 26.1316 | 278.1004 | 218.1251 | 271.652 | 195.6171 | 82.636 | 1.4469 | 36.536 | 52.8651 | 25.2507 | 2.2855 | 36.4292 | 83.2592 | 15.9392 | 21.68 | 5.6969 | 136.1243 | 256.4204 | 196.4451 | 249.972 | 173.9371
2012-10-30 00:02 | 26.1346 | 278.3085 | 218.2345 | 271.8645 | 195.7223 | 82.6585 | 1.447 | 36.539 | 52.8729 | 25.2511 | 2.2857 | 36.4324 | 83.2737 | 15.9392 | 21.6268 | 5.6985 | 136.1466 | 256.6817 | 196.6077 | 250.2377 | 174.0955
2012-10-30 00:02 | 26.1443 | 278.473 | 218.3544 | 272.1341 | 195.9208 | 82.688 | 1.4471 | 36.5417 | 52.8791 | 25.2519 | 2.2859 | 36.4351 | 83.2859 | 15.9393 | 21.6405 | 5.7005 | 136.1651 | 256.8325 | 196.7139 | 250.4936 | 174.2803
2012-10-30 00:02 | 26.1413 | 278.6944 | 218.5089 | 272.3484 | 196.0514 | 82.6996 | 1.4471 | 36.5421 | 52.8793 | 25.2523 | 2.2859 | 36.4353 | 83.2864 | 15.9394 | 21.6255 | 5.7013 | 136.1657 | 257.0689 | 196.8834 | 250.7229 | 174.4259
2012-10-30 00:02 | 26.1309 | 278.8581 | 218.593 | 272.5552 | 196.1865 | 82.7186 | 1.4471 | 36.5429 | 52.8806 | 25.2527 | 2.2859 | 36.4361 | 83.2889 | 15.9396 | 21.6031 | 5.7026 | 136.1695 | 257.255 | 196.9899 | 250.9521 | 174.5834
2012-10-30 00:02 | 26.1249 | 279.0262 | 218.713 | 272.7509 | 196.1766 | 82.7574 | 1.4471 | 36.5434 | 52.8807 | 25.2534 | 2.2859 | 36.4368 | 83.2921 | 15.9395 | 21.6292 | 5.7053 | 136.1728 | 257.397 | 197.0838 | 251.1217 | 174.5474
2012-10-30 00:02 | 26.1369 | 279.2317 | 218.914 | 272.9607 | 196.5071 | 82.7683 | 1.447 | 36.5415 | 52.874 | 25.2541 | 2.2858 | 36.4349 | 83.2827 | 15.9397 | 21.605 | 5.706 | 136.1568 | 257.6267 | 197.309 | 251.3557 | 174.9021
2012-10-30 00:02 | 26.1398 | 279.4821 | 219.031 | 273.173 | 196.6987 | 82.7787 | 1.4469 | 36.5408 | 52.8711 | 25.2544 | 2.2857 | 36.4341 | 83.2787 | 15.9397 | 21.5919 | 5.7068 | 136.1499 | 257.8902 | 197.4391 | 251.5811 | 175.1068
2012-10-30 00:02 | 26.1369 | 279.6572 | 219.1777 | 273.457 | 196.8807 | 82.8071 | 1.4469 | 36.5407 | 52.8702 | 25.2547 | 2.2857 | 36.434 | 83.2783 | 15.9397 | 21.589 | 5.7087 | 136.1485 | 258.0682 | 197.5887 | 251.868 | 175.2917
2012-10-30 00:02 | 26.1331 | 279.812 | 219.3511 | 273.6238 | 197.0145 | 82.8348 | 1.4468 | 36.5401 | 52.8675 | 25.2552 | 2.2857 | 36.4336 | 83.2765 | 15.9397 | 21.6093 | 5.7106 | 136.144 | 258.2027 | 197.7418 | 252.0145 | 175.4052
2012-10-30 00:02 | 26.1369 | 280.0401 | 219.5045 | 273.7905 | 197.1441 | 82.8475 | 1.4468 | 36.5401 | 52.8665 | 25.2557 | 2.2857 | 36.4334 | 83.2752 | 15.9398 | 21.6331 | 5.7115 | 136.1417 | 258.407 | 197.8714 | 252.1574 | 175.511
2012-10-30 00:02 | 26.1279 | 280.2128 | 219.649 | 274.0335 | 197.1705 | 82.8851 | 1.4467 | 36.5387 | 52.8616 | 25.256 | 2.2856 | 36.4321 | 83.2696 | 15.9397 | 21.5679 | 5.7141 | 136.1312 | 258.6449 | 198.0811 | 252.4656 | 175.6026
2012-10-30 00:02 | 26.1301 | 280.4509 | 219.8127 | 274.2566 | 197.2553 | 82.9047 | 1.4467 | 36.5385 | 52.8601 | 25.2565 | 2.2856 | 36.4319 | 83.2684 | 15.9399 | 21.6263 | 5.7155 | 136.1284 | 258.8246 | 198.1864 | 252.6303 | 175.629
2012-10-30 00:02 | 26.1391 | 280.5256 | 219.9555 | 274.5021 | 197.3233 | 82.9198 | 1.4467 | 36.5383 | 52.8593 | 25.2567 | 2.2856 | 36.4315 | 83.2672 | 15.9397 | 21.6113 | 5.7165 | 136.1264 | 258.9143 | 198.3442 | 252.8908 | 175.712
2012-10-30 00:02 | 26.1294 | 280.7479 | 220.0797 | 274.6706 | 197.3844 | 82.9336 | 1.4466 | 36.538 | 52.8569 | 25.2574 | 2.2856 | 36.4315 | 83.2668 | 15.9397 | 21.5815 | 5.7174 | 136.1237 | 259.1664 | 198.4982 | 253.0891 | 175.8029
2012-10-30 00:02 | 26.1301 | 280.9187 | 220.2378 | 274.8845 | 197.5031 | 82.9504 | 1.4466 | 36.5367 | 52.8527 | 25.2576 | 2.2855 | 36.43 | 83.2602 | 15.9397 | 21.6063 | 5.7186 | 136.1129 | 259.3124 | 198.6315 | 253.2782 | 175.8968
2012-10-30 00:03 | 26.1264 | 281.1199 | 220.3325 | 275.1017 | 197.6567 | 82.9966 | 1.4465 | 36.5349 | 52.846 | 25.2583 | 2.2854 | 36.4283 | 83.252 | 15.9398 | 21.6026 | 5.7218 | 136.098 | 259.5173 | 198.7299 | 253.4991 | 176.0541
2012-10-30 00:03 | 26.1249 | 281.2924 | 220.5082 | 275.321 | 197.8715 | 83.0007 | 1.4464 | 36.5347 | 52.8449 | 25.2585 | 2.2853 | 36.428 | 83.2499 | 15.9399 | 21.589 | 5.7221 | 136.0948 | 259.7034 | 198.9192 | 253.732 | 176.2825
2012-10-30 00:03 | 26.1286 | 281.5087 | 220.6103 | 275.4719 | 198.0876 | 83.0301 | 1.4464 | 36.5335 | 52.8409 | 25.2588 | 2.2853 | 36.4268 | 83.2448 | 15.9399 | 21.6529 | 5.7241 | 136.0857 | 259.8558 | 198.9574 | 253.819 | 176.4347
2012-10-30 00:03 | 26.1256 | 281.6371 | 220.7963 | 275.6588 | 198.0886 | 83.0544 | 1.4463 | 36.5329 | 52.8382 | 25.2592 | 2.2852 | 36.4262 | 83.2421 | 15.9399 | 21.5737 | 5.7258 | 136.0802 | 260.0634 | 199.2226 | 254.0851 | 176.5149
2012-10-30 00:03 | 26.1316 | 281.8556 | 220.9204 | 275.8391 | 198.2322 | 83.0793 | 1.4462 | 36.5318 | 52.8339 | 25.2597 | 2.2851 | 36.425 | 83.2364 | 15.9399 | 21.5716 | 5.7275 | 136.0703 | 260.284 | 199.3488 | 254.2675 | 176.6606
2012-10-30 00:03 | 26.1286 | 281.9802 | 221.0473 | 276.0492 | 198.2686 | 83.0868 | 1.4462 | 36.5302 | 52.8284 | 25.2602 | 2.285 | 36.4235 | 83.229 | 15.94 | 21.6208 | 5.728 | 136.0574 | 260.3594 | 199.4265 | 254.4284 | 176.6478
2012-10-30 00:03 | 26.1361 | 282.177 | 221.1532 | 276.2426 | 198.3587 | 83.0972 | 1.4461 | 36.5286 | 52.8226 | 25.2608 | 2.2849 | 36.4219 | 83.2215 | 15.9401 | 21.5481 | 5.7287 | 136.0441 | 260.6289 | 199.6051 | 254.6945 | 176.8106
2012-10-30 00:03 | 26.1339 | 282.3255 | 221.2454 | 276.3527 | 198.5059 | 83.1342 | 1.446 | 36.5287 | 52.8222 | 25.2611 | 2.2849 | 36.422 | 83.2222 | 15.94 | 21.5619 | 5.7313 | 136.0444 | 260.7636 | 199.6835 | 254.7908 | 176.944
2012-10-30 00:03 | 26.1339 | 282.4955 | 221.3714 | 276.5386 | 198.565 | 83.1515 | 1.446 | 36.5284 | 52.8208 | 25.2614 | 2.2849 | 36.4217 | 83.2209 | 15.94 | 21.5739 | 5.7325 | 136.0417 | 260.9216 | 199.7975 | 254.9647 | 176.9911
2012-10-30 00:03 | 26.1316 | 282.6632 | 221.4518 | 276.6796 | 198.7515 | 83.1648 | 1.446 | 36.5289 | 52.8217 | 25.2616 | 2.2849 | 36.4221 | 83.2218 | 15.9402 | 21.5676 | 5.7334 | 136.0435 | 261.0956 | 199.8842 | 255.112 | 177.1839
2012-10-30 00:03 | 26.1406 | 282.8035 | 221.5946 | 276.9054 | 198.8431 | 83.1769 | 1.4459 | 36.5266 | 52.8142 | 25.262 | 2.2848 | 36.4201 | 83.2126 | 15.9402 | 21.5646 | 5.7342 | 136.0267 | 261.2389 | 200.03 | 255.3408 | 177.2785
2012-10-30 00:03 | 26.1383 | 282.9442 | 221.6828 | 277.0851 | 198.947 | 83.1908 | 1.4459 | 36.5265 | 52.8128 | 25.2625 | 2.2848 | 36.4197 | 83.2109 | 15.9402 | 21.5583 | 5.7352 | 136.0236 | 261.3859 | 200.1245 | 255.5268 | 177.3887
2012-10-30 00:03 | 26.1354 | 283.0185 | 221.7665 | 277.1945 | 198.9952 | 83.2006 | 1.4458 | 36.5253 | 52.8095 | 25.2625 | 2.2847 | 36.4186 | 83.2055 | 15.9402 | 21.5674 | 5.7358 | 136.015 | 261.4511 | 200.1991 | 255.6271 | 177.4278
2012-10-30 00:03 | 26.1354 | 283.1499 | 221.9318 | 277.3184 | 198.9284 | 83.2278 | 1.4458 | 36.5241 | 52.8054 | 25.2627 | 2.2846 | 36.4174 | 83.2005 | 15.9402 | 21.5152 | 5.7377 | 136.0059 | 261.6347 | 200.4166 | 255.8032 | 177.4132
2012-10-30 00:03 | 26.1339 | 283.326 | 222.013 | 277.4717 | 199.1983 | 83.2405 | 1.4457 | 36.524 | 52.8045 | 25.263 | 2.2846 | 36.4173 | 83.1999 | 15.9401 | 21.5539 | 5.7386 | 136.0044 | 261.7721 | 200.4591 | 255.9178 | 177.6444
2012-10-30 00:03 | 26.1331 | 283.499 | 222.1382 | 277.6451 | 199.1818 | 83.2573 | 1.4457 | 36.5234 | 52.8015 | 25.2637 | 2.2846 | 36.4166 | 83.1963 | 15.9403 | 21.517 | 5.7398 | 135.9978 | 261.982 | 200.6212 | 256.1281 | 177.6648
2012-10-30 00:03 | 26.1346 | 283.6627 | 222.2735 | 277.8479 | 199.2777 | 83.2844 | 1.4456 | 36.5213 | 52.7948 | 25.2639 | 2.2844 | 36.4148 | 83.1877 | 15.9403 | 21.5586 | 5.7416 | 135.9825 | 262.1041 | 200.7149 | 256.2893 | 177.7191
2012-10-30 00:03 | 26.1309 | 283.7981 | 222.3603 | 277.9332 | 199.4117 | 83.2995 | 1.4456 | 36.5212 | 52.7937 | 25.2644 | 2.2844 | 36.4145 | 83.1865 | 15.9403 | 21.5589 | 5.7427 | 135.9803 | 262.2392 | 200.8014 | 256.3743 | 177.8528
2012-10-30 00:03 | 26.1346 | 283.9717 | 222.4545 | 278.0918 | 199.4665 | 83.3139 | 1.4455 | 36.5197 | 52.789 | 25.2645 | 2.2844 | 36.4131 | 83.1805 | 15.9402 | 21.5867 | 5.7437 | 135.9695 | 262.385 | 200.8678 | 256.5051 | 177.8798
2012-10-30 00:03 | 26.1376 | 284.0326 | 222.552 | 278.2302 | 199.379 | 83.3243 | 1.4454 | 36.5184 | 52.784 | 25.2651 | 2.2843 | 36.4118 | 83.1742 | 15.9402 | 21.5857 | 5.7444 | 135.9582 | 262.4469 | 200.9663 | 256.6445 | 177.7933
2012-10-30 00:03 | 26.1354 | 284.2157 | 222.6639 | 278.4215 | 199.5224 | 83.3515 | 1.4454 | 36.518 | 52.7824 | 25.2653 | 2.2842 | 36.4114 | 83.172 | 15.9403 | 21.5714 | 5.7463 | 135.9544 | 262.6443 | 201.0925 | 256.8501 | 177.951
2012-10-30 00:03 | 26.1376 | 284.3299 | 222.7567 | 278.5205 | 199.7134 | 83.3619 | 1.4453 | 36.5171 | 52.7797 | 25.2655 | 2.2842 | 36.4106 | 83.1688 | 15.9403 | 21.5576 | 5.747 | 135.9485 | 262.7723 | 201.1991 | 256.9629 | 178.1558
2012-10-30 00:03 | 26.1526 | 284.4877 | 222.8897 | 278.7057 | 199.8347 | 83.3815 | 1.4452 | 36.5151 | 52.7729 | 25.2659 | 2.284 | 36.4085 | 83.1586 | 15.9403 | 21.5766 | 5.7483 | 135.9315 | 262.9111 | 201.3131 | 257.1291 | 178.2581
2012-10-30 00:03 | 26.1443 | 284.5798 | 222.9759 | 278.8793 | 199.9287 | 83.3994 | 1.4437 | 36.4768 | 52.6624 | 25.2658 | 2.2815 | 36.3684 | 82.9758 | 15.9403 | 21.5563 | 5.7496 | 135.6381 | 263.0235 | 201.4196 | 257.323 | 178.3724
2012-10-30 00:03 | 26.1443 | 284.7381 | 223.0507 | 279.0186 | 200.0939 | 83.4174 | 1.4414 | 36.4166 | 52.4901 | 25.2651 | 2.2779 | 36.31 | 82.7109 | 15.9401 | 21.5241 | 5.7508 | 135.2011 | 263.214 | 201.5266 | 257.4945 | 178.5698
2012-10-30 00:04 | 26.1563 | 284.8775 | 223.1847 | 279.1389 | 200.2515 | 83.411 | 1.4409 | 36.4034 | 52.4521 | 25.2651 | 2.2771 | 36.297 | 82.6506 | 15.9402 | 21.5602 | 5.7504 | 135.1027 | 263.3173 | 201.6245 | 257.5787 | 178.6913
2012-10-30 00:04 | 26.1563 | 284.9701 | 223.2673 | 279.282 | 200.102 | 83.4422 | 1.4404 | 36.3925 | 52.4203 | 25.2653 | 2.2764 | 36.2863 | 82.6025 | 15.9401 | 21.5562 | 5.7525 | 135.0228 | 263.4139 | 201.7111 | 257.7258 | 178.5458
2012-10-30 00:04 | 26.1496 | 285.0753 | 223.2921 | 279.4184 | 198.7577 | 83.4491 | 1.4403 | 36.3908 | 52.4147 | 25.2656 | 2.2763 | 36.2846 | 82.5944 | 15.9402 | 21.4853 | 5.753 | 135.0091 | 263.59 | 201.8068 | 257.9331 | 177.2724
2012-10-30 00:04 | 26.1586 | 285.2117 | 223.2263 | 279.5319 | 199.6558 | 83.4549 | 1.4405 | 36.3953 | 52.427 | 25.266 | 2.2766 | 36.2889 | 82.6136 | 15.9403 | 21.5184 | 5.7534 | 135.0406 | 263.6933 | 201.7079 | 258.0135 | 178.1374
2012-10-30 00:04 | 26.1503 | 285.2767 | 223.2652 | 279.6628 | 200.4225 | 83.4728 | 1.4406 | 36.3973 | 52.4326 | 25.2661 | 2.2767 | 36.2909 | 82.6222 | 15.9404 | 21.49 | 5.7546 | 135.0548 | 263.7867 | 201.7752 | 258.1728 | 178.9325
2012-10-30 00:04 | 26.1503 | 285.4273 | 223.3951 | 279.7634 | 200.5524 | 83.4884 | 1.4407 | 36.4017 | 52.4444 | 25.2664 | 2.2769 | 36.2952 | 82.6418 | 15.9404 | 21.5462 | 5.7557 | 135.0862 | 263.8811 | 201.8489 | 258.2172 | 179.0062
2012-10-30 00:04 | 26.1615 | 285.4886 | 223.5165 | 279.9022 | 200.4456 | 83.478 | 1.4419 | 36.4341 | 52.5355 | 25.2675 | 2.2789 | 36.3277 | 82.7888 | 15.9406 | 21.5654 | 5.755 | 135.3244 | 263.9232 | 201.9511 | 258.3368 | 178.8802
2012-10-30 00:04 | 26.1443 | 285.7493 | 223.6174 | 280.1248 | 200.271 | 83.526 | 1.4446 | 36.5047 | 52.736 | 25.2691 | 2.2834 | 36.3984 | 83.1107 | 15.9407 | 21.5803 | 5.7583 | 135.8468 | 264.169 | 202.0371 | 258.5445 | 178.6907
2012-10-30 00:04 | 26.1466 | 285.9638 | 223.7574 | 280.3707 | 200.4425 | 83.5572 | 1.4451 | 36.5169 | 52.7706 | 25.2694 | 2.2841 | 36.4103 | 83.1646 | 15.9408 | 21.6147 | 5.7604 | 135.9352 | 264.3491 | 202.1427 | 258.756 | 178.8278
2012-10-30 00:04 | 26.1398 | 286.092 | 223.8175 | 280.4645 | 200.7859 | 83.5774 | 1.4455 | 36.5267 | 52.7979 | 25.2699 | 2.2847 | 36.4202 | 83.2102 | 15.9407 | 21.5839 | 5.7618 | 136.0081 | 264.5081 | 202.2336 | 258.8806 | 179.202
2012-10-30 00:04 | 26.1331 | 286.1548 | 223.8856 | 280.5351 | 201.1097 | 83.5867 | 1.4457 | 36.5339 | 52.8172 | 25.2707 | 2.2852 | 36.4271 | 83.2417 | 15.9408 | 21.5691 | 5.7625 | 136.059 | 264.5857 | 202.3165 | 258.966 | 179.5406
2012-10-30 00:04 | 26.1361 | 286.2851 | 224.105 | 280.6657 | 201.3527 | 83.6052 | 1.4458 | 36.5377 | 52.828 | 25.2708 | 2.2854 | 36.4309 | 83.2589 | 15.9408 | 21.6083 | 5.7637 | 136.0869 | 264.6768 | 202.4967 | 259.0574 | 179.7444
2012-10-30 00:04 | 26.1294 | 286.3748 | 224.2242 | 280.7325 | 201.4405 | 83.6081 | 1.4459 | 36.5409 | 52.8362 | 25.2713 | 2.2856 | 36.4342 | 83.2726 | 15.941 | 21.5654 | 5.7639 | 136.1089 | 264.8094 | 202.6588 | 259.1671 | 179.8751
2012-10-30 00:04 | 26.1316 | 286.5043 | 224.4153 | 280.8738 | 201.5804 | 83.6271 | 1.4459 | 36.5405 | 52.8345 | 25.2716 | 2.2855 | 36.4335 | 83.2697 | 15.941 | 21.6078 | 5.7653 | 136.1042 | 264.8965 | 202.8075 | 259.266 | 179.9726
2012-10-30 00:04 | 26.1324 | 286.6093 | 224.5066 | 280.9752 | 201.5852 | 83.6479 | 1.4459 | 36.5408 | 52.8355 | 25.2715 | 2.2855 | 36.4338 | 83.271 | 15.941 | 21.5564 | 5.7667 | 136.1065 | 265.0529 | 202.9502 | 259.4188 | 180.0288
2012-10-30 00:04 | 26.1346 | 286.6115 | 224.5442 | 281.0663 | 201.4103 | 83.641 | 1.446 | 36.5421 | 52.8385 | 25.2718 | 2.2856 | 36.4352 | 83.2769 | 15.941 | 21.6068 | 5.7662 | 136.1154 | 265.0047 | 202.9374 | 259.4595 | 179.8035
2012-10-30 00:04 | 26.1361 | 286.775 | 224.6402 | 281.1876 | 201.5023 | 83.6693 | 1.4459 | 36.5415 | 52.8356 | 25.2723 | 2.2856 | 36.4348 | 83.2751 | 15.9411 | 21.6283 | 5.7682 | 136.1107 | 265.1467 | 203.0119 | 259.5593 | 179.874
2012-10-30 00:04 | 26.1294 | 286.7992 | 224.7712 | 281.2969 | 201.5625 | 83.6745 | 1.4459 | 36.5424 | 52.8382 | 25.2724 | 2.2857 | 36.4356 | 83.2792 | 15.941 | 21.6297 | 5.7685 | 136.1173 | 265.1695 | 203.1415 | 259.6672 | 179.9328
2012-10-30 00:04 | 26.1301 | 286.9001 | 224.8152 | 281.3787 | 201.7403 | 83.7052 | 1.4459 | 36.5421 | 52.8364 | 25.2728 | 2.2856 | 36.4354 | 83.2783 | 15.941 | 21.5942 | 5.7706 | 136.1147 | 265.3059 | 203.221 | 259.7845 | 180.1461
2012-10-30 00:04 | 26.1324 | 287.0143 | 224.904 | 281.4429 | 201.7819 | 83.7011 | 1.4459 | 36.5429 | 52.8387 | 25.2728 | 2.2857 | 36.4358 | 83.2801 | 15.941 | 21.5965 | 5.7704 | 136.1188 | 265.4178 | 203.3075 | 259.8464 | 180.1854
2012-10-30 00:04 | 26.1331 | 287.0767 | 224.952 | 281.5286 | 201.7393 | 83.7167 | 1.4459 | 36.5422 | 52.8358 | 25.2732 | 2.2856 | 36.4354 | 83.2776 | 15.9411 | 21.6053 | 5.7714 | 136.1134 | 265.4714 | 203.3467 | 259.9233 | 180.134
2012-10-30 00:04 | 26.1398 | 287.1683 | 225.0216 | 281.6474 | 201.8011 | 83.7283 | 1.4459 | 36.5429 | 52.8379 | 25.2732 | 2.2857 | 36.4359 | 83.2804 | 15.941 | 21.5919 | 5.7722 | 136.1183 | 265.5764 | 203.4297 | 260.0555 | 180.2092
2012-10-30 00:04 | 26.1309 | 287.2828 | 225.115 | 281.739 | 201.9064 | 83.7369 | 1.4459 | 36.5431 | 52.8376 | 25.2736 | 2.2857 | 36.436 | 83.2813 | 15.941 | 21.591 | 5.7728 | 136.1189 | 265.6918 | 203.524 | 260.148 | 180.3154
2012-10-30 00:04 | 26.1413 | 287.4011 | 225.208 | 281.8692 | 202.1057 | 83.7554 | 1.4459 | 36.5428 | 52.8361 | 25.274 | 2.2857 | 36.436 | 83.2807 | 15.9411 | 21.5934 | 5.7741 | 136.1168 | 265.8077 | 203.6146 | 260.2758 | 180.5123
2012-10-30 00:05 | 26.1406 | 287.439 | 225.3018 | 282.0076 | 202.2388 | 83.767 | 1.4458 | 36.5419 | 52.8325 | 25.2744 | 2.2856 | 36.4352 | 83.2765 | 15.9411 | 21.6007 | 5.7749 | 136.109 | 265.8383 | 203.7011 | 260.4069 | 180.6381
2012-10-30 00:05 | 26.1361 | 287.5386 | 225.3878 | 282.0688 | 202.2972 | 83.7786 | 1.4457 | 36.5403 | 52.8273 | 25.2747 | 2.2855 | 36.4335 | 83.2692 | 15.9411 | 21.5842 | 5.7757 | 136.0965 | 265.9544 | 203.8036 | 260.4846 | 180.713
2012-10-30 00:05 | 26.1383 | 287.6565 | 225.4805 | 282.1753 | 202.1971 | 83.7826 | 1.4457 | 36.5396 | 52.8253 | 25.2747 | 2.2855 | 36.4328 | 83.2664 | 15.941 | 21.5663 | 5.776 | 136.0917 | 266.0902 | 203.9142 | 260.609 | 180.6308
2012-10-30 00:05 | 26.1398 | 287.7621 | 225.5882 | 282.2812 | 202.3088 | 83.8173 | 1.4456 | 36.5384 | 52.8209 | 25.2751 | 2.2854 | 36.4315 | 83.2605 | 15.941 | 21.5919 | 5.7784 | 136.0814 | 266.1702 | 203.9963 | 260.6893 | 180.7169
2012-10-30 00:05 | 26.1473 | 287.8583 | 225.7019 | 282.4123 | 202.3557 | 83.8231 | 1.4456 | 36.5375 | 52.8187 | 25.2749 | 2.2854 | 36.4307 | 83.257 | 15.9409 | 21.5793 | 5.7788 | 136.0757 | 266.279 | 204.1226 | 260.833 | 180.7764
2012-10-30 00:05 | 26.1413 | 287.9448 | 225.7274 | 282.4334 | 202.5386 | 83.8254 | 1.4456 | 36.5375 | 52.8176 | 25.2755 | 2.2853 | 36.4306 | 83.2564 | 15.941 | 21.5733 | 5.7789 | 136.074 | 266.3715 | 204.1541 | 260.8601 | 180.9653
2012-10-30 00:05 | 26.1428 | 288.0273 | 225.8312 | 282.601 | 202.5716 | 83.845 | 1.4456 | 36.5379 | 52.8188 | 25.2755 | 2.2854 | 36.4311 | 83.259 | 15.9409 | 21.5828 | 5.7803 | 136.0778 | 266.4445 | 204.2484 | 261.0182 | 180.9888
2012-10-30 00:05 | 26.1436 | 288.1168 | 225.9107 | 282.6713 | 202.6668 | 83.8433 | 1.4455 | 36.5368 | 52.8148 | 25.2759 | 2.2853 | 36.4299 | 83.253 | 15.941 | 21.5636 | 5.7802 | 136.0677 | 266.5532 | 204.3471 | 261.1077 | 181.1032
2012-10-30 00:05 | 26.1563 | 288.2297 | 226.0296 | 282.7845 | 202.7504 | 83.8704 | 1.4455 | 36.5367 | 52.814 | 25.2761 | 2.2853 | 36.4299 | 83.2535 | 15.9409 | 21.6164 | 5.782 | 136.0675 | 266.6133 | 204.4132 | 261.1681 | 181.134
2012-10-30 00:05 | 26.1451 | 288.2994 | 226.1049 | 282.8776 | 202.7982 | 83.882 | 1.4454 | 36.5352 | 52.8096 | 25.2762 | 2.2852 | 36.4285 | 83.2469 | 15.9409 | 21.545 | 5.7828 | 136.0565 | 266.7544 | 204.5599 | 261.3326 | 181.2532
2012-10-30 00:05 | 26.1496 | 288.381 | 226.1645 | 282.9593 | 202.7948 | 83.8913 | 1.4454 | 36.5339 | 52.8048 | 25.2766 | 2.2851 | 36.427 | 83.2401 | 15.941 | 21.5856 | 5.7835 | 136.0449 | 266.7954 | 204.5789 | 261.3737 | 181.2092
2012-10-30 00:05 | 26.1623 | 288.494 | 226.2205 | 283.057 | 202.8233 | 83.9069 | 1.4453 | 36.533 | 52.8022 | 25.2766 | 2.2851 | 36.4261 | 83.2357 | 15.941 | 21.5622 | 5.7845 | 136.0379 | 266.9318 | 204.6583 | 261.4948 | 181.2611
2012-10-30 00:05 | 26.1586 | 288.5712 | 226.2955 | 283.1499 | 202.9652 | 83.9167 | 1.4453 | 36.5317 | 52.7978 | 25.2769 | 2.285 | 36.425 | 83.231 | 15.9409 | 21.5746 | 5.7852 | 136.0289 | 266.9966 | 204.7209 | 261.5753 | 181.3906
2012-10-30 00:05 | 26.1608 | 288.612 | 226.3449 | 283.2602 | 203.2154 | 83.9103 | 1.4452 | 36.5307 | 52.7949 | 25.2769 | 2.2849 | 36.424 | 83.2267 | 15.9409 | 21.5728 | 5.7848 | 136.0215 | 267.0392 | 204.7721 | 261.6874 | 181.6426
2012-10-30 00:05 | 26.163 | 288.7182 | 226.4612 | 283.3397 | 203.3435 | 83.9242 | 1.4452 | 36.5302 | 52.7928 | 25.2772 | 2.2849 | 36.4234 | 83.2236 | 15.9409 | 21.5428 | 5.7857 | 136.0163 | 267.1754 | 204.9184 | 261.7969 | 181.8007
2012-10-30 00:05 | 26.163 | 288.8416 | 226.5084 | 283.4015 | 203.3042 | 83.9421 | 1.4452 | 36.5303 | 52.7931 | 25.2773 | 2.2849 | 36.4234 | 83.2243 | 15.9409 | 21.583 | 5.787 | 136.0173 | 267.2586 | 204.9254 | 261.8185 | 181.7212
2012-10-30 00:05 | 26.1683 | 289.0049 | 226.7379 | 283.6153 | 203.4433 | 83.9664 | 1.4451 | 36.5277 | 52.7847 | 25.2777 | 2.2847 | 36.4208 | 83.2124 | 15.9409 | 21.6164 | 5.7886 | 135.9971 | 267.3885 | 205.1215 | 261.9989 | 181.8269
2012-10-30 00:05 | 26.1668 | 289.1614 | 226.878 | 283.7992 | 203.5048 | 83.9924 | 1.445 | 36.5261 | 52.7785 | 25.2783 | 2.2847 | 36.4194 | 83.2056 | 15.9409 | 21.639 | 5.7904 | 135.9842 | 267.5224 | 205.239 | 262.1602 | 181.8658
2012-10-30 00:05 | 26.1623 | 289.2609 | 226.9443 | 283.9028 | 203.768 | 83.997 | 1.4449 | 36.5246 | 52.7734 | 25.2787 | 2.2846 | 36.4179 | 83.1985 | 15.9409 | 21.6746 | 5.7908 | 135.9719 | 267.5863 | 205.2697 | 262.2282 | 182.0934
2012-10-30 00:05 | 26.169 | 289.3563 | 227.0493 | 283.9675 | 203.8455 | 84.0126 | 1.4448 | 36.5237 | 52.7706 | 25.2788 | 2.2845 | 36.4168 | 83.1945 | 15.9408 | 21.6211 | 5.7918 | 135.9651 | 267.7352 | 205.4282 | 262.3464 | 182.2244
2012-10-30 00:05 | 26.169 | 289.3832 | 227.128 | 284.091 | 203.8613 | 84.019 | 1.4448 | 36.5229 | 52.7689 | 25.2786 | 2.2845 | 36.4162 | 83.1919 | 15.9408 | 21.6131 | 5.7923 | 135.9607 | 267.7701 | 205.5149 | 262.4779 | 182.2482
2012-10-30 00:06 | 26.1728 | 289.6606 | 227.36 | 284.3342 | 203.9123 | 84.0502 | 1.4445 | 36.5149 | 52.7446 | 25.2791 | 2.284 | 36.4088 | 83.1589 | 15.9406 | 21.6249 | 5.7944 | 135.9035 | 268.0357 | 205.7351 | 262.7093 | 182.2874
2012-10-30 00:06 | 26.1698 | 289.8811 | 227.5458 | 284.586 | 204.1848 | 84.0802 | 1.4408 | 36.4196 | 52.4717 | 25.2781 | 2.278 | 36.3127 | 82.7217 | 15.9403 | 21.6339 | 5.7965 | 135.1933 | 268.2472 | 205.9119 | 262.9521 | 182.5509
2012-10-30 00:06 | 26.175 | 289.9672 | 227.6375 | 284.68 | 204.1625 | 84.0641 | 1.4404 | 36.4116 | 52.4486 | 25.2782 | 2.2776 | 36.3053 | 82.6879 | 15.9404 | 21.6351 | 5.7954 | 135.1365 | 268.3321 | 206.0024 | 263.0449 | 182.5274
2012-10-30 00:06 | 26.1645 | 290.0337 | 227.627 | 284.7235 | 204.0339 | 84.0872 | 1.4401 | 36.4027 | 52.4226 | 25.2784 | 2.277 | 36.2962 | 82.6461 | 15.9405 | 21.6206 | 5.797 | 135.0687 | 268.4131 | 206.0064 | 263.1029 | 182.4133
2012-10-30 00:06 | 26.166 | 290.0853 | 227.6088 | 284.7868 | 202.9805 | 84.0918 | 1.4399 | 36.3986 | 52.4103 | 25.2785 | 2.2767 | 36.2919 | 82.6269 | 15.9404 | 21.6542 | 5.7973 | 135.0372 | 268.4311 | 205.9546 | 263.1326 | 181.3263
2012-10-30 00:06 | 26.1675 | 290.1677 | 227.5749 | 284.8655 | 203.8755 | 84.0976 | 1.4399 | 36.3976 | 52.4081 | 25.2783 | 2.2767 | 36.2912 | 82.6242 | 15.9402 | 21.6356 | 5.7977 | 135.0323 | 268.5321 | 205.9393 | 263.2299 | 182.2399
2012-10-30 00:06 | 26.163 | 290.2094 | 227.6334 | 284.9421 | 204.4575 | 84.1022 | 1.4399 | 36.3996 | 52.4134 | 25.2784 | 2.2768 | 36.2932 | 82.6334 | 15.9403 | 21.6673 | 5.798 | 135.0468 | 268.5421 | 205.9661 | 263.2748 | 182.7902
2012-10-30 00:06 | 26.1556 | 290.2829 | 227.7243 | 284.981 | 204.706 | 84.1178 | 1.4401 | 36.4045 | 52.4267 | 25.2789 | 2.2771 | 36.2981 | 82.6559 | 15.9402 | 21.6438 | 5.7991 | 135.0826 | 268.6391 | 206.0805 | 263.3372 | 183.0622
2012-10-30 00:06 | 26.1556 | 290.3369 | 227.8265 | 285.0891 | 204.6233 | 84.1074 | 1.4402 | 36.4079 | 52.4363 | 25.2789 | 2.2773 | 36.3014 | 82.6708 | 15.9402 | 21.6599 | 5.7984 | 135.1071 | 268.677 | 206.1666 | 263.4292 | 182.9634
2012-10-30 00:06 | 26.1518 | 290.3908 | 227.8306 | 285.1277 | 204.3637 | 84.1282 | 1.4403 | 36.41 | 52.4425 | 25.2789 | 2.2775 | 36.3035 | 82.6809 | 15.9401 | 21.6641 | 5.7998 | 135.1235 | 268.7267 | 206.1665 | 263.4636 | 182.6996
2012-10-30 00:06 | 26.1428 | 290.4165 | 227.916 | 285.1959 | 203.8862 | 84.1218 | 1.4404 | 36.4136 | 52.4519 | 25.2793 | 2.2777 | 36.3071 | 82.6972 | 15.9402 | 21.6631 | 5.7994 | 135.149 | 268.7534 | 206.2529 | 263.5328 | 182.2231
2012-10-30 00:06 | 26.1376 | 290.4845 | 227.8125 | 285.2332 | 204.0031 | 84.1299 | 1.4422 | 36.4574 | 52.5772 | 25.2798 | 2.2803 | 36.3484 | 82.8851 | 15.9402 | 21.674 | 5.7999 | 135.4622 | 268.8105 | 206.1385 | 263.5592 | 182.3291
2012-10-30 00:06 | 26.1369 | 290.5916 | 227.8118 | 285.3521 | 204.4432 | 84.1663 | 1.4439 | 36.5028 | 52.7068 | 25.2805 | 2.2832 | 36.3956 | 83.1002 | 15.9403 | 21.6733 | 5.8024 | 135.807 | 268.9183 | 206.1385 | 263.6788 | 182.7699
2012-10-30 00:06 | 26.1361 | 290.6525 | 227.8975 | 285.359 | 204.1984 | 84.1623 | 1.4444 | 36.5165 | 52.7447 | 25.2813 | 2.2841 | 36.4094 | 83.1626 | 15.9404 | 21.6243 | 5.8021 | 135.9074 | 269.0282 | 206.2732 | 263.7347 | 182.5741
2012-10-30 00:06 | 26.1331 | 290.7188 | 227.9142 | 285.4564 | 204.1875 | 84.1588 | 1.4449 | 36.5281 | 52.7785 | 25.2812 | 2.2848 | 36.421 | 83.2158 | 15.9403 | 21.6253 | 5.8019 | 135.9943 | 269.0935 | 206.2889 | 263.8311 | 182.5622
2012-10-30 00:06 | 26.1406 | 290.8149 | 227.9138 | 285.5063 | 203.8683 | 84.1837 | 1.4452 | 36.5367 | 52.8022 | 25.2816 | 2.2854 | 36.4298 | 83.2567 | 15.9403 | 21.6368 | 5.8036 | 136.0589 | 269.1781 | 206.277 | 263.8695 | 182.2315
2012-10-30 00:06 | 26.1271 | 290.8669 | 227.9121 | 285.5314 | 204.2996 | 84.1895 | 1.4454 | 36.5418 | 52.8162 | 25.2821 | 2.2857 | 36.4349 | 83.2795 | 15.9404 | 21.6635 | 5.804 | 136.0957 | 269.2034 | 206.2486 | 263.8679 | 182.6361
2012-10-30 00:06 | 26.1234 | 290.9248 | 227.9477 | 285.628 | 204.5557 | 84.1993 | 1.4455 | 36.5442 | 52.8235 | 25.2819 | 2.2859 | 36.4373 | 83.2904 | 15.9403 | 21.6517 | 5.8047 | 136.1139 | 269.2731 | 206.296 | 263.9763 | 182.904
2012-10-30 00:06 | 26.1181 | 290.9464 | 228.0368 | 285.6845 | 204.8063 | 84.1976 | 1.4455 | 36.547 | 52.8301 | 25.2826 | 2.286 | 36.4398 | 83.3021 | 15.9403 | 21.6143 | 5.8046 | 136.1323 | 269.3321 | 206.4225 | 264.0702 | 183.192
2012-10-30 00:06 | 26.1159 | 291.0174 | 228.0897 | 285.7054 | 204.8946 | 84.2091 | 1.4456 | 36.548 | 52.8328 | 25.2827 | 2.2861 | 36.4413 | 83.3082 | 15.9404 | 21.6483 | 5.8054 | 136.141 | 269.3691 | 206.4414 | 264.0571 | 183.2463
2012-10-30 00:06 | 26.1219 | 291.1081 | 228.2215 | 285.7886 | 205.0148 | 84.2172 | 1.4456 | 36.5481 | 52.8328 | 25.2828 | 2.2861 | 36.4413 | 83.3084 | 15.9404 | 21.6583 | 5.8059 | 136.1412 | 269.4498 | 206.5632 | 264.1303 | 183.3565
2012-10-30 00:06 | 26.1069 | 291.147 | 228.3442 | 285.8739 | 205.0392 | 84.2351 | 1.4456 | 36.5485 | 52.8333 | 25.2831 | 2.2861 | 36.4414 | 83.3091 | 15.9404 | 21.6553 | 5.8072 | 136.1425 | 269.4917 | 206.6889 | 264.2186 | 183.3839
2012-10-30 00:06 | 26.1062 | 291.2118 | 228.3867 | 285.881 | 205.0542 | 84.2409 | 1.4456 | 36.5486 | 52.8332 | 25.2833 | 2.2861 | 36.4414 | 83.3097 | 15.9402 | 21.6426 | 5.8076 | 136.1429 | 269.5692 | 206.7441 | 264.2384 | 183.4116
2012-10-30 00:06 | 26.1137 | 291.3156 | 228.4572 | 285.9695 | 205.0499 | 84.2605 | 1.4455 | 36.5483 | 52.8319 | 25.2835 | 2.2861 | 36.4412 | 83.3076 | 15.9404 | 21.634 | 5.8089 | 136.1395 | 269.6816 | 206.8232 | 264.3355 | 183.4159
2012-10-30 00:06 | 26.1099 | 291.3195 | 228.5359 | 286.0158 | 205.0461 | 84.2449 | 1.4455 | 36.5484 | 52.8318 | 25.2838 | 2.2861 | 36.4415 | 83.3092 | 15.9404 | 21.6543 | 5.8078 | 136.141 | 269.6652 | 206.8816 | 264.3615 | 183.3918
2012-10-30 00:07 | 26.1084 | 291.4142 | 228.6328 | 286.0606 | 205.1981 | 84.2657 | 1.4455 | 36.5489 | 52.8331 | 25.2839 | 2.2862 | 36.4421 | 83.3127 | 15.9403 | 21.6448 | 5.8093 | 136.1458 | 269.7694 | 206.988 | 264.4158 | 183.5533
2012-10-30 00:07 | 26.0979 | 291.4499 | 228.638 | 286.1543 | 205.4002 | 84.2796 | 1.4455 | 36.5485 | 52.8317 | 25.2839 | 2.2861 | 36.4416 | 83.3097 | 15.9404 | 21.6423 | 5.8102 | 136.1414 | 269.8076 | 206.9957 | 264.512 | 183.7579
2012-10-30 00:07 | 26.1084 | 291.5028 | 228.7704 | 286.1764 | 205.4658 | 84.2877 | 1.4454 | 36.5468 | 52.8263 | 25.2841 | 2.286 | 36.4399 | 83.3028 | 15.9402 | 21.6648 | 5.8108 | 136.129 | 269.838 | 207.1056 | 264.5116 | 183.801
2012-10-30 00:07 | 26.1047 | 291.553 | 228.8335 | 286.2614 | 205.474 | 84.3079 | 1.4454 | 36.5459 | 52.8234 | 25.2843 | 2.2859 | 36.4387 | 83.2971 | 15.9403 | 21.6531 | 5.8122 | 136.1205 | 269.8999 | 207.1804 | 264.6083 | 183.8209
2012-10-30 00:07 | 26.1099 | 291.5851 | 228.9253 | 286.2705 | 205.5658 | 84.2883 | 1.4454 | 36.5455 | 52.8222 | 25.2843 | 2.286 | 36.4385 | 83.297 | 15.9401 | 21.6703 | 5.8108 | 136.1192 | 269.9148 | 207.255 | 264.6002 | 183.8955
2012-10-30 00:07 | 26.1039 | 291.6754 | 228.994 | 286.4072 | 205.548 | 84.3172 | 1.4454 | 36.5446 | 52.8199 | 25.2842 | 2.2859 | 36.4378 | 83.2939 | 15.9401 | 21.6242 | 5.8128 | 136.1138 | 270.0512 | 207.3698 | 264.783 | 183.9238
2012-10-30 00:07 | 26.1054 | 291.7385 | 229.0387 | 286.4087 | 205.5495 | 84.3195 | 1.4453 | 36.5439 | 52.8173 | 25.2844 | 2.2858 | 36.4368 | 83.2891 | 15.9402 | 21.6337 | 5.813 | 136.1063 | 270.1048 | 207.405 | 264.775 | 183.9158
2012-10-30 00:07 | 26.1092 | 291.8077 | 229.1133 | 286.4819 | 205.6517 | 84.3155 | 1.4453 | 36.5428 | 52.8142 | 25.2844 | 2.2858 | 36.4358 | 83.2846 | 15.9401 | 21.6215 | 5.8127 | 136.0988 | 270.1862 | 207.4918 | 264.8604 | 184.0302
2012-10-30 00:07 | 26.1024 | 291.8779 | 229.1616 | 286.56 | 205.7158 | 84.3339 | 1.4452 | 36.5421 | 52.8115 | 25.2848 | 2.2858 | 36.435 | 83.2815 | 15.9401 | 21.6307 | 5.814 | 136.0931 | 270.2472 | 207.5309 | 264.9293 | 184.0851
2012-10-30 00:07 | 26.1009 | 291.9188 | 229.2387 | 286.6472 | 205.7615 | 84.3391 | 1.4451 | 36.5399 | 52.8043 | 25.2852 | 2.2856 | 36.433 | 83.272 | 15.9401 | 21.6012 | 5.8143 | 136.0763 | 270.3176 | 207.6375 | 265.046 | 184.1603
2012-10-30 00:07 | 26.0987 | 291.9859 | 229.3427 | 286.6759 | 205.8774 | 84.3519 | 1.4451 | 36.5392 | 52.8024 | 25.285 | 2.2856 | 36.4323 | 83.2687 | 15.9401 | 21.6431 | 5.8152 | 136.0711 | 270.3428 | 207.6996 | 265.0328 | 184.2343
2012-10-30 00:07 | 26.0957 | 292.0329 | 229.4026 | 286.7577 | 205.8783 | 84.3611 | 1.445 | 36.5378 | 52.7978 | 25.2854 | 2.2855 | 36.431 | 83.2631 | 15.94 | 21.6602 | 5.8159 | 136.0609 | 270.3727 | 207.7424 | 265.0975 | 184.2181
2012-10-30 00:07 | 26.0934 | 292.1345 | 229.475 | 286.8326 | 205.8642 | 84.3732 | 1.445 | 36.5376 | 52.7972 | 25.2853 | 2.2855 | 36.4308 | 83.2625 | 15.94 | 21.69 | 5.8167 | 136.0598 | 270.4445 | 207.785 | 265.1426 | 184.1742
2012-10-30 00:07 | 26.0875 | 292.1248 | 229.4966 | 286.846 | 205.9253 | 84.3727 | 1.445 | 36.5378 | 52.7977 | 25.2854 | 2.2855 | 36.4309 | 83.2627 | 15.94 | 21.6439 | 5.8167 | 136.0604 | 270.4809 | 207.8527 | 265.2021 | 184.2814
2012-10-30 00:07 | 26.089 | 292.2033 | 229.5217 | 286.9053 | 205.8362 | 84.3952 | 1.445 | 36.538 | 52.7977 | 25.2856 | 2.2855 | 36.4309 | 83.2629 | 15.94 | 21.6454 | 5.8182 | 136.0605 | 270.5579 | 207.8763 | 265.2599 | 184.1908
2012-10-30 00:07 | 26.0979 | 292.2661 | 229.5581 | 286.9335 | 205.8609 | 84.3883 | 1.445 | 36.537 | 52.7942 | 25.2859 | 2.2854 | 36.43 | 83.2588 | 15.94 | 21.7025 | 5.8177 | 136.0529 | 270.5636 | 207.8556 | 265.231 | 184.1584
2012-10-30 00:07 | 26.0964 | 292.3108 | 229.6156 | 287.0516 | 205.9814 | 84.3946 | 1.4449 | 36.5351 | 52.7891 | 25.2858 | 2.2853 | 36.4282 | 83.2506 | 15.94 | 21.6609 | 5.8182 | 136.0397 | 270.6499 | 207.9547 | 265.3907 | 184.3205
2012-10-30 00:07 | 26.0942 | 292.3317 | 229.6252 | 287.0879 | 206.1052 | 84.4166 | 1.4448 | 36.5344 | 52.7865 | 25.286 | 2.2853 | 36.4275 | 83.2478 | 15.9399 | 21.6506 | 5.8197 | 136.0342 | 270.6811 | 207.9746 | 265.4373 | 184.4546
2012-10-30 00:07 | 26.0927 | 292.411 | 229.7181 | 287.1289 | 206.2533 | 84.4004 | 1.4448 | 36.5332 | 52.7828 | 25.2862 | 2.2852 | 36.4262 | 83.2419 | 15.9399 | 21.6732 | 5.8186 | 136.0247 | 270.7378 | 208.0449 | 265.4557 | 184.5801
2012-10-30 00:07 | 26.0949 | 292.4863 | 229.795 | 287.2082 | 206.1925 | 84.4253 | 1.4447 | 36.5319 | 52.7784 | 25.2865 | 2.2851 | 36.425 | 83.236 | 15.9399 | 21.6714 | 5.8203 | 136.0144 | 270.8149 | 208.1236 | 265.5368 | 184.5211
2012-10-30 00:07 | 26.0949 | 292.4825 | 229.8343 | 287.239 | 206.2909 | 84.4351 | 1.4446 | 36.53 | 52.7724 | 25.2867 | 2.285 | 36.423 | 83.2271 | 15.9399 | 21.6634 | 5.821 | 135.9995 | 270.8191 | 208.1709 | 265.5756 | 184.6275
2012-10-30 00:07 | 26.1024 | 292.6016 | 229.8457 | 287.3584 | 206.2748 | 84.4177 | 1.4446 | 36.5285 | 52.7678 | 25.2869 | 2.2849 | 36.4217 | 83.2211 | 15.9399 | 21.6669 | 5.8198 | 135.9889 | 270.9347 | 208.1788 | 265.6915 | 184.6079
2012-10-30 00:07 | 26.0942 | 292.578 | 229.8926 | 287.3155 | 206.2942 | 84.4339 | 1.4445 | 36.5281 | 52.7666 | 25.2868 | 2.2849 | 36.4212 | 83.2191 | 15.9399 | 21.6747 | 5.8209 | 135.9856 | 270.9033 | 208.2179 | 265.6408 | 184.6195
2012-10-30 00:07 | 26.0949 | 292.6441 | 229.9758 | 287.4049 | 206.3736 | 84.4397 | 1.4445 | 36.5266 | 52.762 | 25.2869 | 2.2848 | 36.4197 | 83.2113 | 15.94 | 21.6594 | 5.8213 | 135.9733 | 270.9847 | 208.3164 | 265.7455 | 184.7142
2012-10-30 00:07 | 26.0987 | 292.721 | 230.0268 | 287.4357 | 206.4247 | 84.4449 | 1.4444 | 36.5258 | 52.7596 | 25.2871 | 2.2848 | 36.4188 | 83.2082 | 15.9399 | 21.6511 | 5.8216 | 135.9677 | 271.0699 | 208.3757 | 265.7846 | 184.7736
2012-10-30 00:07 | 26.1017 | 292.7741 | 230.0809 | 287.5312 | 206.4198 | 84.4553 | 1.4444 | 36.5242 | 52.7546 | 25.2873 | 2.2847 | 36.4174 | 83.2018 | 15.9398 | 21.626 | 5.8223 | 135.9564 | 271.1481 | 208.4549 | 265.9052 | 184.7938
2012-10-30 00:07 | 26.1024 | 292.7594 | 230.1209 | 287.5358 | 206.4008 | 84.4565 | 1.4444 | 36.525 | 52.7569 | 25.2873 | 2.2847 | 36.4179 | 83.2047 | 15.9398 | 21.5986 | 5.8224 | 135.9617 | 271.1608 | 208.5223 | 265.9372 | 184.8022
2012-10-30 00:08 | 26.1032 | 292.8679 | 230.1296 | 287.6021 | 206.4646 | 84.4726 | 1.4444 | 36.5252 | 52.757 | 25.2875 | 2.2847 | 36.4182 | 83.2058 | 15.9398 | 21.6556 | 5.8235 | 135.9628 | 271.2123 | 208.474 | 265.9465 | 184.809
2012-10-30 00:08 | 26.1092 | 292.924 | 230.1946 | 287.666 | 206.4785 | 84.4807 | 1.4444 | 36.5241 | 52.7537 | 25.2875 | 2.2847 | 36.4173 | 83.2022 | 15.9397 | 21.6456 | 5.8241 | 135.9559 | 271.2784 | 208.549 | 266.0204 | 184.8329
2012-10-30 00:08 | 26.1047 | 292.9733 | 230.2176 | 287.7116 | 206.6945 | 84.4935 | 1.4443 | 36.5227 | 52.75 | 25.2874 | 2.2846 | 36.4158 | 83.1958 | 15.9396 | 21.633 | 5.825 | 135.9459 | 271.3403 | 208.5846 | 266.0786 | 185.0615
2012-10-30 00:08 | 26.1107 | 293.0024 | 230.2904 | 287.7908 | 206.791 | 84.4987 | 1.4442 | 36.5213 | 52.7449 | 25.2878 | 2.2845 | 36.4143 | 83.1897 | 15.9395 | 21.6591 | 5.8253 | 135.9347 | 271.3433 | 208.6313 | 266.1317 | 185.1319
2012-10-30 00:08 | 26.1047 | 293.0541 | 230.3434 | 287.8234 | 206.7732 | 84.4992 | 1.4442 | 36.5197 | 52.7403 | 25.2879 | 2.2844 | 36.4128 | 83.1826 | 15.9395 | 21.6491 | 5.8254 | 135.9229 | 271.405 | 208.6943 | 266.1743 | 185.1241
2012-10-30 00:08 | 26.1084 | 293.104 | 230.4061 | 287.9196 | 206.7021 | 84.5039 | 1.4441 | 36.519 | 52.7383 | 25.2879 | 2.2844 | 36.4122 | 83.1793 | 15.9396 | 21.6207 | 5.8257 | 135.9175 | 271.4833 | 208.7854 | 266.2989 | 185.0814
2012-10-30 00:08 | 26.1002 | 293.0958 | 230.4372 | 287.923 | 206.6585 | 84.4992 | 1.4408 | 36.4334 | 52.4935 | 25.2868 | 2.2791 | 36.3269 | 82.7917 | 15.9393 | 21.6165 | 5.8254 | 135.2853 | 271.4793 | 208.8207 | 266.3065 | 185.042
2012-10-30 00:08 | 26.1047 | 293.1465 | 230.4063 | 287.993 | 206.8008 | 84.5044 | 1.4401 | 36.415 | 52.441 | 25.2865 | 2.2779 | 36.3082 | 82.7075 | 15.9391 | 21.6651 | 5.8257 | 135.1485 | 271.4814 | 208.7412 | 266.3279 | 185.1357
2012-10-30 00:08 | 26.1024 | 293.2019 | 230.4354 | 288.0254 | 206.7788 | 84.5148 | 1.4396 | 36.4037 | 52.4085 | 25.2865 | 2.2772 | 36.2972 | 82.6571 | 15.9392 | 21.6508 | 5.8265 | 135.0656 | 271.5511 | 208.7846 | 266.3746 | 185.128
2012-10-30 00:08 | 26.1047 | 293.2427 | 230.5203 | 288.0624 | 206.7811 | 84.4836 | 1.4395 | 36.399 | 52.3954 | 25.2864 | 2.2769 | 36.2925 | 82.636 | 15.9392 | 21.6491 | 5.8243 | 135.0314 | 271.5936 | 208.8712 | 266.4133 | 185.132
2012-10-30 00:08 | 26.1054 | 293.2549 | 230.4817 | 288.0786 | 205.9392 | 84.5125 | 1.4395 | 36.3989 | 52.3947 | 25.2865 | 2.2769 | 36.2922 | 82.6348 | 15.9391 | 21.6699 | 5.8263 | 135.0296 | 271.585 | 208.8118 | 266.4087 | 184.2693
2012-10-30 00:08 | 26.1062 | 293.2981 | 230.4314 | 288.1565 | 206.2511 | 84.4929 | 1.4395 | 36.4009 | 52.4006 | 25.2865 | 2.2771 | 36.2944 | 82.6453 | 15.939 | 21.6666 | 5.8249 | 135.0459 | 271.6315 | 208.7648 | 266.4899 | 184.5845
2012-10-30 00:08 | 26.1002 | 293.3267 | 230.3585 | 288.2083 | 206.371 | 84.5154 | 1.4396 | 36.4036 | 52.4078 | 25.2867 | 2.2772 | 36.2969 | 82.656 | 15.9391 | 21.6727 | 5.8265 | 135.0638 | 271.654 | 208.6858 | 266.5356 | 184.6983
2012-10-30 00:08 | 26.0949 | 293.3599 | 230.3768 | 288.2338 | 206.7792 | 84.5206 | 1.4397 | 36.4063 | 52.4155 | 25.2868 | 2.2774 | 36.3001 | 82.6707 | 15.9391 | 21.6353 | 5.8269 | 135.0863 | 271.7246 | 208.7415 | 266.5985 | 185.1439
2012-10-30 00:08 | 26.1009 | 293.3889 | 230.4379 | 288.2437 | 206.8521 | 84.5015 | 1.4397 | 36.407 | 52.4168 | 25.2871 | 2.2775 | 36.3005 | 82.673 | 15.939 | 21.6975 | 5.8255 | 135.0898 | 271.6914 | 208.7404 | 266.5462 | 185.1546
2012-10-30 00:08 | 26.0957 | 293.4261 | 230.4641 | 288.2423 | 206.9256 | 84.5351 | 1.4399 | 36.4112 | 52.4296 | 25.2868 | 2.2777 | 36.3044 | 82.6914 | 15.9389 | 21.7043 | 5.8278 | 135.121 | 271.7218 | 208.7598 | 266.538 | 185.2213
2012-10-30 00:08 | 26.0867 | 293.3978 | 230.459 | 288.3181 | 207.0781 | 84.5304 | 1.4403 | 36.4194 | 52.4531 | 25.2869 | 2.2781 | 36.3112 | 82.7223 | 15.9389 | 21.6913 | 5.8275 | 135.1754 | 271.7065 | 208.7677 | 266.6268 | 185.3868
2012-10-30 00:08 | 26.0815 | 293.5004 | 230.5167 | 288.3476 | 207.6202 | 84.5333 | 1.4432 | 36.4964 | 52.6723 | 25.2882 | 2.283 | 36.3891 | 83.0764 | 15.9392 | 21.662 | 5.8277 | 135.7487 | 271.8384 | 208.8547 | 266.6856 | 185.9582
2012-10-30 00:08 | 26.071 | 293.4668 | 230.6596 | 288.3718 | 207.2199 | 84.5443 | 1.4438 | 36.51 | 52.7114 | 25.2882 | 2.2839 | 36.4029 | 83.1399 | 15.939 | 21.6956 | 5.8285 | 135.8513 | 271.7712 | 208.964 | 266.6762 | 185.5243
2012-10-30 00:08 | 26.074 | 293.5467 | 230.6547 | 288.3787 | 207.4079 | 84.5501 | 1.4442 | 36.5209 | 52.742 | 25.2887 | 2.2846 | 36.4142 | 83.191 | 15.9391 | 21.6104 | 5.8289 | 135.933 | 271.9363 | 209.0443 | 266.7683 | 185.7975
2012-10-30 00:08 | 26.0807 | 293.615 | 230.6889 | 288.4548 | 207.2138 | 84.5668 | 1.4445 | 36.5303 | 52.769 | 25.2888 | 2.2852 | 36.4236 | 83.2339 | 15.9391 | 21.6732 | 5.83 | 136.0029 | 271.9418 | 209.0157 | 266.7816 | 185.5406
2012-10-30 00:08 | 26.0755 | 293.629 | 230.7034 | 288.465 | 206.996 | 84.5628 | 1.4448 | 36.5377 | 52.7891 | 25.2894 | 2.2856 | 36.4308 | 83.2674 | 15.939 | 21.6199 | 5.8298 | 136.0564 | 272.0091 | 209.0835 | 266.8451 | 185.3761
2012-10-30 00:08 | 26.0755 | 293.6829 | 230.6837 | 288.5074 | 207.0708 | 84.5773 | 1.445 | 36.5438 | 52.8064 | 25.2895 | 2.286 | 36.4367 | 83.2945 | 15.939 | 21.64 | 5.8308 | 136.101 | 272.0429 | 209.0437 | 266.8674 | 185.4308
2012-10-30 00:08 | 26.083 | 293.7519 | 230.7502 | 288.5072 | 207.2161 | 84.5842 | 1.4451 | 36.5469 | 52.8153 | 25.2896 | 2.2862 | 36.4399 | 83.3092 | 15.939 | 21.6555 | 5.8312 | 136.1244 | 272.0964 | 209.0947 | 266.8517 | 185.5606
2012-10-30 00:08 | 26.0755 | 293.7637 | 230.7623 | 288.6115 | 207.4606 | 84.5842 | 1.4452 | 36.5481 | 52.8183 | 25.2897 | 2.2863 | 36.441 | 83.3141 | 15.939 | 21.5998 | 5.8312 | 136.1323 | 272.1639 | 209.1625 | 267.0117 | 185.8608
2012-10-30 00:08 | 26.0747 | 293.8129 | 230.848 | 288.6107 | 207.3338 | 84.5957 | 1.4453 | 36.5497 | 52.8234 | 25.2895 | 2.2864 | 36.4426 | 83.3216 | 15.939 | 21.6392 | 5.832 | 136.145 | 272.1737 | 209.2088 | 266.9715 | 185.6946
2012-10-30 00:09 | 26.0695 | 293.7807 | 230.8389 | 288.6825 | 207.301 | 84.6021 | 1.4452 | 36.5495 | 52.822 | 25.29 | 2.2864 | 36.4425 | 83.3214 | 15.939 | 21.638 | 5.8325 | 136.1434 | 272.1427 | 209.2009 | 267.0445 | 185.663
2012-10-30 00:09 | 26.0687 | 293.88 | 230.8813 | 288.6625 | 207.1428 | 84.605 | 1.4452 | 36.5498 | 52.8226 | 25.2902 | 2.2864 | 36.4426 | 83.3229 | 15.9388 | 21.597 | 5.8327 | 136.1454 | 272.283 | 209.2843 | 267.0655 | 185.5458
2012-10-30 00:09 | 26.071 | 293.9169 | 230.9151 | 288.7419 | 207.2947 | 84.6206 | 1.4452 | 36.5499 | 52.8227 | 25.2901 | 2.2864 | 36.443 | 83.325 | 15.9387 | 21.6194 | 5.8337 | 136.1477 | 272.2975 | 209.2957 | 267.1225 | 185.6753
2012-10-30 00:09 | 26.0643 | 293.9294 | 230.9791 | 288.7737 | 207.1305 | 84.6241 | 1.4452 | 36.5494 | 52.8212 | 25.2903 | 2.2864 | 36.4423 | 83.3215 | 15.9387 | 21.6288 | 5.834 | 136.1426 | 272.3006 | 209.3503 | 267.1449 | 185.5017
2012-10-30 00:09 | 26.059 | 293.9664 | 230.9778 | 288.7645 | 206.9874 | 84.6275 | 1.4452 | 36.5499 | 52.8222 | 25.2904 | 2.2864 | 36.4426 | 83.3229 | 15.9387 | 21.6275 | 5.8342 | 136.1451 | 272.3389 | 209.3503 | 267.137 | 185.3599
2012-10-30 00:09 | 26.068 | 294.0408 | 230.9828 | 288.8352 | 207.1814 | 84.6397 | 1.4452 | 36.5497 | 52.8214 | 25.2905 | 2.2864 | 36.4424 | 83.3227 | 15.9386 | 21.6284 | 5.8351 | 136.1441 | 272.4124 | 209.3544 | 267.2068 | 185.553
2012-10-30 00:09 | 26.062 | 294.0502 | 230.9847 | 288.8832 | 207.3054 | 84.6426 | 1.4451 | 36.5477 | 52.8157 | 25.2905 | 2.2863 | 36.4405 | 83.3144 | 15.9386 | 21.6265 | 5.8353 | 136.1301 | 272.4237 | 209.3582 | 267.2567 | 185.6789
2012-10-30 00:09 | 26.0575 | 294.1073 | 231.0392 | 288.9172 | 207.5922 | 84.6454 | 1.4451 | 36.5472 | 52.8142 | 25.2905 | 2.2863 | 36.44 | 83.312 | 15.9386 | 21.638 | 5.8355 | 136.1261 | 272.4693 | 209.4012 | 267.2792 | 185.9542
2012-10-30 00:09 | 26.062 | 294.1348 | 231.1183 | 288.9602 | 207.6912 | 84.6368 | 1.4451 | 36.5467 | 52.8123 | 25.2907 | 2.2863 | 36.4396 | 83.3105 | 15.9385 | 21.6586 | 5.8349 | 136.1228 | 272.4762 | 209.4597 | 267.3016 | 186.0326
2012-10-30 00:09 | 26.059 | 294.178 | 231.1782 | 288.9842 | 207.5819 | 84.6622 | 1.445 | 36.546 | 52.8105 | 25.2907 | 2.2862 | 36.439 | 83.308 | 15.9385 | 21.6475 | 5.8366 | 136.1185 | 272.5305 | 209.5307 | 267.3367 | 185.9344
2012-10-30 00:09 | 26.0568 | 294.1989 | 231.2232 | 288.9936 | 207.6033 | 84.6576 | 1.445 | 36.5454 | 52.8086 | 25.2907 | 2.2862 | 36.4384 | 83.3056 | 15.9384 | 21.6534 | 5.8363 | 136.1142 | 272.5455 | 209.5698 | 267.3402 | 185.9499
2012-10-30 00:09 | 26.0583 | 294.2273 | 231.2483 | 289.0914 | 207.6285 | 84.6657 | 1.445 | 36.5442 | 52.8053 | 25.2906 | 2.2861 | 36.4374 | 83.3008 | 15.9384 | 21.6388 | 5.8369 | 136.1061 | 272.5885 | 209.6095 | 267.4526 | 185.9897
2012-10-30 00:09 | 26.0628 | 294.2664 | 231.355 | 289.0998 | 207.8456 | 84.668 | 1.4449 | 36.5432 | 52.8023 | 25.2907 | 2.2861 | 36.4361 | 83.2951 | 15.9384 | 21.6634 | 5.837 | 136.0974 | 272.603 | 209.6916 | 267.4364 | 186.1822
2012-10-30 00:09 | 26.0598 | 294.3096 | 231.3363 | 289.1584 | 207.8465 | 84.6732 | 1.4449 | 36.5421 | 52.7984 | 25.291 | 2.286 | 36.4351 | 83.2904 | 15.9384 | 21.6604 | 5.8374 | 136.0887 | 272.6492 | 209.6759 | 267.498 | 186.1861
2012-10-30 00:09 | 26.0643 | 294.3718 | 231.3997 | 289.1822 | 207.9652 | 84.6639 | 1.4448 | 36.5406 | 52.7941 | 25.291 | 2.2859 | 36.4337 | 83.2834 | 15.9385 | 21.6127 | 5.8367 | 136.0775 | 272.7591 | 209.787 | 267.5695 | 186.3525
2012-10-30 00:09 | 26.0643 | 294.3833 | 231.4901 | 289.2246 | 207.8195 | 84.6859 | 1.4447 | 36.5393 | 52.7895 | 25.2914 | 2.2858 | 36.4322 | 83.2776 | 15.9383 | 21.6528 | 5.8382 | 136.0671 | 272.7305 | 209.8373 | 267.5718 | 186.1667
2012-10-30 00:09 | 26.065 | 294.3994 | 231.5301 | 289.2484 | 207.7139 | 84.6917 | 1.4446 | 36.5371 | 52.7833 | 25.2913 | 2.2857 | 36.43 | 83.2672 | 15.9384 | 21.6375 | 5.8386 | 136.0505 | 272.7619 | 209.8926 | 267.6109 | 186.0764
2012-10-30 00:09 | 26.0583 | 294.4119 | 231.492 | 289.2571 | 207.7466 | 84.6934 | 1.4446 | 36.5369 | 52.7826 | 25.2913 | 2.2857 | 36.4298 | 83.2674 | 15.9382 | 21.6228 | 5.8388 | 136.0501 | 272.7891 | 209.8692 | 267.6343 | 186.1238
2012-10-30 00:09 | 26.0673 | 294.5094 | 231.556 | 289.3278 | 207.7832 | 84.6876 | 1.4446 | 36.5364 | 52.7812 | 25.2914 | 2.2857 | 36.4294 | 83.2653 | 15.9382 | 21.6478 | 5.8384 | 136.0465 | 272.8616 | 209.9082 | 267.68 | 186.1354
2012-10-30 00:09 | 26.059 | 294.5088 | 231.5831 | 289.3426 | 207.7473 | 84.7067 | 1.4446 | 36.5362 | 52.7807 | 25.2914 | 2.2857 | 36.4291 | 83.2644 | 15.9381 | 21.6154 | 5.8397 | 136.0451 | 272.8934 | 209.9677 | 267.7272 | 186.1319
2012-10-30 00:09 | 26.0658 | 294.531 | 231.6017 | 289.3725 | 207.8683 | 84.7056 | 1.4446 | 36.5352 | 52.7775 | 25.2915 | 2.2856 | 36.4281 | 83.2604 | 15.9381 | 21.6062 | 5.8396 | 136.0379 | 272.9248 | 209.9955 | 267.7663 | 186.2621
2012-10-30 00:09 | 26.0665 | 294.6086 | 231.5434 | 289.431 | 207.802 | 84.7119 | 1.4445 | 36.5346 | 52.7758 | 25.2915 | 2.2856 | 36.4276 | 83.2585 | 15.938 | 21.6149 | 5.84 | 136.0343 | 272.9937 | 209.9285 | 267.8161 | 186.1871
2012-10-30 00:09 | 26.074 | 294.6045 | 231.6374 | 289.4501 | 207.904 | 84.7206 | 1.4445 | 36.5338 | 52.7728 | 25.2919 | 2.2855 | 36.4268 | 83.2543 | 15.938 | 21.6023 | 5.8406 | 136.027 | 273.0022 | 210.0351 | 267.8478 | 186.3017
2012-10-30 00:09 | 26.0747 | 294.6706 | 231.6342 | 289.5009 | 208.0071 | 84.724 | 1.4444 | 36.532 | 52.7672 | 25.2919 | 2.2854 | 36.4248 | 83.2459 | 15.9379 | 21.599 | 5.8409 | 136.0131 | 273.0716 | 210.0352 | 267.9019 | 186.4081
2012-10-30 00:09 | 26.0747 | 294.6975 | 231.6617 | 289.5163 | 207.9717 | 84.7235 | 1.4444 | 36.5308 | 52.7639 | 25.2919 | 2.2854 | 36.4237 | 83.2411 | 15.9379 | 21.603 | 5.8408 | 136.005 | 273.0945 | 210.0587 | 267.9133 | 186.3687
2012-10-30 00:09 | 26.0747 | 294.6937 | 231.6892 | 289.5548 | 207.9717 | 84.724 | 1.4444 | 36.5305 | 52.7633 | 25.2918 | 2.2853 | 36.4235 | 83.2401 | 15.9379 | 21.579 | 5.8409 | 136.0034 | 273.1147 | 210.1102 | 267.9758 | 186.3927
2012-10-30 00:10 | 26.0807 | 294.7458 | 231.707 | 289.5955 | 208.0328 | 84.7275 | 1.4443 | 36.5295 | 52.7606 | 25.2918 | 2.2853 | 36.4223 | 83.235 | 15.9378 | 21.589 | 5.8411 | 135.9956 | 273.1568 | 210.118 | 268.0065 | 186.4438
2012-10-30 00:10 | 26.0867 | 294.7903 | 231.7483 | 289.6015 | 208.1097 | 84.7327 | 1.4443 | 36.5284 | 52.7574 | 25.2917 | 2.2852 | 36.4211 | 83.2303 | 15.9377 | 21.5669 | 5.8415 | 135.9877 | 273.2234 | 210.1814 | 268.0346 | 186.5428
2012-10-30 00:10 | 26.0785 | 294.7782 | 231.7834 | 289.6164 | 208.003 | 84.7495 | 1.4443 | 36.5284 | 52.7567 | 25.292 | 2.2852 | 36.4214 | 83.2318 | 15.9376 | 21.5587 | 5.8426 | 135.9885 | 273.2195 | 210.2247 | 268.0577 | 186.4443
2012-10-30 00:10 | 26.0807 | 294.8612 | 231.7856 | 289.6533 | 207.9462 | 84.7449 | 1.4442 | 36.5271 | 52.7533 | 25.2918 | 2.2852 | 36.4201 | 83.226 | 15.9376 | 21.6251 | 5.8423 | 135.9793 | 273.2361 | 210.1605 | 268.0282 | 186.3211
2012-10-30 00:10 | 26.0927 | 294.8847 | 231.7858 | 289.7115 | 208.0881 | 84.7449 | 1.4442 | 36.5266 | 52.7511 | 25.2922 | 2.2851 | 36.4195 | 83.2232 | 15.9376 | 21.5809 | 5.8423 | 135.9743 | 273.3038 | 210.2049 | 268.1306 | 186.5072
2012-10-30 00:10 | 26.0957 | 294.88 | 231.8124 | 289.7453 | 208.0163 | 84.7535 | 1.4442 | 36.5259 | 52.7496 | 25.292 | 2.2851 | 36.4189 | 83.2208 | 15.9376 | 21.5799 | 5.8429 | 135.9704 | 273.3001 | 210.2325 | 268.1654 | 186.4364
2012-10-30 00:10 | 26.089 | 294.8964 | 231.8332 | 289.7618 | 207.9151 | 84.7449 | 1.4441 | 36.5245 | 52.7456 | 25.292 | 2.285 | 36.4176 | 83.2158 | 15.9374 | 21.5491 | 5.8423 | 135.9614 | 273.3473 | 210.2841 | 268.2127 | 186.366
2012-10-30 00:10 | 26.0972 | 295.0046 | 231.8257 | 289.7584 | 207.8367 | 84.7599 | 1.4441 | 36.5243 | 52.7447 | 25.2921 | 2.285 | 36.4173 | 83.2144 | 15.9373 | 21.5894 | 5.8433 | 135.9591 | 273.4152 | 210.2363 | 268.169 | 186.2473
2012-10-30 00:10 | 26.1017 | 295.0053 | 231.8459 | 289.8207 | 207.7979 | 84.772 | 1.444 | 36.5226 | 52.7399 | 25.2921 | 2.2849 | 36.4157 | 83.2074 | 15.9373 | 21.5658 | 5.8442 | 135.9472 | 273.4395 | 210.2801 | 268.2549 | 186.2321
2012-10-30 00:10 | 26.0987 | 295.0061 | 231.8389 | 289.9024 | 208.2753 | 84.7616 | 1.444 | 36.5218 | 52.7375 | 25.2921 | 2.2849 | 36.415 | 83.204 | 15.9373 | 21.599 | 5.8435 | 135.9415 | 273.4071 | 210.2399 | 268.3034 | 186.6763
2012-10-30 00:10 | 26.1069 | 295.0681 | 231.9375 | 289.9453 | 208.3898 | 84.7847 | 1.4418 | 36.4648 | 52.5743 | 25.2915 | 2.2814 | 36.3582 | 82.9467 | 15.937 | 21.6192 | 5.8451 | 135.521 | 273.4489 | 210.3183 | 268.3261 | 186.7706
2012-10-30 00:10 | 26.1032 | 295.0798 | 231.926 | 289.957 | 208.2916 | 84.7593 | 1.4402 | 36.4224 | 52.4538 | 25.2906 | 2.2787 | 36.3157 | 82.7533 | 15.9368 | 21.6155 | 5.8433 | 135.2071 | 273.4643 | 210.3105 | 268.3415 | 186.6761
2012-10-30 00:10 | 26.1069 | 295.1104 | 231.9257 | 289.9684 | 208.2401 | 84.7605 | 1.4397 | 36.4098 | 52.4178 | 25.2906 | 2.2779 | 36.303 | 82.6953 | 15.9369 | 21.6031 | 5.8434 | 135.1131 | 273.5073 | 210.3226 | 268.3653 | 186.637
2012-10-30 00:10 | 26.1009 | 295.1198 | 231.9355 | 290.0048 | 207.2222 | 84.768 | 1.4393 | 36.4016 | 52.3941 | 25.2906 | 2.2774 | 36.2949 | 82.6589 | 15.9368 | 21.6292 | 5.8439 | 135.053 | 273.4906 | 210.3063 | 268.3756 | 185.593
2012-10-30 00:10 | 26.1092 | 295.1396 | 231.8023 | 289.9977 | 206.6793 | 84.7576 | 1.4392 | 36.3978 | 52.3836 | 25.2903 | 2.2772 | 36.2912 | 82.6422 | 15.9368 | 21.5773 | 5.8432 | 135.0258 | 273.5623 | 210.225 | 268.4204 | 185.102
2012-10-30 00:10 | 26.1054 | 295.1243 | 231.6216 | 289.9939 | 205.7621 | 84.7795 | 1.4393 | 36.3997 | 52.3888 | 25.2905 | 2.2773 | 36.293 | 82.6504 | 15.9368 | 21.5976 | 5.8447 | 135.0392 | 273.5267 | 210.024 | 268.3963 | 184.1645
2012-10-30 00:10 | 26.1084 | 295.1504 | 231.3416 | 290.0547 | 205.4029 | 84.7726 | 1.4393 | 36.4005 | 52.3914 | 25.2903 | 2.2774 | 36.2938 | 82.6546 | 15.9367 | 21.5806 | 5.8442 | 135.0461 | 273.5698 | 209.761 | 268.4741 | 183.8223
2012-10-30 00:10 | 26.1114 | 295.2034 | 231.1874 | 290.0307 | 205.6539 | 84.7622 | 1.4394 | 36.4028 | 52.3978 | 25.2905 | 2.2775 | 36.2961 | 82.6654 | 15.9366 | 21.5996 | 5.8435 | 135.0631 | 273.6038 | 209.5878 | 268.4311 | 184.0543
2012-10-30 00:10 | 26.1002 | 295.1806 | 231.0583 | 290.0503 | 205.7687 | 84.7789 | 1.4396 | 36.4086 | 52.4142 | 25.2905 | 2.2779 | 36.3019 | 82.691 | 15.9368 | 21.6005 | 5.8447 | 135.1052 | 273.5801 | 209.4578 | 268.4498 | 184.1682
2012-10-30 00:10 | 26.1047 | 295.239 | 230.992 | 290.0548 | 206.5645 | 84.7709 | 1.4398 | 36.4137 | 52.4288 | 25.2907 | 2.2782 | 36.3071 | 82.7149 | 15.9367 | 21.5769 | 5.8441 | 135.1437 | 273.6621 | 209.4151 | 268.4779 | 184.9876
2012-10-30 00:10 | 26.0927 | 295.2616 | 230.98 | 290.1006 | 207.1668 | 84.7737 | 1.4411 | 36.4472 | 52.5243 | 25.2911 | 2.2803 | 36.3399 | 82.865 | 15.9366 | 21.6371 | 5.8443 | 135.3893 | 273.6245 | 209.3429 | 268.4635 | 185.5297
2012-10-30 00:10 | 26.0949 | 295.2176 | 231.0333 | 290.1413 | 207.0902 | 84.7888 | 1.4432 | 36.5007 | 52.6769 | 25.292 | 2.2836 | 36.3936 | 83.1084 | 15.9369 | 21.6072 | 5.8453 | 135.7853 | 273.6104 | 209.4261 | 268.5341 | 185.483
2012-10-30 00:10 | 26.083 | 295.2788 | 231.0764 | 290.1603 | 206.8657 | 84.8026 | 1.4436 | 36.5122 | 52.7095 | 25.2922 | 2.2843 | 36.4051 | 83.1612 | 15.9369 | 21.6154 | 5.8463 | 135.8707 | 273.6634 | 209.461 | 268.5449 | 185.2503
2012-10-30 00:10 | 26.0964 | 295.2999 | 231.1134 | 290.1428 | 207.1862 | 84.7841 | 1.4441 | 36.5256 | 52.7474 | 25.2926 | 2.2852 | 36.4187 | 83.2232 | 15.9369 | 21.6247 | 5.845 | 135.9706 | 273.6752 | 209.4887 | 268.5181 | 185.5615
2012-10-30 00:10 | 26.0867 | 295.3094 | 231.1548 | 290.1601 | 207.2946 | 84.798 | 1.4445 | 36.5356 | 52.7757 | 25.2929 | 2.2858 | 36.4286 | 83.2692 | 15.9368 | 21.599 | 5.846 | 136.045 | 273.7104 | 209.5558 | 268.5611 | 185.6956
2012-10-30 00:10 | 26.071 | 295.3514 | 231.1627 | 290.1637 | 207.2907 | 84.8032 | 1.4447 | 36.5405 | 52.7898 | 25.2929 | 2.2861 | 36.4336 | 83.2925 | 15.9367 | 21.5873 | 5.8463 | 136.0823 | 273.7641 | 209.5754 | 268.5764 | 185.7034
2012-10-30 00:11 | 26.0785 | 295.4165 | 231.2685 | 290.202 | 207.1841 | 84.7974 | 1.4449 | 36.5446 | 52.802 | 25.2928 | 2.2864 | 36.4376 | 83.3106 | 15.9367 | 21.6189 | 5.8459 | 136.1126 | 273.7976 | 209.6496 | 268.5831 | 185.5652
2012-10-30 00:11 | 26.0687 | 295.4183 | 231.1997 | 290.1845 | 207.3711 | 84.8119 | 1.445 | 36.5479 | 52.8106 | 25.2932 | 2.2866 | 36.4406 | 83.3236 | 15.9368 | 21.585 | 5.8469 | 136.1343 | 273.8333 | 209.6147 | 268.5995 | 185.7861
2012-10-30 00:11 | 26.0852 | 295.3925 | 231.2673 | 290.2357 | 207.6357 | 84.8252 | 1.445 | 36.5486 | 52.8124 | 25.2933 | 2.2866 | 36.4415 | 83.3281 | 15.9368 | 21.6176 | 5.8478 | 136.1405 | 273.7749 | 209.6497 | 268.6181 | 186.0181
2012-10-30 00:11 | 26.0815 | 295.4772 | 231.3422 | 290.2165 | 207.7107 | 84.8292 | 1.445 | 36.5484 | 52.8114 | 25.2934 | 2.2866 | 36.4415 | 83.3282 | 15.9367 | 21.5737 | 5.8481 | 136.1396 | 273.9035 | 209.7685 | 268.6428 | 186.137
2012-10-30 00:11 | 26.0762 | 295.4565 | 231.4352 | 290.2267 | 207.6857 | 84.8107 | 1.4449 | 36.5476 | 52.8095 | 25.2934 | 2.2866 | 36.4406 | 83.3245 | 15.9367 | 21.5724 | 5.8468 | 136.134 | 273.8841 | 209.8628 | 268.6543 | 186.1133
2012-10-30 00:11 | 26.0807 | 295.4649 | 231.42 | 290.2234 | 207.5564 | 84.831 | 1.445 | 36.5481 | 52.8106 | 25.2935 | 2.2866 | 36.4409 | 83.3255 | 15.9367 | 21.601 | 5.8482 | 136.1361 | 273.8639 | 209.819 | 268.6224 | 185.9554
2012-10-30 00:11 | 26.0658 | 295.4807 | 231.4562 | 290.2509 | 207.7935 | 84.8188 | 1.445 | 36.5489 | 52.8126 | 25.2936 | 2.2867 | 36.4417 | 83.3301 | 15.9366 | 21.546 | 5.8474 | 136.1427 | 273.9347 | 209.9102 | 268.7049 | 186.2475
2012-10-30 00:11 | 26.0717 | 295.4982 | 231.5054 | 290.2722 | 207.9569 | 84.82 | 1.445 | 36.5484 | 52.8111 | 25.2936 | 2.2866 | 36.4411 | 83.3279 | 15.9365 | 21.5679 | 5.8475 | 136.139 | 273.9303 | 209.9375 | 268.7043 | 186.389
2012-10-30 00:11 | 26.0822 | 295.5779 | 231.5788 | 290.2943 | 208.0028 | 84.8356 | 1.445 | 36.5481 | 52.8107 | 25.2934 | 2.2866 | 36.4408 | 83.3272 | 15.9364 | 21.5825 | 5.8486 | 136.138 | 273.9954 | 209.9963 | 268.7118 | 186.4203
2012-10-30 00:11 | 26.0807 | 295.5456 | 231.6323 | 290.3043 | 208.1588 | 84.8385 | 1.4449 | 36.5476 | 52.8092 | 25.2935 | 2.2866 | 36.4404 | 83.3258 | 15.9363 | 21.5408 | 5.8488 | 136.1351 | 274.0048 | 210.0915 | 268.7635 | 186.618
2012-10-30 00:11 | 26.0845 | 295.5225 | 231.7501 | 290.3544 | 208.1547 | 84.8258 | 1.4449 | 36.5458 | 52.8044 | 25.2933 | 2.2865 | 36.4386 | 83.3183 | 15.9362 | 21.5647 | 5.8479 | 136.1227 | 273.9578 | 210.1854 | 268.7897 | 186.59
2012-10-30 00:11 | 26.0762 | 295.5796 | 231.7811 | 290.3692 | 208.1819 | 84.8552 | 1.4448 | 36.544 | 52.798 | 25.2938 | 2.2864 | 36.4369 | 83.3096 | 15.9363 | 21.5604 | 5.8499 | 136.1076 | 274.0192 | 210.2207 | 268.8088 | 186.6215
2012-10-30 00:11 | 26.083 | 295.6017 | 231.8586 | 290.3952 | 208.2399 | 84.8489 | 1.4448 | 36.544 | 52.7978 | 25.2939 | 2.2864 | 36.4369 | 83.3092 | 15.9364 | 21.5873 | 5.8495 | 136.1069 | 274.0144 | 210.2713 | 268.8079 | 186.6526
2012-10-30 00:11 | 26.086 | 295.6393 | 231.9127 | 290.4021 | 208.1681 | 84.8506 | 1.4448 | 36.5443 | 52.7987 | 25.294 | 2.2864 | 36.4373 | 83.3108 | 15.9365 | 21.5782 | 5.8496 | 136.1095 | 274.0611 | 210.3345 | 268.8239 | 186.5899
2012-10-30 00:11 | 26.0807 | 295.5994 | 231.876 | 290.4315 | 208.1509 | 84.857 | 1.4448 | 36.5447 | 52.7996 | 25.294 | 2.2864 | 36.4375 | 83.3124 | 15.9363 | 21.5328 | 5.85 | 136.112 | 274.0666 | 210.3432 | 268.8987 | 186.6181
2012-10-30 00:11 | 26.083 | 295.6287 | 231.9529 | 290.4646 | 208.3659 | 84.8593 | 1.4448 | 36.5438 | 52.797 | 25.2941 | 2.2864 | 36.4367 | 83.3088 | 15.9363 | 21.5792 | 5.8502 | 136.1057 | 274.0495 | 210.3737 | 268.8854 | 186.7867
2012-10-30 00:11 | 26.0845 | 295.6763 | 231.9702 | 290.4314 | 208.3437 | 84.8674 | 1.4448 | 36.5435 | 52.7963 | 25.2939 | 2.2864 | 36.4363 | 83.3071 | 15.9363 | 21.5567 | 5.8508 | 136.1035 | 274.1196 | 210.4135 | 268.8747 | 186.787
2012-10-30 00:11 | 26.0875 | 295.6985 | 232.0203 | 290.4806 | 208.1617 | 84.8755 | 1.4447 | 36.5417 | 52.7912 | 25.294 | 2.2863 | 36.4347 | 83.3008 | 15.9361 | 21.5356 | 5.8513 | 136.092 | 274.1629 | 210.4847 | 268.945 | 186.6261
2012-10-30 00:11 | 26.0845 | 295.7339 | 232.037 | 290.4892 | 208.2729 | 84.8772 | 1.4447 | 36.5415 | 52.7911 | 25.2937 | 2.2863 | 36.4345 | 83.3002 | 15.936 | 21.5767 | 5.8514 | 136.0913 | 274.1572 | 210.4603 | 268.9125 | 186.6962
2012-10-30 00:11 | 26.0897 | 295.7622 | 232.0933 | 290.4867 | 208.4001 | 84.8807 | 1.4446 | 36.5404 | 52.7876 | 25.2938 | 2.2862 | 36.4334 | 83.2953 | 15.936 | 21.5659 | 5.8517 | 136.0829 | 274.1963 | 210.5274 | 268.9208 | 186.8342
2012-10-30 00:11 | 26.0807 | 295.7878 | 232.1157 | 290.5393 | 208.4659 | 84.8812 | 1.4446 | 36.5399 | 52.7859 | 25.294 | 2.2862 | 36.4327 | 83.2926 | 15.9359 | 21.5769 | 5.8517 | 136.0785 | 274.2109 | 210.5388 | 268.9624 | 186.889
2012-10-30 00:11 | 26.0792 | 295.7902 | 232.1378 | 290.5108 | 208.3975 | 84.8714 | 1.4446 | 36.5386 | 52.7824 | 25.2939 | 2.2861 | 36.4315 | 83.2869 | 15.936 | 21.5353 | 5.851 | 136.0694 | 274.2549 | 210.6025 | 268.9755 | 186.8622
2012-10-30 00:11 | 26.086 | 295.8354 | 232.1485 | 290.5446 | 208.3846 | 84.8772 | 1.4445 | 36.5373 | 52.7778 | 25.2942 | 2.286 | 36.4302 | 83.2805 | 15.936 | 21.5461 | 5.8514 | 136.0584 | 274.2893 | 210.6024 | 268.9985 | 186.8385
2012-10-30 00:11 | 26.0957 | 295.8374 | 232.2172 | 290.6044 | 208.4573 | 84.9044 | 1.4445 | 36.5372 | 52.7773 | 25.2943 | 2.286 | 36.43 | 83.2809 | 15.9358 | 21.5799 | 5.8533 | 136.0582 | 274.2575 | 210.6373 | 269.0245 | 186.8774
2012-10-30 00:11 | 26.0875 | 295.8715 | 232.1972 | 290.5923 | 208.4255 | 84.8911 | 1.4444 | 36.5354 | 52.7716 | 25.2946 | 2.2859 | 36.4284 | 83.2727 | 15.9359 | 21.5356 | 5.8524 | 136.0443 | 274.3359 | 210.6616 | 269.0567 | 186.8899
2012-10-30 00:11 | 26.086 | 295.9046 | 232.2507 | 290.5985 | 208.5106 | 84.8951 | 1.4444 | 36.5348 | 52.7699 | 25.2946 | 2.2859 | 36.4277 | 83.2694 | 15.9359 | 21.522 | 5.8527 | 136.0393 | 274.3826 | 210.7287 | 269.0765 | 186.9886
2012-10-30 00:11 | 26.0972 | 295.8774 | 232.313 | 290.5943 | 208.5336 | 84.9044 | 1.4444 | 36.5342 | 52.7682 | 25.2945 | 2.2858 | 36.427 | 83.2666 | 15.9359 | 21.5453 | 5.8533 | 136.0348 | 274.3321 | 210.7677 | 269.049 | 186.9883
2012-10-30 00:12 | 26.0979 | 295.9242 | 232.2822 | 290.6875 | 208.4831 | 84.902 | 1.4443 | 36.533 | 52.7643 | 25.2947 | 2.2858 | 36.426 | 83.262 | 15.9359 | 21.546 | 5.8531 | 136.0263 | 274.3782 | 210.7362 | 269.1415 | 186.9371
2012-10-30 00:12 | 26.0994 | 295.9372 | 232.3112 | 290.689 | 208.591 | 84.9055 | 1.4443 | 36.5324 | 52.7626 | 25.2948 | 2.2858 | 36.4253 | 83.2595 | 15.9358 | 21.5675 | 5.8534 | 136.022 | 274.3697 | 210.7437 | 269.1215 | 187.0235
2012-10-30 00:12 | 26.1002 | 296.0034 | 232.3828 | 290.7206 | 208.58 | 84.9032 | 1.4442 | 36.5312 | 52.7593 | 25.2947 | 2.2857 | 36.4241 | 83.2546 | 15.9356 | 21.5483 | 5.8532 | 136.0139 | 274.4551 | 210.8345 | 269.1723 | 187.0317
2012-10-30 00:12 | 26.1032 | 295.9564 | 232.3819 | 290.7351 | 208.6026 | 84.9055 | 1.4442 | 36.5301 | 52.7558 | 25.2948 | 2.2856 | 36.4231 | 83.2503 | 15.9356 | 21.6155 | 5.8534 | 136.0061 | 274.3409 | 210.7664 | 269.1196 | 186.9871
2012-10-30 00:12 | 26.0942 | 295.932 | 232.4397 | 290.7377 | 208.6763 | 84.9148 | 1.4441 | 36.5283 | 52.7512 | 25.2945 | 2.2855 | 36.4212 | 83.2421 | 15.9355 | 21.5543 | 5.854 | 135.9933 | 274.3777 | 210.8854 | 269.1834 | 187.122
2012-10-30 00:12 | 26.0957 | 296.0065 | 232.4883 | 290.743 | 208.67 | 84.92 | 1.4441 | 36.5276 | 52.7494 | 25.2944 | 2.2855 | 36.4204 | 83.2385 | 15.9354 | 21.5478 | 5.8544 | 135.9879 | 274.4587 | 210.9405 | 269.1952 | 187.1222
2012-10-30 00:12 | 26.1092 | 296.047 | 232.5411 | 290.7527 | 208.5968 | 84.9217 | 1.4441 | 36.5274 | 52.7481 | 25.2948 | 2.2855 | 36.4203 | 83.2377 | 15.9355 | 21.6295 | 5.8545 | 135.9858 | 274.4175 | 210.9116 | 269.1232 | 186.9673
2012-10-30 00:12 | 26.1054 | 296.0547 | 232.5059 | 290.8105 | 208.6285 | 84.92 | 1.444 | 36.5264 | 52.7454 | 25.2947 | 2.2854 | 36.4193 | 83.2329 | 15.9356 | 21.6137 | 5.8544 | 135.9784 | 274.441 | 210.8922 | 269.1968 | 187.0148
2012-10-30 00:12 | 26.0994 | 296.0487 | 232.5117 | 290.8392 | 208.713 | 84.9327 | 1.4439 | 36.5242 | 52.7391 | 25.2947 | 2.2853 | 36.4173 | 83.2244 | 15.9354 | 21.5876 | 5.8553 | 135.9635 | 274.4611 | 210.9241 | 269.2516 | 187.1254
2012-10-30 00:12 | 26.1077 | 296.0801 | 232.5789 | 290.8513 | 208.7371 | 84.9286 | 1.4439 | 36.5236 | 52.7369 | 25.2949 | 2.2853 | 36.4166 | 83.2217 | 15.9354 | 21.5678 | 5.855 | 135.9585 | 274.5123 | 211.0111 | 269.2835 | 187.1693
2012-10-30 00:12 | 26.1122 | 296.1153 | 232.5991 | 290.8288 | 208.8125 | 84.935 | 1.4439 | 36.5226 | 52.7334 | 25.2952 | 2.2852 | 36.4154 | 83.2162 | 15.9354 | 21.5723 | 5.8554 | 135.9496 | 274.543 | 211.0268 | 269.2565 | 187.2402
2012-10-30 00:12 | 26.1166 | 296.1274 | 232.635 | 290.8756 | 208.9232 | 84.9246 | 1.4438 | 36.5213 | 52.73 | 25.2951 | 2.2851 | 36.4143 | 83.2107 | 15.9355 | 21.5807 | 5.8547 | 135.9407 | 274.5467 | 211.0543 | 269.2949 | 187.3425
2012-10-30 00:12 | 26.1122 | 296.1038 | 232.6581 | 290.9136 | 208.7652 | 84.9286 | 1.4431 | 36.5036 | 52.6787 | 25.2951 | 2.284 | 36.3958 | 83.127 | 15.9353 | 21.5643 | 5.855 | 135.8057 | 274.5395 | 211.0938 | 269.3493 | 187.2009
2012-10-30 00:12 | 26.1159 | 296.1613 | 232.6461 | 290.9134 | 208.7965 | 84.924 | 1.4399 | 36.421 | 52.4431 | 25.2938 | 2.2789 | 36.3148 | 82.7592 | 15.9349 | 21.5399 | 5.8547 | 135.2023 | 274.6214 | 211.1062 | 269.3735 | 187.2566
2012-10-30 00:12 | 26.1204 | 296.1773 | 232.6388 | 290.9564 | 208.675 | 84.9333 | 1.4397 | 36.4156 | 52.4282 | 25.2937 | 2.2786 | 36.3092 | 82.7335 | 15.935 | 21.5364 | 5.8553 | 135.1616 | 274.6409 | 211.1024 | 269.42 | 187.1386
2012-10-30 00:12 | 26.1211 | 296.1473 | 232.6513 | 290.9571 | 208.6875 | 84.9477 | 1.4393 | 36.405 | 52.398 | 25.2934 | 2.2779 | 36.2983 | 82.6843 | 15.9349 | 21.5531 | 5.8563 | 135.0822 | 274.5942 | 211.0982 | 269.404 | 187.1344
2012-10-30 00:12 | 26.1174 | 296.1974 | 232.6083 | 290.9804 | 208.3727 | 84.9367 | 1.4392 | 36.4014 | 52.3874 | 25.2935 | 2.2777 | 36.2945 | 82.6671 | 15.9349 | 21.5574 | 5.8555 | 135.0544 | 274.64 | 211.0509 | 269.423 | 186.8153
2012-10-30 00:12 | 26.1166 | 296.1735 | 232.5878 | 291.0027 | 207.8915 | 84.9205 | 1.4391 | 36.3991 | 52.3812 | 25.2933 | 2.2776 | 36.2924 | 82.658 | 15.9348 | 21.5245 | 5.8544 | 135.0391 | 274.649 | 211.0633 | 269.4782 | 186.367
2012-10-30 00:12 | 26.1129 | 296.1775 | 232.5094 | 290.999 | 208.0808 | 84.9182 | 1.4391 | 36.4008 | 52.386 | 25.2933 | 2.2777 | 36.294 | 82.6656 | 15.9348 | 21.5168 | 5.8543 | 135.0516 | 274.6607 | 210.9926 | 269.4822 | 186.564
2012-10-30 00:12 | 26.1024 | 296.2093 | 232.4636 | 290.9731 | 208.4207 | 84.9309 | 1.4393 | 36.4044 | 52.3963 | 25.2934 | 2.2779 | 36.2978 | 82.6831 | 15.9347 | 21.5144 | 5.8551 | 135.0794 | 274.6949 | 210.9492 | 269.4587 | 186.9063
2012-10-30 00:12 | 26.1129 | 296.2083 | 232.4898 | 291.026 | 208.8959 | 84.9315 | 1.4395 | 36.4094 | 52.4104 | 25.2935 | 2.2782 | 36.3024 | 82.7037 | 15.9348 | 21.5249 | 5.8552 | 135.1142 | 274.6834 | 210.9649 | 269.5011 | 187.371
2012-10-30 00:12 | 26.1092 | 296.2238 | 232.5883 | 290.9876 | 208.9945 | 84.939 | 1.4396 | 36.4139 | 52.4232 | 25.2937 | 2.2785 | 36.307 | 82.7249 | 15.9348 | 21.5452 | 5.8557 | 135.1481 | 274.6786 | 211.0431 | 269.4424 | 187.4493
2012-10-30 00:12 | 26.1032 | 296.2255 | 232.5862 | 291.024 | 208.9688 | 84.9269 | 1.4399 | 36.4192 | 52.4382 | 25.2937 | 2.2788 | 36.3123 | 82.7491 | 15.9347 | 21.5633 | 5.8549 | 135.1873 | 274.6622 | 211.0229 | 269.4607 | 187.4055
2012-10-30 00:12 | 26.0987 | 296.2286 | 232.621 | 291.0503 | 208.9289 | 84.9275 | 1.4406 | 36.4381 | 52.4926 | 25.2937 | 2.28 | 36.3312 | 82.8358 | 15.9346 | 21.5428 | 5.8549 | 135.3284 | 274.6858 | 211.0782 | 269.5075 | 187.3861
2012-10-30 00:12 | 26.0927 | 296.2803 | 232.6739 | 291.0597 | 208.4386 | 84.9454 | 1.4432 | 36.5044 | 52.6815 | 25.2949 | 2.2841 | 36.3973 | 83.1362 | 15.9349 | 21.5649 | 5.8561 | 135.8177 | 274.7154 | 211.109 | 269.4948 | 186.8737
2012-10-30 00:13 | 26.089 | 296.2843 | 232.6113 | 291.0444 | 208.49 | 84.9552 | 1.4441 | 36.5296 | 52.7529 | 25.2955 | 2.2857 | 36.4223 | 83.2504 | 15.9348 | 21.5692 | 5.8568 | 136.0033 | 274.7151 | 211.0421 | 269.4752 | 186.9208
2012-10-30 00:13 | 26.0837 | 296.302 | 232.6217 | 291.0776 | 208.7682 | 84.9668 | 1.4448 | 36.5454 | 52.799 | 25.2953 | 2.2867 | 36.4384 | 83.3248 | 15.9347 | 21.604 | 5.8576 | 136.1238 | 274.698 | 211.0177 | 269.4736 | 187.1642
2012-10-30 00:13 | 26.0807 | 296.2914 | 232.6305 | 291.0746 | 208.7967 | 84.9598 | 1.4448 | 36.5474 | 52.8044 | 25.2955 | 2.2868 | 36.4401 | 83.332 | 15.9348 | 21.5729 | 5.8571 | 136.1364 | 274.7185 | 211.0576 | 269.5017 | 187.2238
2012-10-30 00:13 | 26.0807 | 296.3144 | 232.6619 | 291.0708 | 208.7298 | 84.961 | 1.4449 | 36.5495 | 52.8104 | 25.2956 | 2.287 | 36.4422 | 83.3426 | 15.9347 | 21.6131 | 5.8572 | 136.1529 | 274.7013 | 211.0488 | 269.4577 | 187.1167
2012-10-30 00:13 | 26.0867 | 296.2781 | 232.6562 | 291.1037 | 208.909 | 84.9662 | 1.4449 | 36.5502 | 52.8117 | 25.2959 | 2.287 | 36.443 | 83.3464 | 15.9346 | 21.5749 | 5.8576 | 136.1581 | 274.7032 | 211.0813 | 269.5288 | 187.3341
2012-10-30 00:13 | 26.083 | 296.3513 | 232.6721 | 291.0962 | 209.055 | 85.0119 | 1.4449 | 36.5499 | 52.8113 | 25.2957 | 2.287 | 36.4426 | 83.3448 | 15.9346 | 21.5471 | 5.8607 | 136.156 | 274.8042 | 211.125 | 269.5491 | 187.5079
2012-10-30 00:13 | 26.0867 | 296.332 | 232.6954 | 291.1268 | 209.0665 | 84.9708 | 1.4449 | 36.5503 | 52.8119 | 25.2959 | 2.287 | 36.4429 | 83.3465 | 15.9345 | 21.5508 | 5.8579 | 136.1584 | 274.7812 | 211.1446 | 269.576 | 187.5157
2012-10-30 00:13 | 26.0762 | 296.3138 | 232.7675 | 291.1125 | 209.0797 | 84.9783 | 1.4449 | 36.5499 | 52.8109 | 25.2958 | 2.287 | 36.4427 | 83.345 | 15.9346 | 21.5283 | 5.8584 | 136.1559 | 274.7855 | 211.2392 | 269.5842 | 187.5514
2012-10-30 00:13 | 26.0837 | 296.3328 | 232.7514 | 291.12 | 208.9533 | 84.9951 | 1.4449 | 36.5505 | 52.8122 | 25.296 | 2.2871 | 36.4431 | 83.3478 | 15.9345 | 21.5518 | 5.8596 | 136.16 | 274.781 | 211.1996 | 269.5682 | 187.4015
2012-10-30 00:13 | 26.083 | 296.4013 | 232.79 | 291.1193 | 208.9723 | 84.9818 | 1.4449 | 36.5496 | 52.8092 | 25.2962 | 2.287 | 36.4422 | 83.3442 | 15.9343 | 21.5431 | 5.8586 | 136.1533 | 274.8582 | 211.2469 | 269.5762 | 187.4292
2012-10-30 00:13 | 26.0815 | 296.3575 | 232.8042 | 291.1409 | 208.9038 | 84.976 | 1.4449 | 36.5491 | 52.808 | 25.2961 | 2.287 | 36.4418 | 83.3421 | 15.9344 | 21.5537 | 5.8582 | 136.1501 | 274.8038 | 211.2505 | 269.5872 | 187.3501
2012-10-30 00:13 | 26.0867 | 296.4011 | 232.7937 | 291.123 | 208.9563 | 84.9882 | 1.4448 | 36.548 | 52.805 | 25.296 | 2.2869 | 36.4406 | 83.337 | 15.9343 | 21.4906 | 5.8591 | 136.142 | 274.9105 | 211.3031 | 269.6324 | 187.4657
2012-10-30 00:13 | 26.0882 | 296.4026 | 232.8305 | 291.1707 | 209.0365 | 84.9899 | 1.4448 | 36.5484 | 52.8064 | 25.296 | 2.287 | 36.4413 | 83.3403 | 15.9343 | 21.5162 | 5.8592 | 136.1466 | 274.8864 | 211.3143 | 269.6545 | 187.5203
2012-10-30 00:13 | 26.0897 | 296.4349 | 232.777 | 291.1606 | 208.8451 | 84.9957 | 1.4448 | 36.5487 | 52.8066 | 25.2963 | 2.287 | 36.4414 | 83.3415 | 15.9341 | 21.5177 | 5.8596 | 136.1481 | 274.9172 | 211.2593 | 269.6429 | 187.3274
2012-10-30 00:13 | 26.0815 | 296.3959 | 232.816 | 291.2025 | 208.7542 | 85.0032 | 1.4448 | 36.5483 | 52.8053 | 25.2962 | 2.287 | 36.441 | 83.3396 | 15.9341 | 21.5376 | 5.8601 | 136.1449 | 274.8583 | 211.2784 | 269.6649 | 187.2166
2012-10-30 00:13 | 26.0815 | 296.4497 | 232.7728 | 291.2217 | 208.829 | 84.9916 | 1.4448 | 36.5472 | 52.8027 | 25.2961 | 2.2869 | 36.4399 | 83.3352 | 15.934 | 21.5496 | 5.8593 | 136.1378 | 274.9001 | 211.2232 | 269.6721 | 187.2794
2012-10-30 00:13 | 26.0882 | 296.491 | 232.8109 | 291.1938 | 208.7885 | 84.9957 | 1.4447 | 36.5456 | 52.7975 | 25.2963 | 2.2868 | 36.4385 | 83.3284 | 15.9341 | 21.5323 | 5.8596 | 136.1258 | 274.9587 | 211.2786 | 269.6615 | 187.2562
2012-10-30 00:13 | 26.0912 | 296.4863 | 232.8375 | 291.2083 | 208.6812 | 84.9899 | 1.4447 | 36.5457 | 52.7979 | 25.2962 | 2.2868 | 36.4384 | 83.3289 | 15.9339 | 21.5473 | 5.8592 | 136.1268 | 274.939 | 211.2902 | 269.661 | 187.1339
2012-10-30 00:13 | 26.0957 | 296.4678 | 232.7988 | 291.2321 | 208.5676 | 84.9951 | 1.4447 | 36.5447 | 52.7949 | 25.2963 | 2.2868 | 36.4373 | 83.3242 | 15.9339 | 21.5478 | 5.8596 | 136.1191 | 274.92 | 211.251 | 269.6843 | 187.0198
2012-10-30 00:13 | 26.0964 | 296.4915 | 232.7798 | 291.2675 | 208.8518 | 84.9939 | 1.4446 | 36.5436 | 52.7923 | 25.2961 | 2.2867 | 36.4363 | 83.3195 | 15.9339 | 21.5525 | 5.8595 | 136.1117 | 274.939 | 211.2273 | 269.715 | 187.2993
2012-10-30 00:13 | 26.1002 | 296.5146 | 232.8425 | 291.2559 | 208.8556 | 85.0043 | 1.4446 | 36.543 | 52.7896 | 25.2964 | 2.2867 | 36.4357 | 83.3172 | 15.9338 | 21.5724 | 5.8602 | 136.1068 | 274.9422 | 211.2701 | 269.6835 | 187.2832
2012-10-30 00:13 | 26.0912 | 296.471 | 232.8139 | 291.2469 | 208.7875 | 84.998 | 1.4446 | 36.5422 | 52.7881 | 25.296 | 2.2866 | 36.4349 | 83.3136 | 15.9338 | 21.5473 | 5.8598 | 136.1017 | 274.9237 | 211.2666 | 269.6996 | 187.2402
2012-10-30 00:13 | 26.0905 | 296.4895 | 232.7739 | 291.2962 | 208.8459 | 85.0101 | 1.4445 | 36.5412 | 52.7849 | 25.2963 | 2.2866 | 36.4339 | 83.3098 | 15.9337 | 21.5426 | 5.8606 | 136.0948 | 274.9469 | 211.2313 | 269.7536 | 187.3033
2012-10-30 00:13 | 26.0927 | 296.5032 | 232.8075 | 291.2561 | 209.0922 | 84.9887 | 1.4445 | 36.5404 | 52.7826 | 25.2962 | 2.2865 | 36.4332 | 83.3062 | 15.9338 | 21.5689 | 5.8591 | 136.0888 | 274.9343 | 211.2386 | 269.6872 | 187.5233
2012-10-30 00:13 | 26.1084 | 296.492 | 232.8822 | 291.3757 | 208.9662 | 85.0055 | 1.4445 | 36.5399 | 52.7812 | 25.2962 | 2.2865 | 36.4325 | 83.304 | 15.9336 | 21.5765 | 5.8603 | 136.0852 | 274.9155 | 211.3057 | 269.7992 | 187.3897
2012-10-30 00:14 | 26.086 | 296.5119 | 232.8559 | 291.3225 | 208.8453 | 85.0067 | 1.4445 | 36.5388 | 52.7786 | 25.2959 | 2.2865 | 36.4316 | 83.3003 | 15.9335 | 21.5461 | 5.8604 | 136.0789 | 274.9658 | 211.3098 | 269.7764 | 187.2992
2012-10-30 00:14 | 26.1002 | 296.5722 | 232.8818 | 291.3521 | 208.8674 | 85.0003 | 1.4444 | 36.5382 | 52.7767 | 25.296 | 2.2865 | 36.4309 | 83.2976 | 15.9334 | 21.5523 | 5.8599 | 136.0743 | 275.0199 | 211.3295 | 269.7998 | 187.3151
2012-10-30 00:14 | 26.0994 | 296.556 | 232.885 | 291.3937 | 208.9572 | 85.0211 | 1.4444 | 36.5366 | 52.7721 | 25.296 | 2.2864 | 36.4296 | 83.2913 | 15.9335 | 21.5756 | 5.8614 | 136.0634 | 274.9804 | 211.3094 | 269.8181 | 187.3816
2012-10-30 00:14 | 26.1032 | 296.5675 | 232.9202 | 291.4283 | 208.9255 | 85.0142 | 1.4443 | 36.5362 | 52.7706 | 25.2962 | 2.2863 | 36.4289 | 83.2882 | 15.9334 | 21.5633 | 5.8609 | 136.0588 | 275.0042 | 211.3569 | 269.865 | 187.3622
2012-10-30 00:14 | 26.1084 | 296.5919 | 232.9097 | 291.4142 | 208.8953 | 85.0165 | 1.4443 | 36.5346 | 52.7656 | 25.2963 | 2.2862 | 36.4273 | 83.2817 | 15.9332 | 21.6087 | 5.861 | 136.0473 | 274.9832 | 211.301 | 269.8055 | 187.2866
2012-10-30 00:14 | 26.1032 | 296.5867 | 232.8731 | 291.3936 | 208.8271 | 85.0257 | 1.4442 | 36.534 | 52.7636 | 25.2964 | 2.2862 | 36.4268 | 83.2792 | 15.9333 | 21.6075 | 5.8617 | 136.0428 | 274.9792 | 211.2656 | 269.7861 | 187.2196
2012-10-30 00:14 | 26.1092 | 296.562 | 232.8987 | 291.4112 | 208.9355 | 85.0009 | 1.4442 | 36.5325 | 52.7601 | 25.2961 | 2.2861 | 36.4252 | 83.2723 | 15.9332 | 21.6095 | 5.86 | 136.0325 | 274.9525 | 211.2892 | 269.8017 | 187.326
2012-10-30 00:14 | 26.1084 | 296.5804 | 232.8979 | 291.3834 | 208.9937 | 85.0176 | 1.4441 | 36.531 | 52.7551 | 25.2963 | 2.286 | 36.4238 | 83.266 | 15.9331 | 21.6247 | 5.8611 | 136.0212 | 274.9557 | 211.2732 | 269.7587 | 187.369
2012-10-30 00:14 | 26.1114 | 296.5988 | 232.952 | 291.4172 | 209.0676 | 85.0194 | 1.4441 | 36.5305 | 52.7535 | 25.2964 | 2.286 | 36.4232 | 83.2643 | 15.933 | 21.5916 | 5.8612 | 136.0178 | 275.0072 | 211.3604 | 269.8256 | 187.476
2012-10-30 00:14 | 26.1054 | 296.5889 | 232.946 | 291.4266 | 209.2113 | 85.0159 | 1.4441 | 36.5293 | 52.7505 | 25.2963 | 2.286 | 36.4223 | 83.2596 | 15.9331 | 21.6337 | 5.861 | 136.0101 | 274.9552 | 211.3123 | 269.7929 | 187.5776
2012-10-30 00:14 | 26.1054 | 296.612 | 233.0049 | 291.4536 | 209.1443 | 85.028 | 1.444 | 36.5283 | 52.7473 | 25.2963 | 2.2859 | 36.421 | 83.2544 | 15.933 | 21.6578 | 5.8618 | 136.0017 | 274.9542 | 211.3471 | 269.7958 | 187.4865
2012-10-30 00:14 | 26.1024 | 296.6167 | 233.0294 | 291.4429 | 209.0744 | 85.0153 | 1.4439 | 36.5269 | 52.7429 | 25.2965 | 2.2858 | 36.4197 | 83.2482 | 15.9331 | 21.6227 | 5.861 | 135.991 | 274.994 | 211.4067 | 269.8202 | 187.4517
2012-10-30 00:14 | 26.0987 | 296.6937 | 233.065 | 291.4546 | 209.1455 | 85.0263 | 1.444 | 36.5267 | 52.7431 | 25.2963 | 2.2858 | 36.4194 | 83.248 | 15.9328 | 21.623 | 5.8617 | 135.9911 | 275.0707 | 211.442 | 269.8316 | 187.5225
2012-10-30 00:14 | 26.0979 | 296.6506 | 233.1074 | 291.4615 | 209.2077 | 85.0321 | 1.4439 | 36.5253 | 52.7385 | 25.2964 | 2.2857 | 36.4181 | 83.2422 | 15.9328 | 21.6222 | 5.8621 | 135.9807 | 275.0284 | 211.4852 | 269.8393 | 187.5855
2012-10-30 00:14 | 26.0972 | 296.6499 | 233.0753 | 291.5185 | 209.2109 | 85.0269 | 1.4438 | 36.5237 | 52.7341 | 25.2964 | 2.2856 | 36.4165 | 83.2348 | 15.9328 | 21.6095 | 5.8618 | 135.9689 | 275.0404 | 211.4658 | 269.909 | 187.6014
2012-10-30 00:14 | 26.1017 | 296.6544 | 233.0995 | 291.5115 | 209.054 | 85.0332 | 1.4438 | 36.5233 | 52.7333 | 25.2962 | 2.2856 | 36.416 | 83.2321 | 15.9329 | 21.6702 | 5.8622 | 135.9654 | 274.9842 | 211.4293 | 269.8413 | 187.3838
2012-10-30 00:14 | 26.1009 | 296.6613 | 233.0947 | 291.5107 | 209.0374 | 85.035 | 1.4419 | 36.4729 | 52.589 | 25.2957 | 2.2825 | 36.3661 | 83.005 | 15.9327 | 21.6292 | 5.8623 | 135.594 | 275.0321 | 211.4655 | 269.8815 | 187.4082
2012-10-30 00:14 | 26.0837 | 296.6787 | 233.1443 | 291.5089 | 209.3195 | 85.0292 | 1.44 | 36.4253 | 52.4532 | 25.2951 | 2.2795 | 36.3183 | 82.7874 | 15.9326 | 21.5799 | 5.8619 | 135.2406 | 275.0988 | 211.5644 | 269.929 | 187.7396
2012-10-30 00:14 | 26.086 | 296.6887 | 233.1427 | 291.5381 | 209.18 | 85.0298 | 1.4396 | 36.4137 | 52.4194 | 25.2951 | 2.2788 | 36.3069 | 82.736 | 15.9325 | 21.6103 | 5.862 | 135.1554 | 275.0784 | 211.5324 | 269.9278 | 187.5697
2012-10-30 00:14 | 26.0875 | 296.6633 | 233.1403 | 291.5435 | 208.9295 | 85.0194 | 1.4393 | 36.4056 | 52.3978 | 25.2944 | 2.2783 | 36.299 | 82.7001 | 15.9324 | 21.664 | 5.8612 | 135.0979 | 274.9993 | 211.4763 | 269.8795 | 187.2655
2012-10-30 00:14 | 26.0777 | 296.6804 | 233.099 | 291.5221 | 208.7583 | 85.0223 | 1.4391 | 36.4013 | 52.3848 | 25.2947 | 2.278 | 36.2946 | 82.6809 | 15.9323 | 21.6221 | 5.8614 | 135.0657 | 275.0583 | 211.4769 | 269.9 | 187.1362
2012-10-30 00:14 | 26.0837 | 296.6979 | 233.0854 | 291.5551 | 208.2169 | 85.0355 | 1.4392 | 36.4029 | 52.3896 | 25.2945 | 2.2782 | 36.2959 | 82.6878 | 15.9322 | 21.6722 | 5.8623 | 135.0774 | 275.0257 | 211.4132 | 269.8829 | 186.5447
2012-10-30 00:14 | 26.074 | 296.6882 | 232.9971 | 291.5377 | 208.5026 | 85.0223 | 1.4393 | 36.4061 | 52.3987 | 25.2946 | 2.2784 | 36.2993 | 82.7026 | 15.9323 | 21.6264 | 5.8614 | 135.1013 | 275.0618 | 211.3707 | 269.9113 | 186.8762
2012-10-30 00:14 | 26.0732 | 296.6759 | 232.9216 | 291.5099 | 208.6553 | 85.0396 | 1.4393 | 36.4072 | 52.4024 | 25.2943 | 2.2784 | 36.3002 | 82.7069 | 15.9322 | 21.6617 | 5.8626 | 135.1094 | 275.0142 | 211.2599 | 269.8482 | 186.9936
2012-10-30 00:14 | 26.0673 | 296.6584 | 232.904 | 291.5156 | 208.7794 | 85.0217 | 1.4394 | 36.4092 | 52.4084 | 25.2942 | 2.2785 | 36.3022 | 82.7159 | 15.9322 | 21.6157 | 5.8614 | 135.1243 | 275.0427 | 211.2883 | 269.8999 | 187.1637
2012-10-30 00:14 | 26.0702 | 296.7036 | 232.8636 | 291.5262 | 208.55 | 85.0315 | 1.4395 | 36.4131 | 52.4185 | 25.2948 | 2.2788 | 36.3063 | 82.7347 | 15.9322 | 21.6226 | 5.8621 | 135.1532 | 275.081 | 211.241 | 269.9036 | 186.9274
2012-10-30 00:14 | 26.0643 | 296.7246 | 232.9324 | 291.5626 | 208.5165 | 85.0263 | 1.4398 | 36.4185 | 52.434 | 25.2948 | 2.2791 | 36.3116 | 82.7588 | 15.9322 | 21.6368 | 5.8617 | 135.1929 | 275.0878 | 211.2956 | 269.9258 | 186.8797
2012-10-30 00:15 | 26.062 | 296.6455 | 232.8947 | 291.568 | 208.9867 | 85.0171 | 1.4422 | 36.4825 | 52.6167 | 25.2956 | 2.2831 | 36.3745 | 83.0451 | 15.9324 | 21.6586 | 5.8611 | 135.6617 | 274.9869 | 211.2361 | 269.9094 | 187.3281
2012-10-30 00:15 | 26.0583 | 296.7109 | 232.946 | 291.5335 | 209.0027 | 85.0558 | 1.4433 | 36.5097 | 52.6941 | 25.2961 | 2.2848 | 36.402 | 83.1699 | 15.9326 | 21.6348 | 5.8637 | 135.864 | 275.0761 | 211.3112 | 269.8987 | 187.3679
2012-10-30 00:15 | 26.053 | 296.7018 | 232.9368 | 291.5282 | 209.1155 | 85.0407 | 1.4438 | 36.5217 | 52.7283 | 25.2964 | 2.2855 | 36.4141 | 83.2253 | 15.9325 | 21.6295 | 5.8627 | 135.9536 | 275.0723 | 211.3073 | 269.8987 | 187.486
2012-10-30 00:15 | 26.0553 | 296.6926 | 232.9705 | 291.5382 | 208.9879 | 85.0419 | 1.4442 | 36.5321 | 52.7583 | 25.2963 | 2.2862 | 36.4248 | 83.2752 | 15.9323 | 21.6198 | 5.8628 | 136.0335 | 275.0728 | 211.3507 | 269.9184 | 187.3681
2012-10-30 00:15 | 26.0545 | 296.7302 | 232.9855 | 291.5682 | 208.865 | 85.0488 | 1.4444 | 36.5398 | 52.7795 | 25.2969 | 2.2867 | 36.4326 | 83.3104 | 15.9324 | 21.619 | 5.8633 | 136.0899 | 275.1112 | 211.3665 | 269.9492 | 187.246
2012-10-30 00:15 | 26.0493 | 296.7557 | 232.9803 | 291.5592 | 208.8598 | 85.0512 | 1.4446 | 36.5449 | 52.7944 | 25.2968 | 2.287 | 36.4374 | 83.3321 | 15.9325 | 21.6017 | 5.8634 | 136.1265 | 275.154 | 211.3786 | 269.9575 | 187.2581
2012-10-30 00:15 | 26.0426 | 296.7337 | 232.9775 | 291.5255 | 209.1918 | 85.0512 | 1.4448 | 36.548 | 52.8035 | 25.2968 | 2.2872 | 36.4406 | 83.3469 | 15.9324 | 21.6512 | 5.8634 | 136.1505 | 275.0825 | 211.3263 | 269.8743 | 187.5406
2012-10-30 00:15 | 26.0403 | 296.7928 | 233.0184 | 291.5117 | 209.4415 | 85.0616 | 1.4448 | 36.5499 | 52.8082 | 25.2971 | 2.2873 | 36.4425 | 83.3561 | 15.9323 | 21.6007 | 5.8641 | 136.1643 | 275.1921 | 211.4177 | 269.911 | 187.8408
2012-10-30 00:15 | 26.0411 | 296.7706 | 232.9996 | 291.5356 | 209.2139 | 85.061 | 1.4449 | 36.5508 | 52.8108 | 25.2971 | 2.2874 | 36.4434 | 83.3606 | 15.9323 | 21.6176 | 5.8641 | 136.1714 | 275.153 | 211.382 | 269.918 | 187.5963
2012-10-30 00:15 | 26.0328 | 296.7777 | 233.0581 | 291.5619 | 209.2686 | 85.0598 | 1.4449 | 36.5519 | 52.8145 | 25.2969 | 2.2875 | 36.4445 | 83.3653 | 15.9323 | 21.5892 | 5.864 | 136.1798 | 275.1885 | 211.4689 | 269.9727 | 187.6794
2012-10-30 00:15 | 26.0351 | 296.78 | 233.08 | 291.5488 | 209.3693 | 85.0621 | 1.4449 | 36.5529 | 52.8171 | 25.297 | 2.2876 | 36.4455 | 83.3711 | 15.9321 | 21.5795 | 5.8642 | 136.1882 | 275.2005 | 211.5005 | 269.9693 | 187.7898
2012-10-30 00:15 | 26.0321 | 296.7923 | 233.1202 | 291.5189 | 209.4372 | 85.0506 | 1.4449 | 36.552 | 52.8144 | 25.2971 | 2.2875 | 36.4448 | 83.3674 | 15.9322 | 21.6206 | 5.8634 | 136.1818 | 275.1717 | 211.4996 | 269.8983 | 187.8166
2012-10-30 00:15 | 26.0306 | 296.7755 | 233.1462 | 291.5251 | 209.4042 | 85.0679 | 1.4448 | 36.5504 | 52.8093 | 25.2974 | 2.2874 | 36.4433 | 83.3607 | 15.9321 | 21.6352 | 5.8646 | 136.17 | 275.1403 | 211.511 | 269.8899 | 187.769
2012-10-30 00:15 | 26.0261 | 296.771 | 233.1535 | 291.5706 | 209.4982 | 85.0725 | 1.4448 | 36.5495 | 52.8067 | 25.2973 | 2.2873 | 36.4422 | 83.3561 | 15.9321 | 21.6307 | 5.8649 | 136.1628 | 275.1403 | 211.5228 | 269.9399 | 187.8675
2012-10-30 00:15 | 26.0306 | 296.8139 | 233.1737 | 291.5674 | 209.1365 | 85.0743 | 1.4448 | 36.5505 | 52.8098 | 25.2972 | 2.2874 | 36.4432 | 83.3608 | 15.9321 | 21.6071 | 5.865 | 136.1706 | 275.2068 | 211.5666 | 269.9603 | 187.5294
2012-10-30 00:15 | 26.0238 | 296.7917 | 233.1591 | 291.5683 | 209.043 | 85.0668 | 1.4448 | 36.5499 | 52.8079 | 25.2972 | 2.2874 | 36.4426 | 83.3578 | 15.932 | 21.5923 | 5.8645 | 136.1657 | 275.1994 | 211.5668 | 269.976 | 187.4507
2012-10-30 00:15 | 26.0268 | 296.7832 | 233.166 | 291.5752 | 209.3925 | 85.072 | 1.4448 | 36.5495 | 52.8072 | 25.297 | 2.2874 | 36.4423 | 83.3566 | 15.932 | 21.5913 | 5.8649 | 136.1638 | 275.1919 | 211.5747 | 269.9839 | 187.8012
2012-10-30 00:15 | 26.0223 | 296.7864 | 233.1615 | 291.5938 | 209.4353 | 85.0639 | 1.4448 | 36.5486 | 52.8045 | 25.2971 | 2.2873 | 36.4415 | 83.3529 | 15.932 | 21.5868 | 5.8643 | 136.1573 | 275.1996 | 211.5747 | 270.007 | 187.8485
2012-10-30 00:15 | 26.0261 | 296.7902 | 233.2164 | 291.5668 | 209.4352 | 85.0829 | 1.4447 | 36.5465 | 52.7981 | 25.2973 | 2.2872 | 36.4392 | 83.3433 | 15.9319 | 21.6267 | 5.8656 | 136.1413 | 275.1635 | 211.5897 | 269.9401 | 187.8085
2012-10-30 00:15 | 26.0246 | 296.8079 | 233.2149 | 291.6038 | 209.4376 | 85.0795 | 1.4447 | 36.5469 | 52.7993 | 25.2972 | 2.2872 | 36.4395 | 83.3451 | 15.9318 | 21.577 | 5.8654 | 136.1444 | 275.2309 | 211.6379 | 270.0268 | 187.8606
2012-10-30 00:15 | 26.0381 | 296.8598 | 233.252 | 291.6019 | 209.4708 | 85.0795 | 1.4446 | 36.5447 | 52.7929 | 25.2973 | 2.2871 | 36.4376 | 83.3365 | 15.9317 | 21.5865 | 5.8654 | 136.1294 | 275.2733 | 211.6655 | 270.0154 | 187.8843
2012-10-30 00:15 | 26.0276 | 296.8147 | 233.2493 | 291.6299 | 209.3697 | 85.0922 | 1.4446 | 36.5447 | 52.7924 | 25.2974 | 2.2871 | 36.4373 | 83.3356 | 15.9317 | 21.5519 | 5.8663 | 136.128 | 275.2628 | 211.6974 | 270.078 | 187.8178
2012-10-30 00:15 | 26.0306 | 296.8177 | 233.2523 | 291.6213 | 209.2388 | 85.0985 | 1.4446 | 36.5435 | 52.789 | 25.2975 | 2.287 | 36.4361 | 83.3304 | 15.9316 | 21.6071 | 5.8667 | 136.1195 | 275.2106 | 211.6452 | 270.0142 | 187.6317
2012-10-30 00:15 | 26.0381 | 296.8214 | 233.2598 | 291.6327 | 209.3408 | 85.0922 | 1.4445 | 36.5423 | 52.7858 | 25.2973 | 2.287 | 36.4351 | 83.3255 | 15.9317 | 21.5664 | 5.8663 | 136.1113 | 275.255 | 211.6934 | 270.0663 | 187.7744
2012-10-30 00:15 | 26.0463 | 296.8642 | 233.2837 | 291.6486 | 209.3963 | 85.1049 | 1.4445 | 36.5411 | 52.782 | 25.2975 | 2.2869 | 36.4339 | 83.3205 | 15.9316 | 21.5706 | 5.8671 | 136.1025 | 275.2936 | 211.7131 | 270.078 | 187.8257
2012-10-30 00:15 | 26.0396 | 296.8383 | 233.3124 | 291.6573 | 209.3542 | 85.0962 | 1.4444 | 36.5406 | 52.7807 | 25.2974 | 2.2869 | 36.4333 | 83.3181 | 15.9315 | 21.576 | 5.8665 | 136.0988 | 275.2623 | 211.7364 | 270.0813 | 187.7782
2012-10-30 00:16 | 26.0396 | 296.8459 | 233.2849 | 291.6958 | 209.3187 | 85.091 | 1.4444 | 36.5405 | 52.7806 | 25.2973 | 2.2869 | 36.4332 | 83.3176 | 15.9315 | 21.5479 | 5.8662 | 136.0983 | 275.298 | 211.737 | 270.1479 | 187.7708
2012-10-30 00:16 | 26.0381 | 296.8521 | 233.2677 | 291.7174 | 209.3251 | 85.1084 | 1.4444 | 36.5387 | 52.7752 | 25.2974 | 2.2868 | 36.4313 | 83.3103 | 15.9313 | 21.5865 | 5.8674 | 136.0856 | 275.2656 | 211.6812 | 270.1309 | 187.7386
2012-10-30 00:16 | 26.0605 | 296.8784 | 233.2979 | 291.6666 | 209.3357 | 85.0881 | 1.4443 | 36.5369 | 52.7701 | 25.2973 | 2.2867 | 36.4295 | 83.3018 | 15.9313 | 21.5929 | 5.866 | 136.0719 | 275.2855 | 211.705 | 270.0737 | 187.7428
2012-10-30 00:16 | 26.0493 | 296.871 | 233.3182 | 291.7247 | 209.1827 | 85.1032 | 1.4443 | 36.5365 | 52.7688 | 25.2974 | 2.2866 | 36.429 | 83.3002 | 15.9312 | 21.5295 | 5.867 | 136.069 | 275.3415 | 211.7887 | 270.1952 | 187.6532
2012-10-30 00:16 | 26.0538 | 296.9024 | 233.2559 | 291.6984 | 209.3132 | 85.0991 | 1.4443 | 36.5364 | 52.7685 | 25.2975 | 2.2866 | 36.4292 | 83.3007 | 15.9312 | 21.6022 | 5.8667 | 136.0692 | 275.3002 | 211.6537 | 270.0962 | 187.711
2012-10-30 00:16 | 26.0598 | 296.9007 | 233.2972 | 291.7391 | 209.4177 | 85.1165 | 1.4442 | 36.5347 | 52.7634 | 25.2975 | 2.2865 | 36.4274 | 83.2923 | 15.9313 | 21.552 | 5.8679 | 136.0557 | 275.3487 | 211.7452 | 270.1871 | 187.8657
2012-10-30 00:16 | 26.0613 | 296.9214 | 233.2987 | 291.7213 | 209.4664 | 85.1118 | 1.4442 | 36.5345 | 52.7629 | 25.2975 | 2.2865 | 36.4273 | 83.2927 | 15.9312 | 21.5736 | 5.8676 | 136.0556 | 275.3478 | 211.7251 | 270.1477 | 187.8928
2012-10-30 00:16 | 26.062 | 296.9183 | 233.3151 | 291.7413 | 209.5419 | 85.0939 | 1.4441 | 36.5331 | 52.7589 | 25.2975 | 2.2864 | 36.4258 | 83.2855 | 15.9312 | 21.5342 | 5.8664 | 136.0444 | 275.3841 | 211.7809 | 270.2071 | 188.0077
2012-10-30 00:16 | 26.062 | 296.8952 | 233.3937 | 291.7567 | 209.7703 | 85.1072 | 1.4441 | 36.5318 | 52.7551 | 25.2975 | 2.2864 | 36.4245 | 83.2793 | 15.9312 | 21.5663 | 5.8673 | 136.0344 | 275.3289 | 211.8274 | 270.1904 | 188.204
2012-10-30 00:16 | 26.0613 | 296.9138 | 233.4205 | 291.8099 | 209.6869 | 85.1043 | 1.4441 | 36.5326 | 52.7566 | 25.2979 | 2.2864 | 36.4253 | 83.2832 | 15.9312 | 21.5656 | 5.8671 | 136.0397 | 275.3482 | 211.8549 | 270.2443 | 188.1213
2012-10-30 00:16 | 26.068 | 296.9973 | 233.4744 | 291.7742 | 209.7527 | 85.1072 | 1.444 | 36.5311 | 52.7527 | 25.2977 | 2.2863 | 36.424 | 83.2778 | 15.9311 | 21.5723 | 5.8673 | 136.0305 | 275.425 | 211.9021 | 270.2019 | 188.1804
2012-10-30 00:16 | 26.0643 | 296.9283 | 233.5021 | 291.8206 | 209.6703 | 85.1286 | 1.444 | 36.53 | 52.7496 | 25.2976 | 2.2863 | 36.4228 | 83.2726 | 15.931 | 21.5405 | 5.8688 | 136.0222 | 275.3878 | 211.9616 | 270.2801 | 188.1298
2012-10-30 00:16 | 26.0658 | 296.9528 | 233.5311 | 291.8144 | 209.6048 | 85.1101 | 1.444 | 36.5303 | 52.7504 | 25.2977 | 2.2863 | 36.4232 | 83.2742 | 15.9311 | 21.554 | 5.8675 | 136.0246 | 275.3988 | 211.9771 | 270.2604 | 188.0508
2012-10-30 00:16 | 26.071 | 296.9696 | 233.5363 | 291.8388 | 209.7281 | 85.1084 | 1.444 | 36.5296 | 52.7482 | 25.2978 | 2.2863 | 36.4222 | 83.271 | 15.9308 | 21.5753 | 5.8674 | 136.0193 | 275.3943 | 211.961 | 270.2635 | 188.1528
2012-10-30 00:16 | 26.0613 | 296.9752 | 233.5659 | 291.8022 | 209.6239 | 85.1228 | 1.4439 | 36.5281 | 52.7443 | 25.2976 | 2.2862 | 36.4208 | 83.2645 | 15.9308 | 21.5656 | 5.8684 | 136.0088 | 275.4096 | 212.0003 | 270.2366 | 188.0583
2012-10-30 00:16 | 26.0702 | 297.0149 | 233.5866 | 291.838 | 209.6762 | 85.1061 | 1.4439 | 36.5271 | 52.7411 | 25.2977 | 2.2861 | 36.4199 | 83.2607 | 15.9308 | 21.5464 | 5.8672 | 136.0018 | 275.4685 | 212.0402 | 270.2916 | 188.1298
2012-10-30 00:16 | 26.074 | 297.0302 | 233.555 | 291.888 | 209.6169 | 85.1292 | 1.4439 | 36.5262 | 52.7385 | 25.2977 | 2.2861 | 36.4191 | 83.2569 | 15.9308 | 21.5542 | 5.8688 | 135.9954 | 275.476 | 212.0008 | 270.3338 | 188.0627
2012-10-30 00:16 | 26.0732 | 296.9987 | 233.5621 | 291.8641 | 209.6477 | 85.1165 | 1.4438 | 36.5259 | 52.7379 | 25.2976 | 2.2861 | 36.4187 | 83.2558 | 15.9307 | 21.5855 | 5.8679 | 135.9937 | 275.4132 | 211.9766 | 270.2786 | 188.0622
2012-10-30 00:16 | 26.0725 | 297.0172 | 233.526 | 291.8827 | 209.6824 | 85.124 | 1.4438 | 36.5249 | 52.7348 | 25.2977 | 2.286 | 36.4177 | 83.251 | 15.9307 | 21.6008 | 5.8684 | 135.9858 | 275.4164 | 211.9252 | 270.2819 | 188.0816
2012-10-30 00:16 | 26.0852 | 297.1144 | 233.5702 | 291.957 | 209.699 | 85.1234 | 1.4438 | 36.5233 | 52.7307 | 25.2975 | 2.2859 | 36.4163 | 83.2451 | 15.9306 | 21.5774 | 5.8684 | 135.9758 | 275.537 | 211.9928 | 270.3796 | 188.1216
2012-10-30 00:16 | 26.0777 | 297.057 | 233.5941 | 291.8763 | 209.534 | 85.1147 | 1.4432 | 36.5104 | 52.6931 | 25.2976 | 2.2849 | 36.4001 | 83.1716 | 15.9306 | 21.5739 | 5.8678 | 135.8648 | 275.4831 | 212.0202 | 270.3024 | 187.9601
2012-10-30 00:16 | 26.0747 | 297.077 | 233.5164 | 291.9311 | 209.279 | 85.1136 | 1.4392 | 36.4069 | 52.3984 | 25.2959 | 2.2787 | 36.3 | 82.7167 | 15.9302 | 21.587 | 5.8677 | 135.1152 | 275.49 | 211.9294 | 270.3441 | 187.692
2012-10-30 00:16 | 26.0717 | 297.0049 | 233.4466 | 291.9396 | 209.3272 | 85.1078 | 1.4391 | 36.4037 | 52.3894 | 25.2958 | 2.2785 | 36.2968 | 82.7021 | 15.9301 | 21.6281 | 5.8673 | 135.0915 | 275.3768 | 211.8185 | 270.3115 | 187.6991
2012-10-30 00:17 | 26.074 | 297.0571 | 233.3782 | 291.9458 | 209.4712 | 85.1118 | 1.4392 | 36.407 | 52.3988 | 25.2958 | 2.2787 | 36.3003 | 82.7181 | 15.9301 | 21.6264 | 5.8676 | 135.1169 | 275.4307 | 211.7518 | 270.3194 | 187.8448
2012-10-30 00:17 | 26.0658 | 297.0566 | 233.3857 | 291.9183 | 209.5457 | 85.1113 | 1.4395 | 36.4129 | 52.415 | 25.2963 | 2.2791 | 36.306 | 82.7446 | 15.9301 | 21.6423 | 5.8676 | 135.1596 | 275.4143 | 211.7434 | 270.276 | 187.9034
2012-10-30 00:17 | 26.0755 | 297.0471 | 233.4347 | 291.9819 | 209.898 | 85.1107 | 1.4396 | 36.4158 | 52.4234 | 25.2961 | 2.2793 | 36.309 | 82.7582 | 15.93 | 21.6319 | 5.8675 | 135.1816 | 275.4152 | 211.8028 | 270.35 | 188.2661
2012-10-30 00:17 | 26.0732 | 297.064 | 233.4874 | 291.9334 | 209.9902 | 85.1188 | 1.44 | 36.4257 | 52.4515 | 25.2964 | 2.2799 | 36.3191 | 82.8036 | 15.9302 | 21.5815 | 5.8681 | 135.255 | 275.4825 | 211.9059 | 270.3519 | 188.4087
2012-10-30 00:17 | 26.0643 | 297.0781 | 233.5414 | 291.9361 | 209.623 | 85.1014 | 1.4427 | 36.4963 | 52.6537 | 25.2971 | 2.2843 | 36.3891 | 83.1223 | 15.9304 | 21.5686 | 5.8669 | 135.7759 | 275.5095 | 211.9728 | 270.3675 | 188.0544
2012-10-30 00:17 | 26.068 | 297.0665 | 233.498 | 291.9167 | 209.6464 | 85.1269 | 1.4432 | 36.5102 | 52.6929 | 25.2975 | 2.2851 | 36.4025 | 83.1841 | 15.9303 | 21.5963 | 5.8686 | 135.877 | 275.4702 | 211.9017 | 270.3204 | 188.0501
2012-10-30 00:17 | 26.062 | 297.1027 | 233.4762 | 291.9376 | 209.7231 | 85.1072 | 1.4437 | 36.5214 | 52.7253 | 25.2974 | 2.2858 | 36.4143 | 83.2375 | 15.9303 | 21.5582 | 5.8673 | 135.9628 | 275.5445 | 211.918 | 270.3794 | 188.1649
2012-10-30 00:17 | 26.0568 | 297.1129 | 233.4828 | 291.9401 | 209.7691 | 85.1211 | 1.4441 | 36.5325 | 52.7571 | 25.2976 | 2.2866 | 36.4256 | 83.2893 | 15.9303 | 21.5851 | 5.8682 | 136.0463 | 275.5278 | 211.8977 | 270.355 | 188.184
2012-10-30 00:17 | 26.0545 | 297.1029 | 233.5316 | 291.934 | 209.814 | 85.1257 | 1.4444 | 36.5401 | 52.7782 | 25.2979 | 2.287 | 36.4329 | 83.3226 | 15.9303 | 21.5989 | 5.8686 | 136.1008 | 275.504 | 211.9327 | 270.3351 | 188.2151
2012-10-30 00:17 | 26.056 | 297.1044 | 233.5802 | 291.9701 | 209.6895 | 85.1159 | 1.4445 | 36.5434 | 52.7874 | 25.2981 | 2.2872 | 36.4361 | 83.338 | 15.9302 | 21.6044 | 5.8679 | 136.1254 | 275.5 | 211.9758 | 270.3657 | 188.0851
2012-10-30 00:17 | 26.0545 | 297.1221 | 233.6102 | 291.9378 | 209.7747 | 85.1245 | 1.4446 | 36.5458 | 52.7941 | 25.2982 | 2.2874 | 36.4385 | 83.3489 | 15.9302 | 21.6109 | 5.8685 | 136.143 | 275.5112 | 211.9993 | 270.3269 | 188.1638
2012-10-30 00:17 | 26.0523 | 297.1161 | 233.5883 | 291.9318 | 209.7843 | 85.1419 | 1.4447 | 36.5476 | 52.7995 | 25.2981 | 2.2875 | 36.4402 | 83.3565 | 15.9302 | 21.6047 | 5.8697 | 136.1561 | 275.5114 | 211.9836 | 270.3271 | 188.1796
2012-10-30 00:17 | 26.0553 | 297.1729 | 233.6188 | 291.9694 | 209.74 | 85.1222 | 1.4447 | 36.5486 | 52.8024 | 25.2981 | 2.2876 | 36.4412 | 83.3611 | 15.9302 | 21.5796 | 5.8683 | 136.1635 | 275.5933 | 212.0392 | 270.3898 | 188.1604
2012-10-30 00:17 | 26.0575 | 297.1559 | 233.6092 | 291.9601 | 209.7265 | 85.1338 | 1.4448 | 36.55 | 52.8065 | 25.298 | 2.2877 | 36.4426 | 83.3679 | 15.9301 | 21.5618 | 5.8691 | 136.1744 | 275.5941 | 212.0474 | 270.3983 | 188.1647
2012-10-30 00:17 | 26.0583 | 297.1413 | 233.6258 | 291.984 | 209.7824 | 85.139 | 1.4448 | 36.5499 | 52.806 | 25.2982 | 2.2877 | 36.4425 | 83.3683 | 15.93 | 21.6067 | 5.8695 | 136.1743 | 275.5346 | 212.0191 | 270.3773 | 188.1757
2012-10-30 00:17 | 26.059 | 297.1497 | 233.6461 | 291.9616 | 209.8618 | 85.1373 | 1.4448 | 36.5503 | 52.807 | 25.2982 | 2.2877 | 36.4428 | 83.369 | 15.9301 | 21.5793 | 5.8694 | 136.176 | 275.5704 | 212.0668 | 270.3823 | 188.2825
2012-10-30 00:17 | 26.0538 | 297.1445 | 233.6763 | 291.9718 | 209.9748 | 85.1402 | 1.4447 | 36.5496 | 52.8045 | 25.2985 | 2.2876 | 36.4422 | 83.3669 | 15.93 | 21.5661 | 5.8696 | 136.1713 | 275.5784 | 212.1102 | 270.4057 | 188.4087
2012-10-30 00:17 | 26.0515 | 297.1499 | 233.6857 | 291.9502 | 209.7717 | 85.1575 | 1.4447 | 36.5492 | 52.8035 | 25.2984 | 2.2876 | 36.4416 | 83.3647 | 15.9299 | 21.5678 | 5.8708 | 136.1682 | 275.5821 | 212.1179 | 270.3824 | 188.2039
2012-10-30 00:17 | 26.05 | 297.1369 | 233.6882 | 291.9834 | 209.6756 | 85.1448 | 1.4447 | 36.5496 | 52.8045 | 25.2985 | 2.2877 | 36.4421 | 83.3673 | 15.9299 | 21.5703 | 5.8699 | 136.1718 | 275.5666 | 212.1179 | 270.4131 | 188.1053
2012-10-30 00:17 | 26.0403 | 297.154 | 233.6785 | 291.9698 | 209.8707 | 85.161 | 1.4447 | 36.5496 | 52.8047 | 25.2985 | 2.2877 | 36.4423 | 83.3677 | 15.9299 | 21.5486 | 5.871 | 136.1723 | 275.6054 | 212.1299 | 270.4212 | 188.3221
2012-10-30 00:17 | 26.0508 | 297.1645 | 233.7086 | 291.9957 | 209.897 | 85.1667 | 1.4447 | 36.5482 | 52.8004 | 25.2985 | 2.2876 | 36.4408 | 83.3612 | 15.9298 | 21.6193 | 5.8714 | 136.1616 | 275.5452 | 212.0893 | 270.3764 | 188.2777
2012-10-30 00:17 | 26.0523 | 297.1699 | 233.769 | 292.0434 | 209.7685 | 85.1494 | 1.4447 | 36.5475 | 52.7985 | 25.2984 | 2.2875 | 36.4401 | 83.3585 | 15.9298 | 21.5927 | 5.8702 | 136.1571 | 275.5772 | 212.1763 | 270.4507 | 188.1758
2012-10-30 00:17 | 26.0583 | 297.1913 | 233.779 | 292.0533 | 209.7903 | 85.1604 | 1.4446 | 36.5458 | 52.7928 | 25.2989 | 2.2874 | 36.4384 | 83.3507 | 15.9297 | 21.6228 | 5.871 | 136.1436 | 275.5685 | 212.1562 | 270.4305 | 188.1675
2012-10-30 00:17 | 26.0538 | 297.1944 | 233.7509 | 292.068 | 209.8212 | 85.1638 | 1.4445 | 36.5449 | 52.7902 | 25.2988 | 2.2874 | 36.4375 | 83.3467 | 15.9297 | 21.5942 | 5.8712 | 136.1369 | 275.6002 | 212.1567 | 270.4738 | 188.227
2012-10-30 00:17 | 26.0545 | 297.2105 | 233.7359 | 292.0379 | 209.8377 | 85.1569 | 1.4445 | 36.5444 | 52.7892 | 25.2986 | 2.2874 | 36.4371 | 83.3451 | 15.9297 | 21.5828 | 5.8707 | 136.1342 | 275.6277 | 212.1531 | 270.4551 | 188.2549
2012-10-30 00:17 | 26.053 | 297.1898 | 233.7855 | 292.0787 | 209.9346 | 85.1754 | 1.4445 | 36.5436 | 52.787 | 25.2986 | 2.2873 | 36.4363 | 83.3417 | 15.9296 | 21.5854 | 5.872 | 136.1287 | 275.6044 | 212.2001 | 270.4933 | 188.3492
2012-10-30 00:18 | 26.0493 | 297.2207 | 233.7975 | 292.0635 | 209.8915 | 85.1673 | 1.4445 | 36.5424 | 52.7838 | 25.2984 | 2.2872 | 36.4351 | 83.3361 | 15.9297 | 21.5817 | 5.8714 | 136.1199 | 275.639 | 212.2158 | 270.4818 | 188.3098
2012-10-30 00:18 | 26.0583 | 297.2297 | 233.775 | 292.0648 | 209.6603 | 85.165 | 1.4444 | 36.5418 | 52.7809 | 25.2991 | 2.2872 | 36.4345 | 83.3336 | 15.9296 | 21.5586 | 5.8713 | 136.1144 | 275.6711 | 212.2164 | 270.5062 | 188.1017
2012-10-30 00:18 | 26.0605 | 297.2549 | 233.7301 | 292.0785 | 209.698 | 85.1569 | 1.4444 | 36.5408 | 52.7785 | 25.2988 | 2.2872 | 36.4336 | 83.33 | 15.9295 | 21.6169 | 5.8707 | 136.1085 | 275.638 | 212.1132 | 270.4616 | 188.0811
2012-10-30 00:18 | 26.0463 | 297.2407 | 233.7355 | 292.099 | 209.8846 | 85.1783 | 1.4444 | 36.5401 | 52.777 | 25.2985 | 2.2871 | 36.4328 | 83.3267 | 15.9295 | 21.5546 | 5.8722 | 136.1036 | 275.6861 | 212.1809 | 270.5444 | 188.33
2012-10-30 00:18 | 26.0665 | 297.234 | 233.8147 | 292.1269 | 209.8615 | 85.1708 | 1.4443 | 36.5394 | 52.7746 | 25.2987 | 2.2871 | 36.4318 | 83.3229 | 15.9293 | 21.5627 | 5.8717 | 136.0975 | 275.6713 | 212.252 | 270.5642 | 188.2988
2012-10-30 00:18 | 26.068 | 297.224 | 233.7612 | 292.1476 | 209.7409 | 85.1586 | 1.4443 | 36.5383 | 52.7721 | 25.2984 | 2.287 | 36.431 | 83.3189 | 15.9293 | 21.5562 | 5.8708 | 136.091 | 275.6678 | 212.205 | 270.5914 | 188.1847
2012-10-30 00:18 | 26.0605 | 297.2511 | 233.7969 | 292.1132 | 209.8043 | 85.1754 | 1.4442 | 36.537 | 52.7679 | 25.2986 | 2.287 | 36.4297 | 83.3129 | 15.9294 | 21.5688 | 5.872 | 136.0807 | 275.6823 | 212.2281 | 270.5444 | 188.2355
2012-10-30 00:18 | 26.065 | 297.2556 | 233.7778 | 292.1254 | 209.8127 | 85.1558 | 1.4441 | 36.5349 | 52.7615 | 25.2987 | 2.2868 | 36.4274 | 83.3029 | 15.9293 | 21.6094 | 5.8706 | 136.0644 | 275.6462 | 212.1684 | 270.516 | 188.2033
2012-10-30 00:18 | 26.0568 | 297.2436 | 233.7893 | 292.1287 | 209.7888 | 85.1656 | 1.4441 | 36.5339 | 52.7584 | 25.2989 | 2.2868 | 36.4267 | 83.2997 | 15.9293 | 21.5571 | 5.8713 | 136.058 | 275.6865 | 212.2322 | 270.5716 | 188.2317
2012-10-30 00:18 | 26.0725 | 297.2977 | 233.805 | 292.156 | 209.7572 | 85.1754 | 1.4441 | 36.5331 | 52.7561 | 25.2988 | 2.2867 | 36.4258 | 83.296 | 15.9292 | 21.5647 | 5.872 | 136.0521 | 275.733 | 212.2403 | 270.5913 | 188.1925
2012-10-30 00:18 | 26.0762 | 297.3052 | 233.8283 | 292.1905 | 210.0799 | 85.1818 | 1.4439 | 36.5298 | 52.7465 | 25.2988 | 2.2865 | 36.4225 | 83.2813 | 15.9291 | 21.5363 | 5.8724 | 136.0279 | 275.7689 | 212.292 | 270.6542 | 188.5436
2012-10-30 00:18 | 26.08 | 297.3359 | 233.8989 | 292.1827 | 210.1821 | 85.1852 | 1.4438 | 36.527 | 52.7388 | 25.2987 | 2.2864 | 36.4197 | 83.2693 | 15.9289 | 21.5762 | 5.8727 | 136.0081 | 275.7597 | 212.3227 | 270.6065 | 188.6059
2012-10-30 00:18 | 26.0912 | 297.351 | 234.0162 | 292.2247 | 210.0358 | 85.1985 | 1.4437 | 36.5232 | 52.728 | 25.2985 | 2.2862 | 36.416 | 83.2535 | 15.9288 | 21.5915 | 5.8736 | 135.9816 | 275.7595 | 212.4247 | 270.6332 | 188.4443
2012-10-30 00:18 | 26.0979 | 297.3769 | 234.0229 | 292.2699 | 210.2118 | 85.2003 | 1.4427 | 36.4986 | 52.6571 | 25.2985 | 2.2846 | 36.3907 | 83.1384 | 15.9286 | 21.562 | 5.8737 | 135.7955 | 275.8149 | 212.4609 | 270.7079 | 188.6498
2012-10-30 00:19 | 26.1062 | 297.3813 | 234.0508 | 292.259 | 209.8736 | 85.1881 | 1.4391 | 36.4056 | 52.3923 | 25.297 | 2.2789 | 36.2986 | 82.7203 | 15.9283 | 21.5543 | 5.8729 | 135.1126 | 275.827 | 212.4965 | 270.7047 | 188.3193
2012-10-30 00:19 | 26.1024 | 297.3391 | 233.7013 | 292.2629 | 208.4404 | 85.1789 | 1.4389 | 36.4 | 52.3762 | 25.297 | 2.2786 | 36.293 | 82.6956 | 15.9281 | 21.5746 | 5.8722 | 135.0718 | 275.7645 | 212.1267 | 270.6883 | 186.8658
2012-10-30 00:19 | 26.0979 | 297.3577 | 233.4571 | 292.2622 | 209.0541 | 85.1638 | 1.4396 | 36.417 | 52.4252 | 25.297 | 2.2796 | 36.3098 | 82.7718 | 15.9282 | 21.5982 | 5.8712 | 135.197 | 275.7595 | 211.8589 | 270.664 | 187.4559
2012-10-30 00:19 | 26.0919 | 297.367 | 233.3726 | 292.2755 | 208.8079 | 85.1812 | 1.4432 | 36.5091 | 52.6883 | 25.2981 | 2.2854 | 36.4017 | 83.1908 | 15.9283 | 21.5922 | 5.8724 | 135.8791 | 275.7748 | 211.7804 | 270.6833 | 187.2157
2012-10-30 00:19 | 26.089 | 297.391 | 233.3932 | 292.2764 | 209.3327 | 85.1893 | 1.444 | 36.53 | 52.7479 | 25.2985 | 2.2867 | 36.4227 | 83.2868 | 15.9282 | 21.5973 | 5.873 | 136.0346 | 275.7937 | 211.7959 | 270.6791 | 187.7354
2012-10-30 00:19 | 26.0845 | 297.4173 | 233.4005 | 292.2758 | 208.9344 | 85.1968 | 1.4443 | 36.5384 | 52.7715 | 25.2989 | 2.2872 | 36.4309 | 83.3249 | 15.9282 | 21.6048 | 5.8735 | 136.0964 | 275.8125 | 211.7957 | 270.671 | 187.3296
2012-10-30 00:19 | 26.0822 | 297.438 | 233.3314 | 292.2388 | 208.8218 | 85.2003 | 1.4445 | 36.5437 | 52.7869 | 25.2988 | 2.2875 | 36.4363 | 83.3499 | 15.9281 | 21.6025 | 5.8737 | 136.1367 | 275.8355 | 211.7289 | 270.6363 | 187.2193
2012-10-30 00:19 | 26.0934 | 297.4108 | 233.3465 | 292.2385 | 209.404 | 85.2008 | 1.4446 | 36.5469 | 52.7957 | 25.2989 | 2.2877 | 36.4394 | 83.3637 | 15.9281 | 21.5896 | 5.8737 | 136.1594 | 275.8212 | 211.7569 | 270.6489 | 187.8144
2012-10-30 00:19 | 26.086 | 297.4072 | 233.3824 | 292.2619 | 209.4636 | 85.213 | 1.4447 | 36.5483 | 52.8 | 25.2989 | 2.2878 | 36.441 | 83.3714 | 15.9281 | 21.6184 | 5.8746 | 136.1714 | 275.7888 | 211.764 | 270.6435 | 187.8452
2012-10-30 00:19 | 26.086 | 297.4341 | 233.3981 | 292.2118 | 209.5463 | 85.2176 | 1.4447 | 36.5502 | 52.8047 | 25.2992 | 2.2879 | 36.4427 | 83.3789 | 15.9281 | 21.6344 | 5.8749 | 136.1836 | 275.7997 | 211.7637 | 270.5774 | 187.9119
2012-10-30 00:19 | 26.089 | 297.4025 | 233.5032 | 292.211 | 209.6241 | 85.2037 | 1.4447 | 36.5504 | 52.8051 | 25.2993 | 2.288 | 36.4428 | 83.3793 | 15.9281 | 21.6053 | 5.8739 | 136.1844 | 275.7972 | 211.8979 | 270.6057 | 188.0188
2012-10-30 00:19 | 26.0792 | 297.3889 | 233.5524 | 292.2281 | 209.4449 | 85.2268 | 1.4447 | 36.5508 | 52.8066 | 25.2992 | 2.288 | 36.4435 | 83.3834 | 15.9279 | 21.5875 | 5.8755 | 136.19 | 275.8014 | 211.9649 | 270.6406 | 187.8574
2012-10-30 00:19 | 26.0852 | 297.4141 | 233.5152 | 292.3034 | 209.6124 | 85.2159 | 1.4447 | 36.5505 | 52.8059 | 25.299 | 2.288 | 36.4431 | 83.3812 | 15.928 | 21.6015 | 5.8748 | 136.1871 | 275.8126 | 211.9137 | 270.7019 | 188.0109
2012-10-30 00:19 | 26.0785 | 297.4074 | 233.587 | 292.2351 | 209.7632 | 85.2193 | 1.4447 | 36.55 | 52.8037 | 25.2994 | 2.288 | 36.4424 | 83.3787 | 15.9279 | 21.5788 | 5.875 | 136.1824 | 275.8286 | 212.0082 | 270.6563 | 188.1844
2012-10-30 00:19 | 26.0777 | 297.4066 | 233.598 | 292.2651 | 209.7782 | 85.2182 | 1.4447 | 36.5493 | 52.8019 | 25.2993 | 2.2879 | 36.4418 | 83.3753 | 15.928 | 21.59 | 5.8749 | 136.1773 | 275.8166 | 212.008 | 270.6751 | 188.1882
2012-10-30 00:19 | 26.086 | 297.4111 | 233.6456 | 292.2773 | 209.7668 | 85.2211 | 1.4446 | 36.5484 | 52.7994 | 25.2992 | 2.2879 | 36.4409 | 83.372 | 15.9279 | 21.6063 | 5.8751 | 136.1714 | 275.8048 | 212.0393 | 270.671 | 188.1605
2012-10-30 00:19 | 26.0845 | 297.4134 | 233.6834 | 292.2604 | 209.7298 | 85.2182 | 1.4446 | 36.5479 | 52.7982 | 25.2991 | 2.2879 | 36.4404 | 83.3704 | 15.9278 | 21.5928 | 5.8749 | 136.1686 | 275.8206 | 212.0906 | 270.6676 | 188.137
2012-10-30 00:19 | 26.0747 | 297.3767 | 233.705 | 292.266 | 209.8106 | 85.2188 | 1.4445 | 36.546 | 52.7922 | 25.2993 | 2.2877 | 36.4386 | 83.3611 | 15.9279 | 21.599 | 5.875 | 136.1533 | 275.7777 | 212.106 | 270.667 | 188.2116
2012-10-30 00:20 | 26.0822 | 297.4034 | 233.7322 | 292.2889 | 209.9796 | 85.2205 | 1.4445 | 36.5455 | 52.7902 | 25.2996 | 2.2877 | 36.4382 | 83.3586 | 15.928 | 21.5423 | 5.8751 | 136.1488 | 275.8611 | 212.1899 | 270.7466 | 188.4373
2012-10-30 00:20 | 26.08 | 297.4589 | 233.781 | 292.2713 | 209.9301 | 85.2245 | 1.4445 | 36.5447 | 52.7882 | 25.2995 | 2.2876 | 36.4373 | 83.3554 | 15.9279 | 21.6204 | 5.8754 | 136.1435 | 275.8385 | 212.1606 | 270.6509 | 188.3097
2012-10-30 00:20 | 26.0717 | 297.4083 | 233.7727 | 292.2668 | 209.8509 | 85.2245 | 1.4445 | 36.5445 | 52.7876 | 25.2995 | 2.2876 | 36.437 | 83.3543 | 15.9279 | 21.5639 | 5.8754 | 136.1418 | 275.8444 | 212.2088 | 270.7029 | 188.287
2012-10-30 00:20 | 26.083 | 297.4234 | 233.8272 | 292.2858 | 210.0434 | 85.2436 | 1.4445 | 36.5448 | 52.7883 | 25.2997 | 2.2877 | 36.4373 | 83.3559 | 15.9278 | 21.5632 | 5.8767 | 136.1442 | 275.8602 | 212.264 | 270.7226 | 188.4802
2012-10-30 00:20 | 26.0867 | 297.4617 | 233.8427 | 292.2972 | 210.1416 | 85.239 | 1.4444 | 36.5439 | 52.7851 | 25.2999 | 2.2876 | 36.4365 | 83.352 | 15.9279 | 21.587 | 5.8764 | 136.1371 | 275.8747 | 212.2557 | 270.7102 | 188.5546
2012-10-30 00:20 | 26.0905 | 297.4233 | 233.8976 | 292.2895 | 210.106 | 85.2407 | 1.4444 | 36.5423 | 52.7808 | 25.2997 | 2.2875 | 36.4349 | 83.3451 | 15.9278 | 21.5827 | 5.8765 | 136.1259 | 275.8406 | 212.3149 | 270.7068 | 188.5233
2012-10-30 00:20 | 26.0942 | 297.4154 | 233.9367 | 292.3201 | 210.0743 | 85.2372 | 1.4443 | 36.5403 | 52.7746 | 25.2999 | 2.2874 | 36.4327 | 83.336 | 15.9276 | 21.5543 | 5.8763 | 136.1105 | 275.8611 | 212.3824 | 270.7658 | 188.52
2012-10-30 00:20 | 26.0957 | 297.4822 | 233.9382 | 292.3024 | 209.9497 | 85.2494 | 1.4442 | 36.539 | 52.771 | 25.2998 | 2.2873 | 36.4316 | 83.3306 | 15.9277 | 21.5478 | 5.8771 | 136.1015 | 275.9344 | 212.3904 | 270.7546 | 188.4019
2012-10-30 00:20 | 26.0934 | 297.4723 | 233.9398 | 292.3386 | 210.0026 | 85.2528 | 1.4442 | 36.539 | 52.7714 | 25.2997 | 2.2873 | 36.4317 | 83.3318 | 15.9275 | 21.5856 | 5.8773 | 136.1032 | 275.8867 | 212.3542 | 270.753 | 188.417
2012-10-30 00:20 | 26.0994 | 297.5205 | 233.9419 | 292.3676 | 209.918 | 85.2448 | 1.4442 | 36.5385 | 52.7701 | 25.2996 | 2.2873 | 36.4312 | 83.329 | 15.9276 | 21.6037 | 5.8768 | 136.0992 | 275.9168 | 212.3382 | 270.7639 | 188.3143
2012-10-30 00:20 | 26.0919 | 297.5399 | 233.9658 | 292.3524 | 209.9971 | 85.2661 | 1.4442 | 36.5378 | 52.7683 | 25.2995 | 2.2873 | 36.4307 | 83.3265 | 15.9276 | 21.5721 | 5.8782 | 136.0947 | 275.9678 | 212.3937 | 270.7803 | 188.425
2012-10-30 00:20 | 26.0979 | 297.5036 | 233.9836 | 292.37 | 210.0149 | 85.2633 | 1.4441 | 36.5366 | 52.764 | 25.2998 | 2.2872 | 36.4292 | 83.3202 | 15.9276 | 21.5781 | 5.8781 | 136.0842 | 275.9255 | 212.4055 | 270.7919 | 188.4368
2012-10-30 00:20 | 26.0942 | 297.5192 | 234.0349 | 292.3932 | 210.0939 | 85.2575 | 1.4441 | 36.5362 | 52.7625 | 25.3001 | 2.2872 | 36.4288 | 83.3192 | 15.9274 | 21.5985 | 5.8777 | 136.0816 | 275.9207 | 212.4364 | 270.7947 | 188.4954
2012-10-30 00:20 | 26.1002 | 297.5405 | 234.0252 | 292.3377 | 210.1157 | 85.2638 | 1.444 | 36.5344 | 52.7574 | 25.3 | 2.2871 | 36.427 | 83.3108 | 15.9274 | 21.5603 | 5.8781 | 136.0682 | 275.9802 | 212.4649 | 270.7774 | 188.5554
2012-10-30 00:20 | 26.0972 | 297.5337 | 234.0182 | 292.377 | 209.9709 | 85.2488 | 1.444 | 36.5336 | 52.7554 | 25.2999 | 2.287 | 36.4262 | 83.3082 | 15.9272 | 21.5493 | 5.8771 | 136.0636 | 275.9844 | 212.4689 | 270.8277 | 188.4216
2012-10-30 00:20 | 26.0964 | 297.5021 | 233.986 | 292.3839 | 209.9544 | 85.2534 | 1.444 | 36.5325 | 52.7526 | 25.2997 | 2.287 | 36.4251 | 83.303 | 15.9273 | 21.5686 | 5.8774 | 136.0556 | 275.9335 | 212.4174 | 270.8153 | 188.3858
2012-10-30 00:20 | 26.1069 | 297.5588 | 234.0358 | 292.4059 | 209.9058 | 85.2453 | 1.4439 | 36.5318 | 52.75 | 25.2999 | 2.2869 | 36.4244 | 83.2999 | 15.9272 | 21.5831 | 5.8768 | 136.0499 | 275.9757 | 212.4527 | 270.8228 | 188.3227
2012-10-30 00:20 | 26.1114 | 297.4979 | 234.0089 | 292.4297 | 210.0009 | 85.2517 | 1.4439 | 36.5296 | 52.7438 | 25.2999 | 2.2868 | 36.4223 | 83.2902 | 15.9273 | 21.5635 | 5.8773 | 136.034 | 275.9344 | 212.4454 | 270.8662 | 188.4374
2012-10-30 00:20 | 26.1077 | 297.5557 | 234.0287 | 292.4144 | 210.0602 | 85.2621 | 1.4438 | 36.5285 | 52.7408 | 25.2998 | 2.2867 | 36.4211 | 83.2849 | 15.9272 | 21.5317 | 5.878 | 136.0257 | 276.024 | 212.497 | 270.8827 | 188.5285
2012-10-30 00:20 | 26.1084 | 297.5756 | 234.0373 | 292.4575 | 210.0136 | 85.2696 | 1.4438 | 36.527 | 52.7364 | 25.2998 | 2.2866 | 36.4196 | 83.2778 | 15.9272 | 21.5565 | 5.8785 | 136.0142 | 276.0191 | 212.4808 | 270.901 | 188.4571
2012-10-30 00:20 | 26.1084 | 297.5526 | 234.0137 | 292.4228 | 209.7459 | 85.2494 | 1.4438 | 36.5271 | 52.7362 | 25.3 | 2.2866 | 36.4193 | 83.2778 | 15.927 | 21.5886 | 5.8771 | 136.0139 | 275.964 | 212.4251 | 270.8342 | 188.1573
2012-10-30 00:20 | 26.1129 | 297.5993 | 234.0064 | 292.4273 | 209.8134 | 85.2586 | 1.4437 | 36.5251 | 52.7308 | 25.2998 | 2.2865 | 36.4177 | 83.2701 | 15.927 | 21.569 | 5.8777 | 136.001 | 276.0303 | 212.4374 | 270.8583 | 188.2444
2012-10-30 00:20 | 26.1092 | 297.5149 | 233.9752 | 292.4467 | 209.9908 | 85.2702 | 1.4437 | 36.5245 | 52.7294 | 25.2997 | 2.2865 | 36.4169 | 83.2671 | 15.9269 | 21.5492 | 5.8785 | 135.9965 | 275.9657 | 212.426 | 270.8975 | 188.4416
2012-10-30 00:20 | 26.1114 | 297.5633 | 233.9814 | 292.4682 | 209.9418 | 85.2552 | 1.4436 | 36.5234 | 52.7256 | 25.3 | 2.2864 | 36.4162 | 83.2636 | 15.927 | 21.5876 | 5.8775 | 135.9892 | 275.9757 | 212.3938 | 270.8806 | 188.3542
2012-10-30 00:20 | 26.1092 | 297.5418 | 234.0106 | 292.4621 | 209.9435 | 85.2765 | 1.4436 | 36.5223 | 52.7231 | 25.2997 | 2.2864 | 36.4148 | 83.2577 | 15.9269 | 21.5212 | 5.879 | 135.9807 | 276.0206 | 212.4894 | 270.9409 | 188.4223
2012-10-30 00:21 | 26.1039 | 297.5711 | 233.9935 | 292.4453 | 209.8595 | 85.2754 | 1.4435 | 36.521 | 52.7193 | 25.2998 | 2.2863 | 36.4137 | 83.2531 | 15.9268 | 21.6042 | 5.8789 | 135.9724 | 275.9669 | 212.3893 | 270.8411 | 188.2553
2012-10-30 00:21 | 26.1122 | 297.5833 | 234.0018 | 292.4728 | 209.8363 | 85.2673 | 1.4422 | 36.4859 | 52.6191 | 25.2992 | 2.2842 | 36.3797 | 83.0985 | 15.9267 | 21.5322 | 5.8783 | 135.7176 | 276.0511 | 212.4696 | 270.9406 | 188.3041
2012-10-30 00:21 | 26.1151 | 297.5977 | 234.0243 | 292.4565 | 209.9928 | 85.2534 | 1.4397 | 36.4214 | 52.4355 | 25.2982 | 2.2801 | 36.3147 | 82.8028 | 15.9265 | 21.5712 | 5.8774 | 135.2383 | 276.0265 | 212.4531 | 270.8853 | 188.4216
2012-10-30 00:21 | 26.1151 | 297.567 | 233.989 | 292.495 | 209.9534 | 85.2575 | 1.4392 | 36.4087 | 52.3992 | 25.2979 | 2.2793 | 36.3018 | 82.7443 | 15.9264 | 21.5471 | 5.8777 | 135.1435 | 276.0199 | 212.4419 | 270.9479 | 188.4063
2012-10-30 00:21 | 26.1114 | 297.5479 | 233.9696 | 292.4913 | 210.0521 | 85.2731 | 1.4389 | 36.4007 | 52.3759 | 25.298 | 2.2789 | 36.2939 | 82.7083 | 15.9264 | 21.5555 | 5.8787 | 135.0842 | 275.9924 | 212.4141 | 270.9358 | 188.4966
2012-10-30 00:21 | 26.1122 | 297.5756 | 233.9272 | 292.5075 | 209.6433 | 85.2586 | 1.4387 | 36.3965 | 52.3647 | 25.2977 | 2.2786 | 36.2895 | 82.6892 | 15.9263 | 21.5603 | 5.8777 | 135.054 | 276.0153 | 212.3669 | 270.9472 | 188.083
2012-10-30 00:21 | 26.1069 | 297.5703 | 233.9101 | 292.4714 | 209.7837 | 85.2401 | 1.4387 | 36.3962 | 52.364 | 25.2976 | 2.2786 | 36.2892 | 82.6879 | 15.9262 | 21.5429 | 5.8765 | 135.0519 | 276.0274 | 212.3672 | 270.9285 | 188.2408
2012-10-30 00:21 | 26.1084 | 297.5756 | 233.9037 | 292.4806 | 209.738 | 85.2546 | 1.4389 | 36.4018 | 52.3802 | 25.2976 | 2.2789 | 36.2949 | 82.7138 | 15.9262 | 21.5444 | 5.8775 | 135.094 | 276.0312 | 212.3593 | 270.9362 | 188.1936
2012-10-30 00:21 | 26.1069 | 297.5703 | 233.8865 | 292.4906 | 209.8074 | 85.254 | 1.4392 | 36.4079 | 52.3972 | 25.2978 | 2.2793 | 36.301 | 82.7414 | 15.9263 | 21.5269 | 5.8774 | 135.1386 | 276.0434 | 212.3596 | 270.9637 | 188.2805
2012-10-30 00:21 | 26.1099 | 297.5425 | 233.8581 | 292.459 | 209.96 | 85.2453 | 1.4394 | 36.4129 | 52.4113 | 25.298 | 2.2796 | 36.3057 | 82.7629 | 15.9263 | 21.558 | 5.8768 | 135.1742 | 275.9845 | 212.3001 | 270.901 | 188.402
2012-10-30 00:21 | 26.1002 | 297.5636 | 233.8444 | 292.4531 | 210.033 | 85.2413 | 1.4404 | 36.4388 | 52.4857 | 25.298 | 2.2813 | 36.332 | 82.8824 | 15.9263 | 21.5844 | 5.8765 | 135.3681 | 275.9792 | 212.26 | 270.8687 | 188.4486
2012-10-30 00:21 | 26.0964 | 297.5713 | 233.9035 | 292.4609 | 209.9032 | 85.2609 | 1.4428 | 36.5023 | 52.6661 | 25.2993 | 2.2852 | 36.3952 | 83.1705 | 15.9264 | 21.6167 | 5.8779 | 135.8367 | 275.9546 | 212.2868 | 270.8442 | 188.2865
2012-10-30 00:21 | 26.0912 | 297.6045 | 233.8865 | 292.4711 | 209.5908 | 85.269 | 1.4433 | 36.5156 | 52.7044 | 25.2994 | 2.286 | 36.4083 | 83.2304 | 15.9265 | 21.5874 | 5.8784 | 135.9348 | 276.0171 | 212.2991 | 270.8837 | 188.0034
2012-10-30 00:21 | 26.0882 | 297.5785 | 233.8207 | 292.4758 | 209.6784 | 85.2719 | 1.4437 | 36.5266 | 52.7352 | 25.2999 | 2.2867 | 36.4189 | 83.278 | 15.9266 | 21.5644 | 5.8786 | 136.0132 | 276.0141 | 212.2563 | 270.9114 | 188.114
2012-10-30 00:21 | 26.0875 | 297.5893 | 233.8789 | 292.4058 | 209.7171 | 85.2858 | 1.4441 | 36.5347 | 52.7584 | 25.3 | 2.2872 | 36.427 | 83.3158 | 15.9265 | 21.5677 | 5.8796 | 136.0741 | 276.0216 | 212.3112 | 270.8381 | 188.1494
2012-10-30 00:21 | 26.0807 | 297.6363 | 233.8642 | 292.4529 | 209.7497 | 85.2708 | 1.4443 | 36.5408 | 52.7759 | 25.3 | 2.2876 | 36.4335 | 83.3451 | 15.9265 | 21.5649 | 5.8786 | 136.121 | 276.0714 | 212.2993 | 270.888 | 188.1848
2012-10-30 00:21 | 26.077 | 297.6172 | 233.8959 | 292.4222 | 209.6908 | 85.2742 | 1.4445 | 36.5453 | 52.7887 | 25.3001 | 2.2879 | 36.4378 | 83.3647 | 15.9265 | 21.5692 | 5.8788 | 136.1534 | 276.048 | 212.3267 | 270.853 | 188.1216
2012-10-30 00:21 | 26.0732 | 297.6057 | 233.8057 | 292.4107 | 209.7107 | 85.2864 | 1.4446 | 36.5476 | 52.7957 | 25.3 | 2.288 | 36.4401 | 83.3759 | 15.9264 | 21.5654 | 5.8796 | 136.1716 | 276.0403 | 212.2403 | 270.8453 | 188.1453
2012-10-30 00:21 | 26.0747 | 297.5573 | 233.815 | 292.4353 | 209.8303 | 85.2823 | 1.4446 | 36.5488 | 52.7987 | 25.3001 | 2.2881 | 36.4411 | 83.3804 | 15.9264 | 21.5228 | 5.8794 | 136.1791 | 276.0345 | 212.2922 | 270.9125 | 188.3075
2012-10-30 00:21 | 26.0747 | 297.5957 | 233.8386 | 292.4392 | 209.8697 | 85.2973 | 1.4446 | 36.5488 | 52.7985 | 25.3002 | 2.2881 | 36.4414 | 83.3826 | 15.9263 | 21.5147 | 5.8804 | 136.181 | 276.081 | 212.3239 | 270.9245 | 188.355
2012-10-30 00:21 | 26.0665 | 297.5875 | 233.8382 | 292.4502 | 209.9284 | 85.291 | 1.4446 | 36.5498 | 52.8016 | 25.3001 | 2.2882 | 36.4423 | 83.387 | 15.9262 | 21.5507 | 5.88 | 136.1887 | 276.0368 | 212.2875 | 270.8995 | 188.3777
2012-10-30 00:21 | 26.0665 | 297.5683 | 233.8775 | 292.4502 | 209.9639 | 85.2869 | 1.4446 | 36.5498 | 52.8015 | 25.3002 | 2.2882 | 36.4423 | 83.3872 | 15.9262 | 21.5146 | 5.8797 | 136.1888 | 276.0537 | 212.3629 | 270.9356 | 188.4493
2012-10-30 00:21 | 26.0725 | 297.5897 | 233.8992 | 292.4023 | 209.9856 | 85.2875 | 1.4446 | 36.5484 | 52.7975 | 25.3001 | 2.2881 | 36.4409 | 83.3809 | 15.9262 | 21.5487 | 5.8797 | 136.1784 | 276.041 | 212.3505 | 270.8536 | 188.4369
2012-10-30 00:21 | 26.0605 | 297.5623 | 233.8951 | 292.4519 | 210.0248 | 85.2973 | 1.4446 | 36.5479 | 52.796 | 25.3002 | 2.2881 | 36.4403 | 83.3783 | 15.9262 | 21.5246 | 5.8804 | 136.1742 | 276.0377 | 212.3705 | 270.9273 | 188.5002
2012-10-30 00:22 | 26.0628 | 297.5876 | 233.8895 | 292.4273 | 210.0901 | 85.2869 | 1.4446 | 36.5478 | 52.7961 | 25.3 | 2.2881 | 36.4405 | 83.3796 | 15.9261 | 21.543 | 5.8797 | 136.1758 | 276.0446 | 212.3465 | 270.8843 | 188.5471
2012-10-30 00:22 | 26.0702 | 297.5605 | 233.948 | 292.4347 | 209.9754 | 85.2973 | 1.4445 | 36.5457 | 52.7896 | 25.3002 | 2.2879 | 36.4381 | 83.3679 | 15.9262 | 21.5664 | 5.8804 | 136.1575 | 275.9941 | 212.3816 | 270.8683 | 188.409
2012-10-30 00:22 | 26.0658 | 297.6099 | 233.9672 | 292.4457 | 210.0065 | 85.2893 | 1.4445 | 36.5447 | 52.7871 | 25.3 | 2.2879 | 36.4374 | 83.3655 | 15.926 | 21.542 | 5.8798 | 136.1526 | 276.0679 | 212.4252 | 270.9037 | 188.4645
2012-10-30 00:22 | 26.0628 | 297.6261 | 233.9524 | 292.4504 | 210.2122 | 85.291 | 1.4444 | 36.5447 | 52.787 | 25.3001 | 2.2879 | 36.4372 | 83.3645 | 15.9261 | 21.555 | 5.88 | 136.1515 | 276.0711 | 212.3974 | 270.8954 | 188.6572
2012-10-30 00:22 | 26.0673 | 297.6229 | 234.0119 | 292.4241 | 210.0985 | 85.2973 | 1.4444 | 36.5436 | 52.7834 | 25.3002 | 2.2879 | 36.4362 | 83.3606 | 15.9259 | 21.5154 | 5.8804 | 136.1441 | 276.1075 | 212.4965 | 270.9087 | 188.5831
2012-10-30 00:22 | 26.0605 | 297.6084 | 234.0051 | 292.4635 | 210.2296 | 85.3083 | 1.4443 | 36.5417 | 52.7782 | 25.3002 | 2.2877 | 36.4342 | 83.3506 | 15.9261 | 21.5608 | 5.8812 | 136.1288 | 276.0476 | 212.4443 | 270.9027 | 188.6688
2012-10-30 00:22 | 26.0665 | 297.6221 | 234.07 | 292.4541 | 210.1883 | 85.2985 | 1.4442 | 36.5393 | 52.771 | 25.3002 | 2.2876 | 36.4319 | 83.3407 | 15.926 | 21.5828 | 5.8805 | 136.1118 | 276.0393 | 212.4872 | 270.8713 | 188.6055
2012-10-30 00:22 | 26.062 | 297.5945 | 234.0891 | 292.4342 | 210.2271 | 85.2939 | 1.4442 | 36.5395 | 52.7716 | 25.3002 | 2.2876 | 36.4321 | 83.3413 | 15.926 | 21.5663 | 5.8802 | 136.113 | 276.0282 | 212.5228 | 270.8679 | 188.6608
2012-10-30 00:22 | 26.0598 | 297.5885 | 234.0869 | 292.4628 | 210.2052 | 85.2979 | 1.4443 | 36.5402 | 52.7733 | 25.3004 | 2.2876 | 36.4327 | 83.3449 | 15.9259 | 21.552 | 5.8804 | 136.1182 | 276.0365 | 212.5349 | 270.9108 | 188.6532
2012-10-30 00:22 | 26.062 | 297.5753 | 234.1127 | 292.4457 | 210.1917 | 85.2956 | 1.4443 | 36.5405 | 52.7747 | 25.3002 | 2.2877 | 36.4332 | 83.3471 | 15.9259 | 21.5623 | 5.8803 | 136.1218 | 276.013 | 212.5504 | 270.8834 | 188.6294
2012-10-30 00:22 | 26.0635 | 297.6191 | 234.126 | 292.4626 | 210.0672 | 85.3193 | 1.4442 | 36.539 | 52.7701 | 25.3003 | 2.2876 | 36.4316 | 83.3401 | 15.9259 | 21.5718 | 5.8819 | 136.1102 | 276.0473 | 212.5542 | 270.8908 | 188.4954
2012-10-30 00:22 | 26.0635 | 297.6114 | 234.1063 | 292.4703 | 210.2759 | 85.3078 | 1.4442 | 36.5383 | 52.7683 | 25.3002 | 2.2875 | 36.4311 | 83.3377 | 15.9258 | 21.5517 | 5.8811 | 136.106 | 276.0597 | 212.5546 | 270.9186 | 188.7242
2012-10-30 00:22 | 26.0635 | 297.596 | 234.1142 | 292.4742 | 210.3271 | 85.3205 | 1.4441 | 36.5368 | 52.7641 | 25.3001 | 2.2874 | 36.4292 | 83.3299 | 15.9257 | 21.5035 | 5.882 | 136.094 | 276.0925 | 212.6107 | 270.9707 | 188.8236
2012-10-30 00:22 | 26.0628 | 297.6414 | 234.1645 | 292.5158 | 210.287 | 85.3049 | 1.4441 | 36.5359 | 52.7612 | 25.3002 | 2.2874 | 36.4284 | 83.3264 | 15.9257 | 21.5631 | 5.8809 | 136.0876 | 276.0783 | 212.6014 | 270.9527 | 188.7239
2012-10-30 00:22 | 26.0717 | 297.6311 | 234.1695 | 292.5247 | 210.0793 | 85.3182 | 1.444 | 36.5335 | 52.754 | 25.3004 | 2.2873 | 36.4263 | 83.3164 | 15.9257 | 21.596 | 5.8818 | 136.0704 | 276.0351 | 212.5735 | 270.9287 | 188.4833
2012-10-30 00:22 | 26.065 | 297.6475 | 234.2139 | 292.4872 | 210.238 | 85.3066 | 1.444 | 36.5326 | 52.7515 | 25.3004 | 2.2872 | 36.4251 | 83.3117 | 15.9256 | 21.5452 | 5.881 | 136.0632 | 276.1023 | 212.6687 | 270.942 | 188.6928
2012-10-30 00:22 | 26.062 | 297.6291 | 234.203 | 292.5304 | 210.2153 | 85.317 | 1.4439 | 36.5302 | 52.7449 | 25.3002 | 2.2871 | 36.4227 | 83.3006 | 15.9256 | 21.6064 | 5.8818 | 136.0455 | 276.0227 | 212.5966 | 270.924 | 188.6089
2012-10-30 00:22 | 26.074 | 297.6795 | 234.1993 | 292.5347 | 210.2352 | 85.3187 | 1.4439 | 36.5304 | 52.7456 | 25.3001 | 2.2871 | 36.4231 | 83.3022 | 15.9256 | 21.5903 | 5.8819 | 136.0478 | 276.0892 | 212.609 | 270.9444 | 188.6449
2012-10-30 00:22 | 26.074 | 297.6795 | 234.215 | 292.5039 | 210.2391 | 85.3078 | 1.4439 | 36.5301 | 52.7445 | 25.3002 | 2.2871 | 36.4225 | 83.3008 | 15.9254 | 21.6064 | 5.8811 | 136.0453 | 276.0731 | 212.6086 | 270.8975 | 188.6327
2012-10-30 00:22 | 26.0687 | 297.6281 | 234.1901 | 292.5101 | 210.0763 | 85.3095 | 1.4438 | 36.5298 | 52.7435 | 25.3003 | 2.2871 | 36.4225 | 83.3006 | 15.9255 | 21.5609 | 5.8812 | 136.0442 | 276.0672 | 212.6292 | 270.9492 | 188.5154
2012-10-30 00:22 | 26.0747 | 297.6994 | 234.1804 | 292.52 | 210.1335 | 85.3193 | 1.4438 | 36.529 | 52.7418 | 25.3001 | 2.287 | 36.4218 | 83.2973 | 15.9254 | 21.5509 | 5.8819 | 136.0391 | 276.1485 | 212.6295 | 270.9691 | 188.5826
2012-10-30 00:22 | 26.077 | 297.6825 | 234.1709 | 292.5146 | 210.1319 | 85.3135 | 1.4437 | 36.5274 | 52.7357 | 25.3007 | 2.2869 | 36.42 | 83.2894 | 15.9254 | 21.5331 | 5.8815 | 136.0251 | 276.1494 | 212.6378 | 270.9815 | 188.5988
2012-10-30 00:22 | 26.0807 | 297.6824 | 234.1824 | 292.5414 | 210.1238 | 85.3153 | 1.4437 | 36.5269 | 52.7354 | 25.3002 | 2.2869 | 36.4196 | 83.2874 | 15.9254 | 21.589 | 5.8816 | 136.0228 | 276.0934 | 212.5934 | 270.9524 | 188.5348
2012-10-30 00:22 | 26.083 | 297.677 | 234.1965 | 292.5206 | 210.2757 | 85.3286 | 1.4437 | 36.5255 | 52.7311 | 25.3003 | 2.2868 | 36.4181 | 83.2811 | 15.9253 | 21.5431 | 5.8826 | 136.0122 | 276.1339 | 212.6534 | 270.9775 | 188.7326
2012-10-30 00:22 | 26.0815 | 297.6601 | 234.1675 | 292.5422 | 210.2821 | 85.3343 | 1.4436 | 36.5245 | 52.7282 | 25.3004 | 2.2868 | 36.4173 | 83.2777 | 15.9252 | 21.5617 | 5.8829 | 136.0059 | 276.0984 | 212.6058 | 270.9805 | 188.7204
2012-10-30 00:23 | 26.0845 | 297.6708 | 234.198 | 292.5221 | 210.2418 | 85.3153 | 1.4436 | 36.5232 | 52.7247 | 25.3002 | 2.2867 | 36.4159 | 83.2712 | 15.9253 | 21.5567 | 5.8816 | 135.996 | 276.1141 | 212.6413 | 270.9654 | 188.6851
2012-10-30 00:23 | 26.0747 | 297.6802 | 234.2118 | 292.5277 | 210.1847 | 85.3245 | 1.4435 | 36.5219 | 52.7209 | 25.3001 | 2.2866 | 36.4145 | 83.2655 | 15.9251 | 21.5388 | 5.8823 | 135.9864 | 276.1414 | 212.673 | 270.9889 | 188.6459
2012-10-30 00:23 | 26.0822 | 297.7185 | 234.1918 | 292.5467 | 210.1607 | 85.3268 | 1.4435 | 36.5207 | 52.7178 | 25.3 | 2.2866 | 36.4134 | 83.2615 | 15.925 | 21.5303 | 5.8824 | 135.9793 | 276.1882 | 212.6615 | 271.0164 | 188.6304
2012-10-30 00:23 | 26.0837 | 297.7007 | 234.2326 | 292.5867 | 210.3197 | 85.3216 | 1.4435 | 36.5201 | 52.7161 | 25.3001 | 2.2865 | 36.4127 | 83.258 | 15.925 | 21.5438 | 5.8821 | 135.974 | 276.1569 | 212.6888 | 271.0429 | 188.7759
2012-10-30 00:23 | 26.0867 | 297.7037 | 234.2238 | 292.6013 | 210.3424 | 85.3222 | 1.4432 | 36.5141 | 52.6988 | 25.3001 | 2.2859 | 36.4022 | 83.2104 | 15.9249 | 21.5308 | 5.8821 | 135.9092 | 276.1729 | 212.693 | 271.0705 | 188.8116
2012-10-30 00:23 | 26.0875 | 297.766 | 234.2246 | 292.5713 | 210.1739 | 85.3234 | 1.4395 | 36.4185 | 52.426 | 25.2987 | 2.2803 | 36.3121 | 82.8013 | 15.9245 | 21.5275 | 5.8822 | 135.2273 | 276.2385 | 212.6971 | 271.0438 | 188.6464
2012-10-30 00:23 | 26.0815 | 297.7062 | 234.1557 | 292.5422 | 210.1049 | 85.3182 | 1.4394 | 36.4136 | 52.4126 | 25.2983 | 2.2799 | 36.3069 | 82.7774 | 15.9245 | 21.5055 | 5.8818 | 135.19 | 276.2007 | 212.6502 | 271.0367 | 188.5994
2012-10-30 00:23 | 26.0972 | 297.7488 | 234.1872 | 292.5733 | 209.459 | 85.3187 | 1.439 | 36.4037 | 52.3845 | 25.2982 | 2.2793 | 36.2969 | 82.7323 | 15.9244 | 21.5413 | 5.8819 | 135.1168 | 276.2075 | 212.6459 | 271.032 | 187.9177
2012-10-30 00:23 | 26.0942 | 297.7189 | 234.0152 | 292.5818 | 209.4127 | 85.3187 | 1.4388 | 36.3997 | 52.3729 | 25.2982 | 2.2791 | 36.2928 | 82.7136 | 15.9244 | 21.5503 | 5.8819 | 135.0865 | 276.1686 | 212.4649 | 271.0315 | 187.8624
2012-10-30 00:23 | 26.0905 | 297.7152 | 233.9722 | 292.5743 | 209.728 | 85.313 | 1.4389 | 36.4017 | 52.3785 | 25.2983 | 2.2792 | 36.2947 | 82.7222 | 15.9245 | 21.5025 | 5.8815 | 135.1006 | 276.2127 | 212.4697 | 271.0718 | 188.2255
2012-10-30 00:23 | 26.0957 | 297.7051 | 233.9696 | 292.5525 | 209.9182 | 85.3187 | 1.439 | 36.4043 | 52.3853 | 25.2986 | 2.2793 | 36.2973 | 82.7335 | 15.9246 | 21.5237 | 5.8819 | 135.1188 | 276.1814 | 212.4459 | 271.0288 | 188.3945
2012-10-30 00:23 | 26.0845 | 297.7246 | 233.9702 | 292.5259 | 209.9307 | 85.2933 | 1.4391 | 36.4057 | 52.3898 | 25.2984 | 2.2794 | 36.2987 | 82.74 | 15.9245 | 21.5406 | 5.8801 | 135.1298 | 276.184 | 212.4296 | 270.9853 | 188.3901
2012-10-30 00:23 | 26.0837 | 297.7123 | 233.9537 | 292.5405 | 210.0913 | 85.3078 | 1.4392 | 36.41 | 52.4021 | 25.2984 | 2.2797 | 36.303 | 82.76 | 15.9245 | 21.5639 | 5.8811 | 135.1621 | 276.1484 | 212.3898 | 270.9766 | 188.5274
2012-10-30 00:23 | 26.083 | 297.6655 | 234.0355 | 292.5668 | 210.0434 | 85.3008 | 1.4395 | 36.4165 | 52.4207 | 25.2984 | 2.2801 | 36.3092 | 82.7884 | 15.9244 | 21.519 | 5.8806 | 135.2092 | 276.1465 | 212.5165 | 271.0478 | 188.5244
2012-10-30 00:23 | 26.0762 | 297.6664 | 233.9933 | 292.5677 | 209.7058 | 85.2916 | 1.4402 | 36.4351 | 52.4737 | 25.2987 | 2.2813 | 36.3277 | 82.8728 | 15.9244 | 21.5082 | 5.88 | 135.3465 | 276.1582 | 212.4851 | 271.0595 | 188.1976
2012-10-30 00:23 | 26.0755 | 297.6925 | 234.0005 | 292.5785 | 209.6381 | 85.313 | 1.4426 | 36.4978 | 52.6523 | 25.2997 | 2.2852 | 36.3906 | 83.1585 | 15.9247 | 21.5798 | 5.8815 | 135.8108 | 276.1127 | 212.4207 | 270.9987 | 188.0583
2012-10-30 00:23 | 26.0755 | 297.6541 | 233.9415 | 292.5631 | 209.2837 | 85.3286 | 1.4432 | 36.5123 | 52.6944 | 25.2996 | 2.2861 | 36.4047 | 83.2236 | 15.9246 | 21.5196 | 5.8826 | 135.918 | 276.1345 | 212.4219 | 271.0435 | 187.7641
2012-10-30 00:23 | 26.0717 | 297.6657 | 233.8395 | 292.5555 | 209.0594 | 85.3153 | 1.4437 | 36.5255 | 52.7317 | 25.3 | 2.2869 | 36.4183 | 83.2856 | 15.9246 | 21.5238 | 5.8816 | 136.0173 | 276.1419 | 212.3157 | 271.0317 | 187.5356
2012-10-30 00:23 | 26.065 | 297.6744 | 233.7542 | 292.4718 | 209.5371 | 85.3153 | 1.4441 | 36.5352 | 52.7594 | 25.3001 | 2.2875 | 36.4277 | 83.3286 | 15.9246 | 21.5211 | 5.8816 | 136.088 | 276.1533 | 212.2331 | 270.9507 | 188.016
2012-10-30 00:23 | 26.0695 | 297.7173 | 233.798 | 292.5263 | 209.8684 | 85.343 | 1.4444 | 36.5425 | 52.7806 | 25.3001 | 2.288 | 36.4348 | 83.3611 | 15.9246 | 21.5296 | 5.8835 | 136.1417 | 276.1877 | 212.2684 | 270.9967 | 188.3388
2012-10-30 00:23 | 26.0725 | 297.6895 | 233.7892 | 292.5293 | 209.9029 | 85.3291 | 1.4445 | 36.5457 | 52.7889 | 25.3005 | 2.2882 | 36.4381 | 83.3761 | 15.9247 | 21.5647 | 5.8826 | 136.1649 | 276.1248 | 212.2245 | 270.9646 | 188.3382
2012-10-30 00:23 | 26.0725 | 297.7203 | 233.801 | 292.5293 | 209.7533 | 85.3291 | 1.4446 | 36.5482 | 52.7961 | 25.3005 | 2.2883 | 36.4407 | 83.3884 | 15.9246 | 21.5246 | 5.8826 | 136.1845 | 276.1957 | 212.2764 | 271.0047 | 188.2287
2012-10-30 00:23 | 26.0643 | 297.7236 | 233.8478 | 292.5481 | 209.9813 | 85.343 | 1.4446 | 36.5487 | 52.7978 | 25.3005 | 2.2884 | 36.4413 | 83.3911 | 15.9246 | 21.5525 | 5.8835 | 136.1889 | 276.1711 | 212.2953 | 270.9956 | 188.4288
2012-10-30 00:23 | 26.068 | 297.6889 | 233.8594 | 292.5017 | 210.0284 | 85.3303 | 1.4446 | 36.5496 | 52.8003 | 25.3005 | 2.2884 | 36.4421 | 83.3945 | 15.9246 | 21.5562 | 5.8827 | 136.1948 | 276.1327 | 212.3032 | 270.9455 | 188.4722
2012-10-30 00:24 | 26.071 | 297.6804 | 233.9331 | 292.524 | 210.2834 | 85.3239 | 1.4446 | 36.5495 | 52.7998 | 25.3006 | 2.2884 | 36.442 | 83.3938 | 15.9247 | 21.5432 | 5.8822 | 136.1936 | 276.1372 | 212.3899 | 270.9808 | 188.7402
2012-10-30 00:24 | 26.0628 | 297.6491 | 233.9328 | 292.5273 | 210.0665 | 85.3314 | 1.4445 | 36.5477 | 52.7948 | 25.3006 | 2.2883 | 36.4403 | 83.3857 | 15.9248 | 21.535 | 5.8827 | 136.1805 | 276.1141 | 212.3978 | 270.9923 | 188.5315
2012-10-30 00:24 | 26.071 | 297.7226 | 233.9763 | 292.5201 | 210.0865 | 85.343 | 1.4445 | 36.547 | 52.7921 | 25.3007 | 2.2882 | 36.4394 | 83.3821 | 15.9246 | 21.5873 | 5.8835 | 136.1743 | 276.1353 | 212.389 | 270.9328 | 188.4992
2012-10-30 00:24 | 26.0665 | 297.672 | 233.9875 | 292.5272 | 210.0111 | 85.3355 | 1.4445 | 36.5475 | 52.7931 | 25.301 | 2.2883 | 36.4399 | 83.3846 | 15.9246 | 21.5467 | 5.883 | 136.1777 | 276.1253 | 212.4408 | 270.9805 | 188.4644
2012-10-30 00:24 | 26.0643 | 297.666 | 234.0246 | 292.5365 | 209.9538 | 85.3424 | 1.4445 | 36.5468 | 52.7917 | 25.3007 | 2.2882 | 36.4392 | 83.3819 | 15.9245 | 21.5967 | 5.8835 | 136.1736 | 276.0693 | 212.4279 | 270.9398 | 188.3571
2012-10-30 00:24 | 26.0665 | 297.7143 | 233.9757 | 292.5195 | 209.9205 | 85.3262 | 1.4445 | 36.5466 | 52.791 | 25.3007 | 2.2882 | 36.439 | 83.3812 | 15.9244 | 21.6029 | 5.8824 | 136.1722 | 276.1114 | 212.3728 | 270.9166 | 188.3176
2012-10-30 00:24 | 26.0598 | 297.6961 | 233.9848 | 292.532 | 209.9532 | 85.3286 | 1.4445 | 36.5462 | 52.7903 | 25.3006 | 2.2882 | 36.4385 | 83.3797 | 15.9243 | 21.5721 | 5.8826 | 136.1699 | 276.124 | 212.4127 | 270.9599 | 188.3811
2012-10-30 00:24 | 26.068 | 297.6735 | 233.993 | 292.5364 | 210.1071 | 85.3459 | 1.4445 | 36.5459 | 52.7888 | 25.3009 | 2.2882 | 36.4385 | 83.3791 | 15.9244 | 21.5963 | 5.8837 | 136.1679 | 276.0772 | 212.3967 | 270.9401 | 188.5108
2012-10-30 00:24 | 26.062 | 297.7059 | 234.0577 | 292.5458 | 210.1366 | 85.3349 | 1.4444 | 36.5455 | 52.7875 | 25.3009 | 2.2882 | 36.438 | 83.3775 | 15.9243 | 21.5663 | 5.883 | 136.165 | 276.1396 | 212.4914 | 270.9795 | 188.5703
2012-10-30 00:24 | 26.059 | 297.6953 | 234.0704 | 292.5351 | 210.1572 | 85.3286 | 1.4444 | 36.5442 | 52.7842 | 25.3008 | 2.2881 | 36.4367 | 83.3718 | 15.9242 | 21.5633 | 5.8826 | 136.156 | 276.132 | 212.5071 | 270.9718 | 188.5939
2012-10-30 00:24 | 26.071 | 297.7188 | 234.0706 | 292.574 | 210.1731 | 85.3482 | 1.4443 | 36.5435 | 52.7814 | 25.3011 | 2.2881 | 36.4358 | 83.3675 | 15.9243 | 21.5873 | 5.8839 | 136.1489 | 276.1315 | 212.4833 | 270.9867 | 188.5858
2012-10-30 00:24 | 26.0658 | 297.6905 | 234.1125 | 292.5342 | 210.1404 | 85.3465 | 1.4443 | 36.5427 | 52.7796 | 25.3009 | 2.288 | 36.4351 | 83.3644 | 15.9243 | 21.5781 | 5.8838 | 136.144 | 276.1124 | 212.5344 | 270.9561 | 188.5623
2012-10-30 00:24 | 26.0785 | 297.7301 | 234.1135 | 292.5661 | 210.1058 | 85.3349 | 1.4443 | 36.5411 | 52.7746 | 25.301 | 2.2879 | 36.4336 | 83.3574 | 15.9243 | 21.5667 | 5.883 | 136.132 | 276.1634 | 212.5468 | 270.9994 | 188.5391
2012-10-30 00:24 | 26.0792 | 297.7616 | 234.1259 | 292.5514 | 210.075 | 85.3447 | 1.4442 | 36.5403 | 52.7729 | 25.3008 | 2.2879 | 36.4328 | 83.3548 | 15.9241 | 21.5714 | 5.8837 | 136.1277 | 276.1902 | 212.5545 | 270.98 | 188.5036
2012-10-30 00:24 | 26.0792 | 297.7423 | 234.0945 | 292.5861 | 210.1971 | 85.3517 | 1.4442 | 36.5388 | 52.7682 | 25.3009 | 2.2878 | 36.431 | 83.3466 | 15.924 | 21.5754 | 5.8841 | 136.1149 | 276.1669 | 212.5191 | 271.0107 | 188.6217
2012-10-30 00:24 | 26.0785 | 297.7301 | 234.0899 | 292.5815 | 210.1294 | 85.3563 | 1.4441 | 36.5381 | 52.7664 | 25.3008 | 2.2878 | 36.4307 | 83.345 | 15.9241 | 21.5386 | 5.8845 | 136.1115 | 276.1915 | 212.5513 | 271.0429 | 188.5908
2012-10-30 00:24 | 26.086 | 297.753 | 234.1092 | 292.589 | 210.1212 | 85.3511 | 1.4441 | 36.5362 | 52.7609 | 25.3008 | 2.2876 | 36.4288 | 83.3362 | 15.9241 | 21.5622 | 5.8841 | 136.0971 | 276.1908 | 212.547 | 271.0268 | 188.559
2012-10-30 00:24 | 26.0755 | 297.7386 | 234.1026 | 292.5708 | 210.1698 | 85.3413 | 1.444 | 36.5352 | 52.7582 | 25.3008 | 2.2876 | 36.4279 | 83.3331 | 15.9239 | 21.5436 | 5.8834 | 136.0913 | 276.195 | 212.559 | 271.0272 | 188.6262
2012-10-30 00:24 | 26.0852 | 297.7829 | 234.1123 | 292.5767 | 210.144 | 85.3459 | 1.4441 | 36.5355 | 52.7592 | 25.3006 | 2.2876 | 36.4281 | 83.3346 | 15.9239 | 21.5413 | 5.8837 | 136.0938 | 276.2416 | 212.571 | 271.0354 | 188.6027
2012-10-30 00:24 | 26.0822 | 297.753 | 234.1132 | 292.616 | 210.208 | 85.3627 | 1.444 | 36.5351 | 52.7576 | 25.3009 | 2.2876 | 36.4278 | 83.333 | 15.9239 | 21.5263 | 5.8849 | 136.0905 | 276.2267 | 212.5869 | 271.0897 | 188.6817
2012-10-30 00:24 | 26.08 | 297.7662 | 234.1307 | 292.5907 | 210.1388 | 85.3569 | 1.444 | 36.5341 | 52.7549 | 25.3008 | 2.2875 | 36.4265 | 83.3272 | 15.9239 | 21.512 | 5.8845 | 136.0821 | 276.2542 | 212.6187 | 271.0787 | 188.6268
2012-10-30 00:24 | 26.0815 | 297.7869 | 234.1204 | 292.6269 | 210.223 | 85.3523 | 1.444 | 36.5333 | 52.7528 | 25.3007 | 2.2875 | 36.4256 | 83.3238 | 15.9238 | 21.5496 | 5.8842 | 136.0766 | 276.2373 | 212.5708 | 271.0773 | 188.6734
2012-10-30 00:24 | 26.0777 | 297.7869 | 234.1166 | 292.5923 | 210.3137 | 85.3621 | 1.4439 | 36.5309 | 52.7453 | 25.301 | 2.2873 | 36.4236 | 83.313 | 15.924 | 21.4736 | 5.8849 | 136.0584 | 276.3133 | 212.643 | 271.1187 | 188.8401
2012-10-30 00:24 | 26.0919 | 297.7935 | 234.1544 | 292.6373 | 210.3122 | 85.3603 | 1.4438 | 36.5292 | 52.7406 | 25.3008 | 2.2873 | 36.4218 | 83.3059 | 15.9238 | 21.556 | 5.8847 | 136.0465 | 276.2375 | 212.5984 | 271.0813 | 188.7562
2012-10-30 00:25 | 26.0882 | 297.809 | 234.1782 | 292.6297 | 210.2376 | 85.3702 | 1.4437 | 36.5267 | 52.7337 | 25.3007 | 2.2871 | 36.4194 | 83.2941 | 15.9239 | 21.5523 | 5.8854 | 136.0278 | 276.2567 | 212.6259 | 271.0774 | 188.6853
2012-10-30 00:25 | 26.0867 | 297.7883 | 234.2042 | 292.6321 | 210.2361 | 85.3603 | 1.4437 | 36.5258 | 52.731 | 25.3008 | 2.287 | 36.4185 | 83.2905 | 15.9238 | 21.4987 | 5.8847 | 136.0215 | 276.2896 | 212.7055 | 271.1334 | 188.7374
2012-10-30 00:25 | 26.0867 | 297.7729 | 234.157 | 292.6051 | 210.2597 | 85.3655 | 1.4437 | 36.5259 | 52.731 | 25.3009 | 2.287 | 36.4184 | 83.2907 | 15.9238 | 21.5388 | 5.8851 | 136.0217 | 276.2341 | 212.6182 | 271.0663 | 188.7209
2012-10-30 00:25 | 26.0822 | 297.7914 | 234.1957 | 292.643 | 210.2316 | 85.3644 | 1.4437 | 36.5262 | 52.7318 | 25.3009 | 2.2871 | 36.4187 | 83.2928 | 15.9236 | 21.5423 | 5.885 | 136.0245 | 276.2491 | 212.6534 | 271.1007 | 188.6893
2012-10-30 00:25 | 26.0822 | 297.8145 | 234.2036 | 292.643 | 210.1568 | 85.3592 | 1.4436 | 36.5249 | 52.7281 | 25.3009 | 2.287 | 36.4175 | 83.2871 | 15.9236 | 21.5142 | 5.8847 | 136.0152 | 276.3003 | 212.6894 | 271.1288 | 188.6426
2012-10-30 00:25 | 26.0875 | 297.862 | 234.201 | 292.6329 | 210.0164 | 85.38 | 1.4436 | 36.5236 | 52.725 | 25.3006 | 2.287 | 36.4162 | 83.2823 | 15.9234 | 21.5035 | 5.8861 | 136.0073 | 276.3585 | 212.6975 | 271.1294 | 188.5129
2012-10-30 00:25 | 26.0934 | 297.8564 | 234.1441 | 292.6349 | 210.0656 | 85.3592 | 1.4435 | 36.5229 | 52.7224 | 25.3008 | 2.2869 | 36.4155 | 83.2783 | 15.9236 | 21.5094 | 5.8847 | 136.0007 | 276.347 | 212.6347 | 271.1255 | 188.5562
2012-10-30 00:25 | 26.0905 | 297.8843 | 234.1962 | 292.6551 | 210.2911 | 85.358 | 1.4435 | 36.5214 | 52.7183 | 25.3007 | 2.2868 | 36.414 | 83.2721 | 15.9234 | 21.5305 | 5.8846 | 135.9905 | 276.3538 | 212.6657 | 271.1246 | 188.7606
2012-10-30 00:25 | 26.0949 | 297.8272 | 234.1849 | 292.6864 | 210.2128 | 85.3592 | 1.4434 | 36.5185 | 52.7102 | 25.3007 | 2.2867 | 36.4112 | 83.2601 | 15.9233 | 21.5349 | 5.8847 | 135.9703 | 276.2923 | 212.65 | 271.1515 | 188.6779
2012-10-30 00:25 | 26.1047 | 297.8869 | 234.2457 | 292.7078 | 210.1556 | 85.3748 | 1.4419 | 36.4804 | 52.6006 | 25.3005 | 2.2843 | 36.3738 | 83.0894 | 15.9233 | 21.5528 | 5.8857 | 135.6901 | 276.3341 | 212.6929 | 271.155 | 188.6028
2012-10-30 00:25 | 26.0957 | 297.8242 | 234.1464 | 292.6795 | 210.2215 | 85.3569 | 1.4396 | 36.4213 | 52.4324 | 25.2995 | 2.2806 | 36.3139 | 82.8174 | 15.923 | 21.5277 | 5.8845 | 135.2499 | 276.2965 | 212.6187 | 271.1518 | 188.6938
2012-10-30 00:25 | 26.0912 | 297.8197 | 234.1812 | 292.6173 | 210.3469 | 85.3569 | 1.4391 | 36.4076 | 52.3937 | 25.299 | 2.2798 | 36.3002 | 82.7557 | 15.9228 | 21.4751 | 5.8845 | 135.1494 | 276.3446 | 212.7061 | 271.1422 | 188.8718
2012-10-30 00:25 | 26.0919 | 297.8204 | 234.1976 | 292.6757 | 209.6388 | 85.3569 | 1.4387 | 36.3983 | 52.3673 | 25.2989 | 2.2792 | 36.2913 | 82.7139 | 15.9231 | 21.5641 | 5.8845 | 135.0812 | 276.2563 | 212.6335 | 271.1116 | 188.0747
2012-10-30 00:25 | 26.0912 | 297.812 | 234.1379 | 292.675 | 209.0829 | 85.3447 | 1.4386 | 36.3959 | 52.3608 | 25.2987 | 2.279 | 36.2888 | 82.7037 | 15.9228 | 21.5312 | 5.8837 | 135.0645 | 276.2808 | 212.6067 | 271.1438 | 187.5517
2012-10-30 00:25 | 26.0875 | 297.8006 | 233.9496 | 292.6983 | 208.1381 | 85.3569 | 1.4387 | 36.3965 | 52.3629 | 25.2986 | 2.2791 | 36.2894 | 82.7058 | 15.9229 | 21.5195 | 5.8845 | 135.0686 | 276.2811 | 212.4301 | 271.1788 | 186.6186
2012-10-30 00:25 | 26.0942 | 297.8188 | 233.7638 | 292.6627 | 208.8614 | 85.3476 | 1.4388 | 36.3996 | 52.3717 | 25.2986 | 2.2793 | 36.2923 | 82.7193 | 15.9229 | 21.5262 | 5.8839 | 135.091 | 276.2926 | 212.2376 | 271.1365 | 187.3352
2012-10-30 00:25 | 26.0867 | 297.7998 | 233.6463 | 292.6667 | 209.3619 | 85.3424 | 1.4392 | 36.4097 | 52.4 | 25.2989 | 2.2799 | 36.3024 | 82.7654 | 15.9229 | 21.5027 | 5.8835 | 135.1654 | 276.2971 | 212.1436 | 271.164 | 187.8592
2012-10-30 00:25 | 26.0807 | 297.8015 | 233.6874 | 292.6338 | 209.3047 | 85.3286 | 1.4409 | 36.4548 | 52.5293 | 25.2992 | 2.2826 | 36.3453 | 82.9605 | 15.923 | 21.5529 | 5.8826 | 135.4898 | 276.2486 | 212.1345 | 271.0809 | 187.7518
2012-10-30 00:25 | 26.08 | 297.8123 | 233.6671 | 292.6485 | 209.8868 | 85.3413 | 1.4428 | 36.5028 | 52.6657 | 25.3002 | 2.2857 | 36.3952 | 83.1884 | 15.923 | 21.5281 | 5.8834 | 135.8541 | 276.2842 | 212.139 | 271.1204 | 188.3587
2012-10-30 00:25 | 26.0747 | 297.7532 | 233.7325 | 292.647 | 209.5035 | 85.3725 | 1.4433 | 36.5161 | 52.7047 | 25.2999 | 2.2866 | 36.409 | 83.2514 | 15.923 | 21.4987 | 5.8856 | 135.9562 | 276.2545 | 212.2338 | 271.1483 | 188.0048
2012-10-30 00:25 | 26.0695 | 297.7711 | 233.6998 | 292.6533 | 209.2265 | 85.3511 | 1.4438 | 36.5279 | 52.7381 | 25.3002 | 2.2873 | 36.4205 | 83.3045 | 15.923 | 21.5055 | 5.8841 | 136.0427 | 276.2656 | 212.1943 | 271.1478 | 187.721
2012-10-30 00:25 | 26.0673 | 297.7919 | 233.6701 | 292.6319 | 209.0196 | 85.369 | 1.4441 | 36.5359 | 52.7607 | 25.3004 | 2.2878 | 36.4283 | 83.3403 | 15.923 | 21.5194 | 5.8853 | 136.101 | 276.2725 | 212.1507 | 271.1125 | 187.5002
2012-10-30 00:26 | 26.071 | 297.8033 | 233.6345 | 292.5894 | 209.5549 | 85.3627 | 1.4445 | 36.5464 | 52.7908 | 25.3007 | 2.2885 | 36.4389 | 83.3892 | 15.9229 | 21.5151 | 5.8849 | 136.18 | 276.2882 | 212.1194 | 271.0743 | 188.0398
2012-10-30 00:26 | 26.0658 | 297.7674 | 233.7079 | 292.5881 | 209.6284 | 85.3661 | 1.4446 | 36.5496 | 52.7996 | 25.3009 | 2.2887 | 36.4419 | 83.4028 | 15.9229 | 21.4898 | 5.8851 | 136.2024 | 276.2776 | 212.2181 | 271.0983 | 188.1386
2012-10-30 00:26 | 26.0702 | 297.7833 | 233.7044 | 292.6156 | 209.5462 | 85.3754 | 1.4446 | 36.5494 | 52.8 | 25.3004 | 2.2887 | 36.4418 | 83.403 | 15.9227 | 21.4661 | 5.8858 | 136.203 | 276.3172 | 212.2383 | 271.1495 | 188.0801
2012-10-30 00:26 | 26.0695 | 297.821 | 233.688 | 292.5725 | 209.6479 | 85.3736 | 1.4446 | 36.5497 | 52.7999 | 25.3008 | 2.2887 | 36.4421 | 83.4045 | 15.9227 | 21.4453 | 5.8857 | 136.2043 | 276.3757 | 212.2427 | 271.1272 | 188.2026
2012-10-30 00:26 | 26.0695 | 297.7787 | 233.7155 | 292.5571 | 209.6715 | 85.3655 | 1.4446 | 36.5487 | 52.7972 | 25.3007 | 2.2886 | 36.4411 | 83.4005 | 15.9226 | 21.5256 | 5.8851 | 136.1977 | 276.2531 | 212.1899 | 271.0315 | 188.1459
2012-10-30 00:26 | 26.0673 | 297.7727 | 233.7133 | 292.6127 | 209.6614 | 85.369 | 1.4446 | 36.5483 | 52.7964 | 25.3005 | 2.2886 | 36.4406 | 83.3988 | 15.9225 | 21.4833 | 5.8853 | 136.1952 | 276.2894 | 212.23 | 271.1294 | 188.1781
2012-10-30 00:26 | 26.071 | 297.7726 | 233.7445 | 292.5625 | 209.8463 | 85.3661 | 1.4445 | 36.5477 | 52.7942 | 25.3007 | 2.2886 | 36.44 | 83.3949 | 15.9227 | 21.5472 | 5.8851 | 136.1891 | 276.2254 | 212.1973 | 271.0153 | 188.2991
2012-10-30 00:26 | 26.0732 | 297.8093 | 233.8135 | 292.5493 | 209.8721 | 85.3679 | 1.4445 | 36.5467 | 52.791 | 25.3009 | 2.2885 | 36.4392 | 83.3916 | 15.9226 | 21.4932 | 5.8853 | 136.1825 | 276.3161 | 212.3203 | 271.0561 | 188.3789
2012-10-30 00:26 | 26.077 | 297.7517 | 233.7741 | 292.5685 | 209.9035 | 85.3702 | 1.4445 | 36.5461 | 52.7894 | 25.3008 | 2.2885 | 36.4385 | 83.3886 | 15.9226 | 21.5211 | 5.8854 | 136.178 | 276.2306 | 212.253 | 271.0474 | 188.3824
2012-10-30 00:26 | 26.068 | 297.8233 | 233.8397 | 292.5787 | 209.8787 | 85.3713 | 1.4444 | 36.5442 | 52.7835 | 25.301 | 2.2884 | 36.4366 | 83.3798 | 15.9226 | 21.508 | 5.8855 | 136.1633 | 276.3153 | 212.3317 | 271.0707 | 188.3707
2012-10-30 00:26 | 26.0732 | 297.8093 | 233.8724 | 292.5608 | 209.8839 | 85.3748 | 1.4444 | 36.5445 | 52.7843 | 25.301 | 2.2884 | 36.437 | 83.3817 | 15.9226 | 21.5012 | 5.8857 | 136.166 | 276.3081 | 212.3712 | 271.0596 | 188.3827
2012-10-30 00:26 | 26.0687 | 297.824 | 233.8326 | 292.5833 | 209.8597 | 85.3811 | 1.4444 | 36.5436 | 52.7819 | 25.301 | 2.2883 | 36.4358 | 83.3762 | 15.9226 | 21.5328 | 5.8862 | 136.1581 | 276.2912 | 212.2998 | 271.0505 | 188.3269
2012-10-30 00:26 | 26.077 | 297.8016 | 233.888 | 292.5839 | 209.9389 | 85.3679 | 1.4443 | 36.543 | 52.7795 | 25.3013 | 2.2883 | 36.4355 | 83.3751 | 15.9226 | 21.5411 | 5.8853 | 136.1546 | 276.2605 | 212.3469 | 271.0428 | 188.3978
2012-10-30 00:26 | 26.074 | 297.7909 | 233.9047 | 292.5886 | 209.9438 | 85.3921 | 1.4443 | 36.5419 | 52.7768 | 25.3011 | 2.2882 | 36.4343 | 83.3692 | 15.9226 | 21.5181 | 5.8869 | 136.146 | 276.2728 | 212.3866 | 271.0705 | 188.4257
2012-10-30 00:26 | 26.0725 | 297.824 | 233.9346 | 292.5909 | 210.1746 | 85.3627 | 1.4443 | 36.5417 | 52.776 | 25.3011 | 2.2882 | 36.4342 | 83.3696 | 15.9225 | 21.5326 | 5.8849 | 136.1455 | 276.2914 | 212.402 | 271.0583 | 188.642
2012-10-30 00:26 | 26.0665 | 297.8103 | 233.9718 | 292.6157 | 210.1647 | 85.3817 | 1.4443 | 36.5414 | 52.7753 | 25.3011 | 2.2882 | 36.4338 | 83.3668 | 15.9227 | 21.5025 | 5.8862 | 136.1421 | 276.3078 | 212.4693 | 271.1132 | 188.6622
2012-10-30 00:26 | 26.0687 | 297.8164 | 234.0054 | 292.5948 | 210.0251 | 85.3916 | 1.4442 | 36.5391 | 52.7685 | 25.3012 | 2.288 | 36.4315 | 83.3569 | 15.9226 | 21.5529 | 5.8869 | 136.1254 | 276.2635 | 212.4525 | 271.0419 | 188.4722
2012-10-30 00:26 | 26.0665 | 297.8142 | 234.0072 | 292.6234 | 210.1686 | 85.365 | 1.4441 | 36.5378 | 52.7651 | 25.3011 | 2.288 | 36.4302 | 83.3506 | 15.9226 | 21.4905 | 5.8851 | 136.1157 | 276.3237 | 212.5167 | 271.1329 | 188.6781
2012-10-30 00:26 | 26.0702 | 297.8601 | 234.0305 | 292.604 | 210.0818 | 85.38 | 1.4441 | 36.5369 | 52.7626 | 25.301 | 2.2879 | 36.4295 | 83.3484 | 15.9224 | 21.5383 | 5.8861 | 136.1109 | 276.3218 | 212.4922 | 271.0657 | 188.5435
2012-10-30 00:26 | 26.0687 | 297.8164 | 234.0644 | 292.5948 | 210.0724 | 85.3858 | 1.4441 | 36.5376 | 52.7644 | 25.3011 | 2.288 | 36.43 | 83.351 | 15.9223 | 21.5409 | 5.8865 | 136.1154 | 276.2755 | 212.5235 | 271.0539 | 188.5315
2012-10-30 00:26 | 26.071 | 297.8455 | 234.102 | 292.6548 | 210.1141 | 85.3846 | 1.4441 | 36.5374 | 52.7641 | 25.3009 | 2.288 | 36.4299 | 83.3506 | 15.9223 | 21.5512 | 5.8864 | 136.1147 | 276.2943 | 212.5508 | 271.1036 | 188.5629
2012-10-30 00:26 | 26.0755 | 297.85 | 234.0947 | 292.6132 | 209.9571 | 85.3852 | 1.444 | 36.5355 | 52.7582 | 25.3011 | 2.2879 | 36.428 | 83.342 | 15.9223 | 21.5758 | 5.8865 | 136.1002 | 276.2742 | 212.5189 | 271.0374 | 188.3813
2012-10-30 00:26 | 26.08 | 297.8738 | 234.115 | 292.6523 | 209.9734 | 85.3944 | 1.444 | 36.5338 | 52.7537 | 25.301 | 2.2877 | 36.4261 | 83.3338 | 15.9222 | 21.5843 | 5.8871 | 136.0874 | 276.2895 | 212.5307 | 271.068 | 188.3891
2012-10-30 00:27 | 26.074 | 297.8332 | 234.109 | 292.654 | 209.9753 | 85.3858 | 1.4439 | 36.5325 | 52.7502 | 25.3009 | 2.2877 | 36.4251 | 83.3296 | 15.9221 | 21.5622 | 5.8865 | 136.0797 | 276.271 | 212.5468 | 271.0918 | 188.4131
2012-10-30 00:27 | 26.0673 | 297.8495 | 234.1023 | 292.6473 | 210.1261 | 85.3933 | 1.4439 | 36.5317 | 52.7479 | 25.3009 | 2.2877 | 36.4243 | 83.3265 | 15.922 | 21.5836 | 5.887 | 136.0744 | 276.2659 | 212.5187 | 271.0637 | 188.5425
2012-10-30 00:27 | 26.0777 | 297.8292 | 234.093 | 292.6115 | 210.2035 | 85.3887 | 1.4439 | 36.5307 | 52.7453 | 25.3006 | 2.2876 | 36.4232 | 83.3213 | 15.9221 | 21.5338 | 5.8867 | 136.0666 | 276.2954 | 212.5592 | 271.0777 | 188.6697
2012-10-30 00:27 | 26.0695 | 297.8671 | 234.1202 | 292.611 | 210.148 | 85.3927 | 1.4438 | 36.5294 | 52.7412 | 25.3008 | 2.2875 | 36.4218 | 83.3161 | 15.9219 | 21.5457 | 5.887 | 136.0573 | 276.3214 | 212.5745 | 271.0653 | 188.6023
2012-10-30 00:27 | 26.068 | 297.8464 | 234.1187 | 292.6057 | 210.0087 | 85.3835 | 1.4438 | 36.5289 | 52.7395 | 25.301 | 2.2875 | 36.4214 | 83.3136 | 15.922 | 21.5683 | 5.8863 | 136.0531 | 276.2781 | 212.5504 | 271.0374 | 188.4404
2012-10-30 00:27 | 26.0687 | 297.8202 | 234.1154 | 292.6487 | 210.1275 | 85.4002 | 1.4437 | 36.5278 | 52.7367 | 25.3008 | 2.2875 | 36.4205 | 83.3101 | 15.9218 | 21.581 | 5.8875 | 136.0468 | 276.2392 | 212.5344 | 271.0677 | 188.5465
2012-10-30 00:27 | 26.0725 | 297.8663 | 234.1035 | 292.6294 | 210.1628 | 85.3875 | 1.4437 | 36.5266 | 52.7331 | 25.3008 | 2.2874 | 36.4191 | 83.3035 | 15.9219 | 21.5406 | 5.8866 | 136.0367 | 276.3257 | 212.5629 | 271.0888 | 188.6222
2012-10-30 00:27 | 26.074 | 297.8716 | 234.1482 | 292.6655 | 210.2273 | 85.3939 | 1.4436 | 36.5251 | 52.7284 | 25.301 | 2.2873 | 36.4177 | 83.2969 | 15.9219 | 21.5783 | 5.8871 | 136.0254 | 276.2933 | 212.5699 | 271.0872 | 188.649
2012-10-30 00:27 | 26.077 | 297.8784 | 234.163 | 292.6647 | 210.1122 | 85.3921 | 1.4436 | 36.5238 | 52.7249 | 25.3009 | 2.2872 | 36.4161 | 83.2898 | 15.9219 | 21.5451 | 5.8869 | 136.0147 | 276.3333 | 212.6179 | 271.1196 | 188.5671
2012-10-30 00:27 | 26.0702 | 297.8524 | 234.1641 | 292.6463 | 209.9754 | 85.4066 | 1.4436 | 36.5232 | 52.7233 | 25.3009 | 2.2872 | 36.4157 | 83.2884 | 15.9218 | 21.5705 | 5.8879 | 136.0117 | 276.2819 | 212.5936 | 271.0758 | 188.4049
2012-10-30 00:27 | 26.074 | 297.8485 | 234.1836 | 292.6463 | 210.0225 | 85.3968 | 1.4435 | 36.5226 | 52.7216 | 25.3008 | 2.2871 | 36.4152 | 83.2863 | 15.9218 | 21.5542 | 5.8873 | 136.0079 | 276.2943 | 212.6294 | 271.0921 | 188.4683
2012-10-30 00:27 | 26.0658 | 297.8634 | 234.1558 | 292.5765 | 210.1679 | 85.3956 | 1.4435 | 36.5206 | 52.7159 | 25.3009 | 2.287 | 36.4133 | 83.2778 | 15.9217 | 21.5299 | 5.8872 | 135.9937 | 276.3335 | 212.6259 | 271.0466 | 188.638
2012-10-30 00:27 | 26.0755 | 297.8731 | 234.193 | 292.6632 | 210.1264 | 85.4031 | 1.4434 | 36.5194 | 52.7123 | 25.3008 | 2.2869 | 36.412 | 83.2721 | 15.9217 | 21.5396 | 5.8877 | 135.9844 | 276.3335 | 212.6534 | 271.1236 | 188.5868
2012-10-30 00:27 | 26.0732 | 297.8554 | 234.1592 | 292.657 | 210.1556 | 85.4112 | 1.443 | 36.5099 | 52.6844 | 25.301 | 2.2865 | 36.4048 | 83.2394 | 15.9216 | 21.5173 | 5.8882 | 135.9238 | 276.3381 | 212.6419 | 271.1397 | 188.6383
2012-10-30 00:27 | 26.0732 | 297.867 | 234.1985 | 292.6647 | 210.2895 | 85.3921 | 1.4394 | 36.416 | 52.4175 | 25.2993 | 2.2805 | 36.3085 | 82.8011 | 15.9214 | 21.5173 | 5.8869 | 135.2186 | 276.3497 | 212.6812 | 271.1474 | 188.7722
2012-10-30 00:27 | 26.0807 | 297.813 | 234.1432 | 292.6915 | 210.3364 | 85.3863 | 1.4391 | 36.409 | 52.397 | 25.2995 | 2.2801 | 36.3016 | 82.7696 | 15.9214 | 21.5609 | 5.8865 | 135.1666 | 276.2521 | 212.5823 | 271.1306 | 188.7755
2012-10-30 00:27 | 26.0725 | 297.8163 | 234.2057 | 292.6756 | 210.0841 | 85.3991 | 1.4388 | 36.4007 | 52.3744 | 25.2989 | 2.2796 | 36.2934 | 82.7327 | 15.9213 | 21.5125 | 5.8874 | 135.1071 | 276.3038 | 212.6932 | 271.1631 | 188.5716
2012-10-30 00:27 | 26.0717 | 297.8194 | 234.1656 | 292.6517 | 209.2327 | 85.3875 | 1.4388 | 36.4002 | 52.3722 | 25.2992 | 2.2795 | 36.293 | 82.731 | 15.9213 | 21.5318 | 5.8866 | 135.1033 | 276.2876 | 212.6338 | 271.1199 | 187.7009
2012-10-30 00:27 | 26.068 | 297.831 | 234.0205 | 292.6557 | 208.6383 | 85.3777 | 1.4388 | 36.3995 | 52.3703 | 25.2991 | 2.2795 | 36.2924 | 82.7284 | 15.9212 | 21.5562 | 5.8859 | 135.0987 | 276.2748 | 212.4643 | 271.0995 | 187.0821
2012-10-30 00:27 | 26.0665 | 297.8103 | 233.85 | 292.6426 | 208.7786 | 85.3748 | 1.4389 | 36.4021 | 52.3781 | 25.299 | 2.2797 | 36.2951 | 82.7405 | 15.9213 | 21.5427 | 5.8857 | 135.1186 | 276.2676 | 212.3073 | 271.0999 | 187.2359
2012-10-30 00:27 | 26.0702 | 297.8025 | 233.7909 | 292.6194 | 208.9595 | 85.3748 | 1.439 | 36.4045 | 52.3849 | 25.2991 | 2.2798 | 36.2973 | 82.7507 | 15.9212 | 21.5464 | 5.8857 | 135.1356 | 276.2561 | 212.2445 | 271.073 | 187.4131
2012-10-30 00:27 | 26.0605 | 297.8082 | 233.7144 | 292.6174 | 208.6702 | 85.3725 | 1.4391 | 36.4082 | 52.395 | 25.2993 | 2.28 | 36.3011 | 82.7679 | 15.9213 | 21.5166 | 5.8856 | 135.1629 | 276.2916 | 212.1978 | 271.1008 | 187.1536
2012-10-30 00:27 | 26.0658 | 297.775 | 233.6372 | 292.6419 | 208.7976 | 85.3661 | 1.4392 | 36.4112 | 52.404 | 25.2991 | 2.2802 | 36.3043 | 82.7824 | 15.9212 | 21.558 | 5.8851 | 135.1865 | 276.217 | 212.0792 | 271.0839 | 187.2396
2012-10-30 00:27 | 26.062 | 297.7751 | 233.5548 | 292.6343 | 209.7034 | 85.3748 | 1.4401 | 36.4326 | 52.4655 | 25.2993 | 2.2815 | 36.3237 | 82.871 | 15.9213 | 21.5542 | 5.8857 | 135.3364 | 276.2209 | 212.0006 | 271.0801 | 188.1492
2012-10-30 00:27 | 26.0635 | 297.8073 | 233.5445 | 292.6435 | 209.5001 | 85.3771 | 1.4426 | 36.4976 | 52.6505 | 25.3003 | 2.2856 | 36.3899 | 83.1718 | 15.9215 | 21.5397 | 5.8859 | 135.8223 | 276.2676 | 212.0048 | 271.1038 | 187.9604
2012-10-30 00:27 | 26.0598 | 297.8267 | 233.5212 | 292.6321 | 209.2917 | 85.3794 | 1.4431 | 36.5119 | 52.6917 | 25.3004 | 2.2865 | 36.4042 | 83.2377 | 15.9215 | 21.5239 | 5.8861 | 135.9294 | 276.3028 | 211.9973 | 271.1082 | 187.7678
2012-10-30 00:27 | 26.0605 | 297.8466 | 233.5533 | 292.6136 | 209.8712 | 85.3748 | 1.4437 | 36.5258 | 52.7314 | 25.3006 | 2.2874 | 36.4182 | 83.3024 | 15.9213 | 21.5608 | 5.8857 | 136.0337 | 276.2858 | 211.9925 | 271.0528 | 188.3104
2012-10-30 00:28 | 26.0538 | 297.836 | 233.6409 | 292.5915 | 209.8251 | 85.3794 | 1.444 | 36.535 | 52.7581 | 25.3006 | 2.288 | 36.4274 | 83.3444 | 15.9213 | 21.5581 | 5.8861 | 136.1025 | 276.2779 | 212.0828 | 271.0334 | 188.267
2012-10-30 00:28 | 26.0553 | 297.8375 | 233.6935 | 292.6161 | 210.1023 | 85.3863 | 1.4443 | 36.5413 | 52.775 | 25.3011 | 2.2883 | 36.4337 | 83.3729 | 15.9214 | 21.5957 | 5.8865 | 136.1479 | 276.2418 | 212.0978 | 271.0204 | 188.5066
2012-10-30 00:28 | 26.0455 | 297.8469 | 233.7505 | 292.5678 | 210.1319 | 85.3754 | 1.4444 | 36.544 | 52.7828 | 25.3012 | 2.2885 | 36.4363 | 83.3841 | 15.9215 | 21.5779 | 5.8858 | 136.1669 | 276.269 | 212.1726 | 270.9899 | 188.554
2012-10-30 00:28 | 26.0493 | 297.8315 | 233.821 | 292.6101 | 210.2656 | 85.3863 | 1.4445 | 36.5462 | 52.7892 | 25.3011 | 2.2887 | 36.4386 | 83.3956 | 15.9214 | 21.5937 | 5.8865 | 136.1848 | 276.2378 | 212.2273 | 271.0164 | 188.6719
2012-10-30 00:28 | 26.0478 | 297.8454 | 233.8903 | 292.5739 | 210.2484 | 85.3863 | 1.4445 | 36.5482 | 52.795 | 25.3011 | 2.2888 | 36.4405 | 83.4047 | 15.9213 | 21.5721 | 5.8865 | 136.1997 | 276.2733 | 212.3182 | 271.0018 | 188.6763
2012-10-30 00:28 | 26.0493 | 297.8661 | 233.9625 | 292.5562 | 210.1317 | 85.384 | 1.4445 | 36.5483 | 52.7958 | 25.3008 | 2.2888 | 36.4407 | 83.4058 | 15.9212 | 21.5857 | 5.8864 | 136.2016 | 276.2804 | 212.3768 | 270.9705 | 188.546
2012-10-30 00:28 | 26.0448 | 297.8962 | 233.9855 | 292.544 | 209.9422 | 85.406 | 1.4445 | 36.5484 | 52.7957 | 25.301 | 2.2888 | 36.4407 | 83.406 | 15.9212 | 21.6052 | 5.8879 | 136.2017 | 276.291 | 212.3803 | 270.9388 | 188.337
2012-10-30 00:28 | 26.0433 | 297.8601 | 234.0233 | 292.5579 | 209.9328 | 85.3956 | 1.4446 | 36.5491 | 52.7973 | 25.3013 | 2.2889 | 36.4415 | 83.4097 | 15.9212 | 21.5476 | 5.8872 | 136.207 | 276.3125 | 212.4757 | 271.0103 | 188.3852
2012-10-30 00:28 | 26.0426 | 297.8901 | 234.0226 | 292.5841 | 209.9872 | 85.402 | 1.4445 | 36.547 | 52.7904 | 25.3016 | 2.2887 | 36.4393 | 83.3987 | 15.9213 | 21.5348 | 5.8876 | 136.1892 | 276.3553 | 212.4878 | 271.0493 | 188.4524
2012-10-30 00:28 | 26.0396 | 297.8295 | 233.9685 | 292.5542 | 210.0945 | 85.4072 | 1.4445 | 36.5467 | 52.7902 | 25.3012 | 2.2887 | 36.439 | 83.398 | 15.9212 | 21.5037 | 5.888 | 136.1883 | 276.3258 | 212.4648 | 271.0505 | 188.5908
2012-10-30 00:28 | 26.0411 | 297.9002 | 233.9857 | 292.6096 | 210.1669 | 85.4037 | 1.4444 | 36.5459 | 52.7883 | 25.3011 | 2.2886 | 36.4381 | 83.3938 | 15.9213 | 21.5935 | 5.8877 | 136.1822 | 276.3067 | 212.3922 | 271.0161 | 188.5734
2012-10-30 00:28 | 26.0388 | 297.894 | 234.0659 | 292.5611 | 210.3299 | 85.3869 | 1.4444 | 36.5452 | 52.7859 | 25.3014 | 2.2886 | 36.4376 | 83.392 | 15.9212 | 21.5391 | 5.8866 | 136.1779 | 276.3549 | 212.5268 | 271.022 | 188.7908
2012-10-30 00:28 | 26.0358 | 297.868 | 234.0433 | 292.6004 | 210.2639 | 85.4043 | 1.4444 | 36.546 | 52.7878 | 25.3014 | 2.2887 | 36.4382 | 83.3952 | 15.9211 | 21.5361 | 5.8878 | 136.183 | 276.3319 | 212.5072 | 271.0643 | 188.7278
2012-10-30 00:28 | 26.0343 | 297.8742 | 234.0496 | 292.5912 | 210.1522 | 85.4037 | 1.4444 | 36.5448 | 52.785 | 25.3012 | 2.2886 | 36.4371 | 83.3898 | 15.9212 | 21.5265 | 5.8877 | 136.1748 | 276.3477 | 212.5231 | 271.0647 | 188.6257
2012-10-30 00:28 | 26.0396 | 297.9294 | 234.0942 | 292.6465 | 210.2599 | 85.4106 | 1.4444 | 36.545 | 52.7854 | 25.3013 | 2.2886 | 36.4374 | 83.3911 | 15.9211 | 21.5278 | 5.8882 | 136.1764 | 276.4016 | 212.5664 | 271.1187 | 188.7321
2012-10-30 00:28 | 26.0336 | 297.8888 | 234.1511 | 292.5674 | 210.2302 | 85.3939 | 1.4443 | 36.544 | 52.7818 | 25.3016 | 2.2886 | 36.4365 | 83.3872 | 15.9212 | 21.5579 | 5.8871 | 136.1691 | 276.3309 | 212.5932 | 271.0095 | 188.6723
2012-10-30 00:28 | 26.0291 | 297.8766 | 234.123 | 292.6014 | 210.2257 | 85.4193 | 1.4443 | 36.5434 | 52.78 | 25.3016 | 2.2885 | 36.4358 | 83.3839 | 15.9211 | 21.5133 | 5.8888 | 136.1639 | 276.3633 | 212.6097 | 271.0881 | 188.7124
2012-10-30 00:28 | 26.0209 | 297.8761 | 234.1501 | 292.6163 | 210.2372 | 85.406 | 1.4443 | 36.542 | 52.7757 | 25.3017 | 2.2884 | 36.4342 | 83.3767 | 15.9212 | 21.5533 | 5.8879 | 136.1524 | 276.3228 | 212.5968 | 271.063 | 188.6839
2012-10-30 00:28 | 26.0253 | 297.9036 | 234.131 | 292.6207 | 210.155 | 85.4129 | 1.4443 | 36.5426 | 52.7779 | 25.3016 | 2.2885 | 36.4351 | 83.381 | 15.9211 | 21.5536 | 5.8884 | 136.1589 | 276.35 | 212.5774 | 271.0671 | 188.6014
2012-10-30 00:28 | 26.0306 | 297.8781 | 234.1638 | 292.6067 | 210.1052 | 85.4152 | 1.4442 | 36.541 | 52.773 | 25.3017 | 2.2884 | 36.4333 | 83.3733 | 15.9209 | 21.5549 | 5.8885 | 136.1463 | 276.3232 | 212.6089 | 271.0518 | 188.5503
2012-10-30 00:28 | 26.0231 | 297.909 | 234.1248 | 292.6377 | 210.1252 | 85.4245 | 1.4442 | 36.5407 | 52.7724 | 25.3015 | 2.2884 | 36.4331 | 83.3725 | 15.9209 | 21.5274 | 5.8892 | 136.145 | 276.3816 | 212.5974 | 271.1103 | 188.5978
2012-10-30 00:28 | 26.0283 | 297.8682 | 234.1693 | 292.6352 | 210.4297 | 85.4141 | 1.4442 | 36.5398 | 52.7699 | 25.3015 | 2.2883 | 36.4322 | 83.3687 | 15.9209 | 21.5526 | 5.8884 | 136.1386 | 276.3156 | 212.6167 | 271.0826 | 188.8771
2012-10-30 00:28 | 26.0313 | 297.8904 | 234.1959 | 292.6806 | 210.417 | 85.4355 | 1.4441 | 36.5381 | 52.7652 | 25.3014 | 2.2882 | 36.4305 | 83.3618 | 15.9208 | 21.5275 | 5.8899 | 136.1271 | 276.3629 | 212.6684 | 271.1531 | 188.8895
2012-10-30 00:28 | 26.0403 | 297.9224 | 234.3031 | 292.6549 | 210.233 | 85.4077 | 1.4441 | 36.537 | 52.7619 | 25.3014 | 2.2882 | 36.4294 | 83.3571 | 15.9207 | 21.5044 | 5.888 | 136.119 | 276.418 | 212.7987 | 271.1505 | 188.7286
2012-10-30 00:28 | 26.0253 | 297.9074 | 234.2213 | 292.6361 | 210.0644 | 85.4152 | 1.4441 | 36.5367 | 52.7612 | 25.3013 | 2.2881 | 36.4291 | 83.3549 | 15.9208 | 21.5175 | 5.8885 | 136.1161 | 276.3899 | 212.7038 | 271.1186 | 188.5469
2012-10-30 00:28 | 26.0276 | 297.9251 | 234.2354 | 292.673 | 210.2676 | 85.4228 | 1.444 | 36.535 | 52.7563 | 25.3014 | 2.2881 | 36.4276 | 83.349 | 15.9207 | 21.4837 | 5.889 | 136.1052 | 276.4414 | 212.7517 | 271.1893 | 188.7839
2012-10-30 00:29 | 26.0306 | 297.9204 | 234.262 | 292.6683 | 210.1997 | 85.4262 | 1.444 | 36.5344 | 52.7546 | 25.3014 | 2.288 | 36.4268 | 83.346 | 15.9205 | 21.4987 | 5.8893 | 136.1005 | 276.4217 | 212.7633 | 271.1696 | 188.701
2012-10-30 00:29 | 26.0171 | 297.9184 | 234.217 | 292.6587 | 210.2689 | 85.4268 | 1.4439 | 36.5333 | 52.7517 | 25.3013 | 2.288 | 36.4258 | 83.3414 | 15.9205 | 21.4852 | 5.8893 | 136.0931 | 276.4332 | 212.7318 | 271.1735 | 188.7837
2012-10-30 00:29 | 26.0186 | 297.9468 | 234.2578 | 292.7025 | 210.4633 | 85.4204 | 1.4439 | 36.532 | 52.7477 | 25.3013 | 2.2879 | 36.4244 | 83.3346 | 15.9206 | 21.5229 | 5.8889 | 136.0823 | 276.4239 | 212.7349 | 271.1796 | 188.9404
2012-10-30 00:29 | 26.0283 | 297.9603 | 234.3029 | 292.7238 | 210.4612 | 85.4268 | 1.4438 | 36.5309 | 52.7448 | 25.3012 | 2.2878 | 36.4232 | 83.3297 | 15.9205 | 21.5286 | 5.8893 | 136.0745 | 276.4317 | 212.7743 | 271.1952 | 188.9326
2012-10-30 00:29 | 26.0216 | 297.946 | 234.2923 | 292.7286 | 210.3679 | 85.4204 | 1.4438 | 36.5294 | 52.7403 | 25.3012 | 2.2877 | 36.4217 | 83.3234 | 15.9204 | 21.4978 | 5.8889 | 136.0637 | 276.4482 | 212.7945 | 271.2308 | 188.8701
2012-10-30 00:29 | 26.0276 | 297.9558 | 234.3454 | 292.7269 | 210.4802 | 85.4251 | 1.4438 | 36.5289 | 52.7387 | 25.3013 | 2.2877 | 36.4214 | 83.322 | 15.9204 | 21.5078 | 5.8892 | 136.0606 | 276.448 | 212.8376 | 271.2191 | 188.9724
2012-10-30 00:29 | 26.0186 | 297.9852 | 234.3207 | 292.7102 | 210.5106 | 85.4326 | 1.4437 | 36.5279 | 52.736 | 25.3012 | 2.2877 | 36.4204 | 83.3176 | 15.9203 | 21.5108 | 5.8897 | 136.0536 | 276.4744 | 212.8099 | 271.1994 | 188.9998
2012-10-30 00:29 | 26.0231 | 297.9743 | 234.3448 | 292.6454 | 210.4678 | 85.4291 | 1.4437 | 36.5273 | 52.7344 | 25.3011 | 2.2876 | 36.4198 | 83.3149 | 15.9203 | 21.5073 | 5.8895 | 136.0494 | 276.467 | 212.8375 | 271.1381 | 188.9605
2012-10-30 00:29 | 26.0194 | 298.0014 | 234.3293 | 292.788 | 210.4247 | 85.421 | 1.4436 | 36.5256 | 52.7297 | 25.3011 | 2.2875 | 36.4181 | 83.3075 | 15.9203 | 21.5197 | 5.8889 | 136.0372 | 276.4817 | 212.8096 | 271.2683 | 188.905
2012-10-30 00:29 | 26.0291 | 297.9688 | 234.3508 | 292.7053 | 210.3714 | 85.428 | 1.4436 | 36.5253 | 52.7283 | 25.3013 | 2.2875 | 36.4177 | 83.3058 | 15.9203 | 21.5133 | 5.8894 | 136.0341 | 276.4555 | 212.8375 | 271.192 | 188.8581
2012-10-30 00:29 | 26.0276 | 297.9443 | 234.3807 | 292.75 | 210.3187 | 85.4361 | 1.4436 | 36.5252 | 52.7286 | 25.301 | 2.2875 | 36.4176 | 83.3049 | 15.9203 | 21.5399 | 5.89 | 136.0336 | 276.4044 | 212.8408 | 271.2101 | 188.7788
2012-10-30 00:29 | 26.0343 | 297.9548 | 234.3285 | 292.7221 | 210.4003 | 85.4303 | 1.4436 | 36.5244 | 52.7259 | 25.3013 | 2.2874 | 36.4168 | 83.3013 | 15.9203 | 21.5305 | 5.8896 | 136.0272 | 276.4243 | 212.798 | 271.1916 | 188.8698
2012-10-30 00:29 | 26.0261 | 297.9773 | 234.3871 | 292.7216 | 210.5023 | 85.4262 | 1.4435 | 36.5234 | 52.7228 | 25.3013 | 2.2874 | 36.4159 | 83.2971 | 15.9203 | 21.5344 | 5.8893 | 136.02 | 276.4429 | 212.8527 | 271.1872 | 188.9679
2012-10-30 00:29 | 26.0194 | 297.963 | 234.3647 | 292.7418 | 210.3775 | 85.4343 | 1.4435 | 36.5222 | 52.7194 | 25.3014 | 2.2873 | 36.4149 | 83.2935 | 15.9202 | 21.4956 | 5.8898 | 136.0128 | 276.4674 | 212.8691 | 271.2462 | 188.8819
2012-10-30 00:29 | 26.0253 | 297.9765 | 234.3549 | 292.7284 | 210.407 | 85.4308 | 1.4434 | 36.5212 | 52.7162 | 25.3014 | 2.2873 | 36.4137 | 83.2879 | 15.9201 | 21.5175 | 5.8896 | 136.0041 | 276.459 | 212.8374 | 271.2109 | 188.8895
2012-10-30 00:29 | 26.0261 | 297.9658 | 234.3478 | 292.7331 | 210.4196 | 85.4337 | 1.4411 | 36.462 | 52.547 | 25.3008 | 2.2834 | 36.3521 | 83.0077 | 15.9199 | 21.5023 | 5.8898 | 135.5548 | 276.4635 | 212.8455 | 271.2308 | 188.9173
2012-10-30 00:29 | 26.0231 | 297.9398 | 234.3841 | 292.7301 | 210.3733 | 85.4135 | 1.4397 | 36.4228 | 52.4361 | 25.2997 | 2.2812 | 36.3154 | 82.841 | 15.9197 | 21.5354 | 5.8884 | 135.2771 | 276.4044 | 212.8487 | 271.1947 | 188.8379
2012-10-30 00:29 | 26.0253 | 297.942 | 234.3274 | 292.7361 | 210.092 | 85.4268 | 1.4392 | 36.4117 | 52.4045 | 25.2996 | 2.2804 | 36.3043 | 82.7898 | 15.9198 | 21.5296 | 5.8893 | 135.1943 | 276.4124 | 212.7978 | 271.2065 | 188.5624
2012-10-30 00:29 | 26.0231 | 297.9206 | 234.278 | 292.7301 | 209.4637 | 85.4152 | 1.4388 | 36.4019 | 52.3767 | 25.2993 | 2.2799 | 36.2946 | 82.7465 | 15.9197 | 21.5354 | 5.8885 | 135.1232 | 276.3852 | 212.7426 | 271.1947 | 187.9283
2012-10-30 00:29 | 26.0253 | 297.9266 | 234.186 | 292.7438 | 209.6667 | 85.4181 | 1.4388 | 36.4006 | 52.3732 | 25.2992 | 2.2798 | 36.2935 | 82.7417 | 15.9196 | 21.5135 | 5.8887 | 135.115 | 276.4131 | 212.6725 | 271.2303 | 188.1532
2012-10-30 00:29 | 26.0194 | 297.9169 | 234.1015 | 292.761 | 209.8419 | 85.4164 | 1.4389 | 36.4026 | 52.3788 | 25.2993 | 2.2799 | 36.2954 | 82.7503 | 15.9197 | 21.5518 | 5.8886 | 135.1291 | 276.3651 | 212.5497 | 271.2092 | 188.2901
2012-10-30 00:29 | 26.0059 | 297.9111 | 234.033 | 292.7591 | 209.5804 | 85.4048 | 1.439 | 36.4063 | 52.3892 | 25.2995 | 2.2801 | 36.2992 | 82.7675 | 15.9196 | 21.5623 | 5.8878 | 135.1567 | 276.3488 | 212.4707 | 271.1968 | 188.0181
2012-10-30 00:29 | 26.0126 | 297.8909 | 233.9729 | 292.7234 | 209.7958 | 85.406 | 1.4391 | 36.4074 | 52.3931 | 25.2991 | 2.2802 | 36.3 | 82.7714 | 15.9197 | 21.557 | 5.8879 | 135.1645 | 276.3339 | 212.4159 | 271.1664 | 188.2388
2012-10-30 00:29 | 26.0089 | 297.9217 | 233.9967 | 292.7621 | 210.1189 | 85.4077 | 1.4392 | 36.4117 | 52.4052 | 25.2993 | 2.2805 | 36.3045 | 82.7915 | 15.9197 | 21.5413 | 5.888 | 135.1967 | 276.3804 | 212.4554 | 271.2208 | 188.5776
2012-10-30 00:29 | 26.0134 | 297.884 | 234.0562 | 292.7473 | 210.1234 | 85.4129 | 1.4395 | 36.4179 | 52.4227 | 25.2994 | 2.2809 | 36.3108 | 82.8201 | 15.9197 | 21.5417 | 5.8884 | 135.2427 | 276.3423 | 212.5145 | 271.2056 | 188.5817
2012-10-30 00:29 | 26.0029 | 297.9349 | 234.0182 | 292.7407 | 210.0105 | 85.3962 | 1.4413 | 36.4655 | 52.5589 | 25.2998 | 2.2839 | 36.3588 | 83.039 | 15.9197 | 21.5553 | 5.8872 | 135.5979 | 276.3796 | 212.4629 | 271.1854 | 188.4552
2012-10-30 00:30 | 25.9954 | 297.9044 | 234.0107 | 292.7255 | 210.133 | 85.41 | 1.4428 | 36.5033 | 52.6666 | 25.3005 | 2.2862 | 36.3961 | 83.2085 | 15.9199 | 21.5398 | 5.8882 | 135.8752 | 276.3646 | 212.4709 | 271.1857 | 188.5932
2012-10-30 00:30 | 25.9969 | 297.952 | 234.024 | 292.7347 | 210.4731 | 85.4043 | 1.4432 | 36.515 | 52.6999 | 25.3007 | 2.2869 | 36.4075 | 83.2609 | 15.9199 | 21.5092 | 5.8878 | 135.9608 | 276.4428 | 212.5148 | 271.2255 | 188.9639
2012-10-30 00:30 | 25.9917 | 297.9237 | 234.0502 | 292.7064 | 210.4207 | 85.4106 | 1.4437 | 36.5268 | 52.7338 | 25.3009 | 2.2877 | 36.4195 | 83.316 | 15.9199 | 21.5602 | 5.8882 | 136.0498 | 276.3635 | 212.49 | 271.1462 | 188.8605
2012-10-30 00:30 | 25.9857 | 297.9638 | 234.0992 | 292.735 | 210.4029 | 85.4083 | 1.4441 | 36.5366 | 52.762 | 25.3008 | 2.2883 | 36.4292 | 83.3604 | 15.9199 | 21.5662 | 5.888 | 136.1224 | 276.3976 | 212.533 | 271.1688 | 188.8367
2012-10-30 00:30 | 25.9804 | 297.9201 | 234.1214 | 292.6912 | 210.5354 | 85.4025 | 1.4443 | 36.5419 | 52.7768 | 25.3011 | 2.2886 | 36.4345 | 83.3844 | 15.9199 | 21.5649 | 5.8876 | 136.1612 | 276.3552 | 212.5565 | 271.1263 | 188.9705
2012-10-30 00:30 | 25.976 | 297.9426 | 234.1367 | 292.6676 | 210.4168 | 85.4031 | 1.4444 | 36.5454 | 52.7869 | 25.3011 | 2.2888 | 36.4378 | 83.3993 | 15.92 | 21.5284 | 5.8877 | 136.1863 | 276.4142 | 212.6083 | 271.1392 | 188.8884
2012-10-30 00:30 | 25.9707 | 297.9143 | 234.0999 | 292.6738 | 210.38 | 85.443 | 1.4445 | 36.5476 | 52.7927 | 25.3014 | 2.289 | 36.4397 | 83.4088 | 15.9198 | 21.5632 | 5.8904 | 136.2014 | 276.3511 | 212.5367 | 271.1106 | 188.8168
2012-10-30 00:30 | 25.9647 | 297.9736 | 234.1646 | 292.6717 | 210.4134 | 85.4228 | 1.4446 | 36.5504 | 52.8008 | 25.3014 | 2.2891 | 36.4426 | 83.4227 | 15.9197 | 21.5412 | 5.889 | 136.2234 | 276.4324 | 212.6234 | 271.1305 | 188.8722
2012-10-30 00:30 | 25.9722 | 297.9081 | 234.2114 | 292.683 | 210.3224 | 85.4251 | 1.4446 | 36.551 | 52.8022 | 25.3016 | 2.2892 | 36.4433 | 83.425 | 15.9199 | 21.5367 | 5.8892 | 136.2271 | 276.3714 | 212.6747 | 271.1463 | 188.7857
2012-10-30 00:30 | 25.9782 | 297.9563 | 234.2056 | 292.6621 | 210.4505 | 85.4291 | 1.4446 | 36.5514 | 52.8034 | 25.3015 | 2.2892 | 36.4437 | 83.4277 | 15.9197 | 21.5186 | 5.8895 | 136.2311 | 276.4377 | 212.687 | 271.1435 | 188.9319
2012-10-30 00:30 | 25.9625 | 297.9176 | 234.241 | 292.631 | 210.4584 | 85.4332 | 1.4447 | 36.5516 | 52.8044 | 25.3013 | 2.2892 | 36.4438 | 83.428 | 15.9197 | 21.5229 | 5.8898 | 136.2324 | 276.3947 | 212.7181 | 271.1081 | 188.9355
2012-10-30 00:30 | 25.9677 | 297.9497 | 234.203 | 292.6555 | 210.3494 | 85.4314 | 1.4446 | 36.5513 | 52.8032 | 25.3014 | 2.2892 | 36.4435 | 83.4266 | 15.9197 | 21.5522 | 5.8896 | 136.2299 | 276.3975 | 212.6508 | 271.1033 | 188.7972
2012-10-30 00:30 | 25.967 | 297.9182 | 234.2377 | 292.6317 | 210.2857 | 85.4557 | 1.4446 | 36.5513 | 52.803 | 25.3015 | 2.2892 | 36.4436 | 83.427 | 15.9198 | 21.5475 | 5.8913 | 136.2299 | 276.3707 | 212.6902 | 271.0842 | 188.7382
2012-10-30 00:30 | 25.964 | 297.9306 | 234.1993 | 292.6479 | 210.2591 | 85.4407 | 1.4446 | 36.5513 | 52.8026 | 25.3018 | 2.2892 | 36.4436 | 83.4269 | 15.9197 | 21.5164 | 5.8903 | 136.2295 | 276.4142 | 212.6829 | 271.1315 | 188.7427
2012-10-30 00:30 | 25.9617 | 297.8899 | 234.2127 | 292.6764 | 210.2804 | 85.4441 | 1.4445 | 36.5492 | 52.7968 | 25.3016 | 2.2891 | 36.4416 | 83.4183 | 15.9197 | 21.5542 | 5.8905 | 136.2151 | 276.3357 | 212.6585 | 271.1222 | 188.7262
2012-10-30 00:30 | 25.9685 | 297.8967 | 234.1999 | 292.664 | 210.4172 | 85.4274 | 1.4445 | 36.5491 | 52.7963 | 25.3017 | 2.2891 | 36.4413 | 83.416 | 15.9198 | 21.557 | 5.8894 | 136.2123 | 276.3397 | 212.6429 | 271.107 | 188.8602
2012-10-30 00:30 | 25.9745 | 297.9104 | 234.2569 | 292.6892 | 210.2696 | 85.4297 | 1.4445 | 36.5475 | 52.7912 | 25.3019 | 2.289 | 36.4398 | 83.4101 | 15.9196 | 21.539 | 5.8895 | 136.2013 | 276.3714 | 212.7179 | 271.1502 | 188.7306
2012-10-30 00:30 | 25.9745 | 297.9027 | 234.2059 | 292.6892 | 209.9664 | 85.4308 | 1.4445 | 36.5468 | 52.7901 | 25.3015 | 2.2889 | 36.4391 | 83.4073 | 15.9196 | 21.551 | 5.8896 | 136.1973 | 276.3517 | 212.6549 | 271.1382 | 188.4154
2012-10-30 00:30 | 25.9745 | 297.8912 | 234.1116 | 292.6661 | 209.7025 | 85.4326 | 1.4444 | 36.5458 | 52.7868 | 25.3017 | 2.2889 | 36.4383 | 83.4037 | 15.9195 | 21.539 | 5.8897 | 136.1905 | 276.3522 | 212.5726 | 271.1271 | 188.1635
2012-10-30 00:30 | 25.9827 | 297.9224 | 234.108 | 292.682 | 209.7383 | 85.4291 | 1.4443 | 36.5444 | 52.7826 | 25.3018 | 2.2888 | 36.4366 | 83.3955 | 15.9196 | 21.5472 | 5.8895 | 136.1781 | 276.3752 | 212.5608 | 271.1348 | 188.1911
2012-10-30 00:30 | 25.9707 | 297.9219 | 234.0607 | 292.6854 | 209.5845 | 85.4366 | 1.4443 | 36.5439 | 52.7814 | 25.3016 | 2.2888 | 36.4364 | 83.3953 | 15.9195 | 21.5311 | 5.89 | 136.1767 | 276.3908 | 212.5296 | 271.1543 | 188.0534
2012-10-30 00:30 | 25.9745 | 297.8912 | 233.9898 | 292.6853 | 209.3796 | 85.4505 | 1.4443 | 36.5431 | 52.7796 | 25.3015 | 2.2888 | 36.4355 | 83.3919 | 15.9193 | 21.5189 | 5.891 | 136.1715 | 276.3723 | 212.4709 | 271.1664 | 187.8607
2012-10-30 00:30 | 25.9834 | 297.9001 | 233.9869 | 292.6827 | 209.8965 | 85.4337 | 1.4443 | 36.5426 | 52.7783 | 25.3014 | 2.2887 | 36.4349 | 83.3887 | 15.9195 | 21.5318 | 5.8898 | 136.1669 | 276.3683 | 212.4551 | 271.1509 | 188.3647
2012-10-30 00:30 | 25.9842 | 297.9009 | 233.9367 | 292.6797 | 209.8697 | 85.4418 | 1.4442 | 36.5417 | 52.7752 | 25.3016 | 2.2887 | 36.4339 | 83.3852 | 15.9193 | 21.5286 | 5.8904 | 136.1604 | 276.3723 | 212.4081 | 271.1511 | 188.3411
2012-10-30 00:30 | 25.9827 | 297.9532 | 233.9587 | 292.6512 | 209.9155 | 85.4447 | 1.4442 | 36.5414 | 52.7741 | 25.3017 | 2.2886 | 36.4337 | 83.3839 | 15.9193 | 21.5271 | 5.8906 | 136.158 | 276.4261 | 212.4316 | 271.1241 | 188.3884
2012-10-30 00:30 | 25.9917 | 297.9084 | 233.9756 | 292.7333 | 209.8969 | 85.4361 | 1.4442 | 36.5409 | 52.7731 | 25.3015 | 2.2886 | 36.4332 | 83.3824 | 15.9192 | 21.5401 | 5.89 | 136.1554 | 276.3683 | 212.4355 | 271.1932 | 188.3568
2012-10-30 00:31 | 25.9872 | 297.9116 | 233.9514 | 292.7134 | 209.9082 | 85.4407 | 1.4441 | 36.5386 | 52.7663 | 25.3016 | 2.2885 | 36.4309 | 83.3718 | 15.9192 | 21.5356 | 5.8903 | 136.1381 | 276.376 | 212.4158 | 271.1778 | 188.3726
2012-10-30 00:31 | 25.9797 | 297.9002 | 234.0068 | 292.7021 | 209.9834 | 85.4517 | 1.4441 | 36.5388 | 52.7663 | 25.3018 | 2.2885 | 36.4312 | 83.3734 | 15.9191 | 21.504 | 5.891 | 136.1398 | 276.3962 | 212.5028 | 271.1981 | 188.4794
2012-10-30 00:31 | 25.9917 | 297.8853 | 234.0149 | 292.7218 | 210.1017 | 85.4378 | 1.4442 | 36.5393 | 52.7686 | 25.3014 | 2.2885 | 36.4317 | 83.3751 | 15.9192 | 21.5241 | 5.8901 | 136.1437 | 276.3612 | 212.4908 | 271.1977 | 188.5776
2012-10-30 00:31 | 25.9902 | 297.9376 | 234.0448 | 292.7087 | 210.2105 | 85.4407 | 1.4441 | 36.537 | 52.7616 | 25.3015 | 2.2884 | 36.4293 | 83.3637 | 15.9193 | 21.5185 | 5.8903 | 136.1254 | 276.4191 | 212.5263 | 271.1902 | 188.692
2012-10-30 00:31 | 25.9902 | 297.9338 | 234.0527 | 292.7549 | 210.2932 | 85.4522 | 1.4441 | 36.5368 | 52.7612 | 25.3015 | 2.2884 | 36.4293 | 83.3645 | 15.9192 | 21.4985 | 5.8911 | 136.1257 | 276.4353 | 212.5542 | 271.2564 | 188.7947
2012-10-30 00:31 | 25.9872 | 297.9154 | 234.1361 | 292.7327 | 210.1917 | 85.4488 | 1.444 | 36.5352 | 52.7563 | 25.3017 | 2.2883 | 36.4275 | 83.3567 | 15.9191 | 21.5557 | 5.8908 | 136.113 | 276.3597 | 212.5804 | 271.177 | 188.636
2012-10-30 00:31 | 25.9902 | 297.9261 | 234.0959 | 292.7434 | 209.8876 | 85.4401 | 1.4439 | 36.5338 | 52.7522 | 25.3017 | 2.2882 | 36.4261 | 83.35 | 15.9192 | 21.5948 | 5.8902 | 136.1022 | 276.3313 | 212.5011 | 271.1486 | 188.2928
2012-10-30 00:31 | 25.9842 | 297.8663 | 234.1252 | 292.7104 | 210.11 | 85.4384 | 1.4439 | 36.5332 | 52.7503 | 25.3017 | 2.2882 | 36.4254 | 83.3474 | 15.919 | 21.5326 | 5.8901 | 136.0977 | 276.3337 | 212.5926 | 271.1778 | 188.5774
2012-10-30 00:31 | 25.9857 | 297.9523 | 234.1228 | 292.7427 | 210.2257 | 85.447 | 1.4439 | 36.5321 | 52.7477 | 25.3014 | 2.2881 | 36.4244 | 83.3436 | 15.9189 | 21.5542 | 5.8907 | 136.0912 | 276.3981 | 212.5686 | 271.1885 | 188.6715
2012-10-30 00:31 | 25.9812 | 297.9248 | 234.1654 | 292.7498 | 210.166 | 85.443 | 1.4438 | 36.5314 | 52.7457 | 25.3015 | 2.2881 | 36.4238 | 83.3405 | 15.919 | 21.5416 | 5.8904 | 136.0861 | 276.3832 | 212.6238 | 271.2082 | 188.6244
2012-10-30 00:31 | 25.9962 | 297.9667 | 234.1726 | 292.7725 | 210.1338 | 85.447 | 1.4438 | 36.5295 | 52.7398 | 25.3017 | 2.288 | 36.4218 | 83.3322 | 15.9188 | 21.5687 | 5.8907 | 136.072 | 276.398 | 212.6039 | 271.2038 | 188.5651
2012-10-30 00:31 | 25.9924 | 297.9744 | 234.1648 | 292.7494 | 210.2954 | 85.4476 | 1.4438 | 36.5287 | 52.7383 | 25.3013 | 2.2879 | 36.4212 | 83.3288 | 15.9189 | 21.5528 | 5.8908 | 136.0671 | 276.4216 | 212.612 | 271.1966 | 188.7426
2012-10-30 00:31 | 26.0014 | 297.9334 | 234.2053 | 292.7392 | 210.4382 | 85.4551 | 1.4437 | 36.5277 | 52.7346 | 25.3016 | 2.2879 | 36.42 | 83.324 | 15.9188 | 21.5257 | 5.8913 | 136.0586 | 276.4077 | 212.6796 | 271.2135 | 188.9125
2012-10-30 00:31 | 25.9984 | 297.9612 | 234.2219 | 292.7593 | 210.4904 | 85.4395 | 1.4436 | 36.5262 | 52.7302 | 25.3017 | 2.2878 | 36.4186 | 83.3177 | 15.9188 | 21.5468 | 5.8902 | 136.048 | 276.4144 | 212.6751 | 271.2125 | 188.9436
2012-10-30 00:31 | 25.9984 | 297.9112 | 234.2455 | 292.7593 | 210.4234 | 85.4551 | 1.4436 | 36.5245 | 52.7259 | 25.3015 | 2.2877 | 36.4169 | 83.3098 | 15.9188 | 21.4866 | 5.8913 | 136.0356 | 276.4246 | 212.7589 | 271.2727 | 188.9368
2012-10-30 00:31 | 26.0081 | 297.9824 | 234.3062 | 292.7536 | 210.3662 | 85.4615 | 1.4435 | 36.5238 | 52.7239 | 25.3014 | 2.2876 | 36.4161 | 83.3063 | 15.9188 | 21.5364 | 5.8917 | 136.0303 | 276.446 | 212.7698 | 271.2172 | 188.8298
2012-10-30 00:31 | 26.0111 | 297.9393 | 234.3132 | 292.7643 | 210.3574 | 85.4696 | 1.4435 | 36.5234 | 52.7228 | 25.3013 | 2.2876 | 36.4158 | 83.3054 | 15.9187 | 21.5354 | 5.8923 | 136.0281 | 276.4039 | 212.7778 | 271.2289 | 188.822
2012-10-30 00:31 | 26.0111 | 297.947 | 234.3328 | 292.7835 | 210.2471 | 85.454 | 1.441 | 36.4595 | 52.5394 | 25.3009 | 2.2836 | 36.351 | 83.0101 | 15.9185 | 21.5435 | 5.8912 | 135.5495 | 276.4035 | 212.7893 | 271.24 | 188.7036
2012-10-30 00:31 | 26.0104 | 297.9424 | 234.3125 | 292.7597 | 210.2149 | 85.4418 | 1.4396 | 36.4209 | 52.4301 | 25.3001 | 2.2813 | 36.3137 | 82.841 | 15.9183 | 21.5307 | 5.8904 | 135.271 | 276.4117 | 212.7818 | 271.229 | 188.6842
2012-10-30 00:31 | 26.0231 | 297.9551 | 234.2623 | 292.7763 | 209.9205 | 85.4453 | 1.439 | 36.4062 | 52.3878 | 25.2999 | 2.2803 | 36.299 | 82.7741 | 15.9182 | 21.5153 | 5.8906 | 135.1619 | 276.4398 | 212.747 | 271.261 | 188.4052
2012-10-30 00:31 | 26.0104 | 297.9232 | 234.2103 | 292.7636 | 209.7345 | 85.4355 | 1.4389 | 36.4029 | 52.3794 | 25.2995 | 2.2801 | 36.2957 | 82.7589 | 15.9183 | 21.5387 | 5.8899 | 135.1383 | 276.3845 | 212.6716 | 271.2249 | 188.1958
2012-10-30 00:31 | 26.0186 | 297.9045 | 234.2028 | 292.7679 | 209.5182 | 85.454 | 1.4388 | 36.4018 | 52.376 | 25.2996 | 2.2801 | 36.2946 | 82.7541 | 15.9182 | 21.563 | 5.8912 | 135.1301 | 276.3415 | 212.6398 | 271.2049 | 187.9552
2012-10-30 00:31 | 26.0036 | 297.9587 | 234.1171 | 292.7568 | 209.5781 | 85.4268 | 1.4388 | 36.4012 | 52.3746 | 25.2993 | 2.2801 | 36.2943 | 82.7528 | 15.9182 | 21.56 | 5.8893 | 135.1275 | 276.3987 | 212.5571 | 271.1968 | 188.0181
2012-10-30 00:31 | 26.0194 | 297.9322 | 234.1369 | 292.7572 | 210.4878 | 85.4343 | 1.4389 | 36.4025 | 52.378 | 25.2995 | 2.2801 | 36.2953 | 82.7573 | 15.9182 | 21.5357 | 5.8898 | 135.1353 | 276.3965 | 212.6012 | 271.2215 | 188.9521
2012-10-30 00:31 | 26.0029 | 297.8812 | 234.1714 | 292.7484 | 210.6327 | 85.4372 | 1.439 | 36.4056 | 52.3868 | 25.2997 | 2.2803 | 36.2983 | 82.7713 | 15.9181 | 21.5513 | 5.89 | 135.1581 | 276.3299 | 212.6201 | 271.1971 | 189.0814
2012-10-30 00:32 | 26.0006 | 297.9058 | 234.2241 | 292.7692 | 210.6343 | 85.4228 | 1.4392 | 36.4115 | 52.4038 | 25.2996 | 2.2807 | 36.3043 | 82.7987 | 15.9181 | 21.541 | 5.889 | 135.2026 | 276.3648 | 212.6831 | 271.2282 | 189.0933
2012-10-30 00:32 | 26.0059 | 297.9072 | 234.2373 | 292.7706 | 210.297 | 85.4378 | 1.4393 | 36.4143 | 52.4113 | 25.2999 | 2.2808 | 36.307 | 82.8104 | 15.9183 | 21.5503 | 5.8901 | 135.2216 | 276.3569 | 212.687 | 271.2203 | 188.7467
2012-10-30 00:32 | 25.9954 | 297.889 | 234.2386 | 292.7486 | 210.0857 | 85.4389 | 1.4399 | 36.4284 | 52.4527 | 25.2996 | 2.2817 | 36.3203 | 82.8711 | 15.9183 | 21.5037 | 5.8902 | 135.3238 | 276.3853 | 212.7349 | 271.2449 | 188.582
2012-10-30 00:32 | 25.9917 | 297.9007 | 234.1838 | 292.7718 | 210.1371 | 85.4245 | 1.4427 | 36.5006 | 52.6582 | 25.3008 | 2.2862 | 36.3926 | 83.2008 | 15.9184 | 21.5682 | 5.8892 | 135.8589 | 276.3325 | 212.6156 | 271.2036 | 188.5689
2012-10-30 00:32 | 25.9879 | 297.9391 | 234.1093 | 292.7488 | 210.3499 | 85.4389 | 1.4433 | 36.5164 | 52.7032 | 25.3011 | 2.2872 | 36.4085 | 83.2731 | 15.9185 | 21.5283 | 5.8902 | 135.9763 | 276.4108 | 212.581 | 271.2205 | 188.8216
2012-10-30 00:32 | 25.9872 | 297.9116 | 234.0889 | 292.7019 | 210.0106 | 85.4493 | 1.4438 | 36.5287 | 52.7384 | 25.3012 | 2.288 | 36.4208 | 83.3295 | 15.9184 | 21.5717 | 5.8909 | 136.0679 | 276.3399 | 212.5172 | 271.1302 | 188.4389
2012-10-30 00:32 | 25.9797 | 297.9655 | 234.0264 | 292.6828 | 210.2315 | 85.4297 | 1.4441 | 36.5379 | 52.7645 | 25.3015 | 2.2885 | 36.43 | 83.3711 | 15.9185 | 21.5722 | 5.8895 | 136.1357 | 276.3933 | 212.4542 | 271.1106 | 188.6593
2012-10-30 00:32 | 25.9857 | 297.8947 | 234.0599 | 292.735 | 210.1272 | 85.4384 | 1.4443 | 36.5433 | 52.7799 | 25.3015 | 2.2889 | 36.4354 | 83.3962 | 15.9184 | 21.5221 | 5.8901 | 136.1761 | 276.3726 | 212.5378 | 271.2129 | 188.6051
2012-10-30 00:32 | 25.9849 | 297.9323 | 234.0827 | 292.7188 | 210.4651 | 85.4476 | 1.4444 | 36.5455 | 52.7863 | 25.3015 | 2.289 | 36.4379 | 83.4071 | 15.9185 | 21.5534 | 5.8908 | 136.1935 | 276.3789 | 212.5293 | 271.1654 | 188.9117
2012-10-30 00:32 | 25.9849 | 297.9169 | 234.1534 | 292.6534 | 210.4021 | 85.443 | 1.4444 | 36.5463 | 52.788 | 25.3017 | 2.2891 | 36.4384 | 83.4105 | 15.9184 | 21.5373 | 5.8904 | 136.1985 | 276.3796 | 212.6161 | 271.1161 | 188.8648
2012-10-30 00:32 | 25.9842 | 297.9124 | 234.1527 | 292.7143 | 210.421 | 85.4597 | 1.4445 | 36.5488 | 52.7954 | 25.3017 | 2.2892 | 36.441 | 83.4224 | 15.9184 | 21.5085 | 5.8916 | 136.2179 | 276.4039 | 212.6442 | 271.2058 | 188.9125
2012-10-30 00:32 | 25.9864 | 297.8839 | 234.1667 | 292.6819 | 210.3169 | 85.454 | 1.4446 | 36.5496 | 52.7978 | 25.3016 | 2.2893 | 36.4418 | 83.4266 | 15.9182 | 21.5709 | 5.8912 | 136.2244 | 276.313 | 212.5958 | 271.111 | 188.746
2012-10-30 00:32 | 25.9834 | 297.9154 | 234.1755 | 292.6673 | 210.3139 | 85.4603 | 1.4446 | 36.5499 | 52.7984 | 25.3017 | 2.2893 | 36.4421 | 83.4276 | 15.9183 | 21.5519 | 5.8916 | 136.2261 | 276.3635 | 212.6236 | 271.1154 | 188.762
2012-10-30 00:32 | 25.9827 | 297.9071 | 234.1512 | 292.6512 | 210.262 | 85.4742 | 1.4446 | 36.5506 | 52.8011 | 25.3015 | 2.2894 | 36.443 | 83.4318 | 15.9183 | 21.5391 | 5.8926 | 136.2329 | 276.368 | 212.6121 | 271.1121 | 188.7229
2012-10-30 00:32 | 25.9842 | 297.8509 | 234.192 | 292.6873 | 210.2635 | 85.4603 | 1.4446 | 36.5512 | 52.8029 | 25.3015 | 2.2894 | 36.4434 | 83.4341 | 15.9182 | 21.5206 | 5.8916 | 136.237 | 276.3303 | 212.6714 | 271.1667 | 188.7429
2012-10-30 00:32 | 25.9842 | 297.9086 | 234.1488 | 292.6989 | 210.3817 | 85.4655 | 1.4446 | 36.5505 | 52.8008 | 25.3015 | 2.2894 | 36.443 | 83.4322 | 15.9183 | 21.4965 | 5.892 | 136.233 | 276.4121 | 212.6523 | 271.2024 | 188.8852
2012-10-30 00:32 | 25.9767 | 297.8703 | 234.1766 | 292.6875 | 210.2915 | 85.4701 | 1.4446 | 36.5501 | 52.7988 | 25.3018 | 2.2893 | 36.4422 | 83.4287 | 15.9182 | 21.489 | 5.8923 | 136.2275 | 276.3813 | 212.6876 | 271.1985 | 188.8025
2012-10-30 00:32 | 25.9864 | 297.8877 | 234.2138 | 292.7049 | 210.2421 | 85.4574 | 1.4446 | 36.5499 | 52.7983 | 25.3019 | 2.2894 | 36.4423 | 83.4291 | 15.9182 | 21.5107 | 5.8914 | 136.2275 | 276.377 | 212.7031 | 271.1942 | 188.7314
2012-10-30 00:32 | 25.9834 | 297.8693 | 234.1833 | 292.6712 | 210.2942 | 85.4661 | 1.4445 | 36.5494 | 52.7972 | 25.3017 | 2.2893 | 36.4415 | 83.4257 | 15.9182 | 21.4877 | 5.892 | 136.2228 | 276.3816 | 212.6956 | 271.1835 | 188.8065
2012-10-30 00:32 | 25.9872 | 297.8924 | 234.1911 | 292.7057 | 210.3492 | 85.4725 | 1.4445 | 36.5491 | 52.7962 | 25.3017 | 2.2893 | 36.4413 | 83.4249 | 15.9182 | 21.5436 | 5.8925 | 136.2212 | 276.3488 | 212.6475 | 271.1621 | 188.8056
2012-10-30 00:32 | 25.9887 | 297.9169 | 234.2044 | 292.6649 | 210.4452 | 85.4667 | 1.4445 | 36.5476 | 52.7921 | 25.3017 | 2.2892 | 36.44 | 83.4187 | 15.9181 | 21.5491 | 5.8921 | 136.2108 | 276.3678 | 212.6553 | 271.1158 | 188.8961
2012-10-30 00:32 | 25.9909 | 297.9114 | 234.2498 | 292.7056 | 210.6798 | 85.4765 | 1.4444 | 36.547 | 52.7896 | 25.302 | 2.2892 | 36.4394 | 83.416 | 15.9182 | 21.5554 | 5.8927 | 136.2056 | 276.356 | 212.6944 | 271.1502 | 189.1244
2012-10-30 00:32 | 25.9887 | 297.9361 | 234.2829 | 292.7034 | 210.7091 | 85.4621 | 1.4444 | 36.5469 | 52.789 | 25.3022 | 2.2891 | 36.4392 | 83.4145 | 15.9183 | 21.517 | 5.8918 | 136.2035 | 276.4191 | 212.7659 | 271.1864 | 189.1921
2012-10-30 00:32 | 25.9894 | 297.8946 | 234.3308 | 292.6849 | 210.572 | 85.4667 | 1.4444 | 36.5458 | 52.7862 | 25.3019 | 2.2891 | 36.4381 | 83.4097 | 15.9182 | 21.5177 | 5.8921 | 136.1959 | 276.3769 | 212.8131 | 271.1672 | 189.0543
2012-10-30 00:32 | 25.9864 | 297.8954 | 234.3238 | 292.7011 | 210.4548 | 85.48 | 1.4443 | 36.5441 | 52.7806 | 25.3022 | 2.289 | 36.4364 | 83.4019 | 15.9183 | 21.4626 | 5.893 | 136.1825 | 276.4328 | 212.8612 | 271.2385 | 188.9922
2012-10-30 00:32 | 25.9924 | 297.8822 | 234.2984 | 292.6879 | 210.3032 | 85.4667 | 1.4443 | 36.5429 | 52.7779 | 25.302 | 2.2889 | 36.4353 | 83.397 | 15.9182 | 21.5047 | 5.8921 | 136.1749 | 276.3775 | 212.7937 | 271.1832 | 188.7985
2012-10-30 00:32 | 25.9902 | 297.8992 | 234.3041 | 292.7203 | 210.3404 | 85.4701 | 1.4442 | 36.5418 | 52.7738 | 25.3024 | 2.2888 | 36.4342 | 83.3916 | 15.9183 | 21.5065 | 5.8923 | 136.1654 | 276.3927 | 212.7976 | 271.2138 | 188.8339
2012-10-30 00:33 | 25.9969 | 297.9213 | 234.3343 | 292.7308 | 210.2526 | 85.4846 | 1.4442 | 36.5407 | 52.7711 | 25.3021 | 2.2888 | 36.4329 | 83.386 | 15.9182 | 21.5052 | 5.8933 | 136.1571 | 276.4161 | 212.8291 | 271.2256 | 188.7474
2012-10-30 00:33 | 25.9999 | 297.8974 | 234.2902 | 292.78 | 210.2359 | 85.454 | 1.4442 | 36.5407 | 52.7709 | 25.3022 | 2.2888 | 36.4328 | 83.3861 | 15.9181 | 21.4761 | 5.8912 | 136.157 | 276.4213 | 212.8141 | 271.3039 | 188.7598
2012-10-30 00:33 | 25.9977 | 297.9105 | 234.2998 | 292.7624 | 210.2495 | 85.4788 | 1.4441 | 36.5383 | 52.7641 | 25.3022 | 2.2886 | 36.4305 | 83.3764 | 15.918 | 21.498 | 5.8929 | 136.1405 | 276.4125 | 212.8018 | 271.2644 | 188.7515
2012-10-30 00:33 | 25.9992 | 297.9543 | 234.2581 | 292.7562 | 210.125 | 85.4742 | 1.444 | 36.536 | 52.7576 | 25.3022 | 2.2885 | 36.4285 | 83.3674 | 15.9179 | 21.4754 | 5.8926 | 136.125 | 276.4789 | 212.7827 | 271.2808 | 188.6496
2012-10-30 00:33 | 26.0014 | 297.9603 | 234.2485 | 292.7854 | 210.2217 | 85.4817 | 1.444 | 36.5349 | 52.7549 | 25.3019 | 2.2885 | 36.4273 | 83.3628 | 15.9177 | 21.5257 | 5.8931 | 136.1177 | 276.4346 | 212.7228 | 271.2597 | 188.696
2012-10-30 00:33 | 26.0066 | 297.9502 | 234.2537 | 292.7482 | 210.0536 | 85.4777 | 1.4439 | 36.5346 | 52.7539 | 25.3019 | 2.2885 | 36.4272 | 83.3619 | 15.9178 | 21.5149 | 5.8928 | 136.1158 | 276.4353 | 212.7388 | 271.2333 | 188.5387
2012-10-30 00:33 | 26.0081 | 297.967 | 234.2159 | 292.8113 | 209.6298 | 85.4875 | 1.4439 | 36.5338 | 52.7515 | 25.3019 | 2.2884 | 36.4261 | 83.3576 | 15.9177 | 21.4963 | 5.8935 | 136.1091 | 276.4707 | 212.7196 | 271.315 | 188.1335
2012-10-30 00:33 | 26.0096 | 297.9416 | 234.1428 | 292.7743 | 210.0448 | 85.4736 | 1.4439 | 36.533 | 52.7491 | 25.302 | 2.2883 | 36.4253 | 83.3533 | 15.9178 | 21.4777 | 5.8925 | 136.1024 | 276.4639 | 212.6651 | 271.2966 | 188.5671
2012-10-30 00:33 | 26.0149 | 298.0199 | 234.1481 | 292.7796 | 210.1997 | 85.4678 | 1.4439 | 36.5329 | 52.7489 | 25.302 | 2.2883 | 36.4253 | 83.3529 | 15.9178 | 21.5192 | 5.8922 | 136.1018 | 276.5007 | 212.6289 | 271.2604 | 188.6805
2012-10-30 00:33 | 26.0104 | 297.9348 | 234.175 | 292.7867 | 210.1716 | 85.4806 | 1.4438 | 36.5314 | 52.744 | 25.3023 | 2.2882 | 36.4237 | 83.3454 | 15.9179 | 21.4906 | 5.893 | 136.0895 | 276.4442 | 212.6844 | 271.2961 | 188.681
2012-10-30 00:33 | 26.0141 | 297.9769 | 234.1826 | 292.8173 | 210.1674 | 85.4875 | 1.4437 | 36.5299 | 52.7394 | 25.3025 | 2.2881 | 36.4222 | 83.3393 | 15.9178 | 21.4822 | 5.8935 | 136.0787 | 276.4947 | 212.7004 | 271.3351 | 188.6852
2012-10-30 00:33 | 26.0223 | 298.0004 | 234.2144 | 292.8024 | 210.3016 | 85.4869 | 1.4437 | 36.5296 | 52.7391 | 25.3022 | 2.2881 | 36.422 | 83.3385 | 15.9178 | 21.5226 | 5.8935 | 136.0776 | 276.4778 | 212.6918 | 271.2798 | 188.779
2012-10-30 00:33 | 26.0089 | 297.9256 | 234.2246 | 292.8121 | 210.3473 | 85.4753 | 1.4437 | 36.5283 | 52.7352 | 25.3022 | 2.2881 | 36.4206 | 83.3323 | 15.9178 | 21.5051 | 5.8927 | 136.0675 | 276.4205 | 212.7195 | 271.307 | 188.8422
2012-10-30 00:33 | 26.0291 | 298.0226 | 234.2526 | 292.8131 | 210.4029 | 85.4858 | 1.4437 | 36.5275 | 52.7331 | 25.3021 | 2.288 | 36.4199 | 83.3292 | 15.9177 | 21.5855 | 5.8934 | 136.0623 | 276.4371 | 212.6671 | 271.2276 | 188.8174
2012-10-30 00:33 | 26.0201 | 298.029 | 234.2908 | 292.8195 | 210.39 | 85.4846 | 1.4436 | 36.5265 | 52.7301 | 25.3022 | 2.288 | 36.4189 | 83.3248 | 15.9176 | 21.5083 | 5.8933 | 136.0549 | 276.5207 | 212.7825 | 271.3112 | 188.8817
2012-10-30 00:33 | 26.0156 | 297.9784 | 234.3255 | 292.7765 | 210.2516 | 85.4788 | 1.4435 | 36.5245 | 52.7245 | 25.3021 | 2.2879 | 36.417 | 83.3166 | 15.9176 | 21.552 | 5.8929 | 136.0411 | 276.4264 | 212.7735 | 271.2245 | 188.6996
2012-10-30 00:33 | 26.0253 | 298.0034 | 234.3352 | 292.8862 | 210.3834 | 85.4823 | 1.4436 | 36.5248 | 52.7257 | 25.302 | 2.2878 | 36.4172 | 83.3166 | 15.9177 | 21.5898 | 5.8931 | 136.0423 | 276.4136 | 212.7454 | 271.2964 | 188.7936
2012-10-30 00:33 | 26.0171 | 297.9952 | 234.3585 | 292.8357 | 210.4579 | 85.4725 | 1.4435 | 36.5236 | 52.7216 | 25.3022 | 2.2878 | 36.416 | 83.3118 | 15.9176 | 21.5374 | 5.8925 | 136.0334 | 276.4578 | 212.8211 | 271.2983 | 188.9205
2012-10-30 00:33 | 26.0209 | 298.0298 | 234.3898 | 292.8549 | 210.4617 | 85.484 | 1.4435 | 36.5222 | 52.718 | 25.302 | 2.2877 | 36.4147 | 83.306 | 15.9176 | 21.5653 | 5.8933 | 136.024 | 276.4645 | 212.8245 | 271.2896 | 188.8964
2012-10-30 00:33 | 26.0186 | 297.9929 | 234.3953 | 292.8257 | 210.4673 | 85.4806 | 1.4434 | 36.5206 | 52.7129 | 25.3022 | 2.2876 | 36.4129 | 83.2975 | 15.9176 | 21.5911 | 5.893 | 136.0103 | 276.4018 | 212.8042 | 271.2346 | 188.8762
2012-10-30 00:33 | 26.0253 | 297.9727 | 234.4059 | 292.8324 | 210.4543 | 85.4852 | 1.4433 | 36.5192 | 52.7081 | 25.3025 | 2.2871 | 36.4058 | 83.2655 | 15.9176 | 21.5496 | 5.8933 | 135.9736 | 276.4231 | 212.8563 | 271.2828 | 188.9047
2012-10-30 00:33 | 26.0149 | 298.0046 | 234.3563 | 292.8643 | 210.4754 | 85.4892 | 1.4393 | 36.4158 | 52.4142 | 25.3006 | 2.2812 | 36.3098 | 82.8283 | 15.9173 | 21.5954 | 5.8936 | 135.2425 | 276.4092 | 212.7609 | 271.2689 | 188.88
2012-10-30 00:33 | 26.0231 | 298.0204 | 234.4116 | 292.8609 | 210.4796 | 85.4834 | 1.4393 | 36.4156 | 52.4138 | 25.3005 | 2.2811 | 36.3082 | 82.8214 | 15.9172 | 21.5595 | 5.8932 | 135.2352 | 276.4609 | 212.8521 | 271.3014 | 188.9201
2012-10-30 00:33 | 26.0313 | 298.0133 | 234.3727 | 292.8537 | 209.7987 | 85.4597 | 1.4389 | 36.4057 | 52.3856 | 25.3004 | 2.2805 | 36.2984 | 82.7773 | 15.9171 | 21.5596 | 5.8916 | 135.1629 | 276.4537 | 212.8131 | 271.2941 | 188.2391
2012-10-30 00:34 | 26.0321 | 297.9987 | 234.287 | 292.8545 | 209.3663 | 85.4753 | 1.4387 | 36.3999 | 52.3693 | 25.3002 | 2.2801 | 36.2926 | 82.7507 | 15.9171 | 21.4962 | 5.8927 | 135.12 | 276.5025 | 212.7908 | 271.3583 | 187.8701
2012-10-30 00:34 | 26.0261 | 297.9351 | 234.1357 | 292.8793 | 209.4509 | 85.4632 | 1.4386 | 36.398 | 52.364 | 25.3001 | 2.28 | 36.2908 | 82.7425 | 15.9171 | 21.5063 | 5.8918 | 135.1065 | 276.4288 | 212.6294 | 271.373 | 187.9446
2012-10-30 00:34 | 26.0306 | 297.9358 | 234.0813 | 292.8415 | 208.9908 | 85.4638 | 1.4388 | 36.4018 | 52.3743 | 25.3004 | 2.2802 | 36.2945 | 82.7593 | 15.9171 | 21.563 | 5.8919 | 135.1336 | 276.3728 | 212.5183 | 271.2785 | 187.4278
2012-10-30 00:34 | 26.0179 | 297.996 | 234.045 | 292.8442 | 209.4388 | 85.4748 | 1.4389 | 36.4056 | 52.3857 | 25.3002 | 2.2805 | 36.2983 | 82.7768 | 15.9171 | 21.5382 | 5.8926 | 135.1625 | 276.4578 | 212.5068 | 271.306 | 187.9006
2012-10-30 00:34 | 26.0201 | 298.0021 | 234.0315 | 292.8425 | 209.886 | 85.4782 | 1.439 | 36.4083 | 52.3933 | 25.3003 | 2.2806 | 36.3008 | 82.7877 | 15.9172 | 21.5003 | 5.8929 | 135.181 | 276.5018 | 212.5312 | 271.3422 | 188.3857
2012-10-30 00:34 | 26.0246 | 297.9758 | 234.0321 | 292.8778 | 210.2094 | 85.4696 | 1.4393 | 36.4136 | 52.4087 | 25.3002 | 2.281 | 36.3064 | 82.8133 | 15.9172 | 21.5489 | 5.8923 | 135.222 | 276.4269 | 212.4832 | 271.3289 | 188.6605
2012-10-30 00:34 | 26.0216 | 297.9114 | 234.1076 | 292.8248 | 209.7024 | 85.4545 | 1.4395 | 36.4186 | 52.4231 | 25.3002 | 2.2813 | 36.3112 | 82.8359 | 15.9171 | 21.5219 | 5.8912 | 135.259 | 276.3895 | 212.5857 | 271.3029 | 188.1805
2012-10-30 00:34 | 26.0238 | 297.9443 | 233.9802 | 292.8539 | 209.8778 | 85.4476 | 1.4417 | 36.4774 | 52.5906 | 25.301 | 2.2848 | 36.3674 | 83.0914 | 15.9173 | 21.5521 | 5.8908 | 135.682 | 276.3922 | 212.4281 | 271.3018 | 188.3257
2012-10-30 00:34 | 26.0216 | 297.9613 | 233.923 | 292.8594 | 210.1474 | 85.4678 | 1.4429 | 36.5074 | 52.6761 | 25.3015 | 2.2868 | 36.3997 | 83.2385 | 15.9174 | 21.5459 | 5.8922 | 135.9146 | 276.4154 | 212.3771 | 271.3135 | 188.6015
2012-10-30 00:34 | 26.0194 | 297.963 | 234.0072 | 292.8572 | 209.8222 | 85.4673 | 1.4434 | 36.5198 | 52.7114 | 25.3018 | 2.2876 | 36.4124 | 83.2966 | 15.9174 | 21.5116 | 5.8921 | 136.008 | 276.4514 | 212.4956 | 271.3456 | 188.3106
2012-10-30 00:34 | 26.0216 | 297.9767 | 233.9544 | 292.8248 | 209.6315 | 85.469 | 1.4438 | 36.5323 | 52.7467 | 25.3022 | 2.2883 | 36.4246 | 83.3523 | 15.9174 | 21.5259 | 5.8922 | 136.0989 | 276.4508 | 212.4285 | 271.2989 | 188.1056
2012-10-30 00:34 | 26.0104 | 298.0001 | 233.9354 | 292.8175 | 209.4549 | 85.4794 | 1.4441 | 36.5395 | 52.7676 | 25.3022 | 2.2888 | 36.4322 | 83.3865 | 15.9175 | 21.5227 | 5.8929 | 136.1541 | 276.4774 | 212.4127 | 271.2948 | 187.9322
2012-10-30 00:34 | 26.0186 | 298.0351 | 233.9632 | 292.7949 | 209.5655 | 85.4748 | 1.4443 | 36.5448 | 52.7826 | 25.3023 | 2.2891 | 36.437 | 83.4091 | 15.9174 | 21.5389 | 5.8926 | 136.1918 | 276.4962 | 212.4243 | 271.256 | 188.0266
2012-10-30 00:34 | 26.0164 | 297.9907 | 233.9414 | 292.8081 | 209.6499 | 85.4644 | 1.4444 | 36.5475 | 52.7898 | 25.3026 | 2.2893 | 36.4398 | 83.4216 | 15.9174 | 21.5688 | 5.8919 | 136.2114 | 276.4219 | 212.3726 | 271.2393 | 188.0811
2012-10-30 00:34 | 26.0081 | 297.9862 | 233.9448 | 292.8113 | 209.6613 | 85.4725 | 1.4445 | 36.5484 | 52.7929 | 25.3023 | 2.2894 | 36.4406 | 83.4256 | 15.9174 | 21.5084 | 5.8925 | 136.2186 | 276.4778 | 212.4364 | 271.3029 | 188.1529
2012-10-30 00:34 | 26.0096 | 298.03 | 233.9385 | 292.8244 | 209.7573 | 85.4759 | 1.4445 | 36.5485 | 52.7934 | 25.3023 | 2.2894 | 36.4408 | 83.4267 | 15.9174 | 21.546 | 5.8927 | 136.2201 | 276.484 | 212.3925 | 271.2784 | 188.2113
2012-10-30 00:34 | 26.0111 | 297.9854 | 233.9203 | 292.7835 | 209.8139 | 85.48 | 1.4445 | 36.5495 | 52.7962 | 25.3023 | 2.2894 | 36.4418 | 83.4315 | 15.9173 | 21.5435 | 5.893 | 136.2277 | 276.4419 | 212.3768 | 271.24 | 188.2704
2012-10-30 00:34 | 26.0134 | 297.9685 | 233.9619 | 292.7627 | 209.9777 | 85.4788 | 1.4445 | 36.5498 | 52.7968 | 25.3025 | 2.2894 | 36.4419 | 83.4318 | 15.9174 | 21.5137 | 5.8929 | 136.2285 | 276.4548 | 212.4482 | 271.249 | 188.464
2012-10-30 00:34 | 25.9999 | 297.9473 | 233.9641 | 292.8147 | 210.2083 | 85.4817 | 1.4445 | 36.5492 | 52.7953 | 25.3023 | 2.2895 | 36.4416 | 83.4314 | 15.9171 | 21.5002 | 5.8931 | 136.2266 | 276.4471 | 212.4639 | 271.3145 | 188.7081
2012-10-30 00:34 | 26.0006 | 297.9557 | 234.0395 | 292.8154 | 210.1775 | 85.4742 | 1.4445 | 36.5489 | 52.7947 | 25.3022 | 2.2894 | 36.4412 | 83.4296 | 15.9171 | 21.5169 | 5.8926 | 136.2244 | 276.4388 | 212.5226 | 271.2985 | 188.6606
2012-10-30 00:34 | 26.0051 | 298.014 | 234.0754 | 292.7852 | 210.0836 | 85.4811 | 1.4445 | 36.5485 | 52.7937 | 25.3021 | 2.2894 | 36.4407 | 83.4277 | 15.9171 | 21.5455 | 5.8931 | 136.2214 | 276.4685 | 212.5299 | 271.2397 | 188.5381
2012-10-30 00:34 | 26.0051 | 297.941 | 234.0833 | 292.8199 | 210.0363 | 85.4829 | 1.4445 | 36.5479 | 52.792 | 25.3021 | 2.2894 | 36.4401 | 83.4249 | 15.9171 | 21.4572 | 5.8932 | 136.217 | 276.4838 | 212.6261 | 271.3627 | 188.5791
2012-10-30 00:34 | 25.9894 | 297.9675 | 234.0911 | 292.7849 | 210.0128 | 85.4701 | 1.4444 | 36.5471 | 52.7895 | 25.3022 | 2.2893 | 36.4396 | 83.4224 | 15.9171 | 21.4776 | 5.8923 | 136.2119 | 276.4899 | 212.6135 | 271.3073 | 188.5352
2012-10-30 00:34 | 26.0021 | 297.9649 | 234.041 | 292.863 | 210.1357 | 85.4852 | 1.4444 | 36.546 | 52.7858 | 25.3024 | 2.2893 | 36.4383 | 83.4164 | 15.9171 | 21.5064 | 5.8933 | 136.2022 | 276.4585 | 212.5346 | 271.3566 | 188.6293
2012-10-30 00:34 | 26.0029 | 298.0079 | 234.085 | 292.81 | 210.0656 | 85.4823 | 1.4443 | 36.5438 | 52.7791 | 25.3027 | 2.2891 | 36.4362 | 83.4058 | 15.9173 | 21.463 | 5.8931 | 136.185 | 276.5449 | 212.622 | 271.347 | 188.6026
2012-10-30 00:34 | 25.9999 | 298.0203 | 234.1016 | 292.8185 | 210.1453 | 85.4834 | 1.4443 | 36.5441 | 52.7801 | 25.3026 | 2.2891 | 36.4363 | 83.4066 | 15.9173 | 21.4961 | 5.8932 | 136.1867 | 276.5242 | 212.6055 | 271.3224 | 188.6492
2012-10-30 00:34 | 25.9947 | 298.0228 | 234.1239 | 292.8479 | 210.2622 | 85.4927 | 1.4443 | 36.5434 | 52.7779 | 25.3026 | 2.2891 | 36.4357 | 83.4039 | 15.9172 | 21.4789 | 5.8939 | 136.1818 | 276.5439 | 212.645 | 271.369 | 188.7833
2012-10-30 00:35 | 26.0006 | 297.9826 | 234.1613 | 292.8346 | 210.15 | 85.4875 | 1.4442 | 36.5424 | 52.7748 | 25.3027 | 2.289 | 36.4346 | 83.399 | 15.9172 | 21.5049 | 5.8935 | 136.1738 | 276.4777 | 212.6564 | 271.3297 | 188.6451
2012-10-30 00:35 | 26.0059 | 298.0109 | 234.1391 | 292.8553 | 210.1041 | 85.4858 | 1.4442 | 36.5425 | 52.7748 | 25.3029 | 2.289 | 36.4346 | 83.3998 | 15.9171 | 21.466 | 5.8934 | 136.1746 | 276.5449 | 212.6731 | 271.3893 | 188.6381
2012-10-30 00:35 | 25.9999 | 297.9972 | 234.1488 | 292.8262 | 210.0981 | 85.4846 | 1.4442 | 36.5413 | 52.7722 | 25.3024 | 2.289 | 36.4337 | 83.3952 | 15.9171 | 21.46 | 5.8933 | 136.1674 | 276.5372 | 212.6888 | 271.3662 | 188.6381
2012-10-30 00:35 | 25.9999 | 298.0049 | 234.1684 | 292.8493 | 210.1808 | 85.499 | 1.4441 | 36.5404 | 52.7695 | 25.3025 | 2.2889 | 36.4327 | 83.3912 | 15.917 | 21.4801 | 5.8943 | 136.1608 | 276.5248 | 212.6883 | 271.3692 | 188.7007
2012-10-30 00:35 | 26.0059 | 298.0378 | 234.1666 | 292.8668 | 210.2419 | 85.4967 | 1.4441 | 36.5392 | 52.7662 | 25.3024 | 2.2888 | 36.4315 | 83.3861 | 15.9169 | 21.4981 | 5.8941 | 136.1523 | 276.5397 | 212.6685 | 271.3687 | 188.7438
2012-10-30 00:35 | 26.0051 | 298.0024 | 234.1775 | 292.8045 | 210.2805 | 85.4829 | 1.444 | 36.5378 | 52.7616 | 25.3026 | 2.2888 | 36.43 | 83.3794 | 15.9169 | 21.4532 | 5.8932 | 136.1411 | 276.5492 | 212.7243 | 271.3513 | 188.8273
2012-10-30 00:35 | 25.9954 | 298.0388 | 234.1639 | 292.814 | 210.3023 | 85.4852 | 1.444 | 36.5373 | 52.7603 | 25.3027 | 2.2887 | 36.4297 | 83.3783 | 15.9169 | 21.4716 | 5.8933 | 136.1386 | 276.5672 | 212.6923 | 271.3424 | 188.8307
2012-10-30 00:35 | 26.0029 | 298.0271 | 234.2421 | 292.8177 | 210.3413 | 85.4933 | 1.444 | 36.5365 | 52.7574 | 25.3029 | 2.2887 | 36.429 | 83.3749 | 15.9169 | 21.4831 | 5.8939 | 136.1324 | 276.544 | 212.759 | 271.3346 | 188.8582
2012-10-30 00:35 | 26.0141 | 298.0537 | 234.328 | 292.8519 | 210.2462 | 85.4985 | 1.444 | 36.5361 | 52.7574 | 25.3024 | 2.2887 | 36.4284 | 83.3729 | 15.9168 | 21.4903 | 5.8943 | 136.1303 | 276.5634 | 212.8377 | 271.3616 | 188.7559
2012-10-30 00:35 | 26.0096 | 298.0607 | 234.2253 | 292.8821 | 210.1826 | 85.4846 | 1.4439 | 36.5347 | 52.7526 | 25.3028 | 2.2886 | 36.4271 | 83.3673 | 15.9167 | 21.4657 | 5.8933 | 136.1198 | 276.595 | 212.7596 | 271.4164 | 188.7169
2012-10-30 00:35 | 26.0036 | 298.0701 | 234.2468 | 292.8607 | 210.3302 | 85.499 | 1.4439 | 36.5344 | 52.7529 | 25.3021 | 2.2886 | 36.4267 | 83.3664 | 15.9166 | 21.4316 | 5.8943 | 136.1193 | 276.6385 | 212.8152 | 271.4291 | 188.8986
2012-10-30 00:35 | 26.0104 | 298.0846 | 234.3046 | 292.8175 | 210.4157 | 85.499 | 1.4439 | 36.5333 | 52.7489 | 25.3025 | 2.2885 | 36.4255 | 83.3613 | 15.9165 | 21.4545 | 5.8943 | 136.1102 | 276.6301 | 212.8501 | 271.363 | 188.9612
2012-10-30 00:35 | 26.0119 | 298.063 | 234.259 | 292.8459 | 210.2361 | 85.4938 | 1.4438 | 36.532 | 52.7456 | 25.3024 | 2.2885 | 36.4243 | 83.3557 | 15.9165 | 21.4921 | 5.8939 | 136.1014 | 276.5709 | 212.7669 | 271.3538 | 188.744
2012-10-30 00:35 | 26.0126 | 298.033 | 234.2675 | 292.8466 | 210.225 | 85.5227 | 1.4438 | 36.5306 | 52.7418 | 25.3023 | 2.2884 | 36.4229 | 83.3494 | 15.9165 | 21.4567 | 5.8959 | 136.0911 | 276.5763 | 212.8108 | 271.3899 | 188.7683
2012-10-30 00:35 | 26.0111 | 298.0814 | 234.2935 | 292.8528 | 210.2392 | 85.5008 | 1.4437 | 36.53 | 52.74 | 25.3022 | 2.2884 | 36.4222 | 83.3471 | 15.9163 | 21.4552 | 5.8944 | 136.0872 | 276.6262 | 212.8383 | 271.3976 | 188.784
2012-10-30 00:35 | 26.0081 | 298.0823 | 234.2788 | 292.8498 | 210.1378 | 85.4962 | 1.4437 | 36.5286 | 52.7362 | 25.3021 | 2.2883 | 36.421 | 83.3415 | 15.9164 | 21.4642 | 5.8941 | 136.0777 | 276.6181 | 212.8146 | 271.3856 | 188.6736
2012-10-30 00:35 | 26.0126 | 298.0791 | 234.2793 | 292.8697 | 210.1895 | 85.5054 | 1.4437 | 36.5278 | 52.7341 | 25.3021 | 2.2882 | 36.4202 | 83.3372 | 15.9164 | 21.4326 | 5.8947 | 136.0713 | 276.6465 | 212.8467 | 271.4371 | 188.7569
2012-10-30 00:35 | 26.0074 | 298.12 | 234.2977 | 292.8606 | 210.3104 | 85.5094 | 1.4436 | 36.5267 | 52.7308 | 25.3022 | 2.2882 | 36.4192 | 83.3328 | 15.9164 | 21.4675 | 5.895 | 136.0636 | 276.6525 | 212.8302 | 271.3931 | 188.8429
2012-10-30 00:35 | 26.0051 | 298.1177 | 234.3111 | 292.9122 | 210.186 | 85.4927 | 1.4436 | 36.5254 | 52.7266 | 25.3024 | 2.2881 | 36.4177 | 83.3269 | 15.9162 | 21.4532 | 5.8939 | 136.0535 | 276.6645 | 212.8579 | 271.459 | 188.7328
2012-10-30 00:35 | 26.0179 | 298.1266 | 234.3121 | 292.8827 | 210.2539 | 85.5106 | 1.4435 | 36.5238 | 52.7222 | 25.3023 | 2.288 | 36.4162 | 83.3196 | 15.9163 | 21.4539 | 5.8951 | 136.0418 | 276.6727 | 212.8582 | 271.4288 | 188.8
2012-10-30 00:35 | 26.0164 | 298.1367 | 234.2792 | 292.9043 | 210.3745 | 85.5025 | 1.4435 | 36.5229 | 52.7193 | 25.3023 | 2.2879 | 36.4152 | 83.3152 | 15.9163 | 21.4484 | 5.8945 | 136.0346 | 276.6883 | 212.8308 | 271.4559 | 188.9261
2012-10-30 00:35 | 26.0156 | 298.132 | 234.3491 | 292.915 | 210.3304 | 85.5031 | 1.4435 | 36.5227 | 52.719 | 25.3022 | 2.2879 | 36.4151 | 83.3143 | 15.9163 | 21.4556 | 5.8946 | 136.0333 | 276.6764 | 212.8935 | 271.4594 | 188.8748
2012-10-30 00:35 | 26.0201 | 298.1365 | 234.3575 | 292.8887 | 210.2758 | 85.506 | 1.4434 | 36.5214 | 52.7155 | 25.3021 | 2.2879 | 36.4139 | 83.3101 | 15.9161 | 21.4682 | 5.8948 | 136.0256 | 276.6683 | 212.8893 | 271.4205 | 188.8076
2012-10-30 00:35 | 26.0231 | 298.0973 | 234.3252 | 292.8802 | 210.1213 | 85.5112 | 1.4402 | 36.4387 | 52.4792 | 25.3011 | 2.2829 | 36.3349 | 82.9497 | 15.916 | 21.4551 | 5.8951 | 135.4289 | 276.6422 | 212.8701 | 271.4251 | 188.6662
2012-10-30 00:36 | 26.0141 | 298.0959 | 234.3201 | 292.8904 | 210.0453 | 85.5071 | 1.4395 | 36.4205 | 52.4274 | 25.3008 | 2.2816 | 36.313 | 82.8504 | 15.9159 | 21.4903 | 5.8949 | 135.2778 | 276.6056 | 212.8298 | 271.4001 | 188.555
2012-10-30 00:36 | 26.0209 | 298.122 | 234.2483 | 292.9049 | 209.8277 | 85.4956 | 1.439 | 36.4085 | 52.3931 | 25.3006 | 2.2808 | 36.301 | 82.796 | 15.9158 | 21.5051 | 5.8941 | 135.1891 | 276.6169 | 212.7432 | 271.3998 | 188.3226
2012-10-30 00:36 | 26.0156 | 298.0552 | 234.1605 | 292.8688 | 208.6686 | 85.4944 | 1.4387 | 36.4008 | 52.3717 | 25.3003 | 2.2804 | 36.2936 | 82.7626 | 15.9158 | 21.4757 | 5.894 | 135.1343 | 276.5795 | 212.6848 | 271.3931 | 187.1929
2012-10-30 00:36 | 26.0164 | 298.0829 | 233.9217 | 292.8889 | 208.7285 | 85.4811 | 1.4387 | 36.3985 | 52.3656 | 25.3001 | 2.2802 | 36.2914 | 82.7521 | 15.9158 | 21.4524 | 5.8931 | 135.1178 | 276.6305 | 212.4693 | 271.4365 | 187.2761
2012-10-30 00:36 | 26.0223 | 298.0465 | 233.8569 | 292.8678 | 209.2069 | 85.4771 | 1.4388 | 36.4033 | 52.3783 | 25.3005 | 2.2805 | 36.296 | 82.7732 | 15.9158 | 21.4704 | 5.8928 | 135.1515 | 276.5761 | 212.3865 | 271.3974 | 187.7365
2012-10-30 00:36 | 26.0134 | 298.0799 | 233.8284 | 292.8435 | 209.4737 | 85.4759 | 1.439 | 36.4083 | 52.3932 | 25.3003 | 2.2808 | 36.3011 | 82.7956 | 15.9159 | 21.4976 | 5.8927 | 135.1888 | 276.5823 | 212.3308 | 271.3459 | 187.9761
2012-10-30 00:36 | 26.0201 | 298.0443 | 233.7722 | 292.8656 | 209.2362 | 85.48 | 1.4391 | 36.4098 | 52.3973 | 25.3004 | 2.2809 | 36.3024 | 82.8021 | 15.9159 | 21.4561 | 5.893 | 135.1995 | 276.5882 | 212.3161 | 271.4095 | 187.7801
2012-10-30 00:36 | 26.0223 | 298.0427 | 233.7587 | 292.8332 | 209.1557 | 85.4834 | 1.4393 | 36.4154 | 52.4135 | 25.3004 | 2.2813 | 36.3081 | 82.8281 | 15.9158 | 21.4543 | 5.8932 | 135.2416 | 276.5884 | 212.3044 | 271.3789 | 187.7014
2012-10-30 00:36 | 26.0171 | 298.026 | 233.6906 | 292.8242 | 209.4852 | 85.4661 | 1.4404 | 36.444 | 52.4947 | 25.3009 | 2.2831 | 36.338 | 82.9633 | 15.916 | 21.4732 | 5.892 | 135.4581 | 276.5528 | 212.2174 | 271.351 | 188.012
2012-10-30 00:36 | 26.0164 | 298.0445 | 233.6781 | 292.7811 | 209.7287 | 85.4713 | 1.4426 | 36.5006 | 52.6569 | 25.3015 | 2.2866 | 36.3932 | 83.2155 | 15.9161 | 21.4966 | 5.8924 | 135.8724 | 276.5479 | 212.1815 | 271.2845 | 188.2321
2012-10-30 00:36 | 26.0149 | 298.0161 | 233.6845 | 292.8297 | 209.546 | 85.4765 | 1.4432 | 36.515 | 52.6981 | 25.3016 | 2.2875 | 36.4074 | 83.2803 | 15.9161 | 21.5392 | 5.8927 | 135.9784 | 276.4769 | 212.1453 | 271.2905 | 188.0068
2012-10-30 00:36 | 26.0051 | 298.0447 | 233.7297 | 292.8083 | 209.6426 | 85.4806 | 1.4437 | 36.5289 | 52.7375 | 25.3019 | 2.2883 | 36.421 | 83.3428 | 15.9161 | 21.5254 | 5.893 | 136.0804 | 276.5193 | 212.2043 | 271.2829 | 188.1172
2012-10-30 00:36 | 26.0036 | 298.0125 | 233.7439 | 292.7645 | 209.1843 | 85.4777 | 1.4441 | 36.5381 | 52.7639 | 25.302 | 2.2889 | 36.4303 | 83.3853 | 15.9161 | 21.54 | 5.8928 | 136.1492 | 276.4725 | 212.2039 | 271.2245 | 187.6443
2012-10-30 00:36 | 26.0044 | 298.0171 | 233.6819 | 292.7768 | 208.7913 | 85.4852 | 1.4443 | 36.5444 | 52.782 | 25.3021 | 2.2893 | 36.4367 | 83.4148 | 15.9161 | 21.5528 | 5.8933 | 136.1967 | 276.4643 | 212.1291 | 271.224 | 187.2385
2012-10-30 00:36 | 26.0006 | 298.0248 | 233.6152 | 292.7923 | 209.3427 | 85.4886 | 1.4445 | 36.5482 | 52.7926 | 25.3022 | 2.2895 | 36.4405 | 83.4321 | 15.916 | 21.557 | 5.8936 | 136.2247 | 276.4678 | 212.0582 | 271.2353 | 187.7857
2012-10-30 00:36 | 26.0021 | 298.0263 | 233.5774 | 292.763 | 209.6277 | 85.4904 | 1.4445 | 36.5498 | 52.7972 | 25.3023 | 2.2897 | 36.4421 | 83.44 | 15.9159 | 21.5304 | 5.8937 | 136.2372 | 276.4959 | 212.047 | 271.2326 | 188.0973
2012-10-30 00:36 | 25.9962 | 297.9974 | 233.6776 | 292.7763 | 209.5234 | 85.4863 | 1.4446 | 36.5506 | 52.7998 | 25.3022 | 2.2897 | 36.4428 | 83.443 | 15.916 | 21.5326 | 5.8934 | 136.2427 | 276.4648 | 212.145 | 271.2437 | 187.9908
2012-10-30 00:36 | 25.9932 | 298.0213 | 233.6589 | 292.7849 | 209.6228 | 85.4927 | 1.4445 | 36.5507 | 52.7992 | 25.3025 | 2.2897 | 36.4429 | 83.4436 | 15.916 | 21.5416 | 5.8939 | 136.2428 | 276.4797 | 212.1173 | 271.2433 | 188.0812
2012-10-30 00:36 | 25.9902 | 298.0183 | 233.703 | 292.7588 | 209.7694 | 85.4938 | 1.4445 | 36.5503 | 52.7983 | 25.3024 | 2.2897 | 36.4426 | 83.442 | 15.916 | 21.5306 | 5.8939 | 136.2403 | 276.4877 | 212.1724 | 271.2282 | 188.2388
2012-10-30 00:36 | 25.9917 | 298.0236 | 233.7399 | 292.7641 | 209.9245 | 85.4979 | 1.4445 | 36.5496 | 52.7961 | 25.3025 | 2.2896 | 36.4417 | 83.4387 | 15.9159 | 21.512 | 5.8942 | 136.2349 | 276.5116 | 212.2279 | 271.2521 | 188.4125
2012-10-30 00:36 | 25.9849 | 297.9976 | 233.8431 | 292.7458 | 209.9531 | 85.5014 | 1.4445 | 36.5494 | 52.7957 | 25.3024 | 2.2896 | 36.4416 | 83.4383 | 15.9159 | 21.5173 | 5.8945 | 136.234 | 276.4803 | 212.3258 | 271.2285 | 188.4358
2012-10-30 00:37 | 25.9812 | 298.0362 | 233.9494 | 292.7459 | 210.0558 | 85.4973 | 1.4444 | 36.546 | 52.7862 | 25.3023 | 2.2895 | 36.4381 | 83.4235 | 15.9156 | 21.5055 | 5.8942 | 136.2097 | 276.5307 | 212.4439 | 271.2404 | 188.5503
2012-10-30 00:37 | 25.9745 | 297.9526 | 233.927 | 292.7315 | 210.0451 | 85.506 | 1.4443 | 36.5453 | 52.7836 | 25.3025 | 2.2894 | 36.4374 | 83.4198 | 15.9157 | 21.4707 | 5.8948 | 136.2034 | 276.4819 | 212.4563 | 271.2608 | 188.5744
2012-10-30 00:37 | 25.9797 | 298.0039 | 233.9597 | 292.756 | 210.1055 | 85.5002 | 1.4443 | 36.5442 | 52.7803 | 25.3026 | 2.2893 | 36.4363 | 83.4148 | 15.9157 | 21.484 | 5.8944 | 136.1951 | 276.5199 | 212.4757 | 271.272 | 188.6215
2012-10-30 00:37 | 25.9842 | 298.0276 | 234.0152 | 292.7528 | 210.1296 | 85.5204 | 1.4443 | 36.5433 | 52.7782 | 25.3024 | 2.2893 | 36.4356 | 83.4116 | 15.9157 | 21.5085 | 5.8958 | 136.1898 | 276.5191 | 212.5067 | 271.2443 | 188.6211
2012-10-30 00:37 | 25.9857 | 298.0215 | 234.0128 | 292.7697 | 210.0248 | 85.5031 | 1.4442 | 36.5424 | 52.7754 | 25.3024 | 2.2892 | 36.4344 | 83.4064 | 15.9157 | 21.5461 | 5.8946 | 136.1818 | 276.4754 | 212.4667 | 271.2236 | 188.4787
2012-10-30 00:37 | 25.9842 | 297.9854 | 234.0231 | 292.7643 | 210.1139 | 85.5048 | 1.4442 | 36.541 | 52.7712 | 25.3025 | 2.2891 | 36.4332 | 83.4003 | 15.9158 | 21.5286 | 5.8947 | 136.1715 | 276.4568 | 212.4945 | 271.2357 | 188.5853
2012-10-30 00:37 | 25.9745 | 297.9834 | 234.0448 | 292.7623 | 210.0136 | 85.5141 | 1.4441 | 36.5401 | 52.7685 | 25.3025 | 2.2891 | 36.4323 | 83.3966 | 15.9157 | 21.5149 | 5.8953 | 136.1651 | 276.4685 | 212.5299 | 271.2474 | 188.4987
2012-10-30 00:37 | 25.9857 | 297.9907 | 234.0521 | 292.7312 | 210.1272 | 85.5014 | 1.4441 | 36.5395 | 52.7668 | 25.3026 | 2.2891 | 36.4318 | 83.395 | 15.9155 | 21.51 | 5.8945 | 136.1618 | 276.4807 | 212.5421 | 271.2212 | 188.6172
2012-10-30 00:37 | 25.9849 | 298.0168 | 234.0395 | 292.765 | 210.1697 | 85.5077 | 1.4441 | 36.5388 | 52.7649 | 25.3026 | 2.289 | 36.4311 | 83.3915 | 15.9156 | 21.4811 | 5.8949 | 136.1564 | 276.5357 | 212.5584 | 271.2839 | 188.6886
2012-10-30 00:37 | 25.9804 | 298.0162 | 234.1293 | 292.7374 | 210.0392 | 85.5112 | 1.444 | 36.5373 | 52.7605 | 25.3025 | 2.2889 | 36.4296 | 83.3844 | 15.9156 | 21.5168 | 5.8951 | 136.1449 | 276.4994 | 212.6125 | 271.2206 | 188.5224
2012-10-30 00:37 | 25.9924 | 298.0128 | 234.1059 | 292.7725 | 210.0315 | 85.5187 | 1.444 | 36.5362 | 52.7574 | 25.3025 | 2.2889 | 36.4285 | 83.3794 | 15.9156 | 21.5328 | 5.8957 | 136.1369 | 276.48 | 212.5731 | 271.2397 | 188.4987
2012-10-30 00:37 | 25.9834 | 297.9999 | 234.0773 | 292.7597 | 210.0186 | 85.5129 | 1.4439 | 36.5351 | 52.7539 | 25.3027 | 2.2888 | 36.4272 | 83.3734 | 15.9156 | 21.5037 | 5.8953 | 136.1273 | 276.4962 | 212.5736 | 271.256 | 188.5149
2012-10-30 00:37 | 25.9924 | 298.0128 | 234.1256 | 292.7687 | 209.9134 | 85.5019 | 1.4439 | 36.5336 | 52.7502 | 25.3024 | 2.2887 | 36.4258 | 83.368 | 15.9155 | 21.4886 | 5.8945 | 136.1182 | 276.5242 | 212.637 | 271.2801 | 188.4248
2012-10-30 00:37 | 25.9879 | 298.0275 | 234.1054 | 292.7873 | 210.0506 | 85.521 | 1.4439 | 36.5333 | 52.7487 | 25.3027 | 2.2887 | 36.4255 | 83.3661 | 15.9156 | 21.4761 | 5.8958 | 136.1148 | 276.5514 | 212.6293 | 271.3112 | 188.5745
2012-10-30 00:37 | 25.9939 | 298.0757 | 234.1192 | 292.7471 | 210.0094 | 85.5152 | 1.4438 | 36.5316 | 52.7433 | 25.3029 | 2.2886 | 36.4238 | 83.3585 | 15.9155 | 21.4901 | 5.8954 | 136.1019 | 276.5856 | 212.6291 | 271.257 | 188.5193
2012-10-30 00:37 | 25.9924 | 298.0358 | 234.0981 | 292.7225 | 210.0276 | 85.5014 | 1.4437 | 36.5305 | 52.7403 | 25.3029 | 2.2885 | 36.4226 | 83.3531 | 15.9155 | 21.5127 | 5.8945 | 136.0933 | 276.5231 | 212.5854 | 271.2098 | 188.5149
2012-10-30 00:37 | 25.9999 | 298.028 | 234.1566 | 292.7685 | 210.232 | 85.5031 | 1.4437 | 36.5297 | 52.7379 | 25.3028 | 2.2885 | 36.4219 | 83.3499 | 15.9155 | 21.4761 | 5.8946 | 136.0878 | 276.5519 | 212.6805 | 271.2924 | 188.7559
2012-10-30 00:37 | 25.9894 | 298.0252 | 234.1461 | 292.7657 | 210.2727 | 85.5222 | 1.4437 | 36.529 | 52.7363 | 25.3026 | 2.2884 | 36.4214 | 83.3476 | 15.9155 | 21.4736 | 5.8959 | 136.0839 | 276.5516 | 212.6725 | 271.2921 | 188.7991
2012-10-30 00:37 | 25.9992 | 298.035 | 234.1952 | 292.7909 | 210.247 | 85.5152 | 1.4436 | 36.5275 | 52.7318 | 25.3027 | 2.2883 | 36.4197 | 83.3397 | 15.9155 | 21.4954 | 5.8954 | 136.0715 | 276.5396 | 212.6998 | 271.2955 | 188.7516
2012-10-30 00:37 | 25.9992 | 298.0426 | 234.2266 | 292.7678 | 210.314 | 85.5083 | 1.4436 | 36.5264 | 52.7289 | 25.3026 | 2.2883 | 36.4188 | 83.3362 | 15.9154 | 21.4714 | 5.8949 | 136.0651 | 276.5712 | 212.7552 | 271.2964 | 188.8426
2012-10-30 00:37 | 26.0014 | 298.0295 | 234.2367 | 292.8008 | 210.3359 | 85.5118 | 1.4435 | 36.5246 | 52.7237 | 25.3026 | 2.2882 | 36.4169 | 83.3281 | 15.9153 | 21.5057 | 5.8952 | 136.0518 | 276.5238 | 212.731 | 271.2951 | 188.8302
2012-10-30 00:37 | 26.0089 | 298.0639 | 234.2835 | 292.816 | 210.4851 | 85.5129 | 1.4435 | 36.5233 | 52.7201 | 25.3025 | 2.2881 | 36.4156 | 83.3231 | 15.9151 | 21.5453 | 5.8953 | 136.0432 | 276.5186 | 212.7382 | 271.2707 | 188.9398
2012-10-30 00:37 | 26.0111 | 298.0545 | 234.2896 | 292.8259 | 210.3377 | 85.5158 | 1.4435 | 36.5233 | 52.7197 | 25.3027 | 2.2881 | 36.4155 | 83.3219 | 15.9153 | 21.5154 | 5.8955 | 136.0416 | 276.5391 | 212.7742 | 271.3105 | 188.8223
2012-10-30 00:37 | 26.0119 | 298.063 | 234.3297 | 292.8036 | 210.3542 | 85.5141 | 1.4434 | 36.5218 | 52.7165 | 25.3022 | 2.288 | 36.414 | 83.3152 | 15.9152 | 21.5242 | 5.8953 | 136.0317 | 276.5388 | 212.8055 | 271.2794 | 188.83
2012-10-30 00:38 | 26.0089 | 297.987 | 234.311 | 292.8044 | 210.2646 | 85.5146 | 1.4434 | 36.5208 | 52.7128 | 25.3025 | 2.2879 | 36.4131 | 83.311 | 15.9152 | 21.5292 | 5.8954 | 136.0238 | 276.4578 | 212.7818 | 271.2752 | 188.7354
2012-10-30 00:38 | 26.0044 | 298.0786 | 234.33 | 292.8153 | 210.2444 | 85.5193 | 1.4424 | 36.4969 | 52.6444 | 25.3023 | 2.2867 | 36.3928 | 83.2189 | 15.9151 | 21.5006 | 5.8957 | 135.8633 | 276.578 | 212.8294 | 271.3147 | 188.7438
2012-10-30 00:38 | 26.0104 | 298.0769 | 234.2653 | 292.8559 | 210.0062 | 85.5019 | 1.4394 | 36.4194 | 52.4232 | 25.3012 | 2.2816 | 36.3118 | 82.8504 | 15.9148 | 21.5387 | 5.8945 | 135.2737 | 276.5382 | 212.7266 | 271.3172 | 188.4675
2012-10-30 00:38 | 26.0036 | 298.0009 | 234.2507 | 292.8299 | 209.8419 | 85.5031 | 1.439 | 36.4088 | 52.3939 | 25.3008 | 2.281 | 36.3017 | 82.8045 | 15.9148 | 21.5199 | 5.8946 | 135.1983 | 276.481 | 212.7308 | 271.31 | 188.322
2012-10-30 00:38 | 26.0021 | 297.9994 | 234.206 | 292.7822 | 209.8955 | 85.5025 | 1.4388 | 36.4024 | 52.3753 | 25.3008 | 2.2806 | 36.2951 | 82.7744 | 15.9148 | 21.5465 | 5.8945 | 135.1497 | 276.4529 | 212.6595 | 271.2357 | 188.349
2012-10-30 00:38 | 26.0044 | 298.0171 | 234.1611 | 292.8153 | 209.5867 | 85.5019 | 1.4387 | 36.4004 | 52.3698 | 25.3006 | 2.2805 | 36.2931 | 82.7656 | 15.9147 | 21.5327 | 5.8945 | 135.1354 | 276.4844 | 212.6284 | 271.2826 | 188.054
2012-10-30 00:38 | 26.0051 | 298.0063 | 234.099 | 292.8045 | 210.3671 | 85.4846 | 1.4388 | 36.4014 | 52.3732 | 25.3004 | 2.2806 | 36.2942 | 82.7707 | 15.9146 | 21.5334 | 5.8933 | 135.1439 | 276.4729 | 212.5656 | 271.2711 | 188.8337
2012-10-30 00:38 | 26.0006 | 297.9326 | 234.1338 | 292.7961 | 210.4611 | 85.4794 | 1.4389 | 36.4045 | 52.3823 | 25.3003 | 2.2807 | 36.2972 | 82.7848 | 15.9146 | 21.5089 | 5.8929 | 135.1671 | 276.4237 | 212.6249 | 271.2872 | 188.9522
2012-10-30 00:38 | 26.0021 | 297.9572 | 234.1628 | 292.8053 | 210.3799 | 85.4869 | 1.439 | 36.4081 | 52.3923 | 25.3005 | 2.281 | 36.3008 | 82.8012 | 15.9146 | 21.5224 | 5.8935 | 135.1935 | 276.4348 | 212.6404 | 271.2829 | 188.8575
2012-10-30 00:38 | 26.0059 | 297.9687 | 234.1666 | 292.7976 | 210.2892 | 85.4834 | 1.4392 | 36.4126 | 52.4049 | 25.3006 | 2.2812 | 36.305 | 82.8208 | 15.9146 | 21.5062 | 5.8932 | 135.2257 | 276.4625 | 212.6604 | 271.2914 | 188.783
2012-10-30 00:38 | 25.9932 | 297.9483 | 234.1067 | 292.7579 | 209.5677 | 85.4707 | 1.4394 | 36.4169 | 52.4174 | 25.3006 | 2.2815 | 36.3094 | 82.8396 | 15.9147 | 21.5456 | 5.8924 | 135.257 | 276.4027 | 212.5611 | 271.2123 | 188.0221
2012-10-30 00:38 | 25.9939 | 297.8798 | 233.9699 | 292.7586 | 209.9661 | 85.4904 | 1.4422 | 36.4902 | 52.627 | 25.3013 | 2.2859 | 36.3797 | 83.1599 | 15.9149 | 21.5624 | 5.8937 | 135.7869 | 276.3174 | 212.4075 | 271.1962 | 188.4037
2012-10-30 00:38 | 25.9917 | 297.9506 | 234.0109 | 292.7603 | 210.019 | 85.4938 | 1.4429 | 36.509 | 52.6802 | 25.3019 | 2.2872 | 36.401 | 83.2574 | 15.9149 | 21.52 | 5.8939 | 135.9375 | 276.4306 | 212.4909 | 271.2403 | 188.499
2012-10-30 00:38 | 25.9872 | 297.9577 | 233.9514 | 292.7288 | 209.7467 | 85.4823 | 1.4434 | 36.5208 | 52.7139 | 25.3021 | 2.288 | 36.4129 | 83.3125 | 15.9148 | 21.5276 | 5.8931 | 136.0263 | 276.4301 | 212.4238 | 271.2012 | 188.2191
2012-10-30 00:38 | 25.9782 | 297.941 | 233.8953 | 292.6929 | 209.5093 | 85.4782 | 1.4439 | 36.5339 | 52.751 | 25.3024 | 2.2888 | 36.4258 | 83.3705 | 15.915 | 21.5306 | 5.8929 | 136.1215 | 276.4104 | 212.3647 | 271.1623 | 187.9787
2012-10-30 00:38 | 25.9819 | 297.9754 | 233.8597 | 292.7043 | 209.8123 | 85.4973 | 1.4442 | 36.5412 | 52.7717 | 25.3026 | 2.2893 | 36.4335 | 83.4053 | 15.915 | 21.5423 | 5.8942 | 136.1771 | 276.4331 | 212.3174 | 271.162 | 188.27
2012-10-30 00:38 | 25.973 | 297.9242 | 233.8312 | 292.6992 | 210.0239 | 85.5031 | 1.4443 | 36.5446 | 52.7813 | 25.3027 | 2.2895 | 36.4368 | 83.4205 | 15.915 | 21.5294 | 5.8946 | 136.2019 | 276.3948 | 212.3018 | 271.1698 | 188.4945
2012-10-30 00:38 | 25.9752 | 297.9264 | 233.9237 | 292.7014 | 210.1915 | 85.4904 | 1.4444 | 36.5476 | 52.7902 | 25.3026 | 2.2896 | 36.4398 | 83.4344 | 15.915 | 21.5236 | 5.8937 | 136.2245 | 276.4028 | 212.4001 | 271.1778 | 188.6679
2012-10-30 00:38 | 25.973 | 297.9358 | 233.9451 | 292.6838 | 210.2248 | 85.4823 | 1.4444 | 36.5486 | 52.7925 | 25.3028 | 2.2897 | 36.4407 | 83.4387 | 15.915 | 21.5334 | 5.8931 | 136.2311 | 276.4024 | 212.4117 | 271.1504 | 188.6914
2012-10-30 00:38 | 25.9715 | 297.9343 | 234.0104 | 292.6708 | 210.1445 | 85.5037 | 1.4445 | 36.5488 | 52.7933 | 25.3027 | 2.2897 | 36.4409 | 83.4394 | 15.915 | 21.536 | 5.8946 | 136.2328 | 276.3983 | 212.4744 | 271.1348 | 188.6085
2012-10-30 00:38 | 25.9685 | 297.9466 | 234.031 | 292.6909 | 210.1061 | 85.4846 | 1.4444 | 36.5477 | 52.7896 | 25.303 | 2.2896 | 36.4399 | 83.4344 | 15.9151 | 21.549 | 5.8933 | 136.224 | 276.3976 | 212.482 | 271.1419 | 188.5571
2012-10-30 00:38 | 25.973 | 297.9204 | 234.0276 | 292.6646 | 210.1618 | 85.4985 | 1.4444 | 36.5472 | 52.7879 | 25.3031 | 2.2896 | 36.4393 | 83.4317 | 15.9151 | 21.5495 | 5.8943 | 136.2196 | 276.3709 | 212.4781 | 271.1151 | 188.6123
2012-10-30 00:38 | 25.9745 | 297.9219 | 234.0487 | 292.7046 | 210.3405 | 85.5042 | 1.4444 | 36.5477 | 52.7897 | 25.303 | 2.2896 | 36.4399 | 83.4344 | 15.9151 | 21.543 | 5.8947 | 136.2242 | 276.3789 | 212.5057 | 271.1616 | 188.7975
2012-10-30 00:38 | 25.9692 | 297.9166 | 234.0788 | 292.6608 | 210.343 | 85.5071 | 1.4444 | 36.5473 | 52.7885 | 25.303 | 2.2896 | 36.4395 | 83.4332 | 15.915 | 21.5296 | 5.8949 | 136.2218 | 276.387 | 212.5492 | 271.1312 | 188.8134
2012-10-30 00:39 | 25.967 | 297.899 | 234.1002 | 292.6855 | 210.2542 | 85.5129 | 1.4444 | 36.5473 | 52.7881 | 25.3031 | 2.2896 | 36.4393 | 83.4321 | 15.915 | 21.5315 | 5.8953 | 136.2202 | 276.3675 | 212.5687 | 271.154 | 188.7227
2012-10-30 00:39 | 25.9685 | 297.889 | 234.1488 | 292.7024 | 210.2321 | 85.5077 | 1.4444 | 36.5471 | 52.7879 | 25.303 | 2.2896 | 36.4391 | 83.4307 | 15.9151 | 21.5209 | 5.8949 | 136.2186 | 276.3681 | 212.6279 | 271.1815 | 188.7112
2012-10-30 00:39 | 25.9662 | 297.9251 | 234.1308 | 292.677 | 210.1746 | 85.5008 | 1.4444 | 36.5479 | 52.7898 | 25.3031 | 2.2897 | 36.4401 | 83.4355 | 15.915 | 21.5307 | 5.8944 | 136.2253 | 276.3944 | 212.6001 | 271.1463 | 188.6439
2012-10-30 00:39 | 25.9647 | 297.8852 | 234.1214 | 292.6909 | 210.1653 | 85.5216 | 1.4444 | 36.5468 | 52.7866 | 25.3031 | 2.2896 | 36.439 | 83.4315 | 15.9149 | 21.4971 | 5.8959 | 136.2181 | 276.3881 | 212.6243 | 271.1938 | 188.6682
2012-10-30 00:39 | 25.9647 | 297.9236 | 234.1686 | 292.6909 | 210.2322 | 85.5152 | 1.4443 | 36.5451 | 52.7813 | 25.3034 | 2.2895 | 36.4373 | 83.4239 | 15.9148 | 21.5051 | 5.8954 | 136.2052 | 276.4185 | 212.6635 | 271.1858 | 188.7271
2012-10-30 00:39 | 25.97 | 297.9366 | 234.1896 | 292.6655 | 210.2808 | 85.5106 | 1.4443 | 36.5438 | 52.7784 | 25.303 | 2.2894 | 36.4362 | 83.4178 | 15.915 | 21.4983 | 5.8951 | 136.1962 | 276.4383 | 212.6913 | 271.1672 | 188.7825
2012-10-30 00:39 | 25.9625 | 297.956 | 234.1349 | 292.6772 | 210.2891 | 85.5135 | 1.4443 | 36.5444 | 52.7797 | 25.3032 | 2.2895 | 36.4365 | 83.42 | 15.9148 | 21.5229 | 5.8953 | 136.1997 | 276.4331 | 212.612 | 271.1543 | 188.7662
2012-10-30 00:39 | 25.9647 | 297.9044 | 234.1764 | 292.7179 | 210.2204 | 85.5204 | 1.4442 | 36.5424 | 52.7739 | 25.3031 | 2.2894 | 36.4345 | 83.4117 | 15.9147 | 21.485 | 5.8958 | 136.1856 | 276.4194 | 212.6914 | 271.2329 | 188.7354
2012-10-30 00:39 | 25.964 | 297.9229 | 234.1914 | 292.6979 | 210.263 | 85.5146 | 1.4441 | 36.5411 | 52.7696 | 25.3034 | 2.2893 | 36.4331 | 83.4056 | 15.9147 | 21.4723 | 5.8954 | 136.1753 | 276.4506 | 212.7191 | 271.2256 | 188.7907
2012-10-30 00:39 | 25.9647 | 297.9428 | 234.2118 | 292.7294 | 210.3267 | 85.5146 | 1.4441 | 36.5396 | 52.7666 | 25.3028 | 2.2892 | 36.4318 | 83.4 | 15.9146 | 21.469 | 5.8954 | 136.1666 | 276.4738 | 212.7428 | 271.2604 | 188.8577
2012-10-30 00:39 | 25.9662 | 297.9943 | 234.2172 | 292.7309 | 210.3597 | 85.5256 | 1.444 | 36.5373 | 52.7598 | 25.3029 | 2.2891 | 36.4296 | 83.39 | 15.9146 | 21.4825 | 5.8961 | 136.1499 | 276.5118 | 212.7347 | 271.2484 | 188.8772
2012-10-30 00:39 | 25.9692 | 297.9473 | 234.2399 | 292.7378 | 210.28 | 85.521 | 1.4439 | 36.5354 | 52.7543 | 25.3028 | 2.289 | 36.4276 | 83.3817 | 15.9144 | 21.5016 | 5.8958 | 136.1361 | 276.4457 | 212.7383 | 271.2362 | 188.7784
2012-10-30 00:39 | 25.973 | 297.978 | 234.2358 | 292.757 | 210.1972 | 85.5216 | 1.4439 | 36.5341 | 52.7505 | 25.3029 | 2.2889 | 36.4263 | 83.3759 | 15.9143 | 21.4532 | 5.8959 | 136.1264 | 276.5248 | 212.7826 | 271.3038 | 188.744
2012-10-30 00:39 | 25.973 | 298.0049 | 234.2476 | 292.73 | 210.4374 | 85.5222 | 1.4439 | 36.5341 | 52.7504 | 25.303 | 2.2889 | 36.4264 | 83.3764 | 15.9144 | 21.4813 | 5.8959 | 136.1268 | 276.5236 | 212.7663 | 271.2487 | 188.9561
2012-10-30 00:39 | 25.9737 | 297.9902 | 234.2679 | 292.7654 | 210.6468 | 85.536 | 1.4439 | 36.5333 | 52.7489 | 25.3026 | 2.2889 | 36.4254 | 83.3726 | 15.9142 | 21.4659 | 5.8969 | 136.1215 | 276.5243 | 212.802 | 271.2995 | 189.1809
2012-10-30 00:39 | 25.9692 | 297.9896 | 234.3381 | 292.7416 | 210.5793 | 85.5204 | 1.4438 | 36.5316 | 52.7442 | 25.3025 | 2.2888 | 36.4239 | 83.3653 | 15.9143 | 21.4614 | 5.8958 | 136.1095 | 276.5282 | 212.8767 | 271.2802 | 189.1179
2012-10-30 00:39 | 25.9722 | 298.008 | 234.3528 | 292.7754 | 210.5272 | 85.5239 | 1.4438 | 36.5307 | 52.7413 | 25.3027 | 2.2887 | 36.423 | 83.3614 | 15.9142 | 21.4604 | 5.896 | 136.1026 | 276.5476 | 212.8924 | 271.315 | 189.0668
2012-10-30 00:39 | 25.9662 | 297.9904 | 234.394 | 292.7617 | 210.47 | 85.5129 | 1.4437 | 36.5306 | 52.7409 | 25.3026 | 2.2887 | 36.4226 | 83.3598 | 15.9142 | 21.4424 | 5.8953 | 136.1007 | 276.548 | 212.9516 | 271.3193 | 189.0276
2012-10-30 00:39 | 25.9685 | 298.0311 | 234.357 | 292.8217 | 210.4172 | 85.5227 | 1.4437 | 36.5294 | 52.7371 | 25.3029 | 2.2886 | 36.4219 | 83.3566 | 15.9142 | 21.4888 | 5.8959 | 136.0936 | 276.5423 | 212.8682 | 271.3329 | 188.9284
2012-10-30 00:39 | 25.9647 | 298.0005 | 234.4121 | 292.7871 | 210.4527 | 85.5204 | 1.4436 | 36.5277 | 52.7318 | 25.303 | 2.2885 | 36.4198 | 83.3464 | 15.9143 | 21.5091 | 5.8958 | 136.0781 | 276.4914 | 212.903 | 271.278 | 188.9436
2012-10-30 00:39 | 25.9647 | 298.0389 | 234.3925 | 292.8141 | 210.2204 | 85.5279 | 1.4435 | 36.5248 | 52.7233 | 25.303 | 2.2883 | 36.417 | 83.3336 | 15.9143 | 21.489 | 5.8963 | 136.0569 | 276.5499 | 212.9035 | 271.3251 | 188.7314
2012-10-30 00:40 | 25.9595 | 298.0606 | 234.3794 | 292.7858 | 210.2113 | 85.5251 | 1.4435 | 36.5241 | 52.7217 | 25.3028 | 2.2883 | 36.4162 | 83.3314 | 15.914 | 21.4999 | 5.8961 | 136.0531 | 276.5607 | 212.8795 | 271.2859 | 188.7114
2012-10-30 00:40 | 25.9587 | 298.0175 | 234.3943 | 292.7888 | 210.4901 | 85.5233 | 1.4435 | 36.524 | 52.7216 | 25.3027 | 2.2883 | 36.4162 | 83.331 | 15.9141 | 21.5151 | 5.896 | 136.0526 | 276.5024 | 212.8792 | 271.2737 | 188.975
2012-10-30 00:40 | 25.9535 | 298.0353 | 234.3813 | 292.7798 | 210.5912 | 85.521 | 1.4434 | 36.5228 | 52.7184 | 25.3027 | 2.2882 | 36.4152 | 83.3268 | 15.914 | 21.4818 | 5.8958 | 136.0453 | 276.5535 | 212.8995 | 271.298 | 189.1094
2012-10-30 00:40 | 25.9625 | 298.0559 | 234.4924 | 292.835 | 210.3993 | 85.5198 | 1.4434 | 36.5222 | 52.7166 | 25.3027 | 2.2882 | 36.4146 | 83.3237 | 15.9141 | 21.5029 | 5.8957 | 136.0403 | 276.553 | 212.9895 | 271.3321 | 188.8964
2012-10-30 00:40 | 25.964 | 298.0612 | 234.4586 | 292.798 | 210.3221 | 85.5314 | 1.4434 | 36.5211 | 52.7134 | 25.3028 | 2.2881 | 36.4135 | 83.3192 | 15.914 | 21.5285 | 5.8965 | 136.0326 | 276.5327 | 212.9301 | 271.2695 | 188.7936
2012-10-30 00:40 | 25.9655 | 298.0512 | 234.4247 | 292.8033 | 210.4299 | 85.5308 | 1.442 | 36.4866 | 52.6148 | 25.3023 | 2.2861 | 36.3803 | 83.1677 | 15.9139 | 21.538 | 5.8965 | 135.7825 | 276.5132 | 212.8867 | 271.2653 | 188.8919
2012-10-30 00:40 | 25.9625 | 298.0175 | 234.4335 | 292.8119 | 210.4702 | 85.5048 | 1.4396 | 36.4228 | 52.4329 | 25.3013 | 2.282 | 36.315 | 82.8707 | 15.9137 | 21.5631 | 5.8947 | 135.3036 | 276.4544 | 212.8704 | 271.2488 | 188.9071
2012-10-30 00:40 | 25.9617 | 298.0243 | 234.4366 | 292.8111 | 210.367 | 85.5077 | 1.439 | 36.408 | 52.3906 | 25.3012 | 2.2811 | 36.3006 | 82.8053 | 15.9136 | 21.486 | 5.8949 | 135.1958 | 276.5383 | 212.9506 | 271.3251 | 188.881
2012-10-30 00:40 | 25.9685 | 298.0427 | 234.4081 | 292.8063 | 209.3973 | 85.5106 | 1.4388 | 36.4021 | 52.3738 | 25.3011 | 2.2807 | 36.2946 | 82.7775 | 15.9138 | 21.5129 | 5.8951 | 135.1513 | 276.5298 | 212.8952 | 271.2934 | 187.8844
2012-10-30 00:40 | 25.9632 | 297.9874 | 234.2103 | 292.8087 | 208.88 | 85.5106 | 1.4387 | 36.4007 | 52.3696 | 25.3012 | 2.2806 | 36.2933 | 82.7713 | 15.9138 | 21.5116 | 5.8951 | 135.1409 | 276.4758 | 212.6987 | 271.2971 | 187.3684
2012-10-30 00:40 | 25.9662 | 297.9981 | 234.0994 | 292.8156 | 209.7415 | 85.4996 | 1.4388 | 36.4031 | 52.3762 | 25.3013 | 2.2808 | 36.2954 | 82.781 | 15.9138 | 21.4986 | 5.8943 | 135.1571 | 276.4995 | 212.6008 | 271.317 | 188.2429
2012-10-30 00:40 | 25.9617 | 297.9667 | 234.087 | 292.8111 | 210.1741 | 85.506 | 1.4389 | 36.405 | 52.3822 | 25.301 | 2.2809 | 36.2976 | 82.7915 | 15.9137 | 21.5061 | 5.8948 | 135.1737 | 276.4606 | 212.5809 | 271.305 | 188.668
2012-10-30 00:40 | 25.9625 | 297.9522 | 234.0957 | 292.808 | 209.8402 | 85.4956 | 1.439 | 36.4081 | 52.3913 | 25.301 | 2.2811 | 36.3005 | 82.8047 | 15.9137 | 21.4788 | 5.8941 | 135.196 | 276.4734 | 212.6169 | 271.3292 | 188.3614
2012-10-30 00:40 | 25.9692 | 297.955 | 234.0238 | 292.7878 | 209.9296 | 85.5031 | 1.4392 | 36.4122 | 52.4031 | 25.3009 | 2.2814 | 36.3046 | 82.8242 | 15.9135 | 21.5056 | 5.8946 | 135.2273 | 276.4494 | 212.5182 | 271.2822 | 188.424
2012-10-30 00:40 | 25.9587 | 297.9522 | 233.9976 | 292.785 | 210.3522 | 85.5071 | 1.4407 | 36.4512 | 52.5146 | 25.3014 | 2.2837 | 36.3425 | 82.9961 | 15.9138 | 21.483 | 5.8949 | 135.5107 | 276.4692 | 212.5146 | 271.302 | 188.8692
2012-10-30 00:40 | 25.958 | 297.9784 | 234.0519 | 292.7804 | 210.2137 | 85.506 | 1.4426 | 36.5008 | 52.6561 | 25.3021 | 2.2869 | 36.393 | 83.2263 | 15.9139 | 21.4863 | 5.8948 | 135.8824 | 276.4921 | 212.5656 | 271.2941 | 188.7274
2012-10-30 00:40 | 25.9542 | 297.9746 | 234.0481 | 292.7843 | 209.8476 | 85.5135 | 1.4433 | 36.5179 | 52.7048 | 25.3024 | 2.288 | 36.4103 | 83.3059 | 15.9138 | 21.5066 | 5.8953 | 136.0107 | 276.468 | 212.5415 | 271.2777 | 188.341
2012-10-30 00:40 | 25.9572 | 297.9699 | 233.9529 | 292.7489 | 209.5868 | 85.5083 | 1.4438 | 36.5313 | 52.7428 | 25.3028 | 2.2888 | 36.4234 | 83.3654 | 15.9139 | 21.4775 | 5.8949 | 136.1082 | 276.4924 | 212.4754 | 271.2714 | 188.1093
2012-10-30 00:40 | 25.952 | 297.9301 | 233.9516 | 292.7744 | 209.5973 | 85.5285 | 1.4441 | 36.5384 | 52.7637 | 25.3025 | 2.2892 | 36.4303 | 83.3972 | 15.9138 | 21.4964 | 5.8963 | 136.1609 | 276.4337 | 212.4552 | 271.278 | 188.1009
2012-10-30 00:40 | 25.9505 | 297.9478 | 233.9147 | 292.746 | 209.8124 | 85.5054 | 1.4442 | 36.5422 | 52.774 | 25.3029 | 2.2895 | 36.4344 | 83.4154 | 15.9139 | 21.4949 | 5.8947 | 136.1894 | 276.4529 | 212.4198 | 271.2511 | 188.3175
2012-10-30 00:40 | 25.9542 | 298.0053 | 233.997 | 292.742 | 210.2572 | 85.5181 | 1.4443 | 36.5457 | 52.7841 | 25.3029 | 2.2897 | 36.4379 | 83.4319 | 15.9138 | 21.4705 | 5.8956 | 136.216 | 276.5348 | 212.5265 | 271.2715 | 188.7867
2012-10-30 00:40 | 25.9453 | 297.9388 | 234.0078 | 292.7293 | 210.2089 | 85.5141 | 1.4444 | 36.5479 | 52.7903 | 25.3029 | 2.2898 | 36.4398 | 83.4409 | 15.9138 | 21.4937 | 5.8953 | 136.2313 | 276.4451 | 212.5141 | 271.2356 | 188.7152
2012-10-30 00:40 | 25.943 | 297.9634 | 234.0801 | 292.7693 | 210.2735 | 85.5233 | 1.4445 | 36.5488 | 52.7931 | 25.3027 | 2.2899 | 36.4409 | 83.4465 | 15.9137 | 21.5115 | 5.896 | 136.2397 | 276.4519 | 212.5686 | 271.2578 | 188.762
2012-10-30 00:40 | 25.9393 | 297.952 | 234.1 | 292.7502 | 210.1359 | 85.5204 | 1.4445 | 36.5495 | 52.7951 | 25.3028 | 2.2899 | 36.4415 | 83.4492 | 15.9137 | 21.5038 | 5.8958 | 136.2443 | 276.4482 | 212.5962 | 271.2464 | 188.6321
2012-10-30 00:41 | 25.9415 | 297.9926 | 234.1218 | 292.7755 | 210.0515 | 85.5256 | 1.4444 | 36.5482 | 52.7912 | 25.3029 | 2.2899 | 36.4401 | 83.443 | 15.9136 | 21.4899 | 5.8961 | 136.2342 | 276.5027 | 212.6319 | 271.2856 | 188.5616
2012-10-30 00:41 | 25.9385 | 297.955 | 234.1188 | 292.7417 | 210.1706 | 85.5187 | 1.4444 | 36.5478 | 52.7902 | 25.3028 | 2.2899 | 36.4399 | 83.4421 | 15.9136 | 21.4668 | 5.8957 | 136.2323 | 276.4882 | 212.652 | 271.2749 | 188.7038
2012-10-30 00:41 | 25.9378 | 297.962 | 234.1142 | 292.741 | 210.1108 | 85.5135 | 1.4444 | 36.5467 | 52.7872 | 25.3027 | 2.2898 | 36.439 | 83.4385 | 15.9135 | 21.4421 | 5.8953 | 136.2257 | 276.5199 | 212.6721 | 271.2989 | 188.6687
2012-10-30 00:41 | 25.9378 | 297.9236 | 234.1456 | 292.7602 | 210.0911 | 85.521 | 1.4444 | 36.5468 | 52.7868 | 25.303 | 2.2898 | 36.4389 | 83.4373 | 15.9136 | 21.43 | 5.8958 | 136.2242 | 276.4936 | 212.7156 | 271.3302 | 188.6611
2012-10-30 00:41 | 25.9355 | 297.9713 | 234.1433 | 292.7695 | 210.2227 | 85.5222 | 1.4443 | 36.5453 | 52.7829 | 25.3029 | 2.2897 | 36.4376 | 83.4318 | 15.9136 | 21.4598 | 5.8959 | 136.2146 | 276.5115 | 212.6835 | 271.3097 | 188.7629
2012-10-30 00:41 | 25.946 | 297.951 | 234.2127 | 292.7608 | 210.2608 | 85.5303 | 1.4443 | 36.5449 | 52.7816 | 25.303 | 2.2897 | 36.437 | 83.4292 | 15.9136 | 21.4743 | 5.8965 | 136.2109 | 276.4767 | 212.7384 | 271.2865 | 188.7865
2012-10-30 00:41 | 25.937 | 297.9958 | 234.1959 | 292.771 | 210.3423 | 85.5349 | 1.4443 | 36.5446 | 52.7808 | 25.3029 | 2.2897 | 36.4369 | 83.429 | 15.9135 | 21.4533 | 5.8968 | 136.2098 | 276.5425 | 212.7426 | 271.3177 | 188.889
2012-10-30 00:41 | 25.9296 | 297.9193 | 234.216 | 292.7482 | 210.3428 | 85.5274 | 1.4442 | 36.5429 | 52.7755 | 25.3031 | 2.2896 | 36.4352 | 83.4207 | 15.9136 | 21.474 | 5.8963 | 136.1962 | 276.4453 | 212.742 | 271.2742 | 188.8688
2012-10-30 00:41 | 25.9288 | 297.9722 | 234.2191 | 292.8013 | 210.3066 | 85.5303 | 1.4442 | 36.543 | 52.7763 | 25.3029 | 2.2896 | 36.4351 | 83.4207 | 15.9135 | 21.4852 | 5.8965 | 136.197 | 276.487 | 212.7339 | 271.3161 | 188.8214
2012-10-30 00:41 | 25.94 | 297.9604 | 234.2499 | 292.7817 | 210.3099 | 85.5274 | 1.4442 | 36.5422 | 52.7736 | 25.3031 | 2.2895 | 36.4343 | 83.4166 | 15.9136 | 21.4844 | 5.8963 | 136.1902 | 276.476 | 212.7655 | 271.2973 | 188.8255
2012-10-30 00:41 | 25.934 | 297.9659 | 234.24 | 292.768 | 210.2527 | 85.5198 | 1.4441 | 36.541 | 52.7706 | 25.3028 | 2.2895 | 36.433 | 83.4116 | 15.9134 | 21.5185 | 5.8957 | 136.1822 | 276.4474 | 212.7215 | 271.2495 | 188.7342
2012-10-30 00:41 | 25.9288 | 297.9991 | 234.2702 | 292.7743 | 210.2475 | 85.5256 | 1.4441 | 36.5388 | 52.7643 | 25.3028 | 2.2893 | 36.431 | 83.403 | 15.9133 | 21.4812 | 5.8961 | 136.1673 | 276.5179 | 212.789 | 271.2931 | 188.7663
2012-10-30 00:41 | 25.9251 | 297.9647 | 234.3136 | 292.7822 | 210.2084 | 85.5193 | 1.444 | 36.5379 | 52.761 | 25.3031 | 2.2893 | 36.4298 | 83.3972 | 15.9134 | 21.4735 | 5.8957 | 136.1582 | 276.4912 | 212.8401 | 271.3087 | 188.7349
2012-10-30 00:41 | 25.9355 | 297.9943 | 234.2965 | 292.8003 | 210.2542 | 85.5118 | 1.444 | 36.5367 | 52.758 | 25.3029 | 2.2892 | 36.4287 | 83.3924 | 15.9133 | 21.4598 | 5.8952 | 136.1504 | 276.5345 | 212.8367 | 271.3405 | 188.7944
2012-10-30 00:41 | 25.9221 | 297.9847 | 234.2517 | 292.7676 | 210.3078 | 85.5343 | 1.4439 | 36.5358 | 52.7556 | 25.3028 | 2.2892 | 36.428 | 83.3899 | 15.9132 | 21.4504 | 5.8967 | 136.1456 | 276.5343 | 212.8013 | 271.3172 | 188.8574
2012-10-30 00:41 | 25.9273 | 297.9746 | 234.2726 | 292.8152 | 210.2421 | 85.5193 | 1.4439 | 36.5348 | 52.7529 | 25.3028 | 2.2891 | 36.4272 | 83.3856 | 15.9133 | 21.4556 | 5.8957 | 136.1386 | 276.519 | 212.817 | 271.3596 | 188.7865
2012-10-30 00:41 | 25.9303 | 297.9584 | 234.2834 | 292.7989 | 210.5247 | 85.532 | 1.4439 | 36.5336 | 52.7498 | 25.3026 | 2.2891 | 36.4258 | 83.3807 | 15.913 | 21.4948 | 5.8966 | 136.1306 | 276.4636 | 212.7886 | 271.3041 | 189.0299
2012-10-30 00:41 | 25.9281 | 297.9907 | 234.293 | 292.8314 | 210.5028 | 85.5245 | 1.4439 | 36.534 | 52.7505 | 25.3027 | 2.2891 | 36.426 | 83.3813 | 15.9131 | 21.4845 | 5.8961 | 136.1318 | 276.5062 | 212.8085 | 271.3469 | 189.0183
2012-10-30 00:41 | 25.9333 | 298.0151 | 234.3414 | 292.8289 | 210.3308 | 85.5233 | 1.4438 | 36.5328 | 52.7474 | 25.3025 | 2.289 | 36.4248 | 83.3757 | 15.9131 | 21.4897 | 5.896 | 136.1231 | 276.5254 | 212.8517 | 271.3392 | 188.8411
2012-10-30 00:41 | 25.9318 | 298.0098 | 234.3871 | 292.8428 | 210.3332 | 85.5233 | 1.4438 | 36.5313 | 52.7425 | 25.3028 | 2.2889 | 36.4234 | 83.3696 | 15.9131 | 21.5123 | 5.896 | 136.1121 | 276.4975 | 212.8748 | 271.3305 | 188.8209
2012-10-30 00:41 | 25.9318 | 297.9906 | 234.3399 | 292.8043 | 210.3332 | 85.517 | 1.4437 | 36.5309 | 52.7411 | 25.3029 | 2.2889 | 36.4231 | 83.3682 | 15.913 | 21.5083 | 5.8955 | 136.1093 | 276.4823 | 212.8316 | 271.296 | 188.8249
2012-10-30 00:41 | 25.9348 | 298.0128 | 234.3822 | 292.8496 | 210.1944 | 85.5279 | 1.4437 | 36.5292 | 52.7366 | 25.3028 | 2.2888 | 36.4216 | 83.3612 | 15.913 | 21.5434 | 5.8963 | 136.0979 | 276.4694 | 212.8388 | 271.3062 | 188.651
2012-10-30 00:41 | 25.9325 | 298.0028 | 234.3367 | 292.8358 | 210.0543 | 85.5378 | 1.4436 | 36.5277 | 52.7326 | 25.3026 | 2.2887 | 36.4199 | 83.354 | 15.913 | 21.5491 | 5.897 | 136.0866 | 276.4537 | 212.7876 | 271.2867 | 188.5052
2012-10-30 00:42 | 25.9288 | 297.9876 | 234.2898 | 292.8628 | 210.1609 | 85.5245 | 1.4436 | 36.5262 | 52.7276 | 25.3029 | 2.2886 | 36.4184 | 83.3465 | 15.9131 | 21.5093 | 5.8961 | 136.0741 | 276.4783 | 212.7805 | 271.3535 | 188.6516
2012-10-30 00:42 | 25.9333 | 297.9806 | 234.3846 | 292.8327 | 210.2638 | 85.5378 | 1.4435 | 36.5242 | 52.7225 | 25.3026 | 2.2885 | 36.4162 | 83.3372 | 15.9129 | 21.5138 | 5.897 | 136.0597 | 276.4668 | 212.8708 | 271.3189 | 188.75
2012-10-30 00:42 | 25.934 | 298.012 | 234.3578 | 292.8411 | 210.3905 | 85.5291 | 1.4434 | 36.5229 | 52.7183 | 25.3028 | 2.2884 | 36.4151 | 83.332 | 15.913 | 21.5105 | 5.8964 | 136.0503 | 276.5015 | 212.8473 | 271.3306 | 188.88
2012-10-30 00:42 | 25.9378 | 298.012 | 234.3695 | 292.8718 | 210.3983 | 85.5378 | 1.4434 | 36.5227 | 52.7182 | 25.3026 | 2.2884 | 36.4149 | 83.3318 | 15.9128 | 21.4902 | 5.897 | 136.05 | 276.5218 | 212.8793 | 271.3816 | 188.9081
2012-10-30 00:42 | 25.9408 | 298.015 | 234.4275 | 292.8441 | 210.3855 | 85.5401 | 1.4434 | 36.5215 | 52.7146 | 25.3026 | 2.2883 | 36.4137 | 83.3257 | 15.9129 | 21.4932 | 5.8971 | 136.0403 | 276.5218 | 212.9343 | 271.3509 | 188.8923
2012-10-30 00:42 | 25.9445 | 298.0148 | 234.4194 | 292.8747 | 210.4325 | 85.5355 | 1.4434 | 36.5206 | 52.7127 | 25.3024 | 2.2883 | 36.413 | 83.3233 | 15.9128 | 21.525 | 5.8968 | 136.0361 | 276.4898 | 212.8944 | 271.3497 | 188.9075
2012-10-30 00:42 | 25.9408 | 298.0034 | 234.4039 | 292.8748 | 210.4328 | 85.5291 | 1.4433 | 36.5199 | 52.7101 | 25.3026 | 2.2882 | 36.412 | 83.3186 | 15.9129 | 21.5333 | 5.8964 | 136.0286 | 276.4701 | 212.8706 | 271.3415 | 188.8995
2012-10-30 00:42 | 25.9325 | 298.0105 | 234.4388 | 292.8473 | 210.4324 | 85.5349 | 1.4394 | 36.4191 | 52.4222 | 25.3014 | 2.2818 | 36.3094 | 82.8515 | 15.9124 | 21.4889 | 5.8968 | 135.2737 | 276.5216 | 212.9499 | 271.3584 | 188.9435
2012-10-30 00:42 | 25.9325 | 298.0067 | 234.3996 | 292.8588 | 210.4127 | 85.5164 | 1.4392 | 36.4136 | 52.4072 | 25.3009 | 2.2816 | 36.306 | 82.8365 | 15.9124 | 21.4809 | 5.8955 | 135.2437 | 276.5258 | 212.9187 | 271.3779 | 188.9318
2012-10-30 00:42 | 25.9415 | 298.0195 | 234.3889 | 292.8832 | 209.9097 | 85.5135 | 1.4388 | 36.4015 | 52.373 | 25.3006 | 2.2809 | 36.2939 | 82.7814 | 15.9124 | 21.4899 | 5.8953 | 135.1544 | 276.5296 | 212.899 | 271.3933 | 188.4198
2012-10-30 00:42 | 25.9311 | 298.0245 | 234.3196 | 292.8651 | 209.5804 | 85.517 | 1.4386 | 36.3975 | 52.3617 | 25.3006 | 2.2806 | 36.2901 | 82.7634 | 15.9125 | 21.5196 | 5.8955 | 135.125 | 276.5049 | 212.8 | 271.3455 | 188.0608
2012-10-30 00:42 | 25.9273 | 298.0168 | 234.2097 | 292.8767 | 209.675 | 85.5233 | 1.4387 | 36.399 | 52.3658 | 25.3006 | 2.2807 | 36.2918 | 82.7718 | 15.9123 | 21.5319 | 5.896 | 135.1376 | 276.4849 | 212.6778 | 271.3448 | 188.1431
2012-10-30 00:42 | 25.9288 | 298.0183 | 234.172 | 292.8782 | 210.1018 | 85.5158 | 1.4388 | 36.4036 | 52.3787 | 25.3007 | 2.281 | 36.2962 | 82.7915 | 15.9124 | 21.4973 | 5.8955 | 135.1702 | 276.521 | 212.6747 | 271.3809 | 188.6045
2012-10-30 00:42 | 25.9273 | 297.9631 | 234.2019 | 292.8806 | 210.5886 | 85.5014 | 1.439 | 36.4087 | 52.3935 | 25.3007 | 2.2813 | 36.3012 | 82.8146 | 15.9124 | 21.4958 | 5.8945 | 135.2082 | 276.4673 | 212.7061 | 271.3848 | 189.0928
2012-10-30 00:42 | 25.9221 | 297.9732 | 234.2399 | 292.8523 | 210.6937 | 85.506 | 1.4392 | 36.4127 | 52.4055 | 25.3005 | 2.2816 | 36.3053 | 82.8331 | 15.9124 | 21.4946 | 5.8948 | 135.2385 | 276.4786 | 212.7453 | 271.3577 | 189.1991
2012-10-30 00:42 | 25.9266 | 298.0123 | 234.3269 | 292.8876 | 210.3044 | 85.5083 | 1.4394 | 36.4177 | 52.419 | 25.3009 | 2.2819 | 36.3102 | 82.8552 | 15.9124 | 21.475 | 5.8949 | 135.2742 | 276.5373 | 212.8519 | 271.4126 | 188.8294
2012-10-30 00:42 | 25.9243 | 297.9946 | 234.2657 | 292.866 | 209.9516 | 85.5054 | 1.4422 | 36.4893 | 52.6236 | 25.3017 | 2.2863 | 36.3807 | 83.176 | 15.9127 | 21.4647 | 5.8947 | 135.7996 | 276.5299 | 212.801 | 271.4013 | 188.4869
2012-10-30 00:42 | 25.9183 | 297.9809 | 234.2439 | 292.81 | 210.0401 | 85.5083 | 1.443 | 36.5104 | 52.6833 | 25.3022 | 2.2876 | 36.4023 | 83.2746 | 15.9127 | 21.4787 | 5.8949 | 135.9579 | 276.5022 | 212.7652 | 271.3313 | 188.5614
2012-10-30 00:42 | 25.9108 | 298.0272 | 234.2089 | 292.8295 | 209.627 | 85.5083 | 1.4435 | 36.524 | 52.7224 | 25.3023 | 2.2885 | 36.416 | 83.3377 | 15.9127 | 21.4833 | 5.8949 | 136.0601 | 276.5439 | 212.7256 | 271.3462 | 188.1437
2012-10-30 00:42 | 25.9153 | 298.0356 | 234.1585 | 292.8455 | 209.3401 | 85.5175 | 1.444 | 36.5358 | 52.7562 | 25.3025 | 2.2892 | 36.428 | 83.3926 | 15.9127 | 21.4918 | 5.8956 | 136.1488 | 276.5438 | 212.6667 | 271.3537 | 187.8483
2012-10-30 00:42 | 25.9161 | 298.0095 | 234.1239 | 292.7809 | 209.8056 | 85.5141 | 1.4442 | 36.5412 | 52.7713 | 25.3027 | 2.2896 | 36.4333 | 83.4168 | 15.9127 | 21.5167 | 5.8953 | 136.1881 | 276.4928 | 212.6072 | 271.2642 | 188.2889
2012-10-30 00:42 | 25.9064 | 298.0228 | 234.1221 | 292.8289 | 210.0203 | 85.5089 | 1.4443 | 36.5459 | 52.7845 | 25.3029 | 2.2899 | 36.4379 | 83.4377 | 15.9127 | 21.503 | 5.895 | 136.2222 | 276.5198 | 212.6191 | 271.3259 | 188.5173
2012-10-30 00:42 | 25.9026 | 298.0037 | 234.0986 | 292.8405 | 209.9378 | 85.5181 | 1.4444 | 36.5475 | 52.7897 | 25.3027 | 2.29 | 36.4396 | 83.4451 | 15.9128 | 21.5112 | 5.8956 | 136.2347 | 276.4925 | 212.5874 | 271.3293 | 188.4266
2012-10-30 00:43 | 25.9064 | 298.0612 | 234.1299 | 292.7866 | 210.0046 | 85.5268 | 1.4444 | 36.5466 | 52.7874 | 25.3025 | 2.2899 | 36.4387 | 83.4411 | 15.9127 | 21.4829 | 5.8962 | 136.2285 | 276.5783 | 212.647 | 271.3037 | 188.5217
2012-10-30 00:43 | 25.9071 | 298.0312 | 234.162 | 292.8334 | 209.9659 | 85.5146 | 1.4444 | 36.5471 | 52.7887 | 25.3026 | 2.2899 | 36.4392 | 83.4439 | 15.9127 | 21.4756 | 5.8954 | 136.2326 | 276.5556 | 212.6864 | 271.3578 | 188.4903
2012-10-30 00:43 | 25.8951 | 298 | 234.1147 | 292.8176 | 210.0642 | 85.5262 | 1.4444 | 36.5484 | 52.7919 | 25.3028 | 2.29 | 36.4405 | 83.4498 | 15.9127 | 21.4997 | 5.8962 | 136.2417 | 276.5003 | 212.615 | 271.3179 | 188.5645
2012-10-30 00:43 | 25.9064 | 298.0267 | 234.1417 | 292.852 | 210.2251 | 85.5227 | 1.4444 | 36.5484 | 52.7917 | 25.3029 | 2.29 | 36.4404 | 83.4487 | 15.9128 | 21.515 | 5.8959 | 136.2404 | 276.5117 | 212.6267 | 271.337 | 188.7101
2012-10-30 00:43 | 25.8959 | 298.0162 | 234.2019 | 292.8107 | 210.443 | 85.5268 | 1.4444 | 36.5481 | 52.7905 | 25.303 | 2.29 | 36.4402 | 83.4473 | 15.9129 | 21.5165 | 5.8962 | 136.2378 | 276.4997 | 212.6854 | 271.2942 | 188.9265
2012-10-30 00:43 | 25.8981 | 298.0068 | 234.2787 | 292.8052 | 210.1656 | 85.5071 | 1.4444 | 36.547 | 52.7871 | 25.3033 | 2.2899 | 36.4391 | 83.4429 | 15.9128 | 21.5067 | 5.8949 | 136.23 | 276.5001 | 212.772 | 271.2985 | 188.6589
2012-10-30 00:43 | 25.8936 | 297.9754 | 234.2192 | 292.843 | 209.9091 | 85.5198 | 1.4443 | 36.5463 | 52.7853 | 25.3032 | 2.2899 | 36.4383 | 83.4398 | 15.9127 | 21.5142 | 5.8957 | 136.2251 | 276.4612 | 212.705 | 271.3288 | 188.3949
2012-10-30 00:43 | 25.9056 | 298.0143 | 234.1763 | 292.8743 | 209.8148 | 85.5216 | 1.4444 | 36.5471 | 52.7876 | 25.3031 | 2.2899 | 36.4393 | 83.4439 | 15.9127 | 21.4901 | 5.8959 | 136.2315 | 276.5242 | 212.6862 | 271.3842 | 188.3247
2012-10-30 00:43 | 25.8974 | 297.9946 | 234.1877 | 292.7891 | 209.9562 | 85.517 | 1.4443 | 36.5461 | 52.7843 | 25.3033 | 2.2899 | 36.438 | 83.4379 | 15.9128 | 21.51 | 5.8955 | 136.2222 | 276.4846 | 212.6777 | 271.2791 | 188.4462
2012-10-30 00:43 | 25.9093 | 297.9988 | 234.2153 | 292.8241 | 210.1808 | 85.532 | 1.4443 | 36.5451 | 52.7811 | 25.3035 | 2.2898 | 36.4373 | 83.4351 | 15.9127 | 21.4778 | 5.8966 | 136.2162 | 276.521 | 212.7375 | 271.3463 | 188.703
2012-10-30 00:43 | 25.9086 | 297.9943 | 234.2028 | 292.8388 | 210.121 | 85.5279 | 1.4443 | 36.5442 | 52.7793 | 25.3031 | 2.2898 | 36.4363 | 83.4306 | 15.9127 | 21.5092 | 5.8963 | 136.2099 | 276.4851 | 212.6936 | 271.3296 | 188.6118
2012-10-30 00:43 | 25.9019 | 298.0106 | 234.2079 | 292.7744 | 210.1182 | 85.5181 | 1.4442 | 36.5428 | 52.7753 | 25.3031 | 2.2897 | 36.4349 | 83.4245 | 15.9126 | 21.5105 | 5.8956 | 136.1998 | 276.5001 | 212.6974 | 271.2639 | 188.6077
2012-10-30 00:43 | 25.9123 | 297.9941 | 234.2262 | 292.8386 | 210.1011 | 85.5256 | 1.4441 | 36.5407 | 52.7692 | 25.3031 | 2.2895 | 36.4329 | 83.4148 | 15.9127 | 21.5048 | 5.8961 | 136.184 | 276.4893 | 212.7214 | 271.3338 | 188.5963
2012-10-30 00:43 | 25.9123 | 297.998 | 234.2419 | 292.8194 | 210.168 | 85.532 | 1.444 | 36.5385 | 52.7626 | 25.3032 | 2.2894 | 36.4307 | 83.4046 | 15.9127 | 21.5369 | 5.8966 | 136.1672 | 276.4611 | 212.705 | 271.2825 | 188.6311
2012-10-30 00:43 | 25.9146 | 297.9427 | 234.2717 | 292.8525 | 210.3121 | 85.5256 | 1.444 | 36.5385 | 52.7621 | 25.3034 | 2.2894 | 36.4305 | 83.404 | 15.9127 | 21.5152 | 5.8961 | 136.1661 | 276.4275 | 212.7565 | 271.3373 | 188.7969
2012-10-30 00:43 | 25.9213 | 297.9916 | 234.3177 | 292.813 | 210.366 | 85.5198 | 1.444 | 36.5387 | 52.7627 | 25.3034 | 2.2894 | 36.4307 | 83.4047 | 15.9128 | 21.5459 | 5.8957 | 136.1674 | 276.4457 | 212.7718 | 271.2671 | 188.8201
2012-10-30 00:43 | 25.9296 | 297.9961 | 234.2945 | 292.8636 | 210.3349 | 85.5424 | 1.444 | 36.5377 | 52.7597 | 25.3034 | 2.2894 | 36.4298 | 83.4015 | 15.9125 | 21.5342 | 5.8973 | 136.1612 | 276.4619 | 212.7603 | 271.3294 | 188.8007
2012-10-30 00:43 | 25.9221 | 297.9732 | 234.3106 | 292.8254 | 210.2526 | 85.5343 | 1.444 | 36.5371 | 52.7584 | 25.3033 | 2.2893 | 36.4292 | 83.3992 | 15.9125 | 21.5146 | 5.8967 | 136.1575 | 276.4586 | 212.796 | 271.3108 | 188.738
2012-10-30 00:43 | 25.9161 | 297.9864 | 234.3125 | 292.8386 | 210.2978 | 85.5372 | 1.4439 | 36.5356 | 52.754 | 25.3033 | 2.2892 | 36.4279 | 83.3921 | 15.9127 | 21.5086 | 5.8969 | 136.1461 | 276.4778 | 212.8039 | 271.33 | 188.7892
2012-10-30 00:43 | 25.9191 | 297.9664 | 234.3076 | 292.8339 | 210.2536 | 85.5435 | 1.4439 | 36.5353 | 52.7533 | 25.3032 | 2.2892 | 36.4274 | 83.3906 | 15.9126 | 21.5116 | 5.8974 | 136.1439 | 276.4548 | 212.796 | 271.3223 | 188.742
2012-10-30 00:43 | 25.9251 | 297.9724 | 234.3529 | 292.8553 | 210.2793 | 85.5349 | 1.4439 | 36.5348 | 52.7516 | 25.3033 | 2.2892 | 36.4266 | 83.3873 | 15.9125 | 21.5337 | 5.8968 | 136.1389 | 276.4387 | 212.8192 | 271.3216 | 188.7456
2012-10-30 00:43 | 25.9228 | 297.9778 | 234.3702 | 292.8607 | 210.3281 | 85.5476 | 1.4438 | 36.5339 | 52.7491 | 25.3033 | 2.2892 | 36.426 | 83.3851 | 15.9124 | 21.5234 | 5.8977 | 136.1342 | 276.4544 | 212.8468 | 271.3373 | 188.8047
2012-10-30 00:43 | 25.9288 | 298.0337 | 234.3487 | 292.8821 | 210.4168 | 85.5337 | 1.4438 | 36.5324 | 52.7446 | 25.3033 | 2.2891 | 36.4247 | 83.3782 | 15.9125 | 21.5494 | 5.8967 | 136.1229 | 276.4843 | 212.7993 | 271.3327 | 188.8674
2012-10-30 00:43 | 25.9258 | 297.9769 | 234.4007 | 292.8521 | 210.3469 | 85.5418 | 1.4437 | 36.5305 | 52.7397 | 25.303 | 2.2889 | 36.4225 | 83.3688 | 15.9124 | 21.5143 | 5.8973 | 136.1086 | 276.4626 | 212.8864 | 271.3378 | 188.8326
2012-10-30 00:43 | 25.9183 | 297.9617 | 234.4246 | 292.86 | 210.4378 | 85.532 | 1.4437 | 36.5293 | 52.7359 | 25.3033 | 2.2889 | 36.4215 | 83.3642 | 15.9125 | 21.4988 | 5.8966 | 136.1 | 276.4629 | 212.9258 | 271.3612 | 188.939
2012-10-30 00:44 | 25.9288 | 297.9914 | 234.4391 | 292.8667 | 210.4129 | 85.5401 | 1.4436 | 36.5276 | 52.7313 | 25.3031 | 2.2888 | 36.4198 | 83.3572 | 15.9123 | 21.4933 | 5.8971 | 136.0885 | 276.4981 | 212.9458 | 271.3734 | 188.9196
2012-10-30 00:44 | 25.9266 | 297.9739 | 234.4604 | 292.8491 | 210.3989 | 85.5435 | 1.4436 | 36.5271 | 52.7296 | 25.3032 | 2.2888 | 36.4192 | 83.355 | 15.9122 | 21.4911 | 5.8974 | 136.0846 | 276.4828 | 212.9693 | 271.358 | 188.9078
2012-10-30 00:44 | 25.9281 | 298.0023 | 234.4227 | 292.8929 | 210.298 | 85.5401 | 1.4435 | 36.5261 | 52.7269 | 25.3031 | 2.2887 | 36.4182 | 83.3501 | 15.9122 | 21.5086 | 5.8971 | 136.077 | 276.4937 | 212.9141 | 271.3843 | 188.7894
2012-10-30 00:44 | 25.9311 | 298.0168 | 234.4571 | 292.8767 | 210.3837 | 85.5418 | 1.4435 | 36.5241 | 52.7211 | 25.3031 | 2.2886 | 36.4163 | 83.3416 | 15.9122 | 21.4915 | 5.8973 | 136.0626 | 276.5253 | 212.9656 | 271.3852 | 188.8922
2012-10-30 00:44 | 25.9281 | 297.9946 | 234.4344 | 292.9045 | 210.4555 | 85.543 | 1.4434 | 36.5231 | 52.718 | 25.3032 | 2.2885 | 36.4153 | 83.3376 | 15.9121 | 21.5487 | 5.8973 | 136.0556 | 276.4459 | 212.8857 | 271.3558 | 188.9068
2012-10-30 00:44 | 25.9281 | 298.0176 | 234.4541 | 292.9314 | 210.3768 | 85.5412 | 1.4434 | 36.5233 | 52.7192 | 25.3029 | 2.2885 | 36.4153 | 83.337 | 15.9122 | 21.5126 | 5.8972 | 136.0562 | 276.505 | 212.9415 | 271.4188 | 188.8642
2012-10-30 00:44 | 25.9311 | 298.0168 | 234.4571 | 292.8728 | 210.4979 | 85.5447 | 1.4434 | 36.5225 | 52.7169 | 25.3029 | 2.2885 | 36.4147 | 83.3348 | 15.9121 | 21.5196 | 5.8975 | 136.0517 | 276.4972 | 212.9375 | 271.3532 | 188.9783
2012-10-30 00:44 | 25.9281 | 298.0407 | 234.4423 | 292.9045 | 210.4595 | 85.5395 | 1.4434 | 36.5222 | 52.7158 | 25.303 | 2.2885 | 36.4144 | 83.3339 | 15.912 | 21.4805 | 5.8971 | 136.0497 | 276.5602 | 212.9618 | 271.424 | 188.979
2012-10-30 00:44 | 25.9355 | 298.0173 | 234.4458 | 292.9273 | 210.5299 | 85.5563 | 1.4433 | 36.5207 | 52.7113 | 25.3031 | 2.2884 | 36.4129 | 83.3269 | 15.912 | 21.4839 | 5.8983 | 136.0382 | 276.5334 | 212.9619 | 271.4434 | 189.046
2012-10-30 00:44 | 25.9363 | 298.0258 | 234.4898 | 292.8896 | 210.4283 | 85.5407 | 1.4412 | 36.4652 | 52.5529 | 25.3023 | 2.2848 | 36.3556 | 83.0664 | 15.9117 | 21.5489 | 5.8972 | 135.6193 | 276.4769 | 212.9409 | 271.3407 | 188.8794
2012-10-30 00:44 | 25.9333 | 298.0228 | 234.4671 | 292.875 | 210.3938 | 85.5355 | 1.4394 | 36.4201 | 52.4248 | 25.3015 | 2.2822 | 36.3125 | 82.8705 | 15.9115 | 21.4978 | 5.8968 | 135.2953 | 276.525 | 212.9693 | 271.3772 | 188.896
2012-10-30 00:44 | 25.9333 | 298.0151 | 234.4279 | 292.9135 | 210.3386 | 85.5245 | 1.4389 | 36.4073 | 52.3882 | 25.3013 | 2.2814 | 36.2998 | 82.8128 | 15.9115 | 21.4857 | 5.8961 | 135.2009 | 276.5294 | 212.9422 | 271.4278 | 188.8529
2012-10-30 00:44 | 25.9348 | 297.9782 | 234.3861 | 292.9189 | 209.0131 | 85.5331 | 1.4386 | 36.3989 | 52.3652 | 25.3009 | 2.2808 | 36.2915 | 82.7751 | 15.9114 | 21.4912 | 5.8967 | 135.1403 | 276.487 | 212.8949 | 271.4277 | 187.5219
2012-10-30 00:44 | 25.9378 | 298.0004 | 234.2085 | 292.8603 | 209.5674 | 85.5164 | 1.4386 | 36.397 | 52.3593 | 25.3009 | 2.2807 | 36.2894 | 82.7651 | 15.9116 | 21.5143 | 5.8955 | 135.1245 | 276.4861 | 212.6942 | 271.346 | 188.0531
2012-10-30 00:44 | 25.9348 | 297.9936 | 234.1819 | 292.8996 | 210.2574 | 85.5239 | 1.4387 | 36.4006 | 52.3695 | 25.301 | 2.2809 | 36.2931 | 82.7826 | 15.9114 | 21.5354 | 5.896 | 135.152 | 276.4582 | 212.6465 | 271.3642 | 188.722
2012-10-30 00:44 | 25.9393 | 297.9712 | 234.2414 | 292.8849 | 210.3171 | 85.5089 | 1.4389 | 36.4046 | 52.3811 | 25.301 | 2.2812 | 36.2971 | 82.8004 | 15.9115 | 21.5078 | 5.895 | 135.1815 | 276.4634 | 212.7336 | 271.3771 | 188.8093
2012-10-30 00:44 | 25.9318 | 297.9752 | 234.2614 | 292.862 | 209.9158 | 85.5106 | 1.4391 | 36.4096 | 52.3958 | 25.3009 | 2.2815 | 36.3022 | 82.8238 | 15.9115 | 21.5123 | 5.8951 | 135.2196 | 276.4629 | 212.7491 | 271.3497 | 188.4035
2012-10-30 00:44 | 25.9393 | 297.9712 | 234.2571 | 292.9118 | 209.9863 | 85.5106 | 1.4393 | 36.4145 | 52.4099 | 25.3009 | 2.2818 | 36.3071 | 82.8463 | 15.9115 | 21.4837 | 5.8951 | 135.2562 | 276.4875 | 212.7734 | 271.4281 | 188.5026
2012-10-30 00:44 | 25.9378 | 297.939 | 234.2399 | 292.9103 | 209.9533 | 85.5303 | 1.4408 | 36.4536 | 52.5216 | 25.3013 | 2.2843 | 36.3466 | 83.0259 | 15.9116 | 21.4982 | 5.8965 | 135.5474 | 276.4408 | 212.7417 | 271.4121 | 188.4551
2012-10-30 00:44 | 25.9438 | 297.9565 | 234.1752 | 292.8163 | 210.0026 | 85.5204 | 1.4427 | 36.5044 | 52.6659 | 25.3024 | 2.2874 | 36.3964 | 83.2525 | 15.9118 | 21.5042 | 5.8958 | 135.9184 | 276.4523 | 212.671 | 271.3121 | 188.4984
2012-10-30 00:44 | 25.9363 | 297.9682 | 234.1205 | 292.8126 | 209.9912 | 85.5297 | 1.4433 | 36.5179 | 52.7048 | 25.3024 | 2.2883 | 36.4105 | 83.3171 | 15.9118 | 21.4566 | 5.8964 | 136.022 | 276.5116 | 212.6639 | 271.356 | 188.5346
2012-10-30 00:44 | 25.9258 | 297.9808 | 234.11 | 292.8329 | 210.0319 | 85.5349 | 1.4438 | 36.5306 | 52.7412 | 25.3024 | 2.2891 | 36.423 | 83.3746 | 15.9117 | 21.5023 | 5.8968 | 136.1158 | 276.4785 | 212.6077 | 271.3306 | 188.5296
2012-10-30 00:44 | 25.9318 | 297.9791 | 234.116 | 292.8081 | 210.0457 | 85.5308 | 1.4441 | 36.5387 | 52.764 | 25.3028 | 2.2896 | 36.4311 | 83.4115 | 15.9117 | 21.5243 | 5.8965 | 136.1755 | 276.4548 | 212.5917 | 271.2838 | 188.5214
2012-10-30 00:45 | 25.9325 | 297.9567 | 234.1167 | 292.8204 | 210.2394 | 85.5314 | 1.4442 | 36.5436 | 52.7779 | 25.3029 | 2.2899 | 36.4357 | 83.433 | 15.9117 | 21.4889 | 5.8965 | 136.2108 | 276.4678 | 212.6278 | 271.3315 | 188.7505
2012-10-30 00:45 | 25.9325 | 297.9913 | 234.1874 | 292.8127 | 210.2591 | 85.5314 | 1.4443 | 36.5462 | 52.7853 | 25.3029 | 2.29 | 36.4382 | 83.4443 | 15.9117 | 21.513 | 5.8965 | 136.2296 | 276.4783 | 212.6744 | 271.2997 | 188.7461
2012-10-30 00:45 | 25.9325 | 297.9875 | 234.2385 | 292.8358 | 210.1252 | 85.5389 | 1.4444 | 36.5487 | 52.7923 | 25.303 | 2.2902 | 36.4408 | 83.4561 | 15.9118 | 21.4769 | 5.8971 | 136.2484 | 276.5106 | 212.7616 | 271.3589 | 188.6483
2012-10-30 00:45 | 25.9333 | 297.9729 | 234.1804 | 292.8058 | 210.2363 | 85.5268 | 1.4445 | 36.5492 | 52.7937 | 25.303 | 2.2902 | 36.4412 | 83.4583 | 15.9116 | 21.4897 | 5.8962 | 136.252 | 276.4832 | 212.6907 | 271.3161 | 188.7466
2012-10-30 00:45 | 25.9318 | 297.9983 | 234.2339 | 292.8043 | 210.1639 | 85.5308 | 1.4445 | 36.5492 | 52.7938 | 25.3031 | 2.2902 | 36.4412 | 83.4587 | 15.9116 | 21.4882 | 5.8965 | 136.2525 | 276.5101 | 212.7457 | 271.3161 | 188.6757
2012-10-30 00:45 | 25.94 | 297.9565 | 234.2146 | 292.8163 | 210.2863 | 85.5366 | 1.4444 | 36.5483 | 52.7915 | 25.3029 | 2.2902 | 36.4404 | 83.4547 | 15.9117 | 21.5366 | 5.8969 | 136.2462 | 276.4199 | 212.678 | 271.2797 | 188.7497
2012-10-30 00:45 | 25.9453 | 297.9618 | 234.2945 | 292.8101 | 210.2758 | 85.5303 | 1.4444 | 36.5487 | 52.7924 | 25.3031 | 2.2902 | 36.4408 | 83.457 | 15.9115 | 21.4536 | 5.8965 | 136.2494 | 276.5082 | 212.8409 | 271.3565 | 188.8222
2012-10-30 00:45 | 25.9445 | 297.9764 | 234.2623 | 292.8093 | 210.2238 | 85.5262 | 1.4444 | 36.548 | 52.7903 | 25.303 | 2.2902 | 36.4401 | 83.4542 | 15.9115 | 21.4448 | 5.8962 | 136.2445 | 276.5316 | 212.8175 | 271.3645 | 188.779
2012-10-30 00:45 | 25.9408 | 297.9766 | 234.2625 | 292.8094 | 210.2241 | 85.5389 | 1.4444 | 36.5478 | 52.79 | 25.3029 | 2.2902 | 36.4397 | 83.4527 | 15.9114 | 21.4651 | 5.8971 | 136.2427 | 276.5115 | 212.7974 | 271.3443 | 188.759
2012-10-30 00:45 | 25.9505 | 297.9863 | 234.2565 | 292.8422 | 210.3165 | 85.536 | 1.4444 | 36.5473 | 52.7885 | 25.3029 | 2.2901 | 36.4393 | 83.4505 | 15.9115 | 21.4989 | 5.8969 | 136.239 | 276.4874 | 212.7576 | 271.3433 | 188.8176
2012-10-30 00:45 | 25.9423 | 297.9665 | 234.2994 | 292.7994 | 210.3083 | 85.5378 | 1.4444 | 36.547 | 52.7876 | 25.303 | 2.2901 | 36.439 | 83.4488 | 15.9115 | 21.4987 | 5.897 | 136.2365 | 276.4678 | 212.8007 | 271.3007 | 188.8096
2012-10-30 00:45 | 25.946 | 297.9472 | 234.3227 | 292.8223 | 210.3435 | 85.5331 | 1.4443 | 36.5462 | 52.7846 | 25.3033 | 2.2901 | 36.4384 | 83.4467 | 15.9114 | 21.4463 | 5.8967 | 136.2313 | 276.5009 | 212.8764 | 271.376 | 188.8972
2012-10-30 00:45 | 25.9542 | 297.9477 | 234.3623 | 292.8267 | 210.4462 | 85.5331 | 1.4443 | 36.5454 | 52.7825 | 25.3032 | 2.29 | 36.4376 | 83.4427 | 15.9115 | 21.4825 | 5.8967 | 136.2252 | 276.4652 | 212.8798 | 271.3442 | 188.9637
2012-10-30 00:45 | 25.946 | 297.9741 | 234.3856 | 292.7954 | 210.308 | 85.5459 | 1.4443 | 36.5442 | 52.7795 | 25.303 | 2.29 | 36.4364 | 83.4382 | 15.9114 | 21.4663 | 5.8975 | 136.2177 | 276.5078 | 212.9193 | 271.3291 | 188.8417
2012-10-30 00:45 | 25.9498 | 297.9471 | 234.4208 | 292.803 | 210.304 | 85.5274 | 1.4442 | 36.5433 | 52.7767 | 25.303 | 2.2899 | 36.4354 | 83.4327 | 15.9115 | 21.4942 | 5.8963 | 136.2094 | 276.4529 | 212.9266 | 271.3088 | 188.8098
2012-10-30 00:45 | 25.9565 | 297.9807 | 234.3882 | 292.8328 | 210.3067 | 85.5303 | 1.4441 | 36.541 | 52.7702 | 25.303 | 2.2897 | 36.433 | 83.4219 | 15.9115 | 21.4808 | 5.8965 | 136.1921 | 276.4999 | 212.9074 | 271.352 | 188.8259
2012-10-30 00:45 | 25.961 | 298.0006 | 234.4163 | 292.8643 | 210.3742 | 85.532 | 1.444 | 36.5374 | 52.7595 | 25.3031 | 2.2895 | 36.4294 | 83.4053 | 15.9115 | 21.4773 | 5.8966 | 136.1648 | 276.5233 | 212.939 | 271.387 | 188.8969
2012-10-30 00:45 | 25.9595 | 297.9722 | 234.4108 | 292.8589 | 210.2821 | 85.5274 | 1.444 | 36.5372 | 52.7589 | 25.3032 | 2.2895 | 36.4293 | 83.4051 | 15.9114 | 21.4798 | 5.8963 | 136.1639 | 276.4924 | 212.931 | 271.3791 | 188.8023
2012-10-30 00:45 | 25.9617 | 297.9821 | 234.4052 | 292.8111 | 210.3434 | 85.5528 | 1.444 | 36.5365 | 52.7571 | 25.3031 | 2.2895 | 36.4286 | 83.4018 | 15.9114 | 21.486 | 5.898 | 136.1589 | 276.4961 | 212.9192 | 271.3251 | 188.8574
2012-10-30 00:45 | 25.9662 | 298.0212 | 234.4254 | 292.8156 | 210.2652 | 85.5343 | 1.4439 | 36.5355 | 52.7539 | 25.3031 | 2.2894 | 36.4276 | 83.3981 | 15.9112 | 21.5066 | 5.8967 | 136.1521 | 276.5146 | 212.9188 | 271.309 | 188.7586
2012-10-30 00:45 | 25.9595 | 298.0068 | 234.3951 | 292.8166 | 210.3885 | 85.5268 | 1.4439 | 36.5357 | 52.7551 | 25.3029 | 2.2894 | 36.4276 | 83.3986 | 15.9112 | 21.4838 | 5.8962 | 136.1537 | 276.523 | 212.9113 | 271.3328 | 188.9047
2012-10-30 00:46 | 25.9707 | 298.0141 | 234.4613 | 292.8586 | 210.4981 | 85.5424 | 1.4438 | 36.5323 | 52.7448 | 25.3032 | 2.2892 | 36.4245 | 83.3839 | 15.9112 | 21.475 | 5.8973 | 136.1287 | 276.5391 | 212.9863 | 271.3836 | 189.0231
2012-10-30 00:46 | 25.9692 | 298.0203 | 234.5305 | 292.884 | 210.5872 | 85.5401 | 1.4437 | 36.529 | 52.7356 | 25.3029 | 2.289 | 36.4209 | 83.3679 | 15.9112 | 21.5457 | 5.8971 | 136.1034 | 276.4746 | 212.9848 | 271.3383 | 189.0415
2012-10-30 00:46 | 25.976 | 298.0924 | 234.5216 | 292.8947 | 210.531 | 85.543 | 1.4436 | 36.5272 | 52.7297 | 25.3034 | 2.2889 | 36.4194 | 83.3611 | 15.9111 | 21.5204 | 5.8973 | 136.0908 | 276.572 | 213.0012 | 271.3743 | 189.0106
2012-10-30 00:46 | 25.9647 | 298.0312 | 234.5457 | 292.891 | 210.5315 | 85.5487 | 1.4435 | 36.526 | 52.7263 | 25.3032 | 2.2888 | 36.4182 | 83.3555 | 15.9112 | 21.5171 | 5.8977 | 136.0819 | 276.5141 | 213.0286 | 271.3739 | 189.0144
2012-10-30 00:46 | 25.9752 | 298.0763 | 234.5522 | 292.8939 | 210.4829 | 85.5401 | 1.4435 | 36.5251 | 52.7244 | 25.3029 | 2.2888 | 36.4171 | 83.3511 | 15.9111 | 21.5116 | 5.8971 | 136.0755 | 276.5647 | 213.0406 | 271.3823 | 188.9713
2012-10-30 00:46 | 25.976 | 298.0617 | 234.5688 | 292.8754 | 210.4128 | 85.5522 | 1.4435 | 36.5243 | 52.7217 | 25.3031 | 2.2887 | 36.4164 | 83.3478 | 15.9111 | 21.4923 | 5.898 | 136.0695 | 276.5694 | 213.0765 | 271.3831 | 188.9205
2012-10-30 00:46 | 25.9797 | 298.0577 | 234.5685 | 292.9176 | 210.5111 | 85.5568 | 1.4434 | 36.5223 | 52.7159 | 25.3032 | 2.2886 | 36.4144 | 83.3389 | 15.911 | 21.5361 | 5.8983 | 136.0547 | 276.5216 | 213.0324 | 271.3815 | 188.975
2012-10-30 00:46 | 25.9774 | 298.0977 | 234.5859 | 292.8691 | 210.7135 | 85.6216 | 1.524 | 38.601 | 58.8285 | 25.3284 | 2.416 | 38.4577 | 92.9145 | 15.9178 | 21.5017 | 5.9028 | 151.743 | 276.596 | 213.0842 | 271.3674 | 189.2118
2012-10-30 00:46 | 25.9819 | 298.3441 | 234.9439 | 292.8698 | 210.4896 | 85.9429 | 1.553 | 39.3643 | 61.1309 | 25.3481 | 2.4648 | 39.2478 | 96.7382 | 15.9233 | 21.5102 | 5.9249 | 157.8691 | 276.8339 | 213.4337 | 271.3596 | 188.9794
2012-10-30 00:46 | 25.9902 | 298.8517 | 235.2074 | 292.9897 | 210.4152 | 85.9984 | 1.5503 | 39.2993 | 60.9265 | 25.3491 | 2.4607 | 39.1861 | 96.4265 | 15.9245 | 21.5105 | 5.9287 | 157.353 | 277.3412 | 213.6969 | 271.4792 | 188.9047
2012-10-30 00:46 | 25.9887 | 299.2687 | 235.3827 | 293.1074 | 210.5516 | 86.0232 | 1.5489 | 39.2628 | 60.813 | 25.3492 | 2.4581 | 39.1468 | 96.2284 | 15.9253 | 21.4809 | 5.9304 | 157.0415 | 277.7878 | 213.9018 | 271.6265 | 189.0707
2012-10-30 00:46 | 25.9917 | 299.5904 | 235.4956 | 293.2567 | 210.7436 | 86.0463 | 1.5483 | 39.2506 | 60.7731 | 25.3502 | 2.4572 | 39.1343 | 96.1612 | 15.9263 | 21.5321 | 5.932 | 156.9343 | 278.0583 | 213.9635 | 271.7246 | 189.2115
2012-10-30 00:46 | 25.9924 | 299.8559 | 235.622 | 293.3959 | 210.5356 | 86.0764 | 1.5481 | 39.2449 | 60.7545 | 25.3506 | 2.4568 | 39.1286 | 96.1295 | 15.9269 | 21.5087 | 5.9341 | 156.884 | 278.3472 | 214.1133 | 271.8872 | 189.0269
2012-10-30 00:46 | 25.9872 | 300.177 | 235.6639 | 293.61 | 208.7229 | 86.1053 | 1.5481 | 39.2454 | 60.7548 | 25.3511 | 2.4567 | 39.1293 | 96.129 | 15.9276 | 21.4834 | 5.9361 | 156.8838 | 278.6936 | 214.1805 | 272.1266 | 187.2395
2012-10-30 00:46 | 25.9999 | 300.539 | 235.5981 | 293.9073 | 209.5428 | 86.1325 | 1.5482 | 39.2497 | 60.7655 | 25.3522 | 2.4568 | 39.1334 | 96.1445 | 15.9283 | 21.5002 | 5.938 | 156.91 | 279.0388 | 214.0979 | 272.4071 | 188.0426
2012-10-30 00:46 | 25.9954 | 300.83 | 235.676 | 294.0605 | 210.4834 | 86.1625 | 1.5483 | 39.2551 | 60.7802 | 25.3531 | 2.4571 | 39.1388 | 96.1671 | 15.929 | 21.4716 | 5.94 | 156.9472 | 279.3584 | 214.2044 | 272.5889 | 189.0118
2012-10-30 00:46 | 25.9917 | 301.1371 | 235.8569 | 294.2453 | 210.4404 | 86.192 | 1.5485 | 39.2616 | 60.7981 | 25.3539 | 2.4574 | 39.1452 | 96.1952 | 15.9295 | 21.4879 | 5.9421 | 156.9933 | 279.6492 | 214.369 | 272.7574 | 188.9525
2012-10-30 00:46 | 25.9894 | 301.388 | 235.992 | 294.5545 | 210.8909 | 86.2134 | 1.5487 | 39.2665 | 60.8117 | 25.3546 | 2.4576 | 39.1502 | 96.2161 | 15.9302 | 21.4856 | 5.9436 | 157.0278 | 279.9024 | 214.5064 | 273.0689 | 189.4053
2012-10-30 00:46 | 25.9932 | 301.7486 | 236.2432 | 294.8122 | 211.6233 | 86.2463 | 1.5508 | 39.3211 | 60.9778 | 25.3559 | 2.4608 | 39.2024 | 96.4676 | 15.931 | 21.4654 | 5.9458 | 157.4454 | 280.2832 | 214.7778 | 273.3468 | 190.1579
2012-10-30 00:46 | 25.9947 | 302.0033 | 236.4528 | 295.0713 | 211.4042 | 86.2735 | 1.5521 | 39.3568 | 61.085 | 25.3574 | 2.463 | 39.2401 | 96.6501 | 15.9316 | 21.4749 | 5.9477 | 157.7351 | 280.5284 | 214.9779 | 273.5964 | 189.9293
2012-10-30 00:47 | 25.9947 | 302.2488 | 236.5902 | 295.3251 | 211.2546 | 86.3185 | 1.5526 | 39.3695 | 61.1233 | 25.358 | 2.4638 | 39.2529 | 96.7101 | 15.9321 | 21.4709 | 5.9508 | 157.8334 | 280.7779 | 215.1193 | 273.8542 | 189.7837
2012-10-30 00:47 | 25.9857 | 302.4699 | 236.75 | 295.5083 | 211.411 | 86.3382 | 1.5529 | 39.3806 | 61.1557 | 25.3587 | 2.4644 | 39.2638 | 96.761 | 15.9326 | 21.4498 | 5.9522 | 157.9166 | 281.0201 | 215.3002 | 274.0585 | 189.9612
2012-10-30 00:47 | 25.9909 | 302.7627 | 236.924 | 295.7787 | 211.8179 | 86.363 | 1.5532 | 39.3893 | 61.1805 | 25.3597 | 2.4648 | 39.2725 | 96.8 | 15.9331 | 21.439 | 5.9539 | 157.9806 | 281.3237 | 215.485 | 274.3397 | 190.3789
2012-10-30 00:47 | 25.9872 | 303.0428 | 237.0891 | 296.0864 | 211.4676 | 86.3966 | 1.5534 | 39.3947 | 61.1957 | 25.3603 | 2.4651 | 39.2779 | 96.8235 | 15.9336 | 21.4754 | 5.9562 | 158.0191 | 281.5674 | 215.6137 | 274.611 | 189.9922
2012-10-30 00:47 | 25.976 | 303.2617 | 237.2546 | 296.2674 | 211.3973 | 86.4405 | 1.5534 | 39.3965 | 61.1996 | 25.361 | 2.4651 | 39.2797 | 96.8303 | 15.934 | 21.4883 | 5.9592 | 158.0298 | 281.7734 | 215.7663 | 274.7791 | 189.909
2012-10-30 00:47 | 25.9924 | 303.5005 | 237.2985 | 296.5605 | 211.8981 | 86.4463 | 1.5534 | 39.3975 | 61.2015 | 25.3616 | 2.4652 | 39.2808 | 96.8337 | 15.9344 | 21.4967 | 5.9596 | 158.0352 | 282.0038 | 215.8018 | 275.0638 | 190.4014
2012-10-30 00:47 | 25.9864 | 303.7628 | 237.528 | 296.8542 | 212.1914 | 86.4676 | 1.5534 | 39.3982 | 61.2017 | 25.3624 | 2.4651 | 39.2814 | 96.8331 | 15.9349 | 21.4505 | 5.9611 | 158.0347 | 282.3123 | 216.0775 | 275.4037 | 190.7409
2012-10-30 00:47 | 25.9909 | 304.0242 | 237.7248 | 297.1315 | 212.507 | 86.5121 | 1.5534 | 39.3977 | 61.1985 | 25.363 | 2.465 | 39.2808 | 96.8278 | 15.9353 | 21.3868 | 5.9641 | 158.0262 | 282.6374 | 216.338 | 275.7447 | 191.1202
2012-10-30 00:47 | 25.9894 | 304.2297 | 237.8999 | 297.3567 | 212.667 | 86.5266 | 1.5533 | 39.3968 | 61.195 | 25.3633 | 2.4649 | 39.2801 | 96.8222 | 15.9356 | 21.4495 | 5.9651 | 158.0172 | 282.7802 | 216.4504 | 275.9072 | 191.2175
2012-10-30 00:47 | 25.9939 | 304.4565 | 238.0693 | 297.5956 | 212.9432 | 86.5416 | 1.5532 | 39.3964 | 61.1922 | 25.364 | 2.4648 | 39.2797 | 96.8177 | 15.9361 | 21.438 | 5.9662 | 158.0099 | 283.0185 | 216.6313 | 276.1576 | 191.5052
2012-10-30 00:47 | 25.9954 | 304.688 | 238.2749 | 297.8237 | 213.244 | 86.5803 | 1.5532 | 39.3968 | 61.1924 | 25.3644 | 2.4648 | 39.28 | 96.8171 | 15.9365 | 21.4154 | 5.9688 | 158.0095 | 283.2726 | 216.8595 | 276.4083 | 191.8286
2012-10-30 00:47 | 25.9999 | 304.8726 | 238.5266 | 298.0741 | 213.3784 | 86.6121 | 1.5532 | 39.3959 | 61.1881 | 25.3651 | 2.4647 | 39.2792 | 96.8109 | 15.9368 | 21.4761 | 5.971 | 157.999 | 283.3965 | 217.0505 | 276.598 | 191.9023
2012-10-30 00:47 | 26.0059 | 305.12 | 238.721 | 298.3796 | 213.5459 | 86.6358 | 1.5531 | 39.3946 | 61.1828 | 25.3656 | 2.4645 | 39.2777 | 96.8013 | 15.9372 | 21.4098 | 5.9727 | 157.9841 | 283.7102 | 217.3112 | 276.9698 | 192.1361
2012-10-30 00:47 | 26.0014 | 305.288 | 238.9244 | 298.5557 | 213.7855 | 86.6509 | 1.5529 | 39.3926 | 61.1747 | 25.3663 | 2.4644 | 39.2758 | 96.7898 | 15.9375 | 21.4535 | 5.9737 | 157.9645 | 283.8345 | 217.4709 | 277.1022 | 192.332
2012-10-30 00:47 | 26.0111 | 305.4931 | 239.0871 | 298.7919 | 214.067 | 86.6797 | 1.5529 | 39.3909 | 61.1685 | 25.3666 | 2.4642 | 39.274 | 96.7785 | 15.9379 | 21.4913 | 5.9757 | 157.947 | 284.0018 | 217.5958 | 277.3006 | 192.5757
2012-10-30 00:47 | 26.0111 | 305.7268 | 239.3069 | 299.0799 | 214.2363 | 86.704 | 1.5527 | 39.3889 | 61.1609 | 25.3673 | 2.464 | 39.2721 | 96.7667 | 15.9383 | 21.4511 | 5.9774 | 157.9276 | 284.2757 | 217.8558 | 277.6288 | 192.7852
2012-10-30 00:47 | 26.0089 | 305.9621 | 239.544 | 299.3004 | 214.2932 | 86.7277 | 1.5527 | 39.3885 | 61.1583 | 25.3678 | 2.4639 | 39.2717 | 96.763 | 15.9386 | 21.473 | 5.979 | 157.9214 | 284.4891 | 218.071 | 277.8274 | 192.8202
2012-10-30 00:47 | 26.0164 | 306.1267 | 239.6849 | 299.5383 | 214.1904 | 86.7451 | 1.5527 | 39.3883 | 61.1571 | 25.3681 | 2.4639 | 39.2715 | 96.7608 | 15.9388 | 21.4685 | 5.9802 | 157.918 | 284.6582 | 218.2164 | 278.0698 | 192.7219
2012-10-30 00:47 | 26.0238 | 306.3218 | 239.81 | 299.6992 | 214.2451 | 86.7664 | 1.5526 | 39.388 | 61.1547 | 25.3687 | 2.4638 | 39.2713 | 96.7582 | 15.9391 | 21.4237 | 5.9817 | 157.9129 | 284.8981 | 218.3863 | 278.2755 | 192.8214
2012-10-30 00:47 | 26.0246 | 306.518 | 239.9363 | 299.9495 | 214.4073 | 86.7826 | 1.5525 | 39.386 | 61.1478 | 25.369 | 2.4637 | 39.2692 | 96.7463 | 15.9393 | 21.4767 | 5.9828 | 157.8942 | 285.0413 | 218.4596 | 278.4728 | 192.9306
2012-10-30 00:47 | 26.0306 | 306.7346 | 240.1306 | 300.1551 | 214.7086 | 86.8057 | 1.5525 | 39.3857 | 61.1454 | 25.3696 | 2.4636 | 39.2688 | 96.7426 | 15.9396 | 21.4506 | 5.9844 | 157.888 | 285.284 | 218.68 | 278.7045 | 193.258
2012-10-30 00:47 | 26.0261 | 306.9216 | 240.2948 | 300.4193 | 214.7829 | 86.8445 | 1.5524 | 39.3851 | 61.1423 | 25.3701 | 2.4635 | 39.2683 | 96.7385 | 15.9399 | 21.43 | 5.9871 | 157.8808 | 285.4916 | 218.8648 | 278.9893 | 193.3529
2012-10-30 00:47 | 26.0351 | 307.1489 | 240.4763 | 300.6586 | 215.0833 | 86.8635 | 1.5523 | 39.3838 | 61.1366 | 25.3708 | 2.4634 | 39.2671 | 96.7298 | 15.9404 | 21.427 | 5.9884 | 157.8663 | 285.7219 | 219.0493 | 279.2316 | 193.6563
2012-10-30 00:47 | 26.0306 | 307.3551 | 240.6287 | 300.8536 | 215.2481 | 86.8838 | 1.5523 | 39.3836 | 61.1351 | 25.3711 | 2.4633 | 39.2668 | 96.7276 | 15.9405 | 21.4225 | 5.9898 | 157.8628 | 285.9326 | 219.2062 | 279.4311 | 193.8256
2012-10-30 00:48 | 26.0328 | 307.5373 | 240.831 | 301.0746 | 215.4197 | 86.9051 | 1.5522 | 39.3823 | 61.1299 | 25.3716 | 2.4632 | 39.2654 | 96.718 | 15.9409 | 21.4889 | 5.9912 | 157.8479 | 286.0484 | 219.3421 | 279.5857 | 193.9308
2012-10-30 00:48 | 26.0455 | 307.7606 | 241.0633 | 301.352 | 215.6214 | 86.934 | 1.5521 | 39.38 | 61.1214 | 25.3722 | 2.463 | 39.2634 | 96.7064 | 15.9412 | 21.4575 | 5.9932 | 157.8279 | 286.3031 | 219.6058 | 279.8945 | 194.1639
2012-10-30 00:48 | 26.0388 | 307.9262 | 241.1781 | 301.5525 | 215.7013 | 86.9572 | 1.5521 | 39.3798 | 61.1198 | 25.3726 | 2.463 | 39.2631 | 96.7034 | 15.9414 | 21.4708 | 5.9948 | 157.8233 | 286.4554 | 219.7073 | 280.0817 | 194.2305
2012-10-30 00:48 | 26.0411 | 308.1084 | 241.353 | 301.7121 | 215.8099 | 86.9832 | 1.552 | 39.379 | 61.1164 | 25.373 | 2.4629 | 39.2623 | 96.6979 | 15.9417 | 21.4611 | 5.9966 | 157.8143 | 286.6473 | 219.8919 | 280.251 | 194.3488
2012-10-30 00:48 | 26.0441 | 308.2492 | 241.5128 | 301.8954 | 215.9153 | 87.0005 | 1.5519 | 39.3767 | 61.1084 | 25.3733 | 2.4627 | 39.2601 | 96.6856 | 15.942 | 21.4922 | 5.9978 | 157.794 | 286.757 | 220.0206 | 280.4032 | 194.4231
2012-10-30 00:48 | 26.0553 | 308.4289 | 241.6338 | 302.1291 | 216.1746 | 87.0173 | 1.5518 | 39.3742 | 61.0996 | 25.3737 | 2.4625 | 39.2575 | 96.671 | 15.9422 | 21.4632 | 5.999 | 157.7707 | 286.9657 | 220.1706 | 280.6659 | 194.7114
2012-10-30 00:48 | 26.0448 | 308.5754 | 241.788 | 302.3142 | 216.2231 | 87.0369 | 1.5517 | 39.373 | 61.096 | 25.3737 | 2.4624 | 39.2564 | 96.6643 | 15.9424 | 21.4648 | 6.0003 | 157.7603 | 287.1106 | 220.3232 | 280.8494 | 194.7583
2012-10-30 00:48 | 26.05 | 308.7375 | 241.9344 | 302.5112 | 216.3701 | 87.0514 | 1.5516 | 39.3716 | 61.0902 | 25.3743 | 2.4623 | 39.2548 | 96.6553 | 15.9426 | 21.5021 | 6.0013 | 157.7455 | 287.2354 | 220.4323 | 281.0091 | 194.868
2012-10-30 00:48 | 26.0485 | 308.9159 | 242.0348 | 302.6362 | 216.5104 | 87.0762 | 1.5516 | 39.3709 | 61.0877 | 25.3745 | 2.4622 | 39.2542 | 96.651 | 15.9429 | 21.4845 | 6.003 | 157.7388 | 287.4314 | 220.5503 | 281.1517 | 195.0259
2012-10-30 00:48 | 26.0545 | 309.1171 | 242.2211 | 302.8378 | 216.4927 | 87.1022 | 1.5515 | 39.3694 | 61.082 | 25.375 | 2.4621 | 39.2527 | 96.6428 | 15.943 | 21.4945 | 6.0048 | 157.7248 | 287.6226 | 220.7266 | 281.3433 | 194.9982
2012-10-30 00:48 | 26.0598 | 309.2717 | 242.3205 | 303.0579 | 216.6319 | 87.1334 | 1.5514 | 39.3682 | 61.0774 | 25.3753 | 2.462 | 39.2517 | 96.6368 | 15.9431 | 21.4718 | 6.007 | 157.7142 | 287.7999 | 220.8487 | 281.5861 | 195.1601
2012-10-30 00:48 | 26.0598 | 309.4401 | 242.4342 | 303.2189 | 216.6949 | 87.1346 | 1.5514 | 39.3668 | 61.072 | 25.3757 | 2.4619 | 39.2501 | 96.6282 | 15.9433 | 21.4637 | 6.0071 | 157.7003 | 287.9764 | 220.9705 | 281.7552 | 195.2312
2012-10-30 00:48 | 26.0553 | 309.5887 | 242.5748 | 303.3563 | 216.7849 | 87.1456 | 1.5494 | 39.3178 | 60.9205 | 25.3756 | 2.4589 | 39.2027 | 96.3949 | 15.9433 | 21.4592 | 6.0078 | 157.3154 | 288.1295 | 221.1156 | 281.8971 | 195.3257
2012-10-30 00:48 | 26.0635 | 309.6772 | 242.6261 | 303.5486 | 216.8955 | 87.1531 | 1.5474 | 39.266 | 60.7613 | 25.375 | 2.4556 | 39.1498 | 96.1352 | 15.9433 | 21.4554 | 6.0083 | 156.8965 | 288.2218 | 221.1707 | 282.0932 | 195.4401
2012-10-30 00:48 | 26.059 | 309.8717 | 242.7549 | 303.6859 | 217.0879 | 87.1721 | 1.547 | 39.2543 | 60.7253 | 25.3749 | 2.4548 | 39.138 | 96.0768 | 15.9433 | 21.4589 | 6.0096 | 156.8021 | 288.4128 | 221.296 | 282.227 | 195.629
2012-10-30 00:48 | 26.0635 | 309.9872 | 242.8691 | 303.8668 | 216.3797 | 87.171 | 1.5467 | 39.2489 | 60.7079 | 25.3752 | 2.4545 | 39.1327 | 96.0495 | 15.9435 | 21.4554 | 6.0096 | 156.7574 | 288.5318 | 221.4137 | 282.4114 | 194.9243
2012-10-30 00:48 | 26.0553 | 310.1321 | 242.8374 | 304.0119 | 216.2809 | 87.1918 | 1.5468 | 39.2509 | 60.7132 | 25.3756 | 2.4546 | 39.1347 | 96.0588 | 15.9436 | 21.4351 | 6.011 | 156.772 | 288.697 | 221.4023 | 282.5768 | 194.8458
2012-10-30 00:48 | 26.0523 | 310.263 | 242.8383 | 304.1737 | 216.9985 | 87.2149 | 1.5469 | 39.2545 | 60.7239 | 25.3757 | 2.4548 | 39.1382 | 96.0748 | 15.9438 | 21.4402 | 6.0126 | 156.7987 | 288.8228 | 221.3981 | 282.7335 | 195.5583
2012-10-30 00:48 | 26.056 | 310.4006 | 242.9753 | 304.3269 | 217.0455 | 87.2195 | 1.5471 | 39.2606 | 60.7416 | 25.3763 | 2.4551 | 39.1443 | 96.1031 | 15.9441 | 21.488 | 6.0129 | 156.8447 | 288.9126 | 221.4873 | 282.8389 | 195.5575
2012-10-30 00:48 | 26.0583 | 310.5559 | 243.056 | 304.5056 | 217.0754 | 87.2346 | 1.5473 | 39.2653 | 60.756 | 25.3764 | 2.4554 | 39.1489 | 96.1247 | 15.9442 | 21.4903 | 6.014 | 156.8807 | 289.0656 | 221.5657 | 283.0153 | 195.5851
2012-10-30 00:48 | 26.053 | 310.6539 | 243.1408 | 304.6344 | 216.885 | 87.2548 | 1.5474 | 39.268 | 60.7632 | 25.3768 | 2.4555 | 39.1516 | 96.1372 | 15.9444 | 21.4529 | 6.0153 | 156.9004 | 289.201 | 221.6879 | 283.1815 | 195.4321
2012-10-30 00:48 | 26.0545 | 310.8046 | 243.1894 | 304.7624 | 216.7211 | 87.2606 | 1.5475 | 39.272 | 60.7744 | 25.3774 | 2.4557 | 39.1554 | 96.1546 | 15.9446 | 21.4945 | 6.0157 | 156.929 | 289.3101 | 221.6949 | 283.2679 | 195.2266
2012-10-30 00:48 | 26.0463 | 310.9418 | 243.1772 | 304.9343 | 216.6381 | 87.2889 | 1.5507 | 39.3551 | 61.0282 | 25.3788 | 2.4607 | 39.2362 | 96.5485 | 15.9451 | 21.4904 | 6.0177 | 157.5767 | 289.4514 | 221.6868 | 283.4439 | 195.1477
2012-10-30 00:49 | 26.0455 | 311.0749 | 243.2392 | 305.037 | 217.35 | 87.3114 | 1.5512 | 39.3689 | 61.0697 | 25.3794 | 2.4616 | 39.2519 | 96.6238 | 15.9454 | 21.4735 | 6.0192 | 157.6935 | 289.6014 | 221.7657 | 283.5635 | 195.8765
2012-10-30 00:49 | 26.0448 | 311.189 | 243.3639 | 305.1819 | 217.5383 | 87.3207 | 1.5516 | 39.3804 | 61.1045 | 25.3798 | 2.4623 | 39.2636 | 96.6797 | 15.9458 | 21.4808 | 6.0199 | 157.7841 | 289.7082 | 221.8831 | 283.7011 | 196.0575
2012-10-30 00:49 | 26.0478 | 311.3909 | 243.5471 | 305.2845 | 217.7342 | 87.3409 | 1.552 | 39.3901 | 61.1333 | 25.3802 | 2.4629 | 39.2733 | 96.7265 | 15.9459 | 21.4798 | 6.0213 | 157.8598 | 289.9111 | 222.0673 | 283.8047 | 196.2544
2012-10-30 00:49 | 26.0388 | 311.489 | 243.6518 | 305.4556 | 218.1032 | 87.3536 | 1.5522 | 39.3962 | 61.1512 | 25.3806 | 2.4633 | 39.2796 | 96.7562 | 15.9461 | 21.4708 | 6.0222 | 157.9074 | 290.0182 | 222.181 | 283.9848 | 196.6324
2012-10-30 00:49 | 26.0336 | 311.6215 | 243.8386 | 305.5999 | 218.2043 | 87.3779 | 1.5523 | 39.4001 | 61.1622 | 25.3811 | 2.4635 | 39.2833 | 96.7733 | 15.9464 | 21.4817 | 6.0238 | 157.9355 | 290.1398 | 222.3569 | 284.1182 | 196.7226
2012-10-30 00:49 | 26.0298 | 311.8127 | 243.9759 | 305.7187 | 218.1611 | 87.3952 | 1.5523 | 39.4007 | 61.1633 | 25.3814 | 2.4635 | 39.2839 | 96.7747 | 15.9465 | 21.4698 | 6.025 | 157.938 | 290.3429 | 222.5061 | 284.2489 | 196.6913
2012-10-30 00:49 | 26.0313 | 311.864 | 244.1067 | 305.8581 | 218.2729 | 87.4154 | 1.5524 | 39.4014 | 61.1648 | 25.3818 | 2.4635 | 39.2847 | 96.7775 | 15.9468 | 21.4633 | 6.0264 | 157.9423 | 290.4007 | 222.6434 | 284.3948 | 196.8096
2012-10-30 00:49 | 26.0298 | 312.0231 | 244.1718 | 305.9524 | 218.232 | 87.4339 | 1.5524 | 39.4018 | 61.1656 | 25.3819 | 2.4635 | 39.2849 | 96.7779 | 15.9469 | 21.4578 | 6.0277 | 157.9435 | 290.5653 | 222.714 | 284.4946 | 196.7742
2012-10-30 00:49 | 26.0343 | 312.1194 | 244.2978 | 306.091 | 218.3586 | 87.4339 | 1.5523 | 39.401 | 61.1624 | 25.3822 | 2.4634 | 39.2842 | 96.7731 | 15.9471 | 21.4543 | 6.0277 | 157.9355 | 290.6651 | 222.8435 | 284.6367 | 196.9043
2012-10-30 00:49 | 26.0238 | 312.2924 | 244.3891 | 306.2146 | 218.5213 | 87.449 | 1.5522 | 39.3999 | 61.158 | 25.3826 | 2.4633 | 39.2832 | 96.7674 | 15.9472 | 21.4638 | 6.0287 | 157.9253 | 290.8286 | 222.9253 | 284.7508 | 197.0575
2012-10-30 00:49 | 26.0283 | 312.3849 | 244.5503 | 306.3953 | 218.6794 | 87.4721 | 1.5522 | 39.4004 | 61.1593 | 25.3828 | 2.4633 | 39.2838 | 96.7692 | 15.9474 | 21.4483 | 6.0303 | 157.9285 | 290.9366 | 223.102 | 284.947 | 197.2311
2012-10-30 00:49 | 26.0261 | 312.4936 | 244.6461 | 306.4544 | 218.7559 | 87.4819 | 1.5522 | 39.3992 | 61.1542 | 25.3833 | 2.4632 | 39.2822 | 96.7602 | 15.9476 | 21.4541 | 6.031 | 157.9144 | 291.0395 | 223.192 | 285.0003 | 197.3018
2012-10-30 00:49 | 26.0201 | 312.5755 | 244.7302 | 306.6131 | 218.8995 | 87.5056 | 1.5521 | 39.3993 | 61.1536 | 25.3837 | 2.4632 | 39.2825 | 96.761 | 15.9477 | 21.4762 | 6.0326 | 157.9146 | 291.0993 | 223.254 | 285.1369 | 197.4233
2012-10-30 00:49 | 26.0298 | 312.7305 | 244.8927 | 306.7492 | 219.0785 | 87.5224 | 1.5521 | 39.3986 | 61.1517 | 25.3836 | 2.4631 | 39.2819 | 96.7568 | 15.9479 | 21.4618 | 6.0338 | 157.9085 | 291.2687 | 223.4309 | 285.2874 | 197.6167
2012-10-30 00:49 | 26.0336 | 312.8452 | 245.0532 | 306.8985 | 219.1099 | 87.5391 | 1.552 | 39.3968 | 61.1454 | 25.3839 | 2.463 | 39.2803 | 96.7479 | 15.9481 | 21.4456 | 6.0349 | 157.8933 | 291.3996 | 223.6076 | 285.4529 | 197.6643
2012-10-30 00:49 | 26.0306 | 312.9798 | 245.1089 | 306.9989 | 219.2171 | 87.5594 | 1.552 | 39.3956 | 61.141 | 25.3841 | 2.4629 | 39.2787 | 96.7401 | 15.9481 | 21.4787 | 6.0363 | 157.8811 | 291.5011 | 223.6302 | 285.5202 | 197.7384
2012-10-30 00:49 | 26.0306 | 313.064 | 245.246 | 307.1406 | 219.5006 | 87.5634 | 1.5519 | 39.3948 | 61.1375 | 25.3846 | 2.4628 | 39.2781 | 96.7354 | 15.9483 | 21.4706 | 6.0366 | 157.8728 | 291.5934 | 223.7754 | 285.67 | 198.03
2012-10-30 00:49 | 26.0321 | 313.2107 | 245.4159 | 307.28 | 219.632 | 87.5773 | 1.5519 | 39.3942 | 61.1349 | 25.3849 | 2.4628 | 39.2774 | 96.7313 | 15.9485 | 21.4922 | 6.0376 | 157.8662 | 291.7185 | 223.9237 | 285.7878 | 198.1398
2012-10-30 00:49 | 26.0388 | 313.3283 | 245.5127 | 307.4322 | 219.6229 | 87.5952 | 1.5518 | 39.3922 | 61.128 | 25.3852 | 2.4626 | 39.2753 | 96.7201 | 15.9486 | 21.4949 | 6.0388 | 157.848 | 291.8334 | 224.0178 | 285.9373 | 198.128
2012-10-30 00:49 | 26.0298 | 313.4416 | 245.6565 | 307.5189 | 219.6572 | 87.6044 | 1.5517 | 39.3908 | 61.1236 | 25.3852 | 2.4625 | 39.2741 | 96.7132 | 15.9487 | 21.4698 | 6.0394 | 157.8369 | 291.9718 | 224.1867 | 286.0491 | 198.1874
2012-10-30 00:49 | 26.0246 | 313.5435 | 245.7218 | 307.6401 | 219.6127 | 87.6247 | 1.5516 | 39.3893 | 61.1182 | 25.3855 | 2.4624 | 39.2726 | 96.705 | 15.9489 | 21.4887 | 6.0408 | 157.8232 | 292.0548 | 224.2331 | 286.1514 | 198.124
2012-10-30 00:49 | 26.0216 | 313.6246 | 245.8088 | 307.7099 | 219.653 | 87.6426 | 1.5516 | 39.3886 | 61.1152 | 25.3859 | 2.4623 | 39.2718 | 96.6995 | 15.9491 | 21.4777 | 6.0421 | 157.8147 | 292.1469 | 224.3311 | 286.2322 | 198.1753
2012-10-30 00:49 | 26.0276 | 313.7758 | 245.9167 | 307.8652 | 219.8952 | 87.6495 | 1.5515 | 39.3872 | 61.1094 | 25.3865 | 2.4622 | 39.2704 | 96.6924 | 15.9492 | 21.4917 | 6.0426 | 157.8018 | 292.2841 | 224.425 | 286.3735 | 198.4035
2012-10-30 00:50 | 26.0366 | 313.8919 | 245.9805 | 308.0235 | 220.1522 | 87.6668 | 1.5515 | 39.386 | 61.106 | 25.3864 | 2.4621 | 39.2693 | 96.6864 | 15.9493 | 21.4887 | 6.0437 | 157.7924 | 292.4032 | 224.4918 | 286.5348 | 198.6635
2012-10-30 00:50 | 26.0291 | 313.9761 | 246.1179 | 308.1002 | 220.125 | 87.6807 | 1.5514 | 39.3841 | 61.0999 | 25.3864 | 2.462 | 39.2675 | 96.6762 | 15.9495 | 21.4611 | 6.0447 | 157.7761 | 292.515 | 224.6568 | 286.6391 | 198.6639
2012-10-30 00:50 | 26.0291 | 314.1443 | 246.1845 | 308.2304 | 220.0896 | 87.6842 | 1.5513 | 39.3832 | 61.0962 | 25.3868 | 2.4619 | 39.2665 | 96.6709 | 15.9496 | 21.4732 | 6.0449 | 157.767 | 292.6711 | 224.7113 | 286.7572 | 198.6164
2012-10-30 00:50 | 26.0358 | 314.2504 | 246.2812 | 308.3405 | 220.1947 | 87.6998 | 1.5513 | 39.3822 | 61.0927 | 25.3869 | 2.4618 | 39.2655 | 96.6654 | 15.9497 | 21.4758 | 6.046 | 157.7581 | 292.7746 | 224.8054 | 286.8647 | 198.7189
2012-10-30 00:50 | 26.0253 | 314.301 | 246.353 | 308.418 | 220.2314 | 87.7154 | 1.5512 | 39.3808 | 61.0875 | 25.3873 | 2.4617 | 39.2641 | 96.6577 | 15.9498 | 21.4934 | 6.0471 | 157.7452 | 292.8076 | 224.8596 | 286.9246 | 198.738
2012-10-30 00:50 | 26.0283 | 314.3575 | 246.4304 | 308.4899 | 220.1596 | 87.7246 | 1.5512 | 39.3809 | 61.0874 | 25.3875 | 2.4617 | 39.2642 | 96.6577 | 15.9499 | 21.4523 | 6.0477 | 157.7451 | 292.9052 | 224.9781 | 287.0376 | 198.7073
2012-10-30 00:50 | 26.0209 | 314.4457 | 246.4778 | 308.605 | 220.3176 | 87.7235 | 1.5511 | 39.3786 | 61.0795 | 25.3878 | 2.4616 | 39.2619 | 96.6464 | 15.9499 | 21.4569 | 6.0477 | 157.7259 | 292.9888 | 225.0209 | 287.1481 | 198.8607
2012-10-30 00:50 | 26.0276 | 314.5403 | 246.5471 | 308.7113 | 220.3715 | 87.7356 | 1.551 | 39.377 | 61.0742 | 25.388 | 2.4615 | 39.2604 | 96.6377 | 15.9501 | 21.4837 | 6.0485 | 157.7119 | 293.0566 | 225.0634 | 287.2276 | 198.8878
2012-10-30 00:50 | 26.0231 | 314.681 | 246.6405 | 308.8063 | 220.6701 | 87.761 | 1.5509 | 39.3753 | 61.0687 | 25.3881 | 2.4613 | 39.2587 | 96.6288 | 15.9502 | 21.4631 | 6.0502 | 157.6975 | 293.2179 | 225.1774 | 287.3432 | 199.207
2012-10-30 00:50 | 26.0298 | 314.8062 | 246.7451 | 308.9508 | 220.7083 | 87.7576 | 1.5509 | 39.3752 | 61.0678 | 25.3883 | 2.4613 | 39.2586 | 96.6276 | 15.9503 | 21.4618 | 6.05 | 157.6954 | 293.3444 | 225.2833 | 287.489 | 199.2465
2012-10-30 00:50 | 26.0201 | 314.8576 | 246.8293 | 309.0062 | 220.7498 | 87.7755 | 1.5509 | 39.3735 | 61.0625 | 25.3884 | 2.4612 | 39.2568 | 96.6174 | 15.9505 | 21.4521 | 6.0512 | 157.6799 | 293.4055 | 225.3772 | 287.5541 | 199.2977
2012-10-30 00:50 | 26.0276 | 314.8957 | 246.9151 | 309.0979 | 220.7809 | 87.7882 | 1.5508 | 39.3717 | 61.0559 | 25.3888 | 2.461 | 39.2552 | 96.6088 | 15.9506 | 21.4837 | 6.0521 | 157.6648 | 293.412 | 225.4314 | 287.6142 | 199.2972
2012-10-30 00:50 | 26.0276 | 314.9912 | 246.9386 | 309.1783 | 220.9463 | 87.7917 | 1.5507 | 39.3699 | 61.0498 | 25.389 | 2.4609 | 39.2532 | 96.5985 | 15.9507 | 21.5399 | 6.0524 | 157.6482 | 293.4513 | 225.3987 | 287.6384 | 199.4064
2012-10-30 00:50 | 26.0328 | 315.0843 | 247.0535 | 309.283 | 221.0302 | 87.8206 | 1.5491 | 39.3306 | 60.9287 | 25.3887 | 2.4584 | 39.2136 | 96.4037 | 15.9507 | 21.4688 | 6.0544 | 157.3324 | 293.6155 | 225.5847 | 287.8142 | 199.5614
2012-10-30 00:50 | 26.0328 | 315.2028 | 247.1239 | 309.3595 | 221.1404 | 87.8142 | 1.5468 | 39.2691 | 60.7405 | 25.3878 | 2.4546 | 39.1528 | 96.1059 | 15.9505 | 21.4809 | 6.0539 | 156.8464 | 293.7219 | 225.643 | 287.8786 | 199.6595
2012-10-30 00:50 | 26.0246 | 315.2481 | 247.194 | 309.4393 | 221.1637 | 87.8194 | 1.5464 | 39.261 | 60.7148 | 25.388 | 2.4541 | 39.1446 | 96.0651 | 15.9506 | 21.4687 | 6.0543 | 156.7799 | 293.7794 | 225.7253 | 287.9706 | 199.695
2012-10-30 00:50 | 26.0328 | 315.2716 | 247.2883 | 309.5279 | 221.1286 | 87.8437 | 1.5461 | 39.2524 | 60.6881 | 25.388 | 2.4536 | 39.1362 | 96.0234 | 15.9507 | 21.4809 | 6.0559 | 156.7115 | 293.7907 | 225.8074 | 288.047 | 199.6477
2012-10-30 00:50 | 26.0261 | 315.3489 | 247.3717 | 309.6092 | 220.803 | 87.824 | 1.546 | 39.2503 | 60.682 | 25.3879 | 2.4534 | 39.1339 | 96.0127 | 15.9506 | 21.4902 | 6.0546 | 156.6947 | 293.8587 | 225.8815 | 288.119 | 199.3128
2012-10-30 00:50 | 26.0223 | 315.4139 | 247.2739 | 309.6437 | 220.4095 | 87.8414 | 1.5461 | 39.2522 | 60.687 | 25.3883 | 2.4536 | 39.1362 | 96.0229 | 15.9508 | 21.4824 | 6.0558 | 156.7098 | 293.9315 | 225.7915 | 288.1613 | 198.9271
2012-10-30 00:50 | 26.0179 | 315.4592 | 247.2969 | 309.7273 | 221.3381 | 87.8599 | 1.5462 | 39.2546 | 60.6937 | 25.3885 | 2.4537 | 39.1381 | 96.0325 | 15.9508 | 21.478 | 6.0571 | 156.7262 | 293.9812 | 225.8189 | 288.2493 | 199.8601
2012-10-30 00:50 | 26.0223 | 315.5476 | 247.3835 | 309.8121 | 221.2795 | 87.846 | 1.5464 | 39.2603 | 60.7113 | 25.3886 | 2.454 | 39.144 | 96.0604 | 15.951 | 21.5025 | 6.0561 | 156.7717 | 294.0451 | 225.881 | 288.3096 | 199.777
2012-10-30 00:50 | 26.0231 | 315.6248 | 247.447 | 309.9124 | 221.3197 | 87.8622 | 1.5465 | 39.2651 | 60.7252 | 25.3889 | 2.4543 | 39.1487 | 96.0829 | 15.951 | 21.4872 | 6.0572 | 156.8081 | 294.1376 | 225.9598 | 288.4252 | 199.8325
2012-10-30 00:50 | 26.0171 | 315.68 | 247.5819 | 309.9906 | 221.7781 | 87.8749 | 1.5467 | 39.2698 | 60.7396 | 25.389 | 2.4546 | 39.1537 | 96.1064 | 15.9512 | 21.4652 | 6.0581 | 156.846 | 294.2148 | 226.1167 | 288.5254 | 200.3129
2012-10-30 00:50 | 26.0141 | 315.761 | 247.7041 | 310.0488 | 221.7476 | 87.8789 | 1.5469 | 39.2755 | 60.7561 | 25.3894 | 2.4549 | 39.159 | 96.1315 | 15.9514 | 21.4742 | 6.0584 | 156.8876 | 294.2868 | 226.2299 | 288.5746 | 200.2734
2012-10-30 00:50 | 26.0044 | 315.8392 | 247.7962 | 310.1156 | 221.5647 | 87.8882 | 1.5491 | 39.3312 | 60.9267 | 25.3902 | 2.4585 | 39.2166 | 96.4129 | 15.9516 | 21.5327 | 6.059 | 157.3396 | 294.3065 | 226.2635 | 288.5829 | 200.032
2012-10-30 00:50 | 26.0096 | 315.8788 | 247.8053 | 310.1897 | 221.8336 | 87.9136 | 1.5503 | 39.3624 | 61.0217 | 25.3909 | 2.4603 | 39.246 | 96.5561 | 15.9519 | 21.562 | 6.0608 | 157.5778 | 294.3168 | 226.2433 | 288.6277 | 200.2716
2012-10-30 00:51 | 25.9969 | 315.9463 | 247.8396 | 310.2191 | 221.2187 | 87.9026 | 1.5507 | 39.3738 | 61.0562 | 25.3912 | 2.461 | 39.257 | 96.6097 | 15.9519 | 21.5212 | 6.06 | 157.6659 | 294.4251 | 226.3184 | 288.6979 | 199.6975
2012-10-30 00:51 | 25.9872 | 316.055 | 247.869 | 310.2744 | 221.3034 | 87.9292 | 1.5511 | 39.3847 | 61.0894 | 25.3915 | 2.4616 | 39.2678 | 96.6614 | 15.9522 | 21.5316 | 6.0618 | 157.7508 | 294.5234 | 226.3374 | 288.7428 | 199.7718
2012-10-30 00:51 | 25.9887 | 316.1329 | 247.8822 | 310.3983 | 220.8562 | 87.9385 | 1.5514 | 39.3929 | 61.1148 | 25.3916 | 2.4621 | 39.2762 | 96.7025 | 15.9522 | 21.5532 | 6.0625 | 157.8173 | 294.5797 | 226.329 | 288.8451 | 199.303
2012-10-30 00:51 | 25.9745 | 316.1913 | 247.8289 | 310.4186 | 221.212 | 87.935 | 1.5516 | 39.3982 | 61.1303 | 25.3919 | 2.4624 | 39.2815 | 96.7283 | 15.9523 | 21.5028 | 6.0622 | 157.8585 | 294.6885 | 226.3261 | 288.9158 | 199.7092
2012-10-30 00:51 | 25.9789 | 316.2836 | 247.8372 | 310.5492 | 222.0706 | 87.9633 | 1.5517 | 39.401 | 61.1385 | 25.3921 | 2.4626 | 39.2842 | 96.7399 | 15.9525 | 21.5474 | 6.0642 | 157.8784 | 294.7362 | 226.2898 | 289.0018 | 200.5232
2012-10-30 00:51 | 25.9625 | 316.313 | 247.9969 | 310.552 | 222.0227 | 87.9691 | 1.5517 | 39.4024 | 61.1426 | 25.3922 | 2.4627 | 39.2856 | 96.7467 | 15.9525 | 21.5109 | 6.0646 | 157.8894 | 294.8021 | 226.486 | 289.0411 | 200.5118
2012-10-30 00:51 | 25.9685 | 316.3954 | 248.0812 | 310.6804 | 221.8712 | 87.9708 | 1.5517 | 39.4025 | 61.1426 | 25.3924 | 2.4626 | 39.2857 | 96.7467 | 15.9527 | 21.557 | 6.0647 | 157.8893 | 294.8384 | 226.5242 | 289.1234 | 200.3142
2012-10-30 00:51 | 25.9602 | 316.4826 | 248.1394 | 310.7601 | 221.9889 | 87.9934 | 1.5518 | 39.4034 | 61.1444 | 25.3928 | 2.4627 | 39.2868 | 96.7507 | 15.9529 | 21.5608 | 6.0663 | 157.8951 | 294.9218 | 226.5786 | 289.1993 | 200.4281
2012-10-30 00:51 | 25.9542 | 316.5148 | 248.1765 | 310.7962 | 222.1049 | 87.9986 | 1.5517 | 39.4026 | 61.1418 | 25.3929 | 2.4626 | 39.2857 | 96.745 | 15.9529 | 21.5026 | 6.0666 | 157.8868 | 295.0122 | 226.6739 | 289.2936 | 200.6023
2012-10-30 00:51 | 25.9565 | 316.5859 | 248.2766 | 310.8597 | 222.0796 | 87.9986 | 1.5517 | 39.4031 | 61.143 | 25.393 | 2.4627 | 39.2863 | 96.7485 | 15.9529 | 21.545 | 6.0666 | 157.8915 | 295.0409 | 226.7316 | 289.3147 | 200.5346
2012-10-30 00:51 | 25.9557 | 316.6538 | 248.3541 | 310.8971 | 222.2481 | 88.0211 | 1.5517 | 39.4024 | 61.1402 | 25.3933 | 2.4626 | 39.2858 | 96.7444 | 15.9531 | 21.5442 | 6.0682 | 157.8846 | 295.1096 | 226.8099 | 289.3529 | 200.7039
2012-10-30 00:51 | 25.943 | 316.7405 | 248.3727 | 310.9915 | 222.2629 | 88.0304 | 1.5516 | 39.4009 | 61.1348 | 25.3935 | 2.4625 | 39.2842 | 96.736 | 15.9532 | 21.5075 | 6.0688 | 157.8708 | 295.233 | 226.8652 | 289.484 | 200.7554
2012-10-30 00:51 | 25.9587 | 316.7867 | 248.4627 | 311.0417 | 222.1645 | 88.0257 | 1.5516 | 39.3997 | 61.1309 | 25.3936 | 2.4624 | 39.283 | 96.7298 | 15.9532 | 21.5633 | 6.0685 | 157.8607 | 295.2234 | 226.8994 | 289.4784 | 200.6012
2012-10-30 00:51 | 25.9505 | 316.8473 | 248.5093 | 311.1253 | 222.2153 | 88.046 | 1.5515 | 39.3986 | 61.1274 | 25.3937 | 2.4623 | 39.2819 | 96.7244 | 15.9532 | 21.5471 | 6.0699 | 157.8518 | 295.3002 | 226.9622 | 289.5782 | 200.6682
2012-10-30 00:51 | 25.9557 | 316.8983 | 248.5889 | 311.1955 | 222.2166 | 88.0442 | 1.5515 | 39.3977 | 61.1239 | 25.3939 | 2.4622 | 39.2809 | 96.7185 | 15.9534 | 21.5242 | 6.0698 | 157.8424 | 295.3741 | 227.0647 | 289.6713 | 200.6924
2012-10-30 00:51 | 25.9528 | 316.9527 | 248.5899 | 311.2615 | 222.379 | 88.0379 | 1.5514 | 39.3964 | 61.1203 | 25.3938 | 2.4621 | 39.2796 | 96.712 | 15.9534 | 21.4651 | 6.0693 | 157.8323 | 295.4876 | 227.1248 | 289.7964 | 200.9139
2012-10-30 00:51 | 25.9453 | 317.0369 | 248.6802 | 311.3458 | 222.4542 | 88.0506 | 1.5514 | 39.3954 | 61.116 | 25.3943 | 2.4621 | 39.2786 | 96.7066 | 15.9535 | 21.5098 | 6.0702 | 157.8227 | 295.5271 | 227.1704 | 289.836 | 200.9444
2012-10-30 00:51 | 25.955 | 317.1 | 248.7603 | 311.4167 | 222.456 | 88.0535 | 1.5513 | 39.3943 | 61.1129 | 25.3942 | 2.462 | 39.2776 | 96.7013 | 15.9536 | 21.4954 | 6.0704 | 157.8142 | 295.6046 | 227.2649 | 289.9213 | 200.9606
2012-10-30 00:51 | 25.946 | 317.156 | 248.7904 | 311.4612 | 222.5336 | 88.0749 | 1.5512 | 39.393 | 61.1078 | 25.3946 | 2.4619 | 39.2762 | 96.6949 | 15.9535 | 21.4824 | 6.0719 | 157.8026 | 295.6736 | 227.308 | 289.9788 | 201.0512
2012-10-30 00:51 | 25.9498 | 317.1827 | 248.8842 | 311.5338 | 222.6397 | 88.0864 | 1.5512 | 39.3924 | 61.1068 | 25.3942 | 2.4619 | 39.2758 | 96.6928 | 15.9535 | 21.5062 | 6.0727 | 157.7996 | 295.6765 | 227.378 | 290.0276 | 201.1335
2012-10-30 00:51 | 25.952 | 317.2383 | 248.9451 | 311.5972 | 222.7954 | 88.0991 | 1.5512 | 39.3927 | 61.1072 | 25.3945 | 2.4619 | 39.2758 | 96.6924 | 15.9536 | 21.4563 | 6.0736 | 157.7996 | 295.782 | 227.4888 | 290.1409 | 201.3391
2012-10-30 00:51 | 25.9513 | 317.3522 | 249.0109 | 311.6807 | 222.8222 | 88.0991 | 1.5512 | 39.3907 | 61.1012 | 25.3943 | 2.4618 | 39.274 | 96.6829 | 15.9537 | 21.4676 | 6.0736 | 157.7842 | 295.8846 | 227.5433 | 290.2131 | 201.3546
2012-10-30 00:51 | 25.961 | 317.4383 | 249.091 | 311.7784 | 222.95 | 88.0957 | 1.5511 | 39.3897 | 61.0969 | 25.3948 | 2.4617 | 39.2728 | 96.6768 | 15.9537 | 21.4332 | 6.0733 | 157.7737 | 296.0051 | 227.6578 | 290.3452 | 201.5168
2012-10-30 00:51 | 25.9535 | 317.4919 | 249.1109 | 311.832 | 222.8835 | 88.102 | 1.5511 | 39.3891 | 61.0946 | 25.3951 | 2.4616 | 39.2724 | 96.674 | 15.9538 | 21.4417 | 6.0738 | 157.7687 | 296.0502 | 227.6692 | 290.3903 | 201.4418
2012-10-30 00:51 | 25.9468 | 317.5387 | 249.159 | 311.8636 | 223.1798 | 88.1159 | 1.551 | 39.3871 | 61.0882 | 25.3952 | 2.4615 | 39.2704 | 96.6641 | 15.9539 | 21.4511 | 6.0747 | 157.7524 | 296.0876 | 227.7079 | 290.4125 | 201.7287
2012-10-30 00:52 | 25.9565 | 317.6056 | 249.2978 | 311.9154 | 223.0006 | 88.1275 | 1.5509 | 39.3853 | 61.0832 | 25.3948 | 2.4614 | 39.2686 | 96.6545 | 15.954 | 21.4728 | 6.0755 | 157.7377 | 296.1328 | 227.825 | 290.4426 | 201.5278
2012-10-30 00:52 | 25.952 | 317.6202 | 249.3402 | 312.0103 | 222.8229 | 88.1419 | 1.5509 | 39.3844 | 61.0796 | 25.3952 | 2.4613 | 39.2676 | 96.6496 | 15.954 | 21.5165 | 6.0765 | 157.7292 | 296.1037 | 227.8237 | 290.4938 | 201.3064
2012-10-30 00:52 | 25.9542 | 317.695 | 249.3072 | 312.0507 | 222.9314 | 88.1419 | 1.5508 | 39.3832 | 61.0762 | 25.3951 | 2.4612 | 39.2664 | 96.6436 | 15.954 | 21.4545 | 6.0765 | 157.7199 | 296.2405 | 227.8527 | 290.5962 | 201.4769
2012-10-30 00:52 | 25.9408 | 317.7351 | 249.3447 | 312.0373 | 223.1699 | 88.1361 | 1.5508 | 39.3832 | 61.0753 | 25.3955 | 2.4612 | 39.2665 | 96.6439 | 15.954 | 21.4852 | 6.0761 | 157.7193 | 296.2499 | 227.8595 | 290.5521 | 201.6847
2012-10-30 00:52 | 25.9535 | 317.8089 | 249.42 | 312.1801 | 223.218 | 88.1668 | 1.5507 | 39.3819 | 61.0714 | 25.3954 | 2.4611 | 39.2652 | 96.637 | 15.9541 | 21.4818 | 6.0782 | 157.7084 | 296.3271 | 227.9382 | 290.6983 | 201.7362
2012-10-30 00:52 | 25.9505 | 317.8517 | 249.4874 | 312.2076 | 223.1324 | 88.1696 | 1.5507 | 39.3804 | 61.0667 | 25.3955 | 2.461 | 39.2639 | 96.6297 | 15.9542 | 21.4949 | 6.0784 | 157.6964 | 296.3568 | 227.9925 | 290.7127 | 201.6375
2012-10-30 00:52 | 25.9468 | 317.89 | 249.5463 | 312.2651 | 223.2034 | 88.1691 | 1.5505 | 39.3773 | 61.056 | 25.396 | 2.4608 | 39.2606 | 96.6124 | 15.9544 | 21.4751 | 6.0784 | 157.6684 | 296.4149 | 228.0712 | 290.79 | 201.7283
2012-10-30 00:52 | 25.9587 | 318.0012 | 249.5856 | 312.3229 | 223.3058 | 88.1945 | 1.5505 | 39.3765 | 61.0532 | 25.396 | 2.4608 | 39.2599 | 96.6099 | 15.9543 | 21.4991 | 6.0801 | 157.6631 | 296.5021 | 228.0865 | 290.8238 | 201.8067
2012-10-30 00:52 | 25.9572 | 317.992 | 249.6154 | 312.3711 | 223.2178 | 88.1876 | 1.5505 | 39.376 | 61.0514 | 25.3961 | 2.4607 | 39.2592 | 96.6059 | 15.9544 | 21.4775 | 6.0797 | 157.6573 | 296.5145 | 228.1379 | 290.8936 | 201.7403
2012-10-30 00:52 | 25.9528 | 318.0525 | 249.6344 | 312.4317 | 223.1779 | 88.1893 | 1.5504 | 39.375 | 61.0479 | 25.3963 | 2.4607 | 39.2582 | 96.6007 | 15.9544 | 21.4852 | 6.0798 | 157.6486 | 296.5673 | 228.1492 | 290.9465 | 201.6927
2012-10-30 00:52 | 25.9513 | 318.074 | 249.6955 | 312.4417 | 223.0505 | 88.2055 | 1.5504 | 39.3747 | 61.0471 | 25.3963 | 2.4606 | 39.2582 | 96.6003 | 15.9544 | 21.5117 | 6.0809 | 157.6474 | 296.5623 | 228.1838 | 290.93 | 201.5388
2012-10-30 00:52 | 25.961 | 318.1371 | 249.6935 | 312.5699 | 223.1035 | 88.1991 | 1.5504 | 39.374 | 61.0447 | 25.3963 | 2.4606 | 39.2572 | 96.5956 | 15.9545 | 21.4653 | 6.0804 | 157.6403 | 296.6718 | 228.2282 | 291.1046 | 201.6382
2012-10-30 00:52 | 25.9632 | 318.208 | 249.7426 | 312.5683 | 223.2041 | 88.2003 | 1.5503 | 39.3731 | 61.042 | 25.3964 | 2.4605 | 39.2564 | 96.5919 | 15.9544 | 21.4915 | 6.0805 | 157.6339 | 296.7165 | 228.2511 | 291.0768 | 201.7126
2012-10-30 00:52 | 25.961 | 318.2249 | 249.7639 | 312.677 | 223.2727 | 88.2141 | 1.5468 | 39.2822 | 60.7628 | 25.3953 | 2.4548 | 39.1647 | 96.1433 | 15.9541 | 21.4773 | 6.0815 | 156.9061 | 296.7476 | 228.2866 | 291.1997 | 201.7954
2012-10-30 00:52 | 25.9505 | 318.287 | 249.7886 | 312.7086 | 223.1678 | 88.2107 | 1.5464 | 39.2719 | 60.7315 | 25.3951 | 2.4543 | 39.1554 | 96.0973 | 15.9541 | 21.4668 | 6.0812 | 156.8288 | 296.8202 | 228.3218 | 291.2418 | 201.701
2012-10-30 00:52 | 25.9565 | 318.335 | 249.7555 | 312.7146 | 223.2722 | 88.1985 | 1.546 | 39.2606 | 60.6971 | 25.3949 | 2.4536 | 39.1444 | 96.043 | 15.9541 | 21.4889 | 6.0804 | 156.7401 | 296.8461 | 228.2666 | 291.2257 | 201.7833
2012-10-30 00:52 | 25.958 | 318.3327 | 249.8039 | 312.8231 | 223.1792 | 88.2251 | 1.5458 | 39.2556 | 60.681 | 25.3952 | 2.4532 | 39.1394 | 96.0183 | 15.9542 | 21.5265 | 6.0822 | 156.6994 | 296.8062 | 228.2774 | 291.2966 | 201.6527
2012-10-30 00:52 | 25.9572 | 318.3739 | 249.8148 | 312.8338 | 223.4106 | 88.2286 | 1.5458 | 39.2548 | 60.6785 | 25.3951 | 2.4532 | 39.1384 | 96.0134 | 15.9542 | 21.4735 | 6.0825 | 156.6919 | 296.9004 | 228.3413 | 291.3603 | 201.9371
2012-10-30 00:52 | 25.9528 | 318.3885 | 249.8613 | 312.8753 | 223.8037 | 88.2286 | 1.5458 | 39.2568 | 60.6845 | 25.3952 | 2.4533 | 39.1406 | 96.0238 | 15.9542 | 21.4892 | 6.0825 | 156.7084 | 296.8993 | 228.3721 | 291.3861 | 202.3145
2012-10-30 00:52 | 25.9595 | 318.4678 | 249.9697 | 312.9699 | 224.015 | 88.2321 | 1.546 | 39.262 | 60.7005 | 25.3953 | 2.4536 | 39.1456 | 96.0473 | 15.9544 | 21.5079 | 6.0827 | 156.7478 | 296.9599 | 228.4618 | 291.462 | 202.5071
2012-10-30 00:52 | 25.9513 | 318.494 | 250.024 | 312.9502 | 224.298 | 88.2349 | 1.5462 | 39.2666 | 60.7137 | 25.3957 | 2.4539 | 39.1502 | 96.0697 | 15.9544 | 21.4877 | 6.0829 | 156.7833 | 297.0063 | 228.5363 | 291.4625 | 202.8103
2012-10-30 00:52 | 25.9557 | 318.5442 | 250.1732 | 313.012 | 224.1489 | 88.2269 | 1.5463 | 39.2697 | 60.7228 | 25.3959 | 2.454 | 39.1532 | 96.0836 | 15.9545 | 21.4921 | 6.0824 | 156.8064 | 297.0521 | 228.6811 | 291.5199 | 202.6568
2012-10-30 00:52 | 25.9513 | 318.6199 | 250.157 | 313.0497 | 223.2512 | 88.2488 | 1.5464 | 39.273 | 60.7333 | 25.3958 | 2.4543 | 39.1566 | 96.1005 | 15.9546 | 21.5238 | 6.0839 | 156.8338 | 297.0961 | 228.6332 | 291.5259 | 201.7274
2012-10-30 00:52 | 25.9438 | 318.6201 | 250.1652 | 313.0574 | 224.1921 | 88.261 | 1.5469 | 39.2843 | 60.7677 | 25.396 | 2.4549 | 39.1662 | 96.1474 | 15.9546 | 21.4681 | 6.0847 | 156.9151 | 297.152 | 228.6971 | 291.5893 | 202.724
2012-10-30 00:52 | 25.9348 | 318.6645 | 250.2539 | 313.0867 | 224.49 | 88.2563 | 1.5497 | 39.3587 | 60.9949 | 25.3973 | 2.4595 | 39.2417 | 96.5151 | 15.9551 | 21.5073 | 6.0844 | 157.51 | 297.1572 | 228.7466 | 291.5794 | 202.9827
2012-10-30 00:52 | 25.94 | 318.7155 | 250.3178 | 313.1607 | 223.6098 | 88.269 | 1.5502 | 39.3718 | 61.0352 | 25.3975 | 2.4603 | 39.2548 | 96.5794 | 15.9551 | 21.4924 | 6.0853 | 157.6146 | 297.2231 | 228.8254 | 291.6683 | 202.1174
2012-10-30 00:52 | 25.9415 | 318.759 | 250.2802 | 313.1928 | 223.8986 | 88.2777 | 1.5507 | 39.3834 | 61.0704 | 25.3978 | 2.4611 | 39.2665 | 96.6372 | 15.9551 | 21.506 | 6.0859 | 157.7077 | 297.253 | 228.7742 | 291.6868 | 202.3926
2012-10-30 00:53 | 25.9266 | 318.7938 | 250.2966 | 313.2123 | 223.7538 | 88.2922 | 1.5509 | 39.3915 | 61.094 | 25.3984 | 2.4615 | 39.2744 | 96.6754 | 15.9552 | 21.487 | 6.0869 | 157.7694 | 297.3068 | 228.8096 | 291.7253 | 202.2668
2012-10-30 00:53 | 25.9378 | 318.8126 | 250.2999 | 313.2082 | 223.706 | 88.2922 | 1.5512 | 39.3969 | 61.1108 | 25.3984 | 2.4619 | 39.28 | 96.7014 | 15.9555 | 21.5063 | 6.0869 | 157.8122 | 297.3063 | 228.7936 | 291.7019 | 202.1997
2012-10-30 00:53 | 25.9385 | 318.8362 | 250.2967 | 313.3465 | 223.8563 | 88.3107 | 1.5513 | 39.4002 | 61.1206 | 25.3986 | 2.462 | 39.2834 | 96.7165 | 15.9558 | 21.527 | 6.0881 | 157.837 | 297.3092 | 228.7697 | 291.8195 | 202.3293
2012-10-30 00:53 | 25.9281 | 318.9403 | 250.302 | 313.3476 | 223.8104 | 88.3083 | 1.5514 | 39.4026 | 61.1274 | 25.3989 | 2.4622 | 39.2857 | 96.7276 | 15.9558 | 21.5046 | 6.088 | 157.855 | 297.4357 | 228.7974 | 291.843 | 202.3058
2012-10-30 00:53 | 25.937 | 318.9569 | 250.35 | 313.3985 | 223.7249 | 88.3055 | 1.5514 | 39.4032 | 61.1289 | 25.399 | 2.4622 | 39.2865 | 96.731 | 15.9559 | 21.4774 | 6.0878 | 157.8598 | 297.4795 | 228.8726 | 291.9211 | 202.2475
2012-10-30 00:53 | 25.9296 | 319.0029 | 250.3973 | 313.437 | 224.0166 | 88.3239 | 1.5514 | 39.4034 | 61.1284 | 25.3994 | 2.4622 | 39.2866 | 96.7302 | 15.9561 | 21.49 | 6.089 | 157.8586 | 297.5129 | 228.9073 | 291.947 | 202.5266
2012-10-30 00:53 | 25.9333 | 319.0372 | 250.4988 | 313.4789 | 224.2564 | 88.3095 | 1.5513 | 39.4032 | 61.1277 | 25.3994 | 2.4621 | 39.2864 | 96.728 | 15.9563 | 21.4817 | 6.0881 | 157.8557 | 297.5555 | 229.0171 | 291.9972 | 202.7747
2012-10-30 00:53 | 25.9318 | 319.1044 | 250.5794 | 313.5692 | 224.3926 | 88.3211 | 1.5513 | 39.4029 | 61.127 | 25.3994 | 2.4621 | 39.2861 | 96.7264 | 15.9563 | 21.4842 | 6.0889 | 157.8534 | 297.6202 | 229.0952 | 292.085 | 202.9084
2012-10-30 00:53 | 25.9288 | 319.1128 | 250.6624 | 313.6006 | 224.4132 | 88.3274 | 1.5512 | 39.4004 | 61.1182 | 25.3998 | 2.462 | 39.2838 | 96.7151 | 15.9563 | 21.4892 | 6.0893 | 157.8333 | 297.6236 | 229.1732 | 292.1114 | 202.924
2012-10-30 00:53 | 25.937 | 319.1936 | 250.6354 | 313.6088 | 224.3545 | 88.3309 | 1.5512 | 39.3999 | 61.1167 | 25.3999 | 2.4619 | 39.2829 | 96.7105 | 15.9563 | 21.4974 | 6.0895 | 157.8272 | 297.6962 | 229.138 | 292.1114 | 202.8571
2012-10-30 00:53 | 25.9333 | 319.228 | 250.7334 | 313.6433 | 224.4295 | 88.3372 | 1.5512 | 39.3998 | 61.1156 | 25.4001 | 2.4619 | 39.283 | 96.7108 | 15.9563 | 21.5138 | 6.09 | 157.8264 | 297.7142 | 229.2196 | 292.1295 | 202.9157
2012-10-30 00:53 | 25.9385 | 319.2256 | 250.7425 | 313.7441 | 224.2222 | 88.3615 | 1.5511 | 39.3993 | 61.1141 | 25.4001 | 2.4619 | 39.2825 | 96.7082 | 15.9564 | 21.511 | 6.0916 | 157.8222 | 297.7146 | 229.2315 | 292.2331 | 202.7112
2012-10-30 00:53 | 25.943 | 319.295 | 250.8057 | 313.7448 | 224.3094 | 88.3419 | 1.5511 | 39.3978 | 61.1094 | 25.4002 | 2.4617 | 39.2811 | 96.6999 | 15.9566 | 21.4673 | 6.0903 | 157.8093 | 297.8277 | 229.3384 | 292.2775 | 202.8421
2012-10-30 00:53 | 25.943 | 319.3255 | 250.8096 | 313.7906 | 224.3487 | 88.3476 | 1.5511 | 39.3975 | 61.1087 | 25.4001 | 2.4617 | 39.2808 | 96.6991 | 15.9565 | 21.4392 | 6.0907 | 157.8078 | 297.8863 | 229.3704 | 292.3514 | 202.9095
2012-10-30 00:53 | 25.9528 | 319.3735 | 250.8389 | 313.8616 | 224.4136 | 88.3644 | 1.551 | 39.3969 | 61.1059 | 25.4005 | 2.4617 | 39.28 | 96.6948 | 15.9566 | 21.4731 | 6.0918 | 157.8007 | 297.9004 | 229.3658 | 292.3885 | 202.9405
2012-10-30 00:53 | 25.946 | 319.382 | 250.8791 | 313.893 | 224.4776 | 88.3748 | 1.5509 | 39.3947 | 61.0984 | 25.4006 | 2.4615 | 39.2779 | 96.6836 | 15.9567 | 21.4463 | 6.0926 | 157.7821 | 297.9357 | 229.4328 | 292.4467 | 203.0313
2012-10-30 00:53 | 25.9468 | 319.4629 | 250.8994 | 313.9167 | 224.4469 | 88.38 | 1.5509 | 39.3944 | 61.0972 | 25.4008 | 2.4615 | 39.2776 | 96.682 | 15.9568 | 21.4511 | 6.0929 | 157.7792 | 298.0118 | 229.4483 | 292.4656 | 202.9958
2012-10-30 00:53 | 25.9438 | 319.4752 | 250.9199 | 313.9443 | 224.4912 | 88.3725 | 1.5509 | 39.3942 | 61.0967 | 25.4008 | 2.4615 | 39.2775 | 96.6815 | 15.9567 | 21.4922 | 6.0924 | 157.7783 | 297.983 | 229.4277 | 292.4521 | 202.999
2012-10-30 00:53 | 25.9445 | 319.4988 | 250.9557 | 313.9947 | 224.5548 | 88.3933 | 1.5509 | 39.393 | 61.0927 | 25.4009 | 2.4614 | 39.2762 | 96.6751 | 15.9568 | 21.5009 | 6.0938 | 157.7678 | 297.9979 | 229.4548 | 292.4938 | 203.0539
2012-10-30 00:53 | 25.9535 | 319.5574 | 251.0312 | 314.0305 | 224.5677 | 88.3823 | 1.5509 | 39.3928 | 61.0924 | 25.4008 | 2.4614 | 39.2758 | 96.6727 | 15.9568 | 21.5862 | 6.0931 | 157.7651 | 297.9712 | 229.445 | 292.4443 | 202.9815
2012-10-30 00:53 | 25.9483 | 319.5904 | 251.0573 | 314.075 | 224.6176 | 88.3858 | 1.5508 | 39.3915 | 61.0873 | 25.4012 | 2.4613 | 39.2748 | 96.6674 | 15.9569 | 21.5368 | 6.0933 | 157.7547 | 298.0536 | 229.5205 | 292.5382 | 203.0808
2012-10-30 00:53 | 25.9475 | 319.6125 | 251.1073 | 314.078 | 224.5342 | 88.4118 | 1.5507 | 39.3905 | 61.0839 | 25.4013 | 2.4612 | 39.2736 | 96.6621 | 15.9568 | 21.54 | 6.0951 | 157.746 | 298.0725 | 229.5673 | 292.538 | 202.9942
2012-10-30 00:53 | 25.9587 | 319.6657 | 251.1224 | 314.1656 | 224.5493 | 88.3933 | 1.5507 | 39.3894 | 61.0815 | 25.4009 | 2.4612 | 39.2727 | 96.6571 | 15.9569 | 21.5392 | 6.0938 | 157.7386 | 298.1265 | 229.5832 | 292.6264 | 203.0101
2012-10-30 00:54 | 25.9513 | 319.6736 | 251.1424 | 314.1659 | 224.5774 | 88.402 | 1.5506 | 39.3873 | 61.0743 | 25.4012 | 2.461 | 39.2706 | 96.6463 | 15.957 | 21.4837 | 6.0944 | 157.7206 | 298.1899 | 229.6587 | 292.6822 | 203.0937
2012-10-30 00:54 | 25.9528 | 319.7018 | 251.1869 | 314.2209 | 224.6812 | 88.4141 | 1.5506 | 39.3858 | 61.07 | 25.4011 | 2.461 | 39.2692 | 96.6396 | 15.9569 | 21.4852 | 6.0953 | 157.7096 | 298.2166 | 229.7017 | 292.7357 | 203.196
2012-10-30 00:54 | 25.9587 | 319.7535 | 251.2045 | 314.223 | 224.7697 | 88.4234 | 1.5505 | 39.3853 | 61.0682 | 25.4012 | 2.4609 | 39.2685 | 96.6354 | 15.957 | 21.471 | 6.0959 | 157.7036 | 298.2825 | 229.7335 | 292.752 | 203.2987
2012-10-30 00:54 | 25.9655 | 319.787 | 251.27 | 314.2909 | 224.8473 | 88.421 | 1.5505 | 39.3843 | 61.0636 | 25.4017 | 2.4608 | 39.2673 | 96.6289 | 15.9572 | 21.4858 | 6.0957 | 157.6925 | 298.3012 | 229.7842 | 292.8051 | 203.3615
2012-10-30 00:54 | 25.9617 | 319.8634 | 251.2623 | 314.3521 | 224.7648 | 88.4407 | 1.5504 | 39.3824 | 61.0586 | 25.4014 | 2.4607 | 39.2656 | 96.621 | 15.9571 | 21.4499 | 6.0971 | 157.6796 | 298.4135 | 229.8124 | 292.9022 | 203.3149
2012-10-30 00:54 | 25.9625 | 319.8413 | 251.31 | 314.3682 | 224.7971 | 88.4447 | 1.5503 | 39.3807 | 61.0535 | 25.4013 | 2.4606 | 39.2639 | 96.6128 | 15.957 | 21.4507 | 6.0974 | 157.6663 | 298.3906 | 229.8593 | 292.9175 | 203.3464
2012-10-30 00:54 | 25.9617 | 319.9092 | 251.3053 | 314.3941 | 224.8356 | 88.4418 | 1.5503 | 39.3799 | 61.0513 | 25.4012 | 2.4606 | 39.2631 | 96.6093 | 15.957 | 21.4539 | 6.0972 | 157.6606 | 298.4553 | 229.8514 | 292.9402 | 203.3817
2012-10-30 00:54 | 25.9692 | 319.9319 | 251.3597 | 314.4169 | 224.8235 | 88.4424 | 1.5502 | 39.3777 | 61.0431 | 25.4018 | 2.4604 | 39.261 | 96.5984 | 15.9571 | 21.4654 | 6.0972 | 157.6415 | 298.4665 | 229.8943 | 292.9515 | 203.3581
2012-10-30 00:54 | 25.97 | 319.9938 | 251.38 | 314.4903 | 224.8912 | 88.4551 | 1.5502 | 39.3768 | 61.0413 | 25.4014 | 2.4603 | 39.2602 | 96.5938 | 15.9572 | 21.4703 | 6.0981 | 157.635 | 298.5235 | 229.9097 | 293.02 | 203.4209
2012-10-30 00:54 | 25.9715 | 320.0029 | 251.4245 | 314.5071 | 225.0933 | 88.4459 | 1.5501 | 39.3754 | 61.0355 | 25.402 | 2.4603 | 39.2588 | 96.5869 | 15.9572 | 21.4396 | 6.0975 | 157.6224 | 298.5633 | 229.9849 | 293.0675 | 203.6537
2012-10-30 00:54 | 25.9745 | 320.0136 | 251.4549 | 314.5407 | 225.1081 | 88.4655 | 1.5501 | 39.3744 | 61.0328 | 25.4018 | 2.4602 | 39.2578 | 96.5819 | 15.9572 | 21.4908 | 6.0988 | 157.6148 | 298.5228 | 229.9641 | 293.0499 | 203.6173
2012-10-30 00:54 | 25.9789 | 320.0638 | 251.5101 | 314.5795 | 225.0732 | 88.4736 | 1.55 | 39.3739 | 61.031 | 25.4019 | 2.4602 | 39.2572 | 96.5791 | 15.9571 | 21.4671 | 6.0994 | 157.6101 | 298.5967 | 230.043 | 293.1124 | 203.6061
2012-10-30 00:54 | 25.9774 | 320.0737 | 251.5438 | 314.6315 | 224.9222 | 88.458 | 1.5473 | 39.304 | 60.8162 | 25.4011 | 2.4559 | 39.1888 | 96.2444 | 15.9569 | 21.4616 | 6.0983 | 157.0606 | 298.6121 | 230.0822 | 293.1699 | 203.4606
2012-10-30 00:54 | 25.9737 | 320.1158 | 251.5127 | 314.6507 | 224.8083 | 88.4511 | 1.546 | 39.2703 | 60.7137 | 25.4005 | 2.4538 | 39.1542 | 96.0753 | 15.9567 | 21.4499 | 6.0978 | 156.789 | 298.6659 | 230.0628 | 293.2008 | 203.3584
2012-10-30 00:54 | 25.97 | 320.1007 | 251.4973 | 314.6814 | 224.718 | 88.465 | 1.5457 | 39.2597 | 60.6819 | 25.4001 | 2.4531 | 39.1434 | 96.0231 | 15.9567 | 21.4542 | 6.0988 | 156.705 | 298.6465 | 230.0431 | 293.2272 | 203.2638
2012-10-30 00:54 | 25.9715 | 320.1327 | 251.5379 | 314.6829 | 224.1608 | 88.4551 | 1.5455 | 39.2565 | 60.6714 | 25.4002 | 2.4529 | 39.1401 | 96.0066 | 15.9567 | 21.5039 | 6.0981 | 156.678 | 298.6288 | 230.034 | 293.179 | 202.6569
2012-10-30 00:54 | 25.9804 | 320.1531 | 251.4334 | 314.7759 | 224.5356 | 88.4517 | 1.5455 | 39.2562 | 60.6706 | 25.4003 | 2.4529 | 39.1398 | 96.0049 | 15.9568 | 21.4646 | 6.0979 | 156.6754 | 298.6885 | 229.9688 | 293.3113 | 203.071
2012-10-30 00:54 | 25.9685 | 320.2099 | 251.4684 | 314.7869 | 225.3697 | 88.4627 | 1.5455 | 39.2568 | 60.6725 | 25.4003 | 2.4529 | 39.1405 | 96.0079 | 15.9568 | 21.5009 | 6.0986 | 156.6804 | 298.709 | 229.9675 | 293.286 | 203.8688
2012-10-30 00:54 | 25.967 | 320.1931 | 251.4982 | 314.816 | 224.8134 | 88.4702 | 1.5457 | 39.2608 | 60.6839 | 25.4006 | 2.4532 | 39.1444 | 96.0277 | 15.9567 | 21.4713 | 6.0991 | 156.7116 | 298.7218 | 230.0269 | 293.3447 | 203.3421
2012-10-30 00:54 | 25.964 | 320.2855 | 251.5578 | 314.8321 | 225.1016 | 88.4598 | 1.5458 | 39.264 | 60.6941 | 25.4005 | 2.4534 | 39.1475 | 96.0426 | 15.9567 | 21.4803 | 6.0984 | 156.7367 | 298.8052 | 230.0775 | 293.3518 | 203.6213
2012-10-30 00:54 | 25.964 | 320.2741 | 251.5187 | 314.855 | 225.1212 | 88.4679 | 1.5461 | 39.2712 | 60.7157 | 25.4007 | 2.4538 | 39.1545 | 96.0772 | 15.9567 | 21.4482 | 6.099 | 156.7929 | 298.8259 | 230.0705 | 293.4068 | 203.673
2012-10-30 00:54 | 25.9602 | 320.3122 | 251.5618 | 314.8589 | 224.8932 | 88.4742 | 1.5462 | 39.2753 | 60.7286 | 25.4008 | 2.4541 | 39.1588 | 96.0978 | 15.9568 | 21.4805 | 6.0994 | 156.8264 | 298.8317 | 230.0813 | 293.3784 | 203.4127
2012-10-30 00:54 | 25.961 | 320.355 | 251.5782 | 314.9552 | 225.5195 | 88.4858 | 1.5491 | 39.35 | 60.9576 | 25.4017 | 2.4586 | 39.2322 | 96.4559 | 15.9572 | 21.4653 | 6.1002 | 157.4135 | 298.8897 | 230.1129 | 293.4899 | 204.0542
2012-10-30 00:54 | 25.9565 | 320.4192 | 251.6206 | 314.9354 | 224.8737 | 88.4835 | 1.5498 | 39.3688 | 61.0144 | 25.4022 | 2.4598 | 39.2519 | 96.5522 | 15.9573 | 21.4768 | 6.1001 | 157.5666 | 298.9424 | 230.1438 | 293.4586 | 203.3969
2012-10-30 00:54 | 25.9557 | 320.3993 | 251.5377 | 314.9575 | 224.3732 | 88.4881 | 1.5503 | 39.3806 | 61.0503 | 25.4025 | 2.4605 | 39.2636 | 96.6091 | 15.9574 | 21.476 | 6.1004 | 157.6595 | 298.9233 | 230.0617 | 293.4815 | 202.8972
2012-10-30 00:54 | 25.9602 | 320.4458 | 251.511 | 314.9964 | 224.3502 | 88.5037 | 1.5506 | 39.3895 | 61.0773 | 25.4028 | 2.4611 | 39.273 | 96.6548 | 15.9575 | 21.5046 | 6.1014 | 157.7321 | 298.9412 | 230.0064 | 293.4918 | 202.8456
2012-10-30 00:55 | 25.9505 | 320.4552 | 251.4856 | 314.9906 | 224.6749 | 88.5002 | 1.5509 | 39.3969 | 61.1 | 25.4029 | 2.4615 | 39.2801 | 96.6895 | 15.9576 | 21.4949 | 6.1012 | 157.7896 | 298.9603 | 229.9907 | 293.4957 | 203.18
2012-10-30 00:55 | 25.9483 | 320.4835 | 251.4912 | 315.0457 | 225.0819 | 88.5031 | 1.5509 | 39.3976 | 61.1014 | 25.4032 | 2.4616 | 39.2809 | 96.6925 | 15.9577 | 21.4606 | 6.1014 | 157.794 | 299.0229 | 230.0306 | 293.5851 | 203.6213
2012-10-30 00:55 | 25.955 | 320.5093 | 251.5409 | 315.0256 | 225.1988 | 88.5077 | 1.5509 | 39.3981 | 61.1029 | 25.4032 | 2.4615 | 39.2811 | 96.6919 | 15.958 | 21.4954 | 6.1017 | 157.7948 | 299.0139 | 230.0455 | 293.5302 | 203.7034
2012-10-30 00:55 | 25.9595 | 320.5062 | 251.6197 | 315.076 | 225.3725 | 88.5285 | 1.5509 | 39.3986 | 61.104 | 25.4034 | 2.4616 | 39.2819 | 96.6956 | 15.958 | 21.5119 | 6.1032 | 157.7996 | 298.9943 | 230.1078 | 293.5641 | 203.8606
2012-10-30 00:55 | 25.9438 | 320.5134 | 251.6665 | 315.0717 | 225.3174 | 88.5216 | 1.5509 | 39.3993 | 61.1063 | 25.4034 | 2.4616 | 39.2824 | 96.6988 | 15.9578 | 21.4842 | 6.1027 | 157.8051 | 299.0292 | 230.1823 | 293.5875 | 203.8332
2012-10-30 00:55 | 25.9483 | 320.5637 | 251.675 | 315.0839 | 225.196 | 88.5314 | 1.551 | 39.4001 | 61.1088 | 25.4034 | 2.4617 | 39.2833 | 96.7029 | 15.9579 | 21.4726 | 6.1034 | 157.8117 | 299.0911 | 230.2024 | 293.6113 | 203.7234
2012-10-30 00:55 | 25.9453 | 320.5683 | 251.6993 | 315.1229 | 225.2914 | 88.532 | 1.551 | 39.4011 | 61.1119 | 25.4033 | 2.4617 | 39.2843 | 96.707 | 15.9581 | 21.4496 | 6.1034 | 157.819 | 299.1187 | 230.2497 | 293.6733 | 203.8418
2012-10-30 00:55 | 25.9498 | 320.5804 | 251.7468 | 315.116 | 224.989 | 88.5372 | 1.551 | 39.4011 | 61.1107 | 25.4038 | 2.4617 | 39.2842 | 96.7076 | 15.9579 | 21.4541 | 6.1038 | 157.8183 | 299.1263 | 230.2927 | 293.6619 | 203.5349
2012-10-30 00:55 | 25.9423 | 320.5729 | 251.7276 | 315.1467 | 224.7769 | 88.5401 | 1.551 | 39.4013 | 61.1125 | 25.4034 | 2.4617 | 39.2845 | 96.7084 | 15.958 | 21.4546 | 6.104 | 157.8209 | 299.1183 | 230.273 | 293.6921 | 203.3223
2012-10-30 00:55 | 25.949 | 320.6216 | 251.7226 | 315.1992 | 225.0197 | 88.5441 | 1.551 | 39.4006 | 61.1092 | 25.4038 | 2.4617 | 39.2838 | 96.7058 | 15.9578 | 21.4452 | 6.1042 | 157.815 | 299.1764 | 230.2774 | 293.754 | 203.5745
2012-10-30 00:55 | 25.9483 | 320.6591 | 251.714 | 315.252 | 225.0819 | 88.558 | 1.551 | 39.4007 | 61.1101 | 25.4036 | 2.4617 | 39.2839 | 96.7061 | 15.9579 | 21.4646 | 6.1052 | 157.8162 | 299.1945 | 230.2494 | 293.7874 | 203.6173
2012-10-30 00:55 | 25.9438 | 320.6927 | 251.7486 | 315.2475 | 225.2348 | 88.5482 | 1.5509 | 39.3998 | 61.1069 | 25.4037 | 2.4617 | 39.2829 | 96.7008 | 15.9579 | 21.4601 | 6.1045 | 157.8076 | 299.2326 | 230.2885 | 293.7874 | 203.7747
2012-10-30 00:55 | 25.9453 | 320.679 | 251.7814 | 315.3063 | 225.3622 | 88.5522 | 1.5509 | 39.3976 | 61.0999 | 25.4038 | 2.4615 | 39.2808 | 96.6894 | 15.9581 | 21.4295 | 6.1048 | 157.7893 | 299.2495 | 230.3519 | 293.8768 | 203.9327
2012-10-30 00:55 | 25.9415 | 320.7362 | 251.8011 | 315.314 | 225.3741 | 88.5736 | 1.5508 | 39.3969 | 61.0978 | 25.4038 | 2.4614 | 39.2798 | 96.6843 | 15.9582 | 21.4377 | 6.1063 | 157.7821 | 299.2985 | 230.3634 | 293.8763 | 203.9364
2012-10-30 00:55 | 25.9505 | 320.7529 | 251.8609 | 315.3115 | 225.3045 | 88.5638 | 1.5507 | 39.3953 | 61.092 | 25.4041 | 2.4613 | 39.2783 | 96.6773 | 15.9581 | 21.4387 | 6.1056 | 157.7692 | 299.3142 | 230.4222 | 293.8728 | 203.8658
2012-10-30 00:55 | 25.9505 | 320.7872 | 251.8687 | 315.365 | 225.269 | 88.5858 | 1.5508 | 39.3956 | 61.0931 | 25.4041 | 2.4614 | 39.2789 | 96.6799 | 15.9581 | 21.4307 | 6.1071 | 157.7731 | 299.3565 | 230.438 | 293.9343 | 203.8383
2012-10-30 00:55 | 25.9468 | 320.8484 | 251.908 | 315.3766 | 225.3952 | 88.5748 | 1.5507 | 39.3953 | 61.0918 | 25.4043 | 2.4614 | 39.2786 | 96.6784 | 15.9582 | 21.4792 | 6.1063 | 157.7703 | 299.3692 | 230.4288 | 293.8974 | 203.916
2012-10-30 00:55 | 25.9475 | 320.8071 | 251.9243 | 315.4384 | 225.3054 | 88.5771 | 1.5507 | 39.394 | 61.0877 | 25.4042 | 2.4613 | 39.2772 | 96.6722 | 15.9581 | 21.4799 | 6.1065 | 157.7599 | 299.3272 | 230.4444 | 293.9585 | 203.8255
2012-10-30 00:55 | 25.9438 | 320.8568 | 251.8933 | 315.45 | 225.4866 | 88.591 | 1.5506 | 39.3918 | 61.0808 | 25.4043 | 2.4611 | 39.275 | 96.6614 | 15.958 | 21.44 | 6.1075 | 157.7423 | 299.4168 | 230.4533 | 294.01 | 204.0466
2012-10-30 00:55 | 25.9483 | 320.9147 | 251.9759 | 315.4469 | 225.495 | 88.5904 | 1.5505 | 39.3904 | 61.0766 | 25.4042 | 2.461 | 39.2736 | 96.654 | 15.9581 | 21.4486 | 6.1074 | 157.7306 | 299.4661 | 230.5273 | 293.9983 | 204.0464
2012-10-30 00:55 | 25.9513 | 320.9139 | 251.9985 | 315.4804 | 225.5571 | 88.5898 | 1.5505 | 39.3889 | 61.0716 | 25.4044 | 2.461 | 39.272 | 96.6466 | 15.9581 | 21.4435 | 6.1074 | 157.7181 | 299.4704 | 230.555 | 294.0369 | 204.1136
2012-10-30 00:55 | 25.946 | 320.9735 | 252.0088 | 315.5325 | 225.5163 | 88.5933 | 1.5505 | 39.3881 | 61.0697 | 25.4042 | 2.4609 | 39.2715 | 96.6434 | 15.9581 | 21.4222 | 6.1076 | 157.7131 | 299.5513 | 230.5866 | 294.1103 | 204.0941
2012-10-30 00:55 | 25.9498 | 320.9659 | 252.0751 | 315.5668 | 225.4651 | 88.5956 | 1.5504 | 39.3874 | 61.0672 | 25.4043 | 2.4609 | 39.2706 | 96.639 | 15.9581 | 21.438 | 6.1078 | 157.7062 | 299.5279 | 230.6371 | 294.1288 | 204.0271
2012-10-30 00:55 | 25.9557 | 321.0404 | 252.0967 | 315.5192 | 225.5024 | 88.5962 | 1.5504 | 39.3868 | 61.0647 | 25.4045 | 2.4608 | 39.27 | 96.6357 | 15.9582 | 21.4279 | 6.1078 | 157.7005 | 299.6125 | 230.6688 | 294.0913 | 204.0745
2012-10-30 00:56 | 25.958 | 321.0542 | 252.1146 | 315.6056 | 225.5087 | 88.6077 | 1.5503 | 39.3859 | 61.0618 | 25.4045 | 2.4607 | 39.2691 | 96.6315 | 15.9582 | 21.4342 | 6.1086 | 157.6933 | 299.62 | 230.6804 | 294.1714 | 204.0745
2012-10-30 00:56 | 25.955 | 321.0474 | 252.1233 | 315.6332 | 225.4899 | 88.6245 | 1.5503 | 39.3847 | 61.0582 | 25.4045 | 2.4607 | 39.268 | 96.6254 | 15.9583 | 21.4231 | 6.1098 | 157.6836 | 299.6243 | 230.7002 | 294.2101 | 204.0668
2012-10-30 00:56 | 25.9542 | 321.0924 | 252.1538 | 315.6247 | 225.3396 | 88.6135 | 1.5502 | 39.3827 | 61.0515 | 25.4047 | 2.4605 | 39.2657 | 96.6146 | 15.9582 | 21.4023 | 6.109 | 157.6661 | 299.6901 | 230.7515 | 294.2224 | 203.9373
2012-10-30 00:56 | 25.9655 | 321.1227 | 252.1026 | 315.6857 | 225.2604 | 88.6164 | 1.5502 | 39.3814 | 61.0482 | 25.4044 | 2.4605 | 39.2647 | 96.6097 | 15.9582 | 21.4216 | 6.1092 | 157.6579 | 299.7011 | 230.681 | 294.2641 | 203.8388
2012-10-30 00:56 | 25.97 | 321.1311 | 252.1227 | 315.7016 | 225.4066 | 88.6349 | 1.5501 | 39.3801 | 61.0441 | 25.4044 | 2.4604 | 39.2632 | 96.6028 | 15.9581 | 21.4221 | 6.1105 | 157.6469 | 299.709 | 230.7006 | 294.2795 | 203.9845
2012-10-30 00:56 | 25.9632 | 321.1624 | 252.1433 | 315.733 | 225.5493 | 88.6233 | 1.5501 | 39.3799 | 61.043 | 25.4047 | 2.4604 | 39.2633 | 96.6029 | 15.9582 | 21.3872 | 6.1097 | 157.6459 | 299.7752 | 230.7561 | 294.3458 | 204.1621
2012-10-30 00:56 | 25.9722 | 321.1409 | 252.2109 | 315.7268 | 225.6842 | 88.6268 | 1.55 | 39.3774 | 61.0356 | 25.4045 | 2.4602 | 39.2608 | 96.5909 | 15.9581 | 21.4363 | 6.1099 | 157.6266 | 299.7046 | 230.7746 | 294.2905 | 204.2479
2012-10-30 00:56 | 25.9685 | 321.1868 | 252.2267 | 315.746 | 225.7316 | 88.6297 | 1.55 | 39.3771 | 61.0345 | 25.4046 | 2.4602 | 39.2602 | 96.5881 | 15.9581 | 21.4447 | 6.1101 | 157.6226 | 299.7421 | 230.782 | 294.3013 | 204.2869
2012-10-30 00:56 | 25.973 | 321.2142 | 252.2742 | 315.8345 | 225.7322 | 88.6383 | 1.55 | 39.3759 | 61.031 | 25.4046 | 2.4601 | 39.2591 | 96.5827 | 15.9581 | 21.4331 | 6.1107 | 157.6137 | 299.7811 | 230.8411 | 294.4014 | 204.2991
2012-10-30 00:56 | 25.9745 | 321.2348 | 252.2679 | 315.7978 | 225.6196 | 88.647 | 1.5499 | 39.3742 | 61.0258 | 25.4045 | 2.46 | 39.2575 | 96.5754 | 15.958 | 21.4587 | 6.1113 | 157.6012 | 299.7761 | 230.8092 | 294.3391 | 204.1609
2012-10-30 00:56 | 25.973 | 321.2714 | 252.2977 | 315.8307 | 225.567 | 88.6343 | 1.5498 | 39.3723 | 61.0197 | 25.4046 | 2.4599 | 39.2557 | 96.5669 | 15.958 | 21.4692 | 6.1104 | 157.5866 | 299.8022 | 230.8285 | 294.3615 | 204.0978
2012-10-30 00:56 | 25.9789 | 321.3079 | 252.2879 | 315.8672 | 225.7499 | 88.6389 | 1.5475 | 39.3132 | 60.8376 | 25.4041 | 2.456 | 39.1921 | 96.2548 | 15.9579 | 21.4751 | 6.1108 | 157.0924 | 299.8328 | 230.8128 | 294.3921 | 204.2748
2012-10-30 00:56 | 25.976 | 321.2859 | 252.3085 | 315.8605 | 225.8099 | 88.6262 | 1.546 | 39.2742 | 60.7196 | 25.4031 | 2.4539 | 39.1579 | 96.0888 | 15.9576 | 21.4281 | 6.1099 | 156.8084 | 299.8578 | 230.8804 | 294.4324 | 204.3818
2012-10-30 00:56 | 25.9737 | 321.3179 | 252.3413 | 315.8582 | 225.8706 | 88.6337 | 1.5455 | 39.2615 | 60.6804 | 25.403 | 2.4531 | 39.1452 | 96.0267 | 15.9575 | 21.4338 | 6.1104 | 156.7071 | 299.8841 | 230.9075 | 294.4244 | 204.4368
2012-10-30 00:56 | 25.9797 | 321.3048 | 252.3903 | 315.9024 | 226.0221 | 88.6337 | 1.5452 | 39.253 | 60.6543 | 25.4029 | 2.4526 | 39.1367 | 95.9852 | 15.9575 | 21.4358 | 6.1104 | 156.6395 | 299.869 | 230.9545 | 294.4666 | 204.5863
2012-10-30 00:56 | 25.9842 | 321.2674 | 252.387 | 315.8763 | 225.2752 | 88.647 | 1.5453 | 39.2543 | 60.6578 | 25.4031 | 2.4526 | 39.1379 | 95.9914 | 15.9574 | 21.4604 | 6.1113 | 156.6492 | 299.807 | 230.9266 | 294.4159 | 203.8148
2012-10-30 00:56 | 25.9774 | 321.314 | 252.3216 | 315.8771 | 224.7294 | 88.636 | 1.5453 | 39.2547 | 60.6598 | 25.4029 | 2.4527 | 39.1383 | 95.9934 | 15.9574 | 21.4375 | 6.1106 | 156.6532 | 299.8765 | 230.8841 | 294.4396 | 203.2919
2012-10-30 00:56 | 25.9752 | 321.3423 | 252.167 | 315.9055 | 224.8098 | 88.6303 | 1.5454 | 39.257 | 60.6669 | 25.4028 | 2.4528 | 39.1405 | 96.0044 | 15.9574 | 21.3992 | 6.1102 | 156.6713 | 299.9431 | 230.7678 | 294.5063 | 203.4106
2012-10-30 00:56 | 25.9872 | 321.3963 | 252.0891 | 315.9252 | 224.7746 | 88.6291 | 1.5457 | 39.2651 | 60.6913 | 25.4032 | 2.4533 | 39.1485 | 96.0431 | 15.9575 | 21.3911 | 6.1101 | 156.7344 | 300.0052 | 230.698 | 294.5341 | 203.3835
2012-10-30 00:56 | 25.9737 | 321.3904 | 251.9466 | 315.9269 | 224.2928 | 88.6343 | 1.5459 | 39.2715 | 60.7109 | 25.4033 | 2.4537 | 39.155 | 96.0753 | 15.9574 | 21.4097 | 6.1104 | 156.7861 | 299.9807 | 230.5369 | 294.5172 | 202.8831
2012-10-30 00:56 | 25.9752 | 321.4072 | 251.7918 | 315.959 | 224.5776 | 88.6401 | 1.5462 | 39.278 | 60.7307 | 25.4033 | 2.4541 | 39.1616 | 96.1067 | 15.9576 | 21.3992 | 6.1108 | 156.8374 | 300.008 | 230.3926 | 294.5598 | 203.1784
2012-10-30 00:57 | 25.9602 | 321.4723 | 251.8784 | 315.9707 | 225.0269 | 88.6476 | 1.5501 | 39.3806 | 61.0446 | 25.4049 | 2.4604 | 39.2634 | 96.6035 | 15.9581 | 21.4203 | 6.1114 | 157.6481 | 300.052 | 230.4581 | 294.5504 | 203.6066
2012-10-30 00:57 | 25.9505 | 321.516 | 251.896 | 315.9954 | 225.2887 | 88.6534 | 1.5509 | 39.4019 | 61.1092 | 25.4054 | 2.4617 | 39.285 | 96.7088 | 15.9583 | 21.4708 | 6.1118 | 157.818 | 300.0452 | 230.4252 | 294.5246 | 203.8179
2012-10-30 00:57 | 25.9505 | 321.5236 | 251.982 | 315.9916 | 225.1982 | 88.6643 | 1.551 | 39.403 | 61.1127 | 25.4054 | 2.4618 | 39.286 | 96.7142 | 15.9583 | 21.4347 | 6.1125 | 157.8268 | 300.0889 | 230.5473 | 294.5569 | 203.7635
2012-10-30 00:57 | 25.9438 | 321.5284 | 251.9675 | 315.9925 | 225.1758 | 88.6719 | 1.551 | 39.4039 | 61.1155 | 25.4055 | 2.4619 | 39.2871 | 96.7194 | 15.9583 | 21.432 | 6.113 | 157.8349 | 300.0964 | 230.5355 | 294.5605 | 203.7438
2012-10-30 00:57 | 25.9475 | 321.6046 | 251.9751 | 316.0612 | 225.03 | 88.6742 | 1.551 | 39.4033 | 61.1133 | 25.4056 | 2.4618 | 39.2864 | 96.7153 | 15.9584 | 21.4237 | 6.1132 | 157.8286 | 300.1809 | 230.5514 | 294.6375 | 203.6063
2012-10-30 00:57 | 25.9348 | 321.5347 | 251.9859 | 315.9988 | 225.0212 | 88.6591 | 1.551 | 39.4037 | 61.115 | 25.4054 | 2.4619 | 39.2869 | 96.719 | 15.9582 | 21.3869 | 6.1122 | 157.834 | 300.1478 | 230.599 | 294.6119 | 203.6343
2012-10-30 00:57 | 25.949 | 321.5832 | 252.0313 | 316.0474 | 225.2557 | 88.6782 | 1.551 | 39.4034 | 61.1141 | 25.4054 | 2.4618 | 39.2865 | 96.7165 | 15.9583 | 21.4653 | 6.1135 | 157.8306 | 300.1179 | 230.566 | 294.5821 | 203.7904
2012-10-30 00:57 | 25.9325 | 321.5629 | 252.0266 | 316.0462 | 225.5343 | 88.6904 | 1.5509 | 39.4024 | 61.1109 | 25.4055 | 2.4617 | 39.2855 | 96.7108 | 15.9584 | 21.513 | 6.1143 | 157.8217 | 300.0499 | 230.5136 | 294.5332 | 204.0213
2012-10-30 00:57 | 25.9318 | 321.5851 | 252.1509 | 316.0837 | 225.7303 | 88.6915 | 1.5509 | 39.4017 | 61.1082 | 25.4057 | 2.4617 | 39.2851 | 96.709 | 15.9583 | 21.4601 | 6.1144 | 157.8172 | 300.125 | 230.6908 | 294.6236 | 204.2702
2012-10-30 00:57 | 25.9228 | 321.5799 | 252.2006 | 316.0785 | 225.8354 | 88.6921 | 1.5509 | 39.4004 | 61.1042 | 25.4056 | 2.4616 | 39.2836 | 96.7015 | 15.9584 | 21.4391 | 6.1144 | 157.8057 | 300.1408 | 230.7615 | 294.6394 | 204.3963
2012-10-30 00:57 | 25.9385 | 321.5994 | 252.2983 | 316.0827 | 225.8078 | 88.6956 | 1.5507 | 39.3982 | 61.0964 | 25.406 | 2.4615 | 39.2814 | 96.6905 | 15.9584 | 21.4829 | 6.1147 | 157.7869 | 300.1165 | 230.8154 | 294.5998 | 204.3249
2012-10-30 00:57 | 25.9355 | 321.6193 | 252.3539 | 316.0912 | 225.7419 | 88.7048 | 1.5507 | 39.3965 | 61.0918 | 25.4057 | 2.4614 | 39.2795 | 96.6813 | 15.9584 | 21.4959 | 6.1153 | 157.7731 | 300.1234 | 230.858 | 294.5953 | 204.246
2012-10-30 00:57 | 25.9333 | 321.6285 | 252.3752 | 316.089 | 225.7712 | 88.7146 | 1.5506 | 39.3953 | 61.0876 | 25.4059 | 2.4613 | 39.2784 | 96.6759 | 15.9584 | 21.4697 | 6.116 | 157.7635 | 300.1588 | 230.9055 | 294.6193 | 204.3015
2012-10-30 00:57 | 25.9243 | 321.6005 | 252.3975 | 316.1297 | 225.7307 | 88.7112 | 1.5507 | 39.3954 | 61.0886 | 25.4057 | 2.4613 | 39.2786 | 96.6768 | 15.9584 | 21.4727 | 6.1158 | 157.7654 | 300.1278 | 230.9248 | 294.657 | 204.258
2012-10-30 00:57 | 25.9258 | 321.6668 | 252.4576 | 316.1656 | 225.8896 | 88.71 | 1.5506 | 39.3943 | 61.0843 | 25.406 | 2.4612 | 39.2775 | 96.6709 | 15.9585 | 21.5063 | 6.1157 | 157.7552 | 300.1605 | 230.9513 | 294.6593 | 204.3833
2012-10-30 00:57 | 25.9281 | 321.6386 | 252.4755 | 316.1488 | 225.9784 | 88.7233 | 1.5506 | 39.3937 | 61.083 | 25.4058 | 2.4612 | 39.2769 | 96.6681 | 15.9585 | 21.4765 | 6.1166 | 157.7511 | 300.1621 | 230.999 | 294.6723 | 204.5019
2012-10-30 00:57 | 25.9221 | 321.6593 | 252.5203 | 316.1771 | 225.9409 | 88.7088 | 1.5505 | 39.3923 | 61.0781 | 25.406 | 2.4611 | 39.2756 | 96.6606 | 15.9586 | 21.4745 | 6.1156 | 157.7387 | 300.1848 | 231.0458 | 294.7026 | 204.4664
2012-10-30 00:57 | 25.9281 | 321.6653 | 252.5498 | 316.1946 | 225.9391 | 88.7198 | 1.5505 | 39.3922 | 61.0779 | 25.406 | 2.4611 | 39.2753 | 96.6593 | 15.9587 | 21.4484 | 6.1163 | 157.7372 | 300.2169 | 231.1014 | 294.7462 | 204.4907
2012-10-30 00:57 | 25.9161 | 321.6533 | 252.5807 | 316.2132 | 226.0687 | 88.7129 | 1.5504 | 39.3911 | 61.0737 | 25.4063 | 2.461 | 39.2741 | 96.6542 | 15.9585 | 21.4846 | 6.1159 | 157.7279 | 300.1687 | 231.0961 | 294.7286 | 204.5841
2012-10-30 00:57 | 25.9258 | 321.7012 | 252.6647 | 316.2229 | 226.1492 | 88.7256 | 1.5504 | 39.3896 | 61.0693 | 25.4062 | 2.4609 | 39.2727 | 96.6476 | 15.9585 | 21.5143 | 6.1167 | 157.7168 | 300.1869 | 231.1504 | 294.7086 | 204.6349
2012-10-30 00:57 | 25.9236 | 321.7257 | 252.7094 | 316.2589 | 226.1077 | 88.7297 | 1.5503 | 39.3881 | 61.0645 | 25.4062 | 2.4608 | 39.2713 | 96.6406 | 15.9584 | 21.5242 | 6.117 | 157.7051 | 300.2015 | 231.1852 | 294.7347 | 204.5835
2012-10-30 00:57 | 25.9198 | 321.7181 | 252.7173 | 316.2703 | 226.0724 | 88.7383 | 1.5502 | 39.3846 | 61.0538 | 25.4062 | 2.4606 | 39.2678 | 96.6222 | 15.9586 | 21.5083 | 6.1176 | 157.676 | 300.2098 | 231.209 | 294.762 | 204.5641
2012-10-30 00:57 | 25.9236 | 321.802 | 252.7094 | 316.2894 | 226.147 | 88.7395 | 1.5502 | 39.3838 | 61.0512 | 25.4062 | 2.4606 | 39.2669 | 96.6184 | 15.9585 | 21.4961 | 6.1177 | 157.6696 | 300.3059 | 231.2133 | 294.7933 | 204.6509
2012-10-30 00:58 | 25.9243 | 321.795 | 252.7648 | 316.2939 | 226.2893 | 88.7406 | 1.5502 | 39.3834 | 61.0502 | 25.4062 | 2.4605 | 39.2666 | 96.6164 | 15.9586 | 21.5289 | 6.1178 | 157.6667 | 300.2661 | 231.2359 | 294.765 | 204.7604
2012-10-30 00:58 | 25.9221 | 321.8234 | 252.8329 | 316.3223 | 226.3422 | 88.7464 | 1.5501 | 39.3826 | 61.047 | 25.4065 | 2.4605 | 39.2658 | 96.6125 | 15.9586 | 21.4464 | 6.1182 | 157.6595 | 300.377 | 231.3865 | 294.8759 | 204.8958
2012-10-30 00:58 | 25.9228 | 321.8317 | 252.8531 | 316.3574 | 226.4255 | 88.7493 | 1.55 | 39.3804 | 61.0402 | 25.4064 | 2.4604 | 39.2638 | 96.6029 | 15.9586 | 21.5073 | 6.1184 | 157.6432 | 300.3244 | 231.3458 | 294.8501 | 204.9182
2012-10-30 00:58 | 25.9303 | 321.8392 | 252.8958 | 316.384 | 226.3976 | 88.7406 | 1.5499 | 39.3787 | 61.0349 | 25.4065 | 2.4602 | 39.2619 | 96.5941 | 15.9585 | 21.4667 | 6.1178 | 157.629 | 300.3725 | 231.4291 | 294.9173 | 204.9309
2012-10-30 00:58 | 25.9318 | 321.8445 | 252.9559 | 316.3817 | 226.2064 | 88.7458 | 1.5499 | 39.3768 | 61.0295 | 25.4062 | 2.4601 | 39.2598 | 96.5838 | 15.9585 | 21.4922 | 6.1181 | 157.6133 | 300.3523 | 231.4637 | 294.8895 | 204.7142
2012-10-30 00:58 | 25.9266 | 321.8698 | 252.9312 | 316.407 | 226.1303 | 88.7447 | 1.5498 | 39.3758 | 61.026 | 25.4064 | 2.4601 | 39.2592 | 96.5803 | 15.9586 | 21.459 | 6.1181 | 157.6063 | 300.4108 | 231.4722 | 294.948 | 204.6713
2012-10-30 00:58 | 25.934 | 321.8925 | 252.9347 | 316.403 | 226.303 | 88.7464 | 1.5498 | 39.3751 | 61.024 | 25.4064 | 2.46 | 39.2584 | 96.5768 | 15.9585 | 21.4583 | 6.1182 | 157.6009 | 300.4342 | 231.4764 | 294.9447 | 204.8447
2012-10-30 00:58 | 25.94 | 321.8871 | 252.9563 | 316.4204 | 226.3837 | 88.7372 | 1.5498 | 39.3745 | 61.0221 | 25.4064 | 2.46 | 39.2576 | 96.5722 | 15.9586 | 21.4804 | 6.1175 | 157.5943 | 300.4067 | 231.4759 | 294.94 | 204.9033
2012-10-30 00:58 | 25.9355 | 321.9054 | 252.987 | 316.4618 | 226.4894 | 88.7586 | 1.5498 | 39.3749 | 61.0226 | 25.4066 | 2.46 | 39.2583 | 96.575 | 15.9587 | 21.4719 | 6.119 | 157.5976 | 300.4335 | 231.5151 | 294.9899 | 205.0175
2012-10-30 00:58 | 25.9296 | 321.9148 | 253.0358 | 316.4444 | 226.5385 | 88.7533 | 1.5487 | 39.3455 | 60.9324 | 25.4063 | 2.4584 | 39.232 | 96.4468 | 15.9585 | 21.4981 | 6.1187 | 157.3792 | 300.4167 | 231.5377 | 294.9463 | 205.0404
2012-10-30 00:58 | 25.9363 | 321.9787 | 253.0698 | 316.4587 | 226.4666 | 88.747 | 1.5459 | 39.2733 | 60.7117 | 25.4052 | 2.4537 | 39.1566 | 96.0788 | 15.9581 | 21.4646 | 6.1182 | 156.7906 | 300.5141 | 231.6052 | 294.9941 | 205.002
2012-10-30 00:58 | 25.94 | 321.929 | 253.1126 | 316.4968 | 226.2932 | 88.7441 | 1.5454 | 39.2606 | 60.6726 | 25.4051 | 2.4529 | 39.1442 | 96.0181 | 15.9581 | 21.4804 | 6.118 | 156.6906 | 300.4486 | 231.6322 | 295.0164 | 204.8128
2012-10-30 00:58 | 25.9363 | 321.9597 | 253.0307 | 316.5122 | 226.0378 | 88.7614 | 1.5452 | 39.2566 | 60.6607 | 25.4049 | 2.4527 | 39.1403 | 95.9995 | 15.958 | 21.4727 | 6.1192 | 156.6601 | 300.487 | 231.558 | 295.0395 | 204.5651
2012-10-30 00:58 | 25.9206 | 321.9325 | 253.0111 | 316.5194 | 226.0496 | 88.7372 | 1.5452 | 39.257 | 60.6616 | 25.4051 | 2.4527 | 39.1405 | 96.0004 | 15.9581 | 21.4489 | 6.1175 | 156.662 | 300.4836 | 231.5622 | 295.0705 | 204.6007
2012-10-30 00:58 | 25.9243 | 321.9591 | 252.9406 | 316.5002 | 225.9904 | 88.7516 | 1.5453 | 39.2586 | 60.6669 | 25.405 | 2.4528 | 39.1423 | 96.0097 | 15.958 | 21.4727 | 6.1185 | 156.6766 | 300.4864 | 231.4679 | 295.0275 | 204.5177
2012-10-30 00:58 | 25.9296 | 321.9911 | 252.9381 | 316.5666 | 226.2278 | 88.7551 | 1.5455 | 39.2621 | 60.6779 | 25.4049 | 2.4531 | 39.1457 | 96.0265 | 15.9579 | 21.49 | 6.1188 | 156.7044 | 300.5011 | 231.4481 | 295.0766 | 204.7378
2012-10-30 00:58 | 25.9198 | 321.966 | 252.9908 | 316.5759 | 226.336 | 88.7591 | 1.5456 | 39.2659 | 60.6894 | 25.405 | 2.4533 | 39.1495 | 96.0452 | 15.9579 | 21.4843 | 6.1191 | 156.7347 | 300.4817 | 231.5065 | 295.0916 | 204.8517
2012-10-30 00:58 | 25.9198 | 321.9775 | 252.9947 | 316.5683 | 226.3989 | 88.7406 | 1.5458 | 39.2704 | 60.7031 | 25.405 | 2.4536 | 39.154 | 96.0667 | 15.9581 | 21.5083 | 6.1178 | 156.7699 | 300.4692 | 231.4864 | 295.06 | 204.8906
2012-10-30 00:58 | 25.9131 | 321.9861 | 252.9411 | 316.5616 | 226.0185 | 88.7424 | 1.546 | 39.276 | 60.72 | 25.4052 | 2.4539 | 39.1595 | 96.0933 | 15.9581 | 21.4856 | 6.1179 | 156.8133 | 300.5005 | 231.4555 | 295.076 | 204.5329
2012-10-30 00:58 | 25.9108 | 321.9952 | 252.9154 | 316.5516 | 226.2129 | 88.7412 | 1.5476 | 39.3188 | 60.8518 | 25.4055 | 2.4566 | 39.2021 | 96.3019 | 15.9582 | 21.4672 | 6.1178 | 157.1537 | 300.528 | 231.4482 | 295.0844 | 204.7457
2012-10-30 00:58 | 25.9049 | 322.0122 | 252.9681 | 316.5496 | 226.2581 | 88.747 | 1.5494 | 39.364 | 60.9895 | 25.4064 | 2.4593 | 39.2473 | 96.5222 | 15.9585 | 21.4734 | 6.1182 | 157.5117 | 300.5388 | 231.4947 | 295.0762 | 204.7847
2012-10-30 00:58 | 25.9093 | 322.0814 | 252.9412 | 316.5731 | 226.0934 | 88.7429 | 1.5499 | 39.3767 | 61.0284 | 25.4066 | 2.4601 | 39.2595 | 96.583 | 15.9584 | 21.4738 | 6.1179 | 157.6114 | 300.6076 | 231.4674 | 295.0993 | 204.6196
2012-10-30 00:58 | 25.9093 | 322.0586 | 252.9256 | 316.5616 | 226.2153 | 88.769 | 1.5503 | 39.3873 | 61.0611 | 25.4067 | 2.4608 | 39.2704 | 96.6362 | 15.9585 | 21.4617 | 6.1197 | 157.6974 | 300.5969 | 231.4639 | 295.0999 | 204.7536
2012-10-30 00:58 | 25.9004 | 322.1031 | 252.9206 | 316.568 | 226.1631 | 88.762 | 1.5505 | 39.3943 | 61.0819 | 25.4071 | 2.4612 | 39.2776 | 96.671 | 15.9586 | 21.4528 | 6.1193 | 157.7529 | 300.6503 | 231.4678 | 295.1152 | 204.7103
2012-10-30 00:59 | 25.8951 | 322.1092 | 252.8997 | 316.5703 | 226.2562 | 88.7643 | 1.5507 | 39.3985 | 61.095 | 25.4071 | 2.4615 | 39.2817 | 96.6906 | 15.9587 | 21.4596 | 6.1194 | 157.7856 | 300.6496 | 231.4401 | 295.1107 | 204.7966
2012-10-30 00:59 | 25.8981 | 322.1313 | 252.9339 | 316.5657 | 226.1884 | 88.762 | 1.5508 | 39.4005 | 61.101 | 25.4072 | 2.4616 | 39.2837 | 96.7008 | 15.9586 | 21.4385 | 6.1193 | 157.8017 | 300.6928 | 231.4954 | 295.1272 | 204.7499
2012-10-30 00:59 | 25.8981 | 322.1122 | 252.9222 | 316.5771 | 226.3103 | 88.7759 | 1.5508 | 39.4024 | 61.106 | 25.4074 | 2.4617 | 39.2854 | 96.7082 | 15.9587 | 21.4345 | 6.1202 | 157.8142 | 300.6777 | 231.4877 | 295.1426 | 204.8758
2012-10-30 00:59 | 25.8921 | 322.0948 | 252.8889 | 316.6055 | 226.3476 | 88.7574 | 1.5509 | 39.4029 | 61.1081 | 25.4072 | 2.4617 | 39.2861 | 96.7123 | 15.9586 | 21.4325 | 6.1189 | 157.8203 | 300.6623 | 231.4564 | 295.173 | 204.9151
2012-10-30 00:59 | 25.8959 | 322.1062 | 252.9317 | 316.594 | 226.4301 | 88.7713 | 1.5509 | 39.4034 | 61.1098 | 25.4071 | 2.4618 | 39.2865 | 96.7141 | 15.9587 | 21.4523 | 6.1199 | 157.8239 | 300.6539 | 231.4794 | 295.1417 | 204.9778
2012-10-30 00:59 | 25.9056 | 322.0968 | 252.9532 | 316.6114 | 226.365 | 88.7811 | 1.5508 | 39.4029 | 61.1079 | 25.4073 | 2.4617 | 39.2859 | 96.7113 | 15.9587 | 21.446 | 6.1206 | 157.8191 | 300.6508 | 231.5072 | 295.1654 | 204.919
2012-10-30 00:59 | 25.9019 | 322.1389 | 252.969 | 316.6344 | 226.4636 | 88.7863 | 1.5508 | 39.4018 | 61.1045 | 25.4074 | 2.4617 | 39.285 | 96.7066 | 15.9587 | 21.4062 | 6.1209 | 157.8111 | 300.7327 | 231.5628 | 295.2282 | 205.0574
2012-10-30 00:59 | 25.8966 | 322.1145 | 253.0379 | 316.6253 | 226.5842 | 88.7643 | 1.5508 | 39.401 | 61.1015 | 25.4076 | 2.4616 | 39.2841 | 96.7012 | 15.9588 | 21.4851 | 6.1194 | 157.8028 | 300.6294 | 231.5528 | 295.1402 | 205.0991
2012-10-30 00:59 | 25.8951 | 322.1435 | 253.0521 | 316.6162 | 226.3585 | 88.7857 | 1.5508 | 39.4015 | 61.1031 | 25.4075 | 2.4616 | 39.2846 | 96.7043 | 15.9587 | 21.4716 | 6.1209 | 157.8073 | 300.6719 | 231.5805 | 295.1446 | 204.8869
2012-10-30 00:59 | 25.8891 | 322.1147 | 253.0187 | 316.6675 | 226.1204 | 88.7886 | 1.5507 | 39.4002 | 61.0988 | 25.4076 | 2.4615 | 39.2832 | 96.6972 | 15.9588 | 21.4174 | 6.1211 | 157.796 | 300.6973 | 231.6013 | 295.2501 | 204.703
2012-10-30 00:59 | 25.8936 | 322.123 | 252.988 | 316.6376 | 226.3452 | 88.7898 | 1.5507 | 39.3987 | 61.0941 | 25.4077 | 2.4614 | 39.2817 | 96.6899 | 15.9588 | 21.454 | 6.1212 | 157.784 | 300.669 | 231.534 | 295.1836 | 204.8912
2012-10-30 00:59 | 25.8936 | 322.1687 | 253.0076 | 316.6796 | 226.4553 | 88.7909 | 1.5505 | 39.3957 | 61.0848 | 25.4076 | 2.4613 | 39.2788 | 96.6751 | 15.9589 | 21.45 | 6.1212 | 157.7598 | 300.7187 | 231.5576 | 295.2296 | 205.0053
2012-10-30 00:59 | 25.8951 | 322.2008 | 253.0442 | 316.6887 | 226.4332 | 88.7984 | 1.5505 | 39.3942 | 61.0796 | 25.4079 | 2.4612 | 39.2774 | 96.6683 | 15.9589 | 21.4395 | 6.1218 | 157.7479 | 300.7613 | 231.6047 | 295.2492 | 204.9937
2012-10-30 00:59 | 25.8996 | 322.2015 | 253.08 | 316.6703 | 226.4771 | 88.8123 | 1.5504 | 39.3923 | 61.0742 | 25.4077 | 2.4611 | 39.2755 | 96.6597 | 15.9587 | 21.456 | 6.1227 | 157.7339 | 300.7455 | 231.624 | 295.2143 | 205.0211
2012-10-30 00:59 | 25.8914 | 322.1933 | 253.0952 | 316.7232 | 226.3705 | 88.8002 | 1.5503 | 39.3906 | 61.0688 | 25.4077 | 2.461 | 39.2738 | 96.6518 | 15.9586 | 21.4157 | 6.1219 | 157.7206 | 300.7776 | 231.6795 | 295.3075 | 204.9548
2012-10-30 00:59 | 25.8974 | 322.2221 | 253.1091 | 316.7063 | 226.3411 | 88.8083 | 1.5503 | 39.3895 | 61.0652 | 25.4078 | 2.4609 | 39.2727 | 96.6456 | 15.9588 | 21.4458 | 6.1224 | 157.7108 | 300.7763 | 231.6633 | 295.2605 | 204.8953
2012-10-30 00:59 | 25.8876 | 322.2162 | 253.1071 | 316.7271 | 226.3274 | 88.8146 | 1.5502 | 39.3873 | 61.0582 | 25.4079 | 2.4607 | 39.2706 | 96.6351 | 15.9588 | 21.444 | 6.1229 | 157.6933 | 300.7722 | 231.6631 | 295.2831 | 204.8834
2012-10-30 00:59 | 25.9019 | 322.2381 | 253.0862 | 316.7566 | 226.5148 | 88.8187 | 1.5501 | 39.385 | 61.0507 | 25.4081 | 2.4606 | 39.2682 | 96.6226 | 15.9589 | 21.4383 | 6.1232 | 157.6733 | 300.7998 | 231.6479 | 295.3183 | 205.0765
2012-10-30 00:59 | 25.9041 | 322.2403 | 253.1626 | 316.7818 | 226.6192 | 88.799 | 1.5501 | 39.384 | 61.0478 | 25.4079 | 2.4605 | 39.2672 | 96.6179 | 15.9589 | 21.4003 | 6.1218 | 157.6657 | 300.84 | 231.7623 | 295.3815 | 205.2189
2012-10-30 00:59 | 25.9176 | 322.25 | 253.1918 | 316.7914 | 226.5147 | 88.8088 | 1.55 | 39.3828 | 61.0443 | 25.4078 | 2.4605 | 39.266 | 96.6123 | 15.9588 | 21.4459 | 6.1225 | 157.6566 | 300.8041 | 231.7459 | 295.3455 | 205.0688
2012-10-30 00:59 | 25.9041 | 322.2327 | 253.1665 | 316.7627 | 226.576 | 88.81 | 1.55 | 39.3824 | 61.043 | 25.4079 | 2.4604 | 39.2656 | 96.6111 | 15.9587 | 21.4084 | 6.1226 | 157.6541 | 300.8243 | 231.7581 | 295.3543 | 205.1676
2012-10-30 00:59 | 25.9153 | 322.2591 | 253.2285 | 316.7968 | 226.658 | 88.8279 | 1.5499 | 39.381 | 61.0383 | 25.408 | 2.4603 | 39.2641 | 96.6024 | 15.9589 | 21.4115 | 6.1238 | 157.6407 | 300.8476 | 231.817 | 295.3853 | 205.2465
2012-10-30 01:00 | 25.9191 | 322.2858 | 253.2167 | 316.7586 | 226.5359 | 88.8198 | 1.5499 | 39.3803 | 61.0362 | 25.408 | 2.4603 | 39.2633 | 96.5988 | 15.9588 | 21.4314 | 6.1232 | 157.635 | 300.8544 | 231.7853 | 295.3272 | 205.1045
2012-10-30 01:00 | 25.9183 | 322.3499 | 253.2198 | 316.8074 | 226.421 | 88.8302 | 1.5499 | 39.3798 | 61.0347 | 25.408 | 2.4603 | 39.263 | 96.5981 | 15.9587 | 21.4386 | 6.124 | 157.6328 | 300.9113 | 231.7812 | 295.3688 | 204.9824
2012-10-30 01:00 | 25.9296 | 322.3039 | 253.2311 | 316.8684 | 226.4992 | 88.8262 | 1.5498 | 39.3781 | 61.0302 | 25.4077 | 2.4602 | 39.2613 | 96.5898 | 15.9587 | 21.4419 | 6.1237 | 157.62 | 300.862 | 231.7892 | 295.4265 | 205.0573
2012-10-30 01:00 | 25.9266 | 322.3276 | 253.2437 | 316.8234 | 226.5631 | 88.8169 | 1.5498 | 39.3764 | 61.0249 | 25.4077 | 2.4601 | 39.2597 | 96.5825 | 15.9586 | 21.4389 | 6.123 | 157.6074 | 300.8887 | 231.8048 | 295.3845 | 205.1242
2012-10-30 01:00 | 25.9273 | 322.2597 | 253.2757 | 316.8241 | 226.493 | 88.8308 | 1.5497 | 39.3756 | 61.022 | 25.4078 | 2.4601 | 39.2588 | 96.5787 | 15.9585 | 21.4235 | 6.124 | 157.6007 | 300.8362 | 231.8522 | 295.4006 | 205.0695
2012-10-30 01:00 | 25.9296 | 322.3497 | 253.2311 | 316.8722 | 226.3497 | 88.8319 | 1.5497 | 39.3751 | 61.021 | 25.4076 | 2.46 | 39.2581 | 96.5746 | 15.9586 | 21.4339 | 6.1241 | 157.5956 | 300.9158 | 231.7972 | 295.4383 | 204.9158
2012-10-30 01:00 | 25.9393 | 322.3747 | 253.1939 | 316.8743 | 226.3358 | 88.8395 | 1.5497 | 39.3737 | 61.0169 | 25.4076 | 2.4599 | 39.257 | 96.5701 | 15.9585 | 21.4476 | 6.1246 | 157.587 | 300.9271 | 231.7463 | 295.4267 | 204.8882
2012-10-30 01:00 | 25.9363 | 322.3679 | 253.2105 | 316.9209 | 226.38 | 88.8389 | 1.5463 | 39.2873 | 60.7512 | 25.4068 | 2.4546 | 39.1699 | 96.1448 | 15.9581 | 21.4125 | 6.1246 | 156.896 | 300.9554 | 231.798 | 295.5084 | 204.9675
2012-10-30 01:00 | 25.934 | 322.3656 | 253.1808 | 316.8919 | 226.4564 | 88.8233 | 1.5456 | 39.2667 | 60.6891 | 25.4062 | 2.4533 | 39.1503 | 96.0489 | 15.9579 | 21.4664 | 6.1235 | 156.738 | 300.8992 | 231.7144 | 295.4255 | 204.99
2012-10-30 01:00 | 25.9393 | 322.3632 | 253.2408 | 316.9201 | 226.682 | 88.8291 | 1.5452 | 39.258 | 60.6627 | 25.406 | 2.4528 | 39.1416 | 96.0073 | 15.9578 | 21.4275 | 6.1239 | 156.6699 | 300.9357 | 231.8133 | 295.4926 | 205.2545
2012-10-30 01:00 | 25.9296 | 322.3497 | 253.2155 | 316.9142 | 226.3536 | 88.8262 | 1.545 | 39.2534 | 60.6481 | 25.406 | 2.4525 | 39.1369 | 95.984 | 15.9578 | 21.4339 | 6.1237 | 156.6321 | 300.9158 | 231.7816 | 295.4803 | 204.9197
2012-10-30 01:00 | 25.9348 | 322.3397 | 253.2676 | 316.885 | 227.1062 | 88.8175 | 1.5452 | 39.2569 | 60.6587 | 25.4062 | 2.4527 | 39.1405 | 96.0018 | 15.9578 | 21.4872 | 6.1231 | 156.6604 | 300.8525 | 231.7804 | 295.3978 | 205.619
2012-10-30 01:00 | 25.937 | 322.3686 | 253.3362 | 316.9369 | 227.0809 | 88.8267 | 1.5454 | 39.2627 | 60.6764 | 25.4062 | 2.4531 | 39.1462 | 96.0292 | 15.9579 | 21.4373 | 6.1237 | 156.7056 | 300.9313 | 231.8989 | 295.4996 | 205.6436
2012-10-30 01:00 | 25.9251 | 322.3223 | 253.3516 | 316.9097 | 227.0769 | 88.8181 | 1.5456 | 39.2671 | 60.6898 | 25.4063 | 2.4534 | 39.1504 | 96.0497 | 15.9579 | 21.4414 | 6.1231 | 156.7395 | 300.8809 | 231.9102 | 295.4683 | 205.6355
2012-10-30 01:00 | 25.9296 | 322.3345 | 253.3757 | 316.918 | 226.7706 | 88.8111 | 1.5457 | 39.2709 | 60.7022 | 25.4061 | 2.4536 | 39.1543 | 96.0693 | 15.9579 | 21.4539 | 6.1226 | 156.7715 | 300.8806 | 231.9218 | 295.4641 | 205.3167
2012-10-30 01:00 | 25.9258 | 322.3688 | 253.3289 | 316.9524 | 226.3341 | 88.8036 | 1.5459 | 39.2751 | 60.7145 | 25.4064 | 2.4539 | 39.1586 | 96.0902 | 15.9579 | 21.4702 | 6.1221 | 156.8047 | 300.8986 | 231.8587 | 295.4822 | 204.8639
2012-10-30 01:00 | 25.9213 | 322.3643 | 253.2541 | 316.9174 | 226.9275 | 88.8221 | 1.5479 | 39.3272 | 60.8747 | 25.4067 | 2.4572 | 39.2118 | 96.35 | 15.9582 | 21.4135 | 6.1234 | 157.2247 | 300.9508 | 231.8406 | 295.5039 | 205.514
2012-10-30 01:00 | 25.9176 | 322.3492 | 253.3519 | 316.9404 | 227.1205 | 88.8343 | 1.5494 | 39.3662 | 60.9937 | 25.4076 | 2.4595 | 39.2495 | 96.5337 | 15.9584 | 21.4018 | 6.1242 | 157.5274 | 300.9474 | 231.9501 | 295.5386 | 205.7187
2012-10-30 01:00 | 25.9198 | 322.3552 | 253.3893 | 316.8891 | 226.7805 | 88.8233 | 1.5498 | 39.3769 | 61.0259 | 25.4079 | 2.4601 | 39.2599 | 96.5839 | 15.9586 | 21.412 | 6.1235 | 157.6099 | 300.9432 | 231.9773 | 295.4771 | 205.3685
2012-10-30 01:00 | 25.9191 | 322.385 | 253.3495 | 316.9152 | 226.1976 | 88.8279 | 1.5502 | 39.3879 | 61.0596 | 25.408 | 2.4608 | 39.2708 | 96.6373 | 15.9586 | 21.3913 | 6.1238 | 157.6969 | 300.9937 | 231.9582 | 295.5239 | 204.8063
2012-10-30 01:00 | 25.9206 | 322.3865 | 253.2846 | 316.9205 | 226.5689 | 88.8441 | 1.5505 | 39.3951 | 61.0817 | 25.4081 | 2.4612 | 39.2781 | 96.673 | 15.9586 | 21.3968 | 6.1249 | 157.7547 | 300.9897 | 231.8878 | 295.5237 | 205.1721
2012-10-30 01:00 | 25.9243 | 322.4245 | 253.2727 | 316.9166 | 226.6984 | 88.8441 | 1.5507 | 39.4001 | 61.0967 | 25.4084 | 2.4616 | 39.2831 | 96.6983 | 15.9585 | 21.4165 | 6.1249 | 157.795 | 301.008 | 231.8562 | 295.5001 | 205.2819
2012-10-30 01:00 | 25.9221 | 322.43 | 253.3057 | 316.9144 | 226.9008 | 88.84 | 1.5507 | 39.4019 | 61.1018 | 25.4085 | 2.4617 | 39.2849 | 96.7063 | 15.9586 | 21.4023 | 6.1246 | 157.8081 | 301.0277 | 231.9034 | 295.5121 | 205.4985
2012-10-30 01:01 | 25.9183 | 322.4185 | 253.3331 | 316.8991 | 227.1212 | 88.8597 | 1.5507 | 39.4023 | 61.103 | 25.4085 | 2.4617 | 39.2854 | 96.7089 | 15.9586 | 21.4226 | 6.126 | 157.8119 | 300.9959 | 231.9105 | 295.4765 | 205.6986
2012-10-30 01:01 | 25.9176 | 322.475 | 253.3871 | 316.9213 | 227.0536 | 88.862 | 1.5507 | 39.4012 | 61.0998 | 25.4085 | 2.4616 | 39.2843 | 96.7032 | 15.9587 | 21.4058 | 6.1261 | 157.8029 | 301.0692 | 231.9813 | 295.5155 | 205.6478
2012-10-30 01:01 | 25.9176 | 322.4636 | 253.4574 | 316.948 | 227.0969 | 88.8608 | 1.5507 | 39.4007 | 61.0981 | 25.4086 | 2.4616 | 39.2837 | 96.7002 | 15.9587 | 21.4058 | 6.1261 | 157.7983 | 301.0578 | 232.0516 | 295.5422 | 205.6911
2012-10-30 01:01 | 25.9236 | 322.4467 | 253.479 | 316.9617 | 227.0675 | 88.8626 | 1.5506 | 39.3993 | 61.0943 | 25.4084 | 2.4615 | 39.2824 | 96.6934 | 15.9587 | 21.3998 | 6.1262 | 157.7876 | 301.0469 | 232.0792 | 295.5619 | 205.6677
2012-10-30 01:01 | 25.9251 | 322.4711 | 253.4688 | 316.967 | 227.0415 | 88.8626 | 1.5507 | 39.4007 | 61.0981 | 25.4086 | 2.4616 | 39.2837 | 96.7005 | 15.9586 | 21.4213 | 6.1262 | 157.7986 | 301.0498 | 232.0475 | 295.5457 | 205.6202
2012-10-30 01:01 | 25.9168 | 322.4666 | 253.4879 | 316.9625 | 226.8522 | 88.8585 | 1.5506 | 39.3994 | 61.0937 | 25.4088 | 2.4615 | 39.2828 | 96.6954 | 15.9587 | 21.4211 | 6.1259 | 157.7891 | 301.0455 | 232.0668 | 295.5414 | 205.4311
2012-10-30 01:01 | 25.9161 | 322.4545 | 253.5028 | 316.9656 | 227.0285 | 88.8689 | 1.5506 | 39.4 | 61.0955 | 25.4087 | 2.4615 | 39.283 | 96.697 | 15.9587 | 21.4204 | 6.1266 | 157.7925 | 301.0341 | 232.0824 | 295.5452 | 205.6081
2012-10-30 01:01 | 25.9221 | 322.5101 | 253.4854 | 317.0442 | 226.9519 | 88.8712 | 1.5506 | 39.399 | 61.0925 | 25.4087 | 2.4615 | 39.2818 | 96.6905 | 15.9588 | 21.4585 | 6.1268 | 157.783 | 301.0516 | 232.0269 | 295.5857 | 205.4934
2012-10-30 01:01 | 25.9153 | 322.4766 | 253.5059 | 316.9954 | 227.0356 | 88.8637 | 1.5505 | 39.3977 | 61.088 | 25.4089 | 2.4614 | 39.2808 | 96.6853 | 15.9588 | 21.4316 | 6.1263 | 157.7733 | 301.045 | 232.0743 | 295.5638 | 205.604
2012-10-30 01:01 | 25.9228 | 322.5108 | 253.529 | 316.9647 | 227.0667 | 88.862 | 1.5505 | 39.3966 | 61.085 | 25.4088 | 2.4614 | 39.2799 | 96.6815 | 15.9587 | 21.4311 | 6.1261 | 157.7665 | 301.0797 | 232.0979 | 295.5336 | 205.6356
2012-10-30 01:01 | 25.9161 | 322.5079 | 253.5536 | 317.0382 | 226.8318 | 88.8845 | 1.5505 | 39.3964 | 61.0848 | 25.4085 | 2.4613 | 39.2795 | 96.6797 | 15.9586 | 21.4364 | 6.1277 | 157.7645 | 301.0715 | 232.1172 | 295.6018 | 205.3954
2012-10-30 01:01 | 25.9236 | 322.5344 | 253.5415 | 317.0724 | 226.9692 | 88.8712 | 1.5505 | 39.3947 | 61.0795 | 25.4086 | 2.4612 | 39.2778 | 96.6721 | 15.9586 | 21.4239 | 6.1268 | 157.7516 | 301.1105 | 232.1176 | 295.6485 | 205.5453
2012-10-30 01:01 | 25.9146 | 322.5407 | 253.5638 | 317.0825 | 227.2198 | 88.8776 | 1.5504 | 39.3937 | 61.0764 | 25.4086 | 2.4612 | 39.2767 | 96.6668 | 15.9586 | 21.3948 | 6.1272 | 157.7432 | 301.1459 | 232.169 | 295.6877 | 205.825
2012-10-30 01:01 | 25.9153 | 322.549 | 253.627 | 317.0756 | 227.1536 | 88.8759 | 1.5504 | 39.3933 | 61.0745 | 25.4088 | 2.4611 | 39.2762 | 96.6642 | 15.9586 | 21.4035 | 6.1271 | 157.7387 | 301.1455 | 232.2235 | 295.6721 | 205.7501
2012-10-30 01:01 | 25.9168 | 322.5582 | 253.609 | 317.0656 | 227.0568 | 88.8799 | 1.5503 | 39.3908 | 61.0672 | 25.4087 | 2.461 | 39.274 | 96.6527 | 15.9586 | 21.393 | 6.1274 | 157.7199 | 301.1652 | 232.216 | 295.6726 | 205.6638
2012-10-30 01:01 | 25.9183 | 322.533 | 253.63 | 317.0824 | 227.0465 | 88.888 | 1.5502 | 39.3897 | 61.0637 | 25.4087 | 2.4609 | 39.273 | 96.648 | 15.9586 | 21.4186 | 6.1279 | 157.7117 | 301.1144 | 232.2114 | 295.6638 | 205.6279
2012-10-30 01:01 | 25.9281 | 322.5962 | 253.7023 | 317.1075 | 227.0405 | 88.8874 | 1.5502 | 39.3878 | 61.0574 | 25.4088 | 2.4608 | 39.2709 | 96.6386 | 15.9585 | 21.4163 | 6.1279 | 157.696 | 301.1799 | 232.286 | 295.6912 | 205.6242
2012-10-30 01:01 | 25.9236 | 322.5993 | 253.6665 | 317.0763 | 226.5601 | 88.8799 | 1.5501 | 39.3863 | 61.0531 | 25.4087 | 2.4607 | 39.2696 | 96.6314 | 15.9586 | 21.4158 | 6.1274 | 157.6846 | 301.1835 | 232.2507 | 295.6605 | 205.1443
2012-10-30 01:01 | 25.9198 | 322.6031 | 253.6081 | 317.1183 | 226.6035 | 88.8949 | 1.55 | 39.3849 | 61.0484 | 25.4089 | 2.4606 | 39.268 | 96.6243 | 15.9585 | 21.4201 | 6.1284 | 157.6726 | 301.183 | 232.188 | 295.6982 | 205.1834
2012-10-30 01:01 | 25.9228 | 322.5909 | 253.615 | 317.0946 | 226.7284 | 88.8903 | 1.55 | 39.3843 | 61.0468 | 25.4088 | 2.4606 | 39.2676 | 96.6226 | 15.9584 | 21.415 | 6.1281 | 157.6694 | 301.1759 | 232.2 | 295.6796 | 205.3134
2012-10-30 01:01 | 25.9281 | 322.5962 | 253.5617 | 317.1381 | 226.5016 | 88.8869 | 1.55 | 39.3833 | 61.0438 | 25.4087 | 2.4606 | 39.2666 | 96.618 | 15.9584 | 21.3842 | 6.1279 | 157.6617 | 301.212 | 232.1775 | 295.7539 | 205.1174
2012-10-30 01:01 | 25.937 | 322.5784 | 253.5354 | 317.1546 | 226.9472 | 88.8996 | 1.55 | 39.3824 | 61.0411 | 25.4087 | 2.4605 | 39.2656 | 96.6126 | 15.9585 | 21.4212 | 6.1287 | 157.6537 | 301.1572 | 232.1142 | 295.7334 | 205.526
2012-10-30 01:01 | 25.9153 | 322.5872 | 253.502 | 317.1023 | 226.894 | 88.8955 | 1.5499 | 39.3807 | 61.0352 | 25.4089 | 2.4604 | 39.2637 | 96.6042 | 15.9583 | 21.3955 | 6.1285 | 157.6394 | 301.1917 | 232.1065 | 295.7068 | 205.4985
2012-10-30 01:02 | 25.9251 | 322.6504 | 253.586 | 317.1618 | 226.7858 | 88.9007 | 1.5498 | 39.3793 | 61.0312 | 25.4088 | 2.4603 | 39.2624 | 96.5976 | 15.9584 | 21.4013 | 6.1288 | 157.6288 | 301.2491 | 232.1847 | 295.7605 | 205.3845
2012-10-30 01:02 | 25.9258 | 322.6168 | 253.5555 | 317.1778 | 226.8101 | 88.8978 | 1.5498 | 39.3782 | 61.0287 | 25.4084 | 2.4602 | 39.2615 | 96.5931 | 15.9583 | 21.4261 | 6.1286 | 157.6217 | 301.1907 | 232.1294 | 295.7517 | 205.384
2012-10-30 01:02 | 25.9303 | 322.6251 | 253.5522 | 317.1593 | 226.9641 | 88.888 | 1.5498 | 39.3773 | 61.0254 | 25.4086 | 2.4602 | 39.2605 | 96.589 | 15.9582 | 21.4025 | 6.1279 | 157.6145 | 301.2226 | 232.1497 | 295.7568 | 205.5616
2012-10-30 01:02 | 25.9236 | 322.5993 | 253.565 | 317.1794 | 227.0046 | 88.9157 | 1.5497 | 39.3758 | 61.0206 | 25.4087 | 2.4601 | 39.259 | 96.5819 | 15.9582 | 21.4118 | 6.1298 | 157.6024 | 301.1875 | 232.1532 | 295.7676 | 205.5928
2012-10-30 01:02 | 25.9236 | 322.6298 | 253.5728 | 317.2099 | 226.7685 | 88.9042 | 1.5497 | 39.3747 | 61.0173 | 25.4086 | 2.46 | 39.2576 | 96.5749 | 15.9581 | 21.4239 | 6.1291 | 157.5922 | 301.2059 | 232.1489 | 295.786 | 205.3446
2012-10-30 01:02 | 25.9258 | 322.6396 | 253.6062 | 317.2274 | 226.7117 | 88.9111 | 1.5496 | 39.3732 | 61.0127 | 25.4086 | 2.46 | 39.2564 | 96.5697 | 15.958 | 21.406 | 6.1295 | 157.5824 | 301.2336 | 232.2002 | 295.8214 | 205.3057
2012-10-30 01:02 | 25.9355 | 322.7218 | 253.5847 | 317.2218 | 226.6664 | 88.8961 | 1.5463 | 39.2865 | 60.7468 | 25.4075 | 2.4546 | 39.1696 | 96.1447 | 15.9578 | 21.4077 | 6.1285 | 156.8915 | 301.3141 | 232.177 | 295.8141 | 205.2587
2012-10-30 01:02 | 25.9363 | 322.6616 | 253.5542 | 317.1883 | 226.9071 | 88.8822 | 1.5455 | 39.2671 | 60.6882 | 25.407 | 2.4534 | 39.1509 | 96.0541 | 15.9576 | 21.4325 | 6.1275 | 156.7423 | 301.2291 | 232.1217 | 295.7558 | 205.4746
2012-10-30 01:02 | 25.9355 | 322.6837 | 253.5534 | 317.2371 | 226.7136 | 88.8967 | 1.5451 | 39.2562 | 60.6542 | 25.4071 | 2.4527 | 39.1397 | 95.9994 | 15.9576 | 21.4237 | 6.1285 | 156.6536 | 301.26 | 232.1297 | 295.8134 | 205.2899
2012-10-30 01:02 | 25.9296 | 322.6625 | 253.5163 | 317.2236 | 226.1019 | 88.8869 | 1.545 | 39.255 | 60.6508 | 25.407 | 2.4527 | 39.1384 | 95.9935 | 15.9575 | 21.4098 | 6.1279 | 156.6442 | 301.2527 | 232.1065 | 295.8138 | 204.6921
2012-10-30 01:02 | 25.9251 | 322.6809 | 253.418 | 317.2114 | 226.239 | 88.8857 | 1.5451 | 39.2548 | 60.6509 | 25.4067 | 2.4527 | 39.1384 | 95.9934 | 15.9575 | 21.4133 | 6.1278 | 156.6443 | 301.2676 | 232.0047 | 295.7981 | 204.8257
2012-10-30 01:02 | 25.9266 | 322.6481 | 253.3844 | 317.2511 | 226.3034 | 88.888 | 1.5451 | 39.2572 | 60.658 | 25.4068 | 2.4528 | 39.1405 | 96.004 | 15.9575 | 21.3947 | 6.1279 | 156.662 | 301.2534 | 231.9897 | 295.8564 | 204.9087
2012-10-30 01:02 | 25.9206 | 322.665 | 253.2729 | 317.2642 | 226.4076 | 88.8811 | 1.5454 | 39.2635 | 60.6771 | 25.407 | 2.4532 | 39.1469 | 96.0355 | 15.9574 | 21.4249 | 6.1275 | 156.7126 | 301.2401 | 231.848 | 295.8393 | 204.9827
2012-10-30 01:02 | 25.9258 | 322.6587 | 253.2508 | 317.216 | 226.7078 | 88.8828 | 1.5455 | 39.2673 | 60.6886 | 25.407 | 2.4534 | 39.1507 | 96.0536 | 15.9576 | 21.422 | 6.1276 | 156.7423 | 301.2367 | 231.8288 | 295.794 | 205.2858
2012-10-30 01:02 | 25.9153 | 322.6215 | 253.3145 | 317.2513 | 227.3699 | 88.8822 | 1.5457 | 39.2717 | 60.7022 | 25.4071 | 2.4537 | 39.1551 | 96.0752 | 15.9575 | 21.4156 | 6.1275 | 156.7774 | 301.2059 | 231.8989 | 295.8357 | 205.9543
2012-10-30 01:02 | 25.9161 | 322.6376 | 253.3973 | 317.2254 | 227.0875 | 88.8869 | 1.5473 | 39.3118 | 60.8253 | 25.4075 | 2.4561 | 39.1941 | 96.2656 | 15.9577 | 21.4485 | 6.1279 | 157.0909 | 301.1891 | 231.9488 | 295.7769 | 205.639
2012-10-30 01:02 | 25.9131 | 322.6689 | 253.3982 | 317.2224 | 226.4119 | 88.8903 | 1.5494 | 39.368 | 60.9969 | 25.4084 | 2.4596 | 39.251 | 96.5435 | 15.958 | 21.4535 | 6.1281 | 157.5404 | 301.2154 | 231.9447 | 295.7689 | 204.9584
2012-10-30 01:02 | 25.9071 | 322.6782 | 253.3453 | 317.2164 | 226.0558 | 88.8926 | 1.5499 | 39.3805 | 61.0354 | 25.4086 | 2.4604 | 39.2637 | 96.6056 | 15.9581 | 21.4475 | 6.1283 | 157.641 | 301.2307 | 231.8978 | 295.7689 | 204.6083
2012-10-30 01:02 | 25.9138 | 322.7077 | 253.3013 | 317.2001 | 226.0428 | 88.8811 | 1.5503 | 39.3901 | 61.065 | 25.4086 | 2.461 | 39.2733 | 96.6527 | 15.9581 | 21.4502 | 6.1275 | 157.7178 | 301.2575 | 231.8511 | 295.7499 | 204.5926
2012-10-30 01:02 | 25.9071 | 322.6972 | 253.193 | 317.1629 | 226.3469 | 88.9152 | 1.5505 | 39.3972 | 61.0859 | 25.4091 | 2.4615 | 39.2802 | 96.6865 | 15.9581 | 21.4074 | 6.1298 | 157.7725 | 301.2898 | 231.7856 | 295.7555 | 204.9395
2012-10-30 01:02 | 25.9034 | 322.6706 | 253.2205 | 317.1859 | 226.8191 | 88.8984 | 1.5506 | 39.4009 | 61.0969 | 25.4093 | 2.4617 | 39.2838 | 96.7038 | 15.9582 | 21.4598 | 6.1287 | 157.8007 | 301.2108 | 231.7607 | 295.7261 | 205.3593
2012-10-30 01:02 | 25.9093 | 322.6842 | 253.371 | 317.1345 | 226.8132 | 88.9152 | 1.5507 | 39.4011 | 61.0977 | 25.4092 | 2.4617 | 39.2842 | 96.7052 | 15.9583 | 21.5059 | 6.1298 | 157.8029 | 301.1783 | 231.8651 | 295.6286 | 205.3073
2012-10-30 01:02 | 25.9101 | 322.7117 | 253.3796 | 317.2346 | 226.8337 | 88.9204 | 1.5507 | 39.4021 | 61.1013 | 25.409 | 2.4618 | 39.2851 | 96.7103 | 15.9581 | 21.5187 | 6.1302 | 157.8116 | 301.193 | 231.8609 | 295.7159 | 205.315
2012-10-30 01:03 | 25.9041 | 322.7133 | 253.3814 | 317.1713 | 226.6035 | 88.9077 | 1.5508 | 39.4034 | 61.1052 | 25.4091 | 2.4618 | 39.2863 | 96.716 | 15.9582 | 21.4726 | 6.1293 | 157.8212 | 301.2407 | 231.9088 | 295.6987 | 205.1309
2012-10-30 01:03 | 25.8951 | 322.6623 | 253.3959 | 317.1662 | 226.8226 | 88.9134 | 1.5507 | 39.403 | 61.1035 | 25.4092 | 2.4618 | 39.2861 | 96.7144 | 15.9583 | 21.4917 | 6.1297 | 157.8179 | 301.1706 | 231.9042 | 295.6745 | 205.3309
2012-10-30 01:03 | 25.8847 | 322.6634 | 253.3894 | 317.1787 | 226.7926 | 88.9169 | 1.5507 | 39.402 | 61.1004 | 25.4093 | 2.4617 | 39.2852 | 96.7102 | 15.9583 | 21.4492 | 6.1299 | 157.8106 | 301.2142 | 231.9402 | 295.7295 | 205.3434
2012-10-30 01:03 | 25.8974 | 322.6341 | 253.5153 | 317.1914 | 226.8643 | 88.9048 | 1.5507 | 39.4017 | 61.0992 | 25.4094 | 2.4617 | 39.2847 | 96.7077 | 15.9582 | 21.502 | 6.1291 | 157.8069 | 301.1321 | 232.0133 | 295.6894 | 205.3623
2012-10-30 01:03 | 25.8906 | 322.6693 | 253.489 | 317.1769 | 226.8968 | 88.9082 | 1.5507 | 39.401 | 61.0974 | 25.4093 | 2.4617 | 39.284 | 96.7049 | 15.9582 | 21.4631 | 6.1293 | 157.8023 | 301.2062 | 232.0259 | 295.7138 | 205.4337
2012-10-30 01:03 | 25.8914 | 322.6434 | 253.4898 | 317.193 | 226.878 | 88.9302 | 1.5506 | 39.3999 | 61.0935 | 25.4094 | 2.4616 | 39.2827 | 96.6983 | 15.9582 | 21.4679 | 6.1308 | 157.7918 | 301.1755 | 232.0219 | 295.7251 | 205.4101
2012-10-30 01:03 | 25.8847 | 322.6481 | 253.5534 | 317.1978 | 226.9342 | 88.9314 | 1.5506 | 39.3987 | 61.0905 | 25.4092 | 2.4615 | 39.2818 | 96.6931 | 15.9583 | 21.5013 | 6.1309 | 157.7836 | 301.1468 | 232.0521 | 295.6965 | 205.4329
2012-10-30 01:03 | 25.8817 | 322.6756 | 253.5661 | 317.191 | 226.821 | 88.9181 | 1.5505 | 39.3967 | 61.0834 | 25.4096 | 2.4614 | 39.2798 | 96.683 | 15.9584 | 21.4702 | 6.13 | 157.7664 | 301.2054 | 232.0959 | 295.7208 | 205.3508
2012-10-30 01:03 | 25.8899 | 322.6304 | 253.5391 | 317.2297 | 226.8607 | 88.9244 | 1.5504 | 39.3956 | 61.0793 | 25.4098 | 2.4613 | 39.2787 | 96.678 | 15.9583 | 21.5065 | 6.1304 | 157.7573 | 301.1239 | 232.0326 | 295.7232 | 205.3542
2012-10-30 01:03 | 25.8906 | 322.6922 | 253.5789 | 317.2228 | 226.9322 | 88.9256 | 1.5504 | 39.3957 | 61.0801 | 25.4096 | 2.4613 | 39.2788 | 96.6785 | 15.9583 | 21.4751 | 6.1305 | 157.7586 | 301.2171 | 232.1038 | 295.7477 | 205.4571
2012-10-30 01:03 | 25.8899 | 322.6762 | 253.5547 | 317.2297 | 226.8922 | 88.9279 | 1.5504 | 39.3941 | 61.0752 | 25.4096 | 2.4612 | 39.2771 | 96.6706 | 15.9582 | 21.4544 | 6.1307 | 157.7458 | 301.2218 | 232.1003 | 295.7753 | 205.4378
2012-10-30 01:03 | 25.8869 | 322.7152 | 253.5908 | 317.2191 | 226.8656 | 88.9539 | 1.5503 | 39.3933 | 61.0727 | 25.4096 | 2.4612 | 39.2764 | 96.6668 | 15.9583 | 21.4714 | 6.1325 | 157.7395 | 301.2438 | 232.1194 | 295.7477 | 205.3942
2012-10-30 01:03 | 25.8936 | 322.7219 | 253.5975 | 317.2487 | 226.9077 | 88.9302 | 1.5502 | 39.3908 | 61.0648 | 25.4097 | 2.461 | 39.274 | 96.6546 | 15.9583 | 21.4982 | 6.1308 | 157.7195 | 301.2237 | 232.0993 | 295.7505 | 205.4095
2012-10-30 01:03 | 25.8817 | 322.6947 | 253.6051 | 317.2444 | 227.0256 | 88.9342 | 1.5502 | 39.3888 | 61.0588 | 25.4096 | 2.4609 | 39.272 | 96.6446 | 15.9584 | 21.4542 | 6.1311 | 157.7034 | 301.2405 | 232.1509 | 295.7902 | 205.5714
2012-10-30 01:03 | 25.8727 | 322.7315 | 253.6078 | 317.2583 | 226.8396 | 88.9366 | 1.5501 | 39.3877 | 61.0551 | 25.4097 | 2.4608 | 39.2709 | 96.6394 | 15.9583 | 21.401 | 6.1313 | 157.6945 | 301.3305 | 232.2068 | 295.8573 | 205.4386
2012-10-30 01:03 | 25.8929 | 322.7517 | 253.6163 | 317.2709 | 226.8401 | 88.9504 | 1.5501 | 39.388 | 61.0564 | 25.4096 | 2.4609 | 39.2711 | 96.6415 | 15.9582 | 21.4533 | 6.1322 | 157.6979 | 301.2984 | 232.163 | 295.8176 | 205.3868
2012-10-30 01:03 | 25.8884 | 322.7396 | 253.6001 | 317.2511 | 226.9458 | 88.9481 | 1.5501 | 39.3877 | 61.0554 | 25.4096 | 2.4608 | 39.2705 | 96.6387 | 15.9582 | 21.4729 | 6.1321 | 157.694 | 301.2667 | 232.1272 | 295.7782 | 205.4729
2012-10-30 01:03 | 25.8891 | 322.7517 | 253.636 | 317.2709 | 227.0291 | 88.9591 | 1.55 | 39.3857 | 61.0494 | 25.4095 | 2.4607 | 39.2689 | 96.6308 | 15.9582 | 21.4576 | 6.1328 | 157.6802 | 301.2941 | 232.1784 | 295.8133 | 205.5715
2012-10-30 01:03 | 25.8832 | 322.7534 | 253.6613 | 317.2612 | 226.9878 | 88.9446 | 1.55 | 39.3845 | 61.0457 | 25.4095 | 2.4607 | 39.2675 | 96.6237 | 15.9581 | 21.4196 | 6.1318 | 157.6694 | 301.3338 | 232.2417 | 295.8416 | 205.5682
2012-10-30 01:03 | 25.8869 | 322.7381 | 253.6455 | 317.3107 | 227.0387 | 88.9504 | 1.5499 | 39.3832 | 61.0411 | 25.4097 | 2.4606 | 39.2663 | 96.6173 | 15.9582 | 21.4514 | 6.1322 | 157.6583 | 301.2867 | 232.1941 | 295.8593 | 205.5873
2012-10-30 01:03 | 25.8936 | 322.7333 | 253.6639 | 317.2869 | 227.1358 | 88.9279 | 1.5499 | 39.382 | 61.0372 | 25.4097 | 2.4605 | 39.265 | 96.6115 | 15.9582 | 21.4059 | 6.1307 | 157.6487 | 301.3274 | 232.258 | 295.881 | 205.7299
2012-10-30 01:03 | 25.8884 | 322.7586 | 253.7173 | 317.3199 | 227.3391 | 88.9539 | 1.5498 | 39.3804 | 61.032 | 25.4099 | 2.4604 | 39.2635 | 96.603 | 15.9584 | 21.4609 | 6.1325 | 157.635 | 301.2977 | 232.2564 | 295.859 | 205.8782
2012-10-30 01:03 | 25.9004 | 322.7745 | 253.7293 | 317.3204 | 227.2606 | 88.9516 | 1.5497 | 39.3788 | 61.027 | 25.4098 | 2.4603 | 39.2618 | 96.595 | 15.9583 | 21.4448 | 6.1323 | 157.6221 | 301.3297 | 232.2845 | 295.8756 | 205.8158
2012-10-30 01:04 | 25.9019 | 322.7836 | 253.762 | 317.3334 | 227.2779 | 88.9464 | 1.5497 | 39.3773 | 61.0221 | 25.41 | 2.4602 | 39.2604 | 96.5877 | 15.9584 | 21.4583 | 6.132 | 157.6097 | 301.3253 | 232.3037 | 295.8751 | 205.8196
2012-10-30 01:04 | 25.8974 | 322.7905 | 253.7458 | 317.3671 | 227.2537 | 88.9475 | 1.5496 | 39.3764 | 61.0196 | 25.4099 | 2.4601 | 39.2596 | 96.5841 | 15.9583 | 21.4578 | 6.132 | 157.6037 | 301.3327 | 232.288 | 295.9093 | 205.7959
2012-10-30 01:04 | 25.9019 | 322.795 | 253.8402 | 317.299 | 227.2228 | 88.9504 | 1.5496 | 39.3756 | 61.0171 | 25.4098 | 2.4601 | 39.2585 | 96.5784 | 15.9583 | 21.4302 | 6.1322 | 157.5955 | 301.3648 | 232.41 | 295.8688 | 205.7926
2012-10-30 01:04 | 25.9049 | 322.8095 | 253.7924 | 317.3325 | 227.214 | 88.9429 | 1.5496 | 39.3751 | 61.015 | 25.4102 | 2.4601 | 39.2585 | 96.578 | 15.9584 | 21.3891 | 6.1317 | 157.593 | 301.4204 | 232.4033 | 295.9434 | 205.8249
2012-10-30 01:04 | 25.9176 | 322.8222 | 253.8207 | 317.3682 | 227.1048 | 88.9602 | 1.5468 | 39.3016 | 60.7898 | 25.4091 | 2.4553 | 39.1816 | 96.2021 | 15.958 | 21.45 | 6.1329 | 156.9919 | 301.3722 | 232.3707 | 295.9182 | 205.6548
2012-10-30 01:04 | 25.9138 | 322.8146 | 253.813 | 317.3147 | 227.3488 | 88.9487 | 1.5454 | 39.2671 | 60.6846 | 25.4086 | 2.4534 | 39.1506 | 96.0512 | 15.9578 | 21.4221 | 6.1321 | 156.7358 | 301.3925 | 232.3909 | 295.8926 | 205.9267
2012-10-30 01:04 | 25.9108 | 322.8192 | 253.8569 | 317.3499 | 227.5935 | 88.9452 | 1.5451 | 39.2575 | 60.6552 | 25.4084 | 2.4528 | 39.141 | 96.0052 | 15.9577 | 21.3789 | 6.1319 | 156.6603 | 301.4403 | 232.478 | 295.971 | 206.2146
2012-10-30 01:04 | 25.9049 | 322.8171 | 253.8744 | 317.3249 | 226.5138 | 88.9325 | 1.545 | 39.2566 | 60.6528 | 25.4083 | 2.4527 | 39.1403 | 96.0011 | 15.9577 | 21.4212 | 6.131 | 156.6539 | 301.3959 | 232.4532 | 295.9037 | 205.0926
2012-10-30 01:04 | 25.9071 | 322.7888 | 253.6969 | 317.3577 | 225.1194 | 88.9493 | 1.5451 | 39.259 | 60.6598 | 25.4084 | 2.4529 | 39.1425 | 96.0123 | 15.9577 | 21.4033 | 6.1322 | 156.6721 | 301.3855 | 232.2936 | 295.9544 | 203.7161
2012-10-30 01:04 | 25.9086 | 322.7979 | 253.4484 | 317.3592 | 225.7779 | 88.9452 | 1.5452 | 39.2617 | 60.6687 | 25.4082 | 2.4531 | 39.1452 | 96.0251 | 15.9577 | 21.4209 | 6.1319 | 156.6938 | 301.377 | 232.0275 | 295.9383 | 204.357
2012-10-30 01:04 | 25.9019 | 322.7912 | 253.3597 | 317.3295 | 226.68 | 88.9389 | 1.5454 | 39.2661 | 60.6815 | 25.4085 | 2.4533 | 39.1494 | 96.0465 | 15.9576 | 21.4423 | 6.1314 | 156.728 | 301.3489 | 231.9174 | 295.8872 | 205.2377
2012-10-30 01:04 | 25.8974 | 322.7982 | 253.4646 | 317.3976 | 226.8328 | 88.9481 | 1.5456 | 39.2717 | 60.6991 | 25.4084 | 2.4537 | 39.155 | 96.0738 | 15.9577 | 21.4338 | 6.1321 | 156.7729 | 301.3644 | 232.0308 | 295.9638 | 205.399
2012-10-30 01:04 | 25.8966 | 322.8164 | 253.4833 | 317.3663 | 226.5016 | 88.9452 | 1.5458 | 39.2771 | 60.7157 | 25.4084 | 2.454 | 39.1606 | 96.1011 | 15.9577 | 21.441 | 6.1319 | 156.8168 | 301.3754 | 232.0423 | 295.9253 | 205.0606
2012-10-30 01:04 | 25.8906 | 322.799 | 253.4968 | 317.3717 | 226.2674 | 88.9429 | 1.5474 | 39.3179 | 60.8406 | 25.4089 | 2.4567 | 39.2032 | 96.309 | 15.9579 | 21.4671 | 6.1317 | 157.1496 | 301.3319 | 232.0297 | 295.9046 | 204.8003
2012-10-30 01:04 | 25.8921 | 322.77 | 253.4397 | 317.3006 | 226.1942 | 88.9446 | 1.5492 | 39.3653 | 60.9847 | 25.4101 | 2.4595 | 39.2485 | 96.5302 | 15.9582 | 21.4967 | 6.1318 | 157.5149 | 301.2733 | 231.943 | 295.8039 | 204.6975
2012-10-30 01:04 | 25.8891 | 322.8013 | 253.4328 | 317.3664 | 226.7616 | 88.9683 | 1.5497 | 39.3779 | 61.0242 | 25.4099 | 2.4602 | 39.261 | 96.5918 | 15.9582 | 21.4937 | 6.1335 | 157.6159 | 301.3076 | 231.9391 | 295.8727 | 205.2679
2012-10-30 01:04 | 25.8959 | 322.8424 | 253.4748 | 317.3885 | 226.9297 | 88.9504 | 1.5501 | 39.3881 | 61.0556 | 25.41 | 2.4609 | 39.2713 | 96.6417 | 15.9583 | 21.4925 | 6.1322 | 157.6973 | 301.3499 | 231.9823 | 295.896 | 205.4372
2012-10-30 01:04 | 25.8876 | 322.8494 | 253.5173 | 317.3267 | 226.9056 | 88.9545 | 1.5504 | 39.3969 | 61.0821 | 25.4104 | 2.4614 | 39.28 | 96.6843 | 15.9583 | 21.5042 | 6.1325 | 157.7665 | 301.3452 | 232.0131 | 295.8225 | 205.4014
2012-10-30 01:04 | 25.8832 | 322.8374 | 253.5285 | 317.3452 | 226.5865 | 88.9637 | 1.5506 | 39.4005 | 61.0927 | 25.4105 | 2.4616 | 39.2836 | 96.7012 | 15.9584 | 21.5119 | 6.1332 | 157.7939 | 301.3255 | 232.0166 | 295.8333 | 205.0746
2012-10-30 01:04 | 25.8832 | 322.8526 | 253.5363 | 317.3681 | 226.9406 | 88.9816 | 1.5507 | 39.4029 | 61.1002 | 25.4105 | 2.4618 | 39.286 | 96.7137 | 15.9583 | 21.4798 | 6.1344 | 157.8139 | 301.3728 | 232.0565 | 295.8883 | 205.4608
2012-10-30 01:04 | 25.8824 | 322.8633 | 253.5628 | 317.3864 | 227.046 | 88.9706 | 1.5507 | 39.4039 | 61.1031 | 25.4106 | 2.4618 | 39.2869 | 96.7173 | 15.9585 | 21.479 | 6.1336 | 157.8204 | 301.3843 | 232.0838 | 295.9074 | 205.567
2012-10-30 01:04 | 25.8779 | 322.9045 | 253.6091 | 317.3628 | 227.1123 | 88.9747 | 1.5507 | 39.4041 | 61.1039 | 25.4105 | 2.4618 | 39.2871 | 96.7183 | 15.9585 | 21.4664 | 6.1339 | 157.8222 | 301.4381 | 232.1427 | 295.8964 | 205.6459
2012-10-30 01:04 | 25.8817 | 322.8702 | 253.6325 | 317.3857 | 227.2341 | 88.9654 | 1.5507 | 39.4035 | 61.1016 | 25.4107 | 2.4618 | 39.2866 | 96.7149 | 15.9586 | 21.4622 | 6.1333 | 157.8165 | 301.408 | 232.1703 | 295.9235 | 205.7719
2012-10-30 01:05 | 25.8899 | 322.8898 | 253.7266 | 317.3939 | 227.1715 | 88.9793 | 1.5506 | 39.402 | 61.0968 | 25.4108 | 2.4617 | 39.2849 | 96.7071 | 15.9585 | 21.4303 | 6.1342 | 157.8039 | 301.4595 | 232.2963 | 295.9636 | 205.7412
2012-10-30 01:05 | 25.8876 | 322.8913 | 253.736 | 317.4031 | 227.2242 | 88.9689 | 1.5506 | 39.4008 | 61.0933 | 25.4107 | 2.4616 | 39.2838 | 96.701 | 15.9586 | 21.4119 | 6.1335 | 157.7943 | 301.4794 | 232.3241 | 295.9912 | 205.8123
2012-10-30 01:05 | 25.8869 | 322.883 | 253.7509 | 317.3718 | 227.2235 | 88.9741 | 1.5506 | 39.4011 | 61.0939 | 25.4108 | 2.4616 | 39.2842 | 96.7032 | 15.9586 | 21.4112 | 6.1339 | 157.7971 | 301.4718 | 232.3397 | 295.9606 | 205.8123
2012-10-30 01:05 | 25.8839 | 322.8876 | 253.7792 | 317.3765 | 227.2048 | 88.992 | 1.5505 | 39.4003 | 61.0916 | 25.4107 | 2.4616 | 39.2833 | 96.6987 | 15.9586 | 21.4002 | 6.1351 | 157.7903 | 301.4874 | 232.379 | 295.9763 | 205.8046
2012-10-30 01:05 | 25.9049 | 322.8896 | 253.8158 | 317.4357 | 227.1825 | 88.9863 | 1.5505 | 39.4005 | 61.0918 | 25.411 | 2.4616 | 39.2834 | 96.6994 | 15.9586 | 21.4453 | 6.1347 | 157.7912 | 301.4443 | 232.3705 | 295.9904 | 205.7372
2012-10-30 01:05 | 25.8989 | 322.9065 | 253.8098 | 317.4449 | 227.1844 | 88.9631 | 1.5505 | 39.3992 | 61.0882 | 25.4107 | 2.4615 | 39.2821 | 96.6928 | 15.9586 | 21.3951 | 6.1331 | 157.781 | 301.5114 | 232.4147 | 296.0498 | 205.7893
2012-10-30 01:05 | 25.9019 | 322.9247 | 253.8362 | 317.4441 | 227.3487 | 88.9845 | 1.5504 | 39.3968 | 61.0809 | 25.4108 | 2.4614 | 39.2797 | 96.6817 | 15.9585 | 21.3901 | 6.1346 | 157.7626 | 301.5346 | 232.4461 | 296.054 | 205.9586
2012-10-30 01:05 | 25.9011 | 322.9201 | 253.8823 | 317.4624 | 227.3715 | 88.9828 | 1.5504 | 39.3967 | 61.0802 | 25.4108 | 2.4614 | 39.2796 | 96.6809 | 15.9586 | 21.4375 | 6.1345 | 157.7611 | 301.4826 | 232.4448 | 296.0249 | 205.934
2012-10-30 01:05 | 25.9071 | 322.9299 | 253.8883 | 317.4646 | 227.3145 | 88.9834 | 1.5503 | 39.3957 | 61.0767 | 25.411 | 2.4613 | 39.2788 | 96.6774 | 15.9585 | 21.3913 | 6.1345 | 157.7541 | 301.5386 | 232.497 | 296.0733 | 205.9232
2012-10-30 01:05 | 25.9093 | 322.9512 | 253.9022 | 317.4973 | 227.3403 | 88.9984 | 1.5503 | 39.3941 | 61.0714 | 25.4112 | 2.4612 | 39.277 | 96.6683 | 15.9585 | 21.3975 | 6.1355 | 157.7397 | 301.5537 | 232.5047 | 296.0998 | 205.9428
2012-10-30 01:05 | 25.9138 | 322.9404 | 253.9185 | 317.5438 | 227.3134 | 88.9857 | 1.5502 | 39.3926 | 61.0675 | 25.4109 | 2.4611 | 39.2757 | 96.6625 | 15.9584 | 21.4141 | 6.1347 | 157.7299 | 301.5263 | 232.5044 | 296.1297 | 205.8993
2012-10-30 01:05 | 25.9176 | 322.9976 | 253.9769 | 317.4865 | 227.2149 | 88.9955 | 1.5502 | 39.3919 | 61.0651 | 25.4109 | 2.4611 | 39.2748 | 96.6583 | 15.9584 | 21.4379 | 6.1353 | 157.7234 | 301.5597 | 232.539 | 296.0486 | 205.777
2012-10-30 01:05 | 25.9131 | 322.9741 | 253.9373 | 317.5279 | 227.2969 | 88.9961 | 1.5501 | 39.3893 | 61.0571 | 25.411 | 2.4609 | 39.2723 | 96.6456 | 15.9584 | 21.3853 | 6.1354 | 157.7026 | 301.5888 | 232.552 | 296.1426 | 205.9116
2012-10-30 01:05 | 25.9153 | 322.9953 | 253.9785 | 317.5759 | 227.2323 | 88.9938 | 1.55 | 39.3872 | 61.0508 | 25.4109 | 2.4608 | 39.2704 | 96.6364 | 15.9584 | 21.3794 | 6.1352 | 157.6872 | 301.6159 | 232.5991 | 296.1965 | 205.8529
2012-10-30 01:05 | 25.9206 | 323.0006 | 253.9995 | 317.5354 | 227.2061 | 88.9949 | 1.5499 | 39.386 | 61.0462 | 25.4112 | 2.4607 | 39.269 | 96.6303 | 15.9583 | 21.3767 | 6.1353 | 157.6765 | 301.6239 | 232.6228 | 296.1587 | 205.8294
2012-10-30 01:05 | 25.9243 | 323.0272 | 253.9797 | 317.5734 | 227.2924 | 89.0013 | 1.5499 | 39.3838 | 61.04 | 25.411 | 2.4606 | 39.2669 | 96.6195 | 15.9583 | 21.3965 | 6.1357 | 157.6595 | 301.6307 | 232.5832 | 296.1769 | 205.8959
2012-10-30 01:05 | 25.9273 | 323.0569 | 253.9905 | 317.5879 | 227.0594 | 89.0076 | 1.5498 | 39.3814 | 61.0321 | 25.4111 | 2.4604 | 39.2644 | 96.6081 | 15.9582 | 21.3874 | 6.1362 | 157.6402 | 301.6695 | 232.6031 | 296.2005 | 205.672
2012-10-30 01:05 | 25.9281 | 323.0348 | 253.9952 | 317.6231 | 227.0641 | 89.0013 | 1.5498 | 39.3811 | 61.0313 | 25.4111 | 2.4604 | 39.2642 | 96.6072 | 15.9582 | 21.4163 | 6.1357 | 157.6385 | 301.6185 | 232.5789 | 296.2068 | 205.6478
2012-10-30 01:05 | 25.9288 | 323.0203 | 253.9725 | 317.639 | 227.0806 | 88.9972 | 1.5497 | 39.3793 | 61.0262 | 25.4108 | 2.4603 | 39.2623 | 96.598 | 15.9582 | 21.4531 | 6.1355 | 157.6242 | 301.5672 | 232.5194 | 296.1859 | 205.6275
2012-10-30 01:05 | 25.9438 | 323.0658 | 253.9563 | 317.5891 | 226.8713 | 89.0019 | 1.5497 | 39.3782 | 61.0226 | 25.411 | 2.4603 | 39.2612 | 96.5933 | 15.9581 | 21.42 | 6.1358 | 157.6159 | 301.6458 | 232.5363 | 296.1691 | 205.4513
2012-10-30 01:05 | 25.9318 | 323.0347 | 253.9052 | 317.6459 | 226.8672 | 88.9938 | 1.5496 | 39.3759 | 61.016 | 25.4107 | 2.4601 | 39.259 | 96.5825 | 15.9581 | 21.4561 | 6.1352 | 157.5984 | 301.5786 | 232.4491 | 296.1898 | 205.4111
2012-10-30 01:06 | 25.9363 | 323.0583 | 253.9019 | 317.6656 | 227.1431 | 89.0059 | 1.5471 | 39.3128 | 60.8225 | 25.4099 | 2.4562 | 39.1957 | 96.2738 | 15.9576 | 21.4486 | 6.1361 | 157.0963 | 301.6097 | 232.4533 | 296.217 | 205.6945
2012-10-30 01:06 | 25.9385 | 323.0452 | 253.8572 | 317.6984 | 226.9447 | 88.992 | 1.5449 | 39.255 | 60.6461 | 25.409 | 2.4527 | 39.1385 | 95.9955 | 15.9572 | 21.4668 | 6.1351 | 156.6415 | 301.5784 | 232.3904 | 296.2316 | 205.4779
2012-10-30 01:06 | 25.9325 | 323.0011 | 253.82 | 317.6733 | 226.8876 | 88.9822 | 1.545 | 39.2575 | 60.654 | 25.4089 | 2.4529 | 39.141 | 96.0075 | 15.9573 | 21.4368 | 6.1344 | 156.6616 | 301.5643 | 232.3832 | 296.2365 | 205.4508
2012-10-30 01:06 | 25.9363 | 322.9973 | 253.8042 | 317.6618 | 226.502 | 88.9828 | 1.5451 | 39.2602 | 60.6623 | 25.4089 | 2.453 | 39.1437 | 96.0214 | 15.9572 | 21.4687 | 6.1345 | 156.6837 | 301.5286 | 232.3355 | 296.1931 | 205.0333
2012-10-30 01:06 | 25.9318 | 322.9889 | 253.5966 | 317.6726 | 226.525 | 88.9764 | 1.5454 | 39.2661 | 60.6804 | 25.409 | 2.4534 | 39.1493 | 96.0481 | 15.9573 | 21.4762 | 6.134 | 156.7285 | 301.5127 | 232.1204 | 296.1964 | 205.0488
2012-10-30 01:06 | 25.9303 | 322.9989 | 253.5404 | 317.6329 | 227.3141 | 88.9834 | 1.5456 | 39.2718 | 60.6978 | 25.409 | 2.4537 | 39.1552 | 96.0768 | 15.9573 | 21.4426 | 6.1345 | 156.7747 | 301.5563 | 232.0978 | 296.1903 | 205.8715
2012-10-30 01:06 | 25.9206 | 323.0235 | 253.6753 | 317.6003 | 227.2651 | 88.9776 | 1.5458 | 39.2775 | 60.7154 | 25.409 | 2.4541 | 39.1605 | 96.1033 | 15.9573 | 21.453 | 6.1341 | 156.8188 | 301.5705 | 232.2223 | 296.1473 | 205.8121
2012-10-30 01:06 | 25.9206 | 323.0311 | 253.73 | 317.6614 | 226.6633 | 88.9828 | 1.5484 | 39.3441 | 60.9193 | 25.41 | 2.4582 | 39.2267 | 96.4267 | 15.9576 | 21.461 | 6.1345 | 157.346 | 301.5701 | 232.269 | 296.2004 | 205.2023
2012-10-30 01:06 | 25.9131 | 323.0427 | 253.6209 | 317.6042 | 226.526 | 88.9955 | 1.5494 | 39.3707 | 61.0008 | 25.4104 | 2.4599 | 39.2538 | 96.5592 | 15.9577 | 21.4575 | 6.1353 | 157.5599 | 301.5852 | 232.1634 | 296.1467 | 205.0685
2012-10-30 01:06 | 25.9138 | 323.0625 | 253.6021 | 317.6851 | 226.4755 | 89.0024 | 1.5498 | 39.382 | 61.0343 | 25.411 | 2.4606 | 39.2652 | 96.6145 | 15.9578 | 21.4622 | 6.1358 | 157.6488 | 301.6003 | 232.1399 | 296.2229 | 205.0133
2012-10-30 01:06 | 25.9086 | 323.0153 | 253.5617 | 317.6227 | 226.0258 | 89.0042 | 1.5502 | 39.3915 | 61.064 | 25.4109 | 2.4611 | 39.2743 | 96.659 | 15.9578 | 21.4731 | 6.1359 | 157.7231 | 301.5422 | 232.0886 | 296.1496 | 204.5527
2012-10-30 01:06 | 25.9064 | 323.0398 | 253.4853 | 317.6357 | 226.5193 | 89.0059 | 1.5504 | 39.3974 | 61.0824 | 25.4109 | 2.4615 | 39.2802 | 96.6882 | 15.9578 | 21.4468 | 6.1361 | 157.7706 | 301.593 | 232.0385 | 296.1889 | 205.0725
2012-10-30 01:06 | 25.9071 | 323.0558 | 253.4469 | 317.6097 | 226.7638 | 89.0024 | 1.5506 | 39.4034 | 61.1004 | 25.411 | 2.4619 | 39.2862 | 96.7171 | 15.958 | 21.4956 | 6.1358 | 157.8175 | 301.5602 | 231.9513 | 296.1141 | 205.2682
2012-10-30 01:06 | 25.9056 | 323.0619 | 253.5235 | 317.6082 | 226.8725 | 88.9984 | 1.5507 | 39.4045 | 61.1033 | 25.4113 | 2.4619 | 39.2875 | 96.7224 | 15.9581 | 21.4701 | 6.1355 | 157.8257 | 301.5918 | 232.0534 | 296.1381 | 205.4024
2012-10-30 01:06 | 25.8951 | 323.0629 | 253.5443 | 317.5901 | 226.9485 | 89.0232 | 1.5507 | 39.4048 | 61.1044 | 25.4113 | 2.4619 | 39.2877 | 96.7239 | 15.958 | 21.4756 | 6.1373 | 157.8283 | 301.5873 | 232.0687 | 296.1145 | 205.4729
2012-10-30 01:06 | 25.8951 | 323.0323 | 253.5873 | 317.6588 | 226.9289 | 89.0134 | 1.5507 | 39.4045 | 61.1038 | 25.4111 | 2.4619 | 39.2875 | 96.7223 | 15.9581 | 21.4555 | 6.1366 | 157.8261 | 301.5768 | 232.1318 | 296.2033 | 205.4734
2012-10-30 01:06 | 25.8966 | 323.0453 | 253.6474 | 317.645 | 227.0484 | 89.003 | 1.5507 | 39.4047 | 61.1041 | 25.4112 | 2.4619 | 39.2876 | 96.7228 | 15.9581 | 21.4731 | 6.1359 | 157.8269 | 301.5722 | 232.1743 | 296.1719 | 205.5753
2012-10-30 01:06 | 25.8944 | 323.0545 | 253.653 | 317.5779 | 227.0108 | 89.0117 | 1.5507 | 39.4042 | 61.1027 | 25.4112 | 2.4619 | 39.2872 | 96.7216 | 15.958 | 21.4348 | 6.1365 | 157.8242 | 301.6197 | 232.2182 | 296.1431 | 205.576
2012-10-30 01:06 | 25.8884 | 323.0371 | 253.6704 | 317.6063 | 227.0677 | 89.0128 | 1.5506 | 39.4039 | 61.1013 | 25.4113 | 2.4619 | 39.2868 | 96.7199 | 15.9579 | 21.4488 | 6.1365 | 157.8213 | 301.5883 | 232.2216 | 296.1575 | 205.6189
2012-10-30 01:07 | 25.8959 | 323.098 | 253.7287 | 317.6443 | 227.3466 | 89.0099 | 1.5506 | 39.4022 | 61.0966 | 25.4111 | 2.4618 | 39.2853 | 96.7128 | 15.9579 | 21.4764 | 6.1363 | 157.8095 | 301.6216 | 232.2523 | 296.1679 | 205.8702
2012-10-30 01:07 | 25.8832 | 323.051 | 253.8293 | 317.6278 | 227.6407 | 89.0175 | 1.5506 | 39.4017 | 61.095 | 25.4112 | 2.4618 | 39.2846 | 96.7094 | 15.9579 | 21.4717 | 6.1369 | 157.8045 | 301.5793 | 232.3576 | 296.1561 | 206.169
2012-10-30 01:07 | 25.8847 | 323.0181 | 253.8503 | 317.6102 | 227.477 | 89.0238 | 1.5505 | 39.4005 | 61.0908 | 25.4113 | 2.4617 | 39.2836 | 96.7042 | 15.958 | 21.4652 | 6.1373 | 157.795 | 301.5529 | 232.3851 | 296.145 | 206.0118
2012-10-30 01:07 | 25.8787 | 323.0693 | 253.9302 | 317.6195 | 227.3884 | 89.0319 | 1.5505 | 39.4003 | 61.0901 | 25.4113 | 2.4617 | 39.2832 | 96.7027 | 15.9579 | 21.4231 | 6.1379 | 157.7928 | 301.6462 | 232.5071 | 296.1964 | 205.9653
2012-10-30 01:07 | 25.8832 | 323.0243 | 253.9425 | 317.624 | 227.385 | 89.0244 | 1.5505 | 39.3994 | 61.0876 | 25.4112 | 2.4616 | 39.2823 | 96.6991 | 15.9578 | 21.4477 | 6.1373 | 157.7867 | 301.5766 | 232.4948 | 296.1763 | 205.9373
2012-10-30 01:07 | 25.8854 | 323.0684 | 253.9799 | 317.6224 | 227.3125 | 89.0232 | 1.5504 | 39.3975 | 61.0818 | 25.4113 | 2.4615 | 39.2805 | 96.689 | 15.9579 | 21.5341 | 6.1373 | 157.7708 | 301.5343 | 232.4458 | 296.0883 | 205.7784
2012-10-30 01:07 | 25.8727 | 323.071 | 253.975 | 317.6173 | 227.1346 | 89.0215 | 1.5503 | 39.3963 | 61.0779 | 25.4113 | 2.4614 | 39.2792 | 96.6832 | 15.9579 | 21.4492 | 6.1371 | 157.7611 | 301.6218 | 232.5258 | 296.1681 | 205.6854
2012-10-30 01:07 | 25.8794 | 323.0929 | 253.9544 | 317.6851 | 227.2239 | 89.0169 | 1.5503 | 39.3945 | 61.0713 | 25.4116 | 2.4613 | 39.2772 | 96.6729 | 15.9579 | 21.4238 | 6.1368 | 157.7442 | 301.6691 | 232.5306 | 296.2613 | 205.8001
2012-10-30 01:07 | 25.8817 | 323.0456 | 253.941 | 317.6454 | 227.2026 | 89.0238 | 1.5502 | 39.3924 | 61.0656 | 25.4113 | 2.4612 | 39.2753 | 96.664 | 15.9579 | 21.4462 | 6.1373 | 157.7297 | 301.5994 | 232.4948 | 296.1992 | 205.7564
2012-10-30 01:07 | 25.8914 | 323.063 | 253.9117 | 317.6475 | 227.0392 | 89.0331 | 1.5501 | 39.3902 | 61.0591 | 25.4112 | 2.4611 | 39.2732 | 96.6538 | 15.9578 | 21.4679 | 6.1379 | 157.7129 | 301.5951 | 232.4438 | 296.1796 | 205.5713
2012-10-30 01:07 | 25.8794 | 323.0739 | 253.9114 | 317.6584 | 227.0154 | 89.0446 | 1.5501 | 39.3895 | 61.0573 | 25.4111 | 2.461 | 39.2725 | 96.6509 | 15.9577 | 21.4238 | 6.1387 | 157.7082 | 301.6501 | 232.4876 | 296.2346 | 205.5916
2012-10-30 01:07 | 25.8862 | 323.0959 | 253.9104 | 317.6461 | 226.861 | 89.0412 | 1.55 | 39.3877 | 61.0515 | 25.4112 | 2.4609 | 39.2706 | 96.6424 | 15.9576 | 21.4386 | 6.1385 | 157.6939 | 301.6573 | 232.4718 | 296.2075 | 205.4224
2012-10-30 01:07 | 25.8742 | 323.0839 | 253.8164 | 317.6685 | 226.6877 | 89.0342 | 1.55 | 39.3874 | 61.0509 | 25.4111 | 2.4609 | 39.2704 | 96.6409 | 15.9577 | 21.4106 | 6.138 | 157.6918 | 301.6733 | 232.4058 | 296.2579 | 205.2771
2012-10-30 01:07 | 25.8824 | 323.0807 | 253.7894 | 317.6767 | 226.8965 | 89.0383 | 1.5499 | 39.386 | 61.0462 | 25.4112 | 2.4608 | 39.269 | 96.6343 | 15.9576 | 21.4348 | 6.1383 | 157.6805 | 301.6459 | 232.3546 | 296.2419 | 205.4617
2012-10-30 01:07 | 25.8809 | 323.0563 | 253.8152 | 317.6637 | 227.0366 | 89.0417 | 1.5499 | 39.3852 | 61.0435 | 25.4112 | 2.4608 | 39.2683 | 96.6302 | 15.9577 | 21.4413 | 6.1385 | 157.6738 | 301.615 | 232.3739 | 296.2224 | 205.5953
2012-10-30 01:07 | 25.8899 | 323.1339 | 253.8282 | 317.6765 | 227.1754 | 89.0377 | 1.5498 | 39.3832 | 61.0374 | 25.4113 | 2.4606 | 39.2661 | 96.62 | 15.9576 | 21.4142 | 6.1383 | 157.6574 | 301.7197 | 232.414 | 296.2623 | 205.7612
2012-10-30 01:07 | 25.8847 | 323.1173 | 253.8503 | 317.6904 | 227.2803 | 89.029 | 1.5497 | 39.3803 | 61.0285 | 25.4112 | 2.4605 | 39.2633 | 96.6061 | 15.9576 | 21.3809 | 6.1377 | 157.6346 | 301.7364 | 232.4694 | 296.3095 | 205.8994
2012-10-30 01:07 | 25.8876 | 323.0897 | 253.8727 | 317.7086 | 227.3108 | 89.04 | 1.5497 | 39.3793 | 61.0249 | 25.4114 | 2.4604 | 39.2623 | 96.6012 | 15.9576 | 21.4159 | 6.1384 | 157.6262 | 301.6738 | 232.4568 | 296.2927 | 205.8949
2012-10-30 01:07 | 25.8929 | 323.0912 | 253.8897 | 317.6833 | 227.2059 | 89.0452 | 1.5497 | 39.3795 | 61.0258 | 25.4113 | 2.4604 | 39.2627 | 96.6035 | 15.9576 | 21.4212 | 6.1388 | 157.6293 | 301.67 | 232.4685 | 296.2621 | 205.7847
2012-10-30 01:07 | 25.8869 | 323.0966 | 253.9345 | 317.6621 | 227.0701 | 89.044 | 1.5496 | 39.3774 | 61.0192 | 25.4113 | 2.4603 | 39.2605 | 96.5932 | 15.9575 | 21.4112 | 6.1387 | 157.6123 | 301.6854 | 232.5233 | 296.2509 | 205.6589
2012-10-30 01:07 | 25.8869 | 323.1271 | 253.8955 | 317.6773 | 227.1134 | 89.0365 | 1.5495 | 39.3756 | 61.0134 | 25.4114 | 2.4602 | 39.2588 | 96.5843 | 15.9576 | 21.4473 | 6.1382 | 157.5977 | 301.6798 | 232.4482 | 296.23 | 205.6661
2012-10-30 01:07 | 25.8944 | 323.1003 | 253.9186 | 317.7001 | 226.881 | 89.0458 | 1.5495 | 39.3749 | 61.0115 | 25.4113 | 2.4601 | 39.2579 | 96.5798 | 15.9576 | 21.4508 | 6.1388 | 157.5913 | 301.6495 | 232.4678 | 296.2493 | 205.4302
2012-10-30 01:07 | 25.8951 | 323.1353 | 253.8412 | 317.7161 | 226.858 | 89.0469 | 1.5494 | 39.3739 | 61.0078 | 25.4115 | 2.4601 | 39.257 | 96.5759 | 15.9575 | 21.4395 | 6.1389 | 157.5838 | 301.6958 | 232.4017 | 296.2766 | 205.4185
2012-10-30 01:08 | 25.9034 | 323.1017 | 253.8612 | 317.7015 | 226.8939 | 89.0336 | 1.5494 | 39.3729 | 61.0048 | 25.4115 | 2.46 | 39.2559 | 96.57 | 15.9576 | 21.4398 | 6.138 | 157.5748 | 301.6619 | 232.4214 | 296.2617 | 205.4541
2012-10-30 01:08 | 25.8966 | 323.1254 | 253.7919 | 317.7138 | 227.131 | 89.0469 | 1.5477 | 39.3275 | 60.8657 | 25.4109 | 2.4571 | 39.2098 | 96.3442 | 15.9574 | 21.4129 | 6.1389 | 157.2098 | 301.7125 | 232.379 | 296.3009 | 205.7181
2012-10-30 01:08 | 25.8966 | 323.133 | 253.8583 | 317.6909 | 227.2608 | 89.0331 | 1.5455 | 39.2711 | 60.6936 | 25.4099 | 2.4537 | 39.1544 | 96.0746 | 15.957 | 21.457 | 6.1379 | 156.7681 | 301.676 | 232.4013 | 296.2339 | 205.8038
2012-10-30 01:08 | 25.8996 | 323.1055 | 253.8613 | 317.6824 | 227.012 | 89.0284 | 1.545 | 39.2586 | 60.6553 | 25.4098 | 2.453 | 39.1416 | 96.0127 | 15.9569 | 21.436 | 6.1376 | 156.668 | 301.6695 | 232.4253 | 296.2464 | 205.576
2012-10-30 01:08 | 25.8981 | 323.0849 | 253.7856 | 317.7076 | 226.1923 | 89.0302 | 1.5447 | 39.2492 | 60.6271 | 25.4095 | 2.4524 | 39.1328 | 95.9697 | 15.9569 | 21.4024 | 6.1377 | 156.5967 | 301.6825 | 232.3832 | 296.3052 | 204.7899
2012-10-30 01:08 | 25.9026 | 323.1161 | 253.6182 | 317.7007 | 224.93 | 89.0267 | 1.5447 | 39.2506 | 60.6311 | 25.4096 | 2.4525 | 39.134 | 95.9757 | 15.9569 | 21.4189 | 6.1375 | 156.6067 | 301.6972 | 232.1993 | 296.2818 | 203.5111
2012-10-30 01:08 | 25.8989 | 323.0934 | 253.302 | 317.7084 | 224.9735 | 89.0296 | 1.5448 | 39.2533 | 60.6402 | 25.4093 | 2.4527 | 39.1366 | 95.9888 | 15.9568 | 21.4393 | 6.1377 | 156.6289 | 301.6541 | 231.8627 | 296.2691 | 203.5342
2012-10-30 01:08 | 25.9011 | 323.1032 | 253.1557 | 317.6495 | 226.1638 | 89.0152 | 1.5451 | 39.2589 | 60.6573 | 25.4093 | 2.453 | 39.1422 | 96.0152 | 15.957 | 21.4495 | 6.1367 | 156.6726 | 301.6537 | 231.7062 | 296.2 | 204.7143
2012-10-30 01:08 | 25.8914 | 323.0706 | 253.2554 | 317.6971 | 226.225 | 89.0261 | 1.5453 | 39.264 | 60.6727 | 25.4095 | 2.4533 | 39.1473 | 96.041 | 15.9568 | 21.4318 | 6.1375 | 156.7137 | 301.6388 | 231.8236 | 296.2653 | 204.7932
2012-10-30 01:08 | 25.8809 | 323.0525 | 253.1824 | 317.637 | 226.002 | 89.0209 | 1.5455 | 39.2708 | 60.6932 | 25.4096 | 2.4538 | 39.1542 | 96.075 | 15.9568 | 21.4614 | 6.1371 | 156.7681 | 301.5911 | 231.721 | 296.1756 | 204.5406
2012-10-30 01:08 | 25.8899 | 323.0805 | 253.1406 | 317.6651 | 225.7986 | 89.0296 | 1.5458 | 39.2778 | 60.7155 | 25.4094 | 2.4542 | 39.1612 | 96.1095 | 15.9568 | 21.4824 | 6.1377 | 156.825 | 301.5981 | 231.6582 | 296.1827 | 204.3162
2012-10-30 01:08 | 25.8779 | 323.0647 | 253.0583 | 317.5881 | 226.1013 | 89.0261 | 1.548 | 39.3339 | 60.8873 | 25.4102 | 2.4578 | 39.2195 | 96.3947 | 15.957 | 21.4865 | 6.1375 | 157.282 | 301.5782 | 231.5718 | 296.1016 | 204.6148
2012-10-30 01:08 | 25.8839 | 323.0364 | 253.0956 | 317.6323 | 225.9539 | 89.0429 | 1.5492 | 39.3663 | 60.9856 | 25.411 | 2.4597 | 39.2496 | 96.5413 | 15.9572 | 21.4724 | 6.1386 | 157.5269 | 301.564 | 231.6232 | 296.1599 | 204.4815
2012-10-30 01:08 | 25.8757 | 323.0129 | 253.0874 | 317.605 | 225.6546 | 89.0157 | 1.5496 | 39.3762 | 61.0157 | 25.4112 | 2.4603 | 39.259 | 96.5878 | 15.9572 | 21.4803 | 6.1367 | 157.6034 | 301.5326 | 231.6071 | 296.1247 | 204.1743
2012-10-30 01:08 | 25.8719 | 323.0206 | 253.021 | 317.5592 | 225.5878 | 89.0296 | 1.55 | 39.3876 | 61.0508 | 25.4114 | 2.461 | 39.2706 | 96.643 | 15.9575 | 21.4765 | 6.1377 | 157.6938 | 301.5441 | 231.5445 | 296.0827 | 204.1113
2012-10-30 01:08 | 25.8719 | 323.0397 | 253.0015 | 317.586 | 225.9616 | 89.0429 | 1.5503 | 39.3941 | 61.0711 | 25.4113 | 2.4614 | 39.2772 | 96.6758 | 15.9574 | 21.4805 | 6.1386 | 157.7469 | 301.5592 | 231.521 | 296.1055 | 204.4811
2012-10-30 01:08 | 25.8682 | 323.0665 | 253.0486 | 317.5708 | 226.367 | 89.0325 | 1.5504 | 39.3987 | 61.0846 | 25.4116 | 2.4617 | 39.2817 | 96.6977 | 15.9575 | 21.4567 | 6.1379 | 157.7823 | 301.6098 | 231.5919 | 296.1141 | 204.9103
2012-10-30 01:08 | 25.8689 | 323.0176 | 253.1157 | 317.5753 | 226.4818 | 89.014 | 1.5505 | 39.401 | 61.0921 | 25.4115 | 2.4618 | 39.2839 | 96.7085 | 15.9574 | 21.4895 | 6.1366 | 157.8006 | 301.5281 | 231.6262 | 296.0858 | 204.9923
2012-10-30 01:08 | 25.8659 | 323.0375 | 253.1557 | 317.5647 | 226.6007 | 89.0336 | 1.5506 | 39.4031 | 61.0983 | 25.4115 | 2.4619 | 39.286 | 96.7186 | 15.9575 | 21.4905 | 6.138 | 157.8169 | 301.547 | 231.6652 | 296.0742 | 205.1102
2012-10-30 01:08 | 25.8644 | 323.0512 | 253.2167 | 317.5288 | 226.6307 | 89.0406 | 1.5506 | 39.4034 | 61.0986 | 25.4118 | 2.4619 | 39.2865 | 96.7213 | 15.9575 | 21.5211 | 6.1385 | 157.8199 | 301.5301 | 231.6956 | 296.0077 | 205.1096
2012-10-30 01:08 | 25.863 | 323.0117 | 253.3208 | 317.516 | 226.7551 | 89.0336 | 1.5506 | 39.4036 | 61.0997 | 25.4117 | 2.4619 | 39.2865 | 96.7213 | 15.9575 | 21.5278 | 6.138 | 157.821 | 301.4839 | 231.793 | 295.9882 | 205.2273
2012-10-30 01:08 | 25.8547 | 322.9614 | 253.3125 | 317.5306 | 226.7862 | 89.0492 | 1.5506 | 39.4047 | 61.1026 | 25.4118 | 2.462 | 39.2874 | 96.7258 | 15.9575 | 21.5074 | 6.1391 | 157.8284 | 301.454 | 231.8051 | 296.0232 | 205.2788
2012-10-30 01:08 | 25.86 | 322.9591 | 253.3686 | 317.5473 | 226.7915 | 89.0602 | 1.5506 | 39.4024 | 61.0959 | 25.4117 | 2.4619 | 39.2855 | 96.7158 | 15.9576 | 21.4606 | 6.1398 | 157.8117 | 301.4985 | 231.908 | 296.0867 | 205.3309
2012-10-30 01:08 | 25.8562 | 322.9782 | 253.3843 | 317.5283 | 226.7051 | 89.051 | 1.5506 | 39.4034 | 61.0987 | 25.4118 | 2.4619 | 39.2863 | 96.7201 | 15.9575 | 21.4528 | 6.1392 | 157.8188 | 301.5254 | 231.9315 | 296.0755 | 205.2523
2012-10-30 01:09 | 25.8592 | 323.0003 | 253.4693 | 317.5236 | 226.7631 | 89.0481 | 1.5505 | 39.402 | 61.094 | 25.4119 | 2.4618 | 39.2848 | 96.7131 | 15.9575 | 21.4919 | 6.139 | 157.8071 | 301.5084 | 231.9774 | 296.0317 | 205.2712
2012-10-30 01:09 | 25.86 | 322.9896 | 253.4584 | 317.5244 | 226.8229 | 89.0487 | 1.5505 | 39.4013 | 61.0927 | 25.4116 | 2.4618 | 39.2842 | 96.7102 | 15.9575 | 21.5087 | 6.139 | 157.8029 | 301.4809 | 231.9497 | 296.0157 | 205.3142
2012-10-30 01:09 | 25.863 | 322.9697 | 253.477 | 317.5274 | 226.885 | 89.0504 | 1.5505 | 39.4006 | 61.09 | 25.4119 | 2.4618 | 39.2837 | 96.7071 | 15.9575 | 21.4555 | 6.1391 | 157.7971 | 301.5142 | 232.0215 | 296.0719 | 205.4295
2012-10-30 01:09 | 25.86 | 322.9744 | 253.4858 | 317.5053 | 226.6538 | 89.0446 | 1.5505 | 39.4002 | 61.0888 | 25.4118 | 2.4618 | 39.2833 | 96.7058 | 15.9574 | 21.4766 | 6.1387 | 157.7946 | 301.4978 | 232.0092 | 296.0287 | 205.1772
2012-10-30 01:09 | 25.8615 | 322.9987 | 253.5224 | 317.5297 | 226.8402 | 89.0649 | 1.5504 | 39.3982 | 61.0828 | 25.4118 | 2.4616 | 39.2814 | 96.6961 | 15.9575 | 21.442 | 6.1401 | 157.7788 | 301.5567 | 232.0804 | 296.0877 | 205.3982
2012-10-30 01:09 | 25.8615 | 322.9911 | 253.5654 | 317.5106 | 226.671 | 89.0724 | 1.5503 | 39.3968 | 61.0783 | 25.4118 | 2.4615 | 39.2798 | 96.6889 | 15.9574 | 21.426 | 6.1407 | 157.7672 | 301.5651 | 232.1394 | 296.0846 | 205.245
2012-10-30 01:09 | 25.86 | 322.9782 | 253.5522 | 317.5206 | 226.4611 | 89.0729 | 1.5503 | 39.3955 | 61.0746 | 25.4116 | 2.4615 | 39.2784 | 96.6829 | 15.9572 | 21.4084 | 6.1407 | 157.7574 | 301.5698 | 232.1438 | 296.1122 | 205.0527
2012-10-30 01:09 | 25.8622 | 323.03 | 253.5075 | 317.5266 | 226.5852 | 89.0706 | 1.5503 | 39.3948 | 61.0722 | 25.4118 | 2.4614 | 39.2778 | 96.6799 | 15.9572 | 21.4307 | 6.1405 | 157.752 | 301.5993 | 232.0768 | 296.0959 | 205.1545
2012-10-30 01:09 | 25.8547 | 323.0263 | 253.4805 | 317.5306 | 226.6918 | 89.0602 | 1.5502 | 39.3931 | 61.067 | 25.4117 | 2.4613 | 39.2761 | 96.6712 | 15.9573 | 21.4312 | 6.1398 | 157.7382 | 301.5951 | 232.0493 | 296.0994 | 205.2606
2012-10-30 01:09 | 25.8592 | 322.9812 | 253.4967 | 317.5465 | 226.7238 | 89.0747 | 1.5502 | 39.3925 | 61.0656 | 25.4115 | 2.4613 | 39.2754 | 96.6687 | 15.9572 | 21.4277 | 6.1408 | 157.7343 | 301.5535 | 232.069 | 296.1188 | 205.2961
2012-10-30 01:09 | 25.8637 | 322.9895 | 253.5051 | 317.5205 | 226.7676 | 89.0747 | 1.5502 | 39.3921 | 61.0639 | 25.4117 | 2.4613 | 39.2753 | 96.6678 | 15.9572 | 21.4121 | 6.1408 | 157.7317 | 301.5774 | 232.093 | 296.1084 | 205.3555
2012-10-30 01:09 | 25.8652 | 323.0101 | 253.53 | 317.5449 | 226.6079 | 89.0689 | 1.5501 | 39.3907 | 61.059 | 25.4119 | 2.4612 | 39.2736 | 96.6592 | 15.9573 | 21.4216 | 6.1404 | 157.7183 | 301.5885 | 232.1084 | 296.1233 | 205.1863
2012-10-30 01:09 | 25.8607 | 322.9712 | 253.5333 | 317.5518 | 226.7332 | 89.0614 | 1.5501 | 39.3899 | 61.0566 | 25.4118 | 2.4611 | 39.273 | 96.6567 | 15.9572 | 21.3971 | 6.1399 | 157.7134 | 301.5741 | 232.1362 | 296.1547 | 205.3361
2012-10-30 01:09 | 25.8615 | 322.972 | 253.5341 | 317.5374 | 226.8677 | 89.0689 | 1.55 | 39.388 | 61.0512 | 25.4117 | 2.461 | 39.2709 | 96.6472 | 15.9571 | 21.3738 | 6.1404 | 157.6984 | 301.5982 | 232.1603 | 296.1636 | 205.4939
2012-10-30 01:09 | 25.8532 | 322.9599 | 253.5922 | 317.552 | 226.8358 | 89.0776 | 1.5499 | 39.3863 | 61.0463 | 25.4115 | 2.4609 | 39.2694 | 96.6396 | 15.9571 | 21.4056 | 6.141 | 157.6859 | 301.5543 | 232.1866 | 296.1464 | 205.4302
2012-10-30 01:09 | 25.8659 | 323.007 | 253.5698 | 317.5723 | 226.8407 | 89.0666 | 1.5499 | 39.3856 | 61.0442 | 25.4115 | 2.4609 | 39.2685 | 96.6348 | 15.9571 | 21.4103 | 6.1403 | 157.679 | 301.5967 | 232.1595 | 296.162 | 205.4304
2012-10-30 01:09 | 25.8644 | 322.9902 | 253.5956 | 317.5479 | 226.5953 | 89.066 | 1.5498 | 39.3841 | 61.0391 | 25.4117 | 2.4608 | 39.2672 | 96.6288 | 15.9571 | 21.4088 | 6.1402 | 157.6679 | 301.5814 | 232.1868 | 296.1391 | 205.1865
2012-10-30 01:09 | 25.8615 | 322.9911 | 253.5888 | 317.5221 | 226.6828 | 89.0845 | 1.5497 | 39.3815 | 61.0314 | 25.4115 | 2.4607 | 39.2645 | 96.6166 | 15.9569 | 21.3858 | 6.1415 | 157.648 | 301.6053 | 232.203 | 296.1363 | 205.297
2012-10-30 01:09 | 25.8592 | 323.0155 | 253.5514 | 317.5809 | 226.5665 | 89.077 | 1.5497 | 39.3807 | 61.0287 | 25.4116 | 2.4606 | 39.2637 | 96.6121 | 15.957 | 21.3996 | 6.141 | 157.6408 | 301.6159 | 232.1518 | 296.1813 | 205.1669
2012-10-30 01:09 | 25.8615 | 323.0369 | 253.5654 | 317.5641 | 226.9346 | 89.0747 | 1.5497 | 39.3797 | 61.026 | 25.4115 | 2.4606 | 39.2629 | 96.6086 | 15.9569 | 21.3898 | 6.1408 | 157.6346 | 301.6471 | 232.1756 | 296.1743 | 205.5448
2012-10-30 01:09 | 25.86 | 323.0583 | 253.6342 | 317.5702 | 227.0708 | 89.0764 | 1.5496 | 39.3783 | 61.0212 | 25.4116 | 2.4605 | 39.2613 | 96.6008 | 15.9569 | 21.3964 | 6.1409 | 157.6221 | 301.6619 | 232.2378 | 296.1738 | 205.6744
2012-10-30 01:09 | 25.8637 | 323.0467 | 253.634 | 317.5587 | 226.9329 | 89.0805 | 1.5496 | 39.3769 | 61.018 | 25.4111 | 2.4604 | 39.2599 | 96.5941 | 15.9568 | 21.4161 | 6.1412 | 157.6122 | 301.6306 | 232.2179 | 296.1426 | 205.5168
2012-10-30 01:09 | 25.8667 | 323.0116 | 253.68 | 317.5426 | 226.9201 | 89.0677 | 1.5495 | 39.375 | 61.0115 | 25.4115 | 2.4603 | 39.2584 | 96.5871 | 15.9568 | 21.4151 | 6.1403 | 157.5986 | 301.5965 | 232.2649 | 296.1275 | 205.505
2012-10-30 01:09 | 25.8585 | 323.0034 | 253.7264 | 317.5726 | 226.9631 | 89.077 | 1.5471 | 39.3139 | 60.8234 | 25.4109 | 2.4567 | 39.2013 | 96.3075 | 15.9566 | 21.4551 | 6.141 | 157.131 | 301.5483 | 232.2713 | 296.1175 | 205.508
2012-10-30 01:10 | 25.8659 | 323.0298 | 253.7182 | 317.6029 | 226.8879 | 89.0672 | 1.5456 | 39.274 | 60.7023 | 25.41 | 2.454 | 39.1575 | 96.0944 | 15.9563 | 21.4544 | 6.1403 | 156.7967 | 301.5754 | 232.2638 | 296.1485 | 205.4335
2012-10-30 01:10 | 25.8697 | 323.0184 | 253.6673 | 317.5838 | 226.8051 | 89.0666 | 1.5452 | 39.2624 | 60.6671 | 25.4098 | 2.4533 | 39.1459 | 96.0372 | 15.9563 | 21.4422 | 6.1403 | 156.7043 | 301.5762 | 232.2251 | 296.1416 | 205.3629
2012-10-30 01:10 | 25.8727 | 322.9832 | 253.7172 | 317.5868 | 226.4777 | 89.0544 | 1.545 | 39.2566 | 60.6497 | 25.4096 | 2.453 | 39.1401 | 96.0094 | 15.9562 | 21.4091 | 6.1394 | 156.6591 | 301.5741 | 232.3081 | 296.1777 | 205.0686
2012-10-30 01:10 | 25.8757 | 323.0015 | 253.5288 | 317.5783 | 225.3713 | 89.0643 | 1.545 | 39.2569 | 60.6501 | 25.4097 | 2.453 | 39.1403 | 96.0104 | 15.9562 | 21.4321 | 6.1401 | 156.6605 | 301.5694 | 232.0967 | 296.1462 | 203.9392
2012-10-30 01:10 | 25.8704 | 322.9771 | 253.3829 | 317.5692 | 225.9128 | 89.0654 | 1.5451 | 39.2617 | 60.6649 | 25.4097 | 2.4533 | 39.1452 | 96.0345 | 15.9562 | 21.471 | 6.1402 | 156.6994 | 301.5061 | 231.9119 | 296.0982 | 204.4418
2012-10-30 01:10 | 25.8689 | 323.0176 | 253.315 | 317.5486 | 226.108 | 89.0377 | 1.5452 | 39.2641 | 60.6723 | 25.4098 | 2.4534 | 39.1474 | 96.0458 | 15.9561 | 21.4374 | 6.1383 | 156.7181 | 301.5802 | 231.8776 | 296.1112 | 204.6706
2012-10-30 01:10 | 25.8637 | 323.0009 | 253.2902 | 317.551 | 226.3703 | 89.0516 | 1.5454 | 39.2675 | 60.6829 | 25.4097 | 2.4536 | 39.151 | 96.0622 | 15.9563 | 21.4562 | 6.1392 | 156.7451 | 301.5447 | 231.834 | 296.0948 | 204.9141
2012-10-30 01:10 | 25.8667 | 322.9772 | 253.2698 | 317.5769 | 226.2671 | 89.0597 | 1.5456 | 39.273 | 60.6991 | 25.41 | 2.454 | 39.1564 | 96.0897 | 15.9562 | 21.3991 | 6.1398 | 156.7888 | 301.5781 | 231.8707 | 296.1778 | 204.868
2012-10-30 01:10 | 25.8682 | 322.994 | 253.2947 | 317.5479 | 226.3748 | 89.04 | 1.5457 | 39.2766 | 60.7109 | 25.4098 | 2.4542 | 39.16 | 96.1065 | 15.9563 | 21.4086 | 6.1384 | 156.8174 | 301.5854 | 231.8861 | 296.1393 | 204.9662
2012-10-30 01:10 | 25.8517 | 323.0118 | 253.2782 | 317.5314 | 225.9571 | 89.029 | 1.5479 | 39.3323 | 60.8814 | 25.4106 | 2.4577 | 39.2157 | 96.3791 | 15.9564 | 21.4282 | 6.1377 | 157.2605 | 301.5836 | 231.85 | 296.1032 | 204.5289
2012-10-30 01:10 | 25.863 | 322.9621 | 253.2505 | 317.5274 | 226.1533 | 89.0602 | 1.5492 | 39.3661 | 60.9851 | 25.4109 | 2.4597 | 39.249 | 96.5414 | 15.9567 | 21.4596 | 6.1398 | 157.5265 | 301.5025 | 231.7909 | 296.0678 | 204.6937
2012-10-30 01:10 | 25.8637 | 323.0048 | 253.2316 | 317.509 | 225.5599 | 89.051 | 1.5496 | 39.3786 | 61.0228 | 25.4114 | 2.4605 | 39.2616 | 96.6032 | 15.9568 | 21.4362 | 6.1392 | 157.6259 | 301.5686 | 231.7954 | 296.0728 | 204.1237
2012-10-30 01:10 | 25.8585 | 322.9805 | 253.0623 | 317.4885 | 225.2951 | 89.0516 | 1.5501 | 39.391 | 61.0606 | 25.4116 | 2.4613 | 39.274 | 96.6636 | 15.9568 | 21.451 | 6.1392 | 157.7242 | 301.5295 | 231.6113 | 296.0375 | 203.8441
2012-10-30 01:10 | 25.851 | 322.9997 | 252.9884 | 317.4772 | 226.2554 | 89.0521 | 1.5504 | 39.3979 | 61.0821 | 25.4116 | 2.4617 | 39.281 | 96.6983 | 15.9568 | 21.4315 | 6.1393 | 157.7805 | 301.5682 | 231.5569 | 296.0457 | 204.8239
2012-10-30 01:10 | 25.851 | 323.0455 | 253.0236 | 317.4772 | 225.8502 | 89.0614 | 1.5505 | 39.4016 | 61.0929 | 25.4119 | 2.4619 | 39.2845 | 96.7159 | 15.9568 | 21.4636 | 6.1399 | 157.8088 | 301.5819 | 231.56 | 296.0136 | 204.3866
2012-10-30 01:10 | 25.8532 | 323.0057 | 253.0023 | 317.51 | 226.0805 | 89.0649 | 1.5506 | 39.4039 | 61.1001 | 25.4118 | 2.4621 | 39.2867 | 96.726 | 15.9569 | 21.4618 | 6.1401 | 157.8261 | 301.5439 | 231.5405 | 296.0482 | 204.6187
2012-10-30 01:10 | 25.8562 | 323.0011 | 252.9936 | 317.4366 | 226.2606 | 89.0597 | 1.5506 | 39.4038 | 61.1 | 25.4117 | 2.4621 | 39.2866 | 96.7255 | 15.9568 | 21.4487 | 6.1398 | 157.8255 | 301.5524 | 231.5449 | 295.9879 | 204.8119
2012-10-30 01:10 | 25.8405 | 322.9816 | 252.9975 | 317.4744 | 226.4337 | 89.055 | 1.5506 | 39.4038 | 61.0996 | 25.4119 | 2.462 | 39.2867 | 96.7256 | 15.957 | 21.421 | 6.1395 | 157.8252 | 301.5606 | 231.5765 | 296.0534 | 205.0127
2012-10-30 01:10 | 25.848 | 322.9547 | 253.0518 | 317.4895 | 226.6457 | 89.0631 | 1.5506 | 39.4035 | 61.0984 | 25.4121 | 2.462 | 39.2863 | 96.7225 | 15.9571 | 21.4606 | 6.14 | 157.8209 | 301.4941 | 231.5912 | 296.0289 | 205.1851
2012-10-30 01:10 | 25.8592 | 322.985 | 253.1373 | 317.4702 | 226.7395 | 89.0637 | 1.5506 | 39.4034 | 61.0977 | 25.4122 | 2.462 | 39.2862 | 96.7229 | 15.957 | 21.4237 | 6.1401 | 157.8206 | 301.5613 | 231.7136 | 296.0465 | 205.3158
2012-10-30 01:10 | 25.857 | 322.9599 | 253.2093 | 317.4832 | 226.7924 | 89.0724 | 1.5506 | 39.4031 | 61.0966 | 25.4122 | 2.462 | 39.2861 | 96.7217 | 15.9571 | 21.3934 | 6.1407 | 157.8183 | 301.5665 | 231.8159 | 296.0898 | 205.399
2012-10-30 01:10 | 25.86 | 322.9896 | 253.2631 | 317.4557 | 226.8308 | 89.0747 | 1.5505 | 39.4011 | 61.0909 | 25.412 | 2.4618 | 39.2837 | 96.7095 | 15.9572 | 21.4164 | 6.1408 | 157.8004 | 301.5732 | 231.8467 | 296.0393 | 205.4144
2012-10-30 01:10 | 25.8644 | 322.9749 | 253.3339 | 317.4792 | 226.8037 | 89.0851 | 1.5504 | 39.3981 | 61.0813 | 25.4122 | 2.4617 | 39.281 | 96.6967 | 15.9571 | 21.4128 | 6.1415 | 157.778 | 301.5621 | 231.9211 | 296.0664 | 205.3909
2012-10-30 01:10 | 25.863 | 322.9583 | 253.3794 | 317.4587 | 226.9282 | 89.0729 | 1.5503 | 39.3976 | 61.08 | 25.4121 | 2.4616 | 39.2805 | 96.6942 | 15.9571 | 21.3793 | 6.1407 | 157.7743 | 301.579 | 232.0001 | 296.0794 | 205.5489
2012-10-30 01:11 | 25.8644 | 322.9444 | 253.4042 | 317.4257 | 226.9336 | 89.0637 | 1.5503 | 39.3975 | 61.0795 | 25.4122 | 2.4617 | 39.2806 | 96.6953 | 15.9569 | 21.4248 | 6.1401 | 157.7749 | 301.5196 | 231.9794 | 296.0009 | 205.5088
2012-10-30 01:11 | 25.8682 | 322.9559 | 253.4979 | 317.4448 | 226.961 | 89.0868 | 1.5503 | 39.3975 | 61.0791 | 25.4123 | 2.4617 | 39.2805 | 96.6949 | 15.957 | 21.4086 | 6.1416 | 157.774 | 301.5473 | 232.0893 | 296.0362 | 205.5524
2012-10-30 01:11 | 25.8764 | 322.9984 | 253.5022 | 317.4683 | 227.0518 | 89.0706 | 1.5503 | 39.3969 | 61.0778 | 25.4121 | 2.4616 | 39.2797 | 96.6902 | 15.9571 | 21.4489 | 6.1405 | 157.768 | 301.5495 | 232.0533 | 296.0194 | 205.6029
2012-10-30 01:11 | 25.8757 | 322.9786 | 253.5405 | 317.5057 | 227.0511 | 89.0845 | 1.5503 | 39.3957 | 61.0741 | 25.4121 | 2.4615 | 39.2785 | 96.6858 | 15.9568 | 21.4321 | 6.1415 | 157.7598 | 301.5465 | 232.1084 | 296.0736 | 205.619
2012-10-30 01:11 | 25.8742 | 323.0191 | 253.5976 | 317.4622 | 226.963 | 89.0845 | 1.5502 | 39.395 | 61.0721 | 25.4121 | 2.4615 | 39.2781 | 96.6831 | 15.9569 | 21.4266 | 6.1415 | 157.7553 | 301.5925 | 232.171 | 296.0356 | 205.5364
2012-10-30 01:11 | 25.8757 | 323.0358 | 253.6577 | 317.4485 | 227.0668 | 89.0672 | 1.5502 | 39.3942 | 61.0693 | 25.4122 | 2.4615 | 39.2774 | 96.6798 | 15.9569 | 21.4361 | 6.1403 | 157.7491 | 301.5997 | 232.2216 | 296.0124 | 205.6307
2012-10-30 01:11 | 25.8794 | 323.0395 | 253.6536 | 317.4827 | 227.1964 | 89.1001 | 1.5502 | 39.3927 | 61.0648 | 25.4121 | 2.4614 | 39.2756 | 96.6714 | 15.9569 | 21.4358 | 6.1426 | 157.7362 | 301.6037 | 232.2178 | 296.0469 | 205.7606
2012-10-30 01:11 | 25.8809 | 323.0181 | 253.6863 | 317.4919 | 227.3159 | 89.0845 | 1.5501 | 39.3906 | 61.0583 | 25.4121 | 2.4612 | 39.2736 | 96.6618 | 15.9569 | 21.4293 | 6.1415 | 157.7201 | 301.5888 | 232.257 | 296.0626 | 205.8866
2012-10-30 01:11 | 25.8884 | 322.9989 | 253.7602 | 317.507 | 227.1975 | 89.0741 | 1.55 | 39.3896 | 61.0548 | 25.4122 | 2.4612 | 39.2726 | 96.6569 | 15.9568 | 21.4288 | 6.1408 | 157.7118 | 301.5701 | 232.3314 | 296.0782 | 205.7687
2012-10-30 01:11 | 25.8862 | 323.0044 | 253.7737 | 317.5392 | 227.2346 | 89.092 | 1.55 | 39.3889 | 61.0533 | 25.412 | 2.4611 | 39.2718 | 96.6521 | 15.957 | 21.4226 | 6.142 | 157.7054 | 301.5818 | 232.3511 | 296.1166 | 205.812
2012-10-30 01:11 | 25.8847 | 322.9762 | 253.8151 | 317.53 | 227.1348 | 89.0961 | 1.5499 | 39.3875 | 61.0482 | 25.4123 | 2.461 | 39.2707 | 96.6462 | 15.957 | 21.4652 | 6.1423 | 157.6944 | 301.511 | 232.3499 | 296.0648 | 205.6696
2012-10-30 01:11 | 25.8951 | 323.0629 | 253.806 | 317.5175 | 227.0036 | 89.0833 | 1.5499 | 39.3857 | 61.0429 | 25.4122 | 2.4609 | 39.2688 | 96.6379 | 15.9568 | 21.4275 | 6.1414 | 157.6808 | 301.6354 | 232.3785 | 296.09 | 205.5761
2012-10-30 01:11 | 25.8884 | 323.0142 | 253.8149 | 317.5337 | 226.8828 | 89.0845 | 1.5498 | 39.3842 | 61.0385 | 25.412 | 2.4608 | 39.2671 | 96.6304 | 15.9568 | 21.4007 | 6.1415 | 157.6689 | 301.6135 | 232.4142 | 296.133 | 205.4821
2012-10-30 01:11 | 25.8921 | 322.9912 | 253.8343 | 317.5412 | 226.8747 | 89.0839 | 1.5498 | 39.3834 | 61.0354 | 25.4124 | 2.4608 | 39.2665 | 96.6265 | 15.9569 | 21.4245 | 6.1414 | 157.6619 | 301.5667 | 232.4098 | 296.1167 | 205.4502
2012-10-30 01:11 | 25.8966 | 323.072 | 253.745 | 317.5496 | 227.1506 | 89.0909 | 1.5497 | 39.3815 | 61.03 | 25.4122 | 2.4607 | 39.2646 | 96.6171 | 15.9569 | 21.4651 | 6.1419 | 157.6471 | 301.6069 | 232.2799 | 296.0845 | 205.6855
2012-10-30 01:11 | 25.8951 | 323.0285 | 253.7982 | 317.5595 | 227.1373 | 89.0972 | 1.5497 | 39.3812 | 61.0287 | 25.4123 | 2.4607 | 39.2643 | 96.6161 | 15.9568 | 21.3913 | 6.1424 | 157.6448 | 301.6372 | 232.4069 | 296.1682 | 205.746
2012-10-30 01:11 | 25.9011 | 323.0765 | 253.7964 | 317.5388 | 227.1237 | 89.0961 | 1.5497 | 39.3806 | 61.0271 | 25.4122 | 2.4606 | 39.2637 | 96.6131 | 15.9568 | 21.4094 | 6.1423 | 157.6402 | 301.6671 | 232.387 | 296.1294 | 205.7143
2012-10-30 01:11 | 25.9064 | 323.0551 | 253.8681 | 317.5708 | 227.1368 | 89.0885 | 1.5496 | 39.3797 | 61.0245 | 25.4121 | 2.4606 | 39.2626 | 96.6085 | 15.9566 | 21.4147 | 6.1418 | 157.633 | 301.6404 | 232.4534 | 296.1561 | 205.7221
2012-10-30 01:11 | 25.9026 | 323.0513 | 253.8057 | 317.5632 | 227.1488 | 89.0909 | 1.5496 | 39.3785 | 61.02 | 25.4125 | 2.4605 | 39.2617 | 96.603 | 15.9569 | 21.4149 | 6.1419 | 157.623 | 301.6364 | 232.3908 | 296.1483 | 205.7339
2012-10-30 01:11 | 25.9049 | 323.0994 | 253.8783 | 317.5731 | 227.1825 | 89.1013 | 1.5496 | 39.3783 | 61.0189 | 25.4126 | 2.4605 | 39.2613 | 96.6009 | 15.9569 | 21.4011 | 6.1426 | 157.6198 | 301.6983 | 232.4772 | 296.172 | 205.7814
2012-10-30 01:11 | 25.9071 | 323.0482 | 253.8219 | 317.5333 | 227.1415 | 89.0932 | 1.5495 | 39.3756 | 61.0113 | 25.4123 | 2.4603 | 39.2587 | 96.5882 | 15.9568 | 21.4033 | 6.1421 | 157.5995 | 301.6449 | 232.4186 | 296.13 | 205.7382
2012-10-30 01:11 | 25.9064 | 323.0475 | 253.8525 | 317.567 | 227.1565 | 89.0914 | 1.5493 | 39.3712 | 60.9977 | 25.4123 | 2.46 | 39.2542 | 96.5662 | 15.9568 | 21.3384 | 6.142 | 157.5639 | 301.7091 | 232.5141 | 296.2286 | 205.8181
2012-10-30 01:11 | 25.9258 | 323.0821 | 253.8406 | 317.5749 | 227.172 | 89.0926 | 1.5493 | 39.3702 | 60.9944 | 25.4124 | 2.4599 | 39.2533 | 96.561 | 15.957 | 21.41 | 6.142 | 157.5553 | 301.6721 | 232.4306 | 296.1649 | 205.762
2012-10-30 01:11 | 25.9206 | 323.035 | 253.8471 | 317.6041 | 227.1235 | 89.0966 | 1.5492 | 39.3694 | 60.9921 | 25.4123 | 2.4599 | 39.2526 | 96.5587 | 15.9568 | 21.4168 | 6.1423 | 157.5508 | 301.6182 | 232.4303 | 296.1873 | 205.7067
2012-10-30 01:11 | 25.9183 | 323.0975 | 253.9347 | 317.5789 | 227.1999 | 89.0978 | 1.5463 | 39.295 | 60.7627 | 25.4119 | 2.4555 | 39.1818 | 96.2118 | 15.9566 | 21.3744 | 6.1424 | 156.9745 | 301.7231 | 232.5603 | 296.2045 | 205.8255
2012-10-30 01:12 | 25.9198 | 323.0876 | 253.8659 | 317.5766 | 227.3626 | 89.0857 | 1.5453 | 39.2685 | 60.6831 | 25.4109 | 2.4537 | 39.1519 | 96.0665 | 15.9563 | 21.3799 | 6.1416 | 156.7496 | 301.7077 | 232.486 | 296.1967 | 205.9827
2012-10-30 01:12 | 25.9146 | 323.029 | 253.8763 | 317.5981 | 227.2709 | 89.0903 | 1.545 | 39.2594 | 60.6553 | 25.4107 | 2.4531 | 39.1426 | 96.0212 | 15.9563 | 21.3707 | 6.1419 | 156.6765 | 301.6583 | 232.5056 | 296.2274 | 205.9002
2012-10-30 01:12 | 25.9198 | 323.0723 | 253.8854 | 317.588 | 226.1471 | 89.0764 | 1.5448 | 39.2551 | 60.6424 | 25.4107 | 2.4529 | 39.1385 | 96.0025 | 15.9561 | 21.3759 | 6.1409 | 156.6449 | 301.6964 | 232.5095 | 296.2121 | 204.7712
2012-10-30 01:12 | 25.9281 | 323.0806 | 253.6945 | 317.604 | 226.4308 | 89.088 | 1.5448 | 39.2541 | 60.6404 | 25.4102 | 2.4528 | 39.1376 | 95.9979 | 15.9561 | 21.3882 | 6.1417 | 156.6384 | 301.6924 | 232.3063 | 296.2158 | 205.0426
2012-10-30 01:12 | 25.9206 | 322.9854 | 253.6518 | 317.6385 | 226.7105 | 89.0747 | 1.5449 | 39.2567 | 60.6478 | 25.4105 | 2.453 | 39.1402 | 96.011 | 15.956 | 21.4128 | 6.1408 | 156.6588 | 301.5726 | 232.239 | 296.2257 | 205.2977
2012-10-30 01:12 | 25.9273 | 323.034 | 253.596 | 317.6452 | 226.3946 | 89.081 | 1.5451 | 39.2612 | 60.6622 | 25.4103 | 2.4533 | 39.1447 | 96.0333 | 15.956 | 21.4235 | 6.1412 | 156.6955 | 301.6105 | 232.1725 | 296.2217 | 204.9711
2012-10-30 01:12 | 25.9086 | 322.9924 | 253.5344 | 317.5997 | 226.4782 | 89.0643 | 1.5453 | 39.2669 | 60.6788 | 25.4106 | 2.4536 | 39.1501 | 96.0595 | 15.956 | 21.4169 | 6.1401 | 156.7383 | 301.5755 | 232.1175 | 296.1828 | 205.0613
2012-10-30 01:12 | 25.9093 | 323.0465 | 253.5233 | 317.6234 | 226.8172 | 89.0845 | 1.5455 | 39.2713 | 60.6924 | 25.4106 | 2.4539 | 39.1545 | 96.0814 | 15.956 | 21.4537 | 6.1415 | 156.7737 | 301.5928 | 232.0696 | 296.1697 | 205.3635
2012-10-30 01:12 | 25.9086 | 323.0344 | 253.507 | 317.6036 | 225.7268 | 89.0724 | 1.5457 | 39.2766 | 60.7091 | 25.4105 | 2.4542 | 39.1595 | 96.1059 | 15.956 | 21.449 | 6.1407 | 156.815 | 301.5854 | 232.058 | 296.1546 | 204.2778
2012-10-30 01:12 | 25.9049 | 323.065 | 253.3939 | 317.5808 | 226.1952 | 89.0683 | 1.5484 | 39.3463 | 60.9228 | 25.4113 | 2.4584 | 39.2276 | 96.4391 | 15.9563 | 21.4774 | 6.1404 | 157.362 | 301.5876 | 231.9165 | 296.1034 | 204.7178
2012-10-30 01:12 | 25.9004 | 323.0453 | 253.4207 | 317.5839 | 226.7965 | 89.0805 | 1.5493 | 39.3707 | 60.9971 | 25.4119 | 2.4601 | 39.2536 | 96.5658 | 15.9564 | 21.4528 | 6.1412 | 157.5629 | 301.5925 | 231.9679 | 296.1311 | 205.3437
2012-10-30 01:12 | 25.9026 | 323.0704 | 253.5088 | 317.5403 | 226.7239 | 89.0776 | 1.5498 | 39.3829 | 61.0345 | 25.412 | 2.4608 | 39.2659 | 96.626 | 15.9565 | 21.4711 | 6.141 | 157.6605 | 301.5993 | 232.0377 | 296.0692 | 205.2528
2012-10-30 01:12 | 25.9019 | 323.062 | 253.4534 | 317.5434 | 227.1992 | 89.0868 | 1.5501 | 39.3917 | 61.0619 | 25.4121 | 2.4614 | 39.2748 | 96.6692 | 15.9566 | 21.4503 | 6.1416 | 157.7312 | 301.6117 | 232.0031 | 296.0931 | 205.7489
2012-10-30 01:12 | 25.8989 | 323.0628 | 253.5325 | 317.5289 | 227.2001 | 89.1013 | 1.5503 | 39.3976 | 61.0795 | 25.4123 | 2.4617 | 39.2804 | 96.6964 | 15.9566 | 21.4032 | 6.1426 | 157.776 | 301.6596 | 232.1293 | 296.1257 | 205.7969
2012-10-30 01:12 | 25.8929 | 323.0912 | 253.5968 | 317.5382 | 227.1745 | 89.0909 | 1.5505 | 39.4023 | 61.0946 | 25.4121 | 2.462 | 39.2853 | 96.7207 | 15.9566 | 21.4333 | 6.1419 | 157.8152 | 301.6579 | 232.1635 | 296.1049 | 205.7412
2012-10-30 01:12 | 25.8974 | 323.1071 | 253.6638 | 317.558 | 227.1868 | 89.0949 | 1.5506 | 39.404 | 61.0988 | 25.4126 | 2.4621 | 39.2868 | 96.7275 | 15.9567 | 21.4538 | 6.1422 | 157.8264 | 301.6533 | 232.21 | 296.1042 | 205.733
2012-10-30 01:12 | 25.8914 | 323.0973 | 253.6773 | 317.5444 | 227.1415 | 89.1151 | 1.5506 | 39.4045 | 61.1001 | 25.4126 | 2.4621 | 39.2875 | 96.7309 | 15.9567 | 21.4358 | 6.1436 | 157.8311 | 301.6615 | 232.2415 | 296.1086 | 205.7057
2012-10-30 01:12 | 25.8974 | 323.1033 | 253.7263 | 317.5389 | 227.1475 | 89.0955 | 1.5505 | 39.4036 | 61.0966 | 25.4129 | 2.4621 | 39.2865 | 96.7257 | 15.9568 | 21.4659 | 6.1422 | 157.8223 | 301.6374 | 232.2604 | 296.073 | 205.6816
2012-10-30 01:12 | 25.8921 | 323.0751 | 253.7093 | 317.5451 | 227.0399 | 89.1013 | 1.5506 | 39.4038 | 61.098 | 25.4127 | 2.4621 | 39.2868 | 96.7277 | 15.9567 | 21.4365 | 6.1426 | 157.8257 | 301.6386 | 232.2728 | 296.1086 | 205.6034
2012-10-30 01:12 | 25.8936 | 323.0957 | 253.7342 | 317.5351 | 227.2263 | 89.1047 | 1.5506 | 39.4038 | 61.0978 | 25.4128 | 2.4621 | 39.2868 | 96.7284 | 15.9566 | 21.442 | 6.1429 | 157.8262 | 301.6537 | 232.2922 | 296.0931 | 205.7843
2012-10-30 01:12 | 25.8876 | 323.0706 | 253.7946 | 317.552 | 227.4169 | 89.1065 | 1.5506 | 39.4034 | 61.0975 | 25.4123 | 2.4621 | 39.2863 | 96.7254 | 15.9567 | 21.44 | 6.143 | 157.8229 | 301.6306 | 232.3546 | 296.112 | 205.9769
2012-10-30 01:12 | 25.8884 | 323.0752 | 253.854 | 317.5528 | 227.229 | 89.0978 | 1.5505 | 39.4029 | 61.0949 | 25.4127 | 2.462 | 39.2857 | 96.7222 | 15.9567 | 21.4408 | 6.1424 | 157.8171 | 301.6344 | 232.4132 | 296.112 | 205.7882
2012-10-30 01:12 | 25.8936 | 323.0614 | 253.8436 | 317.5886 | 227.1201 | 89.1047 | 1.5505 | 39.4015 | 61.091 | 25.4126 | 2.462 | 39.2845 | 96.7168 | 15.9566 | 21.4179 | 6.1429 | 157.8078 | 301.6435 | 232.4257 | 296.1707 | 205.7022
2012-10-30 01:12 | 25.8884 | 323.0867 | 253.8462 | 317.591 | 227.2329 | 89.1076 | 1.5504 | 39.4001 | 61.0863 | 25.4127 | 2.4619 | 39.2831 | 96.7094 | 15.9567 | 21.4047 | 6.1431 | 157.7957 | 301.682 | 232.4415 | 296.1863 | 205.8282
2012-10-30 01:13 | 25.8951 | 323.1086 | 253.8529 | 317.613 | 227.2907 | 89.1163 | 1.5504 | 39.4001 | 61.0863 | 25.4127 | 2.4618 | 39.283 | 96.7084 | 15.9567 | 21.3873 | 6.1437 | 157.7947 | 301.7213 | 232.4656 | 296.2257 | 205.9034
2012-10-30 01:13 | 25.8959 | 323.1094 | 253.9006 | 317.5794 | 227.2325 | 89.1001 | 1.5503 | 39.3981 | 61.0795 | 25.4129 | 2.4617 | 39.281 | 96.6989 | 15.9568 | 21.3962 | 6.1426 | 157.7784 | 301.7132 | 232.5044 | 296.1832 | 205.8363
2012-10-30 01:13 | 25.8921 | 323.0942 | 253.8928 | 317.6176 | 227.0989 | 89.107 | 1.5503 | 39.3973 | 61.0777 | 25.4127 | 2.4617 | 39.2801 | 96.6945 | 15.9567 | 21.4365 | 6.143 | 157.7722 | 301.6577 | 232.4563 | 296.1811 | 205.6624
2012-10-30 01:13 | 25.8899 | 323.1149 | 253.8789 | 317.6192 | 227.1675 | 89.1169 | 1.5502 | 39.3957 | 61.0727 | 25.4127 | 2.4616 | 39.2789 | 96.6889 | 15.9566 | 21.4062 | 6.1437 | 157.7616 | 301.7087 | 232.4727 | 296.213 | 205.7613
2012-10-30 01:13 | 25.8989 | 323.1086 | 253.9348 | 317.6397 | 227.4322 | 89.1053 | 1.5502 | 39.3936 | 61.0661 | 25.4128 | 2.4615 | 39.2766 | 96.6783 | 15.9565 | 21.4232 | 6.1429 | 157.7444 | 301.6854 | 232.5116 | 296.2165 | 206.009
2012-10-30 01:13 | 25.8989 | 323.101 | 253.9543 | 317.6053 | 227.5069 | 89.1013 | 1.5501 | 39.3931 | 61.064 | 25.4129 | 2.4614 | 39.276 | 96.6755 | 15.9566 | 21.4192 | 6.1426 | 157.7394 | 301.6818 | 232.5351 | 296.1861 | 206.0877
2012-10-30 01:13 | 25.8921 | 323.1285 | 253.9944 | 317.6596 | 227.5827 | 89.1099 | 1.5501 | 39.392 | 61.0609 | 25.4128 | 2.4614 | 39.2751 | 96.6709 | 15.9565 | 21.3883 | 6.1432 | 157.7318 | 301.7402 | 232.6061 | 296.2713 | 206.1944
2012-10-30 01:13 | 25.9019 | 323.1574 | 254.055 | 317.6656 | 227.4155 | 89.1146 | 1.55 | 39.3902 | 61.0557 | 25.4127 | 2.4613 | 39.2732 | 96.662 | 15.9565 | 21.3861 | 6.1436 | 157.7177 | 301.7713 | 232.6689 | 296.2795 | 206.0294
2012-10-30 01:13 | 25.9026 | 323.1428 | 254.0986 | 317.6816 | 227.3966 | 89.1111 | 1.5498 | 39.386 | 61.0422 | 25.4129 | 2.461 | 39.2691 | 96.6417 | 15.9565 | 21.4029 | 6.1433 | 157.6839 | 301.7399 | 232.6957 | 296.2787 | 205.9937
2012-10-30 01:13 | 25.9026 | 323.1505 | 254.0713 | 317.6663 | 227.487 | 89.1105 | 1.5498 | 39.3848 | 61.0392 | 25.4126 | 2.4609 | 39.2676 | 96.6342 | 15.9565 | 21.3908 | 6.1433 | 157.6733 | 301.7597 | 232.6805 | 296.2755 | 206.0962
2012-10-30 01:13 | 25.9086 | 323.1984 | 254.0929 | 317.6723 | 227.3986 | 89.1157 | 1.5498 | 39.3843 | 61.0378 | 25.4125 | 2.4609 | 39.267 | 96.6317 | 15.9565 | 21.4329 | 6.1436 | 157.6695 | 301.7655 | 232.66 | 296.2394 | 205.9657
2012-10-30 01:13 | 25.9086 | 323.1526 | 254.1242 | 317.7029 | 227.5284 | 89.1198 | 1.5497 | 39.3815 | 61.0294 | 25.4124 | 2.4607 | 39.2645 | 96.6195 | 15.9564 | 21.453 | 6.1439 | 157.649 | 301.6996 | 232.6712 | 296.2499 | 206.0754
2012-10-30 01:13 | 25.9004 | 323.1444 | 254.0964 | 317.6717 | 227.5949 | 89.1198 | 1.5496 | 39.3796 | 61.0228 | 25.4127 | 2.4606 | 39.2626 | 96.61 | 15.9565 | 21.3966 | 6.1439 | 157.6328 | 301.7478 | 232.6998 | 296.2751 | 206.1983
2012-10-30 01:13 | 25.9026 | 323.1505 | 254.1299 | 317.6434 | 227.4949 | 89.1122 | 1.5496 | 39.379 | 61.021 | 25.4126 | 2.4606 | 39.262 | 96.6065 | 15.9565 | 21.3828 | 6.1434 | 157.6275 | 301.7677 | 232.7471 | 296.2606 | 206.1121
2012-10-30 01:13 | 25.9168 | 323.2066 | 254.1167 | 317.6881 | 227.5484 | 89.1186 | 1.5495 | 39.377 | 61.0145 | 25.4127 | 2.4604 | 39.26 | 96.5972 | 15.9564 | 21.4211 | 6.1438 | 157.6117 | 301.7855 | 232.6956 | 296.267 | 206.1273
2012-10-30 01:13 | 25.9108 | 323.1663 | 254.1459 | 317.6821 | 227.6132 | 89.1174 | 1.5495 | 39.3764 | 61.0128 | 25.4128 | 2.4604 | 39.2594 | 96.5936 | 15.9566 | 21.3709 | 6.1438 | 157.6064 | 301.7954 | 232.775 | 296.3112 | 206.2423
2012-10-30 01:13 | 25.9146 | 323.1968 | 254.1458 | 317.7394 | 227.5383 | 89.1203 | 1.5494 | 39.3752 | 61.0088 | 25.4128 | 2.4603 | 39.258 | 96.5876 | 15.9564 | 21.4309 | 6.144 | 157.5964 | 301.7659 | 232.7149 | 296.3085 | 206.1074
2012-10-30 01:13 | 25.9101 | 323.2228 | 254.1452 | 317.712 | 227.4001 | 89.1226 | 1.5494 | 39.3742 | 61.0061 | 25.4127 | 2.4603 | 39.257 | 96.5832 | 15.9563 | 21.4465 | 6.1441 | 157.5893 | 301.7763 | 232.6987 | 296.2655 | 205.9536
2012-10-30 01:13 | 25.9101 | 323.2114 | 254.153 | 317.7158 | 227.4276 | 89.125 | 1.5494 | 39.3733 | 61.0031 | 25.4127 | 2.4603 | 39.2564 | 96.581 | 15.9562 | 21.4144 | 6.1443 | 157.5841 | 301.797 | 232.7386 | 296.3014 | 206.0132
2012-10-30 01:13 | 25.9079 | 323.2168 | 254.182 | 317.7403 | 227.3783 | 89.1192 | 1.5491 | 39.3664 | 60.9822 | 25.4125 | 2.4598 | 39.2494 | 96.5462 | 15.9562 | 21.4282 | 6.1439 | 157.5284 | 301.7886 | 232.7538 | 296.3121 | 205.9501
2012-10-30 01:13 | 25.9123 | 323.2441 | 254.1747 | 317.7409 | 227.3512 | 89.1111 | 1.5454 | 39.2691 | 60.6846 | 25.4111 | 2.4538 | 39.1529 | 96.075 | 15.9557 | 21.3925 | 6.1433 | 156.7596 | 301.8516 | 232.7822 | 296.3484 | 205.9587
2012-10-30 01:14 | 25.9034 | 323.1894 | 254.1541 | 317.7358 | 227.4996 | 89.1111 | 1.5451 | 39.2622 | 60.6633 | 25.4111 | 2.4534 | 39.1456 | 96.0394 | 15.9558 | 21.3956 | 6.1433 | 156.7027 | 301.7938 | 232.7585 | 296.3402 | 206.104
2012-10-30 01:14 | 25.9086 | 323.1793 | 254.1281 | 317.7258 | 227.0486 | 89.1047 | 1.5449 | 39.2564 | 60.646 | 25.4108 | 2.4531 | 39.1399 | 96.0122 | 15.9556 | 21.4089 | 6.1429 | 156.6582 | 301.7704 | 232.7192 | 296.3169 | 205.6397
2012-10-30 01:14 | 25.9153 | 323.1517 | 254.0254 | 317.7401 | 225.8122 | 89.1076 | 1.5448 | 39.2542 | 60.639 | 25.4109 | 2.4529 | 39.1376 | 96.0004 | 15.9556 | 21.4477 | 6.1431 | 156.6394 | 301.704 | 232.5777 | 296.2924 | 204.3645
2012-10-30 01:14 | 25.9161 | 323.1678 | 253.8622 | 317.7562 | 226.5408 | 89.1065 | 1.5448 | 39.2548 | 60.6408 | 25.4109 | 2.4529 | 39.1381 | 96.0035 | 15.9555 | 21.4284 | 6.143 | 156.6443 | 301.7394 | 232.4338 | 296.3278 | 205.1124
2012-10-30 01:14 | 25.9131 | 323.1571 | 253.8748 | 317.757 | 227.4031 | 89.107 | 1.545 | 39.2597 | 60.6563 | 25.4108 | 2.4533 | 39.143 | 96.0279 | 15.9555 | 21.4214 | 6.143 | 156.6842 | 301.7357 | 232.4534 | 296.3356 | 205.9817
2012-10-30 01:14 | 25.9101 | 323.1656 | 253.9265 | 317.6929 | 227.1366 | 89.0978 | 1.5452 | 39.2647 | 60.6711 | 25.411 | 2.4536 | 39.148 | 96.0529 | 15.9554 | 21.4304 | 6.1424 | 156.724 | 301.7352 | 232.4961 | 296.2625 | 205.7062
2012-10-30 01:14 | 25.9056 | 323.1268 | 253.8595 | 317.6769 | 226.9944 | 89.0897 | 1.5453 | 39.2685 | 60.6832 | 25.4108 | 2.4538 | 39.1518 | 96.0702 | 15.9557 | 21.4259 | 6.1418 | 156.7534 | 301.7009 | 232.4336 | 296.251 | 205.5685
2012-10-30 01:14 | 25.8996 | 323.1055 | 253.8535 | 317.7626 | 226.9845 | 89.1001 | 1.5455 | 39.2735 | 60.6982 | 25.4111 | 2.4541 | 39.1567 | 96.0944 | 15.9556 | 21.4279 | 6.1426 | 156.7927 | 301.6776 | 232.4256 | 296.3347 | 205.5566
2012-10-30 01:14 | 25.9019 | 323.1192 | 253.848 | 317.7 | 227.3251 | 89.0937 | 1.5467 | 39.305 | 60.7948 | 25.4114 | 2.456 | 39.1868 | 96.2417 | 15.9558 | 21.4343 | 6.1421 | 157.0365 | 301.6849 | 232.4137 | 296.2657 | 205.8908
2012-10-30 01:14 | 25.9011 | 323.1299 | 253.9136 | 317.7183 | 227.3439 | 89.0978 | 1.5489 | 39.361 | 60.9661 | 25.4123 | 2.4595 | 39.2437 | 96.5193 | 15.9561 | 21.4615 | 6.1424 | 157.4853 | 301.6684 | 232.4521 | 296.2568 | 205.8824
2012-10-30 01:14 | 25.8981 | 323.1498 | 253.9145 | 317.6809 | 227.1285 | 89.1088 | 1.5494 | 39.3744 | 61.0074 | 25.4124 | 2.4603 | 39.2571 | 96.5851 | 15.9561 | 21.4465 | 6.1432 | 157.5926 | 301.7033 | 232.468 | 296.2344 | 205.682
2012-10-30 01:14 | 25.8959 | 323.1285 | 253.924 | 317.6711 | 226.9375 | 89.1099 | 1.5498 | 39.3858 | 61.0417 | 25.4128 | 2.461 | 39.2687 | 96.642 | 15.9561 | 21.4242 | 6.1432 | 157.6837 | 301.7043 | 232.4998 | 296.2469 | 205.5133
2012-10-30 01:14 | 25.8989 | 323.1162 | 253.8957 | 317.6473 | 227.3299 | 89.1082 | 1.5502 | 39.3949 | 61.0695 | 25.4129 | 2.4616 | 39.2777 | 96.6854 | 15.9563 | 21.4152 | 6.1431 | 157.7549 | 301.701 | 232.4805 | 296.2321 | 205.9147
2012-10-30 01:14 | 25.8921 | 323.1323 | 253.971 | 317.652 | 227.3507 | 89.1255 | 1.5503 | 39.399 | 61.0823 | 25.413 | 2.4618 | 39.2817 | 96.7046 | 15.9564 | 21.4485 | 6.1443 | 157.7868 | 301.6838 | 232.5225 | 296.2035 | 205.9022
2012-10-30 01:14 | 25.8996 | 323.1322 | 253.9199 | 317.6633 | 227.3385 | 89.1134 | 1.5504 | 39.4012 | 61.0887 | 25.4131 | 2.462 | 39.2842 | 96.7172 | 15.9563 | 21.4039 | 6.1435 | 157.8059 | 301.7283 | 232.516 | 296.2594 | 205.9346
2012-10-30 01:14 | 25.8959 | 323.1171 | 253.9279 | 317.6405 | 227.3427 | 89.0978 | 1.5504 | 39.402 | 61.0907 | 25.4133 | 2.462 | 39.2848 | 96.7202 | 15.9563 | 21.4042 | 6.1424 | 157.8109 | 301.7129 | 232.5237 | 296.2363 | 205.9385
2012-10-30 01:14 | 25.9034 | 323.1513 | 253.9627 | 317.6633 | 227.3502 | 89.129 | 1.5505 | 39.4025 | 61.0923 | 25.4133 | 2.462 | 39.2853 | 96.7215 | 15.9565 | 21.3796 | 6.1446 | 157.8138 | 301.7717 | 232.5831 | 296.2837 | 205.9706
2012-10-30 01:14 | 25.8974 | 323.1071 | 253.9372 | 317.6535 | 227.238 | 89.1198 | 1.5505 | 39.4027 | 61.0927 | 25.4133 | 2.4621 | 39.2855 | 96.7231 | 15.9564 | 21.4057 | 6.1439 | 157.8158 | 301.7014 | 232.5315 | 296.2478 | 205.8323
2012-10-30 01:14 | 25.8966 | 323.1101 | 253.9716 | 317.6488 | 227.3748 | 89.118 | 1.5505 | 39.4024 | 61.0923 | 25.4132 | 2.462 | 39.2855 | 96.722 | 15.9566 | 21.4049 | 6.1438 | 157.8143 | 301.7052 | 232.5667 | 296.2439 | 205.9699
2012-10-30 01:14 | 25.8989 | 323.0934 | 253.9309 | 317.6359 | 227.4047 | 89.114 | 1.5504 | 39.4011 | 61.088 | 25.4133 | 2.4619 | 39.2841 | 96.7152 | 15.9565 | 21.4112 | 6.1435 | 157.8032 | 301.6822 | 232.5197 | 296.2247 | 205.9935
2012-10-30 01:14 | 25.9041 | 323.1405 | 253.9751 | 317.6678 | 227.3037 | 89.133 | 1.5504 | 39.4001 | 61.0846 | 25.4135 | 2.4619 | 39.2832 | 96.7103 | 15.9566 | 21.4204 | 6.1448 | 157.7949 | 301.7201 | 232.5547 | 296.2474 | 205.8833
2012-10-30 01:14 | 25.9071 | 323.1435 | 253.9938 | 317.6632 | 227.2752 | 89.129 | 1.5503 | 39.3998 | 61.0832 | 25.4137 | 2.4619 | 39.2827 | 96.7084 | 15.9565 | 21.3993 | 6.1446 | 157.7916 | 301.7442 | 232.5945 | 296.2639 | 205.8759
2012-10-30 01:14 | 25.9064 | 323.1352 | 254.0321 | 317.6357 | 227.3296 | 89.133 | 1.5503 | 39.3996 | 61.0826 | 25.4136 | 2.4618 | 39.2824 | 96.7066 | 15.9566 | 21.4107 | 6.1448 | 157.7892 | 301.7245 | 232.6214 | 296.225 | 205.9189
2012-10-30 01:14 | 25.9056 | 323.1268 | 254.0001 | 317.6578 | 227.376 | 89.1359 | 1.5503 | 39.3981 | 61.0779 | 25.4136 | 2.4617 | 39.2809 | 96.6996 | 15.9565 | 21.4219 | 6.145 | 157.7775 | 301.7049 | 232.5782 | 296.2359 | 205.9541
2012-10-30 01:15 | 25.9056 | 323.1001 | 254.0196 | 317.6693 | 227.4114 | 89.1174 | 1.5502 | 39.3967 | 61.0738 | 25.4135 | 2.4617 | 39.2797 | 96.6944 | 15.9564 | 21.4219 | 6.1438 | 157.7682 | 301.6782 | 232.5977 | 296.2474 | 205.9895
2012-10-30 01:15 | 25.9064 | 323.1581 | 254.0438 | 317.6625 | 227.5616 | 89.1296 | 1.5502 | 39.3955 | 61.0704 | 25.4134 | 2.4616 | 39.2784 | 96.688 | 15.9564 | 21.3986 | 6.1446 | 157.7584 | 301.7595 | 232.6452 | 296.2639 | 206.163
2012-10-30 01:15 | 25.9041 | 323.1291 | 254.0884 | 317.6793 | 227.6222 | 89.1348 | 1.5502 | 39.3947 | 61.0682 | 25.4133 | 2.4616 | 39.2776 | 96.684 | 15.9564 | 21.4365 | 6.145 | 157.7521 | 301.6926 | 232.6519 | 296.2428 | 206.1857
2012-10-30 01:15 | 25.9064 | 323.1466 | 254.1063 | 317.6854 | 227.5498 | 89.125 | 1.5502 | 39.3948 | 61.0682 | 25.4134 | 2.4616 | 39.2778 | 96.6856 | 15.9563 | 21.4267 | 6.1443 | 157.7537 | 301.7199 | 232.6796 | 296.2587 | 206.1231
2012-10-30 01:15 | 25.9079 | 323.1405 | 254.1274 | 317.7518 | 227.8188 | 89.1429 | 1.5501 | 39.3935 | 61.064 | 25.4135 | 2.4615 | 39.2765 | 96.6789 | 15.9563 | 21.4162 | 6.1455 | 157.7429 | 301.7243 | 232.7112 | 296.3356 | 206.4026
2012-10-30 01:15 | 25.9041 | 323.1863 | 254.1665 | 317.7289 | 227.693 | 89.1342 | 1.55 | 39.392 | 61.0596 | 25.4134 | 2.4614 | 39.2751 | 96.673 | 15.9562 | 21.4124 | 6.1449 | 157.7326 | 301.7739 | 232.7541 | 296.3165 | 206.2806
2012-10-30 01:15 | 25.9064 | 323.1581 | 254.2118 | 317.7274 | 227.656 | 89.1296 | 1.55 | 39.3907 | 61.0557 | 25.4133 | 2.4614 | 39.2738 | 96.6665 | 15.9562 | 21.3986 | 6.1446 | 157.7223 | 301.7595 | 232.8132 | 296.3288 | 206.2574
2012-10-30 01:15 | 25.9093 | 323.1724 | 254.1952 | 317.6883 | 227.5566 | 89.1342 | 1.55 | 39.3894 | 61.0517 | 25.4133 | 2.4613 | 39.2722 | 96.6588 | 15.9562 | 21.4336 | 6.1449 | 157.7104 | 301.7388 | 232.7616 | 296.2547 | 206.123
2012-10-30 01:15 | 25.9123 | 323.2021 | 254.2177 | 317.7829 | 227.5636 | 89.1671 | 1.5499 | 39.3884 | 61.0486 | 25.4133 | 2.4612 | 39.2714 | 96.6552 | 15.9561 | 21.4005 | 6.1472 | 157.7038 | 301.8016 | 232.8172 | 296.3824 | 206.1631
2012-10-30 01:15 | 25.9108 | 323.1968 | 254.224 | 317.7432 | 227.5109 | 89.1371 | 1.5498 | 39.3868 | 61.0437 | 25.4133 | 2.4611 | 39.2697 | 96.6467 | 15.9561 | 21.3749 | 6.1451 | 157.6904 | 301.8219 | 232.8491 | 296.3683 | 206.136
2012-10-30 01:15 | 25.9161 | 323.2059 | 254.2176 | 317.7638 | 227.5123 | 89.1296 | 1.5498 | 39.3853 | 61.0389 | 25.4133 | 2.461 | 39.2684 | 96.64 | 15.9562 | 21.4043 | 6.1446 | 157.6789 | 301.8016 | 232.8133 | 296.3595 | 206.108
2012-10-30 01:15 | 25.9123 | 323.1983 | 254.1669 | 317.76 | 227.4967 | 89.129 | 1.5497 | 39.3834 | 61.0332 | 25.4132 | 2.4609 | 39.2663 | 96.6306 | 15.956 | 21.3804 | 6.1446 | 157.6639 | 301.8179 | 232.7865 | 296.3796 | 206.1163
2012-10-30 01:15 | 25.9318 | 323.2102 | 254.2645 | 317.7795 | 227.4376 | 89.14 | 1.5497 | 39.3824 | 61.0301 | 25.4132 | 2.4609 | 39.2652 | 96.626 | 15.9559 | 21.412 | 6.1453 | 157.6562 | 301.7982 | 232.8525 | 296.3675 | 206.0256
2012-10-30 01:15 | 25.9191 | 323.2089 | 254.205 | 317.7592 | 227.4249 | 89.1417 | 1.5496 | 39.3813 | 61.0267 | 25.4132 | 2.4608 | 39.2642 | 96.6203 | 15.9561 | 21.3551 | 6.1454 | 157.647 | 301.8538 | 232.8499 | 296.4041 | 206.0698
2012-10-30 01:15 | 25.9236 | 323.2172 | 254.2095 | 317.7599 | 227.5237 | 89.1365 | 1.5496 | 39.3798 | 61.0221 | 25.4133 | 2.4607 | 39.2628 | 96.6143 | 15.9559 | 21.3476 | 6.1451 | 157.6365 | 301.8696 | 232.8619 | 296.4123 | 206.1761
2012-10-30 01:15 | 25.9333 | 323.1964 | 254.227 | 317.8001 | 227.502 | 89.133 | 1.5495 | 39.379 | 61.0194 | 25.4134 | 2.4606 | 39.262 | 96.6093 | 15.956 | 21.3733 | 6.1448 | 157.6287 | 301.8231 | 232.8537 | 296.4268 | 206.1287
2012-10-30 01:15 | 25.9288 | 323.2034 | 254.2108 | 317.7727 | 227.4581 | 89.1388 | 1.5495 | 39.3769 | 61.0132 | 25.4132 | 2.4605 | 39.2599 | 96.5995 | 15.956 | 21.3849 | 6.1452 | 157.6127 | 301.8185 | 232.8259 | 296.3878 | 206.0732
2012-10-30 01:15 | 25.9266 | 323.1706 | 254.1929 | 317.7705 | 227.5267 | 89.1446 | 1.5494 | 39.3757 | 61.0094 | 25.4132 | 2.4605 | 39.2588 | 96.5944 | 15.9559 | 21.3586 | 6.1456 | 157.6038 | 301.812 | 232.8343 | 296.4119 | 206.1681
2012-10-30 01:15 | 25.937 | 323.2268 | 254.219 | 317.7962 | 227.5843 | 89.1463 | 1.5494 | 39.3748 | 61.0068 | 25.4132 | 2.4604 | 39.2576 | 96.5892 | 15.9558 | 21.4212 | 6.1457 | 157.5961 | 301.8056 | 232.7978 | 296.375 | 206.1631
2012-10-30 01:15 | 25.9333 | 323.2231 | 254.227 | 317.7925 | 227.5413 | 89.1267 | 1.5455 | 39.2748 | 60.7003 | 25.4119 | 2.4544 | 39.1602 | 96.1128 | 15.9554 | 21.3573 | 6.1444 | 156.8131 | 301.8658 | 232.8697 | 296.4352 | 206.184
2012-10-30 01:15 | 25.9236 | 323.2096 | 254.2095 | 317.7713 | 227.4726 | 89.129 | 1.5452 | 39.2664 | 60.675 | 25.4116 | 2.4537 | 39.1497 | 96.0623 | 15.9552 | 21.3556 | 6.1446 | 156.7373 | 301.854 | 232.8539 | 296.4157 | 206.117
2012-10-30 01:15 | 25.9408 | 323.2077 | 254.172 | 317.8191 | 227.3286 | 89.1238 | 1.5449 | 39.2587 | 60.6522 | 25.4112 | 2.4532 | 39.142 | 96.0247 | 15.9552 | 21.3688 | 6.1442 | 156.6769 | 301.8389 | 232.8032 | 296.4503 | 205.9598
2012-10-30 01:16 | 25.9311 | 323.179 | 254.1232 | 317.8017 | 226.2804 | 89.1169 | 1.5448 | 39.2547 | 60.639 | 25.4116 | 2.453 | 39.1381 | 96.0056 | 15.9552 | 21.3952 | 6.1437 | 156.6445 | 301.7838 | 232.728 | 296.4065 | 204.8852
2012-10-30 01:16 | 25.9355 | 323.1948 | 253.9558 | 317.8023 | 225.9032 | 89.1076 | 1.5449 | 39.2587 | 60.6518 | 25.4113 | 2.4533 | 39.142 | 96.0255 | 15.9551 | 21.4077 | 6.1431 | 156.6773 | 301.7871 | 232.5481 | 296.3946 | 204.4955
2012-10-30 01:16 | 25.9303 | 323.1896 | 253.7904 | 317.8047 | 225.8429 | 89.1076 | 1.5451 | 39.2631 | 60.666 | 25.4112 | 2.4535 | 39.1463 | 96.0466 | 15.9551 | 21.4065 | 6.1431 | 156.7126 | 301.7831 | 232.3839 | 296.3982 | 204.4364
2012-10-30 01:16 | 25.9266 | 323.163 | 253.6656 | 317.7858 | 225.9966 | 89.1226 | 1.5452 | 39.2662 | 60.6753 | 25.4113 | 2.4537 | 39.1496 | 96.0618 | 15.9552 | 21.4028 | 6.1441 | 156.7371 | 301.7602 | 232.2628 | 296.383 | 204.5938
2012-10-30 01:16 | 25.934 | 323.1857 | 253.6262 | 317.7855 | 225.6066 | 89.0961 | 1.5454 | 39.2714 | 60.6914 | 25.4113 | 2.454 | 39.1546 | 96.0869 | 15.9552 | 21.4102 | 6.1423 | 156.7783 | 301.7755 | 232.216 | 296.3753 | 204.1964
2012-10-30 01:16 | 25.9281 | 323.1721 | 253.4679 | 317.7529 | 225.3844 | 89.103 | 1.5457 | 39.278 | 60.7111 | 25.4116 | 2.4545 | 39.1615 | 96.1207 | 15.9552 | 21.3962 | 6.1428 | 156.8318 | 301.7759 | 232.0717 | 296.3567 | 203.9882
2012-10-30 01:16 | 25.9281 | 323.1073 | 253.4093 | 317.7644 | 226.124 | 89.1151 | 1.5483 | 39.3458 | 60.919 | 25.4123 | 2.4586 | 39.2284 | 96.4478 | 15.9555 | 21.4083 | 6.1436 | 157.3668 | 301.699 | 232.001 | 296.3561 | 204.7157
2012-10-30 01:16 | 25.9251 | 323.1463 | 253.4259 | 317.769 | 225.9518 | 89.1105 | 1.5493 | 39.3713 | 60.9968 | 25.4128 | 2.4602 | 39.2541 | 96.5731 | 15.9556 | 21.4414 | 6.1433 | 157.5699 | 301.7049 | 231.9845 | 296.3276 | 204.5104
2012-10-30 01:16 | 25.9153 | 323.1326 | 253.4043 | 317.7248 | 225.8869 | 89.1157 | 1.5497 | 39.3837 | 61.0343 | 25.4131 | 2.461 | 39.2668 | 96.6349 | 15.9557 | 21.4075 | 6.1436 | 157.6692 | 301.7251 | 231.9968 | 296.3173 | 204.4794
2012-10-30 01:16 | 25.9281 | 323.1988 | 253.3898 | 317.7338 | 226.4347 | 89.129 | 1.5501 | 39.3929 | 61.0623 | 25.4134 | 2.4615 | 39.2759 | 96.6796 | 15.9558 | 21.4163 | 6.1446 | 157.7418 | 301.7825 | 231.9735 | 296.3175 | 205.0184
2012-10-30 01:16 | 25.9243 | 323.1722 | 253.4251 | 317.73 | 226.7535 | 89.1278 | 1.5503 | 39.3987 | 61.0802 | 25.4134 | 2.4619 | 39.2816 | 96.7074 | 15.9558 | 21.4205 | 6.1445 | 157.7876 | 301.7517 | 232.0046 | 296.3095 | 205.333
2012-10-30 01:16 | 25.9191 | 323.1174 | 253.5019 | 317.6523 | 226.9175 | 89.114 | 1.5504 | 39.4021 | 61.0907 | 25.4135 | 2.4621 | 39.285 | 96.7232 | 15.9559 | 21.4515 | 6.1435 | 157.8138 | 301.6659 | 232.0504 | 296.2008 | 205.466
2012-10-30 01:16 | 25.9243 | 323.1607 | 253.5618 | 317.7071 | 226.962 | 89.125 | 1.5506 | 39.4048 | 61.0993 | 25.4133 | 2.4623 | 39.2876 | 96.7362 | 15.9559 | 21.4165 | 6.1443 | 157.8356 | 301.7442 | 232.1453 | 296.2906 | 205.5455
2012-10-30 01:16 | 25.9258 | 323.1584 | 253.6961 | 317.6971 | 226.995 | 89.1226 | 1.5506 | 39.4053 | 61.1001 | 25.4137 | 2.4623 | 39.2882 | 96.7393 | 15.9559 | 21.4461 | 6.1441 | 157.8393 | 301.7123 | 232.25 | 296.251 | 205.5489
2012-10-30 01:16 | 25.9273 | 323.1599 | 253.6859 | 317.733 | 227.0004 | 89.1128 | 1.5506 | 39.4057 | 61.1013 | 25.4137 | 2.4623 | 39.2885 | 96.7411 | 15.9558 | 21.4476 | 6.1434 | 157.8424 | 301.7123 | 232.2383 | 296.2854 | 205.5528
2012-10-30 01:16 | 25.9296 | 323.1736 | 253.7155 | 317.6971 | 227.0263 | 89.1325 | 1.5506 | 39.4057 | 61.1017 | 25.4135 | 2.4623 | 39.2886 | 96.7413 | 15.9559 | 21.4459 | 6.1448 | 157.843 | 301.7277 | 232.2696 | 296.2512 | 205.5804
2012-10-30 01:16 | 25.9251 | 323.1386 | 253.7813 | 317.7461 | 227.1831 | 89.1226 | 1.5505 | 39.4047 | 61.0981 | 25.4138 | 2.4623 | 39.2875 | 96.7365 | 15.9558 | 21.4294 | 6.1441 | 157.8346 | 301.7092 | 232.3519 | 296.3167 | 205.7537
2012-10-30 01:16 | 25.9213 | 323.1158 | 253.8166 | 317.6888 | 227.2068 | 89.1273 | 1.5505 | 39.4025 | 61.0917 | 25.4135 | 2.4621 | 39.2854 | 96.725 | 15.956 | 21.4256 | 6.1444 | 157.8167 | 301.6902 | 232.391 | 296.2632 | 205.7812
2012-10-30 01:16 | 25.9281 | 323.1836 | 253.8703 | 317.7147 | 227.1113 | 89.1307 | 1.5504 | 39.4016 | 61.0891 | 25.4135 | 2.4621 | 39.2846 | 96.7218 | 15.9559 | 21.4043 | 6.1447 | 157.811 | 301.7793 | 232.466 | 296.3104 | 205.707
2012-10-30 01:16 | 25.9288 | 323.169 | 253.8749 | 317.6849 | 227.0058 | 89.1302 | 1.5504 | 39.4007 | 61.0861 | 25.4135 | 2.462 | 39.2834 | 96.7157 | 15.9559 | 21.417 | 6.1446 | 157.8019 | 301.752 | 232.4579 | 296.2679 | 205.5888
2012-10-30 01:17 | 25.9273 | 323.1713 | 253.9398 | 317.6948 | 227.2128 | 89.1238 | 1.5502 | 39.3956 | 61.0698 | 25.4138 | 2.4617 | 39.2786 | 96.691 | 15.956 | 21.4155 | 6.1442 | 157.7609 | 301.7558 | 232.5243 | 296.2793 | 205.7973
2012-10-30 01:17 | 25.9355 | 323.1376 | 254.0026 | 317.6763 | 227.1817 | 89.107 | 1.5501 | 39.3938 | 61.0641 | 25.4139 | 2.4615 | 39.2766 | 96.6814 | 15.9561 | 21.3836 | 6.143 | 157.7454 | 301.754 | 232.619 | 296.2927 | 205.7981
2012-10-30 01:17 | 25.9318 | 323.172 | 253.9599 | 317.7451 | 227.3196 | 89.1238 | 1.55 | 39.392 | 61.0583 | 25.414 | 2.4615 | 39.275 | 96.6736 | 15.956 | 21.3879 | 6.1442 | 157.7319 | 301.7841 | 232.572 | 296.3572 | 205.9317
2012-10-30 01:17 | 25.9355 | 323.1795 | 254.0378 | 317.703 | 227.4413 | 89.118 | 1.5499 | 39.3899 | 61.0514 | 25.4141 | 2.4613 | 39.273 | 96.6636 | 15.956 | 21.4277 | 6.1438 | 157.715 | 301.7518 | 232.6101 | 296.2753 | 206.0136
2012-10-30 01:17 | 25.9385 | 323.1673 | 254.0759 | 317.7366 | 227.2594 | 89.1169 | 1.5499 | 39.3884 | 61.0475 | 25.4138 | 2.4612 | 39.2713 | 96.656 | 15.9559 | 21.3866 | 6.1437 | 157.7035 | 301.7807 | 232.6893 | 296.35 | 205.8728
2012-10-30 01:17 | 25.9453 | 323.1665 | 254.0749 | 317.7205 | 227.1482 | 89.144 | 1.5498 | 39.3873 | 61.044 | 25.4138 | 2.4612 | 39.2701 | 96.6507 | 15.9558 | 21.4094 | 6.1456 | 157.6947 | 301.7571 | 232.6655 | 296.3111 | 205.7388
2012-10-30 01:17 | 25.9423 | 323.2016 | 254.0719 | 317.7136 | 227.3222 | 89.1382 | 1.5498 | 39.3861 | 61.04 | 25.4139 | 2.4611 | 39.269 | 96.6455 | 15.9558 | 21.4185 | 6.1452 | 157.6855 | 301.7831 | 232.6534 | 296.2951 | 205.9037
2012-10-30 01:17 | 25.9483 | 323.219 | 254.0857 | 317.7616 | 227.4934 | 89.129 | 1.5498 | 39.385 | 61.0371 | 25.4137 | 2.461 | 39.2677 | 96.6386 | 15.9559 | 21.4245 | 6.1446 | 157.6756 | 301.7945 | 232.6612 | 296.3371 | 206.0689
2012-10-30 01:17 | 25.9408 | 323.1848 | 254.1095 | 317.7007 | 227.4544 | 89.1359 | 1.5497 | 39.3825 | 61.0294 | 25.4137 | 2.4609 | 39.2654 | 96.6282 | 15.9557 | 21.4089 | 6.145 | 157.6577 | 301.7759 | 232.7006 | 296.2918 | 206.0455
2012-10-30 01:17 | 25.9438 | 323.1688 | 254.1203 | 317.7304 | 227.2962 | 89.1377 | 1.5496 | 39.3818 | 61.0273 | 25.4136 | 2.4608 | 39.2646 | 96.6237 | 15.9558 | 21.432 | 6.1452 | 157.6511 | 301.7368 | 232.6883 | 296.2984 | 205.8642
2012-10-30 01:17 | 25.94 | 323.1917 | 254.1399 | 317.7533 | 227.4536 | 89.1475 | 1.5496 | 39.3807 | 61.0238 | 25.4136 | 2.4608 | 39.2637 | 96.6198 | 15.9557 | 21.4162 | 6.1458 | 157.6436 | 301.7755 | 232.7237 | 296.3371 | 206.0374
2012-10-30 01:17 | 25.9393 | 323.2177 | 254.2017 | 317.7794 | 227.4608 | 89.14 | 1.5495 | 39.3784 | 61.0173 | 25.4135 | 2.4607 | 39.2614 | 96.6093 | 15.9556 | 21.4355 | 6.1453 | 157.6266 | 301.7822 | 232.7662 | 296.3439 | 206.0253
2012-10-30 01:17 | 25.9408 | 323.1581 | 254.1368 | 317.7618 | 227.4662 | 89.1382 | 1.5495 | 39.3772 | 61.013 | 25.4136 | 2.4606 | 39.2601 | 96.603 | 15.9555 | 21.4772 | 6.1452 | 157.616 | 301.6809 | 232.6596 | 296.2846 | 205.989
2012-10-30 01:17 | 25.9355 | 323.2139 | 254.2253 | 317.7565 | 227.3193 | 89.1284 | 1.5494 | 39.3768 | 61.0123 | 25.4135 | 2.4606 | 39.2598 | 96.6007 | 15.9557 | 21.4237 | 6.1445 | 157.6129 | 301.7902 | 232.8016 | 296.3328 | 205.8956
2012-10-30 01:17 | 25.9378 | 323.1742 | 254.1729 | 317.7473 | 227.2194 | 89.1417 | 1.5494 | 39.375 | 61.0057 | 25.4138 | 2.4604 | 39.2579 | 96.5912 | 15.9557 | 21.434 | 6.1454 | 157.5969 | 301.7402 | 232.7389 | 296.3133 | 205.7854
2012-10-30 01:17 | 25.946 | 323.2053 | 254.1655 | 317.7975 | 227.2158 | 89.1411 | 1.5493 | 39.3739 | 61.0026 | 25.4137 | 2.4604 | 39.2569 | 96.586 | 15.9557 | 21.386 | 6.1454 | 157.5886 | 301.8193 | 232.7795 | 296.4115 | 205.8298
2012-10-30 01:17 | 25.943 | 323.2328 | 254.1781 | 317.7678 | 227.4212 | 89.1504 | 1.5493 | 39.3722 | 60.9977 | 25.4136 | 2.4603 | 39.2551 | 96.5781 | 15.9556 | 21.3991 | 6.146 | 157.5758 | 301.8337 | 232.779 | 296.3687 | 206.0221
2012-10-30 01:17 | 25.9423 | 323.1978 | 254.1344 | 317.7556 | 227.3065 | 89.1463 | 1.5492 | 39.3711 | 60.994 | 25.4137 | 2.4602 | 39.254 | 96.5729 | 15.9556 | 21.3783 | 6.1457 | 157.5669 | 301.8195 | 232.7561 | 296.3773 | 205.9282
2012-10-30 01:17 | 25.9415 | 323.2008 | 254.1453 | 317.7816 | 227.2467 | 89.1487 | 1.5475 | 39.3264 | 60.8573 | 25.413 | 2.4573 | 39.2078 | 96.3467 | 15.9554 | 21.4177 | 6.1459 | 157.2041 | 301.7831 | 232.7276 | 296.3639 | 205.829
2012-10-30 01:17 | 25.9445 | 323.2076 | 254.1327 | 317.7693 | 227.3401 | 89.1371 | 1.5455 | 39.2732 | 60.6955 | 25.4119 | 2.4542 | 39.1566 | 96.0977 | 15.955 | 21.3886 | 6.1451 | 156.7932 | 301.819 | 232.7441 | 296.3807 | 205.9515
2012-10-30 01:17 | 25.9505 | 323.1869 | 254.1231 | 317.8211 | 227.5152 | 89.1336 | 1.5449 | 39.2595 | 60.6533 | 25.4118 | 2.4534 | 39.143 | 96.0319 | 15.9549 | 21.4267 | 6.1449 | 156.6853 | 301.7602 | 232.6964 | 296.3944 | 206.0885
2012-10-30 01:17 | 25.949 | 323.2007 | 254.106 | 317.8196 | 226.7821 | 89.1463 | 1.5447 | 39.2537 | 60.6359 | 25.4116 | 2.453 | 39.1371 | 96.0029 | 15.9549 | 21.4011 | 6.1457 | 156.6388 | 301.7996 | 232.7049 | 296.4185 | 205.381
2012-10-30 01:18 | 25.9415 | 323.2008 | 253.9891 | 317.7892 | 226.1492 | 89.1209 | 1.5448 | 39.2565 | 60.6444 | 25.4117 | 2.4532 | 39.1398 | 96.016 | 15.9549 | 21.4257 | 6.144 | 156.6604 | 301.7751 | 232.5634 | 296.3635 | 204.7235
2012-10-30 01:18 | 25.9393 | 323.1757 | 253.8385 | 317.7679 | 225.966 | 89.1151 | 1.5449 | 39.2575 | 60.647 | 25.4118 | 2.4532 | 39.1407 | 96.0204 | 15.9549 | 21.4155 | 6.1436 | 156.6674 | 301.7602 | 232.423 | 296.3524 | 204.5505
2012-10-30 01:18 | 25.9378 | 323.1666 | 253.6573 | 317.7855 | 225.937 | 89.1359 | 1.5449 | 39.2598 | 60.6537 | 25.412 | 2.4534 | 39.1432 | 96.0328 | 15.9548 | 21.4381 | 6.145 | 156.6865 | 301.7285 | 232.2192 | 296.3474 | 204.4989
2012-10-30 01:18 | 25.9266 | 323.1401 | 253.5992 | 317.782 | 226.6378 | 89.1099 | 1.5451 | 39.2638 | 60.6664 | 25.4119 | 2.4536 | 39.147 | 96.0512 | 15.9549 | 21.3747 | 6.1432 | 156.7177 | 301.7654 | 232.2245 | 296.4073 | 205.2631
2012-10-30 01:18 | 25.9348 | 323.156 | 253.6504 | 317.7481 | 226.7089 | 89.1128 | 1.5453 | 39.2699 | 60.6848 | 25.412 | 2.454 | 39.1531 | 96.0813 | 15.9549 | 21.3989 | 6.1434 | 156.7661 | 301.7571 | 232.2515 | 296.3492 | 205.31
2012-10-30 01:18 | 25.9348 | 323.1712 | 253.6152 | 317.7558 | 226.2094 | 89.1111 | 1.5464 | 39.2984 | 60.7728 | 25.4121 | 2.4559 | 39.1842 | 96.2337 | 15.955 | 21.419 | 6.1433 | 157.0065 | 301.7522 | 232.1962 | 296.3368 | 204.7904
2012-10-30 01:18 | 25.9288 | 323.15 | 253.5585 | 317.7307 | 227.1042 | 89.1215 | 1.5488 | 39.3601 | 60.961 | 25.4132 | 2.4596 | 39.243 | 96.5202 | 15.9553 | 21.4291 | 6.144 | 157.4812 | 301.7209 | 232.1294 | 296.3016 | 205.6751
2012-10-30 01:18 | 25.9161 | 323.1907 | 253.6551 | 317.7409 | 226.2693 | 89.1117 | 1.5493 | 39.3733 | 61.0008 | 25.4137 | 2.4603 | 39.2559 | 96.5825 | 15.9555 | 21.4324 | 6.1434 | 157.5833 | 301.7583 | 232.2227 | 296.3085 | 204.8369
2012-10-30 01:18 | 25.9191 | 323.1822 | 253.541 | 317.7172 | 226.1661 | 89.1134 | 1.5498 | 39.3857 | 61.0394 | 25.4136 | 2.4611 | 39.2683 | 96.644 | 15.9555 | 21.4113 | 6.1435 | 157.6833 | 301.7709 | 232.1297 | 296.3059 | 204.7548
2012-10-30 01:18 | 25.9183 | 323.2119 | 253.5323 | 317.7278 | 226.543 | 89.1296 | 1.5501 | 39.395 | 61.0681 | 25.4137 | 2.4617 | 39.2778 | 96.6908 | 15.9555 | 21.4707 | 6.1446 | 157.7588 | 301.7412 | 232.0616 | 296.2571 | 205.0723
2012-10-30 01:18 | 25.9146 | 323.1892 | 253.513 | 317.7241 | 226.8815 | 89.1296 | 1.5503 | 39.4001 | 61.0835 | 25.4139 | 2.462 | 39.2829 | 96.7154 | 15.9556 | 21.447 | 6.1446 | 157.7989 | 301.7422 | 232.066 | 296.2771 | 205.4345
2012-10-30 01:18 | 25.9071 | 323.1664 | 253.5563 | 317.6937 | 226.9369 | 89.1359 | 1.5504 | 39.4018 | 61.0884 | 25.414 | 2.4621 | 39.2846 | 96.7236 | 15.9556 | 21.4475 | 6.145 | 157.812 | 301.7189 | 232.1088 | 296.2462 | 205.4894
2012-10-30 01:18 | 25.9116 | 323.2052 | 253.6194 | 317.6906 | 226.9414 | 89.1325 | 1.5504 | 39.4024 | 61.0907 | 25.4139 | 2.4622 | 39.2852 | 96.7269 | 15.9555 | 21.4078 | 6.1448 | 157.8175 | 301.7974 | 232.2116 | 296.2828 | 205.5336
2012-10-30 01:18 | 25.9123 | 323.1563 | 253.6201 | 317.6951 | 226.9815 | 89.1163 | 1.5505 | 39.4031 | 61.0925 | 25.414 | 2.4622 | 39.2859 | 96.7301 | 15.9556 | 21.4286 | 6.1437 | 157.8226 | 301.7277 | 232.1915 | 296.2665 | 205.5529
2012-10-30 01:18 | 25.9093 | 323.18 | 253.6522 | 317.6615 | 226.8801 | 89.1128 | 1.5504 | 39.4026 | 61.0908 | 25.414 | 2.4622 | 39.2853 | 96.727 | 15.9556 | 21.3935 | 6.1434 | 157.8178 | 301.7865 | 232.2587 | 296.268 | 205.4866
2012-10-30 01:18 | 25.9049 | 323.1718 | 253.6947 | 317.6724 | 226.7813 | 89.1198 | 1.5505 | 39.4033 | 61.0929 | 25.414 | 2.4622 | 39.2861 | 96.7315 | 15.9555 | 21.4332 | 6.1439 | 157.8244 | 301.7386 | 232.2615 | 296.2392 | 205.3481
2012-10-30 01:18 | 25.9108 | 323.1701 | 253.6928 | 317.6783 | 226.7321 | 89.1267 | 1.5504 | 39.4025 | 61.0909 | 25.4139 | 2.4622 | 39.2853 | 96.7271 | 15.9555 | 21.4191 | 6.1444 | 157.818 | 301.751 | 232.2737 | 296.2592 | 205.313
2012-10-30 01:18 | 25.9079 | 323.1443 | 253.7016 | 317.6754 | 226.9417 | 89.1261 | 1.5504 | 39.4014 | 61.0865 | 25.4143 | 2.4621 | 39.2842 | 96.721 | 15.9557 | 21.3961 | 6.1444 | 157.8076 | 301.7482 | 232.3055 | 296.2793 | 205.5456
2012-10-30 01:18 | 25.9101 | 323.177 | 253.7273 | 317.6814 | 227.1523 | 89.114 | 1.5503 | 39.3996 | 61.081 | 25.4143 | 2.462 | 39.2825 | 96.7125 | 15.9557 | 21.4023 | 6.1435 | 157.7935 | 301.7747 | 232.325 | 296.2791 | 205.75
2012-10-30 01:18 | 25.9086 | 323.1298 | 253.7804 | 317.6685 | 227.2688 | 89.1198 | 1.5503 | 39.3996 | 61.0809 | 25.4143 | 2.462 | 39.2824 | 96.7122 | 15.9557 | 21.3727 | 6.1439 | 157.7931 | 301.7571 | 232.4077 | 296.2958 | 205.8961
2012-10-30 01:18 | 25.9093 | 323.1381 | 253.8202 | 317.6921 | 227.2695 | 89.1388 | 1.5502 | 39.3982 | 61.0767 | 25.4143 | 2.4619 | 39.281 | 96.7058 | 15.9556 | 21.3815 | 6.1452 | 157.7824 | 301.7566 | 232.4387 | 296.3106 | 205.888
2012-10-30 01:18 | 25.9191 | 323.1517 | 253.8456 | 317.6866 | 227.0394 | 89.1169 | 1.5502 | 39.3972 | 61.0735 | 25.4143 | 2.4618 | 39.28 | 96.7005 | 15.9556 | 21.4354 | 6.1437 | 157.7739 | 301.7163 | 232.4102 | 296.2512 | 205.604
2012-10-30 01:19 | 25.9161 | 323.1716 | 253.8504 | 317.6798 | 226.9931 | 89.1169 | 1.5502 | 39.3968 | 61.0724 | 25.4142 | 2.4618 | 39.2797 | 96.6992 | 15.9556 | 21.4244 | 6.1437 | 157.7716 | 301.7472 | 232.426 | 296.2554 | 205.5687
2012-10-30 01:19 | 25.9138 | 323.1159 | 253.813 | 317.6699 | 226.9987 | 89.1307 | 1.5502 | 39.3958 | 61.0694 | 25.4141 | 2.4617 | 39.2786 | 96.6935 | 15.9557 | 21.402 | 6.1447 | 157.7629 | 301.7139 | 232.411 | 296.2679 | 205.5967
2012-10-30 01:19 | 25.9131 | 323.161 | 253.8045 | 317.6806 | 227.0531 | 89.1284 | 1.5501 | 39.3944 | 61.0653 | 25.4141 | 2.4617 | 39.2772 | 96.687 | 15.9556 | 21.3692 | 6.1445 | 157.7523 | 301.7918 | 232.4353 | 296.3114 | 205.6839
2012-10-30 01:19 | 25.9049 | 323.168 | 253.8314 | 317.661 | 227.1393 | 89.1198 | 1.55 | 39.3921 | 61.0582 | 25.4141 | 2.4615 | 39.2749 | 96.6759 | 15.9556 | 21.3931 | 6.1439 | 157.7341 | 301.7749 | 232.4383 | 296.2679 | 205.7462
2012-10-30 01:19 | 25.9168 | 323.1914 | 253.8707 | 317.7263 | 227.2888 | 89.1267 | 1.5499 | 39.3906 | 61.0531 | 25.4142 | 2.4614 | 39.2733 | 96.6682 | 15.9555 | 21.4451 | 6.1444 | 157.7214 | 301.7463 | 232.4256 | 296.2812 | 205.8437
2012-10-30 01:19 | 25.9176 | 323.2036 | 253.8949 | 317.6584 | 227.3526 | 89.1382 | 1.5499 | 39.3893 | 61.0492 | 25.4141 | 2.4614 | 39.2721 | 96.6629 | 15.9554 | 21.4179 | 6.1452 | 157.7122 | 301.7857 | 232.477 | 296.2405 | 205.9347
2012-10-30 01:19 | 25.9146 | 323.1586 | 253.9544 | 317.7012 | 227.0585 | 89.1273 | 1.5499 | 39.3883 | 61.0462 | 25.4141 | 2.4613 | 39.2712 | 96.6582 | 15.9555 | 21.3827 | 6.1444 | 157.7044 | 301.7759 | 232.5717 | 296.3185 | 205.6758
2012-10-30 01:19 | 25.9093 | 323.18 | 253.9061 | 317.6921 | 226.9116 | 89.1209 | 1.5498 | 39.3857 | 61.0384 | 25.414 | 2.4611 | 39.2686 | 96.6456 | 15.9555 | 21.4015 | 6.144 | 157.684 | 301.7785 | 232.5046 | 296.2906 | 205.5101
2012-10-30 01:19 | 25.9049 | 323.168 | 253.9369 | 317.6915 | 227.2022 | 89.1215 | 1.5497 | 39.3846 | 61.0353 | 25.414 | 2.4611 | 39.2675 | 96.6407 | 15.9554 | 21.4052 | 6.144 | 157.676 | 301.7628 | 232.5317 | 296.2863 | 205.797
2012-10-30 01:19 | 25.9049 | 323.2062 | 253.933 | 317.6953 | 227.2376 | 89.1302 | 1.5497 | 39.3834 | 61.0313 | 25.414 | 2.461 | 39.2662 | 96.6347 | 15.9553 | 21.4413 | 6.1446 | 157.666 | 301.7649 | 232.4917 | 296.254 | 205.7963
2012-10-30 01:19 | 25.8981 | 323.2489 | 253.9496 | 317.7229 | 227.0027 | 89.1365 | 1.5496 | 39.3817 | 61.0256 | 25.4142 | 2.4609 | 39.2647 | 96.6281 | 15.9551 | 21.3984 | 6.1451 | 157.6537 | 301.8505 | 232.5512 | 296.3245 | 205.6043
2012-10-30 01:19 | 25.9004 | 323.2093 | 253.9441 | 317.7214 | 226.9263 | 89.1221 | 1.5495 | 39.3803 | 61.0217 | 25.414 | 2.4608 | 39.2632 | 96.6206 | 15.9552 | 21.4127 | 6.1441 | 157.6423 | 301.7966 | 232.5314 | 296.3087 | 205.5136
2012-10-30 01:19 | 25.9004 | 323.236 | 253.9519 | 317.729 | 227.2095 | 89.1278 | 1.5494 | 39.3775 | 61.013 | 25.4141 | 2.4607 | 39.2605 | 96.6074 | 15.9552 | 21.4368 | 6.1445 | 157.6204 | 301.7992 | 232.5151 | 296.2922 | 205.7727
2012-10-30 01:19 | 25.8921 | 323.2315 | 253.971 | 317.6864 | 227.1225 | 89.1244 | 1.5493 | 39.3747 | 61.0047 | 25.414 | 2.4605 | 39.2578 | 96.594 | 15.9552 | 21.4164 | 6.1442 | 157.5987 | 301.8151 | 232.5546 | 296.27 | 205.7061
2012-10-30 01:19 | 25.8921 | 323.1934 | 253.9514 | 317.7245 | 226.9573 | 89.133 | 1.5493 | 39.3737 | 61.0023 | 25.4137 | 2.4604 | 39.2567 | 96.5889 | 15.9552 | 21.3602 | 6.1448 | 157.5911 | 301.8332 | 232.5912 | 296.3643 | 205.5971
2012-10-30 01:19 | 25.8929 | 323.2323 | 253.9679 | 317.7292 | 227.1666 | 89.1336 | 1.5469 | 39.3106 | 60.808 | 25.4132 | 2.4564 | 39.1913 | 96.2685 | 15.9549 | 21.4172 | 6.1449 | 157.0765 | 301.8151 | 232.5507 | 296.312 | 205.7494
2012-10-30 01:19 | 25.8929 | 323.2132 | 253.9718 | 317.7215 | 227.2767 | 89.1255 | 1.5453 | 39.269 | 60.681 | 25.4124 | 2.454 | 39.1522 | 96.0784 | 15.9546 | 21.3932 | 6.1443 | 156.7595 | 301.82 | 232.5786 | 296.3283 | 205.8835
2012-10-30 01:19 | 25.8862 | 323.176 | 253.9572 | 317.6843 | 226.6485 | 89.133 | 1.5449 | 39.26 | 60.6538 | 25.4123 | 2.4535 | 39.1436 | 96.0371 | 15.9545 | 21.3704 | 6.1448 | 156.6909 | 301.8056 | 232.5868 | 296.3139 | 205.2781
2012-10-30 01:19 | 25.8921 | 323.159 | 253.8304 | 317.6749 | 226.3201 | 89.1244 | 1.5448 | 39.2566 | 60.643 | 25.4124 | 2.4532 | 39.1401 | 96.0192 | 15.9546 | 21.3964 | 6.1442 | 156.6621 | 301.7626 | 232.434 | 296.2785 | 204.9237
2012-10-30 01:19 | 25.8966 | 323.1712 | 253.7528 | 317.6794 | 225.9508 | 89.1255 | 1.5448 | 39.2574 | 60.6464 | 25.412 | 2.4533 | 39.1408 | 96.023 | 15.9545 | 21.4169 | 6.1443 | 156.6695 | 301.7543 | 232.3359 | 296.2625 | 204.5339
2012-10-30 01:19 | 25.8891 | 323.1331 | 253.6281 | 317.6795 | 226.5216 | 89.114 | 1.545 | 39.2616 | 60.659 | 25.4121 | 2.4535 | 39.1448 | 96.0427 | 15.9546 | 21.4295 | 6.1435 | 156.7017 | 301.7036 | 232.1986 | 296.25 | 205.0921
2012-10-30 01:20 | 25.8839 | 323.1508 | 253.5917 | 317.682 | 226.1585 | 89.1163 | 1.5451 | 39.2642 | 60.6666 | 25.4123 | 2.4537 | 39.1472 | 96.055 | 15.9545 | 21.4082 | 6.1437 | 156.7217 | 301.7426 | 232.1835 | 296.2738 | 204.7503
2012-10-30 01:20 | 25.8847 | 323.1326 | 253.573 | 317.6904 | 226.4425 | 89.1042 | 1.5453 | 39.2698 | 60.6839 | 25.4123 | 2.454 | 39.1529 | 96.0823 | 15.9545 | 21.4772 | 6.1428 | 156.7662 | 301.6554 | 232.0958 | 296.2132 | 204.9653
2012-10-30 01:20 | 25.8862 | 323.1226 | 253.5081 | 317.6843 | 226.6721 | 89.1053 | 1.5455 | 39.2747 | 60.6988 | 25.4124 | 2.4543 | 39.158 | 96.1073 | 15.9545 | 21.4306 | 6.1429 | 156.8062 | 301.692 | 232.0775 | 296.2537 | 205.2415
2012-10-30 01:20 | 25.8772 | 323.1251 | 253.5576 | 317.6485 | 226.4232 | 89.1169 | 1.5474 | 39.3249 | 60.8527 | 25.4129 | 2.4573 | 39.2059 | 96.3417 | 15.9547 | 21.4336 | 6.1437 | 157.1944 | 301.6915 | 232.124 | 296.2149 | 204.9896
2012-10-30 01:20 | 25.8727 | 323.1625 | 253.5141 | 317.6364 | 226.5996 | 89.1134 | 1.5491 | 39.3676 | 60.9831 | 25.4137 | 2.4601 | 39.25 | 96.5573 | 15.9549 | 21.401 | 6.1435 | 157.5403 | 301.7615 | 232.1131 | 296.2354 | 205.1986
2012-10-30 01:20 | 25.8659 | 323.1557 | 253.5581 | 317.6411 | 226.581 | 89.1186 | 1.5495 | 39.3802 | 61.0214 | 25.4141 | 2.4609 | 39.263 | 96.6205 | 15.955 | 21.4424 | 6.1438 | 157.6419 | 301.7133 | 232.1157 | 296.1987 | 205.1386
2012-10-30 01:20 | 25.8674 | 323.1305 | 253.5049 | 317.6082 | 225.6935 | 89.1215 | 1.5499 | 39.3901 | 61.0519 | 25.4141 | 2.4615 | 39.2729 | 96.6688 | 15.9551 | 21.476 | 6.144 | 157.7208 | 301.6545 | 232.0289 | 296.1322 | 204.2175
2012-10-30 01:20 | 25.8607 | 323.1581 | 253.3575 | 317.6397 | 226.0684 | 89.1359 | 1.5503 | 39.3991 | 61.0794 | 25.4143 | 2.462 | 39.2819 | 96.7132 | 15.9551 | 21.4412 | 6.145 | 157.7926 | 301.7169 | 231.9163 | 296.1985 | 204.6272
2012-10-30 01:20 | 25.8547 | 323.1636 | 253.3281 | 317.6108 | 226.1411 | 89.1215 | 1.5505 | 39.4038 | 61.094 | 25.4143 | 2.4623 | 39.2867 | 96.7364 | 15.9551 | 21.4513 | 6.144 | 157.8305 | 301.7123 | 231.8768 | 296.1595 | 204.6898
2012-10-30 01:20 | 25.8622 | 323.1978 | 253.4098 | 317.6297 | 226.8566 | 89.1388 | 1.5505 | 39.4051 | 61.0978 | 25.4144 | 2.4624 | 39.2878 | 96.7424 | 15.9551 | 21.4387 | 6.1452 | 157.8402 | 301.7591 | 231.9711 | 296.191 | 205.4179
2012-10-30 01:20 | 25.86 | 323.2185 | 253.4858 | 317.5779 | 227.0668 | 89.1319 | 1.5505 | 39.4053 | 61.0989 | 25.4142 | 2.4624 | 39.2882 | 96.744 | 15.9552 | 21.4285 | 6.1448 | 157.8429 | 301.79 | 232.0573 | 296.1494 | 205.6383
2012-10-30 01:20 | 25.8532 | 323.1697 | 253.518 | 317.6016 | 227.1112 | 89.1394 | 1.5505 | 39.4049 | 61.097 | 25.4144 | 2.4624 | 39.2877 | 96.7406 | 15.9553 | 21.4457 | 6.1453 | 157.8376 | 301.724 | 232.0723 | 296.1559 | 205.6655
2012-10-30 01:20 | 25.848 | 323.1454 | 253.6144 | 317.6308 | 226.988 | 89.1319 | 1.5503 | 39.4016 | 61.0861 | 25.4147 | 2.4621 | 39.2845 | 96.7239 | 15.9555 | 21.4526 | 6.1448 | 157.8101 | 301.6928 | 232.1618 | 296.1782 | 205.5354
2012-10-30 01:20 | 25.8577 | 323.1704 | 253.6319 | 317.652 | 227.0252 | 89.1302 | 1.5503 | 39.4003 | 61.0821 | 25.4147 | 2.462 | 39.283 | 96.7163 | 15.9555 | 21.4382 | 6.1446 | 157.7984 | 301.7322 | 232.1937 | 296.2138 | 205.587
2012-10-30 01:20 | 25.8555 | 323.1873 | 253.657 | 317.6154 | 226.9679 | 89.1434 | 1.5503 | 39.4009 | 61.0846 | 25.4144 | 2.4621 | 39.2836 | 96.7195 | 15.9554 | 21.432 | 6.1455 | 157.8041 | 301.7553 | 232.225 | 296.1834 | 205.5359
2012-10-30 01:20 | 25.8555 | 323.1835 | 253.6844 | 317.6192 | 226.9168 | 89.14 | 1.5503 | 39.4009 | 61.084 | 25.4147 | 2.4621 | 39.2837 | 96.7203 | 15.9554 | 21.4641 | 6.1453 | 157.8043 | 301.7194 | 232.2203 | 296.1551 | 205.4527
2012-10-30 01:20 | 25.8495 | 323.1584 | 253.7057 | 317.6323 | 226.9029 | 89.1354 | 1.5503 | 39.4011 | 61.0853 | 25.4145 | 2.4621 | 39.2839 | 96.7216 | 15.9554 | 21.422 | 6.145 | 157.8069 | 301.7364 | 232.2837 | 296.2103 | 205.4809
2012-10-30 01:20 | 25.8592 | 323.1414 | 253.7115 | 317.642 | 226.8969 | 89.1504 | 1.5503 | 39.4011 | 61.085 | 25.4145 | 2.4621 | 39.2837 | 96.7209 | 15.9553 | 21.4638 | 6.146 | 157.806 | 301.6776 | 232.2477 | 296.1782 | 205.4331
2012-10-30 01:20 | 25.8472 | 323.1485 | 253.6878 | 317.6682 | 226.6174 | 89.1382 | 1.5503 | 39.3996 | 61.0804 | 25.4145 | 2.462 | 39.2825 | 96.7153 | 15.9552 | 21.4718 | 6.1452 | 157.7957 | 301.6767 | 232.216 | 296.1964 | 205.1456
2012-10-30 01:20 | 25.86 | 323.1574 | 253.6381 | 317.6581 | 226.6656 | 89.1371 | 1.5502 | 39.3983 | 61.0757 | 25.4148 | 2.4619 | 39.281 | 96.7073 | 15.9553 | 21.4204 | 6.1451 | 157.783 | 301.737 | 232.2177 | 296.2377 | 205.2452
2012-10-30 01:20 | 25.8577 | 323.1361 | 253.6671 | 317.6481 | 226.7774 | 89.1498 | 1.5502 | 39.3975 | 61.0742 | 25.4145 | 2.4619 | 39.2804 | 96.7051 | 15.9552 | 21.4141 | 6.146 | 157.7792 | 301.722 | 232.253 | 296.234 | 205.3633
2012-10-30 01:20 | 25.863 | 323.1376 | 253.688 | 317.6878 | 226.9676 | 89.1406 | 1.5502 | 39.3962 | 61.0701 | 25.4144 | 2.4618 | 39.279 | 96.6985 | 15.9552 | 21.4555 | 6.1454 | 157.7686 | 301.6821 | 232.2325 | 296.2323 | 205.5121
2012-10-30 01:21 | 25.8637 | 323.1535 | 253.7199 | 317.6389 | 226.8148 | 89.1498 | 1.5502 | 39.3964 | 61.0703 | 25.4145 | 2.4618 | 39.2792 | 96.6993 | 15.9552 | 21.4482 | 6.146 | 157.7697 | 301.7053 | 232.2717 | 296.1907 | 205.3666
2012-10-30 01:21 | 25.857 | 323.101 | 253.7054 | 317.6894 | 226.7216 | 89.1463 | 1.5501 | 39.3946 | 61.0652 | 25.4144 | 2.4617 | 39.2775 | 96.6908 | 15.9552 | 21.4295 | 6.1457 | 157.756 | 301.6715 | 232.2759 | 296.2599 | 205.2921
2012-10-30 01:21 | 25.8659 | 323.1252 | 253.7182 | 317.7022 | 226.8367 | 89.1487 | 1.55 | 39.3933 | 61.0606 | 25.4146 | 2.4616 | 39.2761 | 96.6838 | 15.9552 | 21.4183 | 6.1459 | 157.7444 | 301.7069 | 232.2999 | 296.2839 | 205.4184
2012-10-30 01:21 | 25.8727 | 323.132 | 253.7211 | 317.6784 | 226.6783 | 89.1492 | 1.55 | 39.3933 | 61.0609 | 25.4144 | 2.4617 | 39.276 | 96.6837 | 15.9551 | 21.4211 | 6.1459 | 157.7446 | 301.7109 | 232.3 | 296.2573 | 205.2572
2012-10-30 01:21 | 25.8637 | 323.1459 | 253.7004 | 317.6809 | 227.3812 | 89.1394 | 1.55 | 39.3914 | 61.0552 | 25.4144 | 2.4616 | 39.2742 | 96.6758 | 15.955 | 21.4041 | 6.1453 | 157.731 | 301.7418 | 232.2963 | 296.2768 | 205.9771
2012-10-30 01:21 | 25.8644 | 323.1542 | 253.7714 | 317.693 | 227.3269 | 89.151 | 1.5499 | 39.39 | 61.0509 | 25.4145 | 2.4615 | 39.2728 | 96.6686 | 15.9551 | 21.4449 | 6.1461 | 157.7194 | 301.7093 | 232.3265 | 296.2481 | 205.882
2012-10-30 01:21 | 25.8607 | 323.1353 | 253.8888 | 317.7046 | 227.4805 | 89.1417 | 1.5499 | 39.3892 | 61.049 | 25.4142 | 2.4614 | 39.2721 | 96.6648 | 15.9551 | 21.4211 | 6.1454 | 157.7138 | 301.7142 | 232.4677 | 296.2835 | 206.0594
2012-10-30 01:21 | 25.8637 | 323.1802 | 253.9621 | 317.7305 | 227.3341 | 89.14 | 1.5498 | 39.3876 | 61.043 | 25.4145 | 2.4613 | 39.2704 | 96.6566 | 15.9551 | 21.4442 | 6.1453 | 157.6997 | 301.736 | 232.5179 | 296.2863 | 205.8899
2012-10-30 01:21 | 25.8622 | 323.1787 | 253.9606 | 317.7596 | 227.0572 | 89.1498 | 1.5497 | 39.3852 | 61.0357 | 25.4145 | 2.4612 | 39.268 | 96.6455 | 15.955 | 21.4708 | 6.146 | 157.6811 | 301.7079 | 232.4898 | 296.2888 | 205.5864
2012-10-30 01:21 | 25.8622 | 323.1291 | 253.9528 | 317.7557 | 227.0926 | 89.1475 | 1.5495 | 39.3809 | 61.0227 | 25.4144 | 2.4609 | 39.2639 | 96.625 | 15.955 | 21.4547 | 6.1458 | 157.6476 | 301.6744 | 232.4981 | 296.301 | 205.6379
2012-10-30 01:21 | 25.8622 | 323.1368 | 253.9801 | 317.7672 | 227.1674 | 89.1342 | 1.5495 | 39.3789 | 61.0167 | 25.4143 | 2.4608 | 39.2617 | 96.6146 | 15.955 | 21.4588 | 6.1449 | 157.6314 | 301.678 | 232.5213 | 296.3084 | 205.7086
2012-10-30 01:21 | 25.8615 | 323.1437 | 253.9677 | 317.7398 | 227.1824 | 89.1469 | 1.5494 | 39.3777 | 61.0132 | 25.4143 | 2.4607 | 39.2604 | 96.6077 | 15.955 | 21.4661 | 6.1458 | 157.6209 | 301.6776 | 232.5016 | 296.2737 | 205.7163
2012-10-30 01:21 | 25.8742 | 323.1678 | 254.0038 | 317.7525 | 227.2462 | 89.1492 | 1.5494 | 39.3766 | 61.0091 | 25.4145 | 2.4606 | 39.2596 | 96.6037 | 15.955 | 21.4667 | 6.1459 | 157.6128 | 301.7011 | 232.5371 | 296.2858 | 205.7795
2012-10-30 01:21 | 25.8674 | 323.1381 | 253.9931 | 317.7495 | 227.3496 | 89.1417 | 1.5494 | 39.3763 | 61.0084 | 25.4144 | 2.4606 | 39.2591 | 96.6008 | 15.9551 | 21.3877 | 6.1454 | 157.6092 | 301.7504 | 232.6054 | 296.3618 | 205.9619
2012-10-30 01:21 | 25.8727 | 323.174 | 254.0336 | 317.7166 | 227.3116 | 89.14 | 1.5493 | 39.3757 | 61.0061 | 25.4146 | 2.4606 | 39.2587 | 96.5992 | 15.9551 | 21.389 | 6.1453 | 157.6053 | 301.785 | 232.6446 | 296.3276 | 205.9226
2012-10-30 01:21 | 25.8652 | 323.155 | 254.0456 | 317.7435 | 227.4024 | 89.1463 | 1.5493 | 39.3751 | 61.0043 | 25.4145 | 2.4605 | 39.2578 | 96.5942 | 15.9551 | 21.4297 | 6.1457 | 157.5985 | 301.7253 | 232.6159 | 296.3138 | 205.9727
2012-10-30 01:21 | 25.8667 | 323.1603 | 254.1018 | 317.7602 | 227.3213 | 89.1527 | 1.5492 | 39.3721 | 60.9957 | 25.4143 | 2.4604 | 39.2551 | 96.5816 | 15.9551 | 21.4472 | 6.1462 | 157.5773 | 301.7131 | 232.6546 | 296.313 | 205.8741
2012-10-30 01:21 | 25.8764 | 323.151 | 254.0841 | 317.7929 | 227.3114 | 89.1533 | 1.5452 | 39.2688 | 60.679 | 25.4131 | 2.4542 | 39.1555 | 96.0949 | 15.9546 | 21.4288 | 6.1462 | 156.7739 | 301.7222 | 232.6553 | 296.3641 | 205.8826
2012-10-30 01:21 | 25.8772 | 323.1251 | 254.0381 | 317.7707 | 227.2886 | 89.1463 | 1.5452 | 39.268 | 60.6772 | 25.4128 | 2.4539 | 39.1514 | 96.0751 | 15.9545 | 21.4537 | 6.1457 | 156.7523 | 301.6714 | 232.5844 | 296.317 | 205.8349
2012-10-30 01:21 | 25.8749 | 323.1342 | 254.0631 | 317.7761 | 227.3964 | 89.14 | 1.5448 | 39.2574 | 60.6447 | 25.4127 | 2.4533 | 39.1407 | 96.0233 | 15.9544 | 21.4313 | 6.1453 | 156.6679 | 301.7029 | 232.6318 | 296.3448 | 205.9651
2012-10-30 01:21 | 25.8749 | 323.1151 | 254.0241 | 317.7646 | 226.9559 | 89.1469 | 1.5447 | 39.2548 | 60.6369 | 25.4125 | 2.4531 | 39.1381 | 96.0111 | 15.9543 | 21.4394 | 6.1458 | 156.648 | 301.6757 | 232.5847 | 296.3252 | 205.5165
2012-10-30 01:21 | 25.8749 | 323.1075 | 253.9616 | 317.7532 | 226.9873 | 89.1302 | 1.5448 | 39.2561 | 60.641 | 25.4126 | 2.4532 | 39.1394 | 96.0179 | 15.9543 | 21.4554 | 6.1446 | 156.6589 | 301.6521 | 232.5062 | 296.2978 | 205.5319
2012-10-30 01:21 | 25.8667 | 323.1565 | 253.9612 | 317.6915 | 227.0578 | 89.129 | 1.5449 | 39.261 | 60.6562 | 25.4125 | 2.4535 | 39.144 | 96.0403 | 15.9542 | 21.4231 | 6.1446 | 156.6966 | 301.7334 | 232.5381 | 296.2684 | 205.6347
2012-10-30 01:22 | 25.8637 | 323.1192 | 253.8723 | 317.7611 | 226.508 | 89.1186 | 1.5451 | 39.2644 | 60.6657 | 25.4128 | 2.4537 | 39.1475 | 96.0577 | 15.9542 | 21.4201 | 6.1438 | 156.7235 | 301.6991 | 232.4522 | 296.341 | 205.0879
2012-10-30 01:22 | 25.8787 | 323.1227 | 253.7896 | 317.7531 | 226.8613 | 89.1359 | 1.5453 | 39.2714 | 60.6876 | 25.4128 | 2.4542 | 39.1546 | 96.0918 | 15.9544 | 21.4472 | 6.145 | 156.7794 | 301.6755 | 232.3424 | 296.3059 | 205.4141
2012-10-30 01:22 | 25.8704 | 323.1297 | 253.7305 | 317.7448 | 226.9789 | 89.1336 | 1.5455 | 39.275 | 60.6986 | 25.4128 | 2.4544 | 39.1582 | 96.1098 | 15.9543 | 21.4589 | 6.1449 | 156.8084 | 301.6708 | 232.2716 | 296.2859 | 205.52
2012-10-30 01:22 | 25.8682 | 323.1351 | 253.7244 | 317.7885 | 226.4417 | 89.1255 | 1.5475 | 39.3268 | 60.8573 | 25.4135 | 2.4575 | 39.2086 | 96.3557 | 15.9546 | 21.4367 | 6.1443 | 157.213 | 301.6984 | 232.2877 | 296.3518 | 205.005
2012-10-30 01:22 | 25.8562 | 323.1422 | 253.6968 | 317.7306 | 226.7011 | 89.1544 | 1.5491 | 39.3687 | 60.9854 | 25.4142 | 2.4602 | 39.2513 | 96.5646 | 15.9548 | 21.4608 | 6.1463 | 157.55 | 301.6814 | 232.236 | 296.2698 | 205.2403
2012-10-30 01:22 | 25.8442 | 323.1569 | 253.7004 | 317.7339 | 226.5947 | 89.1515 | 1.5496 | 39.3813 | 61.0243 | 25.4143 | 2.461 | 39.2641 | 96.627 | 15.9548 | 21.4327 | 6.1461 | 157.6513 | 301.7242 | 232.2677 | 296.3012 | 205.162
2012-10-30 01:22 | 25.851 | 323.1942 | 253.6447 | 317.7407 | 226.0272 | 89.1411 | 1.55 | 39.3928 | 61.059 | 25.4146 | 2.4617 | 39.2756 | 96.6843 | 15.9547 | 21.4837 | 6.1454 | 157.7433 | 301.7105 | 232.161 | 296.257 | 204.5435
2012-10-30 01:22 | 25.8457 | 323.1813 | 253.5418 | 317.7163 | 226.6238 | 89.133 | 1.5503 | 39.4 | 61.0816 | 25.4145 | 2.4621 | 39.2826 | 96.7189 | 15.9547 | 21.4423 | 6.1448 | 157.8005 | 301.739 | 232.0995 | 296.274 | 205.1815
2012-10-30 01:22 | 25.836 | 323.164 | 253.5516 | 317.7525 | 226.7478 | 89.1556 | 1.5504 | 39.4027 | 61.0902 | 25.4144 | 2.4623 | 39.2854 | 96.7324 | 15.9547 | 21.4326 | 6.1464 | 157.8226 | 301.7314 | 232.119 | 296.3199 | 205.3152
2012-10-30 01:22 | 25.833 | 323.161 | 253.5955 | 317.7342 | 226.8628 | 89.1319 | 1.5505 | 39.4051 | 61.0971 | 25.4146 | 2.4624 | 39.2877 | 96.7434 | 15.9548 | 21.4215 | 6.1448 | 157.8405 | 301.7395 | 232.174 | 296.3127 | 205.4413
2012-10-30 01:22 | 25.8308 | 323.1778 | 253.6402 | 317.6785 | 227.2186 | 89.1469 | 1.5505 | 39.4053 | 61.0976 | 25.4147 | 2.4624 | 39.2881 | 96.7449 | 15.9549 | 21.4314 | 6.1458 | 157.8425 | 301.7464 | 232.2088 | 296.2471 | 205.7872
2012-10-30 01:22 | 25.833 | 323.18 | 253.697 | 317.7227 | 227.1264 | 89.1458 | 1.5505 | 39.4055 | 61.0979 | 25.4148 | 2.4624 | 39.2882 | 96.7452 | 15.9549 | 21.4255 | 6.1457 | 157.8431 | 301.7545 | 232.2715 | 296.2972 | 205.7009
2012-10-30 01:22 | 25.83 | 323.1847 | 253.7917 | 317.7006 | 227.1902 | 89.1654 | 1.5505 | 39.4047 | 61.0958 | 25.4147 | 2.4624 | 39.2875 | 96.7408 | 15.9551 | 21.3904 | 6.1471 | 157.8366 | 301.7943 | 232.4013 | 296.3102 | 205.7998
2012-10-30 01:22 | 25.8293 | 323.1763 | 253.8222 | 317.7037 | 227.2721 | 89.1614 | 1.5504 | 39.404 | 61.0927 | 25.4151 | 2.4624 | 39.2868 | 96.7385 | 15.9549 | 21.3496 | 6.1468 | 157.8312 | 301.8267 | 232.4726 | 296.3541 | 205.9225
2012-10-30 01:22 | 25.8323 | 323.1793 | 253.8252 | 317.6991 | 227.2122 | 89.151 | 1.5504 | 39.4028 | 61.0896 | 25.4148 | 2.4623 | 39.2856 | 96.7318 | 15.955 | 21.3887 | 6.1461 | 157.8214 | 301.7906 | 232.4365 | 296.3104 | 205.8235
2012-10-30 01:22 | 25.8345 | 323.1663 | 253.8899 | 317.7128 | 227.238 | 89.1567 | 1.5504 | 39.4024 | 61.0876 | 25.4151 | 2.4622 | 39.2851 | 96.7292 | 15.955 | 21.3909 | 6.1465 | 157.8168 | 301.7754 | 232.499 | 296.3219 | 205.8471
2012-10-30 01:22 | 25.8308 | 323.1435 | 253.8511 | 317.7129 | 227.1832 | 89.1585 | 1.5503 | 39.401 | 61.0831 | 25.4151 | 2.4622 | 39.2838 | 96.7228 | 15.9551 | 21.3752 | 6.1466 | 157.8059 | 301.7683 | 232.4759 | 296.3377 | 205.808
2012-10-30 01:22 | 25.8308 | 323.2045 | 253.9175 | 317.7091 | 227.0022 | 89.1481 | 1.5503 | 39.4002 | 61.0816 | 25.4148 | 2.4621 | 39.2829 | 96.7186 | 15.955 | 21.3953 | 6.1459 | 157.8002 | 301.8092 | 232.5222 | 296.3138 | 205.6069
2012-10-30 01:22 | 25.8285 | 323.187 | 253.8527 | 317.7106 | 226.8544 | 89.1527 | 1.5502 | 39.3988 | 61.0767 | 25.415 | 2.462 | 39.2817 | 96.7132 | 15.9549 | 21.3849 | 6.1462 | 157.7899 | 301.8021 | 232.4678 | 296.3257 | 205.4695
2012-10-30 01:22 | 25.833 | 323.1877 | 253.8611 | 317.7304 | 227.032 | 89.1481 | 1.5501 | 39.3969 | 61.0711 | 25.4149 | 2.4619 | 39.2795 | 96.7024 | 15.9549 | 21.4336 | 6.1459 | 157.7735 | 301.7541 | 232.4275 | 296.2968 | 205.5984
2012-10-30 01:22 | 25.8338 | 323.1846 | 253.8814 | 317.7693 | 227.2334 | 89.1521 | 1.5501 | 39.3952 | 61.0652 | 25.4152 | 2.4618 | 39.2782 | 96.6967 | 15.9548 | 21.4183 | 6.1461 | 157.7619 | 301.7663 | 232.4631 | 296.351 | 205.8151
2012-10-30 01:22 | 25.842 | 323.1623 | 253.9013 | 317.7508 | 227.0488 | 89.1625 | 1.55 | 39.3942 | 61.0628 | 25.4148 | 2.4617 | 39.2769 | 96.6897 | 15.9549 | 21.3904 | 6.1469 | 157.7525 | 301.7719 | 232.5109 | 296.3604 | 205.6584
2012-10-30 01:22 | 25.8472 | 323.2362 | 253.9182 | 317.7522 | 226.9282 | 89.1475 | 1.55 | 39.3926 | 61.0578 | 25.4149 | 2.4617 | 39.2755 | 96.6841 | 15.9547 | 21.3515 | 6.1458 | 157.742 | 301.8847 | 232.5667 | 296.4007 | 205.5767
2012-10-30 01:23 | 25.8375 | 323.1883 | 253.9164 | 317.7272 | 227.064 | 89.1533 | 1.55 | 39.3918 | 61.0561 | 25.4146 | 2.4616 | 39.2746 | 96.6797 | 15.9547 | 21.3779 | 6.1462 | 157.7359 | 301.8104 | 232.5385 | 296.3493 | 205.6861
2012-10-30 01:23 | 25.845 | 323.1501 | 253.8497 | 317.7156 | 227.1305 | 89.1596 | 1.5499 | 39.3902 | 61.0504 | 25.4148 | 2.4615 | 39.2731 | 96.6724 | 15.9547 | 21.3733 | 6.1467 | 157.7229 | 301.7768 | 232.4764 | 296.3423 | 205.7572
2012-10-30 01:23 | 25.8413 | 323.1387 | 253.9045 | 317.7616 | 227.1347 | 89.1411 | 1.5498 | 39.3886 | 61.0457 | 25.4148 | 2.4614 | 39.2714 | 96.6646 | 15.9546 | 21.3977 | 6.1454 | 157.7103 | 301.741 | 232.5068 | 296.3639 | 205.737
2012-10-30 01:23 | 25.8465 | 323.1935 | 253.9254 | 317.7553 | 227.1713 | 89.1556 | 1.5498 | 39.3874 | 61.0421 | 25.4147 | 2.4613 | 39.2702 | 96.6576 | 15.9548 | 21.4029 | 6.1464 | 157.6998 | 301.7906 | 232.5225 | 296.3524 | 205.7684
2012-10-30 01:23 | 25.86 | 323.1765 | 253.9467 | 317.7841 | 227.1848 | 89.1544 | 1.5497 | 39.3866 | 61.039 | 25.4149 | 2.4613 | 39.2694 | 96.6533 | 15.9548 | 21.3803 | 6.1463 | 157.6924 | 301.7962 | 232.5664 | 296.4038 | 205.8045
2012-10-30 01:23 | 25.851 | 323.1523 | 253.9611 | 317.7484 | 227.168 | 89.1591 | 1.5497 | 39.386 | 61.0371 | 25.4149 | 2.4613 | 39.269 | 96.6512 | 15.9548 | 21.3713 | 6.1466 | 157.6882 | 301.781 | 232.5898 | 296.3771 | 205.7967
2012-10-30 01:23 | 25.857 | 323.1926 | 253.9749 | 317.72 | 227.2998 | 89.1452 | 1.5496 | 39.3844 | 61.032 | 25.4151 | 2.4611 | 39.2672 | 96.6424 | 15.9548 | 21.3693 | 6.1457 | 157.6744 | 301.8233 | 232.6056 | 296.3507 | 205.9305
2012-10-30 01:23 | 25.8607 | 323.2039 | 253.9942 | 317.781 | 227.1816 | 89.1637 | 1.5495 | 39.3818 | 61.0231 | 25.4153 | 2.461 | 39.2644 | 96.6288 | 15.9548 | 21.385 | 6.1469 | 157.6519 | 301.8189 | 232.6092 | 296.396 | 205.7966
2012-10-30 01:23 | 25.8487 | 323.1652 | 253.9783 | 317.7766 | 227.2601 | 89.1614 | 1.5494 | 39.3791 | 61.0157 | 25.4149 | 2.4608 | 39.2619 | 96.6159 | 15.9549 | 21.4212 | 6.1468 | 157.6316 | 301.744 | 232.5571 | 296.3554 | 205.8389
2012-10-30 01:23 | 25.86 | 323.1841 | 253.9974 | 317.7879 | 227.2006 | 89.1515 | 1.5494 | 39.3787 | 61.014 | 25.4151 | 2.4608 | 39.2615 | 96.6135 | 15.955 | 21.4325 | 6.1461 | 157.6275 | 301.7516 | 232.5649 | 296.3554 | 205.7681
2012-10-30 01:23 | 25.854 | 323.1781 | 254.0071 | 317.7323 | 227.1592 | 89.1533 | 1.5494 | 39.3774 | 61.0101 | 25.4151 | 2.4607 | 39.2601 | 96.6072 | 15.9549 | 21.4024 | 6.1462 | 157.6173 | 301.7757 | 232.6047 | 296.3299 | 205.7568
2012-10-30 01:23 | 25.8622 | 323.1902 | 254.0465 | 317.7901 | 227.3719 | 89.1567 | 1.5493 | 39.3758 | 61.0052 | 25.4151 | 2.4606 | 39.2588 | 96.6005 | 15.955 | 21.4588 | 6.1465 | 157.6057 | 301.7314 | 232.5877 | 296.3313 | 205.9131
2012-10-30 01:23 | 25.8652 | 323.2008 | 254.0534 | 317.7893 | 227.4181 | 89.1487 | 1.5493 | 39.3751 | 61.0025 | 25.4154 | 2.4606 | 39.2581 | 96.5968 | 15.955 | 21.4457 | 6.1459 | 157.5993 | 301.7551 | 232.6077 | 296.3436 | 205.9724
2012-10-30 01:23 | 25.8644 | 323.1771 | 254.1151 | 317.7579 | 227.4724 | 89.1591 | 1.5492 | 39.3734 | 60.9981 | 25.415 | 2.4604 | 39.2563 | 96.5882 | 15.955 | 21.4088 | 6.1466 | 157.5863 | 301.7683 | 232.7063 | 296.3491 | 206.0636
2012-10-30 01:23 | 25.8644 | 323.2191 | 254.1151 | 317.7999 | 227.5393 | 89.1539 | 1.5491 | 39.3711 | 60.9907 | 25.4151 | 2.4603 | 39.2541 | 96.5765 | 15.955 | 21.469 | 6.1463 | 157.5672 | 301.7501 | 232.6461 | 296.3309 | 206.0703
2012-10-30 01:23 | 25.8585 | 323.1674 | 254.1756 | 317.7864 | 227.3996 | 89.1556 | 1.548 | 39.3428 | 60.9037 | 25.4148 | 2.4584 | 39.2238 | 96.4296 | 15.9547 | 21.431 | 6.1464 | 157.3333 | 301.7364 | 232.7446 | 296.3554 | 205.9686
2012-10-30 01:23 | 25.8674 | 323.142 | 254.1962 | 317.7991 | 227.2866 | 89.133 | 1.5454 | 39.2746 | 60.6956 | 25.4136 | 2.4544 | 39.1581 | 96.1094 | 15.9543 | 21.4479 | 6.1448 | 156.805 | 301.6941 | 232.7483 | 296.3512 | 205.8387
2012-10-30 01:23 | 25.857 | 323.1697 | 254.1311 | 317.804 | 227.3195 | 89.1417 | 1.5449 | 39.2606 | 60.6524 | 25.4136 | 2.4535 | 39.1438 | 96.0394 | 15.9543 | 21.4816 | 6.1454 | 156.6917 | 301.6881 | 232.6495 | 296.3224 | 205.8379
2012-10-30 01:23 | 25.857 | 323.1811 | 254.0999 | 317.7582 | 226.4345 | 89.1354 | 1.5446 | 39.253 | 60.6294 | 25.4134 | 2.453 | 39.1363 | 96.0023 | 15.9543 | 21.5017 | 6.145 | 156.6317 | 301.6794 | 232.5982 | 296.2565 | 204.9328
2012-10-30 01:23 | 25.8525 | 323.1271 | 253.8806 | 317.7766 | 225.6078 | 89.1371 | 1.5445 | 39.2515 | 60.6245 | 25.4135 | 2.453 | 39.1347 | 95.9956 | 15.9541 | 21.4932 | 6.1451 | 156.62 | 301.6339 | 232.3874 | 296.2834 | 204.1146
2012-10-30 01:23 | 25.857 | 323.1583 | 253.8148 | 317.8155 | 226.1119 | 89.1406 | 1.5444 | 39.2494 | 60.6183 | 25.4134 | 2.4528 | 39.1327 | 95.9856 | 15.9541 | 21.4576 | 6.1454 | 156.6039 | 301.7007 | 232.3572 | 296.3579 | 204.6543
2012-10-30 01:23 | 25.8517 | 323.1453 | 253.7275 | 317.7682 | 226.7242 | 89.1215 | 1.5446 | 39.2535 | 60.6307 | 25.4135 | 2.4531 | 39.1368 | 96.0059 | 15.9541 | 21.4804 | 6.144 | 156.6367 | 301.6649 | 232.2471 | 296.2878 | 205.2438
2012-10-30 01:24 | 25.8442 | 323.1226 | 253.7121 | 317.7721 | 226.2289 | 89.1313 | 1.5449 | 39.2615 | 60.6557 | 25.4133 | 2.4536 | 39.1447 | 96.0448 | 15.9541 | 21.4247 | 6.1447 | 156.7006 | 301.6979 | 232.2874 | 296.3474 | 204.8042
2012-10-30 01:24 | 25.8457 | 323.1508 | 253.5574 | 317.7851 | 226.018 | 89.129 | 1.5452 | 39.2678 | 60.6747 | 25.4136 | 2.454 | 39.1511 | 96.0761 | 15.9541 | 21.4302 | 6.1446 | 156.7508 | 301.7206 | 232.1272 | 296.3549 | 204.5878
2012-10-30 01:24 | 25.8457 | 323.1508 | 253.4441 | 317.7392 | 226.2107 | 89.1221 | 1.5455 | 39.277 | 60.7033 | 25.4134 | 2.4544 | 39.1576 | 96.1081 | 15.9541 | 21.4423 | 6.1441 | 156.8114 | 301.7085 | 232.0018 | 296.2969 | 204.7684
2012-10-30 01:24 | 25.845 | 323.1196 | 253.4434 | 317.7233 | 225.856 | 89.1209 | 1.5486 | 39.3569 | 60.9486 | 25.4142 | 2.4595 | 39.2394 | 96.5084 | 15.9544 | 21.4335 | 6.144 | 157.457 | 301.6861 | 232.0099 | 296.2898 | 204.4225
2012-10-30 01:24 | 25.8375 | 323.1731 | 253.3969 | 317.7119 | 226.3599 | 89.1232 | 1.5492 | 39.3724 | 60.9961 | 25.4145 | 2.4604 | 39.2549 | 96.5837 | 15.9545 | 21.4702 | 6.1442 | 157.5798 | 301.7029 | 231.9267 | 296.2417 | 204.8897
2012-10-30 01:24 | 25.8293 | 323.1229 | 253.4004 | 317.7076 | 226.2573 | 89.1325 | 1.5497 | 39.3845 | 61.033 | 25.4148 | 2.4612 | 39.2672 | 96.6437 | 15.9546 | 21.4459 | 6.1448 | 157.6767 | 301.677 | 231.9545 | 296.2617 | 204.8114
2012-10-30 01:24 | 25.8263 | 323.1657 | 253.3466 | 317.6931 | 226.2936 | 89.1284 | 1.55 | 39.3931 | 61.0591 | 25.415 | 2.4617 | 39.276 | 96.6867 | 15.9547 | 21.4509 | 6.1445 | 157.7458 | 301.7148 | 231.8957 | 296.2422 | 204.8427
2012-10-30 01:24 | 25.8263 | 323.139 | 253.3778 | 317.7046 | 226.506 | 89.1487 | 1.5502 | 39.3983 | 61.0749 | 25.4151 | 2.462 | 39.281 | 96.7114 | 15.9547 | 21.4269 | 6.1459 | 157.7863 | 301.7121 | 231.9509 | 296.2777 | 205.0791
2012-10-30 01:24 | 25.8173 | 323.1643 | 253.4196 | 317.665 | 226.7881 | 89.1273 | 1.5503 | 39.4013 | 61.0841 | 25.4152 | 2.4622 | 39.284 | 96.7267 | 15.9545 | 21.3858 | 6.1444 | 157.8108 | 301.7785 | 232.0338 | 296.2792 | 205.4023
2012-10-30 01:24 | 25.8128 | 323.1598 | 253.5011 | 317.6949 | 226.8544 | 89.1267 | 1.5503 | 39.4024 | 61.0874 | 25.4152 | 2.4623 | 39.2852 | 96.732 | 15.9546 | 21.3652 | 6.1444 | 157.8195 | 301.7946 | 232.1359 | 296.3297 | 205.4892
2012-10-30 01:24 | 25.8203 | 323.1788 | 253.5242 | 317.6527 | 226.8619 | 89.1313 | 1.5504 | 39.4028 | 61.0888 | 25.4151 | 2.4623 | 39.2854 | 96.7332 | 15.9546 | 21.3888 | 6.1447 | 157.822 | 301.79 | 232.1354 | 296.2639 | 205.4731
2012-10-30 01:24 | 25.8091 | 323.118 | 253.5872 | 317.6645 | 226.8429 | 89.1481 | 1.5503 | 39.4024 | 61.087 | 25.4154 | 2.4623 | 39.285 | 96.7315 | 15.9546 | 21.3655 | 6.1459 | 157.8185 | 301.7525 | 232.2217 | 296.299 | 205.4774
2012-10-30 01:24 | 25.8136 | 323.1416 | 253.58 | 317.6499 | 227.0165 | 89.1319 | 1.5504 | 39.4023 | 61.0875 | 25.4151 | 2.4623 | 39.285 | 96.7306 | 15.9547 | 21.37 | 6.1448 | 157.818 | 301.7716 | 232.21 | 296.2799 | 205.6465
2012-10-30 01:24 | 25.8113 | 323.1164 | 253.6636 | 317.6552 | 227.0299 | 89.1336 | 1.5503 | 39.402 | 61.0864 | 25.4151 | 2.4623 | 39.2848 | 96.73 | 15.9546 | 21.4119 | 6.1449 | 157.8164 | 301.7045 | 232.2517 | 296.2433 | 205.618
2012-10-30 01:24 | 25.8083 | 323.1172 | 253.6919 | 317.6637 | 226.7516 | 89.1562 | 1.5503 | 39.4005 | 61.0809 | 25.4155 | 2.4622 | 39.2833 | 96.722 | 15.9547 | 21.3968 | 6.1464 | 157.8029 | 301.7204 | 232.2951 | 296.2669 | 205.3548
2012-10-30 01:24 | 25.8083 | 323.1134 | 253.6684 | 317.6293 | 226.9168 | 89.1469 | 1.5503 | 39.4003 | 61.0806 | 25.4153 | 2.4622 | 39.2832 | 96.7221 | 15.9547 | 21.4049 | 6.1458 | 157.8027 | 301.7085 | 232.2635 | 296.2244 | 205.5119
2012-10-30 01:24 | 25.8121 | 323.1324 | 253.7152 | 317.6522 | 227.0347 | 89.133 | 1.5502 | 39.3989 | 61.0764 | 25.4153 | 2.4621 | 39.2816 | 96.7139 | 15.9547 | 21.4287 | 6.1448 | 157.7903 | 301.7037 | 232.2865 | 296.2235 | 205.606
2012-10-30 01:24 | 25.8121 | 323.0714 | 253.7191 | 317.6369 | 227.1133 | 89.1446 | 1.5501 | 39.3967 | 61.0695 | 25.4153 | 2.4619 | 39.2796 | 96.7044 | 15.9547 | 21.4046 | 6.1456 | 157.774 | 301.6668 | 232.3145 | 296.2323 | 205.7087
2012-10-30 01:24 | 25.8068 | 323.0928 | 253.7958 | 317.6622 | 227.0766 | 89.1458 | 1.5501 | 39.3962 | 61.0682 | 25.4152 | 2.4619 | 39.279 | 96.7011 | 15.9547 | 21.4555 | 6.1457 | 157.7693 | 301.6373 | 232.3403 | 296.2067 | 205.6211
2012-10-30 01:24 | 25.8136 | 323.111 | 253.8065 | 317.6651 | 227.0716 | 89.1556 | 1.5501 | 39.3956 | 61.0657 | 25.4154 | 2.4619 | 39.2782 | 96.6978 | 15.9546 | 21.4102 | 6.1464 | 157.7635 | 301.7008 | 232.3963 | 296.2549 | 205.6614
2012-10-30 01:24 | 25.8083 | 323.0829 | 253.7583 | 317.6369 | 227.0191 | 89.151 | 1.55 | 39.3943 | 61.0616 | 25.4154 | 2.4618 | 39.2771 | 96.6922 | 15.9547 | 21.3888 | 6.1461 | 157.7538 | 301.6941 | 232.3695 | 296.2481 | 205.6303
2012-10-30 01:24 | 25.8023 | 323.115 | 253.7718 | 317.65 | 227.0445 | 89.144 | 1.55 | 39.3933 | 61.0588 | 25.4153 | 2.4617 | 39.2761 | 96.6869 | 15.9547 | 21.4189 | 6.1456 | 157.7457 | 301.6961 | 232.3529 | 296.2311 | 205.6256
2012-10-30 01:25 | 25.8031 | 323.0548 | 253.8195 | 317.6546 | 227.0965 | 89.1585 | 1.5499 | 39.3913 | 61.0531 | 25.4152 | 2.4616 | 39.2742 | 96.678 | 15.9547 | 21.4077 | 6.1466 | 157.7311 | 301.6471 | 232.4118 | 296.2469 | 205.6888
2012-10-30 01:25 | 25.8046 | 323.0906 | 253.8483 | 317.66 | 227.1176 | 89.1406 | 1.5498 | 39.3899 | 61.0485 | 25.4153 | 2.4615 | 39.2727 | 96.671 | 15.9546 | 21.4012 | 6.1454 | 157.7194 | 301.6894 | 232.4471 | 296.2588 | 205.7164
2012-10-30 01:25 | 25.8106 | 323.1042 | 253.8543 | 317.6469 | 227.0607 | 89.1498 | 1.5498 | 39.3888 | 61.0448 | 25.4154 | 2.4614 | 39.2714 | 96.6645 | 15.9546 | 21.3951 | 6.146 | 157.7093 | 301.7091 | 232.4592 | 296.2518 | 205.6656
2012-10-30 01:25 | 25.8166 | 323.1179 | 253.8916 | 317.6529 | 227.2201 | 89.1452 | 1.5497 | 39.3875 | 61.0406 | 25.4155 | 2.4614 | 39.2702 | 96.6586 | 15.9546 | 21.3811 | 6.1457 | 157.6992 | 301.7368 | 232.5105 | 296.2718 | 205.839
2012-10-30 01:25 | 25.8091 | 323.1218 | 253.8567 | 317.6492 | 227.1025 | 89.1498 | 1.5497 | 39.3862 | 61.0368 | 25.4153 | 2.4613 | 39.269 | 96.6528 | 15.9546 | 21.3655 | 6.146 | 157.6896 | 301.7563 | 232.4912 | 296.2837 | 205.737
2012-10-30 01:25 | 25.8128 | 323.0988 | 253.8956 | 317.6834 | 227.1219 | 89.1573 | 1.5497 | 39.3851 | 61.0341 | 25.4151 | 2.4612 | 39.2679 | 96.6479 | 15.9545 | 21.3532 | 6.1465 | 157.682 | 301.7456 | 232.5424 | 296.3302 | 205.7687
2012-10-30 01:25 | 25.8248 | 323.1299 | 253.8841 | 317.6954 | 227.2952 | 89.1562 | 1.5497 | 39.3845 | 61.0323 | 25.415 | 2.4612 | 39.2673 | 96.6451 | 15.9545 | 21.3732 | 6.1464 | 157.6774 | 301.7567 | 232.5109 | 296.3222 | 205.922
2012-10-30 01:25 | 25.8203 | 323.1216 | 253.9304 | 317.6642 | 227.2867 | 89.1544 | 1.5496 | 39.3824 | 61.0258 | 25.4151 | 2.4611 | 39.2652 | 96.6351 | 15.9544 | 21.4088 | 6.1463 | 157.6609 | 301.7128 | 232.5216 | 296.2554 | 205.8779
2012-10-30 01:25 | 25.8225 | 323.1505 | 253.9053 | 317.6473 | 227.2299 | 89.151 | 1.5495 | 39.3808 | 61.0201 | 25.4153 | 2.461 | 39.2636 | 96.6265 | 15.9545 | 21.3308 | 6.1461 | 157.6466 | 301.8197 | 232.5745 | 296.3165 | 205.8991
2012-10-30 01:25 | 25.8173 | 323.13 | 253.9665 | 317.6765 | 227.1697 | 89.1585 | 1.5494 | 39.3788 | 61.0136 | 25.4155 | 2.4608 | 39.2617 | 96.6165 | 15.9546 | 21.3537 | 6.1466 | 157.6301 | 301.7763 | 232.6128 | 296.3228 | 205.816
2012-10-30 01:25 | 25.8255 | 323.1306 | 254.0059 | 317.6847 | 227.0717 | 89.1648 | 1.5493 | 39.3775 | 61.0094 | 25.4156 | 2.4608 | 39.2604 | 96.6112 | 15.9545 | 21.3498 | 6.147 | 157.6206 | 301.7808 | 232.6561 | 296.3349 | 205.7219
2012-10-30 01:25 | 25.8293 | 323.1687 | 253.9707 | 317.6847 | 227.0204 | 89.1585 | 1.5493 | 39.3769 | 61.0082 | 25.4152 | 2.4607 | 39.2597 | 96.6078 | 15.9544 | 21.3897 | 6.1466 | 157.616 | 301.779 | 232.581 | 296.295 | 205.6307
2012-10-30 01:25 | 25.8308 | 323.1588 | 253.9526 | 317.6632 | 227.136 | 89.1463 | 1.5493 | 39.3755 | 61.0041 | 25.4152 | 2.4607 | 39.2583 | 96.6016 | 15.9544 | 21.3953 | 6.1457 | 157.6058 | 301.7635 | 232.5573 | 296.2679 | 205.7407
2012-10-30 01:25 | 25.827 | 323.1778 | 253.9762 | 317.6824 | 227.1912 | 89.1648 | 1.5493 | 39.3748 | 61.0021 | 25.4151 | 2.4606 | 39.2577 | 96.5986 | 15.9543 | 21.3714 | 6.147 | 157.6007 | 301.8064 | 232.6048 | 296.311 | 205.8198
2012-10-30 01:25 | 25.827 | 323.174 | 253.941 | 317.6671 | 226.963 | 89.1631 | 1.5492 | 39.3725 | 60.9956 | 25.4149 | 2.4605 | 39.2553 | 96.5877 | 15.9542 | 21.4075 | 6.1469 | 157.5833 | 301.7665 | 232.5335 | 296.2596 | 205.5555
2012-10-30 01:25 | 25.8293 | 323.1573 | 253.9511 | 317.6885 | 226.9732 | 89.144 | 1.5491 | 39.3704 | 60.9882 | 25.4152 | 2.4604 | 39.2533 | 96.5774 | 15.9542 | 21.3938 | 6.1456 | 157.5656 | 301.7635 | 232.5573 | 296.2947 | 205.5794
2012-10-30 01:25 | 25.827 | 323.1778 | 253.9332 | 317.6747 | 227.1322 | 89.1515 | 1.5458 | 39.2859 | 60.7295 | 25.4141 | 2.4551 | 39.1683 | 96.1616 | 15.9539 | 21.3915 | 6.1461 | 156.891 | 301.7863 | 232.5417 | 296.2832 | 205.7407
2012-10-30 01:25 | 25.8368 | 323.1495 | 253.9508 | 317.6845 | 226.2608 | 89.1221 | 1.5445 | 39.2515 | 60.6248 | 25.4134 | 2.453 | 39.1347 | 95.998 | 15.9538 | 21.4013 | 6.1441 | 156.6227 | 301.7482 | 232.5495 | 296.2832 | 204.8595
2012-10-30 01:25 | 25.8225 | 323.1428 | 253.7764 | 317.6549 | 226.0459 | 89.1382 | 1.5444 | 39.2491 | 60.6175 | 25.4133 | 2.4529 | 39.1323 | 95.9869 | 15.9536 | 21.387 | 6.1452 | 156.6043 | 301.7558 | 232.3894 | 296.2679 | 204.6589
2012-10-30 01:25 | 25.8353 | 323.1404 | 253.6603 | 317.6945 | 226.0587 | 89.1417 | 1.5445 | 39.2518 | 60.6263 | 25.4132 | 2.4531 | 39.1353 | 96.0015 | 15.9536 | 21.4198 | 6.1454 | 156.6278 | 301.7206 | 232.2405 | 296.2747 | 204.6389
2012-10-30 01:25 | 25.8285 | 323.1336 | 253.6222 | 317.6686 | 226.8111 | 89.1394 | 1.5448 | 39.2594 | 60.649 | 25.4135 | 2.4535 | 39.1427 | 96.0375 | 15.9537 | 21.417 | 6.1453 | 156.6865 | 301.7166 | 232.2052 | 296.2516 | 205.3941
2012-10-30 01:26 | 25.8308 | 323.113 | 253.7105 | 317.6709 | 227.254 | 89.1313 | 1.545 | 39.2646 | 60.6646 | 25.4136 | 2.4539 | 39.1478 | 96.063 | 15.9536 | 21.3993 | 6.1447 | 156.7275 | 301.7137 | 232.3112 | 296.2716 | 205.8547
2012-10-30 01:26 | 25.8263 | 323.1123 | 253.5575 | 317.6282 | 225.8019 | 89.1296 | 1.5476 | 39.3303 | 60.8665 | 25.4142 | 2.4579 | 39.2133 | 96.3825 | 15.9539 | 21.3988 | 6.1446 | 157.249 | 301.7135 | 232.1587 | 296.2294 | 204.4031
2012-10-30 01:26 | 25.821 | 323.1108 | 253.228 | 317.6649 | 225.4661 | 89.1313 | 1.5498 | 39.388 | 61.0429 | 25.4152 | 2.4614 | 39.2704 | 96.6622 | 15.9542 | 21.4296 | 6.1447 | 157.705 | 301.6812 | 231.7984 | 296.2353 | 204.0365
2012-10-30 01:26 | 25.8308 | 323.1282 | 253.1597 | 317.606 | 225.7906 | 89.1302 | 1.55 | 39.3945 | 61.0629 | 25.4152 | 2.4619 | 39.2772 | 96.6947 | 15.9543 | 21.4233 | 6.1446 | 157.7576 | 301.7049 | 231.7364 | 296.1827 | 204.3673
2012-10-30 01:26 | 25.824 | 323.0986 | 253.1177 | 317.5877 | 226.3779 | 89.1354 | 1.5502 | 39.3975 | 61.0723 | 25.4152 | 2.4621 | 39.2804 | 96.7103 | 15.9543 | 21.3764 | 6.145 | 157.7826 | 301.7222 | 231.7413 | 296.2113 | 205.0015
2012-10-30 01:26 | 25.824 | 323.1138 | 253.1607 | 317.6183 | 226.4801 | 89.1348 | 1.5503 | 39.4006 | 61.0811 | 25.4155 | 2.4622 | 39.2834 | 96.7247 | 15.9544 | 21.3965 | 6.145 | 157.8059 | 301.7173 | 231.7642 | 296.2218 | 205.0836
2012-10-30 01:26 | 25.8233 | 323.1207 | 253.2186 | 317.5603 | 226.6446 | 89.1406 | 1.5503 | 39.4021 | 61.0859 | 25.4154 | 2.4623 | 39.2848 | 96.7311 | 15.9545 | 21.3717 | 6.1454 | 157.817 | 301.749 | 231.8469 | 296.1886 | 205.2729
2012-10-30 01:26 | 25.8203 | 323.0834 | 253.2976 | 317.5344 | 226.6731 | 89.1192 | 1.5503 | 39.4011 | 61.0828 | 25.4154 | 2.4622 | 39.2839 | 96.7265 | 15.9545 | 21.3647 | 6.1439 | 157.8093 | 301.7187 | 231.9329 | 296.1697 | 205.3084
2012-10-30 01:26 | 25.8233 | 323.075 | 253.3397 | 317.5068 | 226.7312 | 89.1342 | 1.5503 | 39.4013 | 61.0827 | 25.4158 | 2.4622 | 39.2839 | 96.7268 | 15.9545 | 21.3476 | 6.1449 | 157.8095 | 301.7274 | 231.9921 | 296.1592 | 205.3836
2012-10-30 01:26 | 25.8308 | 323.1092 | 253.4214 | 317.5716 | 226.8567 | 89.14 | 1.5502 | 39.4001 | 61.0793 | 25.4157 | 2.4622 | 39.283 | 96.7218 | 15.9545 | 21.3632 | 6.1453 | 157.8011 | 301.746 | 232.0582 | 296.2084 | 205.4935
2012-10-30 01:26 | 25.8263 | 323.1237 | 253.4442 | 317.5289 | 226.8876 | 89.1434 | 1.5503 | 39.4007 | 61.0812 | 25.4156 | 2.4622 | 39.2833 | 96.7238 | 15.9545 | 21.3707 | 6.1455 | 157.805 | 301.753 | 232.0735 | 296.1582 | 205.5169
2012-10-30 01:26 | 25.8293 | 323.1267 | 253.5176 | 317.5128 | 226.8552 | 89.1394 | 1.5502 | 39.399 | 61.076 | 25.4156 | 2.4621 | 39.2817 | 96.7157 | 15.9545 | 21.3857 | 6.1453 | 157.7918 | 301.741 | 232.1319 | 296.1271 | 205.4695
2012-10-30 01:26 | 25.8345 | 323.1243 | 253.5189 | 317.518 | 226.9037 | 89.1452 | 1.5501 | 39.398 | 61.0723 | 25.4158 | 2.462 | 39.2806 | 96.7099 | 15.9546 | 21.3829 | 6.1457 | 157.7822 | 301.7414 | 232.136 | 296.1351 | 205.5208
2012-10-30 01:26 | 25.8278 | 323.1291 | 253.5434 | 317.5151 | 227.0032 | 89.1452 | 1.5501 | 39.396 | 61.0663 | 25.4158 | 2.4619 | 39.2786 | 96.6998 | 15.9546 | 21.3321 | 6.1457 | 157.7661 | 301.797 | 232.2113 | 296.183 | 205.6711
2012-10-30 01:26 | 25.8375 | 323.0968 | 253.6195 | 317.5439 | 226.9263 | 89.1406 | 1.55 | 39.3941 | 61.0603 | 25.4157 | 2.4618 | 39.2768 | 96.6912 | 15.9545 | 21.3699 | 6.1454 | 157.7515 | 301.7269 | 232.2496 | 296.174 | 205.5564
2012-10-30 01:26 | 25.8293 | 323.0886 | 253.623 | 317.5319 | 227.0165 | 89.1446 | 1.55 | 39.3937 | 61.059 | 25.4158 | 2.4618 | 39.2764 | 96.6898 | 15.9545 | 21.3536 | 6.1456 | 157.7488 | 301.735 | 232.2694 | 296.1783 | 205.6629
2012-10-30 01:26 | 25.8383 | 323.1243 | 253.6945 | 317.5256 | 226.9822 | 89.1463 | 1.55 | 39.3936 | 61.059 | 25.4157 | 2.4618 | 39.2762 | 96.6892 | 15.9544 | 21.3707 | 6.1457 | 157.7483 | 301.7536 | 232.3238 | 296.1549 | 205.6115
2012-10-30 01:26 | 25.8398 | 323.1258 | 253.6921 | 317.4813 | 226.9837 | 89.1371 | 1.5499 | 39.3922 | 61.0549 | 25.4156 | 2.4617 | 39.2749 | 96.683 | 15.9544 | 21.3842 | 6.1451 | 157.7378 | 301.7416 | 232.3079 | 296.0971 | 205.5995
2012-10-30 01:26 | 25.8405 | 323.1341 | 253.6967 | 317.5622 | 227.1417 | 89.1487 | 1.5499 | 39.3907 | 61.0505 | 25.4155 | 2.4616 | 39.2734 | 96.6748 | 15.9545 | 21.3769 | 6.1459 | 157.7252 | 301.7572 | 232.3198 | 296.1853 | 205.7648
2012-10-30 01:26 | 25.8375 | 323.1235 | 253.7445 | 317.5516 | 227.2646 | 89.1429 | 1.5498 | 39.3894 | 61.0464 | 25.4155 | 2.4615 | 39.2723 | 96.6704 | 15.9543 | 21.3859 | 6.1455 | 157.7167 | 301.7376 | 232.3586 | 296.1657 | 205.8787
2012-10-30 01:27 | 25.836 | 323.1334 | 253.7977 | 317.5081 | 227.2316 | 89.1382 | 1.5498 | 39.3887 | 61.0443 | 25.4155 | 2.4615 | 39.2715 | 96.6672 | 15.9542 | 21.3724 | 6.1452 | 157.7115 | 301.761 | 232.4253 | 296.1357 | 205.8592
2012-10-30 01:27 | 25.8353 | 323.1671 | 253.8556 | 317.5608 | 227.168 | 89.1487 | 1.5497 | 39.3868 | 61.038 | 25.4157 | 2.4614 | 39.2695 | 96.6571 | 15.9543 | 21.3636 | 6.1459 | 157.6951 | 301.8035 | 232.492 | 296.1972 | 205.8044
2012-10-30 01:27 | 25.836 | 323.164 | 253.825 | 317.5615 | 227.2041 | 89.1458 | 1.5497 | 39.3862 | 61.0363 | 25.4155 | 2.4614 | 39.2689 | 96.6547 | 15.9542 | 21.3844 | 6.1457 | 157.691 | 301.7796 | 232.4406 | 296.1771 | 205.8197
2012-10-30 01:27 | 25.8375 | 323.1426 | 253.8382 | 317.5897 | 227.1623 | 89.1481 | 1.5497 | 39.3856 | 61.0345 | 25.4155 | 2.4613 | 39.2683 | 96.6513 | 15.9543 | 21.3979 | 6.1459 | 157.6859 | 301.7447 | 232.4403 | 296.1918 | 205.7644
2012-10-30 01:27 | 25.8383 | 323.1472 | 253.8664 | 317.5409 | 227.4581 | 89.1591 | 1.5496 | 39.3847 | 61.0313 | 25.4157 | 2.4613 | 39.2675 | 96.6471 | 15.9543 | 21.3305 | 6.1466 | 157.6785 | 301.8167 | 232.5359 | 296.2104 | 206.1276
2012-10-30 01:27 | 25.842 | 323.17 | 253.9287 | 317.579 | 227.2809 | 89.1515 | 1.5496 | 39.3843 | 61.0304 | 25.4156 | 2.4612 | 39.267 | 96.6443 | 15.9543 | 21.3503 | 6.1461 | 157.6747 | 301.8197 | 232.5784 | 296.2287 | 205.9306
2012-10-30 01:27 | 25.8427 | 323.163 | 253.9177 | 317.5759 | 227.2659 | 89.1591 | 1.5496 | 39.3829 | 61.0265 | 25.4155 | 2.4611 | 39.2658 | 96.6389 | 15.9543 | 21.351 | 6.1466 | 157.6654 | 301.812 | 232.5667 | 296.2249 | 205.9149
2012-10-30 01:27 | 25.8472 | 323.1332 | 253.9456 | 317.6224 | 227.2468 | 89.1469 | 1.5495 | 39.3817 | 61.0227 | 25.4154 | 2.4611 | 39.2645 | 96.6333 | 15.9542 | 21.3475 | 6.1458 | 157.656 | 301.7857 | 232.5981 | 296.2749 | 205.8993
2012-10-30 01:27 | 25.8413 | 323.1731 | 253.9905 | 317.5821 | 227.3785 | 89.1602 | 1.5494 | 39.3794 | 61.016 | 25.4152 | 2.461 | 39.2623 | 96.6225 | 15.9541 | 21.3536 | 6.1467 | 157.6385 | 301.8195 | 232.6369 | 296.2285 | 206.0249
2012-10-30 01:27 | 25.8465 | 323.1363 | 253.98 | 317.5949 | 227.3601 | 89.1452 | 1.5494 | 39.3783 | 61.0124 | 25.4154 | 2.4609 | 39.2609 | 96.6166 | 15.954 | 21.3748 | 6.1457 | 157.629 | 301.7615 | 232.6052 | 296.2201 | 205.9853
2012-10-30 01:27 | 25.8562 | 323.1536 | 253.9976 | 317.6275 | 227.2558 | 89.1654 | 1.5494 | 39.3776 | 61.0105 | 25.4152 | 2.4609 | 39.2603 | 96.6139 | 15.9539 | 21.3685 | 6.1471 | 157.6244 | 301.7851 | 232.6291 | 296.259 | 205.8873
2012-10-30 01:27 | 25.86 | 323.1651 | 253.9935 | 317.6046 | 227.2006 | 89.1475 | 1.5493 | 39.3759 | 61.0051 | 25.4153 | 2.4607 | 39.2587 | 96.6058 | 15.954 | 21.3402 | 6.1458 | 157.611 | 301.8249 | 232.6533 | 296.2644 | 205.8604
2012-10-30 01:27 | 25.8585 | 323.1826 | 253.9647 | 317.6184 | 227.2109 | 89.1562 | 1.5492 | 39.3746 | 61.0006 | 25.4154 | 2.4607 | 39.2574 | 96.6003 | 15.9538 | 21.3588 | 6.1464 | 157.6008 | 301.8238 | 232.6059 | 296.2596 | 205.8521
2012-10-30 01:27 | 25.8577 | 323.178 | 253.9991 | 317.6481 | 227.2887 | 89.1643 | 1.5492 | 39.3736 | 60.9981 | 25.4152 | 2.4606 | 39.2563 | 96.5949 | 15.9539 | 21.3459 | 6.147 | 157.593 | 301.8321 | 232.6532 | 296.3022 | 205.9428
2012-10-30 01:27 | 25.8495 | 323.1927 | 253.9674 | 317.6132 | 227.2805 | 89.1487 | 1.5492 | 39.3721 | 60.9935 | 25.4152 | 2.4605 | 39.255 | 96.5881 | 15.9539 | 21.3377 | 6.1459 | 157.5816 | 301.855 | 232.6297 | 296.2755 | 205.9428
2012-10-30 01:27 | 25.8592 | 323.1948 | 254.0474 | 317.6344 | 227.1565 | 89.1579 | 1.5475 | 39.3294 | 60.8624 | 25.4147 | 2.4579 | 39.213 | 96.3822 | 15.9538 | 21.3595 | 6.1465 | 157.2446 | 301.8353 | 232.6879 | 296.2749 | 205.797
2012-10-30 01:27 | 25.8622 | 323.1482 | 253.9645 | 317.6144 | 227.2303 | 89.1417 | 1.5454 | 39.2735 | 60.6918 | 25.4137 | 2.4544 | 39.1569 | 96.1086 | 15.9535 | 21.3785 | 6.1454 | 156.8003 | 301.7697 | 232.586 | 296.2359 | 205.8518
2012-10-30 01:27 | 25.8607 | 323.1581 | 253.9786 | 317.6359 | 227.1619 | 89.133 | 1.5448 | 39.2588 | 60.6469 | 25.4136 | 2.4535 | 39.142 | 96.0365 | 15.9533 | 21.361 | 6.1448 | 156.6833 | 301.7971 | 232.6176 | 296.2749 | 205.8009
2012-10-30 01:27 | 25.8682 | 323.1618 | 253.9314 | 317.6357 | 226.8902 | 89.1134 | 1.5446 | 39.2537 | 60.6316 | 25.4134 | 2.4532 | 39.137 | 96.0124 | 15.9532 | 21.3444 | 6.1435 | 156.6441 | 301.8174 | 232.587 | 296.2913 | 205.5458
2012-10-30 01:27 | 25.8712 | 323.1191 | 253.8876 | 317.6349 | 226.96 | 89.1273 | 1.5446 | 39.2534 | 60.6308 | 25.4133 | 2.4532 | 39.1367 | 96.0109 | 15.9532 | 21.3594 | 6.1444 | 156.6417 | 301.7597 | 232.5282 | 296.2755 | 205.6006
2012-10-30 01:27 | 25.8682 | 323.1618 | 253.8455 | 317.6319 | 227.2756 | 89.1203 | 1.5447 | 39.2566 | 60.6404 | 25.4134 | 2.4534 | 39.1397 | 96.026 | 15.9532 | 21.3444 | 6.144 | 156.6664 | 301.8174 | 232.5011 | 296.2875 | 205.9312
2012-10-30 01:27 | 25.8652 | 323.1398 | 253.9089 | 317.5984 | 228.1536 | 89.118 | 1.5449 | 39.2628 | 60.6589 | 25.4137 | 2.4538 | 39.1459 | 96.0559 | 15.9532 | 21.3695 | 6.1438 | 156.7148 | 301.7703 | 232.5394 | 296.2289 | 206.7841
2012-10-30 01:27 | 25.8615 | 323.117 | 254.0263 | 317.5947 | 227.8156 | 89.1157 | 1.5452 | 39.2699 | 60.6808 | 25.4137 | 2.4542 | 39.1529 | 96.0902 | 15.9532 | 21.3818 | 6.1436 | 156.771 | 301.7352 | 232.6445 | 296.2129 | 206.4338
2012-10-30 01:28 | 25.8674 | 323.1267 | 254.0088 | 317.5967 | 227.1647 | 89.1302 | 1.5454 | 39.2739 | 60.6938 | 25.4135 | 2.4545 | 39.1569 | 96.1102 | 15.9532 | 21.4038 | 6.1446 | 156.804 | 301.7229 | 232.605 | 296.1929 | 205.7609
2012-10-30 01:28 | 25.86 | 323.1422 | 253.9545 | 317.5932 | 227.5073 | 89.1307 | 1.5456 | 39.2788 | 60.7089 | 25.4135 | 2.4548 | 39.1619 | 96.1339 | 15.9533 | 21.3643 | 6.1447 | 156.8428 | 301.7779 | 232.5902 | 296.2289 | 206.143
2012-10-30 01:28 | 25.8525 | 323.1004 | 253.986 | 317.5895 | 227.4369 | 89.1151 | 1.5485 | 39.355 | 60.9422 | 25.4145 | 2.4595 | 39.2372 | 96.5024 | 15.9536 | 21.3568 | 6.1436 | 157.4446 | 301.7436 | 232.6292 | 296.2327 | 206.0801
2012-10-30 01:28 | 25.854 | 323.079 | 253.9719 | 317.5948 | 227.4856 | 89.1313 | 1.5491 | 39.3706 | 60.9898 | 25.4148 | 2.4605 | 39.2532 | 96.5807 | 15.9537 | 21.4024 | 6.1447 | 157.5705 | 301.6766 | 232.5695 | 296.1924 | 206.0832
2012-10-30 01:28 | 25.8398 | 323.0991 | 253.9772 | 317.6073 | 227.0427 | 89.1475 | 1.5496 | 39.3836 | 61.0295 | 25.4151 | 2.4613 | 39.2666 | 96.6468 | 15.9537 | 21.4203 | 6.1458 | 157.6763 | 301.6788 | 232.5569 | 296.187 | 205.6224
2012-10-30 01:28 | 25.8375 | 323.1121 | 253.9125 | 317.5363 | 226.7493 | 89.1174 | 1.55 | 39.3926 | 61.0566 | 25.4153 | 2.4618 | 39.2755 | 96.6901 | 15.9537 | 21.3899 | 6.1438 | 157.7468 | 301.7222 | 232.5226 | 296.1464 | 205.3594
2012-10-30 01:28 | 25.8435 | 323.1181 | 253.8755 | 317.5079 | 226.3147 | 89.1377 | 1.5501 | 39.3977 | 61.0722 | 25.4155 | 2.4621 | 39.2802 | 96.7123 | 15.9538 | 21.416 | 6.1452 | 157.7845 | 301.7021 | 232.4595 | 296.0919 | 204.8987
2012-10-30 01:28 | 25.8315 | 323.1022 | 253.7346 | 317.5608 | 226.464 | 89.1515 | 1.5503 | 39.4019 | 61.0849 | 25.4156 | 2.4624 | 39.2846 | 96.734 | 15.9539 | 21.4 | 6.1461 | 157.8188 | 301.7022 | 232.3346 | 296.1608 | 205.064
2012-10-30 01:28 | 25.8323 | 323.0992 | 253.7471 | 317.5311 | 226.7598 | 89.1423 | 1.5504 | 39.4036 | 61.09 | 25.4157 | 2.4625 | 39.2862 | 96.7417 | 15.9539 | 21.3927 | 6.1455 | 157.8317 | 301.7065 | 232.3544 | 296.1384 | 205.3671
2012-10-30 01:28 | 25.839 | 323.0907 | 253.7655 | 317.5531 | 226.9436 | 89.1429 | 1.5504 | 39.4042 | 61.0927 | 25.4154 | 2.4625 | 39.2868 | 96.745 | 15.9538 | 21.4275 | 6.1455 | 157.8377 | 301.6632 | 232.338 | 296.1256 | 205.5161
2012-10-30 01:28 | 25.8338 | 323.0588 | 253.7603 | 317.5249 | 227.017 | 89.1365 | 1.5504 | 39.4035 | 61.0902 | 25.4155 | 2.4625 | 39.2863 | 96.7417 | 15.9539 | 21.4183 | 6.1451 | 157.8319 | 301.6405 | 232.342 | 296.1066 | 205.5987
2012-10-30 01:28 | 25.8293 | 323.0505 | 253.7871 | 317.5357 | 226.9181 | 89.1492 | 1.5503 | 39.4026 | 61.0873 | 25.4155 | 2.4624 | 39.2851 | 96.7364 | 15.9539 | 21.3938 | 6.1459 | 157.8238 | 301.6567 | 232.3933 | 296.1419 | 205.5243
2012-10-30 01:28 | 25.8375 | 323.0587 | 253.7992 | 317.5401 | 227.0404 | 89.1446 | 1.5503 | 39.4027 | 61.0876 | 25.4155 | 2.4624 | 39.2854 | 96.7379 | 15.9539 | 21.418 | 6.1456 | 157.8256 | 301.6407 | 232.3812 | 296.1221 | 205.6224
2012-10-30 01:28 | 25.8308 | 323.0481 | 253.7964 | 317.5525 | 226.9826 | 89.1562 | 1.5503 | 39.4008 | 61.0814 | 25.4156 | 2.4623 | 39.2836 | 96.7289 | 15.9539 | 21.4113 | 6.1464 | 157.8103 | 301.6368 | 232.3851 | 296.1412 | 205.5713
2012-10-30 01:28 | 25.827 | 323.0634 | 253.8043 | 317.5563 | 227.0063 | 89.1446 | 1.5502 | 39.4007 | 61.0809 | 25.4157 | 2.4623 | 39.2831 | 96.7262 | 15.9539 | 21.3874 | 6.1456 | 157.8071 | 301.676 | 232.4169 | 296.1689 | 205.6189
2012-10-30 01:28 | 25.8315 | 323.045 | 253.8557 | 317.4997 | 227.1642 | 89.1359 | 1.5502 | 39.3983 | 61.0739 | 25.4155 | 2.4622 | 39.2811 | 96.7163 | 15.9539 | 21.3919 | 6.145 | 157.7902 | 301.6531 | 232.4638 | 296.1078 | 205.7723
2012-10-30 01:28 | 25.8308 | 323.0634 | 253.8315 | 317.5296 | 226.9668 | 89.1487 | 1.5501 | 39.3973 | 61.0705 | 25.4156 | 2.4621 | 39.28 | 96.7105 | 15.954 | 21.3792 | 6.1459 | 157.781 | 301.6842 | 232.4523 | 296.1504 | 205.5876
2012-10-30 01:28 | 25.8293 | 323.0581 | 253.8261 | 317.5434 | 226.4618 | 89.1602 | 1.5501 | 39.3965 | 61.0684 | 25.4155 | 2.4621 | 39.2792 | 96.7076 | 15.9539 | 21.4018 | 6.1467 | 157.776 | 301.6563 | 232.4243 | 296.1416 | 205.06
2012-10-30 01:28 | 25.8383 | 323.0442 | 253.7648 | 317.5447 | 226.6518 | 89.1487 | 1.5501 | 39.3954 | 61.065 | 25.4155 | 2.462 | 39.2782 | 96.703 | 15.9538 | 21.3787 | 6.1459 | 157.768 | 301.6655 | 232.3861 | 296.166 | 205.2731
2012-10-30 01:28 | 25.8278 | 323.0413 | 253.7348 | 317.5266 | 226.6137 | 89.1267 | 1.55 | 39.3947 | 61.0622 | 25.4158 | 2.462 | 39.2774 | 96.6991 | 15.9538 | 21.3963 | 6.1444 | 157.7613 | 301.645 | 232.3385 | 296.1303 | 205.2174
2012-10-30 01:28 | 25.8196 | 323.0484 | 253.6719 | 317.4993 | 226.7393 | 89.1429 | 1.55 | 39.3932 | 61.058 | 25.4156 | 2.4619 | 39.276 | 96.6923 | 15.9537 | 21.368 | 6.1455 | 157.7503 | 301.6804 | 232.3039 | 296.1313 | 205.3713
2012-10-30 01:28 | 25.8353 | 323.0259 | 253.6876 | 317.5646 | 227.0107 | 89.1359 | 1.5499 | 39.392 | 61.0543 | 25.4156 | 2.4618 | 39.2747 | 96.6862 | 15.9537 | 21.3797 | 6.145 | 157.7405 | 301.6462 | 232.3079 | 296.1849 | 205.631
2012-10-30 01:28 | 25.8308 | 323.0558 | 253.7066 | 317.5487 | 227.0376 | 89.166 | 1.5499 | 39.3907 | 61.0507 | 25.4153 | 2.4617 | 39.2734 | 96.6806 | 15.9536 | 21.3752 | 6.1471 | 157.7313 | 301.6806 | 232.3314 | 296.1735 | 205.6624
2012-10-30 01:29 | 25.836 | 323.0266 | 253.7313 | 317.5081 | 226.8462 | 89.1463 | 1.5498 | 39.3895 | 61.0464 | 25.4156 | 2.4616 | 39.2723 | 96.6745 | 15.9537 | 21.4165 | 6.1457 | 157.721 | 301.6101 | 232.3148 | 296.0916 | 205.4297
2012-10-30 01:29 | 25.833 | 323.0542 | 253.7205 | 317.528 | 226.8825 | 89.1319 | 1.5498 | 39.3881 | 61.0426 | 25.4154 | 2.4616 | 39.2709 | 96.6679 | 15.9536 | 21.3694 | 6.1448 | 157.7105 | 301.6848 | 232.3511 | 296.1586 | 205.5131
2012-10-30 01:29 | 25.8405 | 323.0807 | 253.7006 | 317.5355 | 226.7602 | 89.1446 | 1.5497 | 39.3864 | 61.0377 | 25.4152 | 2.4615 | 39.2693 | 96.6604 | 15.9535 | 21.4009 | 6.1456 | 157.6981 | 301.6798 | 232.2997 | 296.1346 | 205.3593
2012-10-30 01:29 | 25.8442 | 323.0692 | 253.7043 | 317.5124 | 226.8819 | 89.1434 | 1.5496 | 39.385 | 61.0329 | 25.4154 | 2.4614 | 39.2677 | 96.6532 | 15.9535 | 21.4046 | 6.1455 | 157.6861 | 301.6646 | 232.2997 | 296.1078 | 205.4773
2012-10-30 01:29 | 25.8353 | 323.0526 | 253.7189 | 317.5837 | 226.9871 | 89.1481 | 1.5496 | 39.3835 | 61.0283 | 25.4155 | 2.4613 | 39.2663 | 96.6459 | 15.9535 | 21.4238 | 6.1459 | 157.6742 | 301.6288 | 232.2951 | 296.1599 | 205.5633
2012-10-30 01:29 | 25.8398 | 323.0495 | 253.7195 | 317.5348 | 227.2512 | 89.1475 | 1.5496 | 39.3827 | 61.0262 | 25.4153 | 2.4613 | 39.2654 | 96.6423 | 15.9534 | 21.4123 | 6.1458 | 157.6685 | 301.6372 | 232.3072 | 296.1225 | 205.8389
2012-10-30 01:29 | 25.8368 | 323.077 | 253.779 | 317.547 | 227.2875 | 89.1469 | 1.5495 | 39.3806 | 61.0191 | 25.4155 | 2.4611 | 39.2634 | 96.6321 | 15.9534 | 21.4133 | 6.1458 | 157.6512 | 301.6637 | 232.3657 | 296.1337 | 205.8742
2012-10-30 01:29 | 25.8413 | 323.1082 | 253.8186 | 317.5401 | 227.1189 | 89.1539 | 1.5494 | 39.3784 | 61.0122 | 25.4155 | 2.461 | 39.2611 | 96.6208 | 15.9534 | 21.3777 | 6.1463 | 157.633 | 301.7305 | 232.4409 | 296.1624 | 205.7412
2012-10-30 01:29 | 25.8465 | 323.0715 | 253.8707 | 317.5644 | 227.0415 | 89.1487 | 1.5494 | 39.3773 | 61.0095 | 25.4153 | 2.4609 | 39.2601 | 96.6155 | 15.9535 | 21.4029 | 6.1459 | 157.625 | 301.6686 | 232.4678 | 296.1615 | 205.6386
2012-10-30 01:29 | 25.8442 | 323.0921 | 253.8567 | 317.5774 | 226.7599 | 89.1469 | 1.5493 | 39.3759 | 61.0055 | 25.4151 | 2.4608 | 39.2587 | 96.6093 | 15.9534 | 21.4127 | 6.1458 | 157.6148 | 301.6794 | 232.444 | 296.1647 | 205.3472
2012-10-30 01:29 | 25.8405 | 323.0464 | 253.7944 | 317.5355 | 226.9372 | 89.1515 | 1.5492 | 39.3738 | 60.9985 | 25.4153 | 2.4607 | 39.2568 | 96.5994 | 15.9535 | 21.4531 | 6.1461 | 157.5978 | 301.5933 | 232.3413 | 296.0824 | 205.4841
2012-10-30 01:29 | 25.845 | 323.0471 | 253.8028 | 317.54 | 227.0243 | 89.1487 | 1.5492 | 39.3722 | 60.9937 | 25.4153 | 2.4606 | 39.255 | 96.592 | 15.9533 | 21.4095 | 6.1459 | 157.5857 | 301.6376 | 232.3933 | 296.1305 | 205.6148
2012-10-30 01:29 | 25.8405 | 323.0617 | 253.8373 | 317.5469 | 226.8231 | 89.1533 | 1.5492 | 39.3723 | 60.9942 | 25.4152 | 2.4606 | 39.2551 | 96.592 | 15.9534 | 21.4371 | 6.1462 | 157.5862 | 301.6246 | 232.4002 | 296.1098 | 205.386
2012-10-30 01:29 | 25.8413 | 323.0815 | 253.8069 | 317.5783 | 226.8161 | 89.1515 | 1.5464 | 39.3002 | 60.7729 | 25.4143 | 2.4559 | 39.1788 | 96.2189 | 15.953 | 21.4178 | 6.1461 | 156.9918 | 301.6637 | 232.3891 | 296.1605 | 205.3983
2012-10-30 01:29 | 25.845 | 323.0623 | 253.7676 | 317.6087 | 226.9417 | 89.1388 | 1.5453 | 39.2716 | 60.6864 | 25.4136 | 2.4544 | 39.1547 | 96.1017 | 15.9528 | 21.4054 | 6.1452 | 156.7881 | 301.6569 | 232.3622 | 296.2033 | 205.5363
2012-10-30 01:29 | 25.8368 | 323.0122 | 253.779 | 317.5509 | 227.0436 | 89.1446 | 1.5448 | 39.259 | 60.6475 | 25.4135 | 2.4536 | 39.1421 | 96.0402 | 15.9528 | 21.3972 | 6.1456 | 156.6877 | 301.615 | 232.3818 | 296.1537 | 205.6464
2012-10-30 01:29 | 25.8353 | 322.9878 | 253.8009 | 317.5494 | 226.7943 | 89.1388 | 1.5447 | 39.2556 | 60.6372 | 25.4134 | 2.4534 | 39.1387 | 96.0234 | 15.9528 | 21.3797 | 6.1452 | 156.6606 | 301.6081 | 232.4212 | 296.1697 | 205.4146
2012-10-30 01:29 | 25.8383 | 322.9832 | 253.7062 | 317.5753 | 226.8563 | 89.144 | 1.5447 | 39.2563 | 60.6398 | 25.4134 | 2.4535 | 39.1393 | 96.0267 | 15.9527 | 21.4148 | 6.1456 | 156.6665 | 301.5684 | 232.2914 | 296.1605 | 205.4415
2012-10-30 01:29 | 25.839 | 322.9953 | 253.7265 | 317.5492 | 226.8177 | 89.1105 | 1.5448 | 39.2586 | 60.6463 | 25.4136 | 2.4536 | 39.1417 | 96.0383 | 15.9527 | 21.4275 | 6.1433 | 156.6846 | 301.5678 | 232.299 | 296.1217 | 205.3902
2012-10-30 01:29 | 25.8263 | 323.0322 | 253.7099 | 317.5404 | 227.0489 | 89.1244 | 1.5451 | 39.2663 | 60.6699 | 25.4136 | 2.4541 | 39.1491 | 96.0749 | 15.9527 | 21.4108 | 6.1442 | 156.7448 | 301.6214 | 232.2991 | 296.1296 | 205.6381
2012-10-30 01:29 | 25.8345 | 322.987 | 253.8079 | 317.5677 | 227.592 | 89.1313 | 1.5452 | 39.2699 | 60.681 | 25.4136 | 2.4543 | 39.1529 | 96.0941 | 15.9526 | 21.4351 | 6.1447 | 156.7751 | 301.5519 | 232.3728 | 296.1326 | 206.1569
2012-10-30 01:29 | 25.827 | 323.01 | 253.8473 | 317.5678 | 227.6749 | 89.1238 | 1.5454 | 39.2749 | 60.6966 | 25.4136 | 2.4546 | 39.1579 | 96.1177 | 15.9528 | 21.4276 | 6.1442 | 156.8143 | 301.5824 | 232.4197 | 296.1402 | 206.2473
2012-10-30 01:29 | 25.8345 | 323.0137 | 253.929 | 317.5562 | 227.352 | 89.1255 | 1.5467 | 39.3089 | 60.8005 | 25.4141 | 2.4568 | 39.1924 | 96.2869 | 15.9528 | 21.4431 | 6.1443 | 157.0874 | 301.5706 | 232.4859 | 296.1131 | 205.9089
2012-10-30 01:30 | 25.821 | 322.9964 | 253.8764 | 317.5198 | 227.0396 | 89.1284 | 1.5488 | 39.3638 | 60.9685 | 25.415 | 2.4601 | 39.2464 | 96.5504 | 15.9532 | 21.4336 | 6.1445 | 157.5189 | 301.5628 | 232.4428 | 296.0862 | 205.606
2012-10-30 01:30 | 25.8121 | 323.0027 | 253.8675 | 317.4994 | 227.1133 | 89.1313 | 1.5494 | 39.3788 | 61.0138 | 25.4154 | 2.4611 | 39.2616 | 96.6248 | 15.9532 | 21.4207 | 6.1447 | 157.6386 | 301.582 | 232.4468 | 296.0787 | 205.6926
2012-10-30 01:30 | 25.8181 | 323.0316 | 253.8657 | 317.5092 | 227.2452 | 89.1307 | 1.5498 | 39.3889 | 61.0451 | 25.4155 | 2.4617 | 39.2717 | 96.6736 | 15.9534 | 21.4387 | 6.1447 | 157.7187 | 301.5929 | 232.427 | 296.0705 | 205.8065
2012-10-30 01:30 | 25.8121 | 323.0294 | 253.8207 | 317.488 | 227.0189 | 89.1359 | 1.5501 | 39.3964 | 61.0679 | 25.4155 | 2.4621 | 39.2792 | 96.7101 | 15.9534 | 21.4528 | 6.145 | 157.778 | 301.5766 | 232.3679 | 296.0352 | 205.5661
2012-10-30 01:30 | 25.8068 | 323.0356 | 253.8388 | 317.4941 | 227.108 | 89.1388 | 1.5502 | 39.4 | 61.0788 | 25.4157 | 2.4623 | 39.2826 | 96.7272 | 15.9533 | 21.3953 | 6.1452 | 157.806 | 301.6403 | 232.4435 | 296.0988 | 205.7127
2012-10-30 01:30 | 25.8031 | 323.0166 | 253.8741 | 317.4713 | 227.2263 | 89.1446 | 1.5503 | 39.4024 | 61.0871 | 25.4153 | 2.4625 | 39.2851 | 96.7393 | 15.9534 | 21.4157 | 6.1456 | 157.8265 | 301.6009 | 232.4584 | 296.0556 | 205.8106
2012-10-30 01:30 | 25.8008 | 323.0029 | 253.8836 | 317.4767 | 227.0509 | 89.1429 | 1.5503 | 39.4025 | 61.0869 | 25.4155 | 2.4625 | 39.2852 | 96.7395 | 15.9534 | 21.4736 | 6.1455 | 157.8265 | 301.5293 | 232.41 | 296.0031 | 205.5773
2012-10-30 01:30 | 25.8008 | 323.041 | 253.864 | 317.4805 | 226.87 | 89.1354 | 1.5504 | 39.4029 | 61.0884 | 25.4155 | 2.4625 | 39.2856 | 96.7422 | 15.9533 | 21.4696 | 6.145 | 157.8306 | 301.5714 | 232.3944 | 296.0109 | 205.4004
2012-10-30 01:30 | 25.7993 | 323.0319 | 253.8664 | 317.4981 | 227.0022 | 89.1406 | 1.5503 | 39.4024 | 61.0864 | 25.4156 | 2.4625 | 39.285 | 96.7393 | 15.9533 | 21.436 | 6.1454 | 157.8256 | 301.5959 | 232.4304 | 296.0621 | 205.5662
2012-10-30 01:30 | 25.8008 | 323.0181 | 253.864 | 317.4461 | 227.106 | 89.1452 | 1.5503 | 39.4025 | 61.0866 | 25.4156 | 2.4625 | 39.2852 | 96.7404 | 15.9533 | 21.4335 | 6.1457 | 157.8269 | 301.5846 | 232.4305 | 296.0126 | 205.6725
2012-10-30 01:30 | 25.7971 | 323.0221 | 253.8486 | 317.5264 | 227.0983 | 89.1434 | 1.5503 | 39.403 | 61.0881 | 25.4157 | 2.4625 | 39.2855 | 96.7421 | 15.9533 | 21.3937 | 6.1455 | 157.8302 | 301.6284 | 232.4549 | 296.1327 | 205.7046
2012-10-30 01:30 | 25.7881 | 323.0207 | 253.8435 | 317.4983 | 227.109 | 89.144 | 1.5503 | 39.4023 | 61.0862 | 25.4156 | 2.4625 | 39.2849 | 96.7387 | 15.9534 | 21.4047 | 6.1456 | 157.8248 | 301.616 | 232.4388 | 296.0936 | 205.7043
2012-10-30 01:30 | 25.7934 | 323.0031 | 253.8957 | 317.4769 | 227.1261 | 89.1319 | 1.5503 | 39.4012 | 61.0818 | 25.416 | 2.4624 | 39.2839 | 96.7341 | 15.9533 | 21.3579 | 6.1448 | 157.8159 | 301.6452 | 232.5378 | 296.119 | 205.7682
2012-10-30 01:30 | 25.7949 | 323.0275 | 253.8777 | 317.486 | 227.2259 | 89.144 | 1.5502 | 39.3988 | 61.075 | 25.4158 | 2.4623 | 39.2814 | 96.7207 | 15.9535 | 21.4075 | 6.1456 | 157.7957 | 301.62 | 232.4702 | 296.0785 | 205.8184
2012-10-30 01:30 | 25.7941 | 323.0496 | 253.8964 | 317.5158 | 227.1071 | 89.1388 | 1.5501 | 39.3976 | 61.0713 | 25.4157 | 2.4622 | 39.2803 | 96.7151 | 15.9535 | 21.3987 | 6.1452 | 157.7863 | 301.6509 | 232.4977 | 296.1171 | 205.7084
2012-10-30 01:30 | 25.7941 | 323.0305 | 253.9315 | 317.5349 | 227.2959 | 89.1539 | 1.5501 | 39.3965 | 61.068 | 25.4157 | 2.4621 | 39.2792 | 96.7098 | 15.9535 | 21.3746 | 6.1463 | 157.7777 | 301.6559 | 232.5569 | 296.1603 | 205.9213
2012-10-30 01:30 | 25.7926 | 323.0481 | 253.9261 | 317.5143 | 227.263 | 89.1411 | 1.55 | 39.3951 | 61.0626 | 25.4161 | 2.462 | 39.2781 | 96.704 | 15.9535 | 21.4012 | 6.1454 | 157.7666 | 301.6469 | 232.5249 | 296.1131 | 205.8618
2012-10-30 01:30 | 25.7851 | 323.0291 | 253.9147 | 317.4992 | 227.3223 | 89.1515 | 1.55 | 39.3948 | 61.0619 | 25.416 | 2.462 | 39.2775 | 96.7012 | 15.9535 | 21.3576 | 6.1461 | 157.7631 | 301.6715 | 232.5571 | 296.1416 | 205.9647
2012-10-30 01:30 | 25.7911 | 323.0466 | 253.9403 | 317.5204 | 227.2969 | 89.1452 | 1.5499 | 39.3924 | 61.0547 | 25.416 | 2.4618 | 39.275 | 96.6891 | 15.9535 | 21.3917 | 6.1457 | 157.7439 | 301.6549 | 232.5486 | 296.1287 | 205.9052
2012-10-30 01:30 | 25.7904 | 323.0154 | 253.9708 | 317.4854 | 227.2214 | 89.1608 | 1.5499 | 39.3915 | 61.0517 | 25.416 | 2.4618 | 39.2742 | 96.6845 | 15.9536 | 21.415 | 6.1467 | 157.7361 | 301.6004 | 232.5558 | 296.0704 | 205.8064
2012-10-30 01:30 | 25.7941 | 323.0648 | 253.994 | 317.5425 | 227.3077 | 89.1463 | 1.5498 | 39.3904 | 61.0485 | 25.4159 | 2.4617 | 39.2729 | 96.6782 | 15.9536 | 21.4027 | 6.1457 | 157.7267 | 301.6621 | 232.5913 | 296.1398 | 205.905
2012-10-30 01:30 | 25.7971 | 322.9992 | 253.9775 | 317.5303 | 227.3776 | 89.151 | 1.5497 | 39.3879 | 61.0406 | 25.416 | 2.4616 | 39.2706 | 96.6672 | 15.9535 | 21.3776 | 6.1461 | 157.7078 | 301.6216 | 232.5999 | 296.1527 | 206
2012-10-30 01:30 | 25.7949 | 323.0199 | 254.0222 | 317.5242 | 227.3793 | 89.1591 | 1.5497 | 39.3866 | 61.0366 | 25.416 | 2.4615 | 39.2694 | 96.6616 | 15.9535 | 21.3955 | 6.1466 | 157.6982 | 301.6244 | 232.6267 | 296.1287 | 205.9838
2012-10-30 01:31 | 25.7964 | 323.0633 | 254.0198 | 317.5601 | 227.3808 | 89.1579 | 1.5497 | 39.3862 | 61.0356 | 25.4159 | 2.4615 | 39.269 | 96.659 | 15.9535 | 21.3729 | 6.1465 | 157.6947 | 301.6904 | 232.6469 | 296.1872 | 206.0079
2012-10-30 01:31 | 25.8001 | 323.0251 | 254.0391 | 317.5944 | 227.2744 | 89.1544 | 1.5496 | 39.3843 | 61.0291 | 25.4161 | 2.4614 | 39.2669 | 96.6496 | 15.9534 | 21.4167 | 6.1463 | 157.6787 | 301.6084 | 232.6224 | 296.1777 | 205.8577
2012-10-30 01:31 | 25.8046 | 323.0601 | 254.0241 | 317.6141 | 227.2278 | 89.1591 | 1.5495 | 39.3829 | 61.0252 | 25.4159 | 2.4613 | 39.2655 | 96.6423 | 15.9535 | 21.357 | 6.1466 | 157.6675 | 301.7031 | 232.6671 | 296.2571 | 205.8708
2012-10-30 01:31 | 25.8121 | 323.0905 | 254.055 | 317.5682 | 227.4162 | 89.1539 | 1.5495 | 39.3824 | 61.0235 | 25.416 | 2.4612 | 39.2652 | 96.6414 | 15.9534 | 21.3404 | 6.1463 | 157.6648 | 301.7501 | 232.7146 | 296.2278 | 206.0758
2012-10-30 01:31 | 25.8113 | 323.1126 | 254.0698 | 317.5979 | 227.6238 | 89.1469 | 1.5495 | 39.3814 | 61.021 | 25.4158 | 2.4612 | 39.2642 | 96.6365 | 15.9534 | 21.3798 | 6.1458 | 157.6575 | 301.7328 | 232.69 | 296.2181 | 206.244
2012-10-30 01:31 | 25.8151 | 323.082 | 254.101 | 317.6094 | 227.4821 | 89.1515 | 1.5494 | 39.3801 | 61.017 | 25.4157 | 2.4611 | 39.2628 | 96.6306 | 15.9532 | 21.3796 | 6.1461 | 157.6476 | 301.7024 | 232.7214 | 296.2298 | 206.1025
2012-10-30 01:31 | 25.8106 | 323.0775 | 254.1004 | 317.6239 | 227.5563 | 89.1591 | 1.5494 | 39.379 | 61.0135 | 25.4157 | 2.4611 | 39.2617 | 96.6253 | 15.9532 | 21.3751 | 6.1466 | 157.6387 | 301.7024 | 232.7253 | 296.2488 | 206.1812
2012-10-30 01:31 | 25.8188 | 323.1277 | 254.1164 | 317.5901 | 227.5409 | 89.1527 | 1.5493 | 39.3769 | 61.0072 | 25.4157 | 2.4609 | 39.2597 | 96.6158 | 15.9531 | 21.3792 | 6.1462 | 157.623 | 301.7485 | 232.7372 | 296.2109 | 206.1617
2012-10-30 01:31 | 25.8173 | 323.1147 | 254.1656 | 317.6154 | 227.3152 | 89.1527 | 1.5492 | 39.3751 | 61.0011 | 25.4159 | 2.4608 | 39.2578 | 96.6066 | 15.9531 | 21.4018 | 6.1462 | 157.6077 | 301.7129 | 232.7638 | 296.2136 | 205.9134
2012-10-30 01:31 | 25.8203 | 323.0834 | 254.1452 | 317.626 | 227.2513 | 89.1885 | 1.6169 | 41.1287 | 66.5012 | 25.4367 | 2.5754 | 41.1086 | 105.8711 | 15.962 | 21.3607 | 6.1487 | 172.3722 | 301.7227 | 232.7845 | 296.2653 | 205.8906
2012-10-30 01:31 | 25.8218 | 323.3595 | 254.4708 | 317.6046 | 227.5045 | 89.5515 | 1.6581 | 42.2178 | 70.003 | 25.461 | 2.6349 | 42.089 | 110.9008 | 15.9736 | 21.3622 | 6.1737 | 180.9038 | 301.9973 | 233.1086 | 296.2424 | 206.1423
2012-10-30 01:31 | 25.821 | 323.7744 | 254.7707 | 317.6916 | 227.6217 | 89.5971 | 1.6543 | 42.1205 | 69.6805 | 25.461 | 2.6286 | 41.9953 | 110.3898 | 15.9761 | 21.3574 | 6.1768 | 180.0702 | 302.417 | 233.4133 | 296.3342 | 206.2643
2012-10-30 01:31 | 25.8233 | 324.1352 | 254.894 | 317.7932 | 227.5414 | 89.6214 | 1.6541 | 42.1173 | 69.6672 | 25.462 | 2.628 | 41.9917 | 110.3557 | 15.9783 | 21.3878 | 6.1785 | 180.023 | 302.7474 | 233.5062 | 296.4054 | 206.1536
2012-10-30 01:31 | 25.8278 | 324.5782 | 255.1055 | 317.9886 | 227.6875 | 89.6566 | 1.6538 | 42.1094 | 69.6385 | 25.463 | 2.6272 | 41.9838 | 110.3014 | 15.9802 | 21.3722 | 6.1809 | 179.9399 | 303.206 | 233.7333 | 296.6164 | 206.3153
2012-10-30 01:31 | 25.8315 | 325.009 | 255.3356 | 318.1909 | 227.5496 | 89.6977 | 1.6536 | 42.1079 | 69.6311 | 25.4638 | 2.6269 | 41.9824 | 110.2829 | 15.9818 | 21.3639 | 6.1838 | 179.914 | 303.6451 | 233.9717 | 296.827 | 206.1857
2012-10-30 01:31 | 25.827 | 325.3438 | 255.4404 | 318.3735 | 227.6435 | 89.7341 | 1.6537 | 42.1111 | 69.6391 | 25.4647 | 2.6269 | 41.9857 | 110.2918 | 15.983 | 21.3915 | 6.1863 | 179.931 | 303.9523 | 234.0489 | 296.982 | 206.252
2012-10-30 01:31 | 25.8308 | 325.6145 | 255.6277 | 318.614 | 228.2018 | 89.7607 | 1.6539 | 42.118 | 69.6594 | 25.4657 | 2.6271 | 41.9925 | 110.3196 | 15.9842 | 21.3792 | 6.1881 | 179.979 | 304.2353 | 234.2485 | 297.2348 | 206.8226
2012-10-30 01:31 | 25.83 | 326.0216 | 255.8494 | 318.8881 | 228.0358 | 89.7809 | 1.654 | 42.1223 | 69.6713 | 25.4665 | 2.6272 | 41.9966 | 110.3339 | 15.9853 | 21.4145 | 6.1895 | 180.0052 | 304.6071 | 234.4349 | 297.4736 | 206.6213
2012-10-30 01:31 | 25.83 | 326.3494 | 256.0406 | 319.1133 | 228.7043 | 89.815 | 1.6542 | 42.1288 | 69.691 | 25.4672 | 2.6274 | 42.0031 | 110.3605 | 15.9863 | 21.3864 | 6.1918 | 180.0515 | 304.963 | 234.6542 | 297.7269 | 207.3179
2012-10-30 01:31 | 25.8263 | 326.5973 | 256.2281 | 319.4073 | 229.2393 | 89.8456 | 1.6547 | 42.1432 | 69.7358 | 25.4683 | 2.6283 | 42.0187 | 110.4365 | 15.9872 | 21.4028 | 6.194 | 180.1723 | 305.1945 | 234.8253 | 298.0045 | 207.8365
2012-10-30 01:31 | 25.8293 | 326.9014 | 256.5043 | 319.6584 | 229.0693 | 89.8924 | 1.6576 | 42.2185 | 69.9816 | 25.4696 | 2.6327 | 42.0929 | 110.8178 | 15.9885 | 21.3817 | 6.1972 | 180.7994 | 305.5197 | 235.1226 | 298.2767 | 207.6876
2012-10-30 01:31 | 25.8345 | 327.2039 | 256.7202 | 319.8965 | 229.279 | 89.9075 | 1.658 | 42.2321 | 70.0225 | 25.4711 | 2.6334 | 42.1063 | 110.8809 | 15.9896 | 21.3949 | 6.1982 | 180.9035 | 305.809 | 235.3253 | 298.5016 | 207.8841
2012-10-30 01:32 | 25.8173 | 327.5068 | 256.863 | 320.1922 | 228.6327 | 89.9439 | 1.6584 | 42.2439 | 70.0589 | 25.4721 | 2.6339 | 42.1176 | 110.9339 | 15.9905 | 21.3858 | 6.2007 | 180.9928 | 306.121 | 235.4772 | 298.8064 | 207.2469
2012-10-30 01:32 | 25.821 | 327.8421 | 256.9408 | 320.4898 | 228.3808 | 89.9768 | 1.6587 | 42.252 | 70.0835 | 25.4729 | 2.6343 | 42.1258 | 110.9719 | 15.9913 | 21.4015 | 6.203 | 181.0554 | 306.4406 | 235.5393 | 299.0883 | 206.9793
2012-10-30 01:32 | 25.8083 | 328.0923 | 257.0803 | 320.7633 | 228.9422 | 90.0057 | 1.6588 | 42.2555 | 70.0941 | 25.4732 | 2.6344 | 42.1296 | 110.987 | 15.992 | 21.4089 | 6.205 | 181.0812 | 306.6834 | 235.6714 | 299.3544 | 207.5333
2012-10-30 01:32 | 25.8196 | 328.3475 | 257.232 | 321.0341 | 229.3545 | 90.0306 | 1.6589 | 42.258 | 70.1005 | 25.474 | 2.6345 | 42.1321 | 110.9956 | 15.9926 | 21.4242 | 6.2067 | 181.0961 | 306.9233 | 235.8078 | 299.6099 | 207.9303
2012-10-30 01:32 | 25.8128 | 328.6265 | 257.4749 | 321.3783 | 229.4382 | 90.0583 | 1.6589 | 42.2601 | 70.1053 | 25.4748 | 2.6345 | 42.1342 | 111.0013 | 15.9934 | 21.4254 | 6.2086 | 181.1066 | 307.2011 | 236.0495 | 299.9529 | 208.0128
2012-10-30 01:32 | 25.8173 | 328.9511 | 257.7251 | 321.6728 | 229.9931 | 90.1022 | 1.6589 | 42.2612 | 70.1075 | 25.4753 | 2.6344 | 42.1353 | 111.0021 | 15.9942 | 21.4139 | 6.2116 | 181.1096 | 307.5372 | 236.3112 | 300.2589 | 208.5792
2012-10-30 01:32 | 25.8143 | 329.1805 | 257.9912 | 321.9788 | 230.2653 | 90.1271 | 1.6588 | 42.2604 | 70.1023 | 25.4762 | 2.6343 | 42.1343 | 110.9928 | 15.9948 | 21.4109 | 6.2134 | 181.0951 | 307.7696 | 236.5803 | 300.5679 | 208.8544
2012-10-30 01:32 | 25.8188 | 329.4022 | 258.1361 | 322.2122 | 230.4663 | 90.1531 | 1.6588 | 42.2602 | 70.1003 | 25.4767 | 2.6342 | 42.1343 | 110.9889 | 15.9953 | 21.4314 | 6.2152 | 181.0891 | 307.9708 | 236.7047 | 300.7808 | 209.0349
2012-10-30 01:32 | 25.8083 | 329.6393 | 258.344 | 322.484 | 230.5895 | 90.1733 | 1.6587 | 42.2595 | 70.0962 | 25.4773 | 2.634 | 42.1335 | 110.981 | 15.9958 | 21.445 | 6.2165 | 181.0772 | 308.1943 | 236.899 | 301.039 | 209.1445
2012-10-30 01:32 | 25.8218 | 329.9005 | 258.572 | 322.8027 | 230.6108 | 90.2028 | 1.6587 | 42.2585 | 70.0924 | 25.4775 | 2.6339 | 42.1325 | 110.9714 | 15.9964 | 21.4184 | 6.2186 | 181.0638 | 308.4821 | 237.1536 | 301.3843 | 209.1924
2012-10-30 01:32 | 25.8233 | 330.1648 | 258.6943 | 323.0445 | 230.5809 | 90.2346 | 1.6585 | 42.2568 | 70.0844 | 25.4784 | 2.6337 | 42.1308 | 110.9599 | 15.9968 | 21.3958 | 6.2208 | 181.0443 | 308.769 | 237.2985 | 301.6487 | 209.1851
2012-10-30 01:32 | 25.8225 | 330.3545 | 258.8144 | 323.2458 | 230.8238 | 90.2571 | 1.6585 | 42.2558 | 70.0794 | 25.479 | 2.6336 | 42.1299 | 110.9516 | 15.9973 | 21.411 | 6.2223 | 181.031 | 308.9435 | 237.4034 | 301.8348 | 209.4128
2012-10-30 01:32 | 25.8225 | 330.5335 | 258.9859 | 323.4937 | 230.8592 | 90.282 | 1.6584 | 42.2548 | 70.0752 | 25.4793 | 2.6334 | 42.1291 | 110.9447 | 15.9977 | 21.395 | 6.224 | 181.0198 | 309.1385 | 237.5909 | 302.0987 | 209.4642
2012-10-30 01:32 | 25.8151 | 330.7814 | 259.1423 | 323.7723 | 231.1977 | 90.3045 | 1.6583 | 42.2527 | 70.0673 | 25.4797 | 2.6332 | 42.1269 | 110.9289 | 15.9983 | 21.3715 | 6.2256 | 180.9962 | 309.4099 | 237.7708 | 302.4008 | 209.8262
2012-10-30 01:32 | 25.8181 | 330.9786 | 259.3714 | 323.9622 | 231.4679 | 90.3213 | 1.6582 | 42.2513 | 70.0607 | 25.4804 | 2.633 | 42.1256 | 110.9187 | 15.9988 | 21.3785 | 6.2268 | 180.9793 | 309.6001 | 237.9929 | 302.5837 | 210.0894
2012-10-30 01:32 | 25.8121 | 331.1745 | 259.5057 | 324.2842 | 231.5484 | 90.3467 | 1.6581 | 42.2498 | 70.0538 | 25.4811 | 2.6329 | 42.1239 | 110.907 | 15.9992 | 21.4448 | 6.2285 | 180.9608 | 309.7297 | 238.0609 | 302.8394 | 210.1036
2012-10-30 01:32 | 25.8203 | 331.3693 | 259.7244 | 324.5135 | 231.6156 | 90.3756 | 1.658 | 42.2493 | 70.0511 | 25.4815 | 2.6328 | 42.1236 | 110.9023 | 15.9996 | 21.4329 | 6.2305 | 180.9534 | 309.9364 | 238.2915 | 303.0806 | 210.1827
2012-10-30 01:32 | 25.8166 | 331.6056 | 259.9 | 324.7234 | 231.7651 | 90.3929 | 1.6579 | 42.2471 | 70.0424 | 25.4819 | 2.6326 | 42.1212 | 110.8872 | 16 | 21.4252 | 6.2317 | 180.9297 | 310.1804 | 238.4748 | 303.2982 | 210.3399
2012-10-30 01:32 | 25.8203 | 331.8378 | 260.079 | 324.9635 | 231.8828 | 90.4218 | 1.6579 | 42.246 | 70.0376 | 25.4824 | 2.6324 | 42.1202 | 110.8791 | 16.0004 | 21.3968 | 6.2337 | 180.9167 | 310.441 | 238.6822 | 303.5667 | 210.486
2012-10-30 01:32 | 25.8113 | 332.0497 | 260.2571 | 325.2175 | 232.1568 | 90.4444 | 1.6578 | 42.2443 | 70.0309 | 25.4827 | 2.6323 | 42.1185 | 110.8674 | 16.0008 | 21.4199 | 6.2352 | 180.8984 | 310.6298 | 238.8372 | 303.7976 | 210.7369
2012-10-30 01:32 | 25.8121 | 332.1762 | 260.4021 | 325.4509 | 232.4562 | 90.4548 | 1.6576 | 42.2414 | 70.0201 | 25.4832 | 2.632 | 42.1156 | 110.8495 | 16.0012 | 21.4247 | 6.236 | 180.8696 | 310.7515 | 238.9774 | 304.0262 | 211.0315
2012-10-30 01:32 | 25.8188 | 332.4114 | 260.5841 | 325.6635 | 232.4472 | 90.4681 | 1.6575 | 42.2397 | 70.0137 | 25.4835 | 2.6319 | 42.1139 | 110.8377 | 16.0016 | 21.4194 | 6.2369 | 180.8514 | 310.992 | 239.1647 | 304.2441 | 211.0278
2012-10-30 01:32 | 25.8203 | 332.5919 | 260.7337 | 325.9051 | 232.4841 | 90.5091 | 1.6575 | 42.239 | 70.0105 | 25.4838 | 2.6317 | 42.1131 | 110.8308 | 16.002 | 21.4369 | 6.2397 | 180.8413 | 311.155 | 239.2968 | 304.4682 | 211.0472
2012-10-30 01:33 | 25.8173 | 332.806 | 260.8398 | 326.1118 | 232.4221 | 90.5322 | 1.6574 | 42.2384 | 70.007 | 25.4844 | 2.6317 | 42.1128 | 110.827 | 16.0023 | 21.4379 | 6.2413 | 180.834 | 311.3681 | 239.4019 | 304.6739 | 210.9842
2012-10-30 01:33 | 25.8166 | 332.9843 | 260.9442 | 326.2865 | 232.3939 | 90.549 | 1.6574 | 42.2377 | 70.0037 | 25.4847 | 2.6316 | 42.112 | 110.8202 | 16.0027 | 21.4091 | 6.2424 | 180.8239 | 311.5752 | 239.5351 | 304.8774 | 210.9848
2012-10-30 01:33 | 25.8196 | 333.1624 | 261.0485 | 326.5067 | 232.5148 | 90.575 | 1.6573 | 42.2357 | 69.9965 | 25.4849 | 2.6314 | 42.11 | 110.8079 | 16.0029 | 21.4282 | 6.2442 | 180.8044 | 311.7342 | 239.6203 | 305.0785 | 211.0866
2012-10-30 01:33 | 25.8196 | 333.3566 | 261.1966 | 326.7126 | 232.7348 | 90.5987 | 1.6571 | 42.2334 | 69.987 | 25.4856 | 2.6312 | 42.1076 | 110.7934 | 16.0032 | 21.4121 | 6.2459 | 180.7804 | 311.9445 | 239.7845 | 305.3005 | 211.3227
2012-10-30 01:33 | 25.8188 | 333.5767 | 261.3321 | 326.9519 | 233.0759 | 90.6085 | 1.6571 | 42.2317 | 69.9807 | 25.4859 | 2.631 | 42.1062 | 110.7825 | 16.0037 | 21.4113 | 6.2465 | 180.7632 | 312.1654 | 239.9208 | 305.5406 | 211.6646
2012-10-30 01:33 | 25.8151 | 333.7139 | 261.4842 | 327.1616 | 233.4415 | 90.6206 | 1.657 | 42.2315 | 69.9783 | 25.4864 | 2.631 | 42.1058 | 110.7786 | 16.004 | 21.4157 | 6.2474 | 180.7569 | 312.2982 | 240.0685 | 305.7459 | 212.0258
2012-10-30 01:33 | 25.8255 | 333.9071 | 261.6816 | 327.3778 | 233.5069 | 90.6565 | 1.6569 | 42.2288 | 69.9683 | 25.4869 | 2.6307 | 42.1033 | 110.7625 | 16.0044 | 21.41 | 6.2499 | 180.7308 | 312.4971 | 240.2716 | 305.9678 | 212.0969
2012-10-30 01:33 | 25.821 | 334.0853 | 261.8484 | 327.5753 | 233.4906 | 90.6692 | 1.6568 | 42.2267 | 69.9607 | 25.4871 | 2.6305 | 42.101 | 110.748 | 16.0048 | 21.4456 | 6.2507 | 180.7086 | 312.6397 | 240.4028 | 306.1297 | 212.045
2012-10-30 01:33 | 25.8263 | 334.2201 | 261.9277 | 327.714 | 233.5038 | 90.6888 | 1.6556 | 42.1961 | 69.8584 | 25.4875 | 2.6286 | 42.0702 | 110.5854 | 16.0049 | 21.4309 | 6.2521 | 180.4438 | 312.7892 | 240.4968 | 306.2831 | 212.0729
2012-10-30 01:33 | 25.8233 | 334.5902 | 262.1662 | 328.1035 | 233.6108 | 90.7212 | 1.6525 | 42.1172 | 69.5993 | 25.4867 | 2.6237 | 41.9917 | 110.1735 | 16.0048 | 21.4239 | 6.2543 | 179.7727 | 313.1663 | 240.7423 | 306.6796 | 212.1869
2012-10-30 01:33 | 25.8278 | 334.9069 | 262.4043 | 328.4853 | 233.1792 | 90.746 | 1.6521 | 42.1083 | 69.5681 | 25.4874 | 2.6231 | 41.9828 | 110.1239 | 16.0052 | 21.4043 | 6.256 | 179.692 | 313.5026 | 241 | 307.081 | 211.7749
2012-10-30 01:33 | 25.8263 | 335.0615 | 262.4184 | 328.6591 | 233.5077 | 90.7501 | 1.6522 | 42.1113 | 69.5772 | 25.4877 | 2.6232 | 41.9858 | 110.138 | 16.0054 | 21.467 | 6.2563 | 179.7151 | 313.5945 | 240.9514 | 307.1921 | 212.0407
2012-10-30 01:33 | 25.8285 | 335.1665 | 262.479 | 328.8061 | 234.162 | 90.7894 | 1.6523 | 42.1141 | 69.5863 | 25.4878 | 2.6233 | 41.9888 | 110.1514 | 16.0058 | 21.4251 | 6.259 | 179.7377 | 313.7414 | 241.0539 | 307.381 | 212.7369
2012-10-30 01:33 | 25.8308 | 335.3249 | 262.6059 | 328.9532 | 233.9561 | 90.7894 | 1.6526 | 42.1213 | 69.6092 | 25.4881 | 2.6238 | 41.9959 | 110.1871 | 16.006 | 21.4033 | 6.259 | 179.7963 | 313.9216 | 241.2026 | 307.5499 | 212.5528
2012-10-30 01:33 | 25.833 | 335.4223 | 262.6976 | 329.1306 | 234.3158 | 90.8067 | 1.6529 | 42.1293 | 69.6336 | 25.4888 | 2.6242 | 42.0038 | 110.2257 | 16.0064 | 21.3894 | 6.2602 | 179.8593 | 314.0329 | 241.3082 | 307.7412 | 212.9264
2012-10-30 01:33 | 25.8308 | 335.5952 | 262.8394 | 329.2846 | 234.2272 | 90.8275 | 1.6531 | 42.1353 | 69.6523 | 25.4892 | 2.6245 | 42.0098 | 110.2547 | 16.0068 | 21.3912 | 6.2616 | 179.907 | 314.204 | 241.4482 | 307.8934 | 212.836
2012-10-30 01:33 | 25.8323 | 335.749 | 262.9071 | 329.3814 | 234.0126 | 90.8449 | 1.655 | 42.187 | 69.8206 | 25.4902 | 2.6276 | 42.061 | 110.5191 | 16.0074 | 21.3968 | 6.2628 | 180.3398 | 314.3522 | 241.5103 | 307.9846 | 212.6158
2012-10-30 01:33 | 25.8368 | 335.8943 | 262.9973 | 329.6031 | 233.8285 | 90.868 | 1.6564 | 42.2239 | 69.9404 | 25.4911 | 2.6298 | 42.0983 | 110.712 | 16.0079 | 21.4213 | 6.2644 | 180.6523 | 314.473 | 241.576 | 308.1818 | 212.4072
2012-10-30 01:33 | 25.83 | 335.9789 | 263.0294 | 329.6915 | 233.8492 | 90.8772 | 1.6569 | 42.238 | 69.9853 | 25.4918 | 2.6306 | 42.1122 | 110.7812 | 16.0085 | 21.4466 | 6.2651 | 180.7665 | 314.5323 | 241.5828 | 308.2449 | 212.4026
2012-10-30 01:33 | 25.8315 | 336.1822 | 263.1438 | 329.8454 | 234.3535 | 90.909 | 1.6574 | 42.25 | 70.0239 | 25.4922 | 2.6313 | 42.1241 | 110.8406 | 16.0089 | 21.4521 | 6.2673 | 180.8646 | 314.7301 | 241.6917 | 308.3933 | 212.9014
2012-10-30 01:34 | 25.8338 | 336.3405 | 263.2862 | 330.0267 | 234.9568 | 90.9269 | 1.6575 | 42.2548 | 70.0389 | 25.4925 | 2.6316 | 42.1291 | 110.865 | 16.0092 | 21.4705 | 6.2685 | 180.904 | 314.87 | 241.8157 | 308.5562 | 213.4863
2012-10-30 01:34 | 25.8315 | 336.4601 | 263.4824 | 330.1692 | 234.9466 | 90.9524 | 1.6576 | 42.2577 | 70.0478 | 25.4928 | 2.6317 | 42.1318 | 110.8769 | 16.0096 | 21.4642 | 6.2703 | 180.9246 | 314.9959 | 242.0182 | 308.705 | 213.4824
2012-10-30 01:34 | 25.8293 | 336.5987 | 263.6048 | 330.2775 | 234.7206 | 90.9535 | 1.6576 | 42.2591 | 70.0503 | 25.4936 | 2.6317 | 42.1334 | 110.8829 | 16.0099 | 21.4259 | 6.2703 | 180.9331 | 315.1728 | 242.1789 | 308.8516 | 213.2947
2012-10-30 01:34 | 25.827 | 336.6878 | 263.6842 | 330.4199 | 234.8047 | 90.9697 | 1.6577 | 42.2601 | 70.0528 | 25.4938 | 2.6317 | 42.1342 | 110.8852 | 16.0102 | 21.4195 | 6.2715 | 180.938 | 315.2683 | 242.2647 | 309.0004 | 213.3852
2012-10-30 01:34 | 25.8278 | 336.8561 | 263.7745 | 330.5884 | 234.9861 | 90.9934 | 1.6576 | 42.2598 | 70.0513 | 25.494 | 2.6316 | 42.1339 | 110.8812 | 16.0105 | 21.4364 | 6.2731 | 180.9325 | 315.4197 | 242.3381 | 309.152 | 213.5497
2012-10-30 01:34 | 25.827 | 336.9809 | 263.8827 | 330.7209 | 235.1699 | 91.0125 | 1.6576 | 42.259 | 70.0473 | 25.4945 | 2.6316 | 42.1331 | 110.876 | 16.0107 | 21.4316 | 6.2744 | 180.9233 | 315.5493 | 242.4511 | 309.2893 | 213.7383
2012-10-30 01:34 | 25.833 | 337.1163 | 263.9821 | 330.864 | 235.4115 | 91.0217 | 1.6575 | 42.2581 | 70.0446 | 25.4945 | 2.6315 | 42.1325 | 110.8712 | 16.0109 | 21.4336 | 6.275 | 180.9158 | 315.6827 | 242.5485 | 309.4304 | 213.9779
2012-10-30 01:34 | 25.8263 | 337.2009 | 264.061 | 330.9183 | 235.5501 | 91.0373 | 1.6575 | 42.2579 | 70.0421 | 25.4951 | 2.6314 | 42.1321 | 110.867 | 16.0112 | 21.4269 | 6.2761 | 180.9091 | 315.774 | 242.6341 | 309.4914 | 214.1232
2012-10-30 01:34 | 25.83 | 337.3264 | 264.2359 | 331.1086 | 235.6638 | 91.0512 | 1.6574 | 42.2558 | 70.0351 | 25.4951 | 2.6312 | 42.1301 | 110.8542 | 16.0115 | 21.4266 | 6.2771 | 180.8894 | 315.8998 | 242.8093 | 309.682 | 214.2372
2012-10-30 01:34 | 25.83 | 337.4673 | 264.3837 | 331.2267 | 235.8326 | 91.0674 | 1.6573 | 42.2549 | 70.031 | 25.4955 | 2.6311 | 42.1291 | 110.8478 | 16.0117 | 21.4105 | 6.2782 | 180.8788 | 316.0568 | 242.9732 | 309.8162 | 214.4221
2012-10-30 01:34 | 25.8338 | 337.5548 | 264.4692 | 331.3752 | 235.8953 | 91.0882 | 1.6573 | 42.2543 | 70.0275 | 25.4961 | 2.6311 | 42.1287 | 110.844 | 16.012 | 21.4223 | 6.2796 | 180.8714 | 316.1325 | 243.0469 | 309.9529 | 214.473
2012-10-30 01:34 | 25.833 | 337.672 | 264.5657 | 331.4849 | 235.9456 | 91.0795 | 1.6572 | 42.2521 | 70.0202 | 25.4961 | 2.6309 | 42.1263 | 110.83 | 16.0122 | 21.4055 | 6.279 | 180.8502 | 316.2665 | 243.1602 | 310.0794 | 214.5401
2012-10-30 01:34 | 25.8225 | 337.8176 | 264.703 | 331.5925 | 235.9822 | 91.1171 | 1.6571 | 42.2505 | 70.014 | 25.4964 | 2.6308 | 42.1248 | 110.8198 | 16.0125 | 21.407 | 6.2816 | 180.8338 | 316.4106 | 243.296 | 310.1855 | 214.5752
2012-10-30 01:34 | 25.83 | 337.9621 | 264.8311 | 331.7561 | 236.2253 | 91.1217 | 1.657 | 42.2491 | 70.0081 | 25.4969 | 2.6306 | 42.1233 | 110.8109 | 16.0126 | 21.4707 | 6.2819 | 180.819 | 316.4914 | 243.3604 | 310.2854 | 214.7546
2012-10-30 01:34 | 25.83 | 338.0611 | 264.9361 | 331.8551 | 236.4098 | 91.1437 | 1.657 | 42.2483 | 70.0052 | 25.497 | 2.6306 | 42.1227 | 110.8061 | 16.0128 | 21.4506 | 6.2834 | 180.8112 | 316.6105 | 243.4855 | 310.4045 | 214.9592
2012-10-30 01:34 | 25.8173 | 338.1511 | 265.0751 | 331.9986 | 236.4089 | 91.1685 | 1.6569 | 42.2472 | 70.001 | 25.4972 | 2.6305 | 42.1215 | 110.7993 | 16.0129 | 21.4339 | 6.2852 | 180.8003 | 316.7172 | 243.6412 | 310.5647 | 214.975
2012-10-30 01:34 | 25.8181 | 338.2927 | 265.1771 | 332.1289 | 236.5981 | 91.1749 | 1.6568 | 42.2453 | 69.9941 | 25.4974 | 2.6303 | 42.1197 | 110.7876 | 16.0133 | 21.4147 | 6.2856 | 180.7817 | 316.878 | 243.7624 | 310.7142 | 215.1834
2012-10-30 01:34 | 25.8248 | 338.4289 | 265.3044 | 332.2498 | 236.703 | 91.2136 | 1.6568 | 42.2444 | 69.9903 | 25.4977 | 2.6302 | 42.1186 | 110.7807 | 16.0134 | 21.4294 | 6.2883 | 180.771 | 316.9995 | 243.875 | 310.8204 | 215.2736
2012-10-30 01:34 | 25.821 | 338.524 | 265.4561 | 332.3717 | 236.7227 | 91.202 | 1.6567 | 42.2428 | 69.9844 | 25.4978 | 2.6301 | 42.117 | 110.7717 | 16.0135 | 21.4537 | 6.2875 | 180.7561 | 317.0703 | 244.0024 | 310.918 | 215.269
2012-10-30 01:34 | 25.8143 | 338.6087 | 265.5194 | 332.4945 | 236.8731 | 91.2124 | 1.6567 | 42.2418 | 69.9811 | 25.4979 | 2.63 | 42.1162 | 110.7659 | 16.0137 | 21.447 | 6.2882 | 180.747 | 317.1617 | 244.0724 | 311.0475 | 215.4261
2012-10-30 01:34 | 25.8181 | 338.738 | 265.601 | 332.5478 | 236.7316 | 91.2211 | 1.6566 | 42.2404 | 69.9757 | 25.4982 | 2.6299 | 42.1148 | 110.7574 | 16.0139 | 21.4187 | 6.2888 | 180.7331 | 317.3193 | 244.1823 | 311.1291 | 215.3129
2012-10-30 01:34 | 25.8263 | 338.849 | 265.7064 | 332.7122 | 236.9204 | 91.2459 | 1.6565 | 42.2389 | 69.9702 | 25.4984 | 2.6298 | 42.1132 | 110.7479 | 16.0141 | 21.4509 | 6.2905 | 180.7181 | 317.3981 | 244.2555 | 311.2613 | 215.4695
2012-10-30 01:34 | 25.8166 | 338.9687 | 265.7628 | 332.8129 | 236.9617 | 91.2529 | 1.6564 | 42.2371 | 69.9636 | 25.4986 | 2.6296 | 42.1115 | 110.7371 | 16.0143 | 21.4252 | 6.291 | 180.7007 | 317.5435 | 244.3376 | 311.3877 | 215.5365
2012-10-30 01:34 | 25.8263 | 339.0507 | 265.8775 | 332.9216 | 237.0656 | 91.2592 | 1.6564 | 42.2358 | 69.9582 | 25.499 | 2.6295 | 42.1102 | 110.729 | 16.0145 | 21.4068 | 6.2914 | 180.6872 | 317.6439 | 244.4707 | 311.5148 | 215.6588
2012-10-30 01:35 | 25.8233 | 339.1619 | 265.964 | 333.0557 | 237.1608 | 91.2777 | 1.6563 | 42.2335 | 69.95 | 25.4992 | 2.6294 | 42.108 | 110.7167 | 16.0146 | 21.4038 | 6.2927 | 180.6667 | 317.7581 | 244.5602 | 311.6519 | 215.757
2012-10-30 01:35 | 25.8173 | 339.2739 | 265.9813 | 333.1182 | 237.143 | 91.2881 | 1.6562 | 42.2313 | 69.9424 | 25.4993 | 2.6292 | 42.1058 | 110.7041 | 16.0147 | 21.3416 | 6.2934 | 180.6465 | 317.9323 | 244.6397 | 311.7766 | 215.8014
2012-10-30 01:35 | 25.8278 | 339.3947 | 266.089 | 333.2316 | 237.2045 | 91.3147 | 1.6561 | 42.2308 | 69.9393 | 25.4998 | 2.6291 | 42.105 | 110.6986 | 16.015 | 21.3602 | 6.2952 | 180.6379 | 318.0345 | 244.7288 | 311.8714 | 215.8443
2012-10-30 01:35 | 25.8323 | 339.5248 | 266.1907 | 333.3541 | 237.3464 | 91.3205 | 1.654 | 42.1756 | 69.7584 | 25.4992 | 2.6257 | 42.0497 | 110.4093 | 16.0148 | 21.3566 | 6.2956 | 180.1676 | 318.1682 | 244.8341 | 311.9975 | 215.9898
2012-10-30 01:35 | 25.8285 | 339.5819 | 266.2452 | 333.495 | 237.48 | 91.3124 | 1.6522 | 42.1291 | 69.6064 | 25.4985 | 2.6229 | 42.004 | 110.1706 | 16.0146 | 21.3849 | 6.2951 | 179.7771 | 318.197 | 244.8603 | 312.1101 | 216.0951
2012-10-30 01:35 | 25.8225 | 339.6368 | 266.2742 | 333.588 | 237.3562 | 91.3188 | 1.6517 | 42.1164 | 69.5642 | 25.4987 | 2.6221 | 41.9912 | 110.103 | 16.0146 | 21.3388 | 6.2955 | 179.6673 | 318.298 | 244.9354 | 312.2492 | 216.0174
2012-10-30 01:35 | 25.8173 | 339.7115 | 266.2729 | 333.6856 | 236.613 | 91.3448 | 1.6516 | 42.1129 | 69.5524 | 25.4987 | 2.6218 | 41.9873 | 110.082 | 16.0148 | 21.3416 | 6.2973 | 179.6345 | 318.3699 | 244.9313 | 312.344 | 215.2714
2012-10-30 01:35 | 25.8158 | 339.8204 | 266.2947 | 333.8098 | 236.8549 | 91.3546 | 1.6516 | 42.1146 | 69.5576 | 25.4989 | 2.6219 | 41.9893 | 110.091 | 16.0149 | 21.3963 | 6.298 | 179.6486 | 318.4241 | 244.8984 | 312.4135 | 215.4586
2012-10-30 01:35 | 25.8166 | 339.9087 | 266.2877 | 333.902 | 237.158 | 91.354 | 1.6517 | 42.1183 | 69.5688 | 25.4993 | 2.6221 | 41.9927 | 110.1082 | 16.015 | 21.4051 | 6.2979 | 179.677 | 318.5036 | 244.8826 | 312.4969 | 215.7529
2012-10-30 01:35 | 25.8038 | 339.9606 | 266.3449 | 333.9996 | 237.7693 | 91.3783 | 1.6519 | 42.1235 | 69.5851 | 25.4995 | 2.6224 | 41.9982 | 110.1361 | 16.0152 | 21.3723 | 6.2996 | 179.7212 | 318.5883 | 244.9726 | 312.6273 | 216.397
2012-10-30 01:35 | 25.7993 | 340.0741 | 266.5193 | 334.0522 | 237.4704 | 91.3951 | 1.6522 | 42.13 | 69.6064 | 25.4996 | 2.6227 | 42.0044 | 110.1665 | 16.0155 | 21.3477 | 6.3008 | 179.773 | 318.7264 | 245.1716 | 312.7045 | 216.1227
2012-10-30 01:35 | 25.7926 | 340.212 | 266.5514 | 334.213 | 237.499 | 91.402 | 1.6524 | 42.1372 | 69.6294 | 25.4999 | 2.6232 | 42.0116 | 110.2038 | 16.0156 | 21.353 | 6.3013 | 179.8332 | 318.859 | 245.1984 | 312.86 | 216.146
2012-10-30 01:35 | 25.7889 | 340.2197 | 266.606 | 334.2246 | 237.7034 | 91.402 | 1.6526 | 42.1421 | 69.6446 | 25.5003 | 2.6234 | 42.0165 | 110.2271 | 16.0159 | 21.3493 | 6.3013 | 179.8716 | 318.8704 | 245.2567 | 312.8753 | 216.3541
2012-10-30 01:35 | 25.7844 | 340.3218 | 266.6443 | 334.2924 | 237.64 | 91.4124 | 1.6555 | 42.2184 | 69.8937 | 25.5015 | 2.628 | 42.0916 | 110.6171 | 16.0165 | 21.3609 | 6.302 | 180.5108 | 318.9609 | 245.2834 | 312.9315 | 216.2791
2012-10-30 01:35 | 25.7784 | 340.4528 | 266.6966 | 334.4159 | 237.6222 | 91.4344 | 1.6561 | 42.235 | 69.9472 | 25.502 | 2.629 | 42.1094 | 110.7074 | 16.017 | 21.3629 | 6.3035 | 180.6546 | 319.0899 | 245.3337 | 313.053 | 216.2593
2012-10-30 01:35 | 25.7784 | 340.5251 | 266.7394 | 334.4387 | 237.2493 | 91.4523 | 1.6566 | 42.2471 | 69.986 | 25.5025 | 2.6297 | 42.1213 | 110.7678 | 16.0173 | 21.379 | 6.3047 | 180.7538 | 319.1461 | 245.3604 | 313.0597 | 215.8703
2012-10-30 01:35 | 25.7761 | 340.6217 | 266.7759 | 334.5849 | 237.879 | 91.4598 | 1.6569 | 42.2553 | 70.0124 | 25.5027 | 2.6302 | 42.1294 | 110.8083 | 16.0177 | 21.3486 | 6.3052 | 180.8207 | 319.2731 | 245.4273 | 313.2363 | 216.5304
2012-10-30 01:35 | 25.7664 | 340.7262 | 266.8906 | 334.6209 | 238.2892 | 91.4748 | 1.6571 | 42.2615 | 70.0316 | 25.5032 | 2.6305 | 42.1357 | 110.8395 | 16.0179 | 21.3509 | 6.3063 | 180.8712 | 319.3753 | 245.5397 | 313.27 | 216.9383
2012-10-30 01:35 | 25.7769 | 340.7671 | 267.0061 | 334.7608 | 238.178 | 91.4742 | 1.6572 | 42.2638 | 70.0383 | 25.5036 | 2.6306 | 42.1379 | 110.849 | 16.0182 | 21.4056 | 6.3062 | 180.8874 | 319.3615 | 245.6005 | 313.3552 | 216.7724
2012-10-30 01:35 | 25.7634 | 340.8792 | 267.1014 | 334.8349 | 238.1684 | 91.4904 | 1.6571 | 42.2636 | 70.0371 | 25.5038 | 2.6306 | 42.1378 | 110.8478 | 16.0183 | 21.4081 | 6.3073 | 180.8849 | 319.4711 | 245.6933 | 313.4268 | 216.7603
2012-10-30 01:35 | 25.7679 | 340.9484 | 267.1758 | 334.9041 | 238.2554 | 91.4898 | 1.6571 | 42.2638 | 70.0373 | 25.504 | 2.6306 | 42.1381 | 110.8475 | 16.0186 | 21.4447 | 6.3073 | 180.8848 | 319.5037 | 245.7311 | 313.4594 | 216.8107
2012-10-30 01:35 | 25.7597 | 341.0239 | 267.2337 | 334.9797 | 238.5611 | 91.5187 | 1.6571 | 42.2629 | 70.0338 | 25.5041 | 2.6305 | 42.137 | 110.8413 | 16.0187 | 21.3603 | 6.3093 | 180.8751 | 319.6636 | 245.8734 | 313.6194 | 217.2008
2012-10-30 01:35 | 25.7597 | 341.1 | 267.3581 | 335.0482 | 238.4277 | 91.5343 | 1.6571 | 42.2623 | 70.0313 | 25.5044 | 2.6304 | 42.1366 | 110.8379 | 16.0188 | 21.3803 | 6.3104 | 180.8693 | 319.7197 | 245.9778 | 313.6679 | 217.0474
2012-10-30 01:36 | 25.7544 | 341.129 | 267.4072 | 335.1229 | 238.2419 | 91.5378 | 1.657 | 42.2606 | 70.025 | 25.5045 | 2.6303 | 42.1347 | 110.8263 | 16.019 | 21.3269 | 6.3106 | 180.8513 | 319.8021 | 246.0803 | 313.796 | 216.915
2012-10-30 01:36 | 25.7589 | 341.2438 | 267.4272 | 335.2073 | 238.1168 | 91.543 | 1.657 | 42.2602 | 70.0239 | 25.5046 | 2.6302 | 42.1344 | 110.8238 | 16.0192 | 21.3394 | 6.311 | 180.8476 | 319.9044 | 246.0878 | 313.8679 | 216.7774
2012-10-30 01:36 | 25.7604 | 341.3176 | 267.4637 | 335.2964 | 238.3381 | 91.5592 | 1.6569 | 42.2587 | 70.0178 | 25.5049 | 2.6301 | 42.133 | 110.8153 | 16.0193 | 21.365 | 6.3121 | 180.8331 | 319.9526 | 246.0987 | 313.9314 | 216.9731
2012-10-30 01:36 | 25.7604 | 341.3481 | 267.5725 | 335.3535 | 238.652 | 91.5719 | 1.6569 | 42.2582 | 70.0159 | 25.505 | 2.6301 | 42.1325 | 110.8116 | 16.0195 | 21.385 | 6.313 | 180.8275 | 319.9631 | 246.1875 | 313.9685 | 217.267
2012-10-30 01:36 | 25.75 | 341.4404 | 267.6359 | 335.4725 | 238.5514 | 91.5835 | 1.6568 | 42.2567 | 70.0099 | 25.5054 | 2.63 | 42.1309 | 110.8022 | 16.0196 | 21.3626 | 6.3138 | 180.8121 | 320.0778 | 246.2733 | 314.1099 | 217.1888
2012-10-30 01:36 | 25.75 | 341.5355 | 267.7214 | 335.5296 | 238.5553 | 91.5991 | 1.6567 | 42.2551 | 70.0036 | 25.5057 | 2.6298 | 42.1293 | 110.7927 | 16.0198 | 21.3827 | 6.3148 | 180.7963 | 320.1528 | 246.3387 | 314.1469 | 217.1726
2012-10-30 01:36 | 25.747 | 341.5782 | 267.7573 | 335.637 | 238.6426 | 91.6043 | 1.6566 | 42.2532 | 69.9972 | 25.5058 | 2.6297 | 42.1275 | 110.7825 | 16.0199 | 21.3516 | 6.3152 | 180.7797 | 320.2266 | 246.4057 | 314.2854 | 217.291
2012-10-30 01:36 | 25.7627 | 341.6662 | 267.8274 | 335.7098 | 238.8034 | 91.6077 | 1.6565 | 42.2512 | 69.9905 | 25.5059 | 2.6295 | 42.1255 | 110.7705 | 16.0201 | 21.3914 | 6.3154 | 180.761 | 320.2748 | 246.436 | 314.3184 | 217.412
2012-10-30 01:36 | 25.7574 | 341.7598 | 267.9076 | 335.8035 | 238.6883 | 91.6262 | 1.6565 | 42.25 | 69.9861 | 25.5059 | 2.6295 | 42.1242 | 110.7639 | 16.0201 | 21.37 | 6.3167 | 180.7501 | 320.3898 | 246.5376 | 314.4335 | 217.3183
2012-10-30 01:36 | 25.7664 | 341.8716 | 267.9787 | 335.8582 | 238.8071 | 91.6251 | 1.6564 | 42.2491 | 69.9825 | 25.5062 | 2.6294 | 42.1234 | 110.758 | 16.0203 | 21.363 | 6.3166 | 180.7405 | 320.5086 | 246.6157 | 314.4952 | 217.4441
2012-10-30 01:36 | 25.7559 | 341.8953 | 268.0381 | 335.9391 | 239.0595 | 91.6308 | 1.6564 | 42.2475 | 69.977 | 25.5063 | 2.6293 | 42.122 | 110.7504 | 16.0204 | 21.3444 | 6.317 | 180.7274 | 320.5509 | 246.6937 | 314.5947 | 217.7151
2012-10-30 01:36 | 25.7589 | 341.9706 | 268.1072 | 335.9877 | 239.2038 | 91.6517 | 1.6563 | 42.2465 | 69.9741 | 25.506 | 2.6292 | 42.1208 | 110.743 | 16.0205 | 21.3514 | 6.3185 | 180.7171 | 320.6192 | 246.7558 | 314.6363 | 217.8524
2012-10-30 01:36 | 25.7612 | 342.0034 | 268.195 | 336.0205 | 239.2806 | 91.6534 | 1.6563 | 42.2449 | 69.9686 | 25.5062 | 2.6291 | 42.1192 | 110.7349 | 16.0205 | 21.4019 | 6.3186 | 180.7035 | 320.6015 | 246.7931 | 314.6186 | 217.8787
2012-10-30 01:36 | 25.7627 | 342.0543 | 268.2547 | 336.1476 | 239.2939 | 91.6597 | 1.6561 | 42.2427 | 69.9599 | 25.5067 | 2.6289 | 42.1172 | 110.7229 | 16.0207 | 21.3793 | 6.319 | 180.6828 | 320.675 | 246.8754 | 314.7683 | 217.9146
2012-10-30 01:36 | 25.7582 | 342.1373 | 268.3396 | 336.2078 | 239.3992 | 91.6725 | 1.6561 | 42.2416 | 69.9558 | 25.5068 | 2.6288 | 42.1159 | 110.7157 | 16.0207 | 21.3548 | 6.3199 | 180.6715 | 320.7825 | 246.9848 | 314.853 | 218.0444
2012-10-30 01:36 | 25.75 | 342.1938 | 268.4052 | 336.2834 | 239.4302 | 91.6736 | 1.656 | 42.2394 | 69.9488 | 25.5068 | 2.6287 | 42.1137 | 110.7026 | 16.0209 | 21.3907 | 6.32 | 180.6513 | 320.8031 | 247.0145 | 314.8927 | 218.0395
2012-10-30 01:36 | 25.7619 | 342.2704 | 268.4715 | 336.3676 | 239.2538 | 91.6881 | 1.656 | 42.2387 | 69.9462 | 25.5069 | 2.6286 | 42.1132 | 110.6996 | 16.021 | 21.3785 | 6.321 | 180.6457 | 320.8919 | 247.093 | 314.9891 | 217.8753
2012-10-30 01:36 | 25.7507 | 342.3353 | 268.4952 | 336.4287 | 239.3525 | 91.7031 | 1.6559 | 42.2365 | 69.9388 | 25.5069 | 2.6285 | 42.1109 | 110.6877 | 16.021 | 21.3553 | 6.322 | 180.6266 | 320.98 | 247.1399 | 315.0734 | 217.9972
2012-10-30 01:36 | 25.7559 | 342.4204 | 268.582 | 336.4872 | 239.5852 | 91.6985 | 1.6557 | 42.2331 | 69.9272 | 25.507 | 2.6282 | 42.1074 | 110.6683 | 16.0212 | 21.4006 | 6.3217 | 180.5955 | 321.0198 | 247.1814 | 315.0866 | 218.1846
2012-10-30 01:37 | 25.7582 | 342.6092 | 268.763 | 336.6646 | 239.7366 | 91.7216 | 1.6534 | 42.1731 | 69.729 | 25.5069 | 2.6245 | 42.0473 | 110.3522 | 16.0212 | 21.3748 | 6.3233 | 180.0812 | 321.2344 | 247.3882 | 315.2898 | 218.3618
2012-10-30 01:37 | 25.7649 | 342.7338 | 268.8435 | 336.8426 | 238.8566 | 91.736 | 1.6512 | 42.1157 | 69.5403 | 25.5065 | 2.621 | 41.9905 | 110.0559 | 16.021 | 21.3775 | 6.3243 | 179.5963 | 321.3563 | 247.466 | 315.4651 | 217.4791
2012-10-30 01:37 | 25.7739 | 342.7619 | 268.7709 | 336.8973 | 239.2031 | 91.7256 | 1.6511 | 42.1151 | 69.5381 | 25.5067 | 2.6209 | 41.9898 | 110.0517 | 16.0211 | 21.3143 | 6.3236 | 179.5898 | 321.4476 | 247.4566 | 315.583 | 217.8888
2012-10-30 01:37 | 25.7694 | 342.8144 | 268.8286 | 336.9728 | 239.9557 | 91.7314 | 1.6513 | 42.1195 | 69.5524 | 25.5068 | 2.6212 | 41.9942 | 110.0734 | 16.0212 | 21.3459 | 6.324 | 179.6258 | 321.4685 | 247.4827 | 315.6269 | 218.6098
2012-10-30 01:37 | 25.7567 | 342.8664 | 268.9013 | 337.0095 | 239.5978 | 91.7603 | 1.6516 | 42.1267 | 69.5754 | 25.507 | 2.6216 | 42.0015 | 110.1114 | 16.0213 | 21.3412 | 6.326 | 179.6868 | 321.5252 | 247.5601 | 315.6683 | 218.2566
2012-10-30 01:37 | 25.7672 | 342.9302 | 268.9739 | 337.0847 | 239.8947 | 91.747 | 1.6519 | 42.1343 | 69.5999 | 25.5073 | 2.622 | 42.0088 | 110.1488 | 16.0214 | 21.3437 | 6.325 | 179.7487 | 321.5865 | 247.6302 | 315.741 | 218.551
2012-10-30 01:37 | 25.753 | 342.9502 | 268.9792 | 337.1429 | 240.3747 | 91.7649 | 1.6521 | 42.1419 | 69.6245 | 25.5074 | 2.6225 | 42.0164 | 110.187 | 16.0217 | 21.3696 | 6.3263 | 179.8115 | 321.5806 | 247.6096 | 315.7733 | 219.0051
2012-10-30 01:37 | 25.7544 | 343.0087 | 269.0776 | 337.1899 | 240.3878 | 91.7643 | 1.6533 | 42.173 | 69.7257 | 25.508 | 2.6243 | 42.0465 | 110.3428 | 16.022 | 21.4071 | 6.3262 | 180.0686 | 321.6016 | 247.6705 | 315.7828 | 218.9807
2012-10-30 01:37 | 25.753 | 343.0758 | 269.1384 | 337.2038 | 239.9707 | 91.7707 | 1.6556 | 42.2326 | 69.9202 | 25.509 | 2.628 | 42.1067 | 110.6546 | 16.0226 | 21.4017 | 6.3267 | 180.5748 | 321.6741 | 247.7367 | 315.8021 | 218.569
2012-10-30 01:37 | 25.7417 | 343.1444 | 269.0145 | 337.299 | 239.0022 | 91.7904 | 1.656 | 42.2439 | 69.9564 | 25.5094 | 2.6286 | 42.1181 | 110.7125 | 16.0229 | 21.4105 | 6.328 | 180.6689 | 321.7339 | 247.604 | 315.8885 | 217.5917
2012-10-30 01:37 | 25.741 | 343.216 | 268.9206 | 337.3212 | 239.8606 | 91.7898 | 1.6563 | 42.2526 | 69.9848 | 25.5095 | 2.6291 | 42.127 | 110.7566 | 16.0232 | 21.4459 | 6.328 | 180.7414 | 321.7701 | 247.4747 | 315.8753 | 218.4147
2012-10-30 01:37 | 25.7372 | 343.2426 | 268.975 | 337.3593 | 240.0098 | 91.8083 | 1.6566 | 42.2587 | 70.0042 | 25.5099 | 2.6295 | 42.1329 | 110.7868 | 16.0234 | 21.3859 | 6.3293 | 180.791 | 321.8567 | 247.5891 | 315.9734 | 218.6239
2012-10-30 01:37 | 25.7447 | 343.2844 | 269.0835 | 337.4239 | 240.2526 | 91.8221 | 1.6567 | 42.2629 | 70.0174 | 25.5101 | 2.6297 | 42.137 | 110.8066 | 16.0236 | 21.4215 | 6.3302 | 180.824 | 321.8629 | 247.662 | 316.0024 | 218.8311
2012-10-30 01:37 | 25.7432 | 343.359 | 269.148 | 337.4338 | 240.2747 | 91.8285 | 1.6568 | 42.2644 | 70.0221 | 25.5102 | 2.6298 | 42.1388 | 110.8159 | 16.0237 | 21.3759 | 6.3307 | 180.838 | 321.9831 | 247.7721 | 316.0579 | 218.8988
2012-10-30 01:37 | 25.7447 | 343.4252 | 269.2155 | 337.4848 | 240.4134 | 91.8337 | 1.6567 | 42.265 | 70.0224 | 25.5108 | 2.6298 | 42.139 | 110.8166 | 16.0238 | 21.3894 | 6.331 | 180.839 | 322.0358 | 247.8261 | 316.0954 | 219.024
2012-10-30 01:37 | 25.738 | 343.4108 | 269.2671 | 337.4971 | 240.3832 | 91.8297 | 1.6567 | 42.2647 | 70.0218 | 25.5107 | 2.6298 | 42.139 | 110.8152 | 16.0239 | 21.3827 | 6.3307 | 180.837 | 322.0281 | 247.8844 | 316.1144 | 219.0005
2012-10-30 01:37 | 25.75 | 343.4913 | 269.38 | 337.6004 | 240.3481 | 91.8458 | 1.6567 | 42.2634 | 70.0168 | 25.5109 | 2.6296 | 42.1377 | 110.8066 | 16.0242 | 21.3466 | 6.3318 | 180.8235 | 322.1447 | 248.0334 | 316.2538 | 219.0015
2012-10-30 01:37 | 25.747 | 343.534 | 269.4236 | 337.6507 | 240.3138 | 91.8574 | 1.6566 | 42.2628 | 70.0142 | 25.5111 | 2.6296 | 42.1369 | 110.802 | 16.0243 | 21.3997 | 6.3326 | 180.8162 | 322.1343 | 248.0239 | 316.251 | 218.9141
2012-10-30 01:37 | 25.7477 | 343.5727 | 269.4826 | 337.7161 | 240.3968 | 91.8643 | 1.6566 | 42.2624 | 70.0128 | 25.5112 | 2.6295 | 42.1366 | 110.7981 | 16.0246 | 21.3362 | 6.3331 | 180.8109 | 322.2365 | 248.1464 | 316.3799 | 219.0606
2012-10-30 01:37 | 25.7492 | 343.6161 | 269.5113 | 337.7899 | 240.5474 | 91.8724 | 1.6565 | 42.26 | 70.0034 | 25.5117 | 2.6293 | 42.1341 | 110.7839 | 16.0247 | 21.3498 | 6.3337 | 180.7873 | 322.2663 | 248.1615 | 316.4401 | 219.1976
2012-10-30 01:37 | 25.7455 | 343.6504 | 269.5503 | 337.8167 | 240.6417 | 91.869 | 1.6564 | 42.2585 | 69.9982 | 25.5119 | 2.6292 | 42.1327 | 110.7766 | 16.0247 | 21.3822 | 6.3334 | 180.7748 | 322.2682 | 248.1681 | 316.4345 | 219.2595
2012-10-30 01:38 | 25.7582 | 343.7354 | 269.6367 | 337.8522 | 240.6426 | 91.888 | 1.6564 | 42.2583 | 69.998 | 25.5116 | 2.6292 | 42.1325 | 110.7754 | 16.0248 | 21.3467 | 6.3348 | 180.7734 | 322.3887 | 248.29 | 316.5055 | 219.2959
2012-10-30 01:38 | 25.7582 | 343.7583 | 269.7144 | 337.8865 | 240.7642 | 91.8898 | 1.6563 | 42.2558 | 69.989 | 25.5119 | 2.629 | 42.1301 | 110.7613 | 16.0249 | 21.2906 | 6.3349 | 180.7503 | 322.4677 | 248.4238 | 316.5959 | 219.4736
2012-10-30 01:38 | 25.7634 | 343.7939 | 269.7817 | 337.964 | 240.8282 | 91.8967 | 1.6563 | 42.2547 | 69.9852 | 25.5119 | 2.629 | 42.1289 | 110.755 | 16.025 | 21.36 | 6.3354 | 180.7403 | 322.4339 | 248.4217 | 316.604 | 219.4682
2012-10-30 01:38 | 25.7687 | 343.8639 | 269.853 | 338.0378 | 240.9355 | 91.9146 | 1.6562 | 42.2534 | 69.9808 | 25.512 | 2.6288 | 42.1276 | 110.7471 | 16.0251 | 21.3131 | 6.3366 | 180.728 | 322.5508 | 248.5399 | 316.7247 | 219.6224
2012-10-30 01:38 | 25.7672 | 343.9195 | 269.8826 | 338.0592 | 240.8948 | 91.9129 | 1.6561 | 42.2517 | 69.9745 | 25.5122 | 2.6287 | 42.1259 | 110.737 | 16.0253 | 21.3236 | 6.3365 | 180.7115 | 322.5959 | 248.559 | 316.7356 | 219.5712
2012-10-30 01:38 | 25.7672 | 343.9613 | 269.9136 | 338.0972 | 240.9693 | 91.9094 | 1.656 | 42.2494 | 69.967 | 25.5122 | 2.6286 | 42.1237 | 110.7257 | 16.0253 | 21.3276 | 6.3362 | 180.6927 | 322.6337 | 248.586 | 316.7696 | 219.6417
2012-10-30 01:38 | 25.7687 | 344.0693 | 270.0199 | 338.1901 | 240.8179 | 91.9111 | 1.656 | 42.248 | 69.9612 | 25.5126 | 2.6285 | 42.1222 | 110.7174 | 16.0253 | 21.3733 | 6.3364 | 180.6786 | 322.696 | 248.6466 | 316.8168 | 219.4446
2012-10-30 01:38 | 25.7799 | 344.092 | 270.0311 | 338.2203 | 240.7859 | 91.9227 | 1.6559 | 42.2471 | 69.9578 | 25.5128 | 2.6284 | 42.1214 | 110.7123 | 16.0255 | 21.3484 | 6.3372 | 180.6701 | 322.7436 | 248.6827 | 316.8719 | 219.4375
2012-10-30 01:38 | 25.7799 | 344.1224 | 270.0738 | 338.2431 | 241.0565 | 91.9418 | 1.6558 | 42.2449 | 69.9506 | 25.5128 | 2.6283 | 42.1193 | 110.7008 | 16.0255 | 21.3564 | 6.3385 | 180.6514 | 322.766 | 248.7174 | 316.8867 | 219.7001
2012-10-30 01:38 | 25.7866 | 344.1519 | 270.1155 | 338.3107 | 241.0906 | 91.9406 | 1.6558 | 42.244 | 69.9481 | 25.5126 | 2.6282 | 42.1183 | 110.6961 | 16.0254 | 21.3551 | 6.3384 | 180.6442 | 322.7968 | 248.7604 | 316.9556 | 219.7355
2012-10-30 01:38 | 25.7874 | 344.2402 | 270.1512 | 338.3724 | 240.7699 | 91.9383 | 1.6557 | 42.2417 | 69.94 | 25.5128 | 2.6281 | 42.1159 | 110.6829 | 16.0255 | 21.3759 | 6.3382 | 180.6229 | 322.8643 | 248.7753 | 316.9965 | 219.394
2012-10-30 01:38 | 25.7904 | 344.2813 | 270.1698 | 338.4287 | 241.0043 | 91.9528 | 1.6556 | 42.2402 | 69.9344 | 25.5129 | 2.6279 | 42.1143 | 110.6734 | 16.0257 | 21.3789 | 6.3392 | 180.6078 | 322.9024 | 248.7909 | 317.0498 | 219.6254
2012-10-30 01:38 | 25.7949 | 344.2934 | 270.2053 | 338.4789 | 241.1891 | 91.9574 | 1.6555 | 42.2374 | 69.9255 | 25.5129 | 2.6277 | 42.1118 | 110.6592 | 16.0258 | 21.3794 | 6.3395 | 180.5846 | 322.914 | 248.8259 | 317.0995 | 219.8097
2012-10-30 01:38 | 25.7926 | 344.3406 | 270.2845 | 338.507 | 241.2653 | 91.9545 | 1.6554 | 42.2356 | 69.9184 | 25.5132 | 2.6276 | 42.1099 | 110.6488 | 16.0259 | 21.3651 | 6.3393 | 180.5672 | 322.9755 | 248.9194 | 317.1419 | 219.9002
2012-10-30 01:38 | 25.7911 | 344.3619 | 270.318 | 338.5246 | 241.3892 | 91.9701 | 1.6553 | 42.2335 | 69.9111 | 25.5135 | 2.6275 | 42.1078 | 110.6369 | 16.026 | 21.3475 | 6.3404 | 180.5479 | 323.0144 | 248.9705 | 317.1771 | 220.0417
2012-10-30 01:38 | 25.7949 | 344.4342 | 270.4033 | 338.5854 | 241.4205 | 91.9747 | 1.6552 | 42.2292 | 69.8967 | 25.5134 | 2.6272 | 42.1036 | 110.6141 | 16.0261 | 21.3594 | 6.3407 | 180.5107 | 323.0748 | 249.0439 | 317.226 | 220.0611
2012-10-30 01:38 | 25.8016 | 344.4637 | 270.5109 | 338.6226 | 241.6193 | 91.9816 | 1.6551 | 42.2287 | 69.8939 | 25.5139 | 2.6272 | 42.1032 | 110.6115 | 16.0261 | 21.3701 | 6.3412 | 180.5054 | 323.0936 | 249.1408 | 317.2525 | 220.2492
2012-10-30 01:38 | 25.8008 | 344.5238 | 270.5179 | 338.7207 | 241.9086 | 91.984 | 1.6551 | 42.2288 | 69.8943 | 25.5138 | 2.6271 | 42.1031 | 110.6099 | 16.0263 | 21.3693 | 6.3414 | 180.5043 | 323.1545 | 249.1486 | 317.3514 | 220.5393
2012-10-30 01:38 | 25.8038 | 344.542 | 270.6451 | 338.7352 | 241.5313 | 91.9822 | 1.6533 | 42.1824 | 69.7419 | 25.5134 | 2.6241 | 42.0533 | 110.3506 | 16.026 | 21.3482 | 6.3413 | 180.0925 | 323.1938 | 249.2969 | 317.387 | 220.1831
2012-10-30 01:38 | 25.8113 | 344.5952 | 270.672 | 338.7845 | 241.5388 | 91.9736 | 1.6514 | 42.1301 | 69.5715 | 25.5125 | 2.6211 | 42.0047 | 110.0977 | 16.0257 | 21.3758 | 6.3407 | 179.6691 | 323.2194 | 249.2962 | 317.4087 | 220.163
2012-10-30 01:38 | 25.8121 | 344.615 | 270.6572 | 338.8082 | 241.571 | 91.9811 | 1.6509 | 42.1185 | 69.5338 | 25.5123 | 2.6204 | 41.993 | 110.0374 | 16.0255 | 21.4006 | 6.3412 | 179.5712 | 323.2144 | 249.2566 | 317.4076 | 220.1704
2012-10-30 01:38 | 25.8106 | 344.6629 | 270.7295 | 338.8371 | 241.4244 | 91.9868 | 1.6507 | 42.1117 | 69.5117 | 25.5122 | 2.62 | 41.9863 | 110.003 | 16.0255 | 21.3871 | 6.3416 | 179.5147 | 323.2758 | 249.3424 | 317.45 | 220.0373
2012-10-30 01:39 | 25.8128 | 344.6651 | 270.689 | 338.9078 | 240.8345 | 91.9886 | 1.6507 | 42.1142 | 69.5191 | 25.5124 | 2.6201 | 41.9888 | 110.016 | 16.0255 | 21.4013 | 6.3417 | 179.5351 | 323.2638 | 249.2877 | 317.5065 | 219.4332
2012-10-30 01:39 | 25.8098 | 344.6964 | 270.6006 | 338.8972 | 241.2001 | 91.9834 | 1.6509 | 42.1182 | 69.5323 | 25.5125 | 2.6204 | 41.993 | 110.0382 | 16.0255 | 21.3422 | 6.3413 | 179.5704 | 323.3542 | 249.2584 | 317.555 | 219.8579
2012-10-30 01:39 | 25.8091 | 344.7413 | 270.5727 | 338.965 | 241.4582 | 92.0007 | 1.6511 | 42.1229 | 69.5483 | 25.5124 | 2.6206 | 41.9975 | 110.0606 | 16.0256 | 21.3896 | 6.3425 | 179.6089 | 323.3517 | 249.1831 | 317.5754 | 220.0686
2012-10-30 01:39 | 25.8083 | 344.7367 | 270.5991 | 338.9566 | 241.9749 | 91.9996 | 1.6513 | 42.1281 | 69.5646 | 25.5127 | 2.621 | 42.0029 | 110.0883 | 16.0257 | 21.4249 | 6.3425 | 179.6529 | 323.3118 | 249.1742 | 317.5317 | 220.55
2012-10-30 01:39 | 25.8046 | 344.7863 | 270.6575 | 339.0176 | 241.8379 | 92.0059 | 1.6515 | 42.1352 | 69.5871 | 25.513 | 2.6214 | 42.0099 | 110.124 | 16.0259 | 21.4533 | 6.3429 | 179.7111 | 323.333 | 249.2042 | 317.5643 | 220.3846
2012-10-30 01:39 | 25.8061 | 344.8182 | 270.7017 | 339.0381 | 242.1178 | 92.0129 | 1.6518 | 42.1419 | 69.6092 | 25.513 | 2.6218 | 42.0166 | 110.1586 | 16.0259 | 21.4107 | 6.3434 | 179.7679 | 323.4075 | 249.291 | 317.6274 | 220.7071
2012-10-30 01:39 | 25.8076 | 344.8464 | 270.7498 | 339.0967 | 242.1271 | 92.0186 | 1.6536 | 42.1886 | 69.7621 | 25.5136 | 2.6247 | 42.0641 | 110.4054 | 16.0263 | 21.4202 | 6.3438 | 180.1674 | 323.4262 | 249.3296 | 317.6765 | 220.7069
2012-10-30 01:39 | 25.8091 | 344.8821 | 270.7823 | 339.0792 | 241.6582 | 92.0348 | 1.6551 | 42.2305 | 69.8977 | 25.5146 | 2.6272 | 42.1049 | 110.616 | 16.0268 | 21.4337 | 6.3449 | 180.5137 | 323.4484 | 249.3486 | 317.6455 | 220.2245
2012-10-30 01:39 | 25.8038 | 344.9149 | 270.8042 | 339.0853 | 241.7431 | 92.0186 | 1.6556 | 42.2428 | 69.9379 | 25.5148 | 2.6279 | 42.1169 | 110.6781 | 16.027 | 21.4124 | 6.3438 | 180.616 | 323.5025 | 249.3918 | 317.6729 | 220.3307
2012-10-30 01:39 | 25.8068 | 344.9331 | 270.8344 | 339.1035 | 241.9342 | 92.0383 | 1.6561 | 42.2541 | 69.9752 | 25.5149 | 2.6285 | 42.1281 | 110.7346 | 16.0273 | 21.4314 | 6.3451 | 180.7097 | 323.5017 | 249.403 | 317.6721 | 220.5028
2012-10-30 01:39 | 25.8046 | 344.9765 | 270.9486 | 339.1698 | 242.0026 | 92.0475 | 1.6562 | 42.2584 | 69.9885 | 25.5152 | 2.6288 | 42.1326 | 110.7566 | 16.0275 | 21.4212 | 6.3458 | 180.7451 | 323.5553 | 249.5274 | 317.7486 | 220.5814
2012-10-30 01:39 | 25.8046 | 344.9841 | 270.9486 | 339.2155 | 241.9085 | 92.0527 | 1.6563 | 42.2614 | 69.997 | 25.5157 | 2.6289 | 42.1357 | 110.7717 | 16.0277 | 21.4573 | 6.3461 | 180.7687 | 323.5268 | 249.4913 | 317.7582 | 220.4512
2012-10-30 01:39 | 25.8008 | 345.0412 | 271.034 | 339.1889 | 241.9361 | 92.051 | 1.6564 | 42.2635 | 70.0038 | 25.5158 | 2.6291 | 42.1377 | 110.7824 | 16.0277 | 21.4094 | 6.346 | 180.7863 | 323.6318 | 249.6246 | 317.7795 | 220.5267
2012-10-30 01:39 | 25.8121 | 345.0754 | 271.0453 | 339.2611 | 241.8415 | 92.0556 | 1.6564 | 42.264 | 70.0053 | 25.5158 | 2.6291 | 42.1382 | 110.7848 | 16.0277 | 21.4287 | 6.3463 | 180.7901 | 323.6467 | 249.6166 | 317.8324 | 220.4128
2012-10-30 01:39 | 25.8008 | 345.1059 | 271.034 | 339.284 | 241.9126 | 92.0597 | 1.6564 | 42.2641 | 70.0059 | 25.5157 | 2.6291 | 42.1384 | 110.7845 | 16.0279 | 21.3974 | 6.3466 | 180.7904 | 323.7085 | 249.6366 | 317.8866 | 220.5152
2012-10-30 01:39 | 25.8091 | 345.1484 | 271.0967 | 339.2999 | 242.0855 | 92.0724 | 1.6564 | 42.2634 | 70.0032 | 25.5159 | 2.629 | 42.1378 | 110.7804 | 16.0281 | 21.4137 | 6.3475 | 180.7835 | 323.7347 | 249.683 | 317.8862 | 220.6718
2012-10-30 01:39 | 25.8113 | 345.1354 | 271.1027 | 339.3782 | 242.0799 | 92.0718 | 1.6563 | 42.2614 | 69.9965 | 25.516 | 2.6289 | 42.1358 | 110.7693 | 16.0281 | 21.4239 | 6.3474 | 180.7658 | 323.7115 | 249.6788 | 317.9543 | 220.656
2012-10-30 01:39 | 25.8046 | 345.182 | 271.1348 | 339.4286 | 241.9203 | 92.0735 | 1.6562 | 42.261 | 69.9939 | 25.5163 | 2.6288 | 42.1351 | 110.7661 | 16.0281 | 21.4052 | 6.3475 | 180.76 | 323.7768 | 249.7296 | 318.0234 | 220.5151
2012-10-30 01:39 | 25.8098 | 345.2329 | 271.175 | 339.4491 | 241.996 | 92.0891 | 1.6562 | 42.2598 | 69.9897 | 25.5164 | 2.6287 | 42.1339 | 110.759 | 16.0282 | 21.3983 | 6.3486 | 180.7487 | 323.8346 | 249.7767 | 318.0508 | 220.5977
2012-10-30 01:39 | 25.8031 | 345.2223 | 271.2071 | 339.5109 | 241.9423 | 92.0967 | 1.6562 | 42.259 | 69.988 | 25.5161 | 2.6287 | 42.1333 | 110.7557 | 16.0282 | 21.3555 | 6.3491 | 180.7438 | 323.8668 | 249.8516 | 318.1554 | 220.5868
2012-10-30 01:39 | 25.8106 | 345.2869 | 271.2146 | 339.5031 | 241.8479 | 92.0891 | 1.6561 | 42.2575 | 69.9829 | 25.5162 | 2.6286 | 42.1317 | 110.7477 | 16.0282 | 21.4393 | 6.3486 | 180.7306 | 323.8476 | 249.7753 | 318.0638 | 220.4086
2012-10-30 01:39 | 25.8053 | 345.2892 | 271.1705 | 339.5283 | 241.9484 | 92.0897 | 1.656 | 42.2559 | 69.9773 | 25.5163 | 2.6285 | 42.1302 | 110.739 | 16.0283 | 21.3417 | 6.3487 | 180.7163 | 323.9475 | 249.8288 | 318.1866 | 220.6067
2012-10-30 01:40 | 25.8136 | 345.3318 | 271.2176 | 339.567 | 241.9058 | 92.1001 | 1.656 | 42.2549 | 69.9739 | 25.5163 | 2.6284 | 42.1293 | 110.7343 | 16.0282 | 21.354 | 6.3494 | 180.7082 | 323.9778 | 249.8636 | 318.213 | 220.5518
2012-10-30 01:40 | 25.8083 | 345.3645 | 271.2278 | 339.5846 | 242.0769 | 92.1059 | 1.6559 | 42.2539 | 69.9701 | 25.5165 | 2.6284 | 42.1283 | 110.7295 | 16.0282 | 21.3928 | 6.3498 | 180.6996 | 323.9717 | 249.835 | 318.1918 | 220.6841
2012-10-30 01:40 | 25.8143 | 345.3933 | 271.3036 | 339.6172 | 242.2632 | 92.1128 | 1.6559 | 42.2527 | 69.9657 | 25.5166 | 2.6283 | 42.127 | 110.7219 | 16.0283 | 21.3788 | 6.3503 | 180.6876 | 324.0145 | 249.9248 | 318.2384 | 220.8844
2012-10-30 01:40 | 25.8136 | 345.3812 | 271.3185 | 339.6774 | 242.3723 | 92.1157 | 1.6558 | 42.2503 | 69.9576 | 25.5167 | 2.6281 | 42.1246 | 110.7085 | 16.0284 | 21.3821 | 6.3505 | 180.6661 | 323.9991 | 249.9364 | 318.2953 | 220.9902
2012-10-30 01:40 | 25.8068 | 345.4505 | 271.3893 | 339.6744 | 242.2557 | 92.1209 | 1.6557 | 42.2489 | 69.9523 | 25.517 | 2.628 | 42.1231 | 110.7012 | 16.0284 | 21.3793 | 6.3508 | 180.6535 | 324.0712 | 250.01 | 318.2951 | 220.8764
2012-10-30 01:40 | 25.8113 | 345.5273 | 271.4442 | 339.7208 | 242.3739 | 92.1377 | 1.6556 | 42.2463 | 69.9447 | 25.5166 | 2.6279 | 42.1207 | 110.6882 | 16.0284 | 21.3838 | 6.352 | 180.6329 | 324.1435 | 250.0604 | 318.337 | 220.9901
2012-10-30 01:40 | 25.8038 | 345.5084 | 271.4445 | 339.7095 | 242.2919 | 92.114 | 1.6556 | 42.2446 | 69.9385 | 25.5168 | 2.6278 | 42.119 | 110.6795 | 16.0284 | 21.3482 | 6.3503 | 180.6179 | 324.1602 | 250.0963 | 318.3613 | 220.9437
2012-10-30 01:40 | 25.8106 | 345.5532 | 271.4745 | 339.8 | 242.1693 | 92.1325 | 1.6555 | 42.2442 | 69.9368 | 25.5169 | 2.6277 | 42.1185 | 110.6759 | 16.0285 | 21.359 | 6.3516 | 180.6127 | 324.1942 | 250.1155 | 318.441 | 220.8103
2012-10-30 01:40 | 25.8106 | 345.5875 | 271.5056 | 339.7771 | 242.2595 | 92.1452 | 1.6555 | 42.2422 | 69.9303 | 25.5169 | 2.6276 | 42.1165 | 110.6652 | 16.0285 | 21.3871 | 6.3525 | 180.5955 | 324.2004 | 250.1185 | 318.39 | 220.8724
2012-10-30 01:40 | 25.8143 | 345.5988 | 271.5287 | 339.8531 | 242.2906 | 92.1498 | 1.6553 | 42.2397 | 69.9214 | 25.5171 | 2.6274 | 42.1139 | 110.651 | 16.0286 | 21.3547 | 6.3528 | 180.5724 | 324.2441 | 250.174 | 318.4984 | 220.9359
2012-10-30 01:40 | 25.8128 | 345.6315 | 271.5699 | 339.8555 | 242.4851 | 92.1475 | 1.6553 | 42.238 | 69.9155 | 25.5172 | 2.6273 | 42.1123 | 110.6427 | 16.0286 | 21.3532 | 6.3526 | 180.5582 | 324.2783 | 250.2167 | 318.5023 | 221.1319
2012-10-30 01:40 | 25.8136 | 345.659 | 271.6289 | 339.8867 | 242.3997 | 92.1464 | 1.6552 | 42.2366 | 69.9103 | 25.5174 | 2.6272 | 42.1109 | 110.6345 | 16.0287 | 21.3821 | 6.3526 | 180.5448 | 324.2769 | 250.2468 | 318.5046 | 221.0176
2012-10-30 01:40 | 25.8173 | 345.636 | 271.6287 | 339.9589 | 242.4622 | 92.1585 | 1.6552 | 42.2347 | 69.9049 | 25.5171 | 2.6271 | 42.1091 | 110.6245 | 16.0288 | 21.3617 | 6.3534 | 180.5294 | 324.2743 | 250.267 | 318.5972 | 221.1005
2012-10-30 01:40 | 25.8128 | 345.6962 | 271.7134 | 339.9201 | 242.5008 | 92.1492 | 1.655 | 42.2322 | 69.8961 | 25.5173 | 2.6269 | 42.1065 | 110.6102 | 16.0289 | 21.3612 | 6.3528 | 180.5063 | 324.335 | 250.3522 | 318.5589 | 221.1396
2012-10-30 01:40 | 25.8203 | 345.7265 | 271.7558 | 339.9771 | 242.4299 | 92.1539 | 1.655 | 42.2306 | 69.8904 | 25.5175 | 2.6268 | 42.1049 | 110.6016 | 16.0289 | 21.3647 | 6.3531 | 180.492 | 324.3618 | 250.3911 | 318.6124 | 221.0652
2012-10-30 01:40 | 25.8188 | 345.7631 | 271.7698 | 340.0175 | 242.3618 | 92.1741 | 1.6537 | 42.1965 | 69.7785 | 25.5172 | 2.6249 | 42.0743 | 110.4423 | 16.0287 | 21.3752 | 6.3545 | 180.2207 | 324.3879 | 250.3946 | 318.6423 | 220.9866
2012-10-30 01:40 | 25.8225 | 345.763 | 271.727 | 340.0554 | 242.3929 | 92.1672 | 1.6511 | 42.1294 | 69.5595 | 25.5161 | 2.6206 | 42.0039 | 110.0767 | 16.0282 | 21.3749 | 6.354 | 179.6362 | 324.3881 | 250.3521 | 318.6805 | 221.018
2012-10-30 01:40 | 25.8255 | 345.747 | 271.7106 | 340.0394 | 242.2822 | 92.1573 | 1.6507 | 42.1191 | 69.5254 | 25.5161 | 2.62 | 41.9936 | 110.0234 | 16.0281 | 21.398 | 6.3533 | 179.5488 | 324.349 | 250.3126 | 318.6414 | 220.8842
2012-10-30 01:40 | 25.8285 | 345.7538 | 271.6826 | 340.0767 | 241.2306 | 92.1585 | 1.6505 | 42.1147 | 69.5113 | 25.516 | 2.6197 | 41.9895 | 110.0017 | 16.0281 | 21.3889 | 6.3534 | 179.513 | 324.3649 | 250.2937 | 318.6878 | 219.8417
2012-10-30 01:40 | 25.8233 | 345.8056 | 271.5726 | 340.0943 | 241.4959 | 92.162 | 1.6505 | 42.1156 | 69.5137 | 25.5162 | 2.6198 | 41.9903 | 110.0071 | 16.028 | 21.3797 | 6.3536 | 179.5207 | 324.4259 | 250.1929 | 318.7146 | 220.1162
2012-10-30 01:40 | 25.8285 | 345.807 | 271.5196 | 340.1566 | 241.6344 | 92.1527 | 1.6506 | 42.1175 | 69.5206 | 25.516 | 2.6199 | 41.9923 | 110.0172 | 16.028 | 21.4331 | 6.353 | 179.5377 | 324.3739 | 250.0865 | 318.7235 | 220.2013
2012-10-30 01:40 | 25.8203 | 345.7988 | 271.4532 | 340.1712 | 240.7871 | 92.1423 | 1.6509 | 42.1243 | 69.5429 | 25.516 | 2.6203 | 41.9988 | 110.0509 | 16.028 | 21.4048 | 6.3523 | 179.5938 | 324.394 | 250.0484 | 318.7664 | 219.3823
2012-10-30 01:41 | 25.8128 | 345.8218 | 271.3215 | 340.1447 | 241.1325 | 92.1562 | 1.6511 | 42.1304 | 69.5625 | 25.5162 | 2.6207 | 42.0052 | 110.0834 | 16.0282 | 21.4134 | 6.3532 | 179.6459 | 324.4084 | 249.9081 | 318.7313 | 219.7191
2012-10-30 01:41 | 25.8068 | 345.8272 | 271.2961 | 340.1577 | 241.5814 | 92.1585 | 1.6514 | 42.1367 | 69.5829 | 25.5163 | 2.6211 | 42.0113 | 110.1153 | 16.0282 | 21.3913 | 6.3534 | 179.6982 | 324.4359 | 249.9048 | 318.7664 | 220.1901
2012-10-30 01:41 | 25.8068 | 345.8386 | 271.3505 | 340.1882 | 242.0166 | 92.162 | 1.6516 | 42.143 | 69.6027 | 25.5168 | 2.6214 | 42.0176 | 110.1466 | 16.0284 | 21.4034 | 6.3536 | 179.7493 | 324.4352 | 249.9471 | 318.7848 | 220.6132
2012-10-30 01:41 | 25.8106 | 345.8995 | 271.5056 | 340.1805 | 242.9455 | 92.1631 | 1.6541 | 42.2096 | 69.8202 | 25.5177 | 2.6254 | 42.0822 | 110.4818 | 16.029 | 21.4352 | 6.3537 | 180.302 | 324.4643 | 250.0704 | 318.7453 | 221.5103
2012-10-30 01:41 | 25.8008 | 345.9315 | 271.6199 | 340.2316 | 242.3281 | 92.1862 | 1.6552 | 42.2369 | 69.9096 | 25.5181 | 2.6271 | 42.1111 | 110.6313 | 16.0293 | 21.4134 | 6.3553 | 180.541 | 324.5181 | 250.2065 | 318.8182 | 220.9147
2012-10-30 01:41 | 25.8016 | 345.9247 | 271.5936 | 340.221 | 241.8546 | 92.1955 | 1.6556 | 42.2475 | 69.9436 | 25.5185 | 2.6278 | 42.1218 | 110.686 | 16.0295 | 21.4343 | 6.356 | 180.6296 | 324.4904 | 250.1593 | 318.7867 | 220.4203
2012-10-30 01:41 | 25.7949 | 345.918 | 271.5403 | 340.2067 | 242.3536 | 92.203 | 1.6559 | 42.2548 | 69.9677 | 25.5185 | 2.6282 | 42.1287 | 110.7226 | 16.0295 | 21.4035 | 6.3565 | 180.6903 | 324.5145 | 250.1368 | 318.8032 | 220.9501
2012-10-30 01:41 | 25.7949 | 345.9523 | 271.5675 | 340.2257 | 242.2791 | 92.2094 | 1.6561 | 42.2606 | 69.9864 | 25.5186 | 2.6285 | 42.1347 | 110.7527 | 16.0297 | 21.4195 | 6.3569 | 180.7391 | 324.5328 | 250.148 | 318.8062 | 220.8596
2012-10-30 01:41 | 25.7926 | 345.988 | 271.6273 | 340.2463 | 242.7001 | 92.2053 | 1.6561 | 42.2626 | 69.9932 | 25.5186 | 2.6286 | 42.1367 | 110.7623 | 16.0298 | 21.3932 | 6.3566 | 180.7556 | 324.5948 | 250.2341 | 318.8531 | 221.3069
2012-10-30 01:41 | 25.7964 | 345.9994 | 271.6582 | 340.2653 | 242.6412 | 92.2169 | 1.6561 | 42.2626 | 69.9926 | 25.5188 | 2.6286 | 42.1368 | 110.7619 | 16.03 | 21.413 | 6.3574 | 180.7545 | 324.5864 | 250.2452 | 318.8523 | 221.2282
2012-10-30 01:41 | 25.7904 | 346.0124 | 271.7182 | 340.3126 | 242.7097 | 92.2273 | 1.6561 | 42.2627 | 69.9925 | 25.519 | 2.6287 | 42.1371 | 110.7644 | 16.0299 | 21.3709 | 6.3582 | 180.7569 | 324.6415 | 250.3473 | 318.9417 | 221.3388
2012-10-30 01:41 | 25.8001 | 346.0297 | 271.7473 | 340.3223 | 242.398 | 92.2198 | 1.6561 | 42.2622 | 69.9903 | 25.5192 | 2.6286 | 42.1364 | 110.76 | 16.0299 | 21.3726 | 6.3576 | 180.7503 | 324.6571 | 250.3747 | 318.9497 | 221.0254
2012-10-30 01:41 | 25.7979 | 346.0009 | 271.7373 | 340.3429 | 242.137 | 92.2371 | 1.6561 | 42.2625 | 69.9912 | 25.5192 | 2.6286 | 42.1368 | 110.7621 | 16.0299 | 21.3945 | 6.3588 | 180.7533 | 324.6064 | 250.3428 | 318.9484 | 220.7425
2012-10-30 01:41 | 25.7874 | 346.0475 | 271.7229 | 340.2867 | 242.3617 | 92.229 | 1.6561 | 42.2617 | 69.9891 | 25.519 | 2.6286 | 42.1357 | 110.7563 | 16.03 | 21.3559 | 6.3583 | 180.7454 | 324.6916 | 250.367 | 318.9308 | 221.0058
2012-10-30 01:41 | 25.7934 | 346.0573 | 271.822 | 340.3117 | 242.4618 | 92.2336 | 1.656 | 42.2607 | 69.9857 | 25.519 | 2.6285 | 42.135 | 110.7528 | 16.0299 | 21.3859 | 6.3586 | 180.7385 | 324.6714 | 250.4361 | 318.9258 | 221.0759
2012-10-30 01:41 | 25.7956 | 346.0633 | 271.7622 | 340.3748 | 242.5228 | 92.2394 | 1.656 | 42.2598 | 69.9822 | 25.5191 | 2.6285 | 42.1339 | 110.7472 | 16.0299 | 21.4002 | 6.359 | 180.7294 | 324.6631 | 250.362 | 318.9746 | 221.1226
2012-10-30 01:41 | 25.7964 | 346.0907 | 271.8444 | 340.3794 | 242.5746 | 92.2394 | 1.6559 | 42.2576 | 69.9749 | 25.5193 | 2.6283 | 42.1319 | 110.7363 | 16.0299 | 21.3889 | 6.359 | 180.7112 | 324.7018 | 250.4555 | 318.9905 | 221.1857
2012-10-30 01:41 | 25.7911 | 346.093 | 271.8275 | 340.4236 | 242.5614 | 92.2406 | 1.6558 | 42.2558 | 69.9688 | 25.5192 | 2.6282 | 42.1302 | 110.7267 | 16.03 | 21.4117 | 6.3591 | 180.6955 | 324.6813 | 250.4158 | 319.0119 | 221.1497
2012-10-30 01:41 | 25.7971 | 346.1371 | 271.8568 | 340.4563 | 242.638 | 92.2458 | 1.6556 | 42.2507 | 69.9515 | 25.5194 | 2.6279 | 42.1248 | 110.6983 | 16.0301 | 21.4137 | 6.3594 | 180.6498 | 324.7234 | 250.4431 | 319.0426 | 221.2243
2012-10-30 01:41 | 25.7911 | 346.1159 | 271.9284 | 340.4274 | 242.8515 | 92.2348 | 1.6556 | 42.2491 | 69.946 | 25.5195 | 2.6277 | 42.1233 | 110.6894 | 16.0302 | 21.4438 | 6.3587 | 180.6354 | 324.6721 | 250.4846 | 318.9836 | 221.4077
2012-10-30 01:41 | 25.7926 | 346.1402 | 271.9686 | 340.4822 | 242.8491 | 92.2475 | 1.6555 | 42.2479 | 69.9421 | 25.5195 | 2.6276 | 42.1221 | 110.6821 | 16.0303 | 21.4172 | 6.3595 | 180.6241 | 324.723 | 250.5514 | 319.065 | 221.4319
2012-10-30 01:41 | 25.7919 | 346.1243 | 272.0223 | 340.4967 | 242.9424 | 92.2539 | 1.6554 | 42.246 | 69.9357 | 25.5195 | 2.6275 | 42.1203 | 110.6727 | 16.0303 | 21.4005 | 6.36 | 180.6084 | 324.7238 | 250.6218 | 319.0962 | 221.5419
2012-10-30 01:42 | 25.7934 | 346.1714 | 272.0625 | 340.5363 | 242.9126 | 92.2579 | 1.6554 | 42.2454 | 69.9336 | 25.5195 | 2.6275 | 42.1198 | 110.6693 | 16.0304 | 21.3819 | 6.3603 | 180.6028 | 324.7895 | 250.6806 | 319.1544 | 221.5307
2012-10-30 01:42 | 25.7881 | 346.1699 | 272.0883 | 340.5462 | 242.97 | 92.266 | 1.6553 | 42.2436 | 69.9274 | 25.5196 | 2.6274 | 42.1179 | 110.6601 | 16.0303 | 21.3847 | 6.3608 | 180.5875 | 324.7852 | 250.7036 | 319.1615 | 221.5853
2012-10-30 01:42 | 25.7941 | 346.1873 | 272.1176 | 340.5179 | 242.9015 | 92.2539 | 1.6553 | 42.2427 | 69.9245 | 25.5196 | 2.6273 | 42.117 | 110.6558 | 16.0303 | 21.3987 | 6.36 | 180.5804 | 324.7886 | 250.7189 | 319.1192 | 221.5028
2012-10-30 01:42 | 25.7941 | 346.2216 | 272.1331 | 340.5484 | 242.878 | 92.2625 | 1.6553 | 42.2415 | 69.9205 | 25.5196 | 2.6273 | 42.1158 | 110.6497 | 16.0302 | 21.4268 | 6.3606 | 180.5702 | 324.7948 | 250.7063 | 319.1216 | 221.4512
2012-10-30 01:42 | 25.8016 | 346.2177 | 272.1522 | 340.6206 | 242.9639 | 92.2683 | 1.6552 | 42.2407 | 69.918 | 25.5196 | 2.6272 | 42.1151 | 110.6457 | 16.0303 | 21.3821 | 6.361 | 180.5636 | 324.8356 | 250.7701 | 319.2385 | 221.5818
2012-10-30 01:42 | 25.7986 | 346.2299 | 272.157 | 340.6214 | 242.8394 | 92.2671 | 1.6551 | 42.2381 | 69.909 | 25.5197 | 2.6271 | 42.1123 | 110.6315 | 16.0302 | 21.3871 | 6.3609 | 180.5405 | 324.8428 | 250.7699 | 319.2343 | 221.4523
2012-10-30 01:42 | 25.8068 | 346.2495 | 272.1225 | 340.6258 | 242.4909 | 92.2533 | 1.6551 | 42.2364 | 69.9034 | 25.5196 | 2.627 | 42.1106 | 110.6229 | 16.0301 | 21.3793 | 6.3599 | 180.5263 | 324.8702 | 250.7432 | 319.2465 | 221.1116
2012-10-30 01:42 | 25.8061 | 346.3097 | 272.1334 | 340.6327 | 242.745 | 92.2712 | 1.655 | 42.2342 | 69.8961 | 25.5197 | 2.6268 | 42.1084 | 110.6115 | 16.0301 | 21.3465 | 6.3612 | 180.5076 | 324.9632 | 250.7869 | 319.2862 | 221.3985
2012-10-30 01:42 | 25.8091 | 346.2822 | 272.1442 | 340.6737 | 242.6147 | 92.2723 | 1.6549 | 42.233 | 69.8923 | 25.5196 | 2.6268 | 42.1072 | 110.6054 | 16.0301 | 21.3374 | 6.3613 | 180.4977 | 324.9448 | 250.8068 | 319.3363 | 221.2773
2012-10-30 01:42 | 25.8121 | 346.3271 | 272.1045 | 340.6767 | 242.4139 | 92.2764 | 1.6549 | 42.2314 | 69.8875 | 25.5195 | 2.6267 | 42.1057 | 110.5971 | 16.0302 | 21.3445 | 6.3615 | 180.4846 | 324.9826 | 250.76 | 319.3322 | 221.0694
2012-10-30 01:42 | 25.8083 | 346.3005 | 272.1124 | 340.6615 | 242.9628 | 92.2683 | 1.6525 | 42.1694 | 69.6836 | 25.5191 | 2.623 | 42.0467 | 110.2905 | 16.0298 | 21.3326 | 6.361 | 179.9741 | 324.9679 | 250.7798 | 319.3289 | 221.6302
2012-10-30 01:42 | 25.8136 | 346.2905 | 272.1681 | 340.6782 | 243.2071 | 92.2521 | 1.6511 | 42.1327 | 69.5644 | 25.5183 | 2.6207 | 42.0073 | 110.0868 | 16.0293 | 21.346 | 6.3599 | 179.6512 | 324.9445 | 250.8221 | 319.3322 | 221.8611
2012-10-30 01:42 | 25.8151 | 346.3491 | 272.2006 | 340.7406 | 242.9382 | 92.251 | 1.6507 | 42.1217 | 69.5288 | 25.5181 | 2.62 | 41.9961 | 110.0288 | 16.0292 | 21.3475 | 6.3598 | 179.5577 | 325.0016 | 250.8531 | 319.3931 | 221.5907
2012-10-30 01:42 | 25.8173 | 346.3171 | 272.1912 | 340.7619 | 242.6425 | 92.27 | 1.6506 | 42.12 | 69.5228 | 25.5182 | 2.6199 | 41.9947 | 110.0209 | 16.0293 | 21.3577 | 6.3611 | 179.5437 | 324.9594 | 250.8335 | 319.4042 | 221.2848
2012-10-30 01:42 | 25.8181 | 346.3597 | 272.1338 | 340.7512 | 242.7139 | 92.2614 | 1.6506 | 42.1205 | 69.5251 | 25.5179 | 2.6199 | 41.9953 | 110.0245 | 16.0292 | 21.3705 | 6.3605 | 179.5497 | 324.9892 | 250.7633 | 319.3807 | 221.3434
2012-10-30 01:42 | 25.8173 | 346.3551 | 272.164 | 340.7657 | 242.4152 | 92.2475 | 1.6507 | 42.1211 | 69.5272 | 25.5179 | 2.62 | 41.9958 | 110.0277 | 16.0291 | 21.3456 | 6.3595 | 179.555 | 325.0095 | 250.8184 | 319.4201 | 221.0696
2012-10-30 01:42 | 25.8076 | 346.3302 | 272.1 | 340.7522 | 242.5231 | 92.2521 | 1.6509 | 42.1283 | 69.5504 | 25.5181 | 2.6204 | 42.0029 | 110.065 | 16.0291 | 21.3801 | 6.3599 | 179.6154 | 324.9501 | 250.7199 | 319.3721 | 221.143
2012-10-30 01:42 | 25.8203 | 346.3315 | 272.0856 | 340.7763 | 241.9634 | 92.2533 | 1.6511 | 42.1324 | 69.5638 | 25.5182 | 2.6207 | 42.0071 | 110.086 | 16.0293 | 21.3928 | 6.3599 | 179.6498 | 324.9387 | 250.6928 | 319.3835 | 220.5706
2012-10-30 01:42 | 25.8136 | 346.3704 | 271.9237 | 340.7582 | 241.1725 | 92.2695 | 1.6513 | 42.1382 | 69.5828 | 25.5182 | 2.621 | 42.0127 | 110.1153 | 16.0293 | 21.37 | 6.3611 | 179.6981 | 325.0004 | 250.5537 | 319.3882 | 219.8025
2012-10-30 01:42 | 25.8098 | 346.3628 | 271.7492 | 340.7468 | 241.6942 | 92.2481 | 1.6515 | 42.1433 | 69.5993 | 25.5183 | 2.6213 | 42.0179 | 110.1419 | 16.0293 | 21.3662 | 6.3596 | 179.7412 | 324.9966 | 250.383 | 319.3806 | 220.328
2012-10-30 01:42 | 25.8053 | 346.3469 | 271.6826 | 340.7537 | 242.3169 | 92.2666 | 1.6537 | 42.2005 | 69.7865 | 25.519 | 2.6248 | 42.0757 | 110.442 | 16.0298 | 21.3657 | 6.3609 | 180.2285 | 324.9812 | 250.3169 | 319.388 | 220.9512
2012-10-30 01:42 | 25.8016 | 346.4155 | 271.7759 | 340.7309 | 242.2779 | 92.2712 | 1.6551 | 42.2367 | 69.9044 | 25.5197 | 2.627 | 42.111 | 110.6244 | 16.0302 | 21.3821 | 6.3612 | 180.5289 | 325.0334 | 250.3938 | 319.3488 | 220.8958
2012-10-30 01:42 | 25.8008 | 346.4299 | 271.8062 | 340.7872 | 242.3281 | 92.2631 | 1.6556 | 42.2495 | 69.9468 | 25.5197 | 2.6278 | 42.1238 | 110.6912 | 16.0303 | 21.4014 | 6.3606 | 180.638 | 325.0285 | 250.4048 | 319.3858 | 220.9267
2012-10-30 01:43 | 25.7964 | 346.4141 | 271.8367 | 340.7714 | 242.6059 | 92.2804 | 1.6559 | 42.2594 | 69.9782 | 25.5202 | 2.6283 | 42.1336 | 110.7416 | 16.0304 | 21.397 | 6.3618 | 180.7198 | 325.0171 | 250.4397 | 319.3744 | 221.2089
2012-10-30 01:43 | 25.7896 | 346.4416 | 271.8687 | 340.7608 | 242.5795 | 92.2706 | 1.6561 | 42.2634 | 69.9915 | 25.5202 | 2.6286 | 42.1376 | 110.7612 | 16.0307 | 21.4343 | 6.3611 | 180.7526 | 325.0073 | 250.4344 | 319.3265 | 221.1452
2012-10-30 01:43 | 25.7836 | 346.4508 | 271.8859 | 340.7662 | 242.7734 | 92.2833 | 1.6562 | 42.2671 | 70.003 | 25.5204 | 2.6288 | 42.1412 | 110.7796 | 16.0308 | 21.4042 | 6.362 | 180.7827 | 325.0466 | 250.4817 | 319.362 | 221.3692
2012-10-30 01:43 | 25.7859 | 346.4759 | 272.0046 | 340.7647 | 242.7561 | 92.2926 | 1.6562 | 42.2681 | 70.006 | 25.5205 | 2.6288 | 42.1422 | 110.7842 | 16.0308 | 21.4466 | 6.3627 | 180.7902 | 325.0293 | 250.558 | 319.3181 | 221.3095
2012-10-30 01:43 | 25.7904 | 346.5413 | 271.9936 | 340.7958 | 242.8978 | 92.2989 | 1.6562 | 42.2686 | 70.007 | 25.5208 | 2.6288 | 42.1427 | 110.7865 | 16.0309 | 21.407 | 6.3631 | 180.7934 | 325.1343 | 250.5866 | 319.3888 | 221.4908
2012-10-30 01:43 | 25.7941 | 346.5716 | 272.0438 | 340.7843 | 243.0936 | 92.2897 | 1.6562 | 42.2683 | 70.0057 | 25.5209 | 2.6288 | 42.1425 | 110.7843 | 16.031 | 21.4669 | 6.3625 | 180.79 | 325.1047 | 250.5769 | 319.3174 | 221.6267
2012-10-30 01:43 | 25.7866 | 346.5603 | 272.1023 | 340.8149 | 242.8823 | 92.2995 | 1.6562 | 42.2667 | 70.0002 | 25.521 | 2.6287 | 42.1408 | 110.7753 | 16.0311 | 21.3711 | 6.3631 | 180.7755 | 325.1892 | 250.7312 | 319.4438 | 221.5112
2012-10-30 01:43 | 25.7941 | 346.5982 | 272.1563 | 340.8528 | 242.9211 | 92.3059 | 1.6561 | 42.2655 | 69.9965 | 25.5209 | 2.6286 | 42.1397 | 110.7697 | 16.031 | 21.3826 | 6.3636 | 180.7661 | 325.2156 | 250.7737 | 319.4702 | 221.5385
2012-10-30 01:43 | 25.7964 | 346.5891 | 272.1586 | 340.8171 | 242.8881 | 92.3093 | 1.6561 | 42.2651 | 69.9951 | 25.5209 | 2.6286 | 42.1392 | 110.7667 | 16.0311 | 21.3729 | 6.3638 | 180.7619 | 325.2162 | 250.7857 | 319.4442 | 221.5152
2012-10-30 01:43 | 25.7821 | 346.6319 | 272.1754 | 340.8522 | 242.9248 | 92.3192 | 1.6561 | 42.2645 | 69.9925 | 25.5211 | 2.6286 | 42.1387 | 110.7638 | 16.0312 | 21.3987 | 6.3645 | 180.7562 | 325.2332 | 250.7767 | 319.4535 | 221.5261
2012-10-30 01:43 | 25.7866 | 346.6174 | 272.2109 | 340.8948 | 243.1566 | 92.3203 | 1.656 | 42.2621 | 69.985 | 25.5209 | 2.6284 | 42.1364 | 110.7511 | 16.0312 | 21.4112 | 6.3646 | 180.7361 | 325.2062 | 250.7997 | 319.4836 | 221.7454
2012-10-30 01:43 | 25.7881 | 346.6645 | 272.2396 | 340.8582 | 242.9778 | 92.3168 | 1.6559 | 42.2601 | 69.977 | 25.5214 | 2.6283 | 42.1344 | 110.7403 | 16.0313 | 21.3365 | 6.3643 | 180.7174 | 325.328 | 250.9031 | 319.5217 | 221.6413
2012-10-30 01:43 | 25.7896 | 346.6584 | 272.2721 | 340.9168 | 242.9911 | 92.3088 | 1.6558 | 42.2588 | 69.9736 | 25.5211 | 2.6282 | 42.1331 | 110.7334 | 16.0313 | 21.3902 | 6.3638 | 180.707 | 325.2682 | 250.8819 | 319.5266 | 221.6009
2012-10-30 01:43 | 25.7836 | 346.66 | 272.27 | 340.8842 | 242.9537 | 92.3134 | 1.6557 | 42.2568 | 69.9666 | 25.5213 | 2.6281 | 42.1311 | 110.723 | 16.0312 | 21.3842 | 6.3641 | 180.6896 | 325.2758 | 250.8858 | 319.5 | 221.5695
2012-10-30 01:43 | 25.7829 | 346.6631 | 272.3236 | 340.9482 | 242.953 | 92.3174 | 1.6557 | 42.2565 | 69.9658 | 25.5212 | 2.628 | 42.1307 | 110.7208 | 16.0313 | 21.3674 | 6.3644 | 180.6866 | 325.2957 | 250.9562 | 319.5808 | 221.5856
2012-10-30 01:43 | 25.7911 | 346.6447 | 272.3162 | 340.9183 | 242.9848 | 92.3215 | 1.6557 | 42.2552 | 69.962 | 25.5211 | 2.628 | 42.1294 | 110.7141 | 16.0313 | 21.3636 | 6.3646 | 180.6761 | 325.2811 | 250.9526 | 319.5547 | 221.6212
2012-10-30 01:43 | 25.7844 | 346.6875 | 272.2902 | 340.9573 | 242.8957 | 92.3301 | 1.6556 | 42.2535 | 69.9557 | 25.5212 | 2.6279 | 42.1277 | 110.7056 | 16.0312 | 21.3368 | 6.3652 | 180.6613 | 325.3507 | 250.9534 | 319.6205 | 221.5589
2012-10-30 01:43 | 25.7896 | 346.7573 | 272.3419 | 340.932 | 243.1675 | 92.3249 | 1.6555 | 42.251 | 69.9472 | 25.5213 | 2.6277 | 42.1254 | 110.6936 | 16.0312 | 21.3541 | 6.3649 | 180.6408 | 325.4032 | 250.9878 | 319.5779 | 221.8134
2012-10-30 01:43 | 25.7874 | 346.7399 | 272.3591 | 340.9679 | 243.0124 | 92.3244 | 1.6554 | 42.2488 | 69.94 | 25.5213 | 2.6276 | 42.1232 | 110.6816 | 16.0312 | 21.3318 | 6.3648 | 180.6216 | 325.4081 | 251.0273 | 319.6361 | 221.6806
2012-10-30 01:43 | 25.7919 | 346.7787 | 272.3403 | 341.0028 | 242.6877 | 92.3197 | 1.6553 | 42.2467 | 69.9326 | 25.5215 | 2.6274 | 42.1209 | 110.6698 | 16.0312 | 21.3764 | 6.3645 | 180.6023 | 325.4023 | 250.9639 | 319.6264 | 221.3113
2012-10-30 01:43 | 25.7926 | 346.7565 | 272.3022 | 340.9845 | 242.5904 | 92.3382 | 1.6553 | 42.2443 | 69.9255 | 25.5212 | 2.6273 | 42.1187 | 110.6586 | 16.0311 | 21.3691 | 6.3658 | 180.5841 | 325.3874 | 250.9331 | 319.6154 | 221.2213
2012-10-30 01:43 | 25.7956 | 346.7443 | 272.2858 | 341.0218 | 242.954 | 92.3324 | 1.6552 | 42.2436 | 69.9229 | 25.5213 | 2.6273 | 42.118 | 110.6552 | 16.0311 | 21.332 | 6.3654 | 180.578 | 325.4123 | 250.9538 | 319.6898 | 221.622
2012-10-30 01:44 | 25.8001 | 346.7716 | 272.3252 | 341.0225 | 242.8997 | 92.3324 | 1.6551 | 42.2417 | 69.9163 | 25.5214 | 2.6271 | 42.116 | 110.6436 | 16.0313 | 21.3726 | 6.3654 | 180.5599 | 325.399 | 250.9526 | 319.6499 | 221.5271
2012-10-30 01:44 | 25.7956 | 346.7709 | 272.3363 | 341.0332 | 242.9618 | 92.3411 | 1.6551 | 42.2405 | 69.9122 | 25.5214 | 2.6271 | 42.1147 | 110.6382 | 16.0311 | 21.336 | 6.366 | 180.5504 | 325.4349 | 251.0003 | 319.6972 | 221.6258
2012-10-30 01:44 | 25.8023 | 346.7929 | 272.3352 | 341.0627 | 242.9803 | 92.3463 | 1.655 | 42.2378 | 69.9038 | 25.5213 | 2.6269 | 42.1122 | 110.6254 | 16.031 | 21.3306 | 6.3664 | 180.5292 | 325.4623 | 251.0046 | 319.7321 | 221.6497
2012-10-30 01:44 | 25.8016 | 346.7998 | 272.3539 | 341.0696 | 243.0109 | 92.3307 | 1.655 | 42.2366 | 69.8996 | 25.5213 | 2.6269 | 42.1111 | 110.6197 | 16.031 | 21.3219 | 6.3653 | 180.5193 | 325.4779 | 251.032 | 319.7477 | 221.689
2012-10-30 01:44 | 25.8098 | 346.8156 | 272.3776 | 341.0664 | 242.9015 | 92.3446 | 1.6549 | 42.2355 | 69.8963 | 25.5212 | 2.6268 | 42.1097 | 110.6129 | 16.031 | 21.3341 | 6.3662 | 180.5091 | 325.4815 | 251.0435 | 319.7323 | 221.5674
2012-10-30 01:44 | 25.8091 | 346.8111 | 272.4002 | 341.1152 | 242.9283 | 92.3533 | 1.6524 | 42.1688 | 69.6778 | 25.5205 | 2.6228 | 42.0452 | 110.2758 | 16.0307 | 21.3294 | 6.3668 | 179.9537 | 325.4817 | 251.0708 | 319.7858 | 221.5989
2012-10-30 01:44 | 25.8166 | 346.8414 | 272.3262 | 341.0884 | 243.0847 | 92.3365 | 1.6509 | 42.1318 | 69.5564 | 25.5201 | 2.6204 | 42.0062 | 110.0749 | 16.0302 | 21.3169 | 6.3657 | 179.6313 | 325.5245 | 251.0093 | 319.7715 | 221.7678
2012-10-30 01:44 | 25.8248 | 346.7773 | 272.3732 | 341.1233 | 243.1399 | 92.3394 | 1.6505 | 42.1203 | 69.5194 | 25.5198 | 2.6198 | 41.9948 | 110.0161 | 16.03 | 21.297 | 6.3659 | 179.5355 | 325.4803 | 251.0762 | 319.8263 | 221.8429
2012-10-30 01:44 | 25.8173 | 346.8117 | 272.4045 | 341.1196 | 242.2976 | 92.3307 | 1.6504 | 42.1187 | 69.5148 | 25.5195 | 2.6197 | 41.9933 | 110.009 | 16.03 | 21.3256 | 6.3653 | 179.5237 | 325.4861 | 251.0789 | 319.794 | 220.972
2012-10-30 01:44 | 25.8136 | 346.8042 | 272.2961 | 341.0892 | 241.9058 | 92.3353 | 1.6505 | 42.1217 | 69.5238 | 25.5198 | 2.6198 | 41.996 | 110.0223 | 16.0301 | 21.2858 | 6.3656 | 179.5461 | 325.5184 | 251.0103 | 319.8034 | 220.62
2012-10-30 01:44 | 25.8151 | 346.7638 | 272.1502 | 341.1288 | 242.2483 | 92.3197 | 1.6507 | 42.125 | 69.535 | 25.5197 | 2.6201 | 41.9995 | 110.0416 | 16.0299 | 21.2953 | 6.3645 | 179.5766 | 325.4685 | 250.8549 | 319.8335 | 220.953
2012-10-30 01:44 | 25.7979 | 346.7771 | 272.1446 | 341.0811 | 242.2546 | 92.3324 | 1.6508 | 42.1284 | 69.546 | 25.5198 | 2.6203 | 42.0031 | 110.0592 | 16.0301 | 21.3062 | 6.3654 | 179.6052 | 325.4709 | 250.8384 | 319.7749 | 220.9484
2012-10-30 01:44 | 25.8098 | 346.8004 | 272.1294 | 341.1387 | 242.3293 | 92.3261 | 1.6511 | 42.1365 | 69.5723 | 25.52 | 2.6208 | 42.011 | 110.1008 | 16.0301 | 21.3301 | 6.365 | 179.6731 | 325.4703 | 250.7993 | 319.8086 | 220.9992
2012-10-30 01:44 | 25.7964 | 346.8136 | 272.1625 | 341.1139 | 242.602 | 92.3226 | 1.6513 | 42.1422 | 69.5906 | 25.5202 | 2.6211 | 42.0168 | 110.1304 | 16.0302 | 21.3448 | 6.3647 | 179.721 | 325.4688 | 250.8177 | 319.7691 | 221.2572
2012-10-30 01:44 | 25.7986 | 346.8044 | 272.1686 | 341.1161 | 242.8943 | 92.333 | 1.6515 | 42.1474 | 69.6078 | 25.5202 | 2.6214 | 42.0219 | 110.1572 | 16.0302 | 21.343 | 6.3654 | 179.765 | 325.4614 | 250.8256 | 319.7731 | 221.5513
2012-10-30 01:44 | 25.7896 | 346.8144 | 272.2061 | 341.1109 | 242.9793 | 92.3388 | 1.6541 | 42.2138 | 69.8254 | 25.5209 | 2.6253 | 42.0857 | 110.4893 | 16.0306 | 21.33 | 6.3658 | 180.3147 | 325.4844 | 250.8761 | 319.7809 | 221.6493
2012-10-30 01:44 | 25.7971 | 346.8219 | 272.202 | 341.145 | 242.395 | 92.3394 | 1.6549 | 42.2363 | 69.8979 | 25.5216 | 2.6268 | 42.1106 | 110.6177 | 16.0309 | 21.3495 | 6.3659 | 180.5157 | 325.4724 | 250.8525 | 319.7955 | 221.0455
2012-10-30 01:44 | 25.7881 | 346.8548 | 272.162 | 341.0866 | 242.5036 | 92.3475 | 1.6554 | 42.2485 | 69.9386 | 25.5215 | 2.6276 | 42.1228 | 110.6811 | 16.031 | 21.3566 | 6.3664 | 180.6197 | 325.4982 | 250.8054 | 319.73 | 221.147
2012-10-30 01:44 | 25.7911 | 346.8958 | 272.1301 | 341.12 | 242.3654 | 92.3475 | 1.6558 | 42.2585 | 69.9706 | 25.5219 | 2.6282 | 42.1327 | 110.7314 | 16.0312 | 21.3435 | 6.3664 | 180.702 | 325.5523 | 250.7866 | 319.7765 | 221.0219
2012-10-30 01:44 | 25.7814 | 346.8633 | 272.0505 | 341.1065 | 242.2538 | 92.3417 | 1.656 | 42.2645 | 69.9901 | 25.522 | 2.6285 | 42.1385 | 110.7615 | 16.0313 | 21.3258 | 6.366 | 180.7516 | 325.5375 | 250.7247 | 319.7807 | 220.928
2012-10-30 01:44 | 25.7784 | 346.8907 | 272.0553 | 341.0997 | 242.682 | 92.3556 | 1.6561 | 42.267 | 69.9986 | 25.5219 | 2.6287 | 42.1413 | 110.776 | 16.0314 | 21.3228 | 6.367 | 180.7746 | 325.5679 | 250.7325 | 319.7769 | 221.3592
2012-10-30 01:44 | 25.7851 | 346.8518 | 272.0736 | 341.0988 | 243.0924 | 92.3509 | 1.6561 | 42.2674 | 69.9996 | 25.5221 | 2.6287 | 42.1415 | 110.7756 | 16.0315 | 21.3536 | 6.3667 | 180.7753 | 325.4982 | 250.72 | 319.7452 | 221.7388
2012-10-30 01:44 | 25.7724 | 346.8809 | 272.0842 | 341.0823 | 242.9817 | 92.355 | 1.6561 | 42.2666 | 69.9971 | 25.5221 | 2.6286 | 42.1408 | 110.7725 | 16.0315 | 21.3128 | 6.367 | 180.7695 | 325.5681 | 250.7714 | 319.7695 | 221.6689
2012-10-30 01:45 | 25.7717 | 346.8688 | 272.1611 | 341.112 | 243.0085 | 92.3625 | 1.6561 | 42.2676 | 70 | 25.5221 | 2.6287 | 42.1417 | 110.7767 | 16.0315 | 21.3161 | 6.3675 | 180.7767 | 325.5527 | 250.845 | 319.7959 | 221.6924
2012-10-30 01:45 | 25.7769 | 346.9044 | 272.1934 | 341.0868 | 243.0568 | 92.3729 | 1.6561 | 42.2675 | 69.9996 | 25.5222 | 2.6287 | 42.1419 | 110.7771 | 16.0316 | 21.3614 | 6.3682 | 180.7767 | 325.543 | 250.832 | 319.7254 | 221.6954
2012-10-30 01:45 | 25.7784 | 346.9516 | 272.2686 | 341.1111 | 243.1758 | 92.3631 | 1.6561 | 42.2668 | 69.9969 | 25.5223 | 2.6286 | 42.1409 | 110.7719 | 16.0316 | 21.3549 | 6.3675 | 180.7687 | 325.5967 | 250.9137 | 319.7562 | 221.8209
2012-10-30 01:45 | 25.7687 | 346.9076 | 272.2667 | 341.1433 | 242.9819 | 92.3613 | 1.656 | 42.2661 | 69.9939 | 25.5225 | 2.6286 | 42.1402 | 110.7679 | 16.0317 | 21.3572 | 6.3674 | 180.7617 | 325.5504 | 250.9095 | 319.7861 | 221.6247
2012-10-30 01:45 | 25.7702 | 346.8749 | 272.2721 | 341.1334 | 242.7091 | 92.3683 | 1.656 | 42.2647 | 69.9894 | 25.5225 | 2.6285 | 42.1389 | 110.7611 | 16.0317 | 21.3226 | 6.3679 | 180.7505 | 325.5523 | 250.9495 | 319.8108 | 221.3865
2012-10-30 01:45 | 25.7717 | 346.903 | 272.2464 | 341.1196 | 242.7694 | 92.3758 | 1.6558 | 42.2608 | 69.9768 | 25.5223 | 2.6282 | 42.1349 | 110.74 | 16.0317 | 21.288 | 6.3684 | 180.7168 | 325.615 | 250.9584 | 319.8316 | 221.4814
2012-10-30 01:45 | 25.7657 | 346.8666 | 272.2521 | 341.1555 | 242.8731 | 92.3764 | 1.6558 | 42.2595 | 69.9723 | 25.5224 | 2.6281 | 42.1334 | 110.7324 | 16.0317 | 21.3021 | 6.3684 | 180.7047 | 325.5645 | 250.95 | 319.8534 | 221.571
2012-10-30 01:45 | 25.7702 | 346.9015 | 272.2992 | 341.0991 | 243.0971 | 92.3677 | 1.6557 | 42.2588 | 69.9695 | 25.5227 | 2.6281 | 42.133 | 110.7292 | 16.0318 | 21.3266 | 6.3678 | 180.6987 | 325.5749 | 250.9726 | 319.7725 | 221.7705
2012-10-30 01:45 | 25.7754 | 346.9143 | 272.2967 | 341.15 | 243.1141 | 92.3689 | 1.6557 | 42.2577 | 69.966 | 25.5226 | 2.628 | 42.1318 | 110.7229 | 16.0318 | 21.3639 | 6.3679 | 180.6889 | 325.5504 | 250.9328 | 319.7861 | 221.7502
2012-10-30 01:45 | 25.7709 | 346.8832 | 272.3232 | 341.1302 | 242.8979 | 92.3764 | 1.6557 | 42.257 | 69.963 | 25.5228 | 2.628 | 42.1312 | 110.72 | 16.0318 | 21.3394 | 6.3684 | 180.683 | 325.5438 | 250.9838 | 319.7908 | 221.5585
2012-10-30 01:45 | 25.7761 | 346.8884 | 272.3866 | 341.1278 | 243.0481 | 92.3694 | 1.6556 | 42.256 | 69.9595 | 25.5229 | 2.6279 | 42.1302 | 110.7145 | 16.0318 | 21.3365 | 6.3679 | 180.674 | 325.5519 | 251.0501 | 319.7913 | 221.7116
2012-10-30 01:45 | 25.7776 | 346.8975 | 272.3958 | 341.1674 | 243.0379 | 92.3735 | 1.6555 | 42.2532 | 69.9508 | 25.5227 | 2.6277 | 42.1273 | 110.6989 | 16.0319 | 21.3621 | 6.3682 | 180.6496 | 325.5354 | 251.0337 | 319.8053 | 221.6758
2012-10-30 01:45 | 25.7747 | 346.8908 | 272.3774 | 341.1455 | 242.9644 | 92.3787 | 1.6554 | 42.2513 | 69.9438 | 25.5229 | 2.6276 | 42.1255 | 110.6906 | 16.0317 | 21.3552 | 6.3686 | 180.6344 | 325.5356 | 251.0222 | 319.7903 | 221.6092
2012-10-30 01:45 | 25.7717 | 346.865 | 272.3783 | 341.1615 | 243.0477 | 92.385 | 1.6554 | 42.2501 | 69.9405 | 25.5228 | 2.6275 | 42.1242 | 110.6829 | 16.0319 | 21.3723 | 6.369 | 180.6235 | 325.4927 | 251.006 | 319.7892 | 221.6754
2012-10-30 01:45 | 25.7657 | 346.9046 | 272.4266 | 341.1517 | 243.0534 | 92.381 | 1.6553 | 42.2477 | 69.9323 | 25.5228 | 2.6274 | 42.122 | 110.6706 | 16.0319 | 21.3783 | 6.3687 | 180.6029 | 325.5263 | 251.0483 | 319.7734 | 221.6751
2012-10-30 01:45 | 25.7657 | 346.8856 | 272.4344 | 341.1859 | 243.2023 | 92.4006 | 1.6552 | 42.2458 | 69.9261 | 25.5227 | 2.6273 | 42.1198 | 110.6594 | 16.0319 | 21.3542 | 6.3701 | 180.5855 | 325.5314 | 251.0802 | 319.8317 | 221.8481
2012-10-30 01:45 | 25.7806 | 346.9157 | 272.5074 | 341.1856 | 243.1741 | 92.3943 | 1.6551 | 42.2435 | 69.9181 | 25.5229 | 2.6271 | 42.1176 | 110.649 | 16.0317 | 21.3611 | 6.3697 | 180.5671 | 325.5546 | 251.1463 | 319.8245 | 221.813
2012-10-30 01:45 | 25.7776 | 346.9089 | 272.5199 | 341.2207 | 243.0771 | 92.4024 | 1.655 | 42.2406 | 69.909 | 25.5227 | 2.627 | 42.1148 | 110.6351 | 16.0316 | 21.3541 | 6.3702 | 180.5441 | 325.5548 | 251.1658 | 319.8666 | 221.723
2012-10-30 01:45 | 25.7747 | 346.9098 | 272.5248 | 341.2025 | 242.9527 | 92.3943 | 1.655 | 42.2398 | 69.9065 | 25.5227 | 2.6269 | 42.1141 | 110.6307 | 16.0317 | 21.3392 | 6.3697 | 180.5373 | 325.5706 | 251.1856 | 319.8633 | 221.6135
2012-10-30 01:45 | 25.7874 | 346.9187 | 272.5181 | 341.219 | 243.0712 | 92.3926 | 1.6549 | 42.2381 | 69.9002 | 25.5229 | 2.6268 | 42.1123 | 110.6211 | 16.0317 | 21.3759 | 6.3695 | 180.5213 | 325.5428 | 251.1422 | 319.8431 | 221.6953
2012-10-30 01:45 | 25.7851 | 346.9583 | 272.543 | 341.232 | 243.1512 | 92.4082 | 1.6548 | 42.2354 | 69.8919 | 25.5227 | 2.6267 | 42.1097 | 110.6082 | 16.0316 | 21.3616 | 6.3706 | 180.5001 | 325.5967 | 251.1814 | 319.8704 | 221.7896
2012-10-30 01:46 | 25.7821 | 346.9401 | 272.5826 | 341.2594 | 243.2579 | 92.4041 | 1.6547 | 42.2335 | 69.8857 | 25.5227 | 2.6266 | 42.1079 | 110.5994 | 16.0315 | 21.3546 | 6.3703 | 180.4851 | 325.5855 | 251.228 | 319.9048 | 221.9033
2012-10-30 01:46 | 25.7821 | 346.9363 | 272.5787 | 341.2442 | 243.254 | 92.396 | 1.6518 | 42.1561 | 69.633 | 25.5215 | 2.6218 | 42.0304 | 110.197 | 16.0309 | 21.3506 | 6.3698 | 179.83 | 325.5857 | 251.2281 | 319.8936 | 221.9034
2012-10-30 01:46 | 25.7896 | 346.9171 | 272.6095 | 341.2555 | 242.85 | 92.3989 | 1.6503 | 42.118 | 69.5087 | 25.5209 | 2.6196 | 41.9926 | 110.0021 | 16.0304 | 21.3982 | 6.37 | 179.5108 | 325.5189 | 251.2113 | 319.8573 | 221.4518
2012-10-30 01:46 | 25.7866 | 346.9484 | 272.5445 | 341.2525 | 242.8901 | 92.3949 | 1.6504 | 42.1202 | 69.516 | 25.5209 | 2.6197 | 41.9948 | 110.0134 | 16.0304 | 21.3671 | 6.3697 | 179.5294 | 325.5813 | 251.1774 | 319.8854 | 221.523
2012-10-30 01:46 | 25.7821 | 346.9096 | 272.5322 | 341.2328 | 243.2775 | 92.3758 | 1.6506 | 42.1239 | 69.5276 | 25.5211 | 2.6199 | 41.9984 | 110.0321 | 16.0305 | 21.3626 | 6.3684 | 179.5597 | 325.547 | 251.1696 | 319.8702 | 221.9149
2012-10-30 01:46 | 25.7761 | 346.8732 | 272.5883 | 341.2458 | 243.3107 | 92.3874 | 1.6508 | 42.1296 | 69.5468 | 25.521 | 2.6203 | 42.0042 | 110.0628 | 16.0304 | 21.3967 | 6.3692 | 179.6096 | 325.4765 | 251.1916 | 319.8491 | 221.914
2012-10-30 01:46 | 25.7747 | 346.9174 | 272.486 | 341.2102 | 242.2706 | 92.3746 | 1.6509 | 42.1331 | 69.5581 | 25.5211 | 2.6205 | 42.0076 | 110.0803 | 16.0305 | 21.3713 | 6.3683 | 179.6384 | 325.5461 | 251.1147 | 319.8389 | 220.8993
2012-10-30 01:46 | 25.7747 | 346.8908 | 272.3154 | 341.2292 | 242.741 | 92.3793 | 1.6511 | 42.1392 | 69.5779 | 25.5212 | 2.6209 | 42.0141 | 110.1136 | 16.0306 | 21.3793 | 6.3686 | 179.6915 | 325.5115 | 250.9361 | 319.8499 | 221.3617
2012-10-30 01:46 | 25.7702 | 346.8597 | 272.4117 | 341.2133 | 243.873 | 92.3822 | 1.6527 | 42.1798 | 69.7105 | 25.5217 | 2.6234 | 42.0555 | 110.3276 | 16.031 | 21.3989 | 6.3688 | 180.0381 | 325.4608 | 251.0128 | 319.8144 | 222.4741
2012-10-30 01:46 | 25.7791 | 346.8496 | 272.5641 | 341.2222 | 243.7487 | 92.3879 | 1.6545 | 42.2284 | 69.8688 | 25.5227 | 2.6263 | 42.1026 | 110.573 | 16.0313 | 21.4559 | 6.3692 | 180.4418 | 325.3937 | 251.1082 | 319.7663 | 222.2928
2012-10-30 01:46 | 25.7627 | 346.8636 | 272.5632 | 341.1563 | 243.1954 | 92.385 | 1.655 | 42.24 | 69.9065 | 25.5229 | 2.627 | 42.1142 | 110.6324 | 16.0316 | 21.4074 | 6.369 | 180.5389 | 325.4562 | 251.1558 | 319.7489 | 221.788
2012-10-30 01:46 | 25.7732 | 346.9083 | 272.5466 | 341.201 | 242.767 | 92.3902 | 1.6554 | 42.2506 | 69.942 | 25.5228 | 2.6276 | 42.125 | 110.688 | 16.0317 | 21.45 | 6.3694 | 180.63 | 325.4583 | 251.0966 | 319.751 | 221.317
2012-10-30 01:46 | 25.7679 | 346.9183 | 272.4792 | 341.2072 | 242.9576 | 92.3931 | 1.6556 | 42.2566 | 69.9603 | 25.5233 | 2.628 | 42.1308 | 110.7187 | 16.0317 | 21.4367 | 6.3696 | 180.679 | 325.4816 | 251.0425 | 319.7705 | 221.5209
2012-10-30 01:46 | 25.7589 | 346.8674 | 272.4974 | 341.1944 | 242.9486 | 92.4001 | 1.6558 | 42.2605 | 69.9737 | 25.5231 | 2.6282 | 42.1348 | 110.7388 | 16.0318 | 21.4437 | 6.3701 | 180.7126 | 325.4237 | 251.0537 | 319.7507 | 221.5049
2012-10-30 01:46 | 25.753 | 346.8805 | 272.4798 | 341.1047 | 242.8487 | 92.403 | 1.6558 | 42.2616 | 69.9771 | 25.5232 | 2.6283 | 42.1357 | 110.7436 | 16.0318 | 21.4498 | 6.3703 | 180.7208 | 325.4307 | 251.03 | 319.6549 | 221.3989
2012-10-30 01:46 | 25.7567 | 346.8576 | 272.4254 | 341.1313 | 242.7505 | 92.3908 | 1.6558 | 42.2614 | 69.9762 | 25.5234 | 2.6283 | 42.1356 | 110.743 | 16.0318 | 21.4335 | 6.3694 | 180.7192 | 325.4241 | 250.9919 | 319.6978 | 221.317
2012-10-30 01:46 | 25.7552 | 346.8789 | 272.4161 | 341.1488 | 242.9214 | 92.418 | 1.6558 | 42.2621 | 69.9778 | 25.5236 | 2.6283 | 42.1362 | 110.7458 | 16.0319 | 21.3999 | 6.3713 | 180.7237 | 325.479 | 251.0162 | 319.7489 | 221.5215
2012-10-30 01:46 | 25.7559 | 346.8568 | 272.4129 | 341.1914 | 242.6869 | 92.3978 | 1.6558 | 42.2618 | 69.9779 | 25.5232 | 2.6283 | 42.1361 | 110.7451 | 16.0318 | 21.4648 | 6.3699 | 180.723 | 325.392 | 250.9481 | 319.7266 | 221.2221
2012-10-30 01:46 | 25.7537 | 346.8926 | 272.403 | 341.1777 | 242.9748 | 92.3943 | 1.6558 | 42.2619 | 69.9779 | 25.5233 | 2.6283 | 42.1362 | 110.7459 | 16.0318 | 21.4225 | 6.3697 | 180.7239 | 325.4701 | 250.9805 | 319.7552 | 221.5523
2012-10-30 01:47 | 25.744 | 346.902 | 272.4243 | 341.1566 | 242.9808 | 92.4082 | 1.6558 | 42.2605 | 69.9732 | 25.5233 | 2.6282 | 42.1346 | 110.7376 | 16.0318 | 21.4449 | 6.3706 | 180.7108 | 325.4571 | 250.9794 | 319.7117 | 221.5359
2012-10-30 01:47 | 25.7417 | 346.8806 | 272.4259 | 341.1543 | 242.8139 | 92.4082 | 1.6557 | 42.2592 | 69.9692 | 25.5232 | 2.6281 | 42.1334 | 110.7321 | 16.0317 | 21.3784 | 6.3706 | 180.7013 | 325.5022 | 251.0475 | 319.7759 | 221.4355
2012-10-30 01:47 | 25.7365 | 346.8602 | 272.3935 | 341.1491 | 243.0673 | 92.4053 | 1.6557 | 42.2588 | 69.9676 | 25.5233 | 2.6281 | 42.1327 | 110.7286 | 16.0317 | 21.4013 | 6.3704 | 180.6962 | 325.4589 | 250.9922 | 319.7478 | 221.666
2012-10-30 01:47 | 25.732 | 346.8633 | 272.3774 | 341.1636 | 242.9452 | 92.4139 | 1.6557 | 42.2577 | 69.9644 | 25.5232 | 2.6281 | 42.132 | 110.7251 | 16.0317 | 21.3968 | 6.371 | 180.6895 | 325.4665 | 250.9806 | 319.7668 | 221.5484
2012-10-30 01:47 | 25.7305 | 346.839 | 272.4263 | 341.1507 | 242.8614 | 92.4053 | 1.6556 | 42.257 | 69.962 | 25.5232 | 2.628 | 42.1311 | 110.7195 | 16.0317 | 21.3351 | 6.3704 | 180.6815 | 325.5039 | 251.0912 | 319.8156 | 221.5263
2012-10-30 01:47 | 25.7275 | 346.9045 | 272.4 | 341.1553 | 243.0074 | 92.418 | 1.6556 | 42.2551 | 69.9556 | 25.5232 | 2.6279 | 42.1292 | 110.7101 | 16.0317 | 21.3481 | 6.3713 | 180.6657 | 325.5564 | 251.0519 | 319.8072 | 221.6593
2012-10-30 01:47 | 25.7215 | 346.8642 | 272.4367 | 341.1798 | 243.26 | 92.4295 | 1.6555 | 42.2536 | 69.951 | 25.5232 | 2.6278 | 42.1279 | 110.7027 | 16.0318 | 21.3461 | 6.3721 | 180.6536 | 325.5181 | 251.0906 | 319.8337 | 221.9139
2012-10-30 01:47 | 25.729 | 346.8984 | 272.4403 | 341.2139 | 243.3851 | 92.433 | 1.6554 | 42.2524 | 69.946 | 25.5235 | 2.6277 | 42.1267 | 110.6967 | 16.0317 | 21.3536 | 6.3723 | 180.6427 | 325.5448 | 251.0867 | 319.8603 | 222.0315
2012-10-30 01:47 | 25.7283 | 346.8406 | 272.49 | 341.1447 | 243.2159 | 92.4151 | 1.6554 | 42.2513 | 69.9425 | 25.5234 | 2.6276 | 42.1253 | 110.6898 | 16.0317 | 21.3329 | 6.3711 | 180.6323 | 325.5077 | 251.1571 | 319.8118 | 221.883
2012-10-30 01:47 | 25.7215 | 346.9137 | 272.4987 | 341.1608 | 243.1385 | 92.422 | 1.6553 | 42.25 | 69.9381 | 25.5234 | 2.6276 | 42.1242 | 110.6838 | 16.0317 | 21.31 | 6.3716 | 180.6219 | 325.6037 | 251.1887 | 319.8508 | 221.8285
2012-10-30 01:47 | 25.723 | 346.8848 | 272.5041 | 341.1851 | 243.1596 | 92.422 | 1.6552 | 42.2465 | 69.9266 | 25.5234 | 2.6273 | 42.1208 | 110.665 | 16.0318 | 21.3356 | 6.3716 | 180.5916 | 325.5492 | 251.1685 | 319.8495 | 221.824
2012-10-30 01:47 | 25.7313 | 346.9083 | 272.5551 | 341.182 | 243.1052 | 92.4347 | 1.6551 | 42.2452 | 69.9221 | 25.5235 | 2.6272 | 42.1195 | 110.6575 | 16.0319 | 21.3319 | 6.3724 | 180.5796 | 325.5764 | 251.2232 | 319.8501 | 221.7733
2012-10-30 01:47 | 25.726 | 346.8992 | 272.5459 | 341.2071 | 243.3312 | 92.4371 | 1.6551 | 42.2435 | 69.9168 | 25.5234 | 2.6271 | 42.1177 | 110.6484 | 16.0318 | 21.3065 | 6.3726 | 180.5652 | 325.5927 | 251.2394 | 319.9006 | 222.0247
2012-10-30 01:47 | 25.735 | 346.9234 | 272.5898 | 341.1857 | 243.2539 | 92.4272 | 1.655 | 42.2409 | 69.9076 | 25.5237 | 2.627 | 42.1151 | 110.6346 | 16.0319 | 21.3195 | 6.3719 | 180.5422 | 325.6039 | 251.2703 | 319.8662 | 221.9344
2012-10-30 01:47 | 25.7372 | 346.9294 | 272.5649 | 341.2145 | 242.9348 | 92.4371 | 1.6549 | 42.238 | 69.8979 | 25.5237 | 2.6268 | 42.1122 | 110.6187 | 16.032 | 21.2936 | 6.3726 | 180.5165 | 325.6358 | 251.2713 | 319.9209 | 221.6412
2012-10-30 01:47 | 25.7425 | 346.9499 | 272.5624 | 341.2198 | 243.3085 | 92.4376 | 1.6548 | 42.2361 | 69.8912 | 25.5238 | 2.6267 | 42.1102 | 110.6089 | 16.0319 | 21.3029 | 6.3726 | 180.5001 | 325.647 | 251.2595 | 319.9169 | 222.0056
2012-10-30 01:47 | 25.7485 | 346.9141 | 272.5917 | 341.2258 | 243.2714 | 92.4272 | 1.6547 | 42.235 | 69.8876 | 25.5237 | 2.6266 | 42.1091 | 110.6028 | 16.032 | 21.2929 | 6.3719 | 180.4904 | 325.6212 | 251.2988 | 319.9329 | 221.9785
2012-10-30 01:47 | 25.7544 | 346.9314 | 272.617 | 341.2355 | 243.3321 | 92.4423 | 1.6545 | 42.2301 | 69.8716 | 25.5237 | 2.6263 | 42.1043 | 110.5773 | 16.032 | 21.3269 | 6.373 | 180.4489 | 325.6045 | 251.2901 | 319.9086 | 222.0052
2012-10-30 01:47 | 25.7574 | 346.9306 | 272.6083 | 341.2157 | 243.147 | 92.4451 | 1.6545 | 42.2288 | 69.8672 | 25.5237 | 2.6262 | 42.1033 | 110.5725 | 16.0319 | 21.354 | 6.3732 | 180.4397 | 325.5766 | 251.2543 | 319.8617 | 221.793
2012-10-30 01:47 | 25.7604 | 346.9412 | 272.5958 | 341.2301 | 243.1618 | 92.4538 | 1.6545 | 42.2279 | 69.8652 | 25.5234 | 2.6261 | 42.1023 | 110.5668 | 16.032 | 21.3289 | 6.3738 | 180.432 | 325.6123 | 251.2669 | 319.9012 | 221.8329
2012-10-30 01:47 | 25.7619 | 346.9541 | 272.6051 | 341.2468 | 243.3396 | 92.4457 | 1.6544 | 42.2275 | 69.8627 | 25.5238 | 2.6261 | 42.1019 | 110.5657 | 16.0318 | 21.3464 | 6.3732 | 180.4284 | 325.6077 | 251.2587 | 319.9004 | 221.9932
2012-10-30 01:47 | 25.7694 | 346.9502 | 272.5777 | 341.2467 | 243.3354 | 92.4469 | 1.652 | 42.1629 | 69.6517 | 25.5228 | 2.6224 | 42.0403 | 110.2449 | 16.0315 | 21.3619 | 6.3733 | 179.8966 | 325.5883 | 251.2158 | 319.8848 | 221.9735
2012-10-30 01:48 | 25.7754 | 346.9296 | 272.6263 | 341.2717 | 243.3335 | 92.4417 | 1.6506 | 42.126 | 69.5312 | 25.5224 | 2.6199 | 42.0006 | 110.0395 | 16.0311 | 21.4121 | 6.3729 | 179.5707 | 325.5175 | 251.2142 | 319.8596 | 221.9214
2012-10-30 01:48 | 25.7724 | 346.9608 | 272.6233 | 341.2611 | 243.1111 | 92.4319 | 1.6502 | 42.1172 | 69.5026 | 25.5221 | 2.6194 | 41.9915 | 109.9916 | 16.0311 | 21.377 | 6.3723 | 179.4942 | 325.5838 | 251.2463 | 319.8841 | 221.7341
2012-10-30 01:48 | 25.7821 | 346.9553 | 272.6136 | 341.2899 | 242.1604 | 92.4255 | 1.6502 | 42.1156 | 69.4973 | 25.5222 | 2.6193 | 41.9902 | 109.9857 | 16.0309 | 21.3827 | 6.3718 | 179.4831 | 325.5726 | 251.2309 | 319.9072 | 220.7777
2012-10-30 01:48 | 25.7724 | 346.919 | 272.4876 | 341.284 | 242.3977 | 92.4186 | 1.6502 | 42.1157 | 69.4979 | 25.5221 | 2.6193 | 41.9902 | 109.9856 | 16.031 | 21.385 | 6.3713 | 179.4835 | 325.534 | 251.1026 | 319.899 | 221.0127
2012-10-30 01:48 | 25.7664 | 346.9091 | 272.3265 | 341.2627 | 241.6155 | 92.4174 | 1.6502 | 42.1152 | 69.4964 | 25.522 | 2.6193 | 41.9898 | 109.9847 | 16.0308 | 21.3991 | 6.3713 | 179.4812 | 325.51 | 250.9274 | 319.8636 | 220.2164
2012-10-30 01:48 | 25.7679 | 346.9259 | 272.1844 | 341.2795 | 242.2834 | 92.4203 | 1.6504 | 42.1212 | 69.5163 | 25.522 | 2.6197 | 41.9956 | 110.0149 | 16.0308 | 21.3564 | 6.3715 | 179.5312 | 325.5695 | 250.828 | 319.9231 | 220.927
2012-10-30 01:48 | 25.7694 | 346.8703 | 272.1743 | 341.2391 | 242.3437 | 92.4272 | 1.6506 | 42.128 | 69.5383 | 25.5222 | 2.6201 | 42.0025 | 110.0505 | 16.0309 | 21.4101 | 6.3719 | 179.5888 | 325.4602 | 250.7642 | 319.829 | 220.9336
2012-10-30 01:48 | 25.7702 | 346.8863 | 272.1596 | 341.2513 | 242.4935 | 92.4295 | 1.6509 | 42.1341 | 69.5576 | 25.5225 | 2.6205 | 42.0086 | 110.0816 | 16.031 | 21.4149 | 6.3721 | 179.6392 | 325.4714 | 250.7447 | 319.8364 | 221.0786
2012-10-30 01:48 | 25.7642 | 346.8879 | 272.1303 | 341.2491 | 241.7466 | 92.4186 | 1.6511 | 42.1408 | 69.5796 | 25.5226 | 2.6209 | 42.0155 | 110.1179 | 16.031 | 21.3929 | 6.3713 | 179.6975 | 325.495 | 250.7374 | 319.8562 | 220.3537
2012-10-30 01:48 | 25.7619 | 346.8514 | 272.0543 | 341.2392 | 242.0854 | 92.4232 | 1.6537 | 42.2087 | 69.8016 | 25.5234 | 2.6248 | 42.0801 | 110.4533 | 16.0315 | 21.3906 | 6.3717 | 180.2548 | 325.4608 | 250.6637 | 319.8486 | 220.6948
2012-10-30 01:48 | 25.7604 | 346.8613 | 271.952 | 341.2225 | 241.9074 | 92.4232 | 1.6546 | 42.2316 | 69.8756 | 25.524 | 2.6264 | 42.1057 | 110.5853 | 16.0319 | 21.377 | 6.3717 | 180.461 | 325.4843 | 250.575 | 319.8455 | 220.5304
2012-10-30 01:48 | 25.7642 | 346.9184 | 271.9015 | 341.2073 | 241.8956 | 92.4313 | 1.655 | 42.2423 | 69.9102 | 25.5243 | 2.627 | 42.1163 | 110.6408 | 16.0319 | 21.3888 | 6.3722 | 180.551 | 325.5296 | 250.5127 | 319.8185 | 220.5068
2012-10-30 01:48 | 25.7627 | 346.8902 | 271.9194 | 341.2248 | 242.3566 | 92.4347 | 1.6554 | 42.2529 | 69.9459 | 25.5241 | 2.6277 | 42.1271 | 110.6967 | 16.032 | 21.4034 | 6.3724 | 180.6426 | 325.4868 | 250.516 | 319.8214 | 220.9532
2012-10-30 01:48 | 25.7679 | 346.9183 | 271.9866 | 341.211 | 242.4951 | 92.4342 | 1.6556 | 42.2596 | 69.967 | 25.5245 | 2.6281 | 42.1337 | 110.7301 | 16.0322 | 21.4086 | 6.3724 | 180.6971 | 325.5097 | 250.578 | 319.8024 | 221.0865
2012-10-30 01:48 | 25.7649 | 346.9305 | 272.0534 | 341.2004 | 242.6176 | 92.4307 | 1.6558 | 42.2627 | 69.9773 | 25.5245 | 2.6283 | 42.1368 | 110.7462 | 16.0323 | 21.3695 | 6.3722 | 180.7235 | 325.561 | 250.6839 | 319.8309 | 221.2481
2012-10-30 01:48 | 25.7664 | 346.9053 | 272.0899 | 341.2171 | 242.6347 | 92.4469 | 1.6559 | 42.265 | 69.9847 | 25.5246 | 2.6284 | 42.1391 | 110.7585 | 16.0322 | 21.4111 | 6.3733 | 180.7431 | 325.4942 | 250.6788 | 319.806 | 221.2236
2012-10-30 01:48 | 25.7672 | 346.8985 | 272.1372 | 341.1836 | 242.4827 | 92.4475 | 1.6559 | 42.2648 | 69.9844 | 25.5245 | 2.6284 | 42.139 | 110.7578 | 16.0322 | 21.3678 | 6.3733 | 180.7422 | 325.5307 | 250.7694 | 319.8158 | 221.1149
2012-10-30 01:48 | 25.7687 | 346.9191 | 272.1309 | 341.2118 | 242.4724 | 92.4503 | 1.6559 | 42.2656 | 69.9861 | 25.5248 | 2.6284 | 42.1397 | 110.7608 | 16.0323 | 21.3974 | 6.3735 | 180.7469 | 325.5217 | 250.7335 | 319.8144 | 221.075
2012-10-30 01:48 | 25.7709 | 346.9365 | 272.1215 | 341.2102 | 242.3844 | 92.4607 | 1.6558 | 42.2647 | 69.9837 | 25.5245 | 2.6284 | 42.1388 | 110.7564 | 16.0323 | 21.3875 | 6.3742 | 180.74 | 325.549 | 250.734 | 319.8227 | 220.9969
2012-10-30 01:48 | 25.7642 | 346.9298 | 272.0915 | 341.2225 | 242.5541 | 92.4544 | 1.6558 | 42.263 | 69.9782 | 25.5246 | 2.6283 | 42.137 | 110.7468 | 16.0323 | 21.3808 | 6.3738 | 180.725 | 325.549 | 250.7107 | 319.8417 | 221.1733
2012-10-30 01:48 | 25.7717 | 346.8688 | 272.1572 | 341.249 | 242.542 | 92.4561 | 1.6557 | 42.2621 | 69.9749 | 25.5247 | 2.6282 | 42.1361 | 110.743 | 16.0322 | 21.3883 | 6.3739 | 180.7179 | 325.4805 | 250.7689 | 319.8607 | 221.1537
2012-10-30 01:48 | 25.7739 | 346.8444 | 272.1866 | 341.2094 | 242.6775 | 92.4619 | 1.6557 | 42.2618 | 69.9733 | 25.5248 | 2.6282 | 42.136 | 110.7416 | 16.0323 | 21.3825 | 6.3743 | 180.7149 | 325.4619 | 250.8041 | 319.8269 | 221.295
2012-10-30 01:48 | 25.7694 | 346.8627 | 272.1743 | 341.2201 | 242.8964 | 92.47 | 1.6556 | 42.2591 | 69.9644 | 25.5249 | 2.628 | 42.1331 | 110.7263 | 16.0323 | 21.366 | 6.3749 | 180.6907 | 325.4967 | 250.8083 | 319.8541 | 221.5304
2012-10-30 01:49 | 25.7702 | 346.9129 | 272.2488 | 341.2361 | 242.9051 | 92.4648 | 1.6556 | 42.258 | 69.9609 | 25.5248 | 2.6279 | 42.1322 | 110.7207 | 16.0324 | 21.3989 | 6.3745 | 180.6816 | 325.514 | 250.8499 | 319.8372 | 221.5062
2012-10-30 01:49 | 25.7739 | 346.8444 | 272.2409 | 341.2094 | 243.0028 | 92.4584 | 1.6555 | 42.2575 | 69.9588 | 25.525 | 2.6279 | 42.1314 | 110.7163 | 16.0325 | 21.4186 | 6.3741 | 180.6751 | 325.4258 | 250.8223 | 319.7908 | 221.5842
2012-10-30 01:49 | 25.7754 | 346.9296 | 272.3005 | 341.2032 | 243.1297 | 92.4729 | 1.6554 | 42.2552 | 69.9507 | 25.5251 | 2.6277 | 42.1292 | 110.704 | 16.0326 | 21.392 | 6.3751 | 180.6547 | 325.5376 | 250.9085 | 319.8112 | 221.7377
2012-10-30 01:49 | 25.7717 | 346.9183 | 272.3395 | 341.2224 | 242.9771 | 92.4706 | 1.6554 | 42.2541 | 69.9473 | 25.525 | 2.6277 | 42.1283 | 110.6996 | 16.0325 | 21.3843 | 6.3749 | 180.6469 | 325.534 | 250.9552 | 319.8381 | 221.5928
2012-10-30 01:49 | 25.7814 | 346.9622 | 272.4074 | 341.293 | 242.9359 | 92.4683 | 1.6553 | 42.2528 | 69.9421 | 25.5254 | 2.6276 | 42.1268 | 110.6916 | 16.0325 | 21.3739 | 6.3748 | 180.6337 | 325.5883 | 251.0335 | 319.9191 | 221.562
2012-10-30 01:49 | 25.7769 | 346.9577 | 272.3874 | 341.2504 | 242.8255 | 92.4746 | 1.6553 | 42.2516 | 69.9383 | 25.5253 | 2.6275 | 42.1257 | 110.6853 | 16.0326 | 21.3855 | 6.3752 | 180.6236 | 325.5722 | 251.0019 | 319.8649 | 221.44
2012-10-30 01:49 | 25.7776 | 346.9394 | 272.4269 | 341.2435 | 243.0536 | 92.4613 | 1.6552 | 42.25 | 69.9325 | 25.5255 | 2.6274 | 42.1244 | 110.6786 | 16.0326 | 21.3541 | 6.3743 | 180.6111 | 325.5853 | 251.0728 | 319.8894 | 221.6995
2012-10-30 01:49 | 25.7754 | 346.9372 | 272.5061 | 341.2413 | 243.1964 | 92.4775 | 1.6551 | 42.2466 | 69.9219 | 25.5253 | 2.6272 | 42.1208 | 110.6599 | 16.0326 | 21.372 | 6.3754 | 180.5818 | 325.5652 | 251.1341 | 319.8693 | 221.8244
2012-10-30 01:49 | 25.7761 | 346.9569 | 272.5184 | 341.2915 | 243.3852 | 92.4688 | 1.655 | 42.2452 | 69.9175 | 25.5252 | 2.6271 | 42.1193 | 110.6524 | 16.0325 | 21.4088 | 6.3748 | 180.5699 | 325.5481 | 251.1096 | 319.8827 | 221.9764
2012-10-30 01:49 | 25.7814 | 346.9888 | 272.6052 | 341.2968 | 243.2768 | 92.4735 | 1.6549 | 42.2421 | 69.9072 | 25.5252 | 2.6269 | 42.1164 | 110.6365 | 16.0326 | 21.394 | 6.3751 | 180.5437 | 325.5948 | 251.2112 | 319.9028 | 221.8828
2012-10-30 01:49 | 25.7814 | 346.9812 | 272.6091 | 341.2853 | 243.2141 | 92.4856 | 1.6548 | 42.2396 | 69.8983 | 25.5255 | 2.6268 | 42.1138 | 110.6231 | 16.0326 | 21.394 | 6.376 | 180.5214 | 325.5872 | 251.2151 | 319.8913 | 221.8201
2012-10-30 01:49 | 25.7776 | 346.985 | 272.644 | 341.2892 | 243.3279 | 92.4781 | 1.6548 | 42.2394 | 69.8983 | 25.5252 | 2.6268 | 42.1136 | 110.6224 | 16.0325 | 21.3942 | 6.3754 | 180.5207 | 325.5908 | 251.2498 | 319.895 | 221.9337
2012-10-30 01:49 | 25.7821 | 347.0047 | 272.6563 | 341.2784 | 243.2814 | 92.4752 | 1.6547 | 42.2368 | 69.8895 | 25.5252 | 2.6266 | 42.111 | 110.6092 | 16.0325 | 21.3466 | 6.3752 | 180.4988 | 325.6581 | 251.3097 | 319.9318 | 221.9348
2012-10-30 01:49 | 25.7821 | 346.9667 | 272.7067 | 341.3279 | 243.3559 | 92.4746 | 1.6546 | 42.235 | 69.8836 | 25.5252 | 2.6265 | 42.109 | 110.599 | 16.0324 | 21.3546 | 6.3752 | 180.4826 | 325.6121 | 251.3521 | 319.9733 | 222.0013
2012-10-30 01:49 | 25.7874 | 346.9872 | 272.6732 | 341.318 | 243.2711 | 92.4752 | 1.6546 | 42.2336 | 69.8789 | 25.5253 | 2.6264 | 42.1078 | 110.5927 | 16.0324 | 21.396 | 6.3752 | 180.4716 | 325.5912 | 251.2772 | 319.922 | 221.8751
2012-10-30 01:49 | 25.7889 | 347.0001 | 272.7019 | 341.3195 | 243.4097 | 92.4764 | 1.6546 | 42.2333 | 69.8775 | 25.5254 | 2.6264 | 42.1074 | 110.5902 | 16.0325 | 21.3654 | 6.3753 | 180.4676 | 325.6347 | 251.3365 | 319.9541 | 222.0443
2012-10-30 01:49 | 25.7881 | 347.0412 | 272.7205 | 341.3415 | 243.4168 | 92.4914 | 1.6544 | 42.2291 | 69.8639 | 25.5254 | 2.6261 | 42.1034 | 110.5688 | 16.0325 | 21.3485 | 6.3764 | 180.4327 | 325.6927 | 251.372 | 319.993 | 222.0683
2012-10-30 01:49 | 25.7934 | 347.0351 | 272.7607 | 341.3354 | 243.4456 | 92.4937 | 1.6544 | 42.2281 | 69.8607 | 25.5253 | 2.6261 | 42.1024 | 110.5641 | 16.0324 | 21.3298 | 6.3765 | 180.4248 | 325.7053 | 251.4309 | 320.0056 | 222.1158
2012-10-30 01:49 | 25.7971 | 347.0502 | 272.7372 | 341.3619 | 243.4493 | 92.4677 | 1.6506 | 42.1302 | 69.54 | 25.5243 | 2.6199 | 42.0022 | 110.0428 | 16.0318 | 21.3656 | 6.3747 | 179.5828 | 325.6846 | 251.3716 | 319.9963 | 222.0837
2012-10-30 01:49 | 25.7941 | 346.9939 | 272.7846 | 341.3589 | 243.5952 | 92.4665 | 1.6503 | 42.1226 | 69.5161 | 25.5238 | 2.6197 | 41.9973 | 110.0184 | 16.0316 | 21.3465 | 6.3746 | 179.5345 | 325.6474 | 251.4381 | 320.0124 | 222.2487
2012-10-30 01:49 | 25.7964 | 346.9696 | 272.7714 | 341.4069 | 243.2291 | 92.47 | 1.65 | 42.1151 | 69.4914 | 25.5238 | 2.6192 | 41.9897 | 109.9793 | 16.0315 | 21.3448 | 6.3749 | 179.4707 | 325.6248 | 251.4266 | 320.0621 | 221.8843
2012-10-30 01:49 | 25.8053 | 347.0013 | 272.7532 | 341.3739 | 242.2189 | 92.4775 | 1.65 | 42.1139 | 69.4873 | 25.5238 | 2.6191 | 41.9885 | 109.9741 | 16.0314 | 21.3617 | 6.3754 | 179.4614 | 325.6396 | 251.3915 | 320.0122 | 220.8572
2012-10-30 01:50 | 25.7934 | 347.0198 | 272.497 | 341.4115 | 241.07 | 92.47 | 1.6501 | 42.1164 | 69.4957 | 25.5238 | 2.6193 | 41.9907 | 109.9861 | 16.0313 | 21.3579 | 6.3749 | 179.4817 | 325.6619 | 251.1391 | 320.0536 | 219.7121
2012-10-30 01:50 | 25.7911 | 346.9909 | 272.293 | 341.3597 | 242.1734 | 92.4642 | 1.6503 | 42.1209 | 69.5102 | 25.5239 | 2.6196 | 41.9956 | 110.0117 | 16.0313 | 21.3515 | 6.3745 | 179.5219 | 325.6394 | 250.9415 | 320.0082 | 220.8219
2012-10-30 01:50 | 25.7851 | 346.9697 | 272.2598 | 341.3728 | 242.1791 | 92.4515 | 1.6505 | 42.1267 | 69.5291 | 25.524 | 2.6199 | 42.0011 | 110.0405 | 16.0313 | 21.3857 | 6.3736 | 179.5696 | 325.584 | 250.8741 | 319.9871 | 220.7934
2012-10-30 01:50 | 25.7761 | 347.0025 | 272.2276 | 341.3942 | 242.664 | 92.4648 | 1.6507 | 42.1334 | 69.5509 | 25.5241 | 2.6203 | 42.0078 | 110.075 | 16.0314 | 21.3365 | 6.3745 | 179.626 | 325.666 | 250.8911 | 320.0577 | 221.3275
2012-10-30 01:50 | 25.7724 | 347.0445 | 272.2277 | 341.4133 | 242.4918 | 92.4677 | 1.651 | 42.1391 | 69.5697 | 25.5241 | 2.6207 | 42.0135 | 110.1033 | 16.0316 | 21.3369 | 6.3747 | 179.6731 | 325.7076 | 250.8908 | 320.0764 | 221.1549
2012-10-30 01:50 | 25.7687 | 347.0332 | 272.2589 | 341.3982 | 242.2882 | 92.4434 | 1.6525 | 42.1806 | 69.7055 | 25.5246 | 2.623 | 42.0522 | 110.3047 | 16.0318 | 21.3653 | 6.373 | 180.0102 | 325.6679 | 250.8936 | 320.0329 | 220.9229
2012-10-30 01:50 | 25.7687 | 347.0256 | 272.2124 | 341.4058 | 241.8922 | 92.4538 | 1.6544 | 42.2283 | 69.8606 | 25.5255 | 2.6261 | 42.1022 | 110.5636 | 16.0323 | 21.3532 | 6.3738 | 180.4242 | 325.6724 | 250.8592 | 320.0526 | 220.539
2012-10-30 01:50 | 25.7739 | 347.0498 | 272.1517 | 341.411 | 242.3482 | 92.4613 | 1.6548 | 42.2399 | 69.8983 | 25.5258 | 2.6268 | 42.1138 | 110.624 | 16.0325 | 21.3504 | 6.3743 | 180.5224 | 325.6994 | 250.8013 | 320.0606 | 220.9978
2012-10-30 01:50 | 25.7634 | 347.0926 | 272.1606 | 341.3891 | 242.1221 | 92.4527 | 1.6551 | 42.2486 | 69.9271 | 25.5258 | 2.6273 | 42.123 | 110.6715 | 16.0325 | 21.364 | 6.3737 | 180.5986 | 325.7286 | 250.7966 | 320.0251 | 220.7581
2012-10-30 01:50 | 25.7567 | 347.0707 | 272.1073 | 341.3596 | 242.5819 | 92.4654 | 1.6554 | 42.2557 | 69.9502 | 25.5259 | 2.6277 | 42.1297 | 110.7063 | 16.0326 | 21.3332 | 6.3746 | 180.6565 | 325.7375 | 250.7741 | 320.0264 | 221.2487
2012-10-30 01:50 | 25.7544 | 347.0417 | 272.2059 | 341.3801 | 242.7874 | 92.4613 | 1.6556 | 42.26 | 69.9643 | 25.526 | 2.628 | 42.1341 | 110.7284 | 16.0327 | 21.375 | 6.3743 | 180.6927 | 325.6667 | 250.8309 | 320.0051 | 221.4124
2012-10-30 01:50 | 25.7544 | 347.0569 | 272.2679 | 341.3763 | 242.9245 | 92.4752 | 1.6557 | 42.2624 | 69.9721 | 25.526 | 2.6281 | 42.1362 | 110.7397 | 16.0327 | 21.375 | 6.3752 | 180.7117 | 325.6819 | 250.8929 | 320.0013 | 221.5495
2012-10-30 01:50 | 25.7544 | 347.095 | 272.3145 | 341.3611 | 242.7834 | 92.4764 | 1.6557 | 42.2631 | 69.9743 | 25.5261 | 2.6282 | 42.1373 | 110.7448 | 16.0328 | 21.3831 | 6.3753 | 180.7191 | 325.7119 | 250.9314 | 319.978 | 221.4003
2012-10-30 01:50 | 25.7552 | 347.0882 | 272.292 | 341.3961 | 242.7725 | 92.4665 | 1.6556 | 42.2623 | 69.9709 | 25.5264 | 2.6281 | 42.1364 | 110.7401 | 16.0328 | 21.3758 | 6.3746 | 180.7109 | 325.7124 | 250.9162 | 320.0203 | 221.3967
2012-10-30 01:50 | 25.7462 | 347.083 | 272.3295 | 341.3871 | 242.8693 | 92.4677 | 1.6556 | 42.2619 | 69.9699 | 25.5262 | 2.6281 | 42.1361 | 110.738 | 16.0329 | 21.3749 | 6.3747 | 180.7079 | 325.7081 | 250.9546 | 320.0122 | 221.4944
2012-10-30 01:50 | 25.7462 | 347.102 | 272.3722 | 341.3757 | 243.1319 | 92.4694 | 1.6556 | 42.2619 | 69.9697 | 25.5263 | 2.6281 | 42.136 | 110.738 | 16.0328 | 21.3909 | 6.3748 | 180.7076 | 325.7111 | 250.9813 | 319.9848 | 221.741
2012-10-30 01:50 | 25.7425 | 347.0717 | 272.4655 | 341.4139 | 243.1909 | 92.4735 | 1.6556 | 42.2608 | 69.9654 | 25.5265 | 2.628 | 42.1348 | 110.7323 | 16.0327 | 21.3431 | 6.3751 | 180.6977 | 325.7286 | 251.1224 | 320.0708 | 221.8478
2012-10-30 01:50 | 25.744 | 347.1112 | 272.4825 | 341.3811 | 243.1806 | 92.4706 | 1.6555 | 42.259 | 69.9603 | 25.5263 | 2.6279 | 42.1331 | 110.7229 | 16.0328 | 21.3606 | 6.3749 | 180.6832 | 325.7506 | 251.1219 | 320.0205 | 221.82
2012-10-30 01:50 | 25.7342 | 347.1471 | 272.5192 | 341.4474 | 243.2335 | 92.4775 | 1.6554 | 42.2571 | 69.954 | 25.5262 | 2.6278 | 42.1314 | 110.714 | 16.0328 | 21.3027 | 6.3754 | 180.668 | 325.8444 | 251.2165 | 320.1447 | 221.9308
2012-10-30 01:50 | 25.732 | 347.0954 | 272.5286 | 341.3958 | 243.392 | 92.4885 | 1.6554 | 42.2558 | 69.9497 | 25.5262 | 2.6277 | 42.1299 | 110.7066 | 16.0327 | 21.3366 | 6.3762 | 180.6563 | 325.7588 | 251.192 | 320.0592 | 222.0554
2012-10-30 01:50 | 25.7402 | 347.1531 | 272.6028 | 341.4154 | 243.4394 | 92.4816 | 1.6553 | 42.254 | 69.9434 | 25.5263 | 2.6276 | 42.1279 | 110.6956 | 16.0328 | 21.3608 | 6.3757 | 180.639 | 325.7923 | 251.242 | 320.0546 | 222.0786
2012-10-30 01:50 | 25.7387 | 347.144 | 272.6362 | 341.4215 | 243.5163 | 92.4868 | 1.6553 | 42.2525 | 69.9386 | 25.5263 | 2.6275 | 42.1266 | 110.6897 | 16.0327 | 21.3353 | 6.376 | 180.6283 | 325.8087 | 251.3009 | 320.0862 | 222.181
2012-10-30 01:51 | 25.7313 | 347.1366 | 272.6869 | 341.4598 | 243.5089 | 92.4798 | 1.6552 | 42.2512 | 69.9342 | 25.5264 | 2.6274 | 42.1251 | 110.681 | 16.0328 | 21.3158 | 6.3756 | 180.6152 | 325.8208 | 251.3711 | 320.144 | 222.1931
2012-10-30 01:51 | 25.7305 | 347.1662 | 272.7094 | 341.4399 | 243.2573 | 92.4879 | 1.6551 | 42.2499 | 69.9293 | 25.5265 | 2.6274 | 42.124 | 110.675 | 16.0328 | 21.3672 | 6.3761 | 180.6043 | 325.799 | 251.3422 | 320.0727 | 221.8901
2012-10-30 01:51 | 25.7305 | 347.1738 | 272.7404 | 341.4818 | 243.3513 | 92.4821 | 1.6551 | 42.2477 | 69.9225 | 25.5264 | 2.6272 | 42.1217 | 110.6636 | 16.0327 | 21.3391 | 6.3757 | 180.5861 | 325.8347 | 251.4013 | 320.1427 | 222.0122
2012-10-30 01:51 | 25.7268 | 347.1321 | 272.7018 | 341.4362 | 243.1242 | 92.4827 | 1.6551 | 42.2473 | 69.9216 | 25.5262 | 2.6272 | 42.1214 | 110.6622 | 16.0327 | 21.3234 | 6.3758 | 180.5838 | 325.8087 | 251.3784 | 320.1128 | 221.8008
2012-10-30 01:51 | 25.7238 | 347.1215 | 272.7182 | 341.4827 | 243.0194 | 92.4862 | 1.655 | 42.2451 | 69.9143 | 25.5262 | 2.6271 | 42.1194 | 110.6511 | 16.0328 | 21.3685 | 6.376 | 180.5654 | 325.753 | 251.3497 | 320.1142 | 221.6509
2012-10-30 01:51 | 25.7253 | 347.1914 | 272.6266 | 341.4538 | 243.2286 | 92.485 | 1.6549 | 42.244 | 69.9106 | 25.5263 | 2.627 | 42.1183 | 110.6454 | 16.0328 | 21.3339 | 6.3759 | 180.556 | 325.8575 | 251.2927 | 320.1199 | 221.8947
2012-10-30 01:51 | 25.7238 | 347.1405 | 272.7027 | 341.4599 | 243.4191 | 92.4908 | 1.6549 | 42.2424 | 69.905 | 25.5264 | 2.6269 | 42.1166 | 110.637 | 16.0327 | 21.3324 | 6.3763 | 180.542 | 325.8081 | 251.3703 | 320.1275 | 222.0867
2012-10-30 01:51 | 25.7245 | 347.1374 | 272.715 | 341.4948 | 243.4511 | 92.5076 | 1.6548 | 42.2409 | 69.9003 | 25.5263 | 2.6268 | 42.1151 | 110.6294 | 16.0326 | 21.3491 | 6.3775 | 180.5296 | 325.7883 | 251.3659 | 320.1457 | 222.102
2012-10-30 01:51 | 25.7313 | 347.1556 | 272.7374 | 341.4788 | 243.4383 | 92.5012 | 1.6548 | 42.2411 | 69.901 | 25.5262 | 2.6269 | 42.1154 | 110.6309 | 16.0326 | 21.2877 | 6.377 | 180.5319 | 325.8679 | 251.4497 | 320.1911 | 222.1506
2012-10-30 01:51 | 25.732 | 347.1867 | 272.7846 | 341.4643 | 243.2313 | 92.4972 | 1.6547 | 42.2396 | 69.8959 | 25.5263 | 2.6268 | 42.1138 | 110.6228 | 16.0326 | 21.3165 | 6.3768 | 180.5188 | 325.8702 | 251.4681 | 320.1478 | 221.9148
2012-10-30 01:51 | 25.7342 | 347.1699 | 272.7713 | 341.5388 | 243.3472 | 92.4925 | 1.6547 | 42.2379 | 69.8898 | 25.5264 | 2.6266 | 42.112 | 110.6128 | 16.0327 | 21.3147 | 6.3764 | 180.5026 | 325.8552 | 251.4566 | 320.2241 | 222.0325
2012-10-30 01:51 | 25.7342 | 347.2079 | 272.7441 | 341.4931 | 243.3903 | 92.5012 | 1.6546 | 42.2355 | 69.8824 | 25.5263 | 2.6265 | 42.1095 | 110.6004 | 16.0326 | 21.3107 | 6.377 | 180.4828 | 325.8972 | 251.4334 | 320.1824 | 222.0796
2012-10-30 01:51 | 25.7327 | 347.2483 | 272.7659 | 341.5068 | 243.6749 | 92.4931 | 1.6546 | 42.2349 | 69.8802 | 25.5264 | 2.6265 | 42.1089 | 110.5975 | 16.0326 | 21.3012 | 6.3765 | 180.4777 | 325.9471 | 251.4647 | 320.2056 | 222.3737
2012-10-30 01:51 | 25.7425 | 347.2086 | 272.8688 | 341.5204 | 243.5514 | 92.4977 | 1.6545 | 42.233 | 69.8739 | 25.5263 | 2.6263 | 42.1071 | 110.5873 | 16.0326 | 21.319 | 6.3768 | 180.4612 | 325.8896 | 251.5498 | 320.2014 | 222.2324
2012-10-30 01:51 | 25.7432 | 347.2246 | 272.8733 | 341.5211 | 243.4189 | 92.5012 | 1.6544 | 42.23 | 69.8639 | 25.5264 | 2.6262 | 42.1042 | 110.5724 | 16.0326 | 21.3157 | 6.377 | 180.4363 | 325.9089 | 251.5576 | 320.2054 | 222.1032
2012-10-30 01:51 | 25.7432 | 347.255 | 272.8889 | 341.5325 | 243.556 | 92.5018 | 1.6543 | 42.2273 | 69.8548 | 25.5265 | 2.626 | 42.1015 | 110.5582 | 16.0326 | 21.3237 | 6.3771 | 180.413 | 325.9313 | 251.5652 | 320.2088 | 222.2323
2012-10-30 01:51 | 25.7455 | 347.2573 | 272.8834 | 341.5615 | 243.3389 | 92.5006 | 1.6515 | 42.1551 | 69.6181 | 25.5258 | 2.6215 | 42.0286 | 110.1789 | 16.0322 | 21.322 | 6.377 | 179.797 | 325.9353 | 251.5614 | 320.2395 | 222.0169
2012-10-30 01:51 | 25.7589 | 347.2479 | 272.8425 | 341.5444 | 243.4973 | 92.5006 | 1.6503 | 42.1237 | 69.5165 | 25.525 | 2.6197 | 41.9982 | 110.0219 | 16.0318 | 21.2913 | 6.377 | 179.5384 | 325.9566 | 251.5512 | 320.2531 | 222.206
2012-10-30 01:51 | 25.747 | 347.2169 | 272.819 | 341.5287 | 243.6852 | 92.4925 | 1.6499 | 42.1144 | 69.4863 | 25.5247 | 2.6191 | 41.9889 | 109.9745 | 16.0316 | 21.3275 | 6.3764 | 179.4607 | 325.8894 | 251.4915 | 320.2012 | 222.3577
2012-10-30 01:51 | 25.7597 | 347.2525 | 272.8898 | 341.5567 | 243.0278 | 92.4833 | 1.6499 | 42.1142 | 69.4862 | 25.5246 | 2.6191 | 41.9888 | 109.9742 | 16.0315 | 21.3522 | 6.3758 | 179.4604 | 325.9003 | 251.5376 | 320.2045 | 221.6756
2012-10-30 01:51 | 25.7574 | 347.2388 | 272.6704 | 341.5467 | 241.8378 | 92.4856 | 1.6501 | 42.1175 | 69.4964 | 25.5248 | 2.6193 | 41.992 | 109.9917 | 16.0315 | 21.2857 | 6.376 | 179.488 | 325.9531 | 251.3847 | 320.261 | 220.5521
2012-10-30 01:51 | 25.7559 | 347.2411 | 272.4982 | 341.5719 | 242.4518 | 92.4844 | 1.6501 | 42.1195 | 69.5027 | 25.5249 | 2.6195 | 41.994 | 110.0024 | 16.0314 | 21.3204 | 6.3759 | 179.5051 | 325.9207 | 251.1778 | 320.2515 | 221.1314
2012-10-30 01:52 | 25.7574 | 347.2654 | 272.5075 | 341.5011 | 243.4801 | 92.4954 | 1.6504 | 42.1256 | 69.5232 | 25.5248 | 2.6198 | 42.0001 | 110.0338 | 16.0315 | 21.3178 | 6.3766 | 179.5569 | 325.9476 | 251.1897 | 320.1833 | 222.1623
2012-10-30 01:52 | 25.753 | 347.2382 | 272.6078 | 341.55 | 243.5698 | 92.4873 | 1.6506 | 42.1318 | 69.5442 | 25.5247 | 2.6202 | 42.0061 | 110.0653 | 16.0315 | 21.2854 | 6.3761 | 179.6095 | 325.9528 | 251.3224 | 320.2646 | 222.2844
2012-10-30 01:52 | 25.7522 | 347.1917 | 272.6613 | 341.5415 | 243.5219 | 92.4764 | 1.6509 | 42.1383 | 69.5649 | 25.5249 | 2.6206 | 42.0127 | 110.0991 | 16.0316 | 21.3367 | 6.3753 | 179.6639 | 325.855 | 251.3246 | 320.2048 | 222.1852
2012-10-30 01:52 | 25.7559 | 347.2144 | 272.7232 | 341.5072 | 243.6119 | 92.4896 | 1.6519 | 42.1644 | 69.6505 | 25.5251 | 2.6224 | 42.0415 | 110.2493 | 16.0318 | 21.3163 | 6.3762 | 179.8998 | 325.8981 | 251.4069 | 320.1909 | 222.2956
2012-10-30 01:52 | 25.7582 | 347.27 | 272.6751 | 341.5399 | 243.3242 | 92.4862 | 1.6542 | 42.2251 | 69.8482 | 25.5262 | 2.6259 | 42.0994 | 110.5494 | 16.0323 | 21.3307 | 6.376 | 180.3977 | 325.9393 | 251.3444 | 320.2092 | 221.9935
2012-10-30 01:52 | 25.7559 | 347.2791 | 272.6534 | 341.5224 | 243.0436 | 92.4943 | 1.6547 | 42.2387 | 69.8922 | 25.5266 | 2.6267 | 42.1125 | 110.6172 | 16.0325 | 21.3565 | 6.3766 | 180.5094 | 325.9226 | 251.2969 | 320.1659 | 221.6871
2012-10-30 01:52 | 25.753 | 347.3028 | 272.6311 | 341.5195 | 242.6292 | 92.4868 | 1.6552 | 42.2519 | 69.9356 | 25.5267 | 2.6275 | 42.1258 | 110.6854 | 16.0326 | 21.2974 | 6.376 | 180.621 | 326.0054 | 251.3337 | 320.2221 | 221.3318
2012-10-30 01:52 | 25.7522 | 347.2526 | 272.5721 | 341.5187 | 243.0556 | 92.5024 | 1.6554 | 42.2574 | 69.9527 | 25.5271 | 2.6278 | 42.1314 | 110.7143 | 16.0328 | 21.3367 | 6.3771 | 180.6669 | 325.9159 | 251.2354 | 320.182 | 221.7189
2012-10-30 01:52 | 25.7477 | 347.3242 | 272.5405 | 341.5066 | 243.0119 | 92.4995 | 1.6554 | 42.2586 | 69.9566 | 25.5271 | 2.6279 | 42.1327 | 110.7204 | 16.0329 | 21.264 | 6.3769 | 180.677 | 326.0602 | 251.2765 | 320.2426 | 221.7479
2012-10-30 01:52 | 25.7552 | 347.305 | 272.5945 | 341.5179 | 243.3251 | 92.4925 | 1.6555 | 42.2601 | 69.9617 | 25.5271 | 2.628 | 42.1342 | 110.7282 | 16.0329 | 21.3237 | 6.3764 | 180.6899 | 325.9813 | 251.2708 | 320.1942 | 222.0014
2012-10-30 01:52 | 25.75 | 347.277 | 272.6242 | 341.5317 | 243.3512 | 92.5 | 1.6555 | 42.2606 | 69.9629 | 25.5272 | 2.628 | 42.1347 | 110.7304 | 16.0329 | 21.3265 | 6.377 | 180.6933 | 325.9505 | 251.2977 | 320.2052 | 222.0247
2012-10-30 01:52 | 25.7515 | 347.3166 | 272.6567 | 341.5256 | 243.5016 | 92.5058 | 1.6556 | 42.2619 | 69.9672 | 25.5272 | 2.6281 | 42.1359 | 110.7367 | 16.033 | 21.316 | 6.3773 | 180.7039 | 326.0006 | 251.3407 | 320.2096 | 222.1856
2012-10-30 01:52 | 25.7537 | 347.3111 | 272.721 | 341.524 | 243.4176 | 92.5133 | 1.6556 | 42.2616 | 69.9668 | 25.527 | 2.6281 | 42.1357 | 110.7349 | 16.033 | 21.3462 | 6.3779 | 180.7017 | 325.9649 | 251.3748 | 320.1778 | 222.0714
2012-10-30 01:52 | 25.7574 | 347.3263 | 272.7518 | 341.5201 | 243.2999 | 92.5041 | 1.6555 | 42.26 | 69.9604 | 25.5274 | 2.6279 | 42.134 | 110.725 | 16.0331 | 21.4061 | 6.3772 | 180.6854 | 325.9202 | 251.3457 | 320.114 | 221.8938
2012-10-30 01:52 | 25.7507 | 347.2853 | 272.7568 | 341.521 | 243.5557 | 92.5018 | 1.6554 | 42.2591 | 69.9573 | 25.5274 | 2.6279 | 42.1333 | 110.722 | 16.0331 | 21.3432 | 6.3771 | 180.6792 | 325.9421 | 251.4136 | 320.1778 | 222.2125
2012-10-30 01:52 | 25.7537 | 347.3226 | 272.7947 | 341.5507 | 243.5744 | 92.5162 | 1.6554 | 42.2576 | 69.9526 | 25.5273 | 2.6278 | 42.1316 | 110.713 | 16.0331 | 21.3422 | 6.3781 | 180.6656 | 325.9804 | 251.4525 | 320.2085 | 222.2322
2012-10-30 01:52 | 25.7597 | 347.359 | 272.8743 | 341.5376 | 243.455 | 92.5052 | 1.6553 | 42.256 | 69.9476 | 25.5273 | 2.6277 | 42.1301 | 110.7059 | 16.033 | 21.3482 | 6.3773 | 180.6536 | 326.0108 | 251.5261 | 320.1894 | 222.1068
2012-10-30 01:52 | 25.7612 | 347.2806 | 272.8409 | 341.4973 | 243.3781 | 92.5122 | 1.6553 | 42.2549 | 69.944 | 25.5272 | 2.6276 | 42.129 | 110.7002 | 16.033 | 21.3497 | 6.3778 | 180.6442 | 325.9309 | 251.4912 | 320.1476 | 222.0284
2012-10-30 01:52 | 25.7485 | 347.2869 | 272.8321 | 341.5416 | 243.3732 | 92.5064 | 1.6552 | 42.2535 | 69.9396 | 25.5272 | 2.6276 | 42.1277 | 110.6936 | 16.0329 | 21.325 | 6.3774 | 180.6332 | 325.9619 | 251.5071 | 320.2166 | 222.0482
2012-10-30 01:52 | 25.7582 | 347.3233 | 272.8379 | 341.5399 | 243.5123 | 92.5076 | 1.6552 | 42.2523 | 69.9345 | 25.5276 | 2.6275 | 42.1265 | 110.6869 | 16.033 | 21.3387 | 6.3775 | 180.6214 | 325.9846 | 251.4992 | 320.2012 | 222.1736
2012-10-30 01:52 | 25.7597 | 347.2867 | 272.8782 | 341.5833 | 243.4315 | 92.5145 | 1.6551 | 42.2504 | 69.9282 | 25.5276 | 2.6274 | 42.1245 | 110.6771 | 16.0329 | 21.3803 | 6.3779 | 180.6053 | 325.9064 | 251.4979 | 320.203 | 222.0512
2012-10-30 01:52 | 25.7559 | 347.2981 | 272.8744 | 341.5643 | 243.5021 | 92.5151 | 1.655 | 42.2479 | 69.9205 | 25.5274 | 2.6272 | 42.1221 | 110.664 | 16.033 | 21.3204 | 6.378 | 180.5845 | 325.9777 | 251.554 | 320.2439 | 222.1817
2012-10-30 01:52 | 25.7612 | 347.2692 | 272.8642 | 341.5696 | 243.3232 | 92.5128 | 1.6549 | 42.246 | 69.914 | 25.5274 | 2.6271 | 42.1199 | 110.6533 | 16.0329 | 21.3176 | 6.3778 | 180.5674 | 325.9516 | 251.5466 | 320.252 | 222.0056
2012-10-30 01:53 | 25.7619 | 347.3003 | 272.83 | 341.5817 | 243.4925 | 92.5162 | 1.6549 | 42.2445 | 69.9095 | 25.5273 | 2.627 | 42.1187 | 110.647 | 16.0328 | 21.3304 | 6.3781 | 180.5566 | 325.9699 | 251.4996 | 320.2513 | 222.1621
2012-10-30 01:53 | 25.7597 | 347.34 | 272.8627 | 341.5795 | 243.0866 | 92.5232 | 1.6548 | 42.2432 | 69.9052 | 25.5272 | 2.6269 | 42.1174 | 110.6396 | 16.0329 | 21.3884 | 6.3785 | 180.5448 | 325.9516 | 251.4743 | 320.1911 | 221.6982
2012-10-30 01:53 | 25.7634 | 347.2714 | 272.7888 | 341.5756 | 243.2588 | 92.5105 | 1.6548 | 42.2425 | 69.903 | 25.5272 | 2.6269 | 42.1167 | 110.6355 | 16.033 | 21.3559 | 6.3777 | 180.5385 | 325.9155 | 251.4329 | 320.2197 | 221.9029
2012-10-30 01:53 | 25.7747 | 347.2903 | 272.8467 | 341.5945 | 243.4621 | 92.5191 | 1.6547 | 42.2409 | 69.8967 | 25.5275 | 2.6268 | 42.115 | 110.6273 | 16.0329 | 21.3713 | 6.3783 | 180.524 | 325.919 | 251.4754 | 320.2232 | 222.0908
2012-10-30 01:53 | 25.7784 | 347.3397 | 272.862 | 341.6248 | 243.6187 | 92.5243 | 1.6547 | 42.239 | 69.8912 | 25.5274 | 2.6267 | 42.1133 | 110.6188 | 16.0328 | 21.3669 | 6.3786 | 180.51 | 325.9728 | 251.4951 | 320.2579 | 222.2518
2012-10-30 01:53 | 25.7754 | 347.31 | 272.9211 | 341.5914 | 243.7293 | 92.5307 | 1.6545 | 42.2363 | 69.8819 | 25.5274 | 2.6265 | 42.1103 | 110.6035 | 16.0328 | 21.3439 | 6.3791 | 180.4854 | 325.9661 | 251.5772 | 320.2475 | 222.3854
2012-10-30 01:53 | 25.7732 | 347.2812 | 272.9189 | 341.6158 | 243.8133 | 92.518 | 1.6545 | 42.2359 | 69.8807 | 25.5273 | 2.6265 | 42.1101 | 110.6023 | 16.0328 | 21.2975 | 6.3782 | 180.483 | 325.9837 | 251.6214 | 320.3183 | 222.5158
2012-10-30 01:53 | 25.7791 | 347.3099 | 272.9519 | 341.5684 | 243.7408 | 92.5197 | 1.6545 | 42.2352 | 69.8781 | 25.5274 | 2.6264 | 42.1091 | 110.5965 | 16.0329 | 21.3676 | 6.3783 | 180.4746 | 325.9423 | 251.5843 | 320.2008 | 222.3732
2012-10-30 01:53 | 25.7896 | 347.3585 | 273.0128 | 341.617 | 243.6377 | 92.5272 | 1.6544 | 42.2329 | 69.8702 | 25.5276 | 2.6263 | 42.107 | 110.5852 | 16.0329 | 21.35 | 6.3788 | 180.4554 | 326.0085 | 251.6628 | 320.267 | 222.2877
2012-10-30 01:53 | 25.7904 | 347.306 | 273.0136 | 341.5988 | 243.6581 | 92.5301 | 1.6543 | 42.2306 | 69.8622 | 25.5277 | 2.6261 | 42.1048 | 110.5731 | 16.0329 | 21.3709 | 6.379 | 180.4354 | 325.9351 | 251.6427 | 320.2279 | 222.2872
2012-10-30 01:53 | 25.7971 | 347.3431 | 273.0397 | 341.6397 | 243.6452 | 92.5295 | 1.6543 | 42.229 | 69.8576 | 25.5274 | 2.6261 | 42.1031 | 110.566 | 16.0327 | 21.3736 | 6.379 | 180.4236 | 325.9695 | 251.6661 | 320.2661 | 222.2716
2012-10-30 01:53 | 25.7971 | 347.3203 | 273.0281 | 341.5902 | 243.5668 | 92.5151 | 1.6504 | 42.129 | 69.5313 | 25.526 | 2.62 | 42.0029 | 110.0458 | 16.0319 | 21.3656 | 6.378 | 179.5771 | 325.9547 | 251.6625 | 320.2246 | 222.2012
2012-10-30 01:53 | 25.8098 | 347.3178 | 272.9322 | 341.6182 | 242.6389 | 92.5226 | 1.6499 | 42.1153 | 69.4866 | 25.5258 | 2.6192 | 41.99 | 109.9797 | 16.0316 | 21.3622 | 6.3785 | 179.4664 | 325.9556 | 251.57 | 320.256 | 221.2767
2012-10-30 01:53 | 25.8106 | 347.2996 | 272.8477 | 341.6037 | 243.3726 | 92.5133 | 1.6499 | 42.1139 | 69.4821 | 25.5257 | 2.6191 | 41.9883 | 109.9715 | 16.0316 | 21.3751 | 6.3779 | 179.4536 | 325.9245 | 251.4726 | 320.2286 | 221.9975
2012-10-30 01:53 | 25.8158 | 347.2857 | 272.8141 | 341.6318 | 243.9931 | 92.5064 | 1.65 | 42.118 | 69.4954 | 25.5258 | 2.6194 | 41.9925 | 109.9933 | 16.0316 | 21.3762 | 6.3774 | 179.4887 | 325.9095 | 251.4379 | 320.2556 | 222.6169
2012-10-30 01:53 | 25.8121 | 347.3125 | 272.8375 | 341.5786 | 244.0403 | 92.5029 | 1.6502 | 42.124 | 69.5152 | 25.5259 | 2.6197 | 41.9984 | 110.0235 | 16.0317 | 21.3605 | 6.3771 | 179.5386 | 325.952 | 251.477 | 320.2181 | 222.6798
2012-10-30 01:53 | 25.8196 | 347.2667 | 272.8334 | 341.6013 | 243.7147 | 92.5099 | 1.6505 | 42.1307 | 69.5368 | 25.526 | 2.6201 | 42.005 | 110.0585 | 16.0316 | 21.364 | 6.3776 | 179.5953 | 325.9027 | 251.4694 | 320.2373 | 222.3507
2012-10-30 01:53 | 25.8136 | 347.3102 | 272.8041 | 341.5877 | 244.1672 | 92.5139 | 1.6507 | 42.137 | 69.5574 | 25.5261 | 2.6205 | 42.0113 | 110.0902 | 16.0319 | 21.354 | 6.3779 | 179.6476 | 325.9562 | 251.4501 | 320.2337 | 222.8132
2012-10-30 01:53 | 25.8151 | 347.2812 | 272.9142 | 341.5588 | 243.6515 | 92.5012 | 1.6509 | 42.1423 | 69.5742 | 25.5263 | 2.6208 | 42.0168 | 110.1186 | 16.0319 | 21.3033 | 6.377 | 179.6928 | 325.9779 | 251.6109 | 320.2555 | 222.3482
2012-10-30 01:53 | 25.8173 | 347.3329 | 272.8583 | 341.5648 | 243.4381 | 92.5093 | 1.6538 | 42.2182 | 69.8221 | 25.5274 | 2.6255 | 42.0934 | 110.5172 | 16.0324 | 21.3497 | 6.3776 | 180.3393 | 325.9832 | 251.5086 | 320.2151 | 222.0884
2012-10-30 01:54 | 25.8136 | 347.2835 | 272.7925 | 341.5573 | 243.4932 | 92.5168 | 1.6545 | 42.2363 | 69.8804 | 25.528 | 2.6265 | 42.1104 | 110.6034 | 16.0328 | 21.3219 | 6.3781 | 180.4838 | 325.9616 | 251.4706 | 320.2354 | 222.1713
2012-10-30 01:54 | 25.8098 | 347.2874 | 272.8042 | 341.5268 | 243.7089 | 92.5145 | 1.6549 | 42.2464 | 69.913 | 25.5282 | 2.6271 | 42.1206 | 110.6564 | 16.0329 | 21.3381 | 6.3779 | 180.5694 | 325.9493 | 251.4661 | 320.1887 | 222.3708
2012-10-30 01:54 | 25.8121 | 347.3011 | 272.8298 | 341.5862 | 243.4016 | 92.5393 | 1.6552 | 42.2534 | 69.9359 | 25.5283 | 2.6276 | 42.1274 | 110.6919 | 16.033 | 21.3485 | 6.3797 | 180.6278 | 325.9526 | 251.4813 | 320.2377 | 222.0531
2012-10-30 01:54 | 25.8091 | 347.3133 | 272.7725 | 341.5299 | 243.5671 | 92.5266 | 1.6554 | 42.2582 | 69.9524 | 25.5282 | 2.6279 | 42.1322 | 110.7171 | 16.0329 | 21.3415 | 6.3788 | 180.6695 | 325.9718 | 251.431 | 320.1884 | 222.2256
2012-10-30 01:54 | 25.8053 | 347.2905 | 272.8191 | 341.5033 | 243.7083 | 92.5382 | 1.6555 | 42.2618 | 69.9642 | 25.5282 | 2.6281 | 42.1359 | 110.7354 | 16.0331 | 21.3015 | 6.3796 | 180.6996 | 325.989 | 251.5176 | 320.2018 | 222.4068
2012-10-30 01:54 | 25.8158 | 347.3162 | 272.8025 | 341.5366 | 243.6091 | 92.5324 | 1.6556 | 42.2638 | 69.9706 | 25.5283 | 2.6282 | 42.1378 | 110.7453 | 16.0331 | 21.3361 | 6.3792 | 180.7159 | 325.9801 | 251.4664 | 320.2005 | 222.273
2012-10-30 01:54 | 25.8098 | 347.3216 | 272.8314 | 341.466 | 243.7441 | 92.5382 | 1.6556 | 42.2644 | 69.9711 | 25.5288 | 2.6282 | 42.1384 | 110.7472 | 16.0333 | 21.29 | 6.3796 | 180.7184 | 326.0316 | 251.5414 | 320.176 | 222.4541
2012-10-30 01:54 | 25.8136 | 347.3178 | 272.8313 | 341.5421 | 243.6931 | 92.544 | 1.6555 | 42.2627 | 69.9658 | 25.5288 | 2.6281 | 42.1368 | 110.7394 | 16.0333 | 21.3339 | 6.38 | 180.7052 | 325.9839 | 251.4974 | 320.2082 | 222.3592
2012-10-30 01:54 | 25.8091 | 347.2828 | 272.8539 | 341.5299 | 243.6729 | 92.544 | 1.6554 | 42.2609 | 69.96 | 25.5287 | 2.628 | 42.1348 | 110.7284 | 16.0333 | 21.2813 | 6.38 | 180.6884 | 326.0015 | 251.5726 | 320.2486 | 222.3916
2012-10-30 01:54 | 25.8136 | 347.295 | 272.8429 | 341.523 | 243.6539 | 92.5428 | 1.6554 | 42.2603 | 69.9581 | 25.5286 | 2.6279 | 42.1342 | 110.7247 | 16.0334 | 21.3339 | 6.3799 | 180.6828 | 325.9611 | 251.509 | 320.1891 | 222.32
2012-10-30 01:54 | 25.8091 | 347.3095 | 272.8268 | 341.5223 | 243.5397 | 92.5486 | 1.6553 | 42.258 | 69.9507 | 25.5285 | 2.6278 | 42.1321 | 110.7143 | 16.0332 | 21.3254 | 6.3803 | 180.665 | 325.9841 | 251.5014 | 320.1969 | 222.2143
2012-10-30 01:54 | 25.8091 | 347.2867 | 272.8772 | 341.5033 | 243.4691 | 92.5445 | 1.6553 | 42.2568 | 69.9466 | 25.5286 | 2.6277 | 42.1308 | 110.7078 | 16.0333 | 21.3495 | 6.38 | 180.6543 | 325.9372 | 251.5277 | 320.1538 | 222.1196
2012-10-30 01:54 | 25.8203 | 347.3169 | 272.8535 | 341.5602 | 243.2805 | 92.5486 | 1.6552 | 42.2556 | 69.943 | 25.5284 | 2.6276 | 42.1296 | 110.7016 | 16.0332 | 21.3206 | 6.3803 | 180.6446 | 325.9963 | 251.5329 | 320.2396 | 221.9599
2012-10-30 01:54 | 25.8188 | 347.2925 | 272.852 | 341.5282 | 243.2515 | 92.5492 | 1.6552 | 42.2549 | 69.9403 | 25.5286 | 2.6276 | 42.1289 | 110.6982 | 16.0332 | 21.3351 | 6.3803 | 180.6385 | 325.9574 | 251.5169 | 320.1931 | 221.9164
2012-10-30 01:54 | 25.8113 | 347.3003 | 272.8135 | 341.5702 | 243.4948 | 92.5497 | 1.6552 | 42.254 | 69.9371 | 25.5286 | 2.6275 | 42.1281 | 110.6935 | 16.0332 | 21.3557 | 6.3804 | 180.6306 | 325.9446 | 251.4578 | 320.2145 | 222.1391
2012-10-30 01:54 | 25.8128 | 347.3056 | 272.8033 | 341.5108 | 243.559 | 92.533 | 1.6551 | 42.2533 | 69.9349 | 25.5286 | 2.6275 | 42.1273 | 110.6903 | 16.0331 | 21.3171 | 6.3792 | 180.6251 | 325.9885 | 251.4862 | 320.1937 | 222.2419
2012-10-30 01:54 | 25.8128 | 347.3132 | 272.8266 | 341.507 | 243.5786 | 92.5336 | 1.6551 | 42.2524 | 69.9324 | 25.5284 | 2.6275 | 42.1264 | 110.6853 | 16.0332 | 21.285 | 6.3793 | 180.6176 | 326.0282 | 251.5416 | 320.222 | 222.2936
2012-10-30 01:54 | 25.8188 | 347.2849 | 272.852 | 341.5663 | 243.5925 | 92.5422 | 1.655 | 42.2502 | 69.9247 | 25.5286 | 2.6274 | 42.1243 | 110.6756 | 16.0329 | 21.303 | 6.3799 | 180.6003 | 325.9819 | 251.549 | 320.2633 | 222.2895
2012-10-30 01:54 | 25.8255 | 347.2916 | 272.882 | 341.5806 | 243.5012 | 92.548 | 1.6549 | 42.2479 | 69.9177 | 25.5284 | 2.6272 | 42.1221 | 110.663 | 16.0331 | 21.2896 | 6.3803 | 180.5807 | 326.002 | 251.5924 | 320.291 | 222.2116
2012-10-30 01:54 | 25.8278 | 347.3092 | 272.8649 | 341.5791 | 243.6328 | 92.5463 | 1.6549 | 42.2466 | 69.9129 | 25.5285 | 2.6271 | 42.1206 | 110.6553 | 16.0331 | 21.3281 | 6.3801 | 180.5683 | 325.9811 | 251.5368 | 320.251 | 222.3047
2012-10-30 01:54 | 25.83 | 347.2771 | 272.8981 | 341.547 | 243.6193 | 92.5515 | 1.6548 | 42.2451 | 69.9078 | 25.5286 | 2.627 | 42.1192 | 110.648 | 16.0331 | 21.3222 | 6.3805 | 180.5558 | 325.9549 | 251.5759 | 320.2248 | 222.2971
2012-10-30 01:54 | 25.83 | 347.319 | 272.8981 | 341.5661 | 243.635 | 92.5497 | 1.6547 | 42.2425 | 69.8996 | 25.5285 | 2.6268 | 42.1164 | 110.6335 | 16.0331 | 21.3543 | 6.3804 | 180.5331 | 325.9647 | 251.5438 | 320.2118 | 222.2807
2012-10-30 01:55 | 25.8315 | 347.3052 | 272.9112 | 341.5523 | 243.6522 | 92.5503 | 1.6547 | 42.241 | 69.8947 | 25.5285 | 2.6268 | 42.115 | 110.6259 | 16.033 | 21.3197 | 6.3804 | 180.5207 | 325.9855 | 251.5915 | 320.2326 | 222.3325
2012-10-30 01:55 | 25.8398 | 347.3097 | 272.9622 | 341.6101 | 243.6291 | 92.544 | 1.6546 | 42.24 | 69.8906 | 25.5287 | 2.6267 | 42.1143 | 110.6215 | 16.0331 | 21.3601 | 6.38 | 180.5121 | 325.9496 | 251.6021 | 320.25 | 222.269
2012-10-30 01:55 | 25.8398 | 347.3021 | 272.9467 | 341.5721 | 243.5508 | 92.5596 | 1.6546 | 42.2402 | 69.8915 | 25.5286 | 2.6267 | 42.1143 | 110.6222 | 16.033 | 21.324 | 6.3811 | 180.5137 | 325.9781 | 251.6227 | 320.2481 | 222.2268
2012-10-30 01:55 | 25.8495 | 347.289 | 272.9835 | 341.6046 | 243.2901 | 92.5613 | 1.6545 | 42.2377 | 69.883 | 25.5287 | 2.6266 | 42.1118 | 110.61 | 16.033 | 21.3819 | 6.3812 | 180.493 | 325.9071 | 251.6016 | 320.2227 | 221.9082
2012-10-30 01:55 | 25.842 | 347.2701 | 272.9256 | 341.5628 | 243.408 | 92.5515 | 1.6544 | 42.2354 | 69.876 | 25.5285 | 2.6264 | 42.1095 | 110.5974 | 16.033 | 21.3382 | 6.3805 | 180.4734 | 325.9319 | 251.5874 | 320.2246 | 222.0698
2012-10-30 01:55 | 25.8472 | 347.3324 | 272.9463 | 341.5719 | 243.5738 | 92.5422 | 1.6544 | 42.2337 | 69.8713 | 25.5282 | 2.6263 | 42.1077 | 110.5889 | 16.0329 | 21.3394 | 6.3799 | 180.4602 | 325.993 | 251.6069 | 320.2325 | 222.2344
2012-10-30 01:55 | 25.845 | 347.3302 | 272.9247 | 341.5506 | 243.6147 | 92.5561 | 1.6543 | 42.2322 | 69.865 | 25.5286 | 2.6263 | 42.1063 | 110.5824 | 16.0328 | 21.3332 | 6.3808 | 180.4474 | 325.997 | 251.5915 | 320.2174 | 222.2815
2012-10-30 01:55 | 25.8547 | 347.2638 | 272.9306 | 341.625 | 243.8086 | 92.5567 | 1.6543 | 42.2315 | 69.8632 | 25.5285 | 2.6262 | 42.1057 | 110.5779 | 16.0329 | 21.383 | 6.3809 | 180.4411 | 325.8808 | 251.5476 | 320.242 | 222.4256
2012-10-30 01:55 | 25.8577 | 347.32 | 273.0305 | 341.6052 | 243.843 | 92.5324 | 1.6507 | 42.1388 | 69.5597 | 25.5273 | 2.6206 | 42.0145 | 110.1038 | 16.0323 | 21.3459 | 6.3792 | 179.6635 | 325.9741 | 251.6846 | 320.2593 | 222.4971
2012-10-30 01:55 | 25.8562 | 347.2957 | 273.0329 | 341.5961 | 243.7905 | 92.5359 | 1.6501 | 42.1219 | 69.5051 | 25.527 | 2.6195 | 41.996 | 110.0091 | 16.032 | 21.3605 | 6.3794 | 179.5142 | 325.9352 | 251.6724 | 320.2356 | 222.43
2012-10-30 01:55 | 25.8562 | 347.2653 | 273.0096 | 341.6037 | 243.2419 | 92.5313 | 1.6499 | 42.1158 | 69.485 | 25.5269 | 2.6192 | 41.9901 | 109.9795 | 16.0318 | 21.3444 | 6.3791 | 179.4645 | 325.9209 | 251.6652 | 320.2593 | 221.8975
2012-10-30 01:55 | 25.8532 | 347.2737 | 272.9291 | 341.654 | 242.208 | 92.5284 | 1.6498 | 42.115 | 69.4827 | 25.5268 | 2.6191 | 41.9895 | 109.9764 | 16.0318 | 21.3334 | 6.3789 | 179.4591 | 325.9403 | 251.5957 | 320.3206 | 220.8746
2012-10-30 01:55 | 25.8495 | 347.2472 | 272.6384 | 341.597 | 241.7495 | 92.5289 | 1.65 | 42.1181 | 69.4933 | 25.5267 | 2.6193 | 41.9925 | 109.9925 | 16.0317 | 21.3738 | 6.3789 | 179.4859 | 325.8734 | 251.2646 | 320.2232 | 220.3757
2012-10-30 01:55 | 25.8562 | 347.2272 | 272.5559 | 341.657 | 242.7559 | 92.5185 | 1.6501 | 42.1222 | 69.507 | 25.5267 | 2.6196 | 41.9967 | 110.0151 | 16.0317 | 21.3605 | 6.3782 | 179.5221 | 325.8667 | 251.1954 | 320.2965 | 221.3954
2012-10-30 01:55 | 25.8547 | 347.2524 | 272.535 | 341.5832 | 243.0562 | 92.5174 | 1.6503 | 42.1274 | 69.5235 | 25.5269 | 2.6199 | 42.0019 | 110.0416 | 16.0317 | 21.3349 | 6.3781 | 179.5651 | 325.9175 | 251.2001 | 320.2483 | 221.7213
2012-10-30 01:55 | 25.8555 | 347.2379 | 272.5591 | 341.5649 | 242.8885 | 92.5261 | 1.6506 | 42.1346 | 69.5471 | 25.527 | 2.6203 | 42.0089 | 110.0777 | 16.0319 | 21.3758 | 6.3787 | 179.6248 | 325.8621 | 251.1833 | 320.1891 | 221.5127
2012-10-30 01:55 | 25.8495 | 347.2395 | 272.5143 | 341.5856 | 242.7805 | 92.5289 | 1.6508 | 42.1399 | 69.5646 | 25.5269 | 2.6207 | 42.0142 | 110.1057 | 16.0318 | 21.3578 | 6.3789 | 179.6704 | 325.8817 | 251.1565 | 320.2278 | 221.4227
2012-10-30 01:55 | 25.8487 | 347.2121 | 272.4786 | 341.5581 | 242.5681 | 92.5318 | 1.6522 | 42.1755 | 69.6812 | 25.5273 | 2.6228 | 42.0489 | 110.2858 | 16.0321 | 21.3409 | 6.3791 | 179.967 | 325.8712 | 251.1377 | 320.2172 | 221.2272
2012-10-30 01:55 | 25.851 | 347.222 | 272.4266 | 341.5566 | 242.8134 | 92.533 | 1.6541 | 42.2271 | 69.849 | 25.5283 | 2.626 | 42.1013 | 110.5579 | 16.0325 | 21.3272 | 6.3792 | 180.407 | 325.8948 | 251.0994 | 320.2294 | 221.4862
2012-10-30 01:55 | 25.851 | 347.2296 | 272.3723 | 341.5376 | 243.2171 | 92.5266 | 1.6546 | 42.2401 | 69.8923 | 25.5282 | 2.6268 | 42.1144 | 110.6259 | 16.0326 | 21.3513 | 6.3788 | 180.5183 | 325.8783 | 251.021 | 320.1863 | 221.8658
2012-10-30 01:55 | 25.8495 | 347.2434 | 272.4483 | 341.5171 | 242.8158 | 92.5376 | 1.6551 | 42.2511 | 69.9284 | 25.5283 | 2.6275 | 42.1252 | 110.6819 | 16.0327 | 21.3457 | 6.3795 | 180.6102 | 325.8977 | 251.1026 | 320.1714 | 221.4701
2012-10-30 01:55 | 25.8442 | 347.2647 | 272.3771 | 341.5384 | 242.8458 | 92.5388 | 1.6553 | 42.2579 | 69.9497 | 25.5288 | 2.6278 | 42.1319 | 110.716 | 16.0329 | 21.3485 | 6.3796 | 180.6657 | 325.9162 | 251.0286 | 320.1899 | 221.4973
2012-10-30 01:55 | 25.8517 | 347.2684 | 272.4273 | 341.5535 | 243.2139 | 92.5399 | 1.6554 | 42.262 | 69.9622 | 25.5291 | 2.6281 | 42.1359 | 110.7355 | 16.0331 | 21.4001 | 6.3797 | 180.6977 | 325.8683 | 251.0272 | 320.1534 | 221.8138
2012-10-30 01:56 | 25.848 | 347.3027 | 272.4662 | 341.5384 | 243.3669 | 92.5538 | 1.6555 | 42.2641 | 69.9689 | 25.5292 | 2.6282 | 42.1381 | 110.7462 | 16.0332 | 21.3723 | 6.3807 | 180.715 | 325.9304 | 251.0939 | 320.1661 | 221.9946
2012-10-30 01:56 | 25.8495 | 347.2852 | 272.5531 | 341.5513 | 243.4429 | 92.5405 | 1.6555 | 42.2636 | 69.9671 | 25.5293 | 2.6281 | 42.1375 | 110.7429 | 16.0332 | 21.3457 | 6.3797 | 180.7101 | 325.9395 | 251.2074 | 320.2056 | 222.0972
2012-10-30 01:56 | 25.8435 | 347.2488 | 272.5975 | 341.5111 | 243.4761 | 92.559 | 1.6555 | 42.2647 | 69.9709 | 25.5293 | 2.6282 | 42.1388 | 110.749 | 16.0333 | 21.3237 | 6.381 | 180.7199 | 325.9251 | 251.2738 | 320.1874 | 222.1524
2012-10-30 01:56 | 25.8472 | 347.2715 | 272.6206 | 341.549 | 243.5464 | 92.5382 | 1.6555 | 42.2646 | 69.9704 | 25.5293 | 2.6282 | 42.1387 | 110.7483 | 16.0333 | 21.3153 | 6.3796 | 180.7187 | 325.9562 | 251.3053 | 320.2337 | 222.2311
2012-10-30 01:56 | 25.854 | 347.3163 | 272.7321 | 341.514 | 243.5571 | 92.5509 | 1.6556 | 42.2651 | 69.9722 | 25.5293 | 2.6282 | 42.139 | 110.7508 | 16.0333 | 21.3302 | 6.3805 | 180.723 | 325.9861 | 251.4019 | 320.1838 | 222.2269
2012-10-30 01:56 | 25.8457 | 347.2814 | 272.7354 | 341.5589 | 243.6194 | 92.5428 | 1.6555 | 42.2641 | 69.9688 | 25.5293 | 2.6281 | 42.138 | 110.745 | 16.0333 | 21.3219 | 6.3799 | 180.7137 | 325.9595 | 251.4135 | 320.237 | 222.2975
2012-10-30 01:56 | 25.854 | 347.2479 | 272.767 | 341.5444 | 243.5493 | 92.544 | 1.6555 | 42.2627 | 69.9643 | 25.5293 | 2.6281 | 42.1368 | 110.7385 | 16.0333 | 21.294 | 6.38 | 180.7028 | 325.9539 | 251.473 | 320.2504 | 222.2553
2012-10-30 01:56 | 25.8532 | 347.3041 | 272.8011 | 341.574 | 243.5289 | 92.5619 | 1.6554 | 42.261 | 69.9581 | 25.5294 | 2.628 | 42.1351 | 110.7291 | 16.0334 | 21.2973 | 6.3812 | 180.6871 | 326.0068 | 251.5038 | 320.2767 | 222.2316
2012-10-30 01:56 | 25.8562 | 347.2957 | 272.8351 | 341.5618 | 243.5436 | 92.5694 | 1.6553 | 42.2593 | 69.9527 | 25.5294 | 2.6279 | 42.1334 | 110.7215 | 16.0332 | 21.3243 | 6.3817 | 180.6742 | 325.9714 | 251.5108 | 320.2375 | 222.2193
2012-10-30 01:56 | 25.86 | 347.2995 | 272.8932 | 341.5732 | 243.6258 | 92.5561 | 1.6553 | 42.2589 | 69.9516 | 25.5292 | 2.6279 | 42.1328 | 110.7189 | 16.0332 | 21.3161 | 6.3808 | 180.6705 | 325.9834 | 251.5771 | 320.2571 | 222.3097
2012-10-30 01:56 | 25.8577 | 347.32 | 272.9142 | 341.5443 | 243.5491 | 92.559 | 1.6552 | 42.2572 | 69.9461 | 25.5293 | 2.6277 | 42.1311 | 110.7095 | 16.0332 | 21.3299 | 6.381 | 180.6555 | 325.9901 | 251.5843 | 320.2144 | 222.2192
2012-10-30 01:56 | 25.857 | 347.2851 | 272.8824 | 341.6007 | 243.5405 | 92.5607 | 1.6552 | 42.2565 | 69.9442 | 25.5291 | 2.6277 | 42.1306 | 110.7067 | 16.0333 | 21.3372 | 6.3811 | 180.6509 | 325.9479 | 251.5452 | 320.2635 | 222.2033
2012-10-30 01:56 | 25.8577 | 347.3391 | 272.9103 | 341.5747 | 243.7215 | 92.5642 | 1.6551 | 42.2532 | 69.9327 | 25.5293 | 2.6275 | 42.1273 | 110.6895 | 16.0332 | 21.3379 | 6.3814 | 180.6222 | 326.0012 | 251.5724 | 320.2368 | 222.3836
2012-10-30 01:56 | 25.8615 | 347.3505 | 272.9645 | 341.6242 | 243.8232 | 92.5636 | 1.6551 | 42.2523 | 69.9303 | 25.5291 | 2.6275 | 42.1263 | 110.685 | 16.0331 | 21.3337 | 6.3813 | 180.6153 | 326.0168 | 251.6308 | 320.2905 | 222.4895
2012-10-30 01:56 | 25.8577 | 347.3429 | 272.9646 | 341.5976 | 243.843 | 92.5688 | 1.655 | 42.2505 | 69.9238 | 25.5293 | 2.6274 | 42.1246 | 110.6771 | 16.033 | 21.3419 | 6.3817 | 180.6009 | 326.001 | 251.6227 | 320.2557 | 222.5011
2012-10-30 01:56 | 25.86 | 347.3642 | 273.0289 | 341.6037 | 243.8492 | 92.5688 | 1.6549 | 42.249 | 69.9185 | 25.5294 | 2.6273 | 42.1231 | 110.6681 | 16.0332 | 21.3562 | 6.3817 | 180.5866 | 326.008 | 251.6727 | 320.2475 | 222.493
2012-10-30 01:56 | 25.8659 | 347.3739 | 273.062 | 341.6438 | 243.714 | 92.5769 | 1.6549 | 42.2478 | 69.9142 | 25.5295 | 2.6272 | 42.1218 | 110.6608 | 16.0332 | 21.3581 | 6.3823 | 180.575 | 326.0158 | 251.7039 | 320.2857 | 222.3559
2012-10-30 01:56 | 25.8607 | 347.3801 | 273.0568 | 341.612 | 243.5795 | 92.555 | 1.6548 | 42.2463 | 69.9095 | 25.5295 | 2.6271 | 42.1203 | 110.6532 | 16.0332 | 21.3088 | 6.3807 | 180.5627 | 326.0713 | 251.748 | 320.3032 | 222.2707
2012-10-30 01:56 | 25.8734 | 347.3624 | 273.0307 | 341.6095 | 243.6471 | 92.5688 | 1.6547 | 42.2438 | 69.9006 | 25.5296 | 2.6269 | 42.1178 | 110.6397 | 16.0332 | 21.3215 | 6.3817 | 180.5403 | 326.0409 | 251.7092 | 320.288 | 222.3256
2012-10-30 01:56 | 25.8689 | 347.3541 | 273.0456 | 341.662 | 243.6269 | 92.5607 | 1.6547 | 42.2427 | 69.897 | 25.5297 | 2.6268 | 42.1166 | 110.6334 | 16.0332 | 21.329 | 6.3811 | 180.5304 | 326.0251 | 251.7166 | 320.333 | 222.2979
2012-10-30 01:57 | 25.8704 | 347.405 | 273.0587 | 341.675 | 243.5735 | 92.5619 | 1.6544 | 42.2373 | 69.8795 | 25.5295 | 2.6265 | 42.1114 | 110.6068 | 16.0331 | 21.3345 | 6.3812 | 180.4863 | 326.0705 | 251.7242 | 320.3405 | 222.239
2012-10-30 01:57 | 25.8667 | 347.4203 | 273.0473 | 341.6751 | 243.8324 | 92.5729 | 1.6532 | 42.2048 | 69.7728 | 25.5292 | 2.6247 | 42.0818 | 110.4537 | 16.0328 | 21.3348 | 6.382 | 180.2265 | 326.0855 | 251.7125 | 320.3403 | 222.4976
2012-10-30 01:57 | 25.8742 | 347.4393 | 273.1168 | 341.7092 | 243.9182 | 92.5654 | 1.6504 | 42.1322 | 69.5367 | 25.5279 | 2.6201 | 42.0063 | 110.0615 | 16.0322 | 21.3383 | 6.3815 | 179.5983 | 326.101 | 251.7785 | 320.3709 | 222.5799
2012-10-30 01:57 | 25.8727 | 347.4035 | 273.092 | 341.7001 | 243.5954 | 92.5671 | 1.6501 | 42.1243 | 69.5102 | 25.5281 | 2.6197 | 41.9986 | 110.0226 | 16.032 | 21.3208 | 6.3816 | 179.5328 | 326.0827 | 251.7712 | 320.3793 | 222.2746
2012-10-30 01:57 | 25.863 | 347.409 | 272.9893 | 341.7246 | 242.8097 | 92.559 | 1.6499 | 42.1184 | 69.492 | 25.5276 | 2.6193 | 41.9929 | 109.9933 | 16.0319 | 21.3713 | 6.381 | 179.4853 | 326.0377 | 251.618 | 320.3533 | 221.4384
2012-10-30 01:57 | 25.8697 | 347.4043 | 272.8602 | 341.7428 | 243.0516 | 92.5451 | 1.65 | 42.1193 | 69.4951 | 25.5275 | 2.6194 | 41.9938 | 109.9986 | 16.0318 | 21.3619 | 6.3801 | 179.4938 | 326.0424 | 251.4983 | 320.3809 | 221.6897
2012-10-30 01:57 | 25.8644 | 347.4142 | 272.7968 | 341.7337 | 243.8693 | 92.5451 | 1.6501 | 42.1229 | 69.5066 | 25.5277 | 2.6196 | 41.9972 | 110.017 | 16.0318 | 21.3486 | 6.3801 | 179.5235 | 326.0656 | 251.4482 | 320.3851 | 222.5207
2012-10-30 01:57 | 25.8525 | 347.3833 | 272.8663 | 341.7065 | 243.5713 | 92.5521 | 1.6503 | 42.128 | 69.5234 | 25.5276 | 2.6199 | 42.0023 | 110.0439 | 16.0317 | 21.3728 | 6.3805 | 179.5673 | 326.0105 | 251.4935 | 320.3337 | 222.1985
2012-10-30 01:57 | 25.854 | 347.3734 | 272.8562 | 341.7042 | 243.3769 | 92.5486 | 1.6505 | 42.1341 | 69.5431 | 25.5278 | 2.6203 | 42.0086 | 110.0759 | 16.0319 | 21.3743 | 6.3803 | 179.619 | 325.9991 | 251.4819 | 320.3299 | 222.0026
2012-10-30 01:57 | 25.848 | 347.4207 | 272.8192 | 341.7096 | 242.5635 | 92.5301 | 1.6507 | 42.1401 | 69.5621 | 25.5281 | 2.6207 | 42.0145 | 110.1071 | 16.0319 | 21.3884 | 6.379 | 179.6691 | 326.0323 | 251.4308 | 320.3212 | 221.1751
2012-10-30 01:57 | 25.8345 | 347.3425 | 272.7048 | 341.6581 | 241.8757 | 92.5359 | 1.6509 | 42.1439 | 69.575 | 25.5279 | 2.6209 | 42.018 | 110.1248 | 16.0319 | 21.3789 | 6.3794 | 179.6998 | 325.9636 | 251.3259 | 320.2792 | 220.4968
2012-10-30 01:57 | 25.833 | 347.379 | 272.5133 | 341.6794 | 241.7644 | 92.5451 | 1.6535 | 42.2127 | 69.7999 | 25.5289 | 2.6251 | 42.087 | 110.4836 | 16.0324 | 21.3493 | 6.3801 | 180.2836 | 326.0297 | 251.164 | 320.3301 | 220.4151
2012-10-30 01:57 | 25.8308 | 347.4035 | 272.4219 | 341.6924 | 242.5619 | 92.5561 | 1.6545 | 42.2389 | 69.8854 | 25.5293 | 2.6267 | 42.1131 | 110.6188 | 16.0326 | 21.4153 | 6.3808 | 180.5042 | 325.9882 | 251.0066 | 320.2771 | 221.1466
2012-10-30 01:57 | 25.8181 | 347.425 | 272.4363 | 341.6569 | 243.3135 | 92.5613 | 1.655 | 42.2516 | 69.9272 | 25.5294 | 2.6275 | 42.1258 | 110.6844 | 16.0328 | 21.3705 | 6.3812 | 180.6116 | 326.0545 | 251.0658 | 320.2864 | 221.943
2012-10-30 01:57 | 25.8098 | 347.4205 | 272.4862 | 341.6562 | 243.3562 | 92.5492 | 1.6553 | 42.2602 | 69.9544 | 25.5299 | 2.628 | 42.1342 | 110.7277 | 16.033 | 21.3662 | 6.3803 | 180.682 | 326.0543 | 251.12 | 320.29 | 221.99
2012-10-30 01:57 | 25.8076 | 347.4335 | 272.5616 | 341.6312 | 243.5146 | 92.5665 | 1.6555 | 42.264 | 69.9671 | 25.5297 | 2.6282 | 42.1377 | 110.7461 | 16.033 | 21.364 | 6.3815 | 180.7132 | 326.0695 | 251.1976 | 320.2672 | 222.1506
2012-10-30 01:57 | 25.7993 | 347.4138 | 272.6347 | 341.6952 | 243.5612 | 92.5769 | 1.6556 | 42.267 | 69.9767 | 25.5299 | 2.6284 | 42.1409 | 110.7616 | 16.0331 | 21.4159 | 6.3823 | 180.7382 | 325.9979 | 251.2188 | 320.2793 | 222.1453
2012-10-30 01:57 | 25.8001 | 347.4641 | 272.7208 | 341.6503 | 243.4209 | 92.5821 | 1.6556 | 42.2683 | 69.981 | 25.5299 | 2.6284 | 42.1422 | 110.7684 | 16.0331 | 21.3926 | 6.3826 | 180.7494 | 326.0715 | 251.3282 | 320.2577 | 222.0283
2012-10-30 01:57 | 25.7919 | 347.4292 | 272.7204 | 341.6383 | 243.6126 | 92.5734 | 1.6557 | 42.2689 | 69.9832 | 25.5299 | 2.6285 | 42.143 | 110.7727 | 16.0331 | 21.3764 | 6.382 | 180.7559 | 326.0528 | 251.344 | 320.2619 | 222.2362
2012-10-30 01:57 | 25.8053 | 347.4274 | 272.8036 | 341.6746 | 243.5398 | 92.5775 | 1.6556 | 42.2677 | 69.9787 | 25.5301 | 2.6284 | 42.1417 | 110.7658 | 16.0331 | 21.3978 | 6.3823 | 180.7445 | 326.0296 | 251.4058 | 320.2768 | 222.142
2012-10-30 01:57 | 25.7934 | 347.4536 | 272.8499 | 341.6627 | 243.716 | 92.5786 | 1.6555 | 42.2664 | 69.9741 | 25.5302 | 2.6283 | 42.1403 | 110.7581 | 16.0332 | 21.3699 | 6.3824 | 180.7322 | 326.0837 | 251.48 | 320.2928 | 222.3461
2012-10-30 01:58 | 25.7934 | 347.4498 | 272.9042 | 341.6779 | 243.6062 | 92.5671 | 1.6555 | 42.2655 | 69.9715 | 25.53 | 2.6283 | 42.1395 | 110.7533 | 16.0332 | 21.3498 | 6.3816 | 180.7249 | 326.1 | 251.5544 | 320.3281 | 222.2564
2012-10-30 01:58 | 25.7926 | 347.468 | 272.915 | 341.6923 | 243.6211 | 92.5867 | 1.6555 | 42.2638 | 69.9657 | 25.5301 | 2.6282 | 42.1378 | 110.7454 | 16.0331 | 21.3611 | 6.3829 | 180.7112 | 326.1069 | 251.5539 | 320.3312 | 222.26
2012-10-30 01:58 | 25.7911 | 347.4589 | 272.9251 | 341.687 | 243.3453 | 92.5919 | 1.6554 | 42.2628 | 69.9618 | 25.5302 | 2.6281 | 42.1368 | 110.7402 | 16.0331 | 21.3435 | 6.3833 | 180.702 | 326.1154 | 251.5816 | 320.3435 | 222.0018
2012-10-30 01:58 | 25.7934 | 347.4536 | 272.8964 | 341.7159 | 243.6689 | 92.5821 | 1.6553 | 42.2608 | 69.9552 | 25.5302 | 2.628 | 42.1348 | 110.7296 | 16.0331 | 21.3418 | 6.3826 | 180.6848 | 326.1118 | 251.5546 | 320.3741 | 222.3271
2012-10-30 01:58 | 25.7844 | 347.4256 | 272.9805 | 341.7069 | 243.5894 | 92.5758 | 1.6553 | 42.2596 | 69.9514 | 25.5302 | 2.6279 | 42.1336 | 110.724 | 16.033 | 21.3448 | 6.3822 | 180.6754 | 326.0808 | 251.6357 | 320.3621 | 222.2446
2012-10-30 01:58 | 25.7971 | 347.4573 | 272.9466 | 341.6816 | 243.5904 | 92.585 | 1.6552 | 42.2577 | 69.9458 | 25.5299 | 2.6278 | 42.1316 | 110.7139 | 16.033 | 21.3495 | 6.3828 | 180.6598 | 326.1078 | 251.5971 | 320.3321 | 222.2409
2012-10-30 01:58 | 25.7956 | 347.4786 | 272.9296 | 341.6953 | 243.5653 | 92.5919 | 1.6551 | 42.2555 | 69.938 | 25.5301 | 2.6277 | 42.1294 | 110.702 | 16.033 | 21.3119 | 6.3833 | 180.6401 | 326.1667 | 251.6177 | 320.3834 | 222.2534
2012-10-30 01:58 | 25.7971 | 347.4649 | 272.9777 | 341.7158 | 243.7275 | 92.5919 | 1.6551 | 42.2543 | 69.9342 | 25.5301 | 2.6276 | 42.1281 | 110.6953 | 16.033 | 21.3455 | 6.3833 | 180.6295 | 326.1194 | 251.6322 | 320.3703 | 222.382
2012-10-30 01:58 | 25.7836 | 347.4248 | 272.9913 | 341.7137 | 243.424 | 92.5873 | 1.655 | 42.2523 | 69.9277 | 25.53 | 2.6275 | 42.1262 | 110.686 | 16.0329 | 21.2879 | 6.383 | 180.6138 | 326.1369 | 251.7034 | 320.4258 | 222.1361
2012-10-30 01:58 | 25.7926 | 347.4756 | 272.9576 | 341.7342 | 243.4683 | 92.5833 | 1.655 | 42.251 | 69.9235 | 25.53 | 2.6274 | 42.1247 | 110.6785 | 16.0329 | 21.2848 | 6.3827 | 180.602 | 326.1908 | 251.6728 | 320.4494 | 222.1835
2012-10-30 01:58 | 25.7881 | 347.5053 | 272.9531 | 341.7106 | 243.6205 | 92.5798 | 1.6549 | 42.2491 | 69.9167 | 25.5302 | 2.6273 | 42.1232 | 110.6708 | 16.0328 | 21.2843 | 6.3825 | 180.5875 | 326.221 | 251.6688 | 320.4263 | 222.3362
2012-10-30 01:58 | 25.7926 | 347.468 | 272.9925 | 341.757 | 243.6446 | 92.5856 | 1.6548 | 42.2478 | 69.9122 | 25.5303 | 2.6272 | 42.1217 | 110.6631 | 16.0328 | 21.2768 | 6.3829 | 180.5753 | 326.1912 | 251.7157 | 320.4802 | 222.3678
2012-10-30 01:58 | 25.7993 | 347.4671 | 272.9992 | 341.7409 | 243.6161 | 92.585 | 1.6548 | 42.2463 | 69.9086 | 25.5298 | 2.6271 | 42.1203 | 110.6553 | 16.0328 | 21.2594 | 6.3828 | 180.5639 | 326.2077 | 251.7398 | 320.4815 | 222.3567
2012-10-30 01:58 | 25.7949 | 347.4931 | 272.991 | 341.7288 | 243.6077 | 92.5867 | 1.6547 | 42.2443 | 69.9015 | 25.53 | 2.627 | 42.1183 | 110.6449 | 16.0328 | 21.2671 | 6.3829 | 180.5465 | 326.226 | 251.7239 | 320.4617 | 222.3406
2012-10-30 01:58 | 25.7964 | 347.506 | 273.0545 | 341.776 | 243.6171 | 92.5706 | 1.6547 | 42.243 | 69.8978 | 25.5298 | 2.6269 | 42.117 | 110.6386 | 16.0328 | 21.2124 | 6.3818 | 180.5364 | 326.2936 | 251.8421 | 320.5636 | 222.4047
2012-10-30 01:58 | 25.7986 | 347.5006 | 273.0334 | 341.763 | 243.6271 | 92.5867 | 1.6546 | 42.2415 | 69.8921 | 25.53 | 2.6268 | 42.1155 | 110.6304 | 16.0328 | 21.2788 | 6.3829 | 180.5225 | 326.2218 | 251.7546 | 320.4842 | 222.3483
2012-10-30 01:58 | 25.8008 | 347.5485 | 273.0318 | 341.8146 | 243.7665 | 92.5931 | 1.6545 | 42.2403 | 69.8884 | 25.5299 | 2.6268 | 42.1143 | 110.6243 | 16.0328 | 21.2529 | 6.3834 | 180.5127 | 326.2956 | 251.7789 | 320.5617 | 222.5136
2012-10-30 01:58 | 25.7971 | 347.5143 | 273.0707 | 341.7957 | 243.7353 | 92.589 | 1.6545 | 42.2379 | 69.8809 | 25.5298 | 2.6266 | 42.112 | 110.6128 | 16.0327 | 21.2693 | 6.3831 | 180.4937 | 326.245 | 251.8014 | 320.5264 | 222.466
2012-10-30 01:58 | 25.8061 | 347.5195 | 273.0991 | 341.8352 | 243.7443 | 92.6029 | 1.6544 | 42.2366 | 69.8759 | 25.5299 | 2.6265 | 42.1106 | 110.6053 | 16.0327 | 21.3023 | 6.384 | 180.4811 | 326.2172 | 251.7968 | 320.5329 | 222.442
2012-10-30 01:58 | 25.7919 | 347.5053 | 273.1392 | 341.7829 | 243.8477 | 92.5971 | 1.6535 | 42.2126 | 69.797 | 25.5298 | 2.6252 | 42.0886 | 110.4908 | 16.0325 | 21.3243 | 6.3836 | 180.2878 | 326.181 | 251.8149 | 320.4586 | 222.5234
2012-10-30 01:58 | 25.7964 | 347.5669 | 273.1437 | 341.7722 | 243.8796 | 92.5948 | 1.6505 | 42.1354 | 69.5457 | 25.5284 | 2.6204 | 42.0093 | 110.079 | 16.0319 | 21.2846 | 6.3835 | 179.6247 | 326.2823 | 251.8591 | 320.4876 | 222.595
2012-10-30 01:59 | 25.7979 | 347.5342 | 273.1801 | 341.8079 | 243.7361 | 92.5786 | 1.6501 | 42.1253 | 69.5121 | 25.5285 | 2.6198 | 41.9995 | 110.0298 | 16.0317 | 21.3182 | 6.3824 | 179.542 | 326.216 | 251.8619 | 320.4897 | 222.4179
2012-10-30 01:59 | 25.7904 | 347.5305 | 273.1105 | 341.8157 | 242.8704 | 92.5769 | 1.6499 | 42.1199 | 69.495 | 25.5282 | 2.6195 | 41.9942 | 110.0029 | 16.0315 | 21.2746 | 6.3823 | 179.4979 | 326.2559 | 251.8359 | 320.5411 | 221.5958
2012-10-30 01:59 | 25.7896 | 347.5639 | 273.0438 | 341.7996 | 242.9166 | 92.5729 | 1.65 | 42.1201 | 69.4961 | 25.5281 | 2.6195 | 41.9946 | 110.0052 | 16.0315 | 21.2658 | 6.382 | 179.5013 | 326.2981 | 251.778 | 320.5338 | 221.6508
2012-10-30 01:59 | 25.7896 | 347.5183 | 272.9391 | 341.7844 | 243.1714 | 92.5746 | 1.65 | 42.1223 | 69.5029 | 25.5282 | 2.6196 | 41.9967 | 110.0157 | 16.0315 | 21.2778 | 6.3821 | 179.5186 | 326.2405 | 251.6613 | 320.5066 | 221.8936
2012-10-30 01:59 | 25.7889 | 347.4681 | 272.919 | 341.7913 | 243.1119 | 92.5804 | 1.6503 | 42.1284 | 69.5232 | 25.5282 | 2.62 | 42.0027 | 110.0474 | 16.0315 | 21.2972 | 6.3825 | 179.5707 | 326.1709 | 251.6218 | 320.4941 | 221.8147
2012-10-30 01:59 | 25.7836 | 347.5237 | 272.8323 | 341.7936 | 242.793 | 92.5711 | 1.6505 | 42.1341 | 69.5427 | 25.528 | 2.6204 | 42.0085 | 110.0772 | 16.0316 | 21.2839 | 6.3819 | 179.6199 | 326.2398 | 251.5484 | 320.5097 | 221.5091
2012-10-30 01:59 | 25.7806 | 347.5245 | 272.8681 | 341.783 | 243.2603 | 92.5659 | 1.6506 | 42.1371 | 69.5513 | 25.5284 | 2.6205 | 42.0113 | 110.0921 | 16.0316 | 21.2849 | 6.3815 | 179.6434 | 326.2396 | 251.5832 | 320.4981 | 221.9754
2012-10-30 01:59 | 25.7814 | 347.5215 | 272.8495 | 341.7382 | 243.6569 | 92.5561 | 1.6508 | 42.1428 | 69.5696 | 25.5286 | 2.6209 | 42.0172 | 110.122 | 16.0317 | 21.2977 | 6.3808 | 179.6916 | 326.2238 | 251.5518 | 320.4405 | 222.3592
2012-10-30 01:59 | 25.7694 | 347.5019 | 272.9461 | 341.7528 | 243.6998 | 92.5584 | 1.6526 | 42.1901 | 69.7244 | 25.5291 | 2.6239 | 42.0665 | 110.3783 | 16.0321 | 21.2817 | 6.381 | 180.1027 | 326.2202 | 251.6644 | 320.4711 | 222.4181
2012-10-30 01:59 | 25.7724 | 347.4973 | 273.0576 | 341.7786 | 244.0985 | 92.574 | 1.6543 | 42.2342 | 69.8686 | 25.5298 | 2.6264 | 42.1084 | 110.5949 | 16.0325 | 21.3328 | 6.3821 | 180.4634 | 326.1645 | 251.7248 | 320.4458 | 222.7657
2012-10-30 01:59 | 25.7679 | 347.5232 | 273.0842 | 341.7703 | 243.7179 | 92.5775 | 1.6548 | 42.2466 | 69.9086 | 25.5301 | 2.6272 | 42.1207 | 110.6587 | 16.0327 | 21.3003 | 6.3823 | 180.5673 | 326.2229 | 251.7839 | 320.47 | 222.4176
2012-10-30 01:59 | 25.7597 | 347.515 | 273.0992 | 341.7621 | 243.3139 | 92.5937 | 1.6551 | 42.2557 | 69.9379 | 25.5305 | 2.6278 | 42.1301 | 110.7075 | 16.0327 | 21.3041 | 6.3834 | 180.6455 | 326.2109 | 251.7951 | 320.458 | 222.0098
2012-10-30 01:59 | 25.7627 | 347.5332 | 273.0518 | 341.7537 | 243.0387 | 92.5867 | 1.6554 | 42.2622 | 69.9594 | 25.5304 | 2.6281 | 42.1362 | 110.7393 | 16.0328 | 21.283 | 6.3829 | 180.6987 | 326.2502 | 251.7688 | 320.4707 | 221.7557
2012-10-30 01:59 | 25.7619 | 347.5552 | 272.9657 | 341.7491 | 243.5199 | 92.5821 | 1.6555 | 42.2661 | 69.9719 | 25.5305 | 2.6283 | 42.1399 | 110.757 | 16.033 | 21.3063 | 6.3826 | 180.7289 | 326.2489 | 251.6594 | 320.4428 | 222.2136
2012-10-30 01:59 | 25.7567 | 347.5614 | 272.9799 | 341.7059 | 243.7263 | 92.5867 | 1.6555 | 42.2651 | 69.9686 | 25.5306 | 2.6283 | 42.1391 | 110.7525 | 16.0331 | 21.2931 | 6.3829 | 180.7211 | 326.2683 | 251.6868 | 320.4128 | 222.4332
2012-10-30 01:59 | 25.7559 | 347.534 | 273.0528 | 341.7431 | 243.8744 | 92.5902 | 1.6555 | 42.2653 | 69.9691 | 25.5307 | 2.6282 | 42.1391 | 110.7522 | 16.0332 | 21.2923 | 6.3832 | 180.7212 | 326.2417 | 251.7605 | 320.4508 | 222.5821
2012-10-30 01:59 | 25.7582 | 347.5629 | 273.0706 | 341.7568 | 243.7239 | 92.5931 | 1.6554 | 42.2647 | 69.9665 | 25.5309 | 2.6282 | 42.1386 | 110.7495 | 16.0331 | 21.2745 | 6.3834 | 180.716 | 326.2884 | 251.7961 | 320.4823 | 222.4494
2012-10-30 01:59 | 25.7612 | 347.5583 | 273.1201 | 341.7522 | 243.8287 | 92.5971 | 1.6554 | 42.2646 | 69.9657 | 25.531 | 2.6282 | 42.1383 | 110.7482 | 16.0331 | 21.3377 | 6.3836 | 180.7139 | 326.2206 | 251.7824 | 320.4145 | 222.491
2012-10-30 01:59 | 25.7604 | 347.5347 | 273.1077 | 341.702 | 243.8593 | 92.5995 | 1.6554 | 42.2632 | 69.9618 | 25.5307 | 2.6281 | 42.137 | 110.7411 | 16.0332 | 21.2807 | 6.3838 | 180.7029 | 326.254 | 251.827 | 320.4213 | 222.5786
2012-10-30 01:59 | 25.7642 | 347.5651 | 273.127 | 341.7552 | 243.8709 | 92.596 | 1.6553 | 42.2623 | 69.9584 | 25.5309 | 2.6281 | 42.1362 | 110.7368 | 16.0331 | 21.2404 | 6.3836 | 180.6953 | 326.3247 | 251.8866 | 320.5148 | 222.6305
2012-10-30 01:59 | 25.7634 | 347.5567 | 273.1727 | 341.7468 | 243.9054 | 92.6047 | 1.6553 | 42.2613 | 69.9551 | 25.5309 | 2.628 | 42.1351 | 110.7311 | 16.0331 | 21.2436 | 6.3842 | 180.6861 | 326.3131 | 251.9291 | 320.5032 | 222.6618
2012-10-30 02:00 | 25.7604 | 347.5613 | 273.193 | 341.7514 | 243.6986 | 92.6104 | 1.6552 | 42.2592 | 69.948 | 25.531 | 2.6279 | 42.1332 | 110.7209 | 16.0332 | 21.2767 | 6.3846 | 180.6688 | 326.2846 | 251.9163 | 320.4747 | 222.4219
2012-10-30 02:00 | 25.7567 | 347.5729 | 273.135 | 341.7439 | 243.2208 | 92.6075 | 1.6552 | 42.258 | 69.9435 | 25.5312 | 2.6278 | 42.132 | 110.7149 | 16.0331 | 21.2409 | 6.3844 | 180.6584 | 326.332 | 251.8941 | 320.503 | 221.9799
2012-10-30 02:00 | 25.7642 | 347.6032 | 273.0921 | 341.7514 | 243.2322 | 92.6047 | 1.6551 | 42.2569 | 69.9405 | 25.5309 | 2.6277 | 42.1308 | 110.7088 | 16.0331 | 21.2966 | 6.3842 | 180.6493 | 326.3066 | 251.7955 | 320.4548 | 221.9356
2012-10-30 02:00 | 25.7619 | 347.5819 | 273.0898 | 341.7605 | 243.6453 | 92.6029 | 1.6551 | 42.2549 | 69.9341 | 25.5309 | 2.6276 | 42.1289 | 110.6985 | 16.0331 | 21.3023 | 6.384 | 180.6326 | 326.2796 | 251.7875 | 320.4582 | 222.343
2012-10-30 02:00 | 25.7612 | 347.585 | 273.1589 | 341.775 | 243.7739 | 92.6099 | 1.655 | 42.2541 | 69.931 | 25.5309 | 2.6276 | 42.128 | 110.6939 | 16.0331 | 21.2775 | 6.3845 | 180.6249 | 326.3075 | 251.8814 | 320.4975 | 222.4964
2012-10-30 02:00 | 25.7634 | 347.6366 | 273.1727 | 341.8039 | 243.7487 | 92.6018 | 1.655 | 42.2529 | 69.9267 | 25.5311 | 2.6275 | 42.1268 | 110.6878 | 16.0331 | 21.3158 | 6.384 | 180.6145 | 326.3208 | 251.8569 | 320.4881 | 222.4329
2012-10-30 02:00 | 25.7642 | 347.5804 | 273.1541 | 341.8237 | 243.6868 | 92.6122 | 1.6548 | 42.2491 | 69.9144 | 25.5311 | 2.6272 | 42.1229 | 110.6671 | 16.0331 | 21.2885 | 6.3847 | 180.5815 | 326.2919 | 251.8656 | 320.5352 | 222.3983
2012-10-30 02:00 | 25.7642 | 347.5918 | 273.1309 | 341.8313 | 243.5143 | 92.6064 | 1.6547 | 42.2467 | 69.9065 | 25.531 | 2.6271 | 42.1206 | 110.6554 | 16.0331 | 21.3126 | 6.3843 | 180.5619 | 326.2792 | 251.8183 | 320.5187 | 222.2017
2012-10-30 02:00 | 25.7702 | 347.5864 | 273.1601 | 341.8183 | 243.4576 | 92.6035 | 1.6547 | 42.2458 | 69.9041 | 25.5308 | 2.6271 | 42.1198 | 110.6516 | 16.033 | 21.3226 | 6.3841 | 180.5557 | 326.2638 | 251.8375 | 320.4957 | 222.135
2012-10-30 02:00 | 25.7589 | 347.5865 | 273.0946 | 341.7804 | 243.317 | 92.5983 | 1.6547 | 42.2448 | 69.9012 | 25.5306 | 2.627 | 42.1189 | 110.6483 | 16.0328 | 21.3113 | 6.3837 | 180.5495 | 326.2752 | 251.7833 | 320.4691 | 222.0057
2012-10-30 02:00 | 25.7627 | 347.5903 | 273.0402 | 341.8298 | 243.4227 | 92.6012 | 1.6546 | 42.2431 | 69.8951 | 25.5308 | 2.6269 | 42.117 | 110.6377 | 16.0329 | 21.3352 | 6.3839 | 180.5328 | 326.2551 | 251.705 | 320.4946 | 222.0875
2012-10-30 02:00 | 25.7627 | 347.6359 | 273.0828 | 341.807 | 243.6382 | 92.6023 | 1.6545 | 42.2408 | 69.8879 | 25.5307 | 2.6268 | 42.1149 | 110.6268 | 16.0329 | 21.291 | 6.384 | 180.5147 | 326.3449 | 251.7918 | 320.516 | 222.3472
2012-10-30 02:00 | 25.7642 | 347.5994 | 273.0727 | 341.8237 | 243.581 | 92.611 | 1.6545 | 42.2402 | 69.8861 | 25.5307 | 2.6267 | 42.1143 | 110.6236 | 16.0329 | 21.3046 | 6.3846 | 180.5097 | 326.2948 | 251.7681 | 320.5191 | 222.2764
2012-10-30 02:00 | 25.7544 | 347.5934 | 273.0513 | 341.8444 | 243.5516 | 92.611 | 1.6544 | 42.2383 | 69.879 | 25.5309 | 2.6266 | 42.1123 | 110.6132 | 16.0329 | 21.2948 | 6.3846 | 180.4922 | 326.2986 | 251.7565 | 320.5496 | 222.2568
2012-10-30 02:00 | 25.7604 | 347.5766 | 273.0379 | 341.8161 | 243.6046 | 92.6081 | 1.6543 | 42.2366 | 69.8735 | 25.5309 | 2.6265 | 42.1107 | 110.6053 | 16.0328 | 21.2607 | 6.3844 | 180.4788 | 326.3159 | 251.7772 | 320.5554 | 222.3439
2012-10-30 02:00 | 25.7642 | 347.6184 | 273.0533 | 341.8313 | 243.5927 | 92.6122 | 1.6543 | 42.2349 | 69.8685 | 25.5307 | 2.6264 | 42.1089 | 110.5968 | 16.0327 | 21.2805 | 6.3847 | 180.4653 | 326.3379 | 251.7728 | 320.5508 | 222.3122
2012-10-30 02:00 | 25.7664 | 347.6092 | 273.0672 | 341.7841 | 243.6341 | 92.6012 | 1.6531 | 42.2038 | 69.7663 | 25.5304 | 2.6243 | 42.0748 | 110.4187 | 16.0325 | 21.2947 | 6.3839 | 180.1849 | 326.3145 | 251.7725 | 320.4894 | 222.3394
2012-10-30 02:00 | 25.7597 | 347.6025 | 273.0488 | 341.8192 | 243.5569 | 92.5977 | 1.6505 | 42.1351 | 69.5422 | 25.5293 | 2.6204 | 42.0094 | 110.0794 | 16.032 | 21.26 | 6.3837 | 179.6216 | 326.3425 | 251.7888 | 320.5592 | 222.2969
2012-10-30 02:00 | 25.7709 | 347.6175 | 273.0174 | 341.8532 | 243.619 | 92.5983 | 1.6501 | 42.1241 | 69.5071 | 25.5289 | 2.6197 | 41.9986 | 110.0235 | 16.0319 | 21.2431 | 6.3837 | 179.5306 | 326.3744 | 251.7743 | 320.6101 | 222.3759
2012-10-30 02:00 | 25.7604 | 347.5461 | 273.0495 | 341.8047 | 243.056 | 92.5879 | 1.6499 | 42.12 | 69.493 | 25.5292 | 2.6195 | 41.9944 | 110.0026 | 16.0317 | 21.2486 | 6.383 | 179.4957 | 326.2975 | 251.8009 | 320.5561 | 221.8074
2012-10-30 02:00 | 25.7612 | 347.5621 | 272.9301 | 341.8169 | 242.312 | 92.5804 | 1.6499 | 42.1201 | 69.4937 | 25.529 | 2.6195 | 41.9946 | 110.0046 | 16.0316 | 21.2574 | 6.3825 | 179.4983 | 326.3047 | 251.6727 | 320.5595 | 221.0546
2012-10-30 02:00 | 25.7612 | 347.5393 | 272.7983 | 341.8321 | 243.0097 | 92.5706 | 1.6499 | 42.1198 | 69.493 | 25.5288 | 2.6195 | 41.9942 | 110.0031 | 16.0315 | 21.2855 | 6.3818 | 179.4962 | 326.2538 | 251.5128 | 320.5466 | 221.7242
2012-10-30 02:01 | 25.7574 | 347.5241 | 272.81 | 341.8169 | 243.3743 | 92.5706 | 1.65 | 42.1231 | 69.5033 | 25.529 | 2.6197 | 41.9973 | 110.02 | 16.0314 | 21.3098 | 6.3818 | 179.5233 | 326.2143 | 251.5002 | 320.5071 | 222.0645
2012-10-30 02:01 | 25.753 | 347.4855 | 272.7901 | 341.8201 | 242.8643 | 92.5862 | 1.6502 | 42.1289 | 69.5226 | 25.529 | 2.62 | 42.0031 | 110.0495 | 16.0315 | 21.3175 | 6.3829 | 179.572 | 326.168 | 251.4726 | 320.5026 | 221.5468
2012-10-30 02:01 | 25.7425 | 347.4864 | 272.6826 | 341.7754 | 242.8499 | 92.5619 | 1.6505 | 42.1361 | 69.5454 | 25.5294 | 2.6205 | 42.0104 | 110.0868 | 16.0316 | 21.307 | 6.3812 | 179.6321 | 326.1794 | 251.3756 | 320.4684 | 221.5429
2012-10-30 02:01 | 25.7402 | 347.5221 | 272.5524 | 341.754 | 242.977 | 92.5671 | 1.6514 | 42.1592 | 69.6216 | 25.5294 | 2.6218 | 42.0331 | 110.2043 | 16.0319 | 21.3127 | 6.3816 | 179.8259 | 326.2094 | 251.2397 | 320.4413 | 221.6643
2012-10-30 02:01 | 25.7372 | 347.5153 | 272.5494 | 341.7701 | 242.4017 | 92.5723 | 1.654 | 42.2265 | 69.8407 | 25.5306 | 2.626 | 42.1008 | 110.5559 | 16.0324 | 21.3177 | 6.3819 | 180.3967 | 326.1976 | 251.2317 | 320.4524 | 221.084
2012-10-30 02:01 | 25.7298 | 347.5422 | 272.476 | 341.7436 | 242.3473 | 92.5648 | 1.6544 | 42.2389 | 69.8811 | 25.5308 | 2.6267 | 42.1126 | 110.6168 | 16.0326 | 21.2983 | 6.3814 | 180.4979 | 326.2439 | 251.1777 | 320.4453 | 221.049
2012-10-30 02:01 | 25.7298 | 347.5726 | 272.4411 | 341.7627 | 242.1473 | 92.5804 | 1.6548 | 42.2495 | 69.9165 | 25.5308 | 2.6273 | 42.1233 | 110.6725 | 16.0326 | 21.2782 | 6.3825 | 180.589 | 326.2944 | 251.1629 | 320.4845 | 220.8691
2012-10-30 02:01 | 25.7305 | 347.5467 | 272.403 | 341.7558 | 242.3793 | 92.5815 | 1.6552 | 42.2586 | 69.9455 | 25.5312 | 2.6279 | 42.1328 | 110.7209 | 16.0329 | 21.323 | 6.3826 | 180.6664 | 326.2237 | 251.08 | 320.4328 | 221.0563
2012-10-30 02:01 | 25.7283 | 347.5559 | 272.428 | 341.7269 | 243.0513 | 92.5752 | 1.6554 | 42.2638 | 69.9626 | 25.5312 | 2.6282 | 42.1378 | 110.746 | 16.033 | 21.3168 | 6.3821 | 180.7086 | 326.2391 | 251.1112 | 320.4101 | 221.7345
2012-10-30 02:01 | 25.711 | 347.5272 | 272.5193 | 341.7058 | 242.9007 | 92.5798 | 1.6555 | 42.266 | 69.97 | 25.5312 | 2.6283 | 42.1398 | 110.7557 | 16.0332 | 21.3116 | 6.3825 | 180.7257 | 326.2156 | 251.2077 | 320.3942 | 221.5891
2012-10-30 02:01 | 25.7155 | 347.5431 | 272.5199 | 341.6761 | 243.0346 | 92.5821 | 1.6555 | 42.2664 | 69.9701 | 25.5316 | 2.6283 | 42.1402 | 110.7582 | 16.0331 | 21.304 | 6.3826 | 180.7283 | 326.2391 | 251.2159 | 320.3721 | 221.7306
2012-10-30 02:01 | 25.7245 | 347.514 | 272.5832 | 341.7307 | 243.0788 | 92.5885 | 1.6554 | 42.2657 | 69.9678 | 25.5316 | 2.6283 | 42.1395 | 110.7542 | 16.0332 | 21.3251 | 6.3831 | 180.7219 | 326.1889 | 251.2581 | 320.4056 | 221.7537
2012-10-30 02:01 | 25.711 | 347.4967 | 272.5581 | 341.7553 | 243.0693 | 92.5885 | 1.6554 | 42.2658 | 69.9687 | 25.5314 | 2.6283 | 42.1397 | 110.7555 | 16.0331 | 21.2955 | 6.3831 | 180.7242 | 326.2012 | 251.2626 | 320.4598 | 221.7738
2012-10-30 02:01 | 25.7178 | 347.5302 | 272.623 | 341.705 | 243.174 | 92.5792 | 1.6554 | 42.2658 | 69.9682 | 25.5316 | 2.6283 | 42.14 | 110.7567 | 16.0332 | 21.3264 | 6.3824 | 180.7249 | 326.2038 | 251.2966 | 320.3786 | 221.8476
2012-10-30 02:01 | 25.714 | 347.4769 | 272.658 | 341.6974 | 243.0174 | 92.5821 | 1.6554 | 42.2653 | 69.9662 | 25.5316 | 2.6282 | 42.1392 | 110.7519 | 16.0333 | 21.3427 | 6.3826 | 180.7181 | 326.1342 | 251.3153 | 320.3547 | 221.6747
2012-10-30 02:01 | 25.7118 | 347.5165 | 272.6403 | 341.6876 | 242.8035 | 92.596 | 1.6554 | 42.2641 | 69.9624 | 25.5317 | 2.6282 | 42.138 | 110.7456 | 16.0332 | 21.3164 | 6.3836 | 180.708 | 326.2001 | 251.3239 | 320.3712 | 221.4871
2012-10-30 02:01 | 25.714 | 347.4883 | 272.5998 | 341.7202 | 243.0174 | 92.589 | 1.6553 | 42.2615 | 69.9536 | 25.5318 | 2.628 | 42.1356 | 110.7335 | 16.0332 | 21.3025 | 6.3831 | 180.687 | 326.1858 | 251.2973 | 320.4177 | 221.7149
2012-10-30 02:01 | 25.7118 | 347.4899 | 272.6364 | 341.7142 | 243.168 | 92.5775 | 1.6552 | 42.2603 | 69.95 | 25.5316 | 2.6279 | 42.1343 | 110.726 | 16.0332 | 21.2803 | 6.3823 | 180.676 | 326.2096 | 251.3561 | 320.4339 | 221.8877
2012-10-30 02:01 | 25.7125 | 347.483 | 272.6526 | 341.7073 | 243.2275 | 92.5902 | 1.6552 | 42.2598 | 69.9475 | 25.5319 | 2.6279 | 42.1337 | 110.7228 | 16.0332 | 21.3091 | 6.3832 | 180.6703 | 326.1739 | 251.3435 | 320.3982 | 221.9184
2012-10-30 02:01 | 25.7133 | 347.4382 | 272.7426 | 341.7119 | 243.675 | 92.5908 | 1.6551 | 42.2578 | 69.9416 | 25.5315 | 2.6278 | 42.1317 | 110.7129 | 16.0332 | 21.3219 | 6.3832 | 180.6546 | 326.1163 | 251.4207 | 320.39 | 222.3531
2012-10-30 02:01 | 25.7096 | 347.4459 | 272.7816 | 341.6968 | 243.5695 | 92.5885 | 1.6551 | 42.2567 | 69.9376 | 25.5318 | 2.6277 | 42.1306 | 110.7066 | 16.0333 | 21.3021 | 6.3831 | 180.6441 | 326.1438 | 251.4795 | 320.3947 | 222.2674
2012-10-30 02:01 | 25.714 | 347.4541 | 272.8364 | 341.6974 | 243.7894 | 92.5919 | 1.6549 | 42.254 | 69.9281 | 25.5319 | 2.6275 | 42.128 | 110.6931 | 16.0332 | 21.3146 | 6.3833 | 180.6212 | 326.1395 | 251.5218 | 320.3828 | 222.4748
2012-10-30 02:01 | 25.714 | 347.4655 | 272.879 | 341.724 | 243.6287 | 92.5867 | 1.6549 | 42.2516 | 69.9205 | 25.5319 | 2.6274 | 42.1256 | 110.6806 | 16.0332 | 21.3346 | 6.3829 | 180.6011 | 326.1309 | 251.5444 | 320.3894 | 222.2941
2012-10-30 02:02 | 25.7088 | 347.4032 | 272.9049 | 341.7112 | 243.4433 | 92.5919 | 1.6548 | 42.2509 | 69.918 | 25.5319 | 2.6273 | 42.1248 | 110.6763 | 16.0333 | 21.3134 | 6.3833 | 180.5943 | 326.0898 | 251.5915 | 320.3978 | 222.1299
2012-10-30 02:02 | 25.7125 | 347.4221 | 272.8737 | 341.7378 | 243.4509 | 92.5856 | 1.6548 | 42.2498 | 69.9142 | 25.532 | 2.6273 | 42.1238 | 110.6713 | 16.0332 | 21.3131 | 6.3829 | 180.5855 | 326.109 | 251.5606 | 320.4247 | 222.1378
2012-10-30 02:02 | 25.7148 | 347.3902 | 272.8605 | 341.721 | 243.5002 | 92.596 | 1.6547 | 42.2476 | 69.907 | 25.532 | 2.6271 | 42.1218 | 110.6602 | 16.0333 | 21.3154 | 6.3836 | 180.5672 | 326.0748 | 251.5451 | 320.4056 | 222.1848
2012-10-30 02:02 | 25.7327 | 347.4157 | 272.9404 | 341.6933 | 243.7101 | 92.5942 | 1.6546 | 42.2453 | 69.8993 | 25.5319 | 2.627 | 42.1193 | 110.6468 | 16.0333 | 21.3453 | 6.3834 | 180.5461 | 326.0704 | 251.5951 | 320.348 | 222.3648
2012-10-30 02:02 | 25.7223 | 347.4281 | 272.9261 | 341.7171 | 243.6723 | 92.5879 | 1.6545 | 42.2436 | 69.8933 | 25.5321 | 2.6269 | 42.1176 | 110.6375 | 16.0333 | 21.3028 | 6.383 | 180.5309 | 326.1253 | 251.6233 | 320.4143 | 222.3695
2012-10-30 02:02 | 25.732 | 347.3998 | 272.9862 | 341.7497 | 243.4194 | 92.6006 | 1.6545 | 42.2426 | 69.8899 | 25.5321 | 2.6268 | 42.1164 | 110.6325 | 16.0332 | 21.2924 | 6.3839 | 180.5224 | 326.1074 | 251.6938 | 320.4573 | 222.127
2012-10-30 02:02 | 25.732 | 347.4226 | 272.9707 | 341.7344 | 243.4234 | 92.5896 | 1.6544 | 42.2403 | 69.8824 | 25.5321 | 2.6267 | 42.1144 | 110.6232 | 16.033 | 21.3045 | 6.3831 | 180.5056 | 326.1181 | 251.6662 | 320.4299 | 222.1189
2012-10-30 02:02 | 25.7253 | 347.4273 | 272.9834 | 341.7201 | 243.6714 | 92.6 | 1.6544 | 42.2398 | 69.8808 | 25.532 | 2.6267 | 42.1137 | 110.6193 | 16.033 | 21.2657 | 6.3838 | 180.5 | 326.1616 | 251.7177 | 320.4544 | 222.4057
2012-10-30 02:02 | 25.726 | 347.409 | 273.019 | 341.7589 | 243.723 | 92.5931 | 1.6543 | 42.2375 | 69.8736 | 25.5319 | 2.6265 | 42.1115 | 110.6075 | 16.0331 | 21.2503 | 6.3834 | 180.4811 | 326.1587 | 251.7687 | 320.5086 | 222.4727
2012-10-30 02:02 | 25.7335 | 347.3785 | 273.0343 | 341.7473 | 243.6639 | 92.6029 | 1.6542 | 42.2354 | 69.866 | 25.5321 | 2.6264 | 42.1092 | 110.5956 | 16.0331 | 21.302 | 6.384 | 180.4616 | 326.0765 | 251.7323 | 320.4453 | 222.3619
2012-10-30 02:02 | 25.7342 | 347.4096 | 273.0233 | 341.7442 | 243.6215 | 92.6087 | 1.6542 | 42.2341 | 69.8624 | 25.5319 | 2.6263 | 42.1081 | 110.5894 | 16.0331 | 21.2866 | 6.3844 | 180.4519 | 326.123 | 251.7367 | 320.4576 | 222.3349
2012-10-30 02:02 | 25.7372 | 347.4088 | 273.0457 | 341.7434 | 243.5226 | 92.5885 | 1.6541 | 42.2322 | 69.8563 | 25.5319 | 2.6262 | 42.1062 | 110.5804 | 16.033 | 21.3097 | 6.3831 | 180.4366 | 326.0991 | 251.736 | 320.4337 | 222.2129
2012-10-30 02:02 | 25.7477 | 347.4155 | 273.0678 | 341.7463 | 243.7094 | 92.6 | 1.6527 | 42.1975 | 69.741 | 25.532 | 2.624 | 42.0703 | 110.392 | 16.033 | 21.3242 | 6.3838 | 180.133 | 326.0913 | 251.7436 | 320.4221 | 222.3852
2012-10-30 02:02 | 25.744 | 347.4308 | 273.1029 | 341.7693 | 243.933 | 92.5862 | 1.6503 | 42.1324 | 69.5299 | 25.5306 | 2.6202 | 42.007 | 110.0649 | 16.0322 | 21.2804 | 6.3829 | 179.5948 | 326.1504 | 251.8225 | 320.4889 | 222.6526
2012-10-30 02:02 | 25.7417 | 347.3752 | 273.0812 | 341.7632 | 243.727 | 92.5815 | 1.6498 | 42.1206 | 69.4918 | 25.5303 | 2.6194 | 41.9951 | 110.0034 | 16.0321 | 21.262 | 6.3826 | 179.4952 | 326.1132 | 251.8192 | 320.5012 | 222.465
2012-10-30 02:02 | 25.7522 | 347.3819 | 273.0568 | 341.728 | 242.6911 | 92.5867 | 1.6496 | 42.1133 | 69.4682 | 25.5301 | 2.619 | 41.9879 | 109.9669 | 16.0319 | 21.2725 | 6.3829 | 179.435 | 326.1094 | 251.7843 | 320.4555 | 221.4186
2012-10-30 02:02 | 25.7515 | 347.3698 | 272.8429 | 341.7425 | 241.2281 | 92.5873 | 1.6495 | 42.1121 | 69.464 | 25.5302 | 2.619 | 41.9866 | 109.9615 | 16.0318 | 21.2839 | 6.383 | 179.4255 | 326.0859 | 251.559 | 320.4586 | 219.9442
2012-10-30 02:02 | 25.7537 | 347.3682 | 272.655 | 341.7523 | 242.0576 | 92.5746 | 1.6497 | 42.1168 | 69.4798 | 25.5301 | 2.6193 | 41.9913 | 109.9859 | 16.0318 | 21.282 | 6.3821 | 179.4656 | 326.0862 | 251.373 | 320.4703 | 220.7756
2012-10-30 02:02 | 25.7544 | 347.3195 | 272.5782 | 341.7759 | 242.7756 | 92.5833 | 1.6498 | 42.1202 | 69.491 | 25.5301 | 2.6195 | 41.9948 | 110.0048 | 16.0317 | 21.3189 | 6.3827 | 179.4958 | 326.0006 | 251.2593 | 320.457 | 221.4567
2012-10-30 02:02 | 25.7559 | 347.3362 | 272.5293 | 341.7507 | 242.9456 | 92.5625 | 1.65 | 42.1255 | 69.508 | 25.5303 | 2.6198 | 41.9997 | 110.0302 | 16.0318 | 21.3324 | 6.3813 | 179.5382 | 326.0038 | 251.1969 | 320.4183 | 221.6132
2012-10-30 02:02 | 25.7657 | 347.3346 | 272.5662 | 341.7339 | 243.0652 | 92.5798 | 1.6502 | 42.1298 | 69.5221 | 25.5303 | 2.62 | 42.0042 | 110.053 | 16.0318 | 21.29 | 6.3825 | 179.5751 | 326.0446 | 251.2762 | 320.4439 | 221.7752
2012-10-30 02:02 | 25.7574 | 347.311 | 272.5812 | 341.7446 | 243.437 | 92.5775 | 1.6505 | 42.1379 | 69.5488 | 25.5303 | 2.6205 | 42.0121 | 110.0946 | 16.0319 | 21.3299 | 6.3823 | 179.6434 | 325.9811 | 251.2513 | 320.4147 | 222.1071
2012-10-30 02:03 | 25.7604 | 347.3102 | 272.6424 | 341.7552 | 243.7143 | 92.5833 | 1.652 | 42.1773 | 69.6778 | 25.5308 | 2.6229 | 42.0505 | 110.2949 | 16.032 | 21.3008 | 6.3827 | 179.9727 | 326.0094 | 251.3416 | 320.4544 | 222.4135
2012-10-30 02:03 | 25.7559 | 347.3324 | 272.6805 | 341.7203 | 243.7843 | 92.574 | 1.6539 | 42.2272 | 69.8399 | 25.5317 | 2.626 | 42.1009 | 110.5556 | 16.0326 | 21.3123 | 6.3821 | 180.3955 | 326.0201 | 251.3682 | 320.408 | 222.472
2012-10-30 02:03 | 25.7537 | 347.3226 | 272.7598 | 341.7257 | 243.4999 | 92.5833 | 1.6545 | 42.2415 | 69.8872 | 25.5317 | 2.6268 | 42.1154 | 110.6301 | 16.0328 | 21.3222 | 6.3827 | 180.5172 | 326.0004 | 251.4376 | 320.4035 | 222.1777
2012-10-30 02:03 | 25.7597 | 347.3704 | 272.7386 | 341.7241 | 242.9573 | 92.5821 | 1.6549 | 42.2543 | 69.9288 | 25.5321 | 2.6276 | 42.1285 | 110.6982 | 16.0328 | 21.3442 | 6.3826 | 180.627 | 326.0262 | 251.3944 | 320.3799 | 221.6131
2012-10-30 02:03 | 25.7559 | 347.3172 | 272.6805 | 341.6784 | 243.3454 | 92.5937 | 1.6552 | 42.261 | 69.9503 | 25.5323 | 2.628 | 42.1348 | 110.7304 | 16.033 | 21.3204 | 6.3834 | 180.6807 | 325.9968 | 251.3601 | 320.358 | 222.025
2012-10-30 02:03 | 25.7537 | 347.372 | 272.6744 | 341.6876 | 243.4686 | 92.6023 | 1.6552 | 42.2621 | 69.9539 | 25.5323 | 2.6281 | 42.1358 | 110.7356 | 16.033 | 21.2981 | 6.384 | 180.6894 | 326.0739 | 251.3763 | 320.3895 | 222.1705
2012-10-30 02:03 | 25.7537 | 347.3492 | 272.7365 | 341.7029 | 243.4882 | 92.5977 | 1.6552 | 42.2625 | 69.9542 | 25.5327 | 2.6281 | 42.1363 | 110.737 | 16.0332 | 21.2861 | 6.3837 | 180.6912 | 326.0631 | 251.4504 | 320.4168 | 222.2021
2012-10-30 02:03 | 25.7657 | 347.3574 | 272.7834 | 341.6692 | 243.6217 | 92.5873 | 1.6553 | 42.2636 | 69.9579 | 25.5327 | 2.6281 | 42.1376 | 110.7434 | 16.0332 | 21.3101 | 6.383 | 180.7012 | 326.0473 | 251.4733 | 320.3591 | 222.3116
2012-10-30 02:03 | 25.7582 | 347.3271 | 272.8341 | 341.6807 | 243.622 | 92.6093 | 1.6552 | 42.2613 | 69.9504 | 25.5326 | 2.628 | 42.1351 | 110.73 | 16.0333 | 21.3186 | 6.3845 | 180.6805 | 326.0085 | 251.5155 | 320.3621 | 222.3034
2012-10-30 02:03 | 25.7574 | 347.292 | 272.841 | 341.6495 | 243.5585 | 92.5966 | 1.6552 | 42.2608 | 69.949 | 25.5325 | 2.628 | 42.1346 | 110.7281 | 16.0332 | 21.2777 | 6.3836 | 180.677 | 326.0143 | 251.5633 | 320.3718 | 222.2808
2012-10-30 02:03 | 25.7567 | 347.356 | 272.9179 | 341.683 | 243.6518 | 92.5954 | 1.6551 | 42.2599 | 69.9462 | 25.5325 | 2.6279 | 42.1338 | 110.7237 | 16.0332 | 21.3212 | 6.3835 | 180.6699 | 326.0348 | 251.5967 | 320.3618 | 222.3306
2012-10-30 02:03 | 25.7552 | 347.3469 | 272.9629 | 341.7006 | 243.7757 | 92.5977 | 1.6551 | 42.2603 | 69.9469 | 25.5328 | 2.6279 | 42.1341 | 110.7235 | 16.0334 | 21.3036 | 6.3837 | 180.6704 | 326.0433 | 251.6593 | 320.397 | 222.4721
2012-10-30 02:03 | 25.7642 | 347.3331 | 272.9913 | 341.7019 | 243.7338 | 92.6047 | 1.6551 | 42.2594 | 69.944 | 25.5327 | 2.6279 | 42.1333 | 110.72 | 16.0334 | 21.2885 | 6.3842 | 180.664 | 326.0446 | 251.7028 | 320.4134 | 222.4453
2012-10-30 02:03 | 25.7604 | 347.3293 | 272.9681 | 341.6677 | 243.7182 | 92.6064 | 1.6551 | 42.2583 | 69.9403 | 25.5327 | 2.6278 | 42.1322 | 110.7143 | 16.0333 | 21.3128 | 6.3843 | 180.6546 | 326.0165 | 251.6553 | 320.3549 | 222.4054
2012-10-30 02:03 | 25.7604 | 347.3635 | 273.0069 | 341.6905 | 243.6869 | 92.6006 | 1.655 | 42.2573 | 69.9368 | 25.5328 | 2.6277 | 42.1311 | 110.7097 | 16.0332 | 21.3048 | 6.3839 | 180.6464 | 326.0587 | 251.7021 | 320.3857 | 222.3821
2012-10-30 02:03 | 25.7657 | 347.3536 | 272.9812 | 341.6958 | 243.653 | 92.6075 | 1.655 | 42.2569 | 69.9362 | 25.5325 | 2.6277 | 42.1307 | 110.7078 | 16.0332 | 21.3021 | 6.3844 | 180.644 | 326.0515 | 251.6791 | 320.3937 | 222.3509
2012-10-30 02:03 | 25.7634 | 347.3855 | 272.9905 | 341.6897 | 243.6546 | 92.607 | 1.655 | 42.2557 | 69.9323 | 25.5325 | 2.6277 | 42.1294 | 110.7018 | 16.033 | 21.3038 | 6.3843 | 180.6342 | 326.0817 | 251.6867 | 320.3859 | 222.3508
2012-10-30 02:03 | 25.7574 | 347.4024 | 273.0504 | 341.7256 | 243.7152 | 92.6081 | 1.6549 | 42.2542 | 69.9276 | 25.5323 | 2.6276 | 42.1283 | 110.6955 | 16.0331 | 21.3058 | 6.3844 | 180.6231 | 326.0966 | 251.7446 | 320.4198 | 222.4094
2012-10-30 02:03 | 25.7544 | 347.3879 | 273.0513 | 341.7036 | 243.6613 | 92.607 | 1.6549 | 42.2527 | 69.9221 | 25.5326 | 2.6275 | 42.1267 | 110.6875 | 16.0331 | 21.2988 | 6.3843 | 180.6096 | 326.0891 | 251.7525 | 320.4048 | 222.3625
2012-10-30 02:03 | 25.7642 | 347.3977 | 273.0339 | 341.7324 | 243.7416 | 92.6064 | 1.6548 | 42.2507 | 69.9153 | 25.5327 | 2.6273 | 42.1246 | 110.6759 | 16.0331 | 21.2966 | 6.3843 | 180.5912 | 326.1011 | 251.7373 | 320.4358 | 222.445
2012-10-30 02:03 | 25.7634 | 347.3855 | 273.0564 | 341.7468 | 243.8936 | 92.6006 | 1.6547 | 42.2487 | 69.9089 | 25.5326 | 2.6272 | 42.1226 | 110.6652 | 16.0332 | 21.2958 | 6.3839 | 180.5741 | 326.0897 | 251.7606 | 320.451 | 222.5978
2012-10-30 02:03 | 25.7679 | 347.4167 | 273.0803 | 341.7323 | 243.9373 | 92.6162 | 1.6546 | 42.2474 | 69.9044 | 25.5326 | 2.6271 | 42.1213 | 110.6585 | 16.0332 | 21.2561 | 6.385 | 180.5629 | 326.1606 | 251.8242 | 320.4762 | 222.6812
2012-10-30 02:04 | 25.7664 | 347.4152 | 273.0982 | 341.7308 | 243.8457 | 92.611 | 1.6545 | 42.2448 | 69.8956 | 25.5327 | 2.627 | 42.1186 | 110.6448 | 16.0331 | 21.2586 | 6.3846 | 180.5404 | 326.1566 | 251.8396 | 320.4722 | 222.5871
2012-10-30 02:04 | 25.7619 | 347.4449 | 273.1247 | 341.7643 | 243.661 | 92.6151 | 1.6545 | 42.2432 | 69.8903 | 25.5327 | 2.6269 | 42.1173 | 110.6379 | 16.0331 | 21.2822 | 6.3849 | 180.5282 | 326.1627 | 251.8425 | 320.4821 | 222.3788
2012-10-30 02:04 | 25.7619 | 347.4449 | 273.1363 | 341.7491 | 243.9666 | 92.6093 | 1.6544 | 42.2416 | 69.885 | 25.5327 | 2.6268 | 42.1155 | 110.6288 | 16.033 | 21.2662 | 6.3845 | 180.5138 | 326.1787 | 251.8701 | 320.4829 | 222.7004
2012-10-30 02:04 | 25.7634 | 347.4616 | 273.165 | 341.7506 | 243.8074 | 92.6156 | 1.6543 | 42.2397 | 69.8789 | 25.5327 | 2.6267 | 42.1137 | 110.6199 | 16.033 | 21.2958 | 6.3849 | 180.4988 | 326.1658 | 251.8692 | 320.4548 | 222.5116
2012-10-30 02:04 | 25.7724 | 347.4744 | 273.1352 | 341.7634 | 243.7341 | 92.6104 | 1.6543 | 42.2384 | 69.8751 | 25.5325 | 2.6266 | 42.1124 | 110.6138 | 16.0329 | 21.2767 | 6.3846 | 180.4889 | 326.1977 | 251.8585 | 320.4867 | 222.4574
2012-10-30 02:04 | 25.7687 | 347.4517 | 273.1082 | 341.7825 | 243.7461 | 92.6122 | 1.6543 | 42.2375 | 69.8719 | 25.5325 | 2.6266 | 42.1116 | 110.6087 | 16.033 | 21.3051 | 6.3847 | 180.4806 | 326.1466 | 251.8031 | 320.4774 | 222.441
2012-10-30 02:04 | 25.7664 | 347.4684 | 273.1757 | 341.7879 | 243.5401 | 92.6168 | 1.6542 | 42.2359 | 69.8666 | 25.5325 | 2.6265 | 42.1098 | 110.6005 | 16.0328 | 21.2867 | 6.385 | 180.4671 | 326.1817 | 251.889 | 320.5012 | 222.2534
2012-10-30 02:04 | 25.7709 | 347.4805 | 273.091 | 341.7847 | 243.282 | 92.6075 | 1.6541 | 42.234 | 69.8608 | 25.5324 | 2.6264 | 42.1081 | 110.5912 | 16.0328 | 21.2952 | 6.3844 | 180.452 | 326.1853 | 251.7958 | 320.4895 | 221.9868
2012-10-30 02:04 | 25.7717 | 347.4661 | 273.0414 | 341.7551 | 243.1182 | 92.6185 | 1.6541 | 42.2329 | 69.8567 | 25.5325 | 2.6263 | 42.1067 | 110.5839 | 16.0328 | 21.2559 | 6.3851 | 180.4406 | 326.2102 | 251.7855 | 320.4992 | 221.8623
2012-10-30 02:04 | 25.7702 | 347.4646 | 272.9818 | 341.765 | 243.0305 | 92.6214 | 1.654 | 42.2312 | 69.8508 | 25.5326 | 2.6262 | 42.1053 | 110.5761 | 16.0329 | 21.2865 | 6.3853 | 180.4269 | 326.1781 | 251.6953 | 320.4785 | 221.744
2012-10-30 02:04 | 25.7709 | 347.4653 | 272.9747 | 341.7733 | 243.5994 | 92.6029 | 1.6517 | 42.1708 | 69.6529 | 25.5319 | 2.6226 | 42.0465 | 110.2707 | 16.0325 | 21.2591 | 6.384 | 179.9236 | 326.2062 | 251.7156 | 320.5142 | 222.3403
2012-10-30 02:04 | 25.7702 | 347.4874 | 272.9701 | 341.7993 | 243.6614 | 92.5937 | 1.6502 | 42.132 | 69.5274 | 25.531 | 2.6201 | 42.0061 | 110.0616 | 16.032 | 21.2624 | 6.3834 | 179.589 | 326.225 | 251.7077 | 320.5369 | 222.399
2012-10-30 02:04 | 25.7732 | 347.479 | 273.0468 | 341.8289 | 243.8055 | 92.6064 | 1.6497 | 42.1193 | 69.4857 | 25.5309 | 2.6194 | 41.994 | 109.9995 | 16.0318 | 21.2694 | 6.3843 | 179.4852 | 326.2096 | 251.7774 | 320.5595 | 222.5361
2012-10-30 02:04 | 25.7732 | 347.441 | 273.0391 | 341.7947 | 243.2412 | 92.6133 | 1.6496 | 42.1154 | 69.4726 | 25.531 | 2.6192 | 41.9899 | 109.978 | 16.0318 | 21.2534 | 6.3848 | 179.4507 | 326.1876 | 251.7857 | 320.5413 | 221.9878
2012-10-30 02:04 | 25.7717 | 347.4433 | 272.9678 | 341.7855 | 243.2201 | 92.596 | 1.6497 | 42.1192 | 69.485 | 25.5311 | 2.6194 | 41.9936 | 109.9974 | 16.0318 | 21.28 | 6.3836 | 179.4824 | 326.1633 | 251.6878 | 320.5055 | 221.9401
2012-10-30 02:04 | 25.7747 | 347.4387 | 272.8544 | 341.8114 | 243.4426 | 92.5925 | 1.6498 | 42.1202 | 69.4888 | 25.5309 | 2.6195 | 41.9946 | 110.004 | 16.0317 | 21.2669 | 6.3833 | 179.4928 | 326.1718 | 251.5875 | 320.5445 | 222.1757
2012-10-30 02:04 | 25.7672 | 347.4426 | 272.8818 | 341.762 | 242.6904 | 92.5833 | 1.6499 | 42.1232 | 69.4994 | 25.5307 | 2.6197 | 41.9974 | 110.0189 | 16.0316 | 21.2634 | 6.3827 | 179.5183 | 326.1792 | 251.6184 | 320.4986 | 221.427
2012-10-30 02:04 | 25.7724 | 347.4097 | 272.697 | 341.7672 | 243.2051 | 92.5879 | 1.65 | 42.1251 | 69.5049 | 25.531 | 2.6198 | 41.9996 | 110.0294 | 16.0317 | 21.2526 | 6.383 | 179.5342 | 326.1571 | 251.4444 | 320.5146 | 221.9525
2012-10-30 02:04 | 25.7672 | 347.4312 | 272.5949 | 341.7392 | 241.3575 | 92.581 | 1.6501 | 42.1298 | 69.5197 | 25.5312 | 2.6201 | 42.0041 | 110.0528 | 16.0318 | 21.2554 | 6.3825 | 179.5725 | 326.1758 | 251.3395 | 320.4838 | 220.1021
2012-10-30 02:04 | 25.7709 | 347.355 | 272.3736 | 341.7467 | 241.0318 | 92.5827 | 1.6504 | 42.1372 | 69.5446 | 25.5311 | 2.6205 | 42.0114 | 110.0906 | 16.0319 | 21.2832 | 6.3827 | 179.6353 | 326.0718 | 251.0904 | 320.4635 | 219.7486
2012-10-30 02:04 | 25.7664 | 347.3733 | 272.148 | 341.7003 | 241.2861 | 92.5989 | 1.653 | 42.2055 | 69.7678 | 25.5319 | 2.6245 | 42.0765 | 110.4292 | 16.0323 | 21.3068 | 6.3838 | 180.197 | 326.0665 | 250.8412 | 320.3935 | 219.9793
2012-10-30 02:04 | 25.7589 | 347.3962 | 272.0086 | 341.7043 | 241.4982 | 92.5977 | 1.654 | 42.2302 | 69.8482 | 25.5324 | 2.6262 | 42.104 | 110.5719 | 16.0325 | 21.3073 | 6.3837 | 180.4201 | 326.0889 | 250.7013 | 320.397 | 220.1909
2012-10-30 02:04 | 25.7619 | 347.3954 | 271.9806 | 341.7187 | 241.6227 | 92.5989 | 1.6544 | 42.2416 | 69.8856 | 25.5325 | 2.6269 | 42.1156 | 110.6317 | 16.0327 | 21.2902 | 6.3838 | 180.5174 | 326.1052 | 250.6904 | 320.4285 | 220.3325
2012-10-30 02:04 | 25.7672 | 347.3931 | 271.9665 | 341.743 | 242.1416 | 92.5983 | 1.6548 | 42.2514 | 69.9174 | 25.5328 | 2.6275 | 42.1252 | 110.6821 | 16.0327 | 21.2755 | 6.3837 | 180.5995 | 326.1176 | 250.691 | 320.4675 | 220.8661
2012-10-30 02:05 | 25.7574 | 347.3643 | 272.0343 | 341.6685 | 242.3788 | 92.6029 | 1.6551 | 42.2589 | 69.9422 | 25.5327 | 2.6279 | 42.1328 | 110.7206 | 16.0329 | 21.2697 | 6.384 | 180.6628 | 326.0946 | 250.7646 | 320.3988 | 221.1091
2012-10-30 02:05 | 25.7694 | 347.3801 | 272.0347 | 341.6577 | 242.3398 | 92.6012 | 1.6552 | 42.2618 | 69.9515 | 25.5329 | 2.6281 | 42.1357 | 110.7355 | 16.033 | 21.3298 | 6.3839 | 180.687 | 326.0503 | 250.7049 | 320.3279 | 221.01
2012-10-30 02:05 | 25.7597 | 347.3742 | 272.0832 | 341.6442 | 242.8201 | 92.5977 | 1.6553 | 42.2634 | 69.9566 | 25.5329 | 2.6282 | 42.1372 | 110.7432 | 16.033 | 21.3522 | 6.3837 | 180.6998 | 326.022 | 250.731 | 320.292 | 221.4679
2012-10-30 02:05 | 25.7559 | 347.3742 | 272.1531 | 341.6556 | 242.6752 | 92.6041 | 1.6552 | 42.2634 | 69.9564 | 25.533 | 2.6281 | 42.1373 | 110.742 | 16.0332 | 21.3244 | 6.3841 | 180.6984 | 326.0498 | 250.8287 | 320.3312 | 221.3508
2012-10-30 02:05 | 25.7582 | 347.3651 | 272.1437 | 341.6617 | 242.6383 | 92.5942 | 1.6553 | 42.264 | 69.9585 | 25.533 | 2.6282 | 42.1381 | 110.7468 | 16.0332 | 21.2946 | 6.3834 | 180.7053 | 326.0705 | 250.8491 | 320.3671 | 221.3437
2012-10-30 02:05 | 25.7642 | 347.3331 | 272.235 | 341.6715 | 242.7109 | 92.5948 | 1.6552 | 42.2633 | 69.9556 | 25.5331 | 2.6281 | 42.1371 | 110.742 | 16.0331 | 21.3287 | 6.3835 | 180.6976 | 326.0044 | 250.9063 | 320.3428 | 221.3822
2012-10-30 02:05 | 25.753 | 347.2876 | 272.2277 | 341.6527 | 242.7899 | 92.5914 | 1.6553 | 42.264 | 69.9576 | 25.5333 | 2.6282 | 42.1378 | 110.7448 | 16.0332 | 21.2854 | 6.3833 | 180.7024 | 326.0022 | 250.9423 | 320.3673 | 221.5045
2012-10-30 02:05 | 25.7597 | 347.3248 | 272.3042 | 341.6594 | 242.8671 | 92.6041 | 1.6552 | 42.2628 | 69.9535 | 25.5333 | 2.6281 | 42.1365 | 110.738 | 16.0332 | 21.3242 | 6.3841 | 180.6915 | 326.0006 | 250.98 | 320.3352 | 221.5429
2012-10-30 02:05 | 25.7612 | 347.3491 | 272.4143 | 341.6457 | 243.0685 | 92.6018 | 1.655 | 42.2587 | 69.9399 | 25.5333 | 2.6278 | 42.1326 | 110.718 | 16.0331 | 21.2936 | 6.384 | 180.6579 | 326.0555 | 251.1207 | 320.3521 | 221.7749
2012-10-30 02:05 | 25.7582 | 347.3841 | 272.5005 | 341.6655 | 243.1204 | 92.6018 | 1.6549 | 42.2552 | 69.9291 | 25.5331 | 2.6277 | 42.1292 | 110.7014 | 16.033 | 21.3146 | 6.384 | 180.6305 | 326.0695 | 251.1859 | 320.3509 | 221.8058
2012-10-30 02:05 | 25.7582 | 347.3499 | 272.5432 | 341.6655 | 243.1713 | 92.607 | 1.6549 | 42.2542 | 69.9254 | 25.5332 | 2.6276 | 42.128 | 110.6946 | 16.033 | 21.3106 | 6.3843 | 180.6201 | 326.0393 | 251.2326 | 320.3549 | 221.8607
2012-10-30 02:05 | 25.7657 | 347.3574 | 272.6283 | 341.6844 | 243.3591 | 92.6203 | 1.6549 | 42.2538 | 69.9238 | 25.5332 | 2.6275 | 42.1275 | 110.6919 | 16.033 | 21.3261 | 6.3852 | 180.6157 | 326.0313 | 251.3022 | 320.3583 | 222.033
2012-10-30 02:05 | 25.7522 | 347.3667 | 272.6148 | 341.6785 | 243.3456 | 92.6145 | 1.6548 | 42.252 | 69.9179 | 25.5332 | 2.6275 | 42.1259 | 110.6841 | 16.033 | 21.3006 | 6.3848 | 180.602 | 326.0661 | 251.3142 | 320.3779 | 222.045
2012-10-30 02:05 | 25.7544 | 347.3803 | 272.7178 | 341.6731 | 243.6064 | 92.6168 | 1.6547 | 42.2503 | 69.9123 | 25.5332 | 2.6274 | 42.1242 | 110.675 | 16.0329 | 21.3108 | 6.385 | 180.5873 | 326.0695 | 251.407 | 320.3623 | 222.2956
2012-10-30 02:05 | 25.7649 | 347.3946 | 272.8369 | 341.6798 | 243.4249 | 92.6133 | 1.6546 | 42.248 | 69.9051 | 25.5331 | 2.6272 | 42.1219 | 110.6634 | 16.0329 | 21.3253 | 6.3848 | 180.5685 | 326.0693 | 251.5116 | 320.3545 | 222.0996
2012-10-30 02:05 | 25.7649 | 347.3794 | 272.771 | 341.7065 | 243.1153 | 92.6099 | 1.6545 | 42.2456 | 69.897 | 25.5331 | 2.6271 | 42.1195 | 110.6515 | 16.0328 | 21.3053 | 6.3845 | 180.5485 | 326.0741 | 251.4657 | 320.4012 | 221.81
2012-10-30 02:05 | 25.7642 | 347.413 | 272.8323 | 341.6715 | 243.193 | 92.6185 | 1.6544 | 42.2433 | 69.889 | 25.5333 | 2.6269 | 42.1173 | 110.6389 | 16.033 | 21.3206 | 6.3851 | 180.5279 | 326.0924 | 251.5117 | 320.3509 | 221.8724
2012-10-30 02:05 | 25.7724 | 347.4059 | 272.8172 | 341.7025 | 243.3697 | 92.6029 | 1.6543 | 42.2405 | 69.8804 | 25.5331 | 2.6268 | 42.1145 | 110.6243 | 16.0329 | 21.3369 | 6.384 | 180.5047 | 326.069 | 251.4803 | 320.3656 | 222.0328
2012-10-30 02:05 | 25.7634 | 347.4007 | 272.8315 | 341.6897 | 243.3059 | 92.6272 | 1.6543 | 42.2393 | 69.8765 | 25.533 | 2.6267 | 42.1133 | 110.6188 | 16.0328 | 21.3158 | 6.3857 | 180.4953 | 326.0849 | 251.5157 | 320.3739 | 221.9901
2012-10-30 02:06 | 25.7724 | 347.4212 | 272.8483 | 341.7178 | 243.4167 | 92.6151 | 1.6542 | 42.2372 | 69.8696 | 25.533 | 2.6266 | 42.1113 | 110.6091 | 16.0327 | 21.3248 | 6.3849 | 180.4787 | 326.0964 | 251.5235 | 320.393 | 222.0919
2012-10-30 02:06 | 25.7664 | 347.3847 | 272.8151 | 341.6851 | 243.2226 | 92.611 | 1.6541 | 42.2353 | 69.8632 | 25.533 | 2.6265 | 42.1093 | 110.5985 | 16.0327 | 21.3028 | 6.3846 | 180.4617 | 326.0819 | 251.5123 | 320.3823 | 221.9198
2012-10-30 02:06 | 25.7747 | 347.4311 | 272.8583 | 341.7201 | 243.3172 | 92.6145 | 1.6531 | 42.2081 | 69.7749 | 25.5324 | 2.6248 | 42.0821 | 110.4584 | 16.0323 | 21.299 | 6.3848 | 180.2333 | 326.1321 | 251.5593 | 320.4211 | 222.0182
2012-10-30 02:06 | 25.7732 | 347.4029 | 272.8491 | 341.7452 | 243.1628 | 92.6 | 1.6499 | 42.1235 | 69.4989 | 25.5311 | 2.6197 | 41.9983 | 110.024 | 16.0315 | 21.3216 | 6.3838 | 179.5229 | 326.0813 | 251.5275 | 320.4236 | 221.8412
2012-10-30 02:06 | 25.7634 | 347.3779 | 272.7811 | 341.7506 | 243.0002 | 92.6023 | 1.6496 | 42.1169 | 69.4778 | 25.5309 | 2.6193 | 41.9912 | 109.9885 | 16.0313 | 21.3038 | 6.384 | 179.4663 | 326.0741 | 251.4773 | 320.4468 | 221.6964
2012-10-30 02:06 | 25.7732 | 347.4105 | 272.7288 | 341.7414 | 242.4455 | 92.5914 | 1.6496 | 42.1157 | 69.4736 | 25.5311 | 2.6192 | 41.99 | 109.9823 | 16.0313 | 21.3136 | 6.3833 | 179.4559 | 326.0969 | 251.4152 | 320.4278 | 221.1319
2012-10-30 02:06 | 25.7702 | 347.4037 | 272.6056 | 341.7232 | 242.7875 | 92.596 | 1.6497 | 42.1176 | 69.4805 | 25.5309 | 2.6194 | 41.9919 | 109.9921 | 16.0313 | 21.2945 | 6.3836 | 179.4726 | 326.1092 | 251.3111 | 320.4287 | 221.493
2012-10-30 02:06 | 25.7717 | 347.3824 | 272.5839 | 341.7361 | 243.224 | 92.5775 | 1.6499 | 42.1222 | 69.4954 | 25.5308 | 2.6197 | 41.9965 | 110.0165 | 16.0313 | 21.3081 | 6.3823 | 179.5119 | 326.0743 | 251.2758 | 320.428 | 221.9159
2012-10-30 02:06 | 25.7694 | 347.4067 | 272.6048 | 341.7071 | 243.4059 | 92.585 | 1.65 | 42.125 | 69.5053 | 25.5307 | 2.6198 | 41.9991 | 110.0299 | 16.0313 | 21.3298 | 6.3828 | 179.5352 | 326.0769 | 251.275 | 320.3773 | 222.0761
2012-10-30 02:06 | 25.7664 | 347.4304 | 272.6212 | 341.7384 | 243.3755 | 92.5804 | 1.6501 | 42.1298 | 69.5195 | 25.5312 | 2.6201 | 42.004 | 110.0544 | 16.0315 | 21.2988 | 6.3825 | 179.5739 | 326.1316 | 251.3224 | 320.4396 | 222.0767
2012-10-30 02:06 | 25.7619 | 347.3802 | 272.6206 | 341.7149 | 243.7276 | 92.5821 | 1.6503 | 42.1344 | 69.5352 | 25.5311 | 2.6204 | 42.0087 | 110.0793 | 16.0314 | 21.3504 | 6.3826 | 179.6145 | 326.0298 | 251.2702 | 320.3645 | 222.3772
2012-10-30 02:06 | 25.7619 | 347.3536 | 272.6826 | 341.6882 | 243.8569 | 92.5781 | 1.6506 | 42.141 | 69.5563 | 25.5313 | 2.6208 | 42.0152 | 110.113 | 16.0315 | 21.3103 | 6.3823 | 179.6693 | 326.0433 | 251.3723 | 320.3779 | 222.5466
2012-10-30 02:06 | 25.7574 | 347.3491 | 272.7596 | 341.699 | 243.6565 | 92.5792 | 1.6526 | 42.1947 | 69.7322 | 25.5319 | 2.6242 | 42.07 | 110.3983 | 16.0318 | 21.358 | 6.3824 | 180.1305 | 325.9911 | 251.4016 | 320.341 | 222.2985
2012-10-30 02:06 | 25.7619 | 347.3308 | 272.7835 | 341.6996 | 242.9477 | 92.5729 | 1.654 | 42.2297 | 69.8462 | 25.5325 | 2.6262 | 42.1039 | 110.5735 | 16.0322 | 21.3103 | 6.382 | 180.4198 | 326.0205 | 251.4732 | 320.3893 | 221.6374
2012-10-30 02:06 | 25.753 | 347.3181 | 272.7203 | 341.6223 | 242.9545 | 92.5931 | 1.6544 | 42.241 | 69.8839 | 25.5324 | 2.6269 | 42.1148 | 110.6305 | 16.0323 | 21.3134 | 6.3834 | 180.5144 | 326.0047 | 251.4069 | 320.3089 | 221.6411
2012-10-30 02:06 | 25.7515 | 347.3546 | 272.7033 | 341.655 | 242.906 | 92.5983 | 1.6548 | 42.2504 | 69.9139 | 25.5327 | 2.6274 | 42.1239 | 110.6778 | 16.0323 | 21.3119 | 6.3837 | 180.5916 | 326.0427 | 251.3914 | 320.3431 | 221.5941
2012-10-30 02:06 | 25.7567 | 347.3826 | 272.6774 | 341.6488 | 243.2208 | 92.5908 | 1.655 | 42.2575 | 69.9373 | 25.5329 | 2.6279 | 42.1313 | 110.7152 | 16.0325 | 21.3131 | 6.3832 | 180.6525 | 326.0695 | 251.3643 | 320.3357 | 221.9077
2012-10-30 02:06 | 25.7552 | 347.4192 | 272.6953 | 341.6359 | 243.4348 | 92.5977 | 1.6552 | 42.2622 | 69.9522 | 25.533 | 2.6281 | 42.136 | 110.7382 | 16.0328 | 21.2795 | 6.3837 | 180.6904 | 326.1397 | 251.4158 | 320.3564 | 222.1553
2012-10-30 02:06 | 25.7507 | 347.388 | 272.7141 | 341.6428 | 243.5008 | 92.5937 | 1.6552 | 42.2638 | 69.957 | 25.5332 | 2.6282 | 42.1377 | 110.7483 | 16.0326 | 21.3432 | 6.3834 | 180.7052 | 326.0448 | 251.3709 | 320.2996 | 222.1576
2012-10-30 02:06 | 25.7537 | 347.3644 | 272.783 | 341.6153 | 243.2844 | 92.5856 | 1.6553 | 42.2639 | 69.958 | 25.5329 | 2.6283 | 42.1378 | 110.7488 | 16.0326 | 21.3141 | 6.3829 | 180.7068 | 326.0503 | 251.4689 | 320.3012 | 221.9703
2012-10-30 02:07 | 25.7492 | 347.398 | 272.8018 | 341.6261 | 243.4092 | 92.5948 | 1.6553 | 42.2641 | 69.958 | 25.5333 | 2.6283 | 42.1379 | 110.7491 | 16.0326 | 21.3137 | 6.3835 | 180.7071 | 326.0843 | 251.4881 | 320.3124 | 222.0955
2012-10-30 02:07 | 25.744 | 347.389 | 272.7578 | 341.6209 | 243.5255 | 92.5885 | 1.6552 | 42.2627 | 69.954 | 25.533 | 2.6282 | 42.1366 | 110.7421 | 16.0327 | 21.2764 | 6.3831 | 180.6962 | 326.1126 | 251.4814 | 320.3445 | 222.2491
2012-10-30 02:07 | 25.7395 | 347.3845 | 272.7766 | 341.6202 | 243.6581 | 92.5948 | 1.6551 | 42.26 | 69.944 | 25.5334 | 2.628 | 42.1336 | 110.7264 | 16.0327 | 21.2478 | 6.3835 | 180.6704 | 326.1367 | 251.5288 | 320.3724 | 222.4103
2012-10-30 02:07 | 25.7432 | 347.3767 | 272.8384 | 341.6239 | 243.6031 | 92.5971 | 1.6551 | 42.2598 | 69.9439 | 25.5332 | 2.628 | 42.1334 | 110.7255 | 16.0327 | 21.2715 | 6.3836 | 180.6693 | 326.1052 | 251.5669 | 320.3524 | 222.3316
2012-10-30 02:07 | 25.7425 | 347.3532 | 272.861 | 341.6574 | 243.7591 | 92.5919 | 1.655 | 42.2585 | 69.9388 | 25.5335 | 2.6279 | 42.1323 | 110.7196 | 16.0326 | 21.2869 | 6.3833 | 180.6584 | 326.0663 | 251.5741 | 320.3705 | 222.4722
2012-10-30 02:07 | 25.744 | 347.3852 | 272.8974 | 341.6513 | 243.7057 | 92.6023 | 1.655 | 42.2581 | 69.9383 | 25.5331 | 2.6279 | 42.1317 | 110.7173 | 16.0326 | 21.2764 | 6.384 | 180.6556 | 326.1088 | 251.621 | 320.3749 | 222.4293
2012-10-30 02:07 | 25.7372 | 347.3936 | 272.9255 | 341.6559 | 243.6363 | 92.5919 | 1.655 | 42.2569 | 69.9345 | 25.5331 | 2.6278 | 42.1307 | 110.7121 | 16.0325 | 21.2696 | 6.3833 | 180.6466 | 326.124 | 251.6559 | 320.3863 | 222.3667
2012-10-30 02:07 | 25.7357 | 347.4263 | 272.9395 | 341.6544 | 243.721 | 92.6023 | 1.6549 | 42.2553 | 69.9293 | 25.5331 | 2.6277 | 42.1292 | 110.7036 | 16.0326 | 21.26 | 6.384 | 180.6329 | 326.1663 | 251.6795 | 320.3944 | 222.461
2012-10-30 02:07 | 25.7372 | 347.4164 | 272.9837 | 341.6559 | 243.8165 | 92.5954 | 1.6549 | 42.2538 | 69.9243 | 25.5331 | 2.6276 | 42.1275 | 110.6953 | 16.0326 | 21.2696 | 6.3835 | 180.6196 | 326.1468 | 251.7141 | 320.3863 | 222.5469
2012-10-30 02:07 | 25.7327 | 347.37 | 273.0296 | 341.659 | 243.8551 | 92.6099 | 1.6548 | 42.2518 | 69.9177 | 25.5331 | 2.6275 | 42.1259 | 110.6869 | 16.0325 | 21.2851 | 6.3845 | 180.6046 | 326.0849 | 251.7445 | 320.3739 | 222.57
2012-10-30 02:07 | 25.7395 | 347.3921 | 273.0674 | 341.6582 | 243.7561 | 92.6029 | 1.6548 | 42.2516 | 69.9172 | 25.533 | 2.6275 | 42.1254 | 110.6851 | 16.0324 | 21.2678 | 6.384 | 180.6023 | 326.1243 | 251.7996 | 320.3904 | 222.4883
2012-10-30 02:07 | 25.735 | 347.3838 | 273.1366 | 341.6651 | 243.9397 | 92.5896 | 1.6547 | 42.2498 | 69.9117 | 25.5329 | 2.6274 | 42.1238 | 110.6773 | 16.0323 | 21.3235 | 6.3831 | 180.5889 | 326.0603 | 251.8131 | 320.3416 | 222.6162
2012-10-30 02:07 | 25.7357 | 347.4644 | 273.1334 | 341.6468 | 243.7131 | 92.6058 | 1.6547 | 42.2492 | 69.9091 | 25.5331 | 2.6274 | 42.1232 | 110.6735 | 16.0324 | 21.2881 | 6.3842 | 180.5826 | 326.1763 | 251.8453 | 320.3587 | 222.425
2012-10-30 02:07 | 25.732 | 347.434 | 273.1064 | 341.6812 | 243.8622 | 92.5954 | 1.6546 | 42.2472 | 69.9018 | 25.5333 | 2.6272 | 42.1209 | 110.6609 | 16.0325 | 21.2924 | 6.3835 | 180.5627 | 326.1416 | 251.814 | 320.3888 | 222.5698
2012-10-30 02:07 | 25.7313 | 347.4257 | 273.1135 | 341.6424 | 243.8184 | 92.6087 | 1.6545 | 42.2445 | 69.8932 | 25.5332 | 2.6271 | 42.1183 | 110.6475 | 16.0325 | 21.2797 | 6.3844 | 180.5407 | 326.146 | 251.8338 | 320.3627 | 222.5387
2012-10-30 02:07 | 25.7283 | 347.4075 | 273.1066 | 341.6851 | 243.7292 | 92.5942 | 1.6545 | 42.2439 | 69.8909 | 25.5333 | 2.627 | 42.1177 | 110.6433 | 16.0326 | 21.2366 | 6.3834 | 180.5342 | 326.1709 | 251.87 | 320.4485 | 222.4926
2012-10-30 02:07 | 25.7357 | 347.4263 | 273.1334 | 341.6468 | 243.7641 | 92.5971 | 1.6544 | 42.2422 | 69.8863 | 25.533 | 2.6269 | 42.1162 | 110.6361 | 16.0325 | 21.2681 | 6.3836 | 180.5225 | 326.1582 | 251.8653 | 320.3787 | 222.496
2012-10-30 02:07 | 25.7342 | 347.4286 | 273.1629 | 341.6529 | 243.8409 | 92.6018 | 1.6543 | 42.2401 | 69.8785 | 25.5332 | 2.6268 | 42.1139 | 110.6239 | 16.0326 | 21.2465 | 6.384 | 180.5024 | 326.1821 | 251.9164 | 320.4064 | 222.5944
2012-10-30 02:07 | 25.7515 | 347.3888 | 273.1686 | 341.674 | 243.8151 | 92.6029 | 1.6542 | 42.2377 | 69.8705 | 25.5333 | 2.6266 | 42.1115 | 110.6119 | 16.0324 | 21.2879 | 6.384 | 180.4824 | 326.1009 | 251.8807 | 320.3861 | 222.5272
2012-10-30 02:07 | 25.7485 | 347.4467 | 273.1462 | 341.671 | 243.6515 | 92.6122 | 1.6542 | 42.2366 | 69.867 | 25.5332 | 2.6266 | 42.1103 | 110.606 | 16.0324 | 21.2367 | 6.3847 | 180.4731 | 326.21 | 251.9095 | 320.4343 | 222.4148
2012-10-30 02:07 | 25.7559 | 347.4389 | 273.204 | 341.686 | 243.9684 | 92.6018 | 1.6542 | 42.2367 | 69.8677 | 25.5331 | 2.6266 | 42.1107 | 110.6075 | 16.0324 | 21.2602 | 6.384 | 180.4752 | 326.1787 | 251.9438 | 320.4258 | 222.7082
2012-10-30 02:07 | 25.7485 | 347.4391 | 273.2044 | 341.7129 | 244.012 | 92.6047 | 1.6541 | 42.2335 | 69.8567 | 25.5332 | 2.6264 | 42.1076 | 110.591 | 16.0325 | 21.2728 | 6.3842 | 180.4476 | 326.1663 | 251.9316 | 320.4401 | 222.7392
2012-10-30 02:08 | 25.7612 | 347.4518 | 273.2364 | 341.7028 | 243.719 | 92.6099 | 1.654 | 42.2316 | 69.8504 | 25.5333 | 2.6263 | 42.1055 | 110.5802 | 16.0324 | 21.1932 | 6.3845 | 180.4306 | 326.2586 | 252.0432 | 320.5096 | 222.5258
2012-10-30 02:08 | 25.7552 | 347.4458 | 273.1839 | 341.7158 | 243.8267 | 92.6162 | 1.6539 | 42.23 | 69.8459 | 25.533 | 2.6262 | 42.1039 | 110.5733 | 16.0323 | 21.2675 | 6.385 | 180.4192 | 326.1783 | 251.9164 | 320.4483 | 222.5592
2012-10-30 02:08 | 25.7619 | 347.4639 | 273.2294 | 341.6844 | 243.8059 | 92.6116 | 1.6538 | 42.2273 | 69.8364 | 25.5332 | 2.626 | 42.1014 | 110.5603 | 16.0322 | 21.2541 | 6.3846 | 180.3966 | 326.2098 | 251.9753 | 320.4303 | 222.5518
2012-10-30 02:08 | 25.7634 | 347.4388 | 273.2425 | 341.6745 | 243.8819 | 92.5989 | 1.65 | 42.128 | 69.5123 | 25.5318 | 2.6202 | 42.0062 | 110.0652 | 16.0316 | 21.2436 | 6.3838 | 179.5776 | 326.1952 | 251.9989 | 320.4309 | 222.6383
2012-10-30 02:08 | 25.7717 | 347.4471 | 273.2469 | 341.7018 | 243.6433 | 92.611 | 1.6497 | 42.1208 | 69.4886 | 25.5317 | 2.6195 | 41.9949 | 110.0076 | 16.0314 | 21.272 | 6.3846 | 179.4962 | 326.1751 | 251.9749 | 320.4298 | 222.3713
2012-10-30 02:08 | 25.7687 | 347.387 | 273.1548 | 341.7141 | 243.2367 | 92.5914 | 1.6494 | 42.1113 | 69.458 | 25.5315 | 2.619 | 41.9856 | 109.9598 | 16.0312 | 21.3171 | 6.3833 | 179.4178 | 326.0699 | 251.8377 | 320.397 | 221.9196
2012-10-30 02:08 | 25.7761 | 347.4477 | 273.073 | 341.6872 | 242.9854 | 92.5931 | 1.6494 | 42.1122 | 69.4609 | 25.5315 | 2.6191 | 41.9865 | 109.9648 | 16.0312 | 21.3044 | 6.3834 | 179.4257 | 326.1433 | 251.7686 | 320.3828 | 221.681
2012-10-30 02:08 | 25.7784 | 347.3967 | 272.9667 | 341.7352 | 243.2111 | 92.581 | 1.6496 | 42.116 | 69.4735 | 25.5314 | 2.6193 | 41.9903 | 109.985 | 16.0312 | 21.3108 | 6.3825 | 179.4586 | 326.0859 | 251.6559 | 320.4244 | 221.9003
2012-10-30 02:08 | 25.7761 | 347.3906 | 272.9179 | 341.6948 | 243.5458 | 92.5879 | 1.6498 | 42.1207 | 69.489 | 25.5315 | 2.6196 | 41.9949 | 110.0095 | 16.0311 | 21.3044 | 6.383 | 179.4986 | 326.0862 | 251.6135 | 320.3904 | 222.2414
2012-10-30 02:08 | 25.7806 | 347.4027 | 272.9495 | 341.6765 | 243.7032 | 92.5798 | 1.6499 | 42.1243 | 69.5012 | 25.5313 | 2.6198 | 41.9985 | 110.0278 | 16.0311 | 21.3491 | 6.3825 | 179.5291 | 326.0536 | 251.6004 | 320.3274 | 222.3541
2012-10-30 02:08 | 25.7761 | 347.3716 | 272.9799 | 341.7177 | 243.4596 | 92.5867 | 1.6501 | 42.1298 | 69.5187 | 25.5316 | 2.6201 | 42.004 | 110.0567 | 16.0312 | 21.3727 | 6.3829 | 179.5753 | 325.9989 | 251.6072 | 320.345 | 222.0869
2012-10-30 02:08 | 25.7747 | 347.3778 | 272.9863 | 341.6934 | 243.5405 | 92.5763 | 1.6503 | 42.1335 | 69.531 | 25.5315 | 2.6204 | 42.0077 | 110.0758 | 16.0312 | 21.3592 | 6.3822 | 179.6067 | 326.0186 | 251.6271 | 320.3342 | 222.1813
2012-10-30 02:08 | 25.7769 | 347.399 | 272.9536 | 341.6728 | 243.8758 | 92.574 | 1.6508 | 42.1487 | 69.5804 | 25.5318 | 2.6214 | 42.0245 | 110.1635 | 16.0312 | 21.3935 | 6.3821 | 179.744 | 326.0055 | 251.5601 | 320.2793 | 222.4823
2012-10-30 02:08 | 25.7687 | 347.3414 | 273.0268 | 341.6265 | 243.6168 | 92.585 | 1.6536 | 42.2225 | 69.8208 | 25.533 | 2.6258 | 42.0964 | 110.537 | 16.0318 | 21.4014 | 6.3828 | 180.3578 | 325.94 | 251.6254 | 320.2251 | 222.2154
2012-10-30 02:08 | 25.7829 | 347.4012 | 273.0681 | 341.7016 | 242.6591 | 92.5954 | 1.6542 | 42.2364 | 69.8668 | 25.5331 | 2.6267 | 42.1105 | 110.6096 | 16.032 | 21.3995 | 6.3835 | 180.4763 | 326.0017 | 251.6686 | 320.3021 | 221.2596
2012-10-30 02:08 | 25.7769 | 347.4219 | 272.9148 | 341.6804 | 243.1116 | 92.5885 | 1.6546 | 42.2467 | 69.9012 | 25.5329 | 2.6273 | 42.1206 | 110.6619 | 16.0321 | 21.3975 | 6.3831 | 180.5631 | 326.0244 | 251.5173 | 320.2829 | 221.7141
2012-10-30 02:08 | 25.7739 | 347.358 | 272.8187 | 341.6203 | 242.5599 | 92.596 | 1.6549 | 42.2546 | 69.9265 | 25.5333 | 2.6277 | 42.1282 | 110.7023 | 16.0321 | 21.3785 | 6.3836 | 180.6288 | 325.9795 | 251.4402 | 320.2418 | 221.1814
2012-10-30 02:08 | 25.7709 | 347.3588 | 272.711 | 341.6211 | 242.945 | 92.5896 | 1.6551 | 42.2606 | 69.9463 | 25.5333 | 2.6281 | 42.1345 | 110.7343 | 16.0323 | 21.3996 | 6.3831 | 180.6805 | 325.9592 | 251.3114 | 320.2215 | 221.5454
2012-10-30 02:08 | 25.7634 | 347.317 | 272.6803 | 341.6327 | 243.1373 | 92.5919 | 1.6552 | 42.2633 | 69.9547 | 25.5335 | 2.6283 | 42.1371 | 110.7481 | 16.0322 | 21.384 | 6.3833 | 180.7028 | 325.933 | 251.2963 | 320.2487 | 221.7533
2012-10-30 02:08 | 25.7739 | 347.3656 | 272.7373 | 341.6546 | 243.0224 | 92.6012 | 1.6552 | 42.2642 | 69.9572 | 25.5336 | 2.6283 | 42.1379 | 110.7511 | 16.0324 | 21.3624 | 6.3839 | 180.7084 | 326.0032 | 251.3749 | 320.2922 | 221.66
2012-10-30 02:08 | 25.7754 | 347.3595 | 272.7155 | 341.6408 | 243.071 | 92.5902 | 1.6552 | 42.264 | 69.9565 | 25.5337 | 2.6283 | 42.1376 | 110.7497 | 16.0324 | 21.392 | 6.3832 | 180.7062 | 325.9675 | 251.3235 | 320.2488 | 221.679
2012-10-30 02:08 | 25.7694 | 347.3573 | 272.6979 | 341.6158 | 243.2687 | 92.5862 | 1.6552 | 42.2633 | 69.9544 | 25.5336 | 2.6282 | 42.1371 | 110.7463 | 16.0325 | 21.3379 | 6.3829 | 180.7007 | 326.0194 | 251.36 | 320.2779 | 221.9308
2012-10-30 02:08 | 25.7694 | 347.3573 | 272.7173 | 341.6272 | 243.3589 | 92.5942 | 1.6552 | 42.2626 | 69.9515 | 25.5338 | 2.6282 | 42.1364 | 110.7422 | 16.0325 | 21.378 | 6.3834 | 180.6937 | 325.9793 | 251.3393 | 320.2492 | 221.9809
2012-10-30 02:08 | 25.7687 | 347.3261 | 272.7786 | 341.5961 | 243.1818 | 92.6052 | 1.6551 | 42.2608 | 69.9457 | 25.5338 | 2.6281 | 42.1347 | 110.7329 | 16.0326 | 21.3653 | 6.3842 | 180.6786 | 325.9608 | 251.4133 | 320.2308 | 221.8165
2012-10-30 02:09 | 25.7604 | 347.3521 | 272.7471 | 341.584 | 243.1187 | 92.6093 | 1.6551 | 42.2597 | 69.942 | 25.5338 | 2.628 | 42.1336 | 110.7276 | 16.0325 | 21.365 | 6.3845 | 180.6696 | 325.9871 | 251.3821 | 320.219 | 221.7537
2012-10-30 02:09 | 25.7664 | 347.3086 | 272.7259 | 341.5938 | 243.1129 | 92.5983 | 1.655 | 42.2587 | 69.939 | 25.5337 | 2.6279 | 42.1326 | 110.722 | 16.0326 | 21.3228 | 6.3837 | 180.661 | 325.9858 | 251.4031 | 320.271 | 221.7901
2012-10-30 02:09 | 25.7702 | 347.3276 | 272.722 | 341.6242 | 243.2617 | 92.611 | 1.655 | 42.2579 | 69.9356 | 25.534 | 2.6279 | 42.1316 | 110.7161 | 16.0326 | 21.3587 | 6.3846 | 180.6517 | 325.9689 | 251.3633 | 320.2655 | 221.903
2012-10-30 02:09 | 25.7687 | 347.349 | 272.7786 | 341.5885 | 243.272 | 92.6035 | 1.6549 | 42.2564 | 69.9301 | 25.5341 | 2.6278 | 42.1301 | 110.7086 | 16.0326 | 21.3612 | 6.3841 | 180.6387 | 325.9878 | 251.4174 | 320.2273 | 221.9108
2012-10-30 02:09 | 25.7739 | 347.3199 | 272.7257 | 341.6241 | 243.2928 | 92.6122 | 1.6549 | 42.2558 | 69.9278 | 25.5342 | 2.6278 | 42.1297 | 110.7067 | 16.0325 | 21.3624 | 6.3847 | 180.6345 | 325.9575 | 251.3633 | 320.2617 | 221.9304
2012-10-30 02:09 | 25.7732 | 347.323 | 272.7482 | 341.6044 | 243.3862 | 92.6099 | 1.6548 | 42.2538 | 69.9214 | 25.5341 | 2.6276 | 42.1278 | 110.6963 | 16.0326 | 21.3336 | 6.3845 | 180.6176 | 325.9894 | 251.4146 | 320.2708 | 222.0526
2012-10-30 02:09 | 25.7776 | 347.3389 | 272.7914 | 341.6202 | 243.371 | 92.6006 | 1.6547 | 42.2514 | 69.914 | 25.534 | 2.6275 | 42.1254 | 110.6846 | 16.0325 | 21.3621 | 6.3839 | 180.5985 | 325.9768 | 251.4293 | 320.2581 | 222.0089
2012-10-30 02:09 | 25.7836 | 347.3068 | 272.809 | 341.6072 | 243.0596 | 92.6052 | 1.6546 | 42.2501 | 69.909 | 25.5342 | 2.6274 | 42.124 | 110.677 | 16.0325 | 21.316 | 6.3842 | 180.586 | 325.9908 | 251.493 | 320.2912 | 221.7436
2012-10-30 02:09 | 25.7806 | 347.3114 | 272.7362 | 341.6346 | 242.8959 | 92.6116 | 1.6546 | 42.2489 | 69.9053 | 25.5341 | 2.6273 | 42.1227 | 110.6705 | 16.0325 | 21.341 | 6.3846 | 180.5757 | 325.9704 | 251.3952 | 320.2936 | 221.5549
2012-10-30 02:09 | 25.7791 | 347.348 | 272.7154 | 341.6179 | 243.2549 | 92.6156 | 1.6546 | 42.2476 | 69.9013 | 25.534 | 2.6273 | 42.1214 | 110.6636 | 16.0325 | 21.3596 | 6.3849 | 180.5649 | 325.9884 | 251.3558 | 320.2583 | 221.8953
2012-10-30 02:09 | 25.7821 | 347.3395 | 272.7494 | 341.6209 | 243.352 | 92.6058 | 1.6545 | 42.2456 | 69.8943 | 25.5342 | 2.6272 | 42.1197 | 110.6547 | 16.0324 | 21.3425 | 6.3842 | 180.549 | 325.997 | 251.4069 | 320.2784 | 222.0095
2012-10-30 02:09 | 25.7814 | 347.3693 | 272.7836 | 341.6354 | 243.2925 | 92.6179 | 1.6544 | 42.2436 | 69.8885 | 25.5339 | 2.627 | 42.1174 | 110.6436 | 16.0323 | 21.3338 | 6.3851 | 180.5321 | 326.0355 | 251.4498 | 320.3016 | 221.9587
2012-10-30 02:09 | 25.7761 | 347.3982 | 272.7434 | 341.6454 | 243.0011 | 92.6064 | 1.6543 | 42.2417 | 69.8817 | 25.534 | 2.6269 | 42.1155 | 110.6333 | 16.0324 | 21.3325 | 6.3843 | 180.515 | 326.0657 | 251.4109 | 320.3129 | 221.6686
2012-10-30 02:09 | 25.7881 | 347.357 | 272.7593 | 341.6574 | 243.1503 | 92.6104 | 1.6543 | 42.2398 | 69.8755 | 25.534 | 2.6268 | 42.1135 | 110.6237 | 16.0323 | 21.3566 | 6.3846 | 180.4992 | 326.0004 | 251.4027 | 320.3008 | 221.7937
2012-10-30 02:09 | 25.7889 | 347.3616 | 272.7252 | 341.6658 | 243.3118 | 92.6237 | 1.6542 | 42.2382 | 69.8705 | 25.5339 | 2.6267 | 42.1123 | 110.6175 | 16.0323 | 21.3453 | 6.3855 | 180.488 | 326.0163 | 251.3799 | 320.3205 | 221.9665
2012-10-30 02:09 | 25.7844 | 347.3723 | 272.7749 | 341.6537 | 243.5933 | 92.6139 | 1.6541 | 42.2364 | 69.8644 | 25.5339 | 2.6266 | 42.1103 | 110.607 | 16.0322 | 21.3127 | 6.3848 | 180.4714 | 326.0596 | 251.4622 | 320.341 | 222.2806
2012-10-30 02:09 | 25.7926 | 347.3805 | 272.8219 | 341.6428 | 243.5623 | 92.6006 | 1.654 | 42.2343 | 69.8572 | 25.5341 | 2.6265 | 42.1082 | 110.5958 | 16.0323 | 21.3089 | 6.3839 | 180.453 | 326.0716 | 251.513 | 320.3339 | 222.2534
2012-10-30 02:09 | 25.7881 | 347.3303 | 272.8368 | 341.6535 | 242.97 | 92.6376 | 1.654 | 42.2329 | 69.8529 | 25.5339 | 2.6264 | 42.1069 | 110.589 | 16.0322 | 21.2924 | 6.3864 | 180.4419 | 326.0379 | 251.5444 | 320.3611 | 221.6776
2012-10-30 02:09 | 25.7859 | 347.3738 | 272.7958 | 341.6552 | 243.152 | 92.6168 | 1.6539 | 42.2317 | 69.8484 | 25.5341 | 2.6263 | 42.1057 | 110.5834 | 16.0322 | 21.3022 | 6.385 | 180.4318 | 326.0716 | 251.4936 | 320.353 | 221.8498
2012-10-30 02:09 | 25.7904 | 347.3402 | 272.8469 | 341.6406 | 243.2819 | 92.6127 | 1.6539 | 42.23 | 69.843 | 25.534 | 2.6262 | 42.1039 | 110.5737 | 16.0322 | 21.3027 | 6.3847 | 180.4168 | 326.0375 | 251.5442 | 320.3379 | 221.9792
2012-10-30 02:09 | 25.7934 | 347.347 | 272.8848 | 341.6284 | 243.3868 | 92.6168 | 1.6517 | 42.1732 | 69.6567 | 25.5334 | 2.6229 | 42.0505 | 110.2958 | 16.0319 | 21.3217 | 6.385 | 179.9525 | 326.0253 | 251.5631 | 320.3067 | 222.0651
2012-10-30 02:09 | 25.7956 | 347.3416 | 272.8831 | 341.6687 | 243.2009 | 92.6041 | 1.6501 | 42.131 | 69.5196 | 25.5327 | 2.6202 | 42.0051 | 110.061 | 16.0313 | 21.348 | 6.3841 | 179.5806 | 325.9936 | 251.5351 | 320.3207 | 221.8529
2012-10-30 02:10 | 25.8031 | 347.3377 | 272.8712 | 341.6571 | 243.4278 | 92.6023 | 1.6496 | 42.1186 | 69.4797 | 25.5322 | 2.6194 | 41.9929 | 109.9982 | 16.0312 | 21.3114 | 6.384 | 179.4779 | 326.0263 | 251.5598 | 320.3457 | 222.1164
2012-10-30 02:10 | 25.7964 | 347.35 | 272.9033 | 341.6695 | 243.1272 | 92.589 | 1.6494 | 42.1126 | 69.4604 | 25.5321 | 2.6191 | 41.9868 | 109.9679 | 16.031 | 21.3448 | 6.3831 | 179.4283 | 326.0052 | 251.5585 | 320.3247 | 221.7824
2012-10-30 02:10 | 25.7949 | 347.3029 | 272.7854 | 341.6565 | 242.6162 | 92.5948 | 1.6495 | 42.1138 | 69.4648 | 25.5319 | 2.6192 | 41.9882 | 109.9745 | 16.031 | 21.3032 | 6.3835 | 179.4394 | 325.9997 | 251.4822 | 320.3533 | 221.313
2012-10-30 02:10 | 25.8023 | 347.2989 | 272.6377 | 341.6335 | 241.8905 | 92.5804 | 1.6496 | 42.1174 | 69.4758 | 25.5323 | 2.6194 | 41.9915 | 109.9924 | 16.031 | 21.2745 | 6.3825 | 179.4681 | 326.0244 | 251.3632 | 320.359 | 220.616
2012-10-30 02:10 | 25.7926 | 347.2663 | 272.4884 | 341.6504 | 241.8299 | 92.5966 | 1.6497 | 42.1211 | 69.4887 | 25.532 | 2.6196 | 41.9953 | 110.0123 | 16.031 | 21.3129 | 6.3836 | 179.501 | 325.9534 | 251.1755 | 320.3375 | 220.517
2012-10-30 02:10 | 25.7896 | 347.29 | 272.4156 | 341.6284 | 242.5795 | 92.5867 | 1.65 | 42.1277 | 69.5095 | 25.5324 | 2.62 | 42.0017 | 110.046 | 16.031 | 21.2979 | 6.3829 | 179.5555 | 325.9921 | 251.1177 | 320.3305 | 221.2816
2012-10-30 02:10 | 25.7934 | 347.271 | 272.3728 | 341.617 | 242.9635 | 92.57 | 1.6502 | 42.1345 | 69.532 | 25.5324 | 2.6204 | 42.0088 | 110.0818 | 16.0311 | 21.3699 | 6.3818 | 179.6138 | 325.9011 | 251.0029 | 320.2471 | 221.5936
2012-10-30 02:10 | 25.7829 | 347.2681 | 272.4205 | 341.6407 | 242.1691 | 92.5786 | 1.6506 | 42.1424 | 69.5582 | 25.5323 | 2.6209 | 42.0167 | 110.123 | 16.0312 | 21.3112 | 6.3824 | 179.6812 | 325.9569 | 251.1093 | 320.3295 | 220.8579
2012-10-30 02:10 | 25.7911 | 347.242 | 272.3007 | 341.6413 | 242.1106 | 92.5723 | 1.6536 | 42.2207 | 69.8142 | 25.5333 | 2.6256 | 42.0932 | 110.5211 | 16.0317 | 21.3395 | 6.3819 | 180.3354 | 325.9025 | 250.9612 | 320.3018 | 220.7711
2012-10-30 02:10 | 25.7844 | 347.2924 | 272.1893 | 341.6118 | 240.7708 | 92.5919 | 1.654 | 42.2326 | 69.8527 | 25.5337 | 2.6264 | 42.1066 | 110.59 | 16.0319 | 21.2927 | 6.3833 | 180.4427 | 325.9997 | 250.8966 | 320.3191 | 219.4781
2012-10-30 02:10 | 25.7739 | 347.2819 | 272.0624 | 341.567 | 241.5798 | 92.5879 | 1.6545 | 42.2458 | 69.8952 | 25.5341 | 2.6272 | 42.1198 | 110.6572 | 16.0322 | 21.3183 | 6.383 | 180.5524 | 325.9636 | 250.7441 | 320.2487 | 220.2615
2012-10-30 02:10 | 25.7776 | 347.3122 | 272.0778 | 341.5936 | 242.2422 | 92.5804 | 1.6549 | 42.2561 | 69.9283 | 25.5344 | 2.6278 | 42.1297 | 110.7083 | 16.0323 | 21.3059 | 6.3825 | 180.6366 | 326.0063 | 250.7719 | 320.2877 | 220.9363
2012-10-30 02:10 | 25.7769 | 347.3496 | 272.0616 | 341.551 | 242.5629 | 92.5856 | 1.655 | 42.2587 | 69.9374 | 25.5342 | 2.628 | 42.1326 | 110.723 | 16.0324 | 21.2812 | 6.3829 | 180.6603 | 326.0684 | 250.7804 | 320.2698 | 221.2817
2012-10-30 02:10 | 25.7717 | 347.3406 | 272.1495 | 341.5572 | 242.8243 | 92.6018 | 1.6551 | 42.2613 | 69.945 | 25.5346 | 2.6281 | 42.135 | 110.735 | 16.0325 | 21.3081 | 6.384 | 180.68 | 326.0325 | 250.8414 | 320.2491 | 221.5162
2012-10-30 02:10 | 25.7649 | 347.3033 | 272.2357 | 341.5733 | 242.8958 | 92.5948 | 1.655 | 42.2595 | 69.939 | 25.5346 | 2.628 | 42.133 | 110.7237 | 16.0326 | 21.2772 | 6.3835 | 180.6626 | 326.0261 | 250.9585 | 320.2961 | 221.6186
2012-10-30 02:10 | 25.7634 | 347.3399 | 272.2847 | 341.5527 | 242.8787 | 92.5995 | 1.655 | 42.2602 | 69.941 | 25.5347 | 2.628 | 42.134 | 110.7296 | 16.0325 | 21.3158 | 6.3838 | 180.6706 | 326.0241 | 250.9689 | 320.2369 | 221.5629
2012-10-30 02:10 | 25.7447 | 347.306 | 272.3164 | 341.5378 | 243.0207 | 92.6058 | 1.6551 | 42.2611 | 69.9445 | 25.5345 | 2.6281 | 42.1348 | 110.7335 | 16.0326 | 21.2971 | 6.3842 | 180.6779 | 326.0089 | 251.0193 | 320.2407 | 221.7236
2012-10-30 02:10 | 25.7559 | 347.3476 | 272.3586 | 341.5567 | 242.9652 | 92.5966 | 1.6551 | 42.2617 | 69.9458 | 25.5348 | 2.6281 | 42.1354 | 110.7359 | 16.0327 | 21.2963 | 6.3836 | 180.6817 | 326.0513 | 251.0623 | 320.2604 | 221.6689
2012-10-30 02:10 | 25.7582 | 347.3271 | 272.3997 | 341.5399 | 243.1165 | 92.6012 | 1.6551 | 42.2615 | 69.946 | 25.5345 | 2.6281 | 42.1353 | 110.7352 | 16.0327 | 21.3347 | 6.3839 | 180.6812 | 325.9924 | 251.065 | 320.2052 | 221.7818
2012-10-30 02:10 | 25.7582 | 347.3195 | 272.4618 | 341.578 | 243.14 | 92.6151 | 1.655 | 42.2605 | 69.9418 | 25.5348 | 2.628 | 42.1343 | 110.7298 | 16.0327 | 21.3588 | 6.3849 | 180.6716 | 325.9607 | 251.103 | 320.2192 | 221.7812
2012-10-30 02:10 | 25.7574 | 347.3148 | 272.5114 | 341.5962 | 243.0804 | 92.6081 | 1.655 | 42.2594 | 69.9376 | 25.535 | 2.628 | 42.1332 | 110.7247 | 16.0326 | 21.3018 | 6.3844 | 180.6623 | 326.013 | 251.2096 | 320.2944 | 221.7786
2012-10-30 02:10 | 25.7492 | 347.3371 | 272.5497 | 341.569 | 243.2564 | 92.611 | 1.6549 | 42.2576 | 69.9322 | 25.5349 | 2.6279 | 42.1316 | 110.7156 | 16.0327 | 21.3016 | 6.3846 | 180.6478 | 326.0355 | 251.2481 | 320.2674 | 221.9548
2012-10-30 02:10 | 25.7649 | 347.3756 | 272.6042 | 341.5733 | 243.2329 | 92.6087 | 1.6548 | 42.2559 | 69.927 | 25.5346 | 2.6277 | 42.1296 | 110.7055 | 16.0326 | 21.2932 | 6.3844 | 180.6325 | 326.0824 | 251.311 | 320.2801 | 221.9397
2012-10-30 02:11 | 25.7589 | 347.3582 | 272.6602 | 341.5711 | 243.2426 | 92.6208 | 1.6548 | 42.2547 | 69.9229 | 25.5347 | 2.6277 | 42.1286 | 110.7007 | 16.0326 | 21.2913 | 6.3853 | 180.6236 | 326.0669 | 251.3689 | 320.2798 | 221.9513
2012-10-30 02:11 | 25.7537 | 347.372 | 272.6318 | 341.6268 | 243.257 | 92.6191 | 1.6548 | 42.2544 | 69.9223 | 25.5346 | 2.6277 | 42.1282 | 110.6987 | 16.0326 | 21.282 | 6.3852 | 180.621 | 326.09 | 251.3498 | 320.3448 | 221.975
2012-10-30 02:11 | 25.7574 | 347.3491 | 272.6277 | 341.6381 | 243.2176 | 92.6064 | 1.6547 | 42.2529 | 69.9164 | 25.5348 | 2.6276 | 42.1266 | 110.6907 | 16.0326 | 21.2376 | 6.3843 | 180.6071 | 326.1115 | 251.3901 | 320.4005 | 221.98
2012-10-30 02:11 | 25.7634 | 347.4046 | 272.6958 | 341.6174 | 243.3764 | 92.6162 | 1.6547 | 42.2511 | 69.9112 | 25.5347 | 2.6275 | 42.1252 | 110.6834 | 16.0325 | 21.2436 | 6.385 | 180.5945 | 326.161 | 251.4522 | 320.3738 | 222.1328
2012-10-30 02:11 | 25.7612 | 347.4138 | 272.7479 | 341.6228 | 243.4095 | 92.6174 | 1.6546 | 42.2498 | 69.9064 | 25.5348 | 2.6274 | 42.1236 | 110.6752 | 16.0325 | 21.2414 | 6.385 | 180.5815 | 326.1724 | 251.5065 | 320.3814 | 222.1681
2012-10-30 02:11 | 25.7642 | 347.413 | 272.7431 | 341.6335 | 243.4948 | 92.6139 | 1.6545 | 42.2485 | 69.9018 | 25.5349 | 2.6273 | 42.1226 | 110.6691 | 16.0326 | 21.2203 | 6.3848 | 180.5708 | 326.1927 | 251.5228 | 320.4132 | 222.2745
2012-10-30 02:11 | 25.7679 | 347.4357 | 272.8205 | 341.6295 | 243.4906 | 92.6064 | 1.6545 | 42.2472 | 69.8971 | 25.5351 | 2.6272 | 42.121 | 110.661 | 16.0326 | 21.2401 | 6.3843 | 180.5581 | 326.1956 | 251.5804 | 320.3894 | 222.2505
2012-10-30 02:11 | 25.7732 | 347.4144 | 272.8297 | 341.6539 | 243.4567 | 92.6151 | 1.6544 | 42.2449 | 69.8899 | 25.5349 | 2.6271 | 42.1188 | 110.6489 | 16.0327 | 21.2373 | 6.3849 | 180.5389 | 326.1771 | 251.5924 | 320.4166 | 222.2194
2012-10-30 02:11 | 25.7627 | 347.4685 | 272.8579 | 341.6548 | 243.3052 | 92.6162 | 1.6544 | 42.2453 | 69.8909 | 25.535 | 2.6271 | 42.1193 | 110.6517 | 16.0326 | 21.2308 | 6.385 | 180.5426 | 326.2377 | 251.6271 | 320.424 | 222.0744
2012-10-30 02:11 | 25.7702 | 347.4608 | 272.8848 | 341.6547 | 243.344 | 92.6226 | 1.6543 | 42.2437 | 69.8855 | 25.535 | 2.627 | 42.1173 | 110.6413 | 16.0326 | 21.2343 | 6.3854 | 180.5268 | 326.2265 | 251.6505 | 320.4204 | 222.1097
2012-10-30 02:11 | 25.7754 | 347.4584 | 272.8396 | 341.7055 | 243.553 | 92.6179 | 1.6542 | 42.2411 | 69.8766 | 25.5351 | 2.6268 | 42.1149 | 110.628 | 16.0327 | 21.2154 | 6.3851 | 180.5046 | 326.243 | 251.6242 | 320.4901 | 222.3376
2012-10-30 02:11 | 25.7754 | 347.4698 | 272.9404 | 341.6789 | 243.6314 | 92.611 | 1.6542 | 42.2394 | 69.8712 | 25.5351 | 2.6267 | 42.1132 | 110.6194 | 16.0326 | 21.2636 | 6.3846 | 180.4906 | 326.2062 | 251.6768 | 320.4153 | 222.3678
2012-10-30 02:11 | 25.7784 | 347.469 | 272.9008 | 341.7009 | 243.7088 | 92.6214 | 1.6541 | 42.238 | 69.8666 | 25.5351 | 2.6266 | 42.1118 | 110.6116 | 16.0327 | 21.2506 | 6.3853 | 180.4782 | 326.2184 | 251.6502 | 320.4503 | 222.4582
2012-10-30 02:11 | 25.7836 | 347.459 | 272.9991 | 341.6947 | 243.667 | 92.6168 | 1.654 | 42.2355 | 69.8583 | 25.5351 | 2.6265 | 42.1095 | 110.5989 | 16.0328 | 21.2517 | 6.385 | 180.4573 | 326.2073 | 251.7474 | 320.443 | 222.4153
2012-10-30 02:11 | 25.7821 | 347.4385 | 273.0208 | 341.7008 | 243.7556 | 92.6203 | 1.654 | 42.2355 | 69.8577 | 25.5353 | 2.6265 | 42.1093 | 110.5988 | 16.0327 | 21.2703 | 6.3852 | 180.4565 | 326.1682 | 251.7505 | 320.4305 | 222.4853
2012-10-30 02:11 | 25.7784 | 347.4652 | 273.0365 | 341.6857 | 243.7872 | 92.6099 | 1.654 | 42.2343 | 69.854 | 25.5353 | 2.6264 | 42.1082 | 110.5923 | 16.0328 | 21.2586 | 6.3845 | 180.4463 | 326.2066 | 251.7779 | 320.4271 | 222.5286
2012-10-30 02:11 | 25.7791 | 347.4697 | 273.0837 | 341.7168 | 243.6076 | 92.6168 | 1.6539 | 42.2333 | 69.8498 | 25.5355 | 2.6263 | 42.1071 | 110.5867 | 16.0328 | 21.2753 | 6.385 | 180.4365 | 326.1944 | 251.8084 | 320.4415 | 222.3323
2012-10-30 02:11 | 25.7784 | 347.4804 | 273.0908 | 341.758 | 243.6735 | 92.6226 | 1.6538 | 42.2307 | 69.842 | 25.5352 | 2.6262 | 42.1048 | 110.5746 | 16.0327 | 21.2787 | 6.3854 | 180.4167 | 326.2017 | 251.8121 | 320.4793 | 222.3948
2012-10-30 02:11 | 25.7761 | 347.4895 | 273.135 | 341.7443 | 243.6125 | 92.6133 | 1.6538 | 42.2299 | 69.8389 | 25.5354 | 2.6261 | 42.1037 | 110.5677 | 16.0329 | 21.2924 | 6.3848 | 180.4067 | 326.1971 | 251.8426 | 320.4519 | 222.3201
2012-10-30 02:11 | 25.7904 | 347.4962 | 273.1183 | 341.7662 | 243.7757 | 92.6035 | 1.6502 | 42.1377 | 69.5377 | 25.5342 | 2.6205 | 42.0116 | 110.0903 | 16.032 | 21.3067 | 6.3841 | 179.628 | 326.1895 | 251.8116 | 320.4595 | 222.469
2012-10-30 02:11 | 25.7851 | 347.5061 | 273.144 | 341.7723 | 243.9271 | 92.611 | 1.6498 | 42.1246 | 69.496 | 25.5336 | 2.6197 | 41.9989 | 110.0259 | 16.0318 | 21.3335 | 6.3846 | 179.5219 | 326.1726 | 251.8105 | 320.4388 | 222.5936
2012-10-30 02:11 | 25.7866 | 347.4886 | 273.1727 | 341.7282 | 243.5015 | 92.5995 | 1.6494 | 42.1152 | 69.4647 | 25.5336 | 2.6192 | 41.9896 | 109.9781 | 16.0316 | 21.339 | 6.3838 | 179.4428 | 326.1496 | 251.8337 | 320.3892 | 222.1625
2012-10-30 02:11 | 25.7844 | 347.4826 | 273.1084 | 341.7488 | 242.3901 | 92.607 | 1.6494 | 42.1152 | 69.4652 | 25.5335 | 2.6192 | 41.9895 | 109.9782 | 16.0315 | 21.405 | 6.3843 | 179.4433 | 326.0776 | 251.7034 | 320.3438 | 220.9851
2012-10-30 02:11 | 25.7844 | 347.4788 | 272.9805 | 341.783 | 242.8957 | 92.6052 | 1.6495 | 42.1169 | 69.4712 | 25.5334 | 2.6193 | 41.9912 | 109.9868 | 16.0316 | 21.401 | 6.3842 | 179.458 | 326.0778 | 251.5795 | 320.382 | 221.4947
2012-10-30 02:11 | 25.7717 | 347.4357 | 272.8786 | 341.7665 | 243.0163 | 92.5971 | 1.6496 | 42.1199 | 69.4807 | 25.5335 | 2.6195 | 41.994 | 110.0016 | 16.0316 | 21.3843 | 6.3836 | 179.4824 | 326.0514 | 251.4943 | 320.3822 | 221.632
2012-10-30 02:12 | 25.7747 | 347.4197 | 272.9048 | 341.7657 | 243.8501 | 92.5989 | 1.6497 | 42.1234 | 69.492 | 25.5335 | 2.6197 | 41.9974 | 110.0195 | 16.0315 | 21.3713 | 6.3838 | 179.5115 | 326.0484 | 251.5335 | 320.3944 | 222.4788
2012-10-30 02:12 | 25.7649 | 347.4365 | 272.9842 | 341.7826 | 243.9147 | 92.5948 | 1.6499 | 42.1286 | 69.5094 | 25.5335 | 2.62 | 42.0025 | 110.0464 | 16.0315 | 21.3895 | 6.3835 | 179.5558 | 326.047 | 251.5947 | 320.3931 | 222.5252
2012-10-30 02:12 | 25.7582 | 347.445 | 272.993 | 341.772 | 243.3516 | 92.5885 | 1.6501 | 42.1332 | 69.524 | 25.5337 | 2.6203 | 42.0076 | 110.0726 | 16.0316 | 21.3909 | 6.3831 | 179.5966 | 326.0541 | 251.6021 | 320.3811 | 221.9607
2012-10-30 02:12 | 25.7537 | 347.4557 | 272.9032 | 341.7523 | 242.1164 | 92.589 | 1.6503 | 42.1376 | 69.5387 | 25.5336 | 2.6206 | 42.0118 | 110.0944 | 16.0316 | 21.4265 | 6.3831 | 179.633 | 326.0292 | 251.4767 | 320.3258 | 220.6899
2012-10-30 02:12 | 25.7597 | 347.4465 | 272.7658 | 341.7203 | 242.8044 | 92.585 | 1.6524 | 42.1926 | 69.7182 | 25.5344 | 2.624 | 42.0687 | 110.3896 | 16.032 | 21.4124 | 6.3828 | 180.1079 | 326.0341 | 251.3534 | 320.3079 | 221.392
2012-10-30 02:12 | 25.7455 | 347.3943 | 272.7438 | 341.767 | 242.5002 | 92.6029 | 1.6537 | 42.228 | 69.834 | 25.5349 | 2.6261 | 42.1021 | 110.5629 | 16.0324 | 21.4143 | 6.384 | 180.3969 | 325.98 | 251.3295 | 320.3527 | 221.0859
2012-10-30 02:12 | 25.741 | 347.4164 | 272.6191 | 341.7206 | 242.527 | 92.6052 | 1.6541 | 42.2387 | 69.8684 | 25.5353 | 2.6267 | 42.1127 | 110.6179 | 16.0325 | 21.4338 | 6.3842 | 180.4864 | 325.9826 | 251.1853 | 320.2868 | 221.0932
2012-10-30 02:12 | 25.744 | 347.4156 | 272.657 | 341.735 | 243.1689 | 92.5838 | 1.6545 | 42.2479 | 69.8992 | 25.5352 | 2.6273 | 42.1218 | 110.6655 | 16.0325 | 21.4288 | 6.3827 | 180.5647 | 325.9868 | 251.2282 | 320.3062 | 221.7401
2012-10-30 02:12 | 25.7417 | 347.3981 | 272.6624 | 341.7099 | 243.2959 | 92.5902 | 1.6548 | 42.2562 | 69.9258 | 25.5354 | 2.6278 | 42.1299 | 110.7076 | 16.0325 | 21.4466 | 6.3832 | 180.6334 | 325.9515 | 251.2158 | 320.2633 | 221.8493
2012-10-30 02:12 | 25.738 | 347.421 | 272.7518 | 341.691 | 243.351 | 92.5971 | 1.655 | 42.2606 | 69.9401 | 25.5355 | 2.628 | 42.1344 | 110.7301 | 16.0327 | 21.4228 | 6.3836 | 180.6702 | 325.9982 | 251.329 | 320.2682 | 221.9282
2012-10-30 02:12 | 25.7298 | 347.4128 | 272.7746 | 341.7246 | 243.3545 | 92.611 | 1.655 | 42.2618 | 69.9434 | 25.5358 | 2.6281 | 42.1354 | 110.7348 | 16.0328 | 21.3946 | 6.3846 | 180.6781 | 326.0182 | 251.38 | 320.33 | 221.9599
2012-10-30 02:12 | 25.7245 | 347.4075 | 272.7926 | 341.6927 | 243.4119 | 92.6035 | 1.655 | 42.2611 | 69.9413 | 25.5357 | 2.628 | 42.1347 | 110.7307 | 16.0329 | 21.4133 | 6.3841 | 180.6719 | 325.9942 | 251.3793 | 320.2794 | 221.9986
2012-10-30 02:12 | 25.723 | 347.4402 | 272.8493 | 341.7102 | 243.4927 | 92.6075 | 1.6549 | 42.2595 | 69.9367 | 25.5354 | 2.6279 | 42.1332 | 110.7236 | 16.0328 | 21.3597 | 6.3844 | 180.6602 | 326.0805 | 251.4896 | 320.3505 | 222.133
2012-10-30 02:12 | 25.729 | 347.3968 | 272.863 | 341.7467 | 243.5105 | 92.6047 | 1.6549 | 42.2593 | 69.9353 | 25.5357 | 2.6279 | 42.1329 | 110.722 | 16.0328 | 21.4058 | 6.3842 | 180.6573 | 325.991 | 251.4572 | 320.3409 | 222.1047
2012-10-30 02:12 | 25.7387 | 347.4179 | 272.9076 | 341.7335 | 243.5907 | 92.6023 | 1.6549 | 42.2589 | 69.9349 | 25.5355 | 2.6279 | 42.133 | 110.7211 | 16.033 | 21.4115 | 6.384 | 180.656 | 326.0064 | 251.4961 | 320.322 | 222.1792
2012-10-30 02:12 | 25.7208 | 347.381 | 272.9091 | 341.6966 | 243.4905 | 92.6023 | 1.6549 | 42.2589 | 69.9332 | 25.536 | 2.6279 | 42.1328 | 110.7198 | 16.033 | 21.3815 | 6.384 | 180.653 | 325.9995 | 251.5276 | 320.3151 | 222.109
2012-10-30 02:12 | 25.7305 | 347.3831 | 272.9033 | 341.7329 | 243.5394 | 92.6127 | 1.6549 | 42.2582 | 69.9312 | 25.5359 | 2.6279 | 42.1322 | 110.7173 | 16.0329 | 21.3752 | 6.3847 | 180.6484 | 326.0079 | 251.5281 | 320.3577 | 222.1642
2012-10-30 02:12 | 25.7245 | 347.3657 | 272.9283 | 341.7079 | 243.6471 | 92.6058 | 1.6548 | 42.2566 | 69.9252 | 25.5361 | 2.6277 | 42.1304 | 110.7065 | 16.0331 | 21.3933 | 6.3842 | 180.6318 | 325.9724 | 251.535 | 320.3146 | 222.2538
2012-10-30 02:12 | 25.7238 | 347.3916 | 272.9625 | 341.7034 | 243.8423 | 92.6099 | 1.6547 | 42.2551 | 69.9207 | 25.536 | 2.6276 | 42.129 | 110.7001 | 16.033 | 21.3765 | 6.3845 | 180.6208 | 326.0151 | 251.586 | 320.3269 | 222.4658
2012-10-30 02:12 | 25.7283 | 347.358 | 273.0407 | 341.7155 | 243.6665 | 92.6162 | 1.6546 | 42.2529 | 69.9132 | 25.5361 | 2.6275 | 42.1269 | 110.689 | 16.033 | 21.4091 | 6.385 | 180.6022 | 325.9489 | 251.6316 | 320.3064 | 222.2574
2012-10-30 02:12 | 25.7298 | 347.3976 | 273.0267 | 341.7627 | 243.6641 | 92.6203 | 1.6546 | 42.2517 | 69.9096 | 25.5359 | 2.6274 | 42.1255 | 110.6818 | 16.033 | 21.3946 | 6.3852 | 180.5915 | 326.003 | 251.6321 | 320.3681 | 222.2695
2012-10-30 02:12 | 25.7268 | 347.3984 | 273.0159 | 341.7216 | 243.5318 | 92.6214 | 1.6545 | 42.2506 | 69.9052 | 25.5362 | 2.6274 | 42.1245 | 110.6765 | 16.033 | 21.3193 | 6.3853 | 180.5817 | 326.0791 | 251.6966 | 320.4023 | 222.2125
2012-10-30 02:12 | 25.7298 | 347.3671 | 273.0422 | 341.7627 | 243.6249 | 92.622 | 1.6545 | 42.2493 | 69.9007 | 25.5363 | 2.6273 | 42.1231 | 110.6692 | 16.033 | 21.3745 | 6.3854 | 180.57 | 325.9926 | 251.6677 | 320.3882 | 222.2504
2012-10-30 02:13 | 25.7253 | 347.4235 | 273.0881 | 341.7544 | 243.6675 | 92.6307 | 1.6544 | 42.2481 | 69.8974 | 25.5361 | 2.6272 | 42.1219 | 110.6634 | 16.0329 | 21.3379 | 6.386 | 180.5609 | 326.0856 | 251.7502 | 320.4165 | 222.3296
2012-10-30 02:13 | 25.729 | 347.431 | 273.084 | 341.7505 | 243.7339 | 92.6122 | 1.6544 | 42.2471 | 69.8937 | 25.5362 | 2.6272 | 42.121 | 110.6583 | 16.0329 | 21.4018 | 6.3847 | 180.552 | 326.0292 | 251.6822 | 320.3487 | 222.3321
2012-10-30 02:13 | 25.7305 | 347.4477 | 273.1049 | 341.7253 | 243.6766 | 92.6075 | 1.6544 | 42.2465 | 69.8915 | 25.5362 | 2.6271 | 42.1204 | 110.6559 | 16.0328 | 21.3752 | 6.3844 | 180.5474 | 326.0725 | 251.7297 | 320.3501 | 222.3014
2012-10-30 02:13 | 25.723 | 347.4174 | 273.1013 | 341.7711 | 243.7984 | 92.6283 | 1.6543 | 42.2449 | 69.8873 | 25.5358 | 2.627 | 42.1188 | 110.6475 | 16.0329 | 21.3837 | 6.3858 | 180.5348 | 326.0337 | 251.7176 | 320.3874 | 222.4147
2012-10-30 02:13 | 25.7305 | 347.4097 | 273.1243 | 341.7748 | 243.7589 | 92.6156 | 1.6543 | 42.2432 | 69.8811 | 25.5361 | 2.6269 | 42.117 | 110.6381 | 16.0329 | 21.3672 | 6.3849 | 180.5192 | 326.0425 | 251.7571 | 320.4076 | 222.3917
2012-10-30 02:13 | 25.729 | 347.4044 | 273.1344 | 341.7733 | 243.7926 | 92.6243 | 1.6542 | 42.2415 | 69.8755 | 25.5361 | 2.6268 | 42.1152 | 110.6283 | 16.0329 | 21.3456 | 6.3855 | 180.5038 | 326.0588 | 251.7888 | 320.4277 | 222.447
2012-10-30 02:13 | 25.7365 | 347.4157 | 273.0993 | 341.7884 | 243.8589 | 92.6335 | 1.6541 | 42.2395 | 69.869 | 25.5361 | 2.6267 | 42.1135 | 110.6191 | 16.0329 | 21.3571 | 6.3862 | 180.4881 | 326.0586 | 251.7422 | 320.4313 | 222.5018
2012-10-30 02:13 | 25.7327 | 347.4461 | 273.2273 | 341.7808 | 243.8473 | 92.6266 | 1.654 | 42.2364 | 69.8591 | 25.5359 | 2.6265 | 42.1104 | 110.6031 | 16.0329 | 21.3774 | 6.3857 | 180.4623 | 326.0687 | 251.8499 | 320.4034 | 222.4699
2012-10-30 02:13 | 25.7357 | 347.4225 | 273.1993 | 341.7952 | 243.6818 | 92.6179 | 1.6539 | 42.2357 | 69.8556 | 25.5363 | 2.6265 | 42.1096 | 110.5995 | 16.0328 | 21.3523 | 6.3851 | 180.455 | 326.0702 | 251.847 | 320.4429 | 222.3295
2012-10-30 02:13 | 25.7365 | 347.4575 | 273.2117 | 341.7922 | 243.8667 | 92.6347 | 1.6539 | 42.2339 | 69.8503 | 25.5361 | 2.6264 | 42.1078 | 110.5899 | 16.0328 | 21.3571 | 6.3862 | 180.4401 | 326.1004 | 251.8546 | 320.4351 | 222.5096
2012-10-30 02:13 | 25.7395 | 347.4682 | 273.1915 | 341.8409 | 243.8501 | 92.6289 | 1.6539 | 42.2332 | 69.8481 | 25.536 | 2.6263 | 42.107 | 110.5852 | 16.0329 | 21.3601 | 6.3858 | 180.4333 | 326.1081 | 251.8314 | 320.4808 | 222.49
2012-10-30 02:13 | 25.7425 | 347.4712 | 273.2177 | 341.7792 | 243.6925 | 92.6243 | 1.6538 | 42.2317 | 69.8433 | 25.536 | 2.6262 | 42.1058 | 110.5797 | 16.0327 | 21.3712 | 6.3855 | 180.4231 | 326.1 | 251.8465 | 320.408 | 222.3213
2012-10-30 02:13 | 25.735 | 347.4637 | 273.1831 | 341.8288 | 243.6536 | 92.6255 | 1.6537 | 42.2301 | 69.8376 | 25.5361 | 2.6261 | 42.1038 | 110.5693 | 16.0327 | 21.3757 | 6.3856 | 180.4069 | 326.088 | 251.8074 | 320.4531 | 222.2779
2012-10-30 02:13 | 25.7455 | 347.4932 | 273.2207 | 341.8393 | 243.6485 | 92.6295 | 1.6537 | 42.2291 | 69.8341 | 25.5361 | 2.6261 | 42.1029 | 110.5644 | 16.0327 | 21.3421 | 6.3859 | 180.3985 | 326.1511 | 251.8786 | 320.4972 | 222.3064
2012-10-30 02:13 | 25.738 | 347.4857 | 273.2404 | 341.8546 | 243.7115 | 92.622 | 1.6537 | 42.2281 | 69.8309 | 25.5361 | 2.626 | 42.1021 | 110.5605 | 16.0327 | 21.3586 | 6.3854 | 180.3914 | 326.1271 | 251.8818 | 320.496 | 222.3529
2012-10-30 02:13 | 25.738 | 347.4819 | 273.2171 | 341.8203 | 243.7585 | 92.6272 | 1.6536 | 42.2257 | 69.8224 | 25.5363 | 2.6258 | 42.0996 | 110.5469 | 16.0328 | 21.2984 | 6.3857 | 180.3693 | 326.1835 | 251.9187 | 320.5219 | 222.4601
2012-10-30 02:13 | 25.7417 | 347.5084 | 273.2131 | 341.8431 | 243.5506 | 92.6203 | 1.651 | 42.1585 | 69.6037 | 25.5351 | 2.6218 | 42.033 | 110.2019 | 16.0321 | 21.3423 | 6.3852 | 179.8056 | 326.1661 | 251.8708 | 320.5008 | 222.2083
2012-10-30 02:13 | 25.7395 | 347.491 | 273.0713 | 341.8751 | 243.278 | 92.6081 | 1.6492 | 42.1101 | 69.4465 | 25.5342 | 2.6189 | 41.9844 | 109.9512 | 16.0316 | 21.3601 | 6.3844 | 179.3977 | 326.1309 | 251.7112 | 320.515 | 221.9179
2012-10-30 02:13 | 25.7395 | 347.4834 | 273.0441 | 341.8409 | 242.2354 | 92.5954 | 1.6491 | 42.1088 | 69.4421 | 25.5342 | 2.6188 | 41.9831 | 109.9451 | 16.0315 | 21.3561 | 6.3835 | 179.3872 | 326.1273 | 251.688 | 320.4848 | 220.8793
2012-10-30 02:13 | 25.7298 | 347.4661 | 272.8793 | 341.8236 | 242.0807 | 92.5919 | 1.6492 | 42.11 | 69.446 | 25.5343 | 2.6189 | 41.9844 | 109.9525 | 16.0314 | 21.3344 | 6.3833 | 179.3985 | 326.1317 | 251.5449 | 320.4892 | 220.7463
2012-10-30 02:13 | 25.7425 | 347.4597 | 272.7641 | 341.8515 | 243.1047 | 92.5989 | 1.6492 | 42.1114 | 69.4507 | 25.5342 | 2.619 | 41.9856 | 109.9588 | 16.0314 | 21.3952 | 6.3838 | 179.4095 | 326.0645 | 251.3689 | 320.4563 | 221.7095
2012-10-30 02:13 | 25.7417 | 347.4627 | 272.7904 | 341.8811 | 242.8962 | 92.5942 | 1.6494 | 42.1174 | 69.4703 | 25.5343 | 2.6193 | 41.9916 | 109.9903 | 16.0314 | 21.3744 | 6.3834 | 179.4606 | 326.0883 | 251.416 | 320.5067 | 221.5218
2012-10-30 02:13 | 25.7313 | 347.4447 | 272.7761 | 341.8365 | 242.7329 | 92.5827 | 1.6497 | 42.1236 | 69.4907 | 25.5343 | 2.6197 | 41.9977 | 110.0222 | 16.0314 | 21.3479 | 6.3827 | 179.5129 | 326.0968 | 251.4282 | 320.4886 | 221.385
2012-10-30 02:14 | 25.7342 | 347.4362 | 272.7325 | 341.8546 | 243.4256 | 92.5942 | 1.6499 | 42.1296 | 69.5107 | 25.5343 | 2.6201 | 42.0035 | 110.0523 | 16.0314 | 21.3508 | 6.3834 | 179.563 | 326.0854 | 251.3817 | 320.5038 | 222.0748
2012-10-30 02:14 | 25.7425 | 347.4826 | 272.6826 | 341.8477 | 242.3404 | 92.5885 | 1.6501 | 42.1332 | 69.5225 | 25.5343 | 2.6203 | 42.0073 | 110.0712 | 16.0316 | 21.3511 | 6.3831 | 179.5937 | 326.1315 | 251.3315 | 320.4966 | 220.9893
2012-10-30 02:14 | 25.7305 | 347.4249 | 272.5814 | 341.752 | 242.8458 | 92.5867 | 1.6513 | 42.1647 | 69.6252 | 25.5347 | 2.6223 | 42.0397 | 110.2405 | 16.0316 | 21.3311 | 6.3829 | 179.8658 | 326.0938 | 251.2503 | 320.4209 | 221.5147
2012-10-30 02:14 | 25.7305 | 347.4135 | 272.6202 | 341.8014 | 243.5512 | 92.5781 | 1.6535 | 42.2227 | 69.8154 | 25.5353 | 2.6258 | 42.0969 | 110.5374 | 16.0321 | 21.3431 | 6.3823 | 180.3528 | 326.0704 | 251.2771 | 320.4583 | 222.2081
2012-10-30 02:14 | 25.7387 | 347.4521 | 272.6905 | 341.8058 | 243.3321 | 92.5908 | 1.6539 | 42.2341 | 69.8526 | 25.5355 | 2.6265 | 42.1082 | 110.5959 | 16.0323 | 21.3433 | 6.3832 | 180.4484 | 326.1088 | 251.3472 | 320.4625 | 221.9888
2012-10-30 02:14 | 25.723 | 347.3908 | 272.6942 | 341.7368 | 243.673 | 92.5821 | 1.6543 | 42.2443 | 69.8854 | 25.5358 | 2.6271 | 42.1182 | 110.6487 | 16.0322 | 21.3476 | 6.3826 | 180.534 | 326.0432 | 251.3466 | 320.3892 | 222.3254
2012-10-30 02:14 | 25.732 | 347.3998 | 272.7574 | 341.7877 | 243.6193 | 92.6029 | 1.6546 | 42.2523 | 69.911 | 25.5362 | 2.6275 | 42.126 | 110.688 | 16.0325 | 21.3165 | 6.384 | 180.5989 | 326.0833 | 251.4409 | 320.4712 | 222.3028
2012-10-30 02:14 | 25.7313 | 347.4295 | 272.718 | 341.7566 | 242.9445 | 92.6006 | 1.6548 | 42.2578 | 69.9296 | 25.536 | 2.6279 | 42.1316 | 110.717 | 16.0325 | 21.3559 | 6.3839 | 180.6466 | 326.0736 | 251.3621 | 320.4007 | 221.5886
2012-10-30 02:14 | 25.7305 | 347.4401 | 272.7017 | 341.7558 | 243.4062 | 92.5937 | 1.6549 | 42.2599 | 69.9361 | 25.5362 | 2.628 | 42.1337 | 110.7277 | 16.0325 | 21.3632 | 6.3834 | 180.6638 | 326.0769 | 251.3385 | 320.3926 | 222.043
2012-10-30 02:14 | 25.7238 | 347.4448 | 272.7454 | 341.7491 | 243.6307 | 92.6104 | 1.6549 | 42.2604 | 69.9375 | 25.5362 | 2.628 | 42.1341 | 110.7295 | 16.0326 | 21.3043 | 6.3846 | 180.667 | 326.1405 | 251.4411 | 320.4448 | 222.3264
2012-10-30 02:14 | 25.7283 | 347.4455 | 272.7964 | 341.7383 | 243.6626 | 92.6099 | 1.6549 | 42.2608 | 69.939 | 25.5362 | 2.628 | 42.1346 | 110.7313 | 16.0328 | 21.3128 | 6.3845 | 180.6704 | 326.1327 | 251.4836 | 320.4255 | 222.3498
2012-10-30 02:14 | 25.7298 | 347.4318 | 272.8444 | 341.7284 | 243.7386 | 92.5989 | 1.6549 | 42.2611 | 69.9396 | 25.5363 | 2.6281 | 42.1348 | 110.7329 | 16.0327 | 21.3304 | 6.3838 | 180.6724 | 326.1014 | 251.514 | 320.398 | 222.4082
2012-10-30 02:14 | 25.7275 | 347.4295 | 272.9042 | 341.7185 | 243.4894 | 92.6018 | 1.6549 | 42.2596 | 69.9346 | 25.5363 | 2.628 | 42.1332 | 110.7244 | 16.0327 | 21.3321 | 6.384 | 180.659 | 326.0974 | 251.5721 | 320.3864 | 222.1573
2012-10-30 02:14 | 25.7268 | 347.4631 | 272.8763 | 341.7673 | 243.4848 | 92.6237 | 1.6549 | 42.2596 | 69.9343 | 25.5364 | 2.628 | 42.1334 | 110.7251 | 16.0327 | 21.3314 | 6.3855 | 180.6594 | 326.1317 | 251.5449 | 320.4359 | 222.1534
2012-10-30 02:14 | 25.7283 | 347.457 | 272.9438 | 341.7345 | 243.6548 | 92.6041 | 1.6549 | 42.2593 | 69.9331 | 25.5366 | 2.628 | 42.1331 | 110.7238 | 16.0327 | 21.389 | 6.3841 | 180.6569 | 326.068 | 251.5548 | 320.3455 | 222.2658
2012-10-30 02:14 | 25.729 | 347.4196 | 272.9716 | 341.7314 | 243.8083 | 92.611 | 1.6548 | 42.2584 | 69.9299 | 25.5366 | 2.6279 | 42.1321 | 110.7171 | 16.0328 | 21.3737 | 6.3846 | 180.647 | 326.0459 | 251.5979 | 320.3577 | 222.4346
2012-10-30 02:14 | 25.7275 | 347.4219 | 273.0166 | 341.7566 | 243.9518 | 92.6116 | 1.6548 | 42.2574 | 69.927 | 25.5365 | 2.6278 | 42.1312 | 110.7143 | 16.0326 | 21.3923 | 6.3846 | 180.6413 | 326.0296 | 251.6243 | 320.3643 | 222.5595
2012-10-30 02:14 | 25.7253 | 347.4463 | 273.0416 | 341.7277 | 243.7497 | 92.596 | 1.6547 | 42.2565 | 69.9237 | 25.5366 | 2.6278 | 42.1302 | 110.7079 | 16.0327 | 21.366 | 6.3836 | 180.6316 | 326.0803 | 251.6756 | 320.3617 | 222.3837
2012-10-30 02:14 | 25.7118 | 347.429 | 273.0513 | 341.7447 | 243.4698 | 92.6041 | 1.6547 | 42.2547 | 69.9174 | 25.5367 | 2.6276 | 42.1283 | 110.6982 | 16.0327 | 21.3685 | 6.3841 | 180.6156 | 326.0605 | 251.6828 | 320.3762 | 222.1013
2012-10-30 02:14 | 25.717 | 347.4076 | 273.0178 | 341.807 | 243.4162 | 92.611 | 1.6546 | 42.2528 | 69.912 | 25.5364 | 2.6276 | 42.1266 | 110.69 | 16.0326 | 21.3737 | 6.3846 | 180.602 | 326.0339 | 251.6441 | 320.4333 | 222.0425
2012-10-30 02:14 | 25.7215 | 347.3931 | 273.0106 | 341.7696 | 243.5069 | 92.6035 | 1.6546 | 42.2513 | 69.9072 | 25.5363 | 2.6275 | 42.1251 | 110.6829 | 16.0325 | 21.3542 | 6.3841 | 180.5901 | 326.0389 | 251.6564 | 320.4154 | 222.1527
2012-10-30 02:14 | 25.72 | 347.4372 | 272.9975 | 341.7909 | 243.5133 | 92.6012 | 1.6545 | 42.25 | 69.9036 | 25.5361 | 2.6274 | 42.1238 | 110.676 | 16.0325 | 21.3647 | 6.3839 | 180.5796 | 326.0725 | 251.6328 | 320.4262 | 222.1486
2012-10-30 02:14 | 25.723 | 347.4288 | 273.0315 | 341.8206 | 243.6769 | 92.6139 | 1.6544 | 42.2482 | 69.8971 | 25.5363 | 2.6273 | 42.1221 | 110.6673 | 16.0325 | 21.3396 | 6.3848 | 180.5643 | 326.0892 | 251.6919 | 320.481 | 222.3373
2012-10-30 02:14 | 25.7193 | 347.4175 | 273.055 | 341.794 | 243.6693 | 92.6035 | 1.6544 | 42.247 | 69.8929 | 25.5363 | 2.6272 | 42.1207 | 110.6605 | 16.0324 | 21.3239 | 6.3841 | 180.5535 | 326.0936 | 251.7311 | 320.4701 | 222.3454
2012-10-30 02:14 | 25.7245 | 347.4189 | 273.0485 | 341.746 | 243.6196 | 92.6214 | 1.6544 | 42.2461 | 69.8913 | 25.5358 | 2.6272 | 42.1198 | 110.6562 | 16.0323 | 21.309 | 6.3853 | 180.5475 | 326.1099 | 251.7395 | 320.437 | 222.3106
2012-10-30 02:15 | 25.7238 | 347.4334 | 273.0439 | 341.8023 | 243.5641 | 92.6127 | 1.6543 | 42.2436 | 69.8827 | 25.5359 | 2.627 | 42.1174 | 110.6428 | 16.0324 | 21.2963 | 6.3847 | 180.5256 | 326.1371 | 251.7476 | 320.506 | 222.2678
2012-10-30 02:15 | 25.726 | 347.4356 | 273.05 | 341.7741 | 243.5153 | 92.6075 | 1.6542 | 42.2429 | 69.8799 | 25.5362 | 2.627 | 42.1167 | 110.64 | 16.0323 | 21.2985 | 6.3844 | 180.5199 | 326.1371 | 251.7515 | 320.4756 | 222.2168
2012-10-30 02:15 | 25.7245 | 347.4227 | 273.0369 | 341.7688 | 243.5844 | 92.611 | 1.6542 | 42.2428 | 69.8798 | 25.536 | 2.627 | 42.1164 | 110.6393 | 16.0322 | 21.317 | 6.3846 | 180.5191 | 326.1057 | 251.7199 | 320.4518 | 222.2674
2012-10-30 02:15 | 25.7298 | 347.4585 | 273.0306 | 341.8159 | 243.5544 | 92.6151 | 1.6542 | 42.2418 | 69.8764 | 25.536 | 2.6269 | 42.1155 | 110.6351 | 16.0321 | 21.2943 | 6.3849 | 180.5115 | 326.1642 | 251.7363 | 320.5216 | 222.2601
2012-10-30 02:15 | 25.72 | 347.422 | 273.0634 | 341.7795 | 243.5368 | 92.6087 | 1.6541 | 42.24 | 69.8711 | 25.5359 | 2.6268 | 42.1138 | 110.6256 | 16.0322 | 21.3286 | 6.3844 | 180.4967 | 326.0934 | 251.7348 | 320.4509 | 222.2082
2012-10-30 02:15 | 25.7313 | 347.4333 | 273.0747 | 341.806 | 243.6029 | 92.6058 | 1.6541 | 42.2391 | 69.8676 | 25.536 | 2.6268 | 42.1129 | 110.621 | 16.0322 | 21.3238 | 6.3842 | 180.4886 | 326.1095 | 251.7509 | 320.4822 | 222.2791
2012-10-30 02:15 | 25.7268 | 347.4326 | 273.0818 | 341.7977 | 243.8179 | 92.6047 | 1.654 | 42.2372 | 69.8619 | 25.5359 | 2.6267 | 42.1113 | 110.6124 | 16.0322 | 21.3354 | 6.3842 | 180.4743 | 326.0972 | 251.7464 | 320.4623 | 222.4825
2012-10-30 02:15 | 25.7357 | 347.4377 | 273.1179 | 341.7952 | 243.913 | 92.6064 | 1.6539 | 42.2353 | 69.8546 | 25.5362 | 2.6265 | 42.1092 | 110.6012 | 16.0322 | 21.3162 | 6.3843 | 180.4559 | 326.1215 | 251.8017 | 320.479 | 222.5968
2012-10-30 02:15 | 25.7372 | 347.4164 | 273.1465 | 341.7663 | 243.7577 | 92.6081 | 1.6539 | 42.2339 | 69.8501 | 25.5362 | 2.6264 | 42.1077 | 110.593 | 16.0323 | 21.3257 | 6.3844 | 180.4431 | 326.0907 | 251.8208 | 320.4406 | 222.432
2012-10-30 02:15 | 25.735 | 347.418 | 273.1133 | 341.8021 | 243.7359 | 92.6093 | 1.6539 | 42.2334 | 69.8488 | 25.536 | 2.6264 | 42.1073 | 110.5917 | 16.0322 | 21.3115 | 6.3845 | 180.4405 | 326.1065 | 251.8018 | 320.4906 | 222.4244
2012-10-30 02:15 | 25.7417 | 347.4361 | 273.2053 | 341.786 | 243.7387 | 92.5995 | 1.6538 | 42.2328 | 69.8463 | 25.5361 | 2.6264 | 42.1064 | 110.5868 | 16.0322 | 21.3423 | 6.3838 | 180.4332 | 326.0938 | 251.863 | 320.4437 | 222.3964
2012-10-30 02:15 | 25.747 | 347.4528 | 273.2145 | 341.8217 | 243.795 | 92.6156 | 1.6538 | 42.2314 | 69.8425 | 25.5359 | 2.6263 | 42.1051 | 110.5793 | 16.0323 | 21.3074 | 6.3849 | 180.4218 | 326.1454 | 251.9071 | 320.5143 | 222.4876
2012-10-30 02:15 | 25.753 | 347.4246 | 273.1933 | 341.8201 | 243.8049 | 92.6058 | 1.6536 | 42.2278 | 69.8298 | 25.5361 | 2.6261 | 42.1016 | 110.5616 | 16.0322 | 21.3496 | 6.3842 | 180.3914 | 326.075 | 251.8437 | 320.4705 | 222.4553
2012-10-30 02:15 | 25.7447 | 347.4734 | 273.2122 | 341.7738 | 243.6438 | 92.5931 | 1.6503 | 42.1405 | 69.5444 | 25.5351 | 2.6205 | 42.0106 | 110.0878 | 16.0317 | 21.2971 | 6.3834 | 179.6321 | 326.1763 | 251.9151 | 320.4767 | 222.3467
2012-10-30 02:15 | 25.7492 | 347.4474 | 273.1818 | 341.8392 | 243.707 | 92.5937 | 1.6498 | 42.1272 | 69.502 | 25.5345 | 2.62 | 42.0017 | 110.0434 | 16.0314 | 21.2856 | 6.3834 | 179.5454 | 326.1618 | 251.8962 | 320.5536 | 222.4214
2012-10-30 02:15 | 25.7559 | 347.4161 | 273.1769 | 341.7888 | 243.8783 | 92.5948 | 1.6494 | 42.1154 | 69.4634 | 25.5344 | 2.6193 | 41.9897 | 109.9818 | 16.0311 | 21.2762 | 6.3835 | 179.4452 | 326.1399 | 251.9007 | 320.5126 | 222.6021
2012-10-30 02:15 | 25.753 | 347.3942 | 273.1546 | 341.8087 | 242.88 | 92.5815 | 1.6492 | 42.1118 | 69.4522 | 25.5341 | 2.619 | 41.986 | 109.9629 | 16.0311 | 21.3295 | 6.3826 | 179.4151 | 326.0647 | 251.8251 | 320.4792 | 221.5505
2012-10-30 02:15 | 25.7582 | 347.3803 | 273.0086 | 341.8253 | 242.262 | 92.5896 | 1.6493 | 42.1141 | 69.4598 | 25.5342 | 2.6192 | 41.9882 | 109.9754 | 16.031 | 21.3186 | 6.3831 | 179.4352 | 326.0617 | 251.69 | 320.5067 | 220.9434
2012-10-30 02:15 | 25.7559 | 347.3894 | 272.8628 | 341.8002 | 242.44 | 92.5879 | 1.6494 | 42.1165 | 69.4671 | 25.5344 | 2.6194 | 41.991 | 109.9897 | 16.031 | 21.3284 | 6.383 | 179.4568 | 326.061 | 251.5344 | 320.4718 | 221.1116
2012-10-30 02:15 | 25.7552 | 347.4268 | 272.7806 | 341.7652 | 242.8352 | 92.5682 | 1.6495 | 42.1187 | 69.4743 | 25.5344 | 2.6195 | 41.9927 | 109.9983 | 16.031 | 21.3197 | 6.3817 | 179.4726 | 326.1071 | 251.4609 | 320.4455 | 221.5155
2012-10-30 02:15 | 25.753 | 347.4284 | 272.7629 | 341.7935 | 243.2759 | 92.581 | 1.6497 | 42.1242 | 69.4933 | 25.5341 | 2.6198 | 41.9984 | 110.0283 | 16.0311 | 21.3215 | 6.3825 | 179.5216 | 326.1069 | 251.4414 | 320.472 | 221.9544
2012-10-30 02:15 | 25.7507 | 347.3842 | 272.8072 | 341.776 | 243.058 | 92.5867 | 1.6501 | 42.1335 | 69.5229 | 25.5345 | 2.6204 | 42.0077 | 110.0765 | 16.0311 | 21.3111 | 6.3829 | 179.5993 | 326.0731 | 251.4961 | 320.4649 | 221.7469
2012-10-30 02:15 | 25.735 | 347.3609 | 272.7566 | 341.7983 | 242.5798 | 92.5798 | 1.6503 | 42.1384 | 69.5391 | 25.5345 | 2.6207 | 42.0125 | 110.1014 | 16.0312 | 21.3556 | 6.3825 | 179.6405 | 326.0053 | 251.401 | 320.4427 | 221.2242
2012-10-30 02:15 | 25.7395 | 347.3654 | 272.6486 | 341.761 | 242.6314 | 92.5659 | 1.6504 | 42.1425 | 69.5531 | 25.5343 | 2.621 | 42.017 | 110.1245 | 16.0312 | 21.3601 | 6.3815 | 179.6776 | 326.0053 | 251.2885 | 320.4009 | 221.2713
2012-10-30 02:15 | 25.7342 | 347.4096 | 272.6394 | 341.7214 | 242.4105 | 92.5717 | 1.6533 | 42.2179 | 69.7993 | 25.5354 | 2.6256 | 42.0921 | 110.5154 | 16.0317 | 21.3629 | 6.3819 | 180.3147 | 326.0467 | 251.2765 | 320.3585 | 221.0476
2012-10-30 02:16 | 25.7327 | 347.4005 | 272.5914 | 341.7085 | 242.8754 | 92.589 | 1.6538 | 42.2324 | 69.8455 | 25.536 | 2.6264 | 42.1058 | 110.5867 | 16.0318 | 21.3453 | 6.3831 | 180.4323 | 326.0552 | 251.2461 | 320.3632 | 221.5301
2012-10-30 02:16 | 25.7298 | 347.3938 | 272.5846 | 341.7284 | 242.7549 | 92.5688 | 1.6543 | 42.244 | 69.8837 | 25.5361 | 2.6271 | 42.1175 | 110.6471 | 16.0319 | 21.3464 | 6.3817 | 180.5308 | 326.0474 | 251.2382 | 320.382 | 221.4085
2012-10-30 02:16 | 25.7185 | 347.3825 | 272.5578 | 341.7628 | 242.5672 | 92.5937 | 1.6547 | 42.2541 | 69.9166 | 25.5363 | 2.6277 | 42.1278 | 110.6998 | 16.0321 | 21.3512 | 6.3834 | 180.6164 | 326.0313 | 251.2066 | 320.4116 | 221.216
2012-10-30 02:16 | 25.7155 | 347.4061 | 272.5742 | 341.7179 | 243.2619 | 92.5995 | 1.6549 | 42.2595 | 69.9343 | 25.5363 | 2.628 | 42.1333 | 110.7279 | 16.0322 | 21.3201 | 6.3838 | 180.6622 | 326.086 | 251.2541 | 320.3978 | 221.9418
2012-10-30 02:16 | 25.7081 | 347.4177 | 272.6094 | 341.6877 | 242.5333 | 92.5752 | 1.655 | 42.2624 | 69.9437 | 25.5365 | 2.6282 | 42.136 | 110.7417 | 16.0323 | 21.2966 | 6.3821 | 180.6854 | 326.1211 | 251.3128 | 320.3911 | 221.2367
2012-10-30 02:16 | 25.7081 | 347.4215 | 272.5513 | 341.7105 | 243.0624 | 92.6012 | 1.6551 | 42.2645 | 69.9502 | 25.5366 | 2.6283 | 42.1381 | 110.753 | 16.0323 | 21.3408 | 6.3839 | 180.7031 | 326.0807 | 251.2105 | 320.3697 | 221.7216
2012-10-30 02:16 | 25.7013 | 347.4109 | 272.56 | 341.6885 | 243.2555 | 92.5977 | 1.6551 | 42.2646 | 69.9512 | 25.5364 | 2.6283 | 42.1383 | 110.7535 | 16.0323 | 21.338 | 6.3837 | 180.7046 | 326.0729 | 251.222 | 320.3505 | 221.9175
2012-10-30 02:16 | 25.7043 | 347.4215 | 272.6018 | 341.6839 | 243.38 | 92.5885 | 1.6551 | 42.2656 | 69.9542 | 25.5365 | 2.6284 | 42.1393 | 110.7581 | 16.0324 | 21.3009 | 6.3831 | 180.7123 | 326.1206 | 251.3009 | 320.383 | 222.0791
2012-10-30 02:16 | 25.6938 | 347.4301 | 272.6223 | 341.6696 | 243.4283 | 92.5908 | 1.6551 | 42.2655 | 69.9534 | 25.5366 | 2.6284 | 42.1391 | 110.7579 | 16.0323 | 21.2663 | 6.3832 | 180.7113 | 326.1638 | 251.356 | 320.4033 | 222.162
2012-10-30 02:16 | 25.6931 | 347.4027 | 272.6255 | 341.6841 | 243.4315 | 92.5925 | 1.6551 | 42.2647 | 69.9512 | 25.5364 | 2.6283 | 42.1383 | 110.7535 | 16.0323 | 21.2977 | 6.3833 | 180.7047 | 326.105 | 251.3278 | 320.3864 | 222.1338
2012-10-30 02:16 | 25.6946 | 347.4385 | 272.689 | 341.678 | 243.4369 | 92.5954 | 1.655 | 42.2633 | 69.9458 | 25.5367 | 2.6283 | 42.1371 | 110.7473 | 16.0323 | 21.3152 | 6.3835 | 180.6931 | 326.1233 | 251.3738 | 320.3628 | 222.1217
2012-10-30 02:16 | 25.6938 | 347.4263 | 272.7076 | 341.6772 | 243.3029 | 92.5983 | 1.655 | 42.2621 | 69.9426 | 25.5365 | 2.6282 | 42.1357 | 110.7404 | 16.0323 | 21.2542 | 6.3837 | 180.683 | 326.1721 | 251.4534 | 320.423 | 222.0487
2012-10-30 02:16 | 25.6953 | 347.4202 | 272.7246 | 341.713 | 243.3632 | 92.5937 | 1.6646 | 42.5098 | 70.7628 | 25.5372 | 2.6392 | 42.3131 | 111.6742 | 16.0324 | 21.2878 | 6.3834 | 182.437 | 326.1324 | 251.4368 | 320.4252 | 222.0754
2012-10-30 02:16 | 25.6976 | 347.4833 | 272.9673 | 341.662 | 243.4791 | 93.0208 | 1.8722 | 47.9558 | 89.7816 | 25.615 | 2.9719 | 47.8127 | 142.0936 | 16.0884 | 21.3062 | 6.4129 | 231.8753 | 326.1771 | 251.6611 | 320.3558 | 222.1729
2012-10-30 02:16 | 25.6916 | 348.2535 | 273.7445 | 341.7397 | 243.528 | 93.4681 | 1.8709 | 47.9508 | 89.7105 | 25.63 | 2.9677 | 47.8077 | 141.8766 | 16.1096 | 21.2962 | 6.4437 | 231.5871 | 326.9573 | 252.4483 | 320.4435 | 222.2318
2012-10-30 02:16 | 25.6938 | 349.0889 | 274.2157 | 341.8865 | 243.4675 | 93.5478 | 1.8705 | 47.9463 | 89.6847 | 25.6325 | 2.9661 | 47.8034 | 141.789 | 16.1167 | 21.2703 | 6.4492 | 231.4737 | 327.8186 | 252.9454 | 320.6162 | 222.1972
2012-10-30 02:16 | 25.6908 | 349.9761 | 274.6816 | 342.146 | 243.6839 | 93.612 | 1.8702 | 47.9421 | 89.6613 | 25.6348 | 2.965 | 47.7992 | 141.7261 | 16.121 | 21.2914 | 6.4536 | 231.3874 | 328.6847 | 253.3902 | 320.8546 | 222.3925
2012-10-30 02:16 | 25.6864 | 350.805 | 275.0995 | 342.526 | 244.134 | 93.6998 | 1.8701 | 47.9415 | 89.6532 | 25.6364 | 2.9645 | 47.7986 | 141.6974 | 16.1238 | 21.2629 | 6.4597 | 231.3506 | 329.5421 | 253.8366 | 321.2631 | 222.8711
2012-10-30 02:16 | 25.7006 | 351.5269 | 275.5088 | 342.9207 | 244.3715 | 93.7628 | 1.8698 | 47.9382 | 89.6356 | 25.638 | 2.9638 | 47.7953 | 141.6556 | 16.1264 | 21.261 | 6.464 | 231.2912 | 330.2659 | 254.2478 | 321.6597 | 223.1105
2012-10-30 02:16 | 25.6998 | 352.1653 | 275.8721 | 343.2851 | 244.6684 | 93.8218 | 1.8697 | 47.9366 | 89.6253 | 25.6392 | 2.9634 | 47.794 | 141.6308 | 16.1283 | 21.2923 | 6.4681 | 231.2561 | 330.873 | 254.5798 | 321.9928 | 223.3761
2012-10-30 02:16 | 25.6916 | 352.9333 | 276.2318 | 343.7602 | 245.1538 | 93.894 | 1.8694 | 47.9342 | 89.6103 | 25.6409 | 2.9628 | 47.7916 | 141.597 | 16.1306 | 21.2561 | 6.4731 | 231.2073 | 331.6772 | 254.9757 | 322.5041 | 223.8977
2012-10-30 02:16 | 25.6983 | 353.6515 | 276.7535 | 344.2272 | 245.4111 | 93.9506 | 1.8693 | 47.9321 | 89.5975 | 25.6423 | 2.9623 | 47.7894 | 141.5675 | 16.1325 | 21.2828 | 6.477 | 231.1649 | 332.3687 | 255.4707 | 322.9444 | 224.1283
2012-10-30 02:16 | 25.6968 | 354.1751 | 277.0501 | 344.6785 | 245.6564 | 94.0206 | 1.8691 | 47.9295 | 89.5838 | 25.6434 | 2.9619 | 47.7871 | 141.5416 | 16.1339 | 21.2653 | 6.4818 | 231.1254 | 332.9098 | 255.7848 | 323.4132 | 224.3911
2012-10-30 02:16 | 25.7013 | 354.7352 | 277.4378 | 345.1586 | 246.0446 | 94.0645 | 1.869 | 47.9291 | 89.5775 | 25.6448 | 2.9616 | 47.7863 | 141.5243 | 16.1352 | 21.2577 | 6.4848 | 231.1018 | 333.4775 | 256.1801 | 323.9009 | 224.7869
2012-10-30 02:17
2012-10-30 02:19
2012-10-30 02:21
2012-10-30 02:24
2012-10-30 02:26
2012-10-30 02:29
2012-10-30 02:31
2012-10-30 02:34
2012-10-30 02:36
2012-10-30 02:39
2012-10-30 02:42
2012-10-30 02:44
2012-10-30 02:47
2012-10-30 02:50
2012-10-30 02:52
2012-10-30 02:55
2012-10-30 02:58
2012-10-30 03:01
2012-10-30 03:03
2012-10-30 03:06
2012-10-30 03:09
2012-10-30 03:11
2012-10-30 03:14
2012-10-30 03:17
2012-10-30 03:19
2012-10-30 03:22
2012-10-30 03:25
2012-10-30 03:27
2012-10-30 03:30 |  |  |  | -8.9218 | -9.1809 | 46.9317 | 0.037731 | 0.90051 | 0.033977 | 23.8666 | 0.051679 | 0.93615 | 0.04838 | 18.1148 | -10.5347 | 3.2355 | 0.082357 | 10.5347 | 10.5347 | 1.6129 | 1.3538
2012-10-30 03:31 | 22.9474 | 13.1913 | 13.9487 | 14.0984 | 13.807 | 46.7624 | 0.029577 | 0.87988 | 0.026024 | 29.7488 | 0.048371 | 0.88304 | 0.042714 | 18.2556 | 12.4248 | 3.2238 | 0.068738 | 0.7665 | 1.5239 | 1.6736 | 1.3822
2012-10-30 03:31 | 22.9586 | 13.2106 | 13.9477 | 14.1784 | 13.8627 | 46.7664 | 0.031028 | 0.915 | 0.028391 | 29.4896 | 0.04981 | 0.90633 | 0.045144 | 18.1958 | 12.436 | 3.2241 | 0.073535 | 0.7746 | 1.5117 | 1.7424 | 1.4267
2012-10-30 03:31 | 22.9646 | 13.1801 | 14.0104 | 14.2087 | 13.8323 | 46.771 | 0.030474 | 0.90148 | 0.027472 | 29.5819 | 0.049375 | 0.89915 | 0.044396 | 18.2107 | 12.4826 | 3.2244 | 0.071868 | 0.6975 | 1.5278 | 1.7261 | 1.3497
2012-10-30 03:31 | 22.9624 | 13.1617 | 13.9475 | 14.2065 | 13.8503 | 46.7618 | 0.030146 | 0.89363 | 0.026939 | 29.6435 | 0.048985 | 0.89277 | 0.043732 | 18.2253 | 12.4236 | 3.2238 | 0.070671 | 0.7381 | 1.5239 | 1.7829 | 1.4267
2012-10-30 03:31 | 22.9796 | 13.1911 | 13.9768 | 14.2277 | 13.8675 | 46.7306 | 0.029684 | 0.88287 | 0.026207 | 29.7423 | 0.048903 | 0.89142 | 0.043593 | 18.2284 | 12.4976 | 3.2216 | 0.0698 | 0.6935 | 1.4792 | 1.7301 | 1.3699
2012-10-30 03:31 | 22.9833 | 13.1583 | 14.0048 | 14.3043 | 13.9158 | 46.7317 | 0.028872 | 0.86256 | 0.024904 | 29.8752 | 0.047786 | 0.87361 | 0.041746 | 18.2817 | 12.4932 | 3.2217 | 0.06665 | 0.6651 | 1.5116 | 1.8111 | 1.4226
2012-10-30 03:31 | 22.9818 | 13.1649 | 14.0276 | 14.2502 | 13.8333 | 46.7624 | 0.028658 | 0.85773 | 0.024581 | 29.9299 | 0.04771 | 0.8726 | 0.041632 | 18.2896 | 12.4795 | 3.2238 | 0.066213 | 0.6854 | 1.5481 | 1.7707 | 1.3538
2012-10-30 03:31 | 22.9871 | 13.1864 | 13.9965 | 14.2474 | 13.8629 | 46.7496 | 0.028129 | 0.84478 | 0.023763 | 30.0325 | 0.04529 | 0.83389 | 0.037767 | 18.4122 | 12.5011 | 3.2229 | 0.06153 | 0.6853 | 1.4954 | 1.7463 | 1.3618
2012-10-30 03:31 | 22.993 | 13.168 | 14.0105 | 14.3382 | 13.8445 | 46.7537 | 0.026571 | 0.80705 | 0.021444 | 30.3732 | 0.043405 | 0.80313 | 0.03486 | 18.5032 | 12.5475 | 3.2232 | 0.056304 | 0.6205 | 1.463 | 1.7907 | 1.297
2012-10-30 03:31 | 23.0072 | 13.166 | 14.0368 | 14.3484 | 13.8992 | 46.745 | 0.026474 | 0.80452 | 0.021299 | 30.3889 | 0.043119 | 0.79864 | 0.034436 | 18.5217 | 12.5252 | 3.2226 | 0.055735 | 0.6408 | 1.5116 | 1.8232 | 1.374
2012-10-30 03:31 | 23.0125 | 13.2077 | 14.0219 | 14.3173 | 13.8883 | 46.7548 | 0.02388 | 0.74194 | 0.017718 | 31.0696 | 0.039575 | 0.74149 | 0.029345 | 18.7364 | 12.5549 | 3.2233 | 0.047063 | 0.6528 | 1.467 | 1.7624 | 1.3334
2012-10-30 03:31 | 23.0199 | 13.1787 | 14.0536 | 14.3328 | 13.8754 | 46.7595 | 0.024371 | 0.75405 | 0.018377 | 30.9405 | 0.040428 | 0.75534 | 0.030537 | 18.6835 | 12.5907 | 3.2236 | 0.048914 | 0.588 | 1.4629 | 1.7421 | 1.2847
2012-10-30 03:31 | 23.0334 | 13.1759 | 14.0225 | 14.3422 | 13.9537 | 46.7439 | 0.024699 | 0.76197 | 0.01882 | 30.8501 | 0.040497 | 0.75638 | 0.030631 | 18.6773 | 12.5798 | 3.2225 | 0.049451 | 0.5961 | 1.4427 | 1.7624 | 1.3739
2012-10-30 03:32 | 23.0401 | 13.1543 | 14.0211 | 14.3692 | 13.9078 | 46.7184 | 0.023679 | 0.73709 | 0.017454 | 31.1283 | 0.039194 | 0.73564 | 0.028833 | 18.7693 | 12.6027 | 3.2208 | 0.046287 | 0.5516 | 1.4184 | 1.7665 | 1.3051
2012-10-30 03:32 | 23.0416 | 13.176 | 14.0672 | 14.3545 | 13.9052 | 46.7363 | 0.023824 | 0.74082 | 0.017649 | 31.0953 | 0.039692 | 0.7435 | 0.029511 | 18.7318 | 12.588 | 3.222 | 0.04716 | 0.588 | 1.4792 | 1.7665 | 1.3172
2012-10-30 03:32 | 23.0573 | 13.1309 | 13.9938 | 14.3823 | 13.9088 | 46.73 | 0.024198 | 0.74954 | 0.018137 | 30.9754 | 0.039968 | 0.74795 | 0.029894 | 18.7138 | 12.6037 | 3.2216 | 0.048031 | 0.5272 | 1.3901 | 1.7786 | 1.3051
2012-10-30 03:32 | 23.0596 | 13.1859 | 14.069 | 14.3684 | 13.9111 | 46.7306 | 0.024642 | 0.76051 | 0.01874 | 30.8623 | 0.039928 | 0.74738 | 0.029841 | 18.7181 | 12.5857 | 3.2216 | 0.048581 | 0.6002 | 1.4833 | 1.7827 | 1.3254
2012-10-30 03:32 | 23.067 | 13.1366 | 14.0399 | 14.3435 | 13.9428 | 46.7086 | 0.023462 | 0.73179 | 0.017169 | 31.1902 | 0.039191 | 0.7357 | 0.028833 | 18.7722 | 12.585 | 3.2201 | 0.046002 | 0.5516 | 1.4549 | 1.7585 | 1.3578
2012-10-30 03:32 | 23.0715 | 13.1978 | 14.0404 | 14.3722 | 13.9432 | 46.7254 | 0.02383 | 0.74051 | 0.017646 | 31.0748 | 0.039053 | 0.73322 | 0.028634 | 18.7749 | 12.5936 | 3.2212 | 0.04628 | 0.6042 | 1.4468 | 1.7786 | 1.3496
2012-10-30 03:32 | 23.0805 | 13.1825 | 14.0777 | 14.3974 | 13.9036 | 46.7387 | 0.02378 | 0.73936 | 0.017582 | 31.0915 | 0.039732 | 0.74423 | 0.02957 | 18.7311 | 12.5782 | 3.2222 | 0.047152 | 0.6043 | 1.4995 | 1.8192 | 1.3254
2012-10-30 03:32 | 23.085 | 13.1748 | 14.0579 | 14.3372 | 13.9446 | 46.7277 | 0.024425 | 0.75516 | 0.018445 | 30.9176 | 0.040198 | 0.75164 | 0.030214 | 18.6984 | 12.6071 | 3.2214 | 0.048659 | 0.5677 | 1.4508 | 1.7301 | 1.3375
2012-10-30 03:32 | 23.0984 | 13.1842 | 14.0916 | 14.3951 | 13.9782 | 46.73 | 0.023814 | 0.73999 | 0.017622 | 31.0738 | 0.039512 | 0.7407 | 0.029267 | 18.7462 | 12.6164 | 3.2216 | 0.046889 | 0.5678 | 1.4752 | 1.7787 | 1.3618
2012-10-30 03:32 | 23.1097 | 13.1833 | 14.0907 | 14.3821 | 13.9328 | 46.7167 | 0.023644 | 0.73602 | 0.017402 | 31.1292 | 0.039251 | 0.73657 | 0.028911 | 18.7656 | 12.6034 | 3.2206 | 0.046313 | 0.5799 | 1.4873 | 1.7787 | 1.3294
2012-10-30 03:32 | 23.1082 | 13.1656 | 14.073 | 14.3766 | 13.9435 | 46.7323 | 0.02343 | 0.73083 | 0.017123 | 31.1921 | 0.039103 | 0.73407 | 0.028704 | 18.7727 | 12.6181 | 3.2217 | 0.045827 | 0.5475 | 1.4549 | 1.7585 | 1.3254
2012-10-30 03:32 | 23.1194 | 13.1647 | 14.1004 | 14.3918 | 13.9668 | 46.7294 | 0.023191 | 0.72552 | 0.016825 | 31.2843 | 0.038729 | 0.72804 | 0.028196 | 18.7984 | 12.5928 | 3.2215 | 0.045021 | 0.5719 | 1.5076 | 1.799 | 1.374
2012-10-30 03:32 | 23.1306 | 13.2042 | 14.0995 | 14.3666 | 13.9699 | 46.7259 | 0.02333 | 0.72839 | 0.016993 | 31.2212 | 0.038961 | 0.73193 | 0.028517 | 18.7862 | 12.6243 | 3.2213 | 0.04551 | 0.5799 | 1.4752 | 1.7423 | 1.3456
2012-10-30 03:32 | 23.1411 | 13.2066 | 14.0533 | 14.4014 | 13.9926 | 46.7144 | 0.023097 | 0.72286 | 0.016696 | 31.2965 | 0.038508 | 0.72446 | 0.027898 | 18.8132 | 12.5658 | 3.2205 | 0.044594 | 0.6408 | 1.4875 | 1.8356 | 1.4268
2012-10-30 03:32 | 23.1463 | 13.1794 | 14.1071 | 14.4268 | 13.9533 | 46.7063 | 0.023518 | 0.7329 | 0.017236 | 31.1633 | 0.039078 | 0.73326 | 0.028654 | 18.764 | 12.5466 | 3.2199 | 0.04589 | 0.6328 | 1.5605 | 1.8802 | 1.4067
2012-10-30 03:32 | 23.1568 | 13.1696 | 14.0771 | 14.3928 | 13.9718 | 46.6953 | 0.023116 | 0.72375 | 0.01673 | 31.3096 | 0.038628 | 0.72627 | 0.028054 | 18.8015 | 12.6261 | 3.2192 | 0.044784 | 0.5435 | 1.451 | 1.7667 | 1.3457
2012-10-30 03:32 | 23.1635 | 13.152 | 14.0352 | 14.3914 | 14.019 | 46.7075 | 0.023024 | 0.7213 | 0.016607 | 31.328 | 0.038499 | 0.72414 | 0.027879 | 18.8094 | 12.6288 | 3.22 | 0.044486 | 0.5232 | 1.4064 | 1.7626 | 1.3902
2012-10-30 03:32 | 23.1717 | 13.2088 | 14.1204 | 14.3996 | 13.9706 | 46.6982 | 0.023005 | 0.72081 | 0.016582 | 31.3325 | 0.038461 | 0.72349 | 0.027826 | 18.8111 | 12.6289 | 3.2194 | 0.044408 | 0.5799 | 1.4915 | 1.7707 | 1.3417
2012-10-30 03:32 | 23.1829 | 13.1309 | 14.087 | 14.4108 | 14.006 | 46.6994 | 0.023034 | 0.7213 | 0.016614 | 31.3144 | 0.03843 | 0.72293 | 0.027782 | 18.8116 | 12.6198 | 3.2195 | 0.044396 | 0.5111 | 1.4672 | 1.791 | 1.3862
2012-10-30 03:32 | 23.1956 | 13.1638 | 14.1604 | 14.375 | 14.0187 | 46.7127 | 0.023078 | 0.72238 | 0.016671 | 31.3016 | 0.038565 | 0.72541 | 0.027976 | 18.8102 | 12.5838 | 3.2204 | 0.044647 | 0.58 | 1.5766 | 1.7912 | 1.4349
2012-10-30 03:32 | 23.2091 | 13.1449 | 14.1011 | 14.3844 | 13.9918 | 46.7098 | 0.023021 | 0.72099 | 0.016598 | 31.3189 | 0.038382 | 0.72216 | 0.027718 | 18.8151 | 12.6419 | 3.2202 | 0.044316 | 0.503 | 1.4592 | 1.7425 | 1.3499
2012-10-30 03:32 | 23.2173 | 13.1288 | 14.1376 | 14.4331 | 14.0283 | 46.6976 | 0.022993 | 0.72039 | 0.016564 | 31.3307 | 0.038354 | 0.72203 | 0.027693 | 18.8255 | 12.5527 | 3.2193 | 0.044257 | 0.5761 | 1.5849 | 1.8804 | 1.4756
2012-10-30 03:32 | 23.2241 | 13.1964 | 14.1039 | 14.4237 | 14.0189 | 46.7121 | 0.023053 | 0.72159 | 0.016635 | 31.3011 | 0.038379 | 0.72229 | 0.027721 | 18.82 | 12.6285 | 3.2203 | 0.044356 | 0.5679 | 1.4754 | 1.7952 | 1.3904
2012-10-30 03:32 | 23.23 | 13.1496 | 14.1341 | 14.4094 | 14.0612 | 46.6988 | 0.022961 | 0.71955 | 0.016522 | 31.3378 | 0.038263 | 0.7205 | 0.027569 | 18.8302 | 12.5451 | 3.2194 | 0.044091 | 0.6045 | 1.589 | 1.8643 | 1.5161
2012-10-30 03:32 | 23.245 | 13.1646 | 14.1208 | 14.4163 | 14.0155 | 46.7231 | 0.022836 | 0.71656 | 0.016363 | 31.3784 | 0.038222 | 0.71976 | 0.027511 | 18.8311 | 12.6088 | 3.2211 | 0.043874 | 0.5558 | 1.512 | 1.8075 | 1.4067
2012-10-30 03:33 | 23.2427 | 13.2028 | 14.0942 | 14.414 | 13.9889 | 46.7236 | 0.022858 | 0.71713 | 0.016392 | 31.3735 | 0.03831 | 0.72106 | 0.027624 | 18.8218 | 12.6065 | 3.2211 | 0.044016 | 0.5963 | 1.4877 | 1.8075 | 1.3824
2012-10-30 03:33 | 23.2517 | 13.2078 | 14.1558 | 14.3542 | 14.006 | 46.7034 | 0.022876 | 0.71783 | 0.016421 | 31.3791 | 0.038175 | 0.71935 | 0.027461 | 18.8434 | 12.6033 | 3.2197 | 0.043882 | 0.6045 | 1.5525 | 1.7509 | 1.4027
2012-10-30 03:33 | 23.2614 | 13.1607 | 14.1574 | 14.4327 | 14.0319 | 46.7202 | 0.022854 | 0.71705 | 0.016387 | 31.3752 | 0.038181 | 0.71913 | 0.027457 | 18.8347 | 12.5887 | 3.2209 | 0.043844 | 0.572 | 1.5687 | 1.844 | 1.4432
2012-10-30 03:33 | 23.2779 | 13.1853 | 14.1253 | 14.3925 | 14.0403 | 46.7144 | 0.022807 | 0.71604 | 0.016331 | 31.3955 | 0.038074 | 0.71763 | 0.027323 | 18.8482 | 12.5849 | 3.2205 | 0.043654 | 0.6004 | 1.5404 | 1.8076 | 1.4554
2012-10-30 03:33 | 23.2831 | 13.2189 | 14.167 | 14.3734 | 14.0172 | 46.7138 | 0.022782 | 0.71556 | 0.016302 | 31.409 | 0.038216 | 0.71953 | 0.027498 | 18.8281 | 12.6347 | 3.2204 | 0.0438 | 0.5842 | 1.5323 | 1.7387 | 1.3825
2012-10-30 03:33 | 23.2958 | 13.187 | 14.1594 | 14.4225 | 14.0461 | 46.7115 | 0.02276 | 0.71503 | 0.016274 | 31.4159 | 0.038055 | 0.71718 | 0.027292 | 18.8458 | 12.6028 | 3.2203 | 0.043566 | 0.5842 | 1.5566 | 1.8197 | 1.4433
2012-10-30 03:33 | 23.3145 | 13.2503 | 14.1579 | 14.4251 | 13.9757 | 46.704 | 0.022747 | 0.71479 | 0.016259 | 31.4237 | 0.038058 | 0.71719 | 0.027295 | 18.8447 | 12.5931 | 3.2198 | 0.043554 | 0.6572 | 1.5648 | 1.832 | 1.3826
2012-10-30 03:33 | 23.331 | 13.1857 | 14.1987 | 14.4294 | 13.9962 | 46.6918 | 0.022754 | 0.71455 | 0.016259 | 31.4032 | 0.03803 | 0.71692 | 0.027264 | 18.8514 | 12.6583 | 3.2189 | 0.043523 | 0.5274 | 1.5404 | 1.7711 | 1.3379
2012-10-30 03:33 | 23.331 | 13.2141 | 14.1582 | 14.4173 | 14.0529 | 46.678 | 0.022747 | 0.71442 | 0.016251 | 31.4071 | 0.038033 | 0.71686 | 0.027264 | 18.8484 | 12.6258 | 3.218 | 0.043515 | 0.5883 | 1.5324 | 1.7915 | 1.4271
2012-10-30 03:33 | 23.3459 | 13.229 | 14.1974 | 14.42 | 14.0314 | 46.6838 | 0.022747 | 0.71437 | 0.01625 | 31.4052 | 0.038024 | 0.71669 | 0.027251 | 18.8483 | 12.6529 | 3.2184 | 0.043501 | 0.5761 | 1.5445 | 1.7671 | 1.3785
2012-10-30 03:33 | 23.3534 | 13.2 | 14.2049 | 14.4073 | 14.047 | 46.6843 | 0.022735 | 0.71426 | 0.016239 | 31.4168 | 0.038017 | 0.71657 | 0.027242 | 18.8487 | 12.632 | 3.2184 | 0.043481 | 0.568 | 1.5729 | 1.7753 | 1.415
2012-10-30 03:33 | 23.3654 | 13.1876 | 14.2128 | 14.4072 | 14.0428 | 46.6872 | 0.022747 | 0.71414 | 0.016245 | 31.3951 | 0.037995 | 0.71649 | 0.027223 | 18.8574 | 12.6562 | 3.2186 | 0.043468 | 0.5314 | 1.5566 | 1.751 | 1.3866
2012-10-30 03:33 | 23.3743 | 13.2249 | 14.1772 | 14.4241 | 14.0395 | 46.6872 | 0.022732 | 0.71398 | 0.01623 | 31.4088 | 0.038002 | 0.7163 | 0.027221 | 18.8489 | 12.6042 | 3.2186 | 0.043451 | 0.6207 | 1.573 | 1.8199 | 1.4353
2012-10-30 03:33 | 23.39 | 13.196 | 14.1808 | 14.4115 | 14.039 | 46.6814 | 0.022722 | 0.71373 | 0.016217 | 31.4112 | 0.037983 | 0.71608 | 0.027199 | 18.8526 | 12.6402 | 3.2182 | 0.043416 | 0.5558 | 1.5406 | 1.7713 | 1.3988
2012-10-30 03:33 | 23.4095 | 13.2034 | 14.1922 | 14.3865 | 14.0504 | 46.6786 | 0.022707 | 0.71351 | 0.016202 | 31.4224 | 0.03798 | 0.71589 | 0.02719 | 18.8492 | 12.6353 | 3.218 | 0.043392 | 0.5681 | 1.5569 | 1.7512 | 1.4151
2012-10-30 03:33 | 23.4199 | 13.1205 | 14.1985 | 14.3766 | 14.0689 | 46.6768 | 0.022688 | 0.71348 | 0.016187 | 31.4474 | 0.037964 | 0.71576 | 0.027173 | 18.8537 | 12.6254 | 3.2179 | 0.04336 | 0.4951 | 1.5731 | 1.7512 | 1.4435
2012-10-30 03:33 | 23.4379 | 13.1831 | 14.1882 | 14.4432 | 14.0302 | 46.6872 | 0.022697 | 0.71332 | 0.01619 | 31.4279 | 0.037964 | 0.7156 | 0.027167 | 18.8495 | 12.6109 | 3.2186 | 0.043357 | 0.5722 | 1.5773 | 1.8323 | 1.4193
2012-10-30 03:33 | 23.4394 | 13.1805 | 14.2302 | 14.4042 | 14.052 | 46.6791 | 0.022697 | 0.71319 | 0.016187 | 31.4222 | 0.037957 | 0.71556 | 0.027161 | 18.8519 | 12.6774 | 3.2181 | 0.043348 | 0.5031 | 1.5528 | 1.7268 | 1.3746
2012-10-30 03:33 | 23.4506 | 13.2404 | 14.1563 | 14.4195 | 14.0389 | 46.6693 | 0.0227 | 0.71312 | 0.016188 | 31.4149 | 0.037945 | 0.71543 | 0.027147 | 18.8544 | 12.648 | 3.2174 | 0.043335 | 0.5924 | 1.5083 | 1.7715 | 1.3909
2012-10-30 03:33 | 23.4551 | 13.1962 | 14.1689 | 14.3916 | 14.0798 | 46.6878 | 0.022684 | 0.7129 | 0.016171 | 31.4275 | 0.037939 | 0.71526 | 0.027136 | 18.8528 | 12.6566 | 3.2187 | 0.043307 | 0.5396 | 1.5123 | 1.735 | 1.4232
2012-10-30 03:33 | 23.4671 | 13.2082 | 14.2376 | 14.44 | 14.0392 | 46.7051 | 0.022697 | 0.71287 | 0.01618 | 31.4082 | 0.037929 | 0.71521 | 0.027127 | 18.8566 | 12.6848 | 3.2198 | 0.043307 | 0.5234 | 1.5528 | 1.7552 | 1.3544
2012-10-30 03:33 | 23.4715 | 13.2167 | 14.1975 | 14.404 | 14.0476 | 46.6774 | 0.022684 | 0.71279 | 0.016169 | 31.4224 | 0.037923 | 0.7151 | 0.027119 | 18.8566 | 12.6283 | 3.2179 | 0.043288 | 0.5884 | 1.5692 | 1.7757 | 1.4193
2012-10-30 03:33 | 23.4813 | 13.1819 | 14.1546 | 14.3733 | 14.0655 | 46.6959 | 0.022678 | 0.71269 | 0.016162 | 31.4263 | 0.03791 | 0.71501 | 0.027106 | 18.8607 | 12.6259 | 3.2192 | 0.043268 | 0.556 | 1.5287 | 1.7474 | 1.4396
2012-10-30 03:34 | 23.4985 | 13.1788 | 14.1678 | 14.4269 | 14.0746 | 46.6976 | 0.022678 | 0.71254 | 0.016159 | 31.4199 | 0.037904 | 0.71481 | 0.027094 | 18.8584 | 12.6715 | 3.2193 | 0.043253 | 0.5073 | 1.4963 | 1.7554 | 1.4031
2012-10-30 03:34 | 23.5142 | 13.2188 | 14.2078 | 14.4143 | 14.0944 | 46.6947 | 0.022656 | 0.71241 | 0.01614 | 31.4446 | 0.037907 | 0.71485 | 0.027098 | 18.858 | 12.6751 | 3.2191 | 0.043238 | 0.5437 | 1.5327 | 1.7392 | 1.4193
2012-10-30 03:34 | 23.5201 | 13.2126 | 14.2096 | 14.408 | 14.0679 | 46.6838 | 0.022647 | 0.71238 | 0.016133 | 31.4559 | 0.037901 | 0.71461 | 0.027084 | 18.8545 | 12.685 | 3.2184 | 0.043217 | 0.5276 | 1.5246 | 1.723 | 1.3829
2012-10-30 03:34 | 23.5366 | 13.225 | 14.2018 | 14.4286 | 14.0844 | 46.7121 | 0.022662 | 0.71227 | 0.016141 | 31.43 | 0.037888 | 0.71452 | 0.027072 | 18.8587 | 12.7015 | 3.2203 | 0.043213 | 0.5235 | 1.5003 | 1.7271 | 1.3829
2012-10-30 03:34 | 23.5448 | 13.2413 | 14.2181 | 14.4165 | 14.0561 | 46.6843 | 0.022656 | 0.71215 | 0.016134 | 31.4332 | 0.037885 | 0.71449 | 0.027068 | 18.8595 | 12.7381 | 3.2184 | 0.043202 | 0.5032 | 1.48 | 1.6784 | 1.318
2012-10-30 03:34 | 23.55 | 13.2181 | 14.1909 | 14.4055 | 14.1059 | 46.6855 | 0.022659 | 0.71203 | 0.016134 | 31.4239 | 0.037882 | 0.71433 | 0.02706 | 18.8568 | 12.7393 | 3.2185 | 0.043194 | 0.4788 | 1.4516 | 1.6662 | 1.3666
2012-10-30 03:34 | 23.5628 | 13.2147 | 14.2199 | 14.4021 | 14.0903 | 46.6942 | 0.02265 | 0.71195 | 0.016126 | 31.4326 | 0.037882 | 0.7143 | 0.027059 | 18.856 | 12.7561 | 3.2191 | 0.043185 | 0.4586 | 1.4638 | 1.646 | 1.3342
2012-10-30 03:34 | 23.574 | 13.2218 | 14.1987 | 14.4376 | 14.0812 | 46.6953 | 0.022644 | 0.71189 | 0.01612 | 31.4383 | 0.037879 | 0.7142 | 0.027053 | 18.8548 | 12.7592 | 3.2192 | 0.043173 | 0.4626 | 1.4395 | 1.6784 | 1.322
2012-10-30 03:34 | 23.5777 | 13.2012 | 14.247 | 14.3927 | 14.0849 | 46.6728 | 0.022618 | 0.7119 | 0.016102 | 31.4751 | 0.037857 | 0.71409 | 0.027033 | 18.8627 | 12.7304 | 3.2176 | 0.043135 | 0.4708 | 1.5166 | 1.6623 | 1.3545
2012-10-30 03:34 | 23.5874 | 13.2231 | 14.204 | 14.4348 | 14.0784 | 46.7017 | 0.022634 | 0.71177 | 0.01611 | 31.4471 | 0.037857 | 0.71397 | 0.027029 | 18.8597 | 12.7401 | 3.2196 | 0.043139 | 0.483 | 1.4639 | 1.6947 | 1.3383
2012-10-30 03:34 | 23.6039 | 13.2315 | 14.2408 | 14.4351 | 14.0787 | 46.6797 | 0.022637 | 0.71166 | 0.01611 | 31.4379 | 0.037854 | 0.71394 | 0.027026 | 18.8604 | 12.7932 | 3.2181 | 0.043136 | 0.4383 | 1.4476 | 1.6419 | 1.2855
2012-10-30 03:34 | 23.5987 | 13.206 | 14.2153 | 14.3976 | 14.0897 | 46.6878 | 0.022637 | 0.71157 | 0.016108 | 31.434 | 0.037857 | 0.71393 | 0.027027 | 18.8585 | 12.7393 | 3.2187 | 0.043135 | 0.4667 | 1.476 | 1.6583 | 1.3504
2012-10-30 03:34 | 23.6173 | 13.1718 | 14.2623 | 14.4202 | 14.0921 | 46.6786 | 0.022618 | 0.71151 | 0.016093 | 31.4579 | 0.037841 | 0.71383 | 0.027012 | 18.8638 | 12.7092 | 3.218 | 0.043105 | 0.4626 | 1.5531 | 1.711 | 1.3829
2012-10-30 03:34 | 23.6308 | 13.2462 | 14.2231 | 14.4661 | 14.0449 | 46.6843 | 0.022631 | 0.71143 | 0.0161 | 31.436 | 0.037838 | 0.71375 | 0.027007 | 18.8634 | 12.7633 | 3.2184 | 0.043107 | 0.4829 | 1.4598 | 1.7028 | 1.2816
2012-10-30 03:34 | 23.6375 | 13.2042 | 14.1812 | 14.3999 | 14.0921 | 46.6953 | 0.022606 | 0.71141 | 0.016082 | 31.4701 | 0.037838 | 0.71368 | 0.027004 | 18.8615 | 12.7618 | 3.2192 | 0.043086 | 0.4424 | 1.4194 | 1.6381 | 1.3303
2012-10-30 03:34 | 23.645 | 13.163 | 14.2414 | 14.3953 | 14.0915 | 46.682 | 0.022622 | 0.71134 | 0.016092 | 31.4447 | 0.037841 | 0.71368 | 0.027006 | 18.86 | 12.7572 | 3.2183 | 0.043098 | 0.4058 | 1.4842 | 1.6381 | 1.3343
2012-10-30 03:34 | 23.6517 | 13.2103 | 14.2157 | 14.406 | 14.0577 | 46.6924 | 0.022612 | 0.71112 | 0.01608 | 31.449 | 0.037832 | 0.71357 | 0.026996 | 18.8614 | 12.7517 | 3.219 | 0.043076 | 0.4586 | 1.464 | 1.6543 | 1.306
2012-10-30 03:34 | 23.6645 | 13.215 | 14.2163 | 14.4148 | 14.0664 | 46.6878 | 0.022606 | 0.71107 | 0.016074 | 31.4547 | 0.037819 | 0.71345 | 0.026982 | 18.8649 | 12.797 | 3.2187 | 0.043056 | 0.418 | 1.4193 | 1.6178 | 1.2694
2012-10-30 03:34 | 23.6831 | 13.2417 | 14.2268 | 14.4172 | 14.1134 | 46.678 | 0.0226 | 0.71108 | 0.01607 | 31.4638 | 0.037838 | 0.71335 | 0.026992 | 18.8527 | 12.7709 | 3.218 | 0.043062 | 0.4708 | 1.4559 | 1.6463 | 1.3425
2012-10-30 03:34 | 23.6824 | 13.168 | 14.2504 | 14.3719 | 14.1086 | 46.7173 | 0.022606 | 0.71114 | 0.016076 | 31.458 | 0.037816 | 0.71336 | 0.026977 | 18.8641 | 12.7905 | 3.2207 | 0.043053 | 0.3775 | 1.4599 | 1.5814 | 1.3181
2012-10-30 03:34 | 23.6936 | 13.2238 | 14.2292 | 14.4155 | 14.055 | 46.6687 | 0.022606 | 0.71102 | 0.016073 | 31.4528 | 0.037819 | 0.71323 | 0.026974 | 18.8591 | 12.8017 | 3.2173 | 0.043047 | 0.4221 | 1.4275 | 1.6138 | 1.2533
2012-10-30 03:34 | 23.6981 | 13.1877 | 14.258 | 14.3998 | 14.0635 | 46.6716 | 0.022612 | 0.71095 | 0.016076 | 31.4413 | 0.037803 | 0.71332 | 0.026966 | 18.8694 | 12.7494 | 3.2175 | 0.043042 | 0.4383 | 1.5086 | 1.6504 | 1.3141
2012-10-30 03:34 | 23.7108 | 13.2166 | 14.2261 | 14.3801 | 14.0843 | 46.6693 | 0.022615 | 0.71086 | 0.016076 | 31.4333 | 0.037794 | 0.7132 | 0.026955 | 18.8709 | 12.7742 | 3.2174 | 0.043031 | 0.4424 | 1.4519 | 1.6059 | 1.3101
2012-10-30 03:34 | 23.7108 | 13.2004 | 14.218 | 14.4125 | 14.06 | 46.645 | 0.02259 | 0.71091 | 0.016059 | 31.47 | 0.037806 | 0.71319 | 0.026963 | 18.8645 | 12.7417 | 3.2157 | 0.043022 | 0.4587 | 1.4763 | 1.6708 | 1.3183
2012-10-30 03:34 | 23.7258 | 13.2276 | 14.1844 | 14.4072 | 14.0629 | 46.6566 | 0.022596 | 0.71082 | 0.016062 | 31.4578 | 0.037803 | 0.71312 | 0.026958 | 18.8641 | 12.7974 | 3.2165 | 0.04302 | 0.4302 | 1.387 | 1.6098 | 1.2655
2012-10-30 03:34 | 23.7325 | 13.2099 | 14.1749 | 14.4018 | 14.106 | 46.6514 | 0.0226 | 0.71072 | 0.016062 | 31.4478 | 0.03781 | 0.71306 | 0.026961 | 18.859 | 12.7675 | 3.2161 | 0.043023 | 0.4424 | 1.4074 | 1.6343 | 1.3385
2012-10-30 03:34 | 23.7519 | 13.2415 | 14.2672 | 14.4495 | 14.0971 | 46.6416 | 0.02259 | 0.71065 | 0.016054 | 31.4585 | 0.037791 | 0.713 | 0.026945 | 18.867 | 12.8316 | 3.2155 | 0.042999 | 0.4099 | 1.4356 | 1.6179 | 1.2655
2012-10-30 03:34 | 23.7587 | 13.1631 | 14.2659 | 14.3956 | 14.0836 | 46.6635 | 0.022577 | 0.71057 | 0.016043 | 31.4734 | 0.037797 | 0.7129 | 0.026946 | 18.8613 | 12.7815 | 3.217 | 0.042989 | 0.3816 | 1.4844 | 1.6141 | 1.3021
2012-10-30 03:34 | 23.7662 | 13.2031 | 14.2734 | 14.3869 | 14.0465 | 46.6734 | 0.022584 | 0.71056 | 0.016047 | 31.463 | 0.037788 | 0.7128 | 0.026935 | 18.8631 | 12.8053 | 3.2177 | 0.042982 | 0.3978 | 1.4681 | 1.5816 | 1.2412
2012-10-30 03:35 | 23.7699 | 13.2027 | 14.2609 | 14.3946 | 14.1029 | 46.641 | 0.022587 | 0.71055 | 0.016049 | 31.4582 | 0.037788 | 0.71287 | 0.026938 | 18.8651 | 12.8293 | 3.2154 | 0.042987 | 0.3734 | 1.4316 | 1.5653 | 1.2736
2012-10-30 03:35 | 23.7893 | 13.1896 | 14.2439 | 14.4181 | 14.0696 | 46.6624 | 0.022581 | 0.71043 | 0.016042 | 31.4614 | 0.037794 | 0.71277 | 0.026938 | 18.8594 | 12.7999 | 3.2169 | 0.04298 | 0.3897 | 1.444 | 1.6182 | 1.2697
2012-10-30 03:35 | 23.7991 | 13.2116 | 14.2861 | 14.3995 | 14.0997 | 46.6404 | 0.022581 | 0.71037 | 0.016041 | 31.4589 | 0.037769 | 0.71276 | 0.02692 | 18.8715 | 12.8382 | 3.2154 | 0.042961 | 0.3734 | 1.4479 | 1.5613 | 1.2615
2012-10-30 03:35 | 23.8058 | 13.1737 | 14.2401 | 14.4184 | 14.1469 | 46.6439 | 0.022571 | 0.71033 | 0.016033 | 31.4709 | 0.037762 | 0.71274 | 0.026915 | 18.8746 | 12.8002 | 3.2156 | 0.042948 | 0.3735 | 1.4399 | 1.6182 | 1.3467
2012-10-30 03:35 | 23.8148 | 13.1989 | 14.2208 | 14.3869 | 14.1033 | 46.6699 | 0.022584 | 0.71029 | 0.016041 | 31.4508 | 0.037766 | 0.71257 | 0.026911 | 18.868 | 12.8254 | 3.2174 | 0.042952 | 0.3735 | 1.3954 | 1.5615 | 1.2779
2012-10-30 03:35 | 23.829 | 13.205 | 14.2755 | 14.3889 | 14.0931 | 46.6554 | 0.022577 | 0.71021 | 0.016035 | 31.4574 | 0.037769 | 0.71263 | 0.026915 | 18.868 | 12.8031 | 3.2164 | 0.04295 | 0.4019 | 1.4724 | 1.5858 | 1.29
2012-10-30 03:35 | 23.8297 | 13.1895 | 14.2195 | 14.3977 | 14.0776 | 46.6462 | 0.022571 | 0.71023 | 0.016031 | 31.4664 | 0.037759 | 0.71256 | 0.026905 | 18.8711 | 12.7876 | 3.2158 | 0.042936 | 0.4019 | 1.4319 | 1.6101 | 1.29
2012-10-30 03:35 | 23.8417 | 13.1731 | 14.2315 | 14.4219 | 14.0653 | 46.6514 | 0.022584 | 0.71014 | 0.016038 | 31.4445 | 0.037759 | 0.71242 | 0.0269 | 18.8677 | 12.828 | 3.2161 | 0.042938 | 0.3451 | 1.4035 | 1.5939 | 1.2373
2012-10-30 03:35 | 23.8544 | 13.206 | 14.2361 | 14.4224 | 14.078 | 46.6676 | 0.022574 | 0.71016 | 0.016031 | 31.459 | 0.037766 | 0.71245 | 0.026907 | 18.865 | 12.8285 | 3.2173 | 0.042938 | 0.3775 | 1.4076 | 1.5939 | 1.2495
2012-10-30 03:35 | 23.8581 | 13.1773 | 14.2438 | 14.3856 | 14.0939 | 46.6491 | 0.022571 | 0.71005 | 0.016027 | 31.4587 | 0.037747 | 0.71238 | 0.02689 | 18.8725 | 12.7185 | 3.216 | 0.042917 | 0.4588 | 1.5253 | 1.6671 | 1.3754
2012-10-30 03:35 | 23.8716 | 13.1705 | 14.2208 | 14.3829 | 14.0628 | 46.6473 | 0.022568 | 0.71003 | 0.016024 | 31.4616 | 0.03774 | 0.71232 | 0.026883 | 18.8745 | 12.8213 | 3.2159 | 0.042907 | 0.3492 | 1.3995 | 1.5616 | 1.2415
2012-10-30 03:35 | 23.8738 | 13.2052 | 14.2109 | 14.4013 | 14.1136 | 46.6641 | 0.022577 | 0.71 | 0.01603 | 31.4478 | 0.037743 | 0.71232 | 0.026885 | 18.873 | 12.7992 | 3.217 | 0.042915 | 0.406 | 1.4117 | 1.6021 | 1.3144
2012-10-30 03:35 | 23.8821 | 13.1769 | 14.2354 | 14.3974 | 14.0773 | 46.6346 | 0.022555 | 0.70985 | 0.016011 | 31.4721 | 0.037747 | 0.71224 | 0.026885 | 18.8687 | 12.7953 | 3.215 | 0.042896 | 0.3816 | 1.4401 | 1.6021 | 1.282
2012-10-30 03:35 | 23.8955 | 13.1619 | 14.1921 | 14.3582 | 14.0624 | 46.6433 | 0.022559 | 0.70991 | 0.016015 | 31.4691 | 0.037747 | 0.71227 | 0.026886 | 18.8695 | 12.8249 | 3.2156 | 0.042901 | 0.337 | 1.3672 | 1.5333 | 1.2375
2012-10-30 03:35 | 23.9097 | 13.2005 | 14.2468 | 14.421 | 14.1212 | 46.6578 | 0.022552 | 0.70993 | 0.01601 | 31.4795 | 0.037734 | 0.71225 | 0.026876 | 18.8756 | 12.8391 | 3.2166 | 0.042886 | 0.3614 | 1.4077 | 1.5819 | 1.2821
2012-10-30 03:35 | 23.915 | 13.1895 | 14.2237 | 14.402 | 14.0981 | 46.6606 | 0.022555 | 0.7099 | 0.016012 | 31.474 | 0.03774 | 0.71217 | 0.026877 | 18.8703 | 12.8322 | 3.2168 | 0.042889 | 0.3573 | 1.3915 | 1.5698 | 1.2659
2012-10-30 03:35 | 23.9299 | 13.2004 | 14.2265 | 14.4007 | 14.0806 | 46.641 | 0.022555 | 0.7099 | 0.016012 | 31.474 | 0.03774 | 0.71218 | 0.026878 | 18.8707 | 12.8065 | 3.2154 | 0.04289 | 0.3939 | 1.42 | 1.5942 | 1.2741
2012-10-30 03:35 | 23.9396 | 13.1695 | 14.2524 | 14.3982 | 14.0659 | 46.6445 | 0.022562 | 0.70977 | 0.016014 | 31.4585 | 0.037747 | 0.71219 | 0.026883 | 18.8675 | 12.8284 | 3.2157 | 0.042897 | 0.3411 | 1.424 | 1.5698 | 1.2375
2012-10-30 03:35 | 23.9456 | 13.1836 | 14.2259 | 14.4204 | 14.1368 | 46.6653 | 0.02254 | 0.70969 | 0.015997 | 31.486 | 0.037731 | 0.71209 | 0.026868 | 18.8729 | 12.8344 | 3.2171 | 0.042865 | 0.3492 | 1.3915 | 1.586 | 1.3024
2012-10-30 03:35 | 23.9539 | 13.2081 | 14.218 | 14.4166 | 14.064 | 46.6468 | 0.022559 | 0.70975 | 0.016011 | 31.462 | 0.037743 | 0.71211 | 0.026877 | 18.8673 | 12.8468 | 3.2158 | 0.042888 | 0.3613 | 1.3712 | 1.5698 | 1.2172
2012-10-30 03:35 | 23.9554 | 13.169 | 14.2276 | 14.414 | 14.0493 | 46.6508 | 0.022552 | 0.70954 | 0.016001 | 31.4622 | 0.037721 | 0.71195 | 0.026855 | 18.874 | 12.8361 | 3.2161 | 0.042856 | 0.3329 | 1.3915 | 1.5779 | 1.2132
2012-10-30 03:35 | 23.9733 | 13.2194 | 14.282 | 14.3914 | 14.0794 | 46.6485 | 0.022543 | 0.70955 | 0.015995 | 31.4754 | 0.037743 | 0.7119 | 0.026869 | 18.8619 | 12.8702 | 3.2159 | 0.042864 | 0.3492 | 1.4118 | 1.5212 | 1.2092
2012-10-30 03:35 | 23.9815 | 13.2195 | 14.2456 | 14.4118 | 14.12 | 46.6335 | 0.022546 | 0.70962 | 0.015999 | 31.4744 | 0.037725 | 0.71185 | 0.026854 | 18.8694 | 12.8703 | 3.2149 | 0.042853 | 0.3492 | 1.3753 | 1.5415 | 1.2497
2012-10-30 03:35 | 23.9912 | 13.1967 | 14.2391 | 14.3931 | 14.0932 | 46.652 | 0.022543 | 0.70955 | 0.015995 | 31.4754 | 0.037737 | 0.71186 | 0.026863 | 18.8637 | 12.8841 | 3.2162 | 0.042858 | 0.3126 | 1.355 | 1.509 | 1.2091
2012-10-30 03:35 | 23.9965 | 13.1898 | 14.2242 | 14.3782 | 14.0904 | 46.6439 | 0.022521 | 0.70943 | 0.015977 | 31.501 | 0.037715 | 0.71185 | 0.026847 | 18.8744 | 12.8 | 3.2156 | 0.042824 | 0.3898 | 1.4242 | 1.5782 | 1.2904
2012-10-30 03:35 | 24.0084 | 13.1774 | 14.2361 | 14.3901 | 14.082 | 46.6335 | 0.022533 | 0.70946 | 0.015986 | 31.4855 | 0.037709 | 0.71172 | 0.026838 | 18.8739 | 12.8932 | 3.2149 | 0.042824 | 0.2842 | 1.3429 | 1.4969 | 1.1888
2012-10-30 03:35 | 24.0212 | 13.1658 | 14.2164 | 14.4312 | 14.0867 | 46.641 | 0.022549 | 0.70943 | 0.015997 | 31.4619 | 0.037712 | 0.71176 | 0.026842 | 18.8736 | 12.8491 | 3.2154 | 0.042839 | 0.3167 | 1.3673 | 1.5821 | 1.2376
2012-10-30 03:36 | 24.0242 | 13.1972 | 14.2478 | 14.3937 | 14.0938 | 46.652 | 0.022543 | 0.70936 | 0.015991 | 31.4671 | 0.037709 | 0.7118 | 0.026841 | 18.8762 | 12.8602 | 3.2162 | 0.042832 | 0.337 | 1.3876 | 1.5335 | 1.2336
2012-10-30 03:36 | 24.0234 | 13.1599 | 14.2835 | 14.3321 | 14.0849 | 46.6161 | 0.022537 | 0.7094 | 0.015988 | 31.4773 | 0.037715 | 0.7117 | 0.026842 | 18.8705 | 12.8838 | 3.2137 | 0.04283 | 0.2761 | 1.3997 | 1.4483 | 1.2011
2012-10-30 03:36 | 24.0346 | 13.1752 | 14.2258 | 14.3919 | 14.088 | 46.6416 | 0.022527 | 0.70929 | 0.015978 | 31.4862 | 0.037718 | 0.71169 | 0.026843 | 18.8687 | 12.8828 | 3.2155 | 0.042821 | 0.2924 | 1.343 | 1.5091 | 1.2052
2012-10-30 03:36 | 24.0421 | 13.1137 | 14.2495 | 14.4075 | 14.0387 | 46.6595 | 0.022518 | 0.7093 | 0.015972 | 31.4994 | 0.037703 | 0.71164 | 0.026831 | 18.875 | 12.8741 | 3.2167 | 0.042803 | 0.2396 | 1.3754 | 1.5334 | 1.1646
2012-10-30 03:36 | 24.0548 | 13.171 | 14.2419 | 14.404 | 14.0919 | 46.6369 | 0.022527 | 0.70926 | 0.015978 | 31.4849 | 0.037703 | 0.71156 | 0.026828 | 18.8727 | 12.9071 | 3.2151 | 0.042806 | 0.2639 | 1.3348 | 1.4969 | 1.1848
2012-10-30 03:36 | 24.0668 | 13.1993 | 14.2418 | 14.3836 | 14.0431 | 46.6393 | 0.022527 | 0.70923 | 0.015977 | 31.4836 | 0.037687 | 0.71151 | 0.026815 | 18.8796 | 12.8906 | 3.2153 | 0.042792 | 0.3087 | 1.3512 | 1.493 | 1.1525
2012-10-30 03:36 | 24.0683 | 13.1723 | 14.2595 | 14.3446 | 14.0568 | 46.6635 | 0.022524 | 0.70923 | 0.015975 | 31.4878 | 0.037699 | 0.71156 | 0.026825 | 18.8747 | 12.9043 | 3.217 | 0.0428 | 0.268 | 1.3552 | 1.4403 | 1.1525
2012-10-30 03:36 | 24.0832 | 13.171 | 14.242 | 14.4122 | 14.0636 | 46.6346 | 0.02253 | 0.70917 | 0.015978 | 31.4769 | 0.037706 | 0.71151 | 0.026828 | 18.87 | 12.8624 | 3.215 | 0.042806 | 0.3086 | 1.3796 | 1.5498 | 1.2012
2012-10-30 03:36 | 24.0922 | 13.1759 | 14.2631 | 14.3847 | 14.101 | 46.6294 | 0.022524 | 0.70909 | 0.015971 | 31.4814 | 0.037687 | 0.71147 | 0.026813 | 18.8784 | 12.8795 | 3.2146 | 0.042784 | 0.2964 | 1.3836 | 1.5052 | 1.2215
2012-10-30 03:36 | 24.1019 | 13.1491 | 14.3012 | 14.4147 | 14.0904 | 46.6381 | 0.022521 | 0.7092 | 0.015972 | 31.4907 | 0.037706 | 0.71146 | 0.026826 | 18.8685 | 12.9095 | 3.2152 | 0.042798 | 0.2396 | 1.3917 | 1.5052 | 1.1809
2012-10-30 03:36 | 24.1102 | 13.2021 | 14.2609 | 14.3987 | 14.046 | 46.6381 | 0.022533 | 0.70904 | 0.015977 | 31.4669 | 0.037699 | 0.71135 | 0.026817 | 18.8693 | 12.9015 | 3.2152 | 0.042794 | 0.3006 | 1.3594 | 1.4972 | 1.1445
2012-10-30 03:36 | 24.1139 | 13.1773 | 14.2443 | 14.4024 | 14.0537 | 46.63 | 0.02253 | 0.709 | 0.015974 | 31.4692 | 0.03769 | 0.71138 | 0.026812 | 18.8746 | 12.9296 | 3.2147 | 0.042786 | 0.2477 | 1.3147 | 1.4728 | 1.1241
2012-10-30 03:36 | 24.1229 | 13.1335 | 14.2614 | 14.3749 | 14.0992 | 46.6433 | 0.022521 | 0.70906 | 0.015969 | 31.4843 | 0.037687 | 0.71137 | 0.026809 | 18.8757 | 12.8939 | 3.2156 | 0.042778 | 0.2396 | 1.3675 | 1.481 | 1.2053
2012-10-30 03:36 | 24.1296 | 13.1362 | 14.2316 | 14.3775 | 14.0816 | 46.6219 | 0.022527 | 0.70903 | 0.015972 | 31.4746 | 0.037706 | 0.71138 | 0.026823 | 18.8666 | 12.9534 | 3.2141 | 0.042795 | 0.1828 | 1.2782 | 1.4241 | 1.1282
2012-10-30 03:36 | 24.1431 | 13.1903 | 14.2532 | 14.3748 | 14.0951 | 46.6375 | 0.022527 | 0.70894 | 0.01597 | 31.4708 | 0.037693 | 0.71138 | 0.026814 | 18.8731 | 12.9547 | 3.2152 | 0.042784 | 0.2356 | 1.2985 | 1.4201 | 1.1404
2012-10-30 03:36 | 24.1528 | 13.2 | 14.271 | 14.3764 | 14.0278 | 46.6433 | 0.022515 | 0.709 | 0.015963 | 31.4901 | 0.037687 | 0.71133 | 0.026808 | 18.8746 | 12.9401 | 3.2156 | 0.042771 | 0.2599 | 1.3309 | 1.4363 | 1.0877
2012-10-30 03:36 | 24.1535 | 13.156 | 14.215 | 14.3487 | 14.1136 | 46.6346 | 0.022518 | 0.70899 | 0.015965 | 31.4853 | 0.03769 | 0.71131 | 0.026809 | 18.8727 | 12.9286 | 3.215 | 0.042774 | 0.2274 | 1.2864 | 1.4201 | 1.185
2012-10-30 03:36 | 24.1662 | 13.1484 | 14.2277 | 14.3493 | 14.1182 | 46.6329 | 0.022515 | 0.70891 | 0.015961 | 31.4863 | 0.03769 | 0.71131 | 0.026809 | 18.8727 | 12.9454 | 3.2149 | 0.04277 | 0.203 | 1.2823 | 1.4039 | 1.1728
2012-10-30 03:36 | 24.1745 | 13.1242 | 14.2684 | 14.4022 | 14.1103 | 46.6387 | 0.022508 | 0.70886 | 0.015955 | 31.4935 | 0.037681 | 0.71121 | 0.026799 | 18.8745 | 12.9049 | 3.2153 | 0.042754 | 0.2193 | 1.3635 | 1.4973 | 1.2054
2012-10-30 03:36 | 24.1849 | 13.1752 | 14.2666 | 14.3964 | 14.1166 | 46.6271 | 0.022518 | 0.70881 | 0.015961 | 31.4776 | 0.037681 | 0.71124 | 0.0268 | 18.8753 | 12.9153 | 3.2145 | 0.042761 | 0.2599 | 1.3513 | 1.4811 | 1.2013
2012-10-30 03:36 | 24.1849 | 13.1306 | 14.2261 | 14.3761 | 14.068 | 46.6398 | 0.022527 | 0.70886 | 0.015968 | 31.467 | 0.037662 | 0.71118 | 0.026785 | 18.8833 | 12.8625 | 3.2153 | 0.042753 | 0.2681 | 1.3636 | 1.5136 | 1.2055
2012-10-30 03:36 | 24.2007 | 13.1139 | 14.2419 | 14.3959 | 14.1081 | 46.6317 | 0.022515 | 0.70881 | 0.015959 | 31.4818 | 0.037665 | 0.71114 | 0.026785 | 18.8806 | 12.9474 | 3.2148 | 0.042744 | 0.1665 | 1.2945 | 1.4485 | 1.1607
2012-10-30 03:36 | 24.1992 | 13.1327 | 14.2363 | 14.3458 | 14.0782 | 46.6416 | 0.022521 | 0.70887 | 0.015964 | 31.476 | 0.037674 | 0.71117 | 0.026793 | 18.8769 | 12.889 | 3.2155 | 0.042757 | 0.2437 | 1.3473 | 1.4568 | 1.1892
2012-10-30 03:36 | 24.2193 | 13.1487 | 14.2362 | 14.3659 | 14.0943 | 46.6161 | 0.022502 | 0.70874 | 0.015948 | 31.4968 | 0.037668 | 0.71117 | 0.026788 | 18.8799 | 12.901 | 3.2137 | 0.042736 | 0.2477 | 1.3352 | 1.4649 | 1.1933
2012-10-30 03:36 | 24.2291 | 13.1423 | 14.2379 | 14.3676 | 14.0838 | 46.6317 | 0.022511 | 0.70873 | 0.015954 | 31.4835 | 0.037677 | 0.71107 | 0.026791 | 18.8727 | 12.8986 | 3.2148 | 0.042745 | 0.2437 | 1.3393 | 1.469 | 1.1852
2012-10-30 03:37 | 24.2523 | 13.1371 | 14.2286 | 14.3746 | 14.0908 | 46.6375 | 0.022493 | 0.70873 | 0.015941 | 31.5087 | 0.037662 | 0.71102 | 0.026779 | 18.8791 | 12.9177 | 3.2152 | 0.04272 | 0.2194 | 1.3109 | 1.4569 | 1.1731
2012-10-30 03:37 | 24.2545 | 13.1758 | 14.2308 | 14.3646 | 14.0808 | 46.63 | 0.022515 | 0.7087 | 0.015956 | 31.4767 | 0.037674 | 0.71112 | 0.026791 | 18.8757 | 12.8915 | 3.2147 | 0.042747 | 0.2843 | 1.3393 | 1.4731 | 1.1893
2012-10-30 03:37 | 24.2695 | 13.134 | 14.2539 | 14.331 | 14.0917 | 46.6225 | 0.022502 | 0.70868 | 0.015947 | 31.4942 | 0.037659 | 0.71094 | 0.026773 | 18.8783 | 12.939 | 3.2142 | 0.04272 | 0.195 | 1.3149 | 1.392 | 1.1527
2012-10-30 03:37 | 24.2702 | 13.155 | 14.2182 | 14.3479 | 14.0722 | 46.6566 | 0.022518 | 0.70873 | 0.015959 | 31.4738 | 0.037671 | 0.71105 | 0.026786 | 18.8753 | 12.9356 | 3.2165 | 0.042745 | 0.2194 | 1.2826 | 1.4123 | 1.1366
2012-10-30 03:37 | 24.2725 | 13.1532 | 14.2245 | 14.334 | 14.1069 | 46.6231 | 0.022505 | 0.70858 | 0.015947 | 31.4855 | 0.037674 | 0.71099 | 0.026786 | 18.8723 | 12.8932 | 3.2142 | 0.042733 | 0.26 | 1.3313 | 1.4408 | 1.2137
2012-10-30 03:37 | 24.2889 | 13.1777 | 14.2733 | 14.3342 | 14.0787 | 46.6283 | 0.022483 | 0.7086 | 0.015931 | 31.517 | 0.037668 | 0.71099 | 0.026782 | 18.8753 | 12.9462 | 3.2146 | 0.042713 | 0.2315 | 1.3271 | 1.388 | 1.1325
2012-10-30 03:37 | 24.3016 | 13.1701 | 14.2374 | 14.3793 | 14.0549 | 46.6294 | 0.022496 | 0.70855 | 0.01594 | 31.4968 | 0.037652 | 0.71094 | 0.026768 | 18.8818 | 12.902 | 3.2146 | 0.042708 | 0.2681 | 1.3354 | 1.4773 | 1.1529
2012-10-30 03:37 | 24.3016 | 13.1173 | 14.2739 | 14.3469 | 14.0995 | 46.619 | 0.022496 | 0.7086 | 0.015941 | 31.4988 | 0.037668 | 0.71098 | 0.026781 | 18.8749 | 12.9183 | 3.2139 | 0.042722 | 0.199 | 1.3556 | 1.4286 | 1.1812
2012-10-30 03:37 | 24.3121 | 13.1278 | 14.2398 | 14.3533 | 14.11 | 46.6208 | 0.022502 | 0.70848 | 0.015942 | 31.4852 | 0.037659 | 0.71098 | 0.026775 | 18.8794 | 12.9613 | 3.214 | 0.042717 | 0.1665 | 1.2785 | 1.392 | 1.1487
2012-10-30 03:37 | 24.3188 | 13.1873 | 14.2303 | 14.3276 | 14.0802 | 46.6369 | 0.022499 | 0.7086 | 0.015943 | 31.4946 | 0.037652 | 0.71092 | 0.026768 | 18.8814 | 12.9355 | 3.2151 | 0.042711 | 0.2518 | 1.2948 | 1.3921 | 1.1447
2012-10-30 03:37 | 24.3278 | 13.1435 | 14.2514 | 14.3325 | 14.0649 | 46.6167 | 0.022508 | 0.70855 | 0.015948 | 31.4801 | 0.037662 | 0.71094 | 0.026775 | 18.8767 | 12.9241 | 3.2138 | 0.042723 | 0.2194 | 1.3273 | 1.4084 | 1.1408
2012-10-30 03:37 | 24.3383 | 13.1093 | 14.2538 | 14.339 | 14.0794 | 46.6219 | 0.022496 | 0.70855 | 0.01594 | 31.4968 | 0.037662 | 0.71082 | 0.026771 | 18.8737 | 12.959 | 3.2141 | 0.042711 | 0.1503 | 1.2948 | 1.38 | 1.1204
2012-10-30 03:37 | 24.3457 | 13.1695 | 14.2247 | 14.3504 | 14.0828 | 46.6375 | 0.022499 | 0.70844 | 0.015939 | 31.4875 | 0.037655 | 0.71083 | 0.026766 | 18.8776 | 12.9258 | 3.2152 | 0.042705 | 0.2437 | 1.2989 | 1.4246 | 1.157
2012-10-30 03:37 | 24.3517 | 13.1674 | 14.2267 | 14.3686 | 14.0766 | 46.6277 | 0.022489 | 0.70845 | 0.015932 | 31.5021 | 0.037643 | 0.71079 | 0.026756 | 18.8824 | 12.9196 | 3.2145 | 0.042688 | 0.2478 | 1.3071 | 1.449 | 1.157
2012-10-30 03:37 | 24.36 | 13.1188 | 14.2147 | 14.3201 | 14.0362 | 46.619 | 0.022496 | 0.70845 | 0.015937 | 31.4923 | 0.037659 | 0.71081 | 0.026768 | 18.8748 | 12.9035 | 3.2139 | 0.042705 | 0.2153 | 1.3112 | 1.4166 | 1.1327
2012-10-30 03:37 | 24.3555 | 13.1143 | 14.2061 | 14.344 | 14.0561 | 46.6317 | 0.022502 | 0.70845 | 0.015942 | 31.4839 | 0.037649 | 0.71076 | 0.02676 | 18.8787 | 12.9031 | 3.2148 | 0.042702 | 0.2112 | 1.303 | 1.4409 | 1.153
2012-10-30 03:37 | 24.3689 | 13.1602 | 14.252 | 14.3128 | 14.0532 | 46.6225 | 0.022499 | 0.70837 | 0.015938 | 31.4843 | 0.037659 | 0.71073 | 0.026766 | 18.8729 | 12.9368 | 3.2142 | 0.042704 | 0.2234 | 1.3152 | 1.376 | 1.1164
2012-10-30 03:37 | 24.3712 | 13.1341 | 14.2502 | 14.3759 | 14.0474 | 46.6277 | 0.022499 | 0.70839 | 0.015938 | 31.4856 | 0.037643 | 0.71078 | 0.026756 | 18.8821 | 12.9391 | 3.2145 | 0.042694 | 0.195 | 1.3111 | 1.4368 | 1.1083
2012-10-30 03:37 | 24.3809 | 13.1601 | 14.2153 | 14.337 | 14.0531 | 46.6317 | 0.022489 | 0.70834 | 0.01593 | 31.497 | 0.037652 | 0.71069 | 0.026759 | 18.8753 | 12.9569 | 3.2148 | 0.042689 | 0.2032 | 1.2584 | 1.3801 | 1.0962
2012-10-30 03:37 | 24.4011 | 13.1031 | 14.2315 | 14.3734 | 14.0733 | 46.6133 | 0.022496 | 0.70837 | 0.015935 | 31.4885 | 0.037643 | 0.71068 | 0.026752 | 18.8794 | 12.9365 | 3.2135 | 0.042687 | 0.1666 | 1.295 | 1.4369 | 1.1368
2012-10-30 03:37 | 24.3959 | 13.0979 | 14.2263 | 14.3601 | 14.0681 | 46.6283 | 0.022493 | 0.70829 | 0.015932 | 31.4894 | 0.037636 | 0.71069 | 0.026748 | 18.8833 | 12.9394 | 3.2146 | 0.04268 | 0.1585 | 1.2869 | 1.4207 | 1.1287
2012-10-30 03:37 | 24.4078 | 13.1016 | 14.1976 | 14.3395 | 14.0475 | 46.6208 | 0.022502 | 0.70828 | 0.015938 | 31.4762 | 0.037649 | 0.71072 | 0.026758 | 18.8775 | 12.9351 | 3.214 | 0.042696 | 0.1665 | 1.2625 | 1.4044 | 1.1124
2012-10-30 03:37 | 24.4101 | 13.1202 | 14.1837 | 14.3297 | 14.062 | 46.6248 | 0.022489 | 0.70829 | 0.015929 | 31.495 | 0.037646 | 0.71063 | 0.026752 | 18.8767 | 12.917 | 3.2143 | 0.042681 | 0.2032 | 1.2667 | 1.4127 | 1.145
2012-10-30 03:37 | 24.4258 | 13.1765 | 14.2318 | 14.3697 | 14.0534 | 46.6237 | 0.02248 | 0.70829 | 0.015922 | 31.5077 | 0.037643 | 0.71068 | 0.026752 | 18.8794 | 12.8921 | 3.2142 | 0.042674 | 0.2844 | 1.3397 | 1.4776 | 1.1613
2012-10-30 03:37 | 24.4265 | 13.1488 | 14.2325 | 14.3501 | 14.0662 | 46.6323 | 0.022483 | 0.70825 | 0.015924 | 31.5015 | 0.037643 | 0.71059 | 0.026749 | 18.8771 | 12.9294 | 3.2148 | 0.042673 | 0.2194 | 1.3031 | 1.4207 | 1.1368
2012-10-30 03:37 | 24.4318 | 13.1581 | 14.2135 | 14.3473 | 14.0472 | 46.6069 | 0.02248 | 0.70826 | 0.015922 | 31.5064 | 0.037649 | 0.71059 | 0.026753 | 18.8741 | 12.9509 | 3.2131 | 0.042675 | 0.2072 | 1.2626 | 1.3964 | 1.0963
2012-10-30 03:37 | 24.4497 | 13.1314 | 14.2476 | 14.3084 | 14.057 | 46.6208 | 0.022486 | 0.70819 | 0.015924 | 31.4948 | 0.037652 | 0.71057 | 0.026755 | 18.8722 | 12.9607 | 3.214 | 0.042679 | 0.1707 | 1.2869 | 1.3477 | 1.0963
2012-10-30 03:37 | 24.4542 | 13.1399 | 14.2156 | 14.3332 | 14.0452 | 46.6265 | 0.02248 | 0.70812 | 0.015919 | 31.5 | 0.03763 | 0.71049 | 0.026736 | 18.8809 | 12.9571 | 3.2144 | 0.042655 | 0.1828 | 1.2585 | 1.3761 | 1.0881
2012-10-30 03:38 | 24.4699 | 13.1475 | 14.2232 | 14.353 | 14.0447 | 46.6196 | 0.022493 | 0.70816 | 0.015929 | 31.4837 | 0.037633 | 0.71053 | 0.026739 | 18.8806 | 12.9525 | 3.214 | 0.042668 | 0.195 | 1.2707 | 1.4005 | 1.0922
2012-10-30 03:38 | 24.4759 | 13.1616 | 14.2251 | 14.3225 | 14.0669 | 46.5942 | 0.022486 | 0.70818 | 0.015924 | 31.4941 | 0.03763 | 0.7105 | 0.026736 | 18.8813 | 12.995 | 3.2122 | 0.04266 | 0.1666 | 1.2301 | 1.3275 | 1.0719
2012-10-30 03:38 | 24.4722 | 13.0807 | 14.2174 | 14.3147 | 14.0267 | 46.6306 | 0.022483 | 0.70815 | 0.015921 | 31.497 | 0.037633 | 0.71052 | 0.026739 | 18.8802 | 12.9873 | 3.2147 | 0.04266 | 0.0934 | 1.2301 | 1.3274 | 1.0394
2012-10-30 03:38 | 24.4901 | 13.1596 | 14.1988 | 14.3042 | 14.0771 | 46.6208 | 0.022486 | 0.70815 | 0.015923 | 31.4928 | 0.03764 | 0.7105 | 0.026743 | 18.8763 | 12.9767 | 3.214 | 0.042666 | 0.1829 | 1.2221 | 1.3275 | 1.1004
2012-10-30 03:38 | 24.4976 | 13.1061 | 14.259 | 14.3442 | 14.04 | 46.6098 | 0.022483 | 0.70821 | 0.015923 | 31.4996 | 0.03764 | 0.71059 | 0.026747 | 18.8786 | 12.9761 | 3.2133 | 0.04267 | 0.13 | 1.2829 | 1.3681 | 1.0639
2012-10-30 03:38 | 24.4961 | 13.1046 | 14.2048 | 14.3143 | 14.0263 | 46.6248 | 0.022486 | 0.70813 | 0.015923 | 31.4922 | 0.037627 | 0.71047 | 0.026733 | 18.882 | 12.9624 | 3.2143 | 0.042656 | 0.1422 | 1.2424 | 1.3519 | 1.0639
2012-10-30 03:38 | 24.514 | 13.1144 | 14.2349 | 14.4133 | 14.0401 | 46.6219 | 0.022477 | 0.70803 | 0.015914 | 31.5003 | 0.03763 | 0.71044 | 0.026734 | 18.8798 | 12.9356 | 3.2141 | 0.042648 | 0.1788 | 1.2993 | 1.4777 | 1.1045
2012-10-30 03:38 | 24.5133 | 13.1096 | 14.2179 | 14.295 | 14.0232 | 46.6075 | 0.022483 | 0.70815 | 0.015921 | 31.497 | 0.037627 | 0.71053 | 0.026735 | 18.8836 | 12.9959 | 3.2131 | 0.042656 | 0.1137 | 1.222 | 1.2991 | 1.0273
2012-10-30 03:38 | 24.5215 | 13.0935 | 14.218 | 14.3235 | 14.003 | 46.6202 | 0.02248 | 0.70822 | 0.015921 | 31.5044 | 0.037646 | 0.71062 | 0.026752 | 18.8763 | 12.9837 | 3.214 | 0.042673 | 0.1098 | 1.2343 | 1.3398 | 1.0193
2012-10-30 03:38 | 24.5238 | 13.0876 | 14.1919 | 14.3136 | 14.0459 | 46.6167 | 0.022483 | 0.70806 | 0.015919 | 31.4932 | 0.037633 | 0.71053 | 0.026739 | 18.8806 | 13.0023 | 3.2138 | 0.042658 | 0.0853 | 1.1896 | 1.3113 | 1.0436
2012-10-30 03:38 | 24.535 | 13.0826 | 14.2031 | 14.2964 | 14.0408 | 46.6115 | 0.022508 | 0.70813 | 0.015939 | 31.4614 | 0.037636 | 0.71049 | 0.02674 | 18.8779 | 12.9363 | 3.2134 | 0.042679 | 0.1463 | 1.2668 | 1.3601 | 1.1045
2012-10-30 03:38 | 24.547 | 13.0783 | 14.1827 | 14.3125 | 14.0366 | 46.608 | 0.02248 | 0.70818 | 0.01592 | 31.5025 | 0.03763 | 0.71046 | 0.026735 | 18.8801 | 12.9727 | 3.2132 | 0.042655 | 0.1056 | 1.21 | 1.3398 | 1.0639
2012-10-30 03:38 | 24.5477 | 13.0912 | 14.1874 | 14.3172 | 14.0373 | 46.6133 | 0.02248 | 0.70819 | 0.01592 | 31.5032 | 0.037636 | 0.71046 | 0.026739 | 18.8771 | 12.9652 | 3.2135 | 0.042659 | 0.126 | 1.2222 | 1.352 | 1.0721
2012-10-30 03:38 | 24.5552 | 13.1109 | 14.2071 | 14.3085 | 14.0326 | 46.6069 | 0.022477 | 0.70812 | 0.015916 | 31.5042 | 0.037633 | 0.7105 | 0.026738 | 18.8798 | 13.0174 | 3.2131 | 0.042654 | 0.0935 | 1.1897 | 1.2911 | 1.0152
2012-10-30 03:38 | 24.5559 | 13.0547 | 14.2037 | 14.2889 | 14.0577 | 46.5948 | 0.022493 | 0.70815 | 0.015928 | 31.483 | 0.037621 | 0.71047 | 0.026729 | 18.885 | 12.9816 | 3.2122 | 0.042657 | 0.0731 | 1.2221 | 1.3073 | 1.0761
2012-10-30 03:38 | 24.5746 | 13.0937 | 14.2062 | 14.2955 | 14.0439 | 46.5976 | 0.022489 | 0.70819 | 0.015927 | 31.4906 | 0.037621 | 0.71047 | 0.026729 | 18.885 | 12.9921 | 3.2124 | 0.042656 | 0.1016 | 1.2141 | 1.3034 | 1.0518
2012-10-30 03:38 | 24.5814 | 13.0721 | 14.2292 | 14.286 | 14.0588 | 46.6121 | 0.022477 | 0.70805 | 0.015915 | 31.501 | 0.037636 | 0.71044 | 0.026738 | 18.8767 | 12.9867 | 3.2134 | 0.042653 | 0.0854 | 1.2425 | 1.2993 | 1.0721
2012-10-30 03:38 | 24.5836 | 13.119 | 14.2558 | 14.3328 | 14.0367 | 46.6011 | 0.022486 | 0.70803 | 0.015921 | 31.4877 | 0.037636 | 0.71044 | 0.026738 | 18.8767 | 12.9808 | 3.2127 | 0.042659 | 0.1382 | 1.275 | 1.352 | 1.0559
2012-10-30 03:38 | 24.5903 | 13.081 | 14.1935 | 14.2706 | 14.0312 | 46.5948 | 0.022477 | 0.70802 | 0.015914 | 31.4997 | 0.037621 | 0.71043 | 0.026727 | 18.8839 | 13.0322 | 3.2122 | 0.042641 | 0.0488 | 1.1613 | 1.2384 | 0.999
2012-10-30 03:38 | 24.5933 | 13.0962 | 14.2087 | 14.3263 | 14.018 | 46.6046 | 0.02247 | 0.70806 | 0.01591 | 31.5114 | 0.037633 | 0.71046 | 0.026737 | 18.8786 | 12.9743 | 3.2129 | 0.042647 | 0.1219 | 1.2344 | 1.352 | 1.0437
2012-10-30 03:38 | 24.6016 | 13.1085 | 14.221 | 14.3103 | 14.0101 | 46.5971 | 0.022477 | 0.708 | 0.015914 | 31.499 | 0.037627 | 0.71029 | 0.026726 | 18.877 | 12.9866 | 3.2124 | 0.04264 | 0.1219 | 1.2344 | 1.3237 | 1.0235
2012-10-30 03:38 | 24.6053 | 13.0554 | 14.1923 | 14.2815 | 14.0543 | 46.5919 | 0.022502 | 0.70802 | 0.015932 | 31.4647 | 0.037633 | 0.71034 | 0.026732 | 18.8756 | 13.0188 | 3.212 | 0.042664 | 0.0366 | 1.1735 | 1.2627 | 1.0355
2012-10-30 03:38 | 24.6218 | 13.1003 | 14.1926 | 14.3061 | 14.0424 | 46.5994 | 0.022464 | 0.70805 | 0.015906 | 31.5192 | 0.037636 | 0.71032 | 0.026733 | 18.8733 | 12.9987 | 3.2126 | 0.042639 | 0.1016 | 1.1939 | 1.3074 | 1.0437
2012-10-30 03:38 | 24.6218 | 13.1166 | 14.2494 | 14.3102 | 13.9897 | 46.5971 | 0.022493 | 0.70803 | 0.015926 | 31.4779 | 0.03763 | 0.7104 | 0.026732 | 18.8786 | 12.9946 | 3.2124 | 0.042658 | 0.122 | 1.2548 | 1.3156 | 0.9951
2012-10-30 03:38 | 24.627 | 13.0892 | 14.1896 | 14.2951 | 14.0071 | 46.5855 | 0.022474 | 0.70798 | 0.015911 | 31.502 | 0.037633 | 0.71042 | 0.026735 | 18.8775 | 12.9876 | 3.2116 | 0.042646 | 0.1016 | 1.202 | 1.3075 | 1.0195
2012-10-30 03:38 | 24.6345 | 13.0927 | 14.1931 | 14.2864 | 14.0714 | 46.5959 | 0.022474 | 0.70798 | 0.015911 | 31.502 | 0.037627 | 0.71033 | 0.026728 | 18.8782 | 12.9382 | 3.2123 | 0.042639 | 0.1545 | 1.2549 | 1.3482 | 1.1332
2012-10-30 03:38 | 24.6337 | 13.0634 | 14.1801 | 14.2572 | 14.0503 | 46.6086 | 0.022483 | 0.70796 | 0.015917 | 31.4887 | 0.037633 | 0.71039 | 0.026734 | 18.8767 | 12.9903 | 3.2132 | 0.042651 | 0.0731 | 1.1898 | 1.2669 | 1.06
2012-10-30 03:38 | 24.6517 | 13.0855 | 14.1575 | 14.2711 | 14.0399 | 46.5913 | 0.02248 | 0.70796 | 0.015915 | 31.4929 | 0.037621 | 0.71034 | 0.026724 | 18.8816 | 12.9636 | 3.212 | 0.042639 | 0.1219 | 1.1939 | 1.3075 | 1.0763
2012-10-30 03:39 | 24.6577 | 13.0752 | 14.196 | 14.2771 | 14.0053 | 46.5965 | 0.022486 | 0.70799 | 0.01592 | 31.4858 | 0.037627 | 0.71032 | 0.026727 | 18.8778 | 12.9696 | 3.2124 | 0.042647 | 0.1056 | 1.2264 | 1.3075 | 1.0357
2012-10-30 03:39 | 24.6651 | 13.0948 | 14.1709 | 14.2805 | 14.0167 | 46.6034 | 0.02247 | 0.70792 | 0.015907 | 31.505 | 0.037621 | 0.7103 | 0.026722 | 18.8804 | 12.9688 | 3.2128 | 0.042629 | 0.126 | 1.2021 | 1.3117 | 1.0479
2012-10-30 03:39 | 24.6674 | 13.0646 | 14.2179 | 14.3477 | 13.9906 | 46.5994 | 0.022489 | 0.70802 | 0.015923 | 31.4829 | 0.03763 | 0.71036 | 0.026731 | 18.8775 | 12.9467 | 3.2126 | 0.042654 | 0.1179 | 1.2712 | 1.401 | 1.0439
2012-10-30 03:39 | 24.6764 | 13.0574 | 14.2025 | 14.2756 | 14.0767 | 46.5884 | 0.022474 | 0.70786 | 0.015908 | 31.4968 | 0.037624 | 0.71029 | 0.026724 | 18.8785 | 12.9476 | 3.2118 | 0.042632 | 0.1098 | 1.2549 | 1.328 | 1.1291
2012-10-30 03:39 | 24.6853 | 13.0906 | 14.1668 | 14.3047 | 14.0288 | 46.589 | 0.022477 | 0.70789 | 0.015911 | 31.4939 | 0.037621 | 0.71026 | 0.026721 | 18.8793 | 12.9931 | 3.2118 | 0.042632 | 0.0975 | 1.1737 | 1.3116 | 1.0357
2012-10-30 03:39 | 24.6928 | 13.0209 | 14.1581 | 14.3285 | 14.0485 | 46.6028 | 0.022483 | 0.70793 | 0.015916 | 31.4874 | 0.037611 | 0.71024 | 0.026713 | 18.8839 | 12.9925 | 3.2128 | 0.042629 | 0.0284 | 1.1656 | 1.336 | 1.056
2012-10-30 03:39 | 24.7025 | 13.0428 | 14.2286 | 14.3138 | 14.0136 | 46.5809 | 0.02248 | 0.70795 | 0.015915 | 31.4923 | 0.037618 | 0.71024 | 0.026718 | 18.8804 | 12.9818 | 3.2113 | 0.042633 | 0.061 | 1.2468 | 1.332 | 1.0318
2012-10-30 03:39 | 24.7033 | 13.0517 | 14.2416 | 14.27 | 13.9941 | 46.5867 | 0.02247 | 0.70783 | 0.015905 | 31.5011 | 0.037621 | 0.71027 | 0.026721 | 18.8797 | 12.9704 | 3.2117 | 0.042626 | 0.0813 | 1.2712 | 1.2996 | 1.0237
2012-10-30 03:39 | 24.7123 | 13.0933 | 14.2019 | 14.2465 | 14.0234 | 46.6075 | 0.022483 | 0.70786 | 0.015915 | 31.4842 | 0.037608 | 0.7102 | 0.026709 | 18.8843 | 12.9551 | 3.2131 | 0.042624 | 0.1382 | 1.2468 | 1.2914 | 1.0683
2012-10-30 03:39 | 24.719 | 13.0512 | 14.2167 | 14.3222 | 13.9976 | 46.5774 | 0.022452 | 0.70767 | 0.015889 | 31.5193 | 0.037614 | 0.71021 | 0.026714 | 18.8816 | 12.974 | 3.211 | 0.042603 | 0.0772 | 1.2427 | 1.3482 | 1.0236
2012-10-30 03:39 | 24.725 | 13.0734 | 14.1537 | 14.2917 | 14.0117 | 46.5855 | 0.02247 | 0.70769 | 0.015902 | 31.4947 | 0.037608 | 0.71004 | 0.026703 | 18.8801 | 13.0043 | 3.2116 | 0.042605 | 0.0691 | 1.1494 | 1.2874 | 1.0074
2012-10-30 03:39 | 24.7302 | 13.0746 | 14.1792 | 14.2644 | 13.9844 | 46.5774 | 0.022474 | 0.7077 | 0.015905 | 31.4897 | 0.037602 | 0.71007 | 0.0267 | 18.8838 | 12.9526 | 3.211 | 0.042605 | 0.122 | 1.2266 | 1.3118 | 1.0318
2012-10-30 03:39 | 24.7369 | 13.0447 | 14.1981 | 14.2914 | 13.9952 | 46.6057 | 0.022464 | 0.70779 | 0.0159 | 31.5076 | 0.037608 | 0.71013 | 0.026706 | 18.8823 | 12.9756 | 3.213 | 0.042606 | 0.0691 | 1.2225 | 1.3158 | 1.0196
2012-10-30 03:39 | 24.7399 | 13.0192 | 14.1524 | 14.2579 | 13.9779 | 46.5855 | 0.022464 | 0.70776 | 0.015899 | 31.5063 | 0.037605 | 0.71005 | 0.026702 | 18.8819 | 13.0192 | 3.2116 | 0.042601 | 0 | 1.1332 | 1.2387 | 0.9587
2012-10-30 03:39 | 24.7542 | 13.0417 | 14.1748 | 14.2682 | 13.98 | 46.5896 | 0.022467 | 0.70771 | 0.0159 | 31.5002 | 0.037599 | 0.71005 | 0.026697 | 18.8849 | 12.9766 | 3.2119 | 0.042597 | 0.0651 | 1.1982 | 1.2916 | 1.0034
2012-10-30 03:39 | 24.7654 | 13.061 | 14.1536 | 14.2794 | 13.9831 | 46.5907 | 0.022474 | 0.70763 | 0.015903 | 31.4865 | 0.037602 | 0.71003 | 0.026698 | 18.8827 | 12.9756 | 3.212 | 0.042601 | 0.0854 | 1.178 | 1.3038 | 1.0075
2012-10-30 03:39 | 24.7624 | 12.997 | 14.1709 | 14.2561 | 14.0045 | 46.578 | 0.022467 | 0.70758 | 0.015897 | 31.4944 | 0.037611 | 0.71004 | 0.026705 | 18.8785 | 12.9686 | 3.2111 | 0.042602 | 0.0284 | 1.2023 | 1.2875 | 1.0359
2012-10-30 03:39 | 24.7729 | 13.0279 | 14.1773 | 14.2666 | 13.9865 | 46.5936 | 0.022464 | 0.70763 | 0.015896 | 31.5005 | 0.037611 | 0.70998 | 0.026703 | 18.877 | 12.9709 | 3.2122 | 0.042599 | 0.057 | 1.2064 | 1.2957 | 1.0156
2012-10-30 03:39 | 24.7766 | 13.0234 | 14.1404 | 14.2824 | 13.9172 | 46.5942 | 0.022448 | 0.70763 | 0.015885 | 31.523 | 0.037599 | 0.70997 | 0.026694 | 18.8826 | 12.938 | 3.2122 | 0.042579 | 0.0854 | 1.2024 | 1.3444 | 0.9792
2012-10-30 03:39 | 24.7953 | 13.0177 | 14.1672 | 14.2525 | 14.0495 | 46.5693 | 0.022464 | 0.7076 | 0.015896 | 31.4992 | 0.037608 | 0.70985 | 0.026696 | 18.8751 | 13.0218 | 3.2105 | 0.042592 | -0.0041 | 1.1454 | 1.2307 | 1.0277
2012-10-30 03:39 | 24.802 | 13.0529 | 14.1942 | 14.2673 | 14.0238 | 46.5953 | 0.022455 | 0.70743 | 0.015885 | 31.5042 | 0.03758 | 0.70984 | 0.026676 | 18.8887 | 13.0285 | 3.2123 | 0.042561 | 0.0244 | 1.1657 | 1.2388 | 0.9953
2012-10-30 03:39 | 24.805 | 12.9705 | 14.1932 | 14.2013 | 14.0349 | 46.582 | 0.022467 | 0.70747 | 0.015895 | 31.4893 | 0.037574 | 0.70979 | 0.02667 | 18.8906 | 13.003 | 3.2114 | 0.042565 | -0.0325 | 1.1902 | 1.1983 | 1.0319
2012-10-30 03:39 | 24.8013 | 13.0278 | 14.1935 | 14.2625 | 13.9743 | 46.5913 | 0.022464 | 0.70744 | 0.015892 | 31.4922 | 0.037592 | 0.7099 | 0.026686 | 18.8842 | 13.0603 | 3.212 | 0.042578 | -0.0325 | 1.1332 | 1.2022 | 0.914
2012-10-30 03:39 | 24.8013 | 13.0563 | 14.1611 | 14.2219 | 13.9743 | 46.5855 | 0.022461 | 0.70751 | 0.015891 | 31.4996 | 0.037596 | 0.70991 | 0.02669 | 18.8826 | 13.0603 | 3.2116 | 0.042581 | -0.004 | 1.1008 | 1.1616 | 0.914
2012-10-30 03:39 | 24.808 | 13.0223 | 14.1515 | 14.253 | 13.9892 | 46.5792 | 0.022464 | 0.70745 | 0.015892 | 31.4928 | 0.037599 | 0.70993 | 0.026692 | 18.8815 | 13.0142 | 3.2112 | 0.042584 | 0.0081 | 1.1373 | 1.2388 | 0.975
2012-10-30 03:39 | 24.8245 | 13.0469 | 14.1761 | 14.2411 | 14.0097 | 46.5872 | 0.022458 | 0.70753 | 0.01589 | 31.5045 | 0.037602 | 0.70988 | 0.026693 | 18.8788 | 13.0673 | 3.2117 | 0.042583 | -0.0204 | 1.1088 | 1.1738 | 0.9424
2012-10-30 03:39 | 24.8275 | 13.0215 | 14.2157 | 14.236 | 13.9802 | 46.5971 | 0.022452 | 0.70757 | 0.015886 | 31.5148 | 0.037596 | 0.70998 | 0.026693 | 18.8845 | 13.0052 | 3.2124 | 0.042579 | 0.0163 | 1.2105 | 1.2308 | 0.975
2012-10-30 03:40 | 24.8372 | 13.0108 | 14.1361 | 14.2416 | 14.0143 | 46.5751 | 0.022467 | 0.7075 | 0.015895 | 31.4905 | 0.037605 | 0.7099 | 0.026696 | 18.8777 | 13.0108 | 3.2109 | 0.042591 | 0 | 1.1253 | 1.2308 | 1.0035
2012-10-30 03:40 | 24.8327 | 13.0551 | 14.1722 | 14.1884 | 13.9773 | 46.578 | 0.022452 | 0.70758 | 0.015887 | 31.5154 | 0.037599 | 0.7099 | 0.026691 | 18.8807 | 13.0104 | 3.2111 | 0.042578 | 0.0447 | 1.1618 | 1.178 | 0.9669
2012-10-30 03:40 | 24.8349 | 13.0248 | 14.146 | 14.2555 | 13.9633 | 46.574 | 0.022477 | 0.70754 | 0.015903 | 31.4784 | 0.037589 | 0.70993 | 0.026685 | 18.8865 | 13.0248 | 3.2108 | 0.042588 | 0 | 1.1212 | 1.2307 | 0.9385
2012-10-30 03:40 | 24.8402 | 13.0098 | 14.131 | 14.2568 | 14.0011 | 46.5705 | 0.022452 | 0.70758 | 0.015887 | 31.5154 | 0.037592 | 0.70987 | 0.026685 | 18.8835 | 13.0504 | 3.2106 | 0.042572 | -0.0406 | 1.0806 | 1.2064 | 0.9507
2012-10-30 03:40 | 24.8529 | 13.0021 | 14.1518 | 14.2451 | 13.9813 | 46.5815 | 0.022461 | 0.70751 | 0.015891 | 31.4996 | 0.037592 | 0.7099 | 0.026686 | 18.8842 | 12.9778 | 3.2113 | 0.042577 | 0.0243 | 1.174 | 1.2673 | 1.0035
2012-10-30 03:40 | 24.8529 | 12.99 | 14.1964 | 14.2573 | 13.9407 | 46.582 | 0.022448 | 0.70741 | 0.01588 | 31.5133 | 0.037605 | 0.70984 | 0.026693 | 18.8762 | 12.9981 | 3.2114 | 0.042573 | -0.0081 | 1.1983 | 1.2592 | 0.9426
2012-10-30 03:40 | 24.8581 | 12.9992 | 14.1286 | 14.1813 | 13.9743 | 46.5861 | 0.022464 | 0.70754 | 0.015894 | 31.4967 | 0.037574 | 0.70985 | 0.026672 | 18.8921 | 12.9504 | 3.2116 | 0.042566 | 0.0488 | 1.1782 | 1.2309 | 1.0239
2012-10-30 03:40 | 24.8649 | 13.006 | 14.1313 | 14.2328 | 13.9852 | 46.5792 | 0.022464 | 0.70745 | 0.015892 | 31.4928 | 0.037608 | 0.70982 | 0.026695 | 18.8743 | 12.9816 | 3.2112 | 0.042587 | 0.0244 | 1.1497 | 1.2512 | 1.0036
2012-10-30 03:40 | 24.8701 | 13.0275 | 14.169 | 14.2339 | 14.0147 | 46.578 | 0.022461 | 0.70743 | 0.015889 | 31.4957 | 0.037599 | 0.70987 | 0.02669 | 18.8799 | 12.995 | 3.2111 | 0.042579 | 0.0325 | 1.174 | 1.2389 | 1.0197
2012-10-30 03:40 | 24.8821 | 13.011 | 14.1485 | 14.2419 | 14.0024 | 46.5664 | 0.022461 | 0.70743 | 0.015889 | 31.4957 | 0.037586 | 0.70981 | 0.026679 | 18.8849 | 12.9988 | 3.2103 | 0.042568 | 0.0122 | 1.1497 | 1.2431 | 1.0036
2012-10-30 03:40 | 24.8836 | 12.9719 | 14.1541 | 14.2596 | 13.9633 | 46.5711 | 0.022461 | 0.70744 | 0.01589 | 31.4964 | 0.037592 | 0.70982 | 0.026684 | 18.8823 | 12.9922 | 3.2106 | 0.042574 | -0.0203 | 1.1619 | 1.2674 | 0.9711
2012-10-30 03:40 | 24.8806 | 13.0177 | 14.143 | 14.1835 | 14.0009 | 46.5965 | 0.022452 | 0.70737 | 0.015882 | 31.5058 | 0.037589 | 0.70977 | 0.026679 | 18.8823 | 13.0339 | 3.2124 | 0.042561 | -0.0162 | 1.1091 | 1.1496 | 0.967
2012-10-30 03:40 | 24.8888 | 12.9974 | 14.1714 | 14.212 | 13.9685 | 46.563 | 0.022461 | 0.7074 | 0.015889 | 31.4945 | 0.037577 | 0.70982 | 0.026673 | 18.8899 | 13.0502 | 3.2101 | 0.042562 | -0.0528 | 1.1212 | 1.1618 | 0.9183
2012-10-30 03:40 | 24.8955 | 12.9797 | 14.1903 | 14.2471 | 14.028 | 46.5734 | 0.022445 | 0.7074 | 0.015878 | 31.5169 | 0.037589 | 0.70979 | 0.02668 | 18.8831 | 12.996 | 3.2108 | 0.042558 | -0.0163 | 1.1943 | 1.2511 | 1.032
2012-10-30 03:40 | 24.8978 | 12.9983 | 14.1439 | 14.2251 | 13.9815 | 46.5682 | 0.022439 | 0.70741 | 0.015874 | 31.526 | 0.037599 | 0.70977 | 0.026686 | 18.8773 | 13.0389 | 3.2104 | 0.04256 | -0.0406 | 1.105 | 1.1862 | 0.9426
2012-10-30 03:40 | 24.9105 | 13.0191 | 14.1647 | 14.2053 | 13.9942 | 46.574 | 0.022452 | 0.70732 | 0.015881 | 31.5038 | 0.037583 | 0.70977 | 0.026675 | 18.8853 | 13.015 | 3.2108 | 0.042556 | 0.0041 | 1.1497 | 1.1903 | 0.9792
2012-10-30 03:40 | 24.909 | 13.0257 | 14.1064 | 14.2038 | 14.0009 | 46.5647 | 0.02247 | 0.70731 | 0.015893 | 31.478 | 0.037577 | 0.70977 | 0.026671 | 18.8883 | 12.9932 | 3.2102 | 0.042564 | 0.0325 | 1.1132 | 1.2106 | 1.0077
2012-10-30 03:40 | 24.9165 | 12.9804 | 14.1261 | 14.191 | 13.9921 | 46.574 | 0.022445 | 0.70728 | 0.015875 | 31.5117 | 0.037564 | 0.70954 | 0.026653 | 18.8887 | 13.0251 | 3.2108 | 0.042528 | -0.0447 | 1.101 | 1.1659 | 0.967
2012-10-30 03:40 | 24.9097 | 12.941 | 14.1315 | 14.1964 | 13.9447 | 46.5688 | 0.022458 | 0.7073 | 0.015884 | 31.4942 | 0.037577 | 0.70961 | 0.026665 | 18.8841 | 13.002 | 3.2105 | 0.042549 | -0.061 | 1.1295 | 1.1944 | 0.9427
2012-10-30 03:40 | 24.9232 | 12.9911 | 14.145 | 14.1774 | 13.9623 | 46.5682 | 0.022445 | 0.70711 | 0.015871 | 31.504 | 0.037577 | 0.70959 | 0.026664 | 18.8837 | 12.9952 | 3.2104 | 0.042535 | -0.0041 | 1.1498 | 1.1822 | 0.9671
2012-10-30 03:40 | 24.9329 | 13.009 | 14.1506 | 14.2237 | 14.0004 | 46.5682 | 0.022439 | 0.70727 | 0.01587 | 31.5195 | 0.03758 | 0.70959 | 0.026667 | 18.8822 | 12.9886 | 3.2104 | 0.042537 | 0.0204 | 1.162 | 1.2351 | 1.0118
2012-10-30 03:40 | 24.9299 | 12.9775 | 14.1638 | 14.1679 | 13.9893 | 46.5693 | 0.022445 | 0.70724 | 0.015874 | 31.5098 | 0.03757 | 0.70948 | 0.026655 | 18.8841 | 12.9694 | 3.2105 | 0.042529 | 0.0081 | 1.1944 | 1.1985 | 1.0199
2012-10-30 03:40 | 24.9449 | 12.9966 | 14.191 | 14.2073 | 13.984 | 46.5711 | 0.022436 | 0.70719 | 0.015867 | 31.5205 | 0.03757 | 0.70954 | 0.026657 | 18.8857 | 12.9884 | 3.2106 | 0.042524 | 0.0082 | 1.2026 | 1.2189 | 0.9956
2012-10-30 03:40 | 24.9486 | 13.0043 | 14.1866 | 14.1338 | 13.9795 | 46.5456 | 0.022442 | 0.70722 | 0.015871 | 31.5134 | 0.037586 | 0.70952 | 0.026668 | 18.8773 | 12.9759 | 3.2089 | 0.042539 | 0.0284 | 1.2107 | 1.1579 | 1.0036
2012-10-30 03:40 | 24.9509 | 12.9863 | 14.1321 | 14.1605 | 13.9372 | 46.5699 | 0.022439 | 0.70727 | 0.01587 | 31.5195 | 0.03757 | 0.70961 | 0.02666 | 18.8876 | 13.0148 | 3.2105 | 0.04253 | -0.0285 | 1.1173 | 1.1457 | 0.9224
2012-10-30 03:40 | 24.9561 | 12.9834 | 14.1373 | 14.1738 | 13.9627 | 46.5664 | 0.022439 | 0.70719 | 0.015869 | 31.5163 | 0.037574 | 0.70959 | 0.026662 | 18.8852 | 12.963 | 3.2103 | 0.042531 | 0.0204 | 1.1743 | 1.2108 | 0.9997
2012-10-30 03:40 | 24.9606 | 12.9716 | 14.1621 | 14.2148 | 13.9306 | 46.563 | 0.022455 | 0.70734 | 0.015883 | 31.5003 | 0.037564 | 0.70954 | 0.026653 | 18.8887 | 13.0326 | 3.2101 | 0.042536 | -0.061 | 1.1295 | 1.1822 | 0.898
2012-10-30 03:40 | 24.9703 | 13.0016 | 14.1149 | 14.2611 | 13.9566 | 46.5566 | 0.022455 | 0.70717 | 0.015879 | 31.4925 | 0.037567 | 0.70952 | 0.026655 | 18.8868 | 13.0016 | 3.2096 | 0.042534 | 0 | 1.1133 | 1.2595 | 0.955
2012-10-30 03:40 | 24.9733 | 12.9762 | 14.1504 | 14.2113 | 13.988 | 46.563 | 0.022442 | 0.70717 | 0.01587 | 31.5108 | 0.037564 | 0.70956 | 0.026654 | 18.8895 | 13.0128 | 3.2101 | 0.042524 | -0.0366 | 1.1376 | 1.1985 | 0.9752
2012-10-30 03:41 | 24.9666 | 12.9817 | 14.1437 | 14.1762 | 13.9123 | 46.5387 | 0.022436 | 0.70711 | 0.015865 | 31.5167 | 0.037574 | 0.70959 | 0.026662 | 18.8852 | 12.9898 | 3.2084 | 0.042527 | -0.0081 | 1.1539 | 1.1864 | 0.9225
2012-10-30 03:41 | 24.9756 | 12.9866 | 14.0999 | 14.2136 | 13.9456 | 46.5641 | 0.022442 | 0.70714 | 0.01587 | 31.5095 | 0.037564 | 0.70974 | 0.026661 | 18.8941 | 13.0395 | 3.2101 | 0.042531 | -0.0529 | 1.0604 | 1.1741 | 0.9061
2012-10-30 03:41 | 24.9928 | 13.0282 | 14.0847 | 14.2349 | 13.9019 | 46.5768 | 0.022448 | 0.7073 | 0.015877 | 31.5082 | 0.037561 | 0.70936 | 0.026644 | 18.8856 | 13.016 | 3.211 | 0.042521 | 0.0122 | 1.0687 | 1.2189 | 0.8859
2012-10-30 03:41 | 24.9905 | 12.9486 | 14.127 | 14.1838 | 13.9443 | 46.5612 | 0.02242 | 0.70696 | 0.01585 | 31.5327 | 0.037589 | 0.70965 | 0.026675 | 18.8792 | 12.9893 | 3.2099 | 0.042525 | -0.0407 | 1.1377 | 1.1945 | 0.955
2012-10-30 03:41 | 24.9958 | 12.9946 | 14.1526 | 14.177 | 13.9212 | 46.5699 | 0.022445 | 0.70738 | 0.015877 | 31.5162 | 0.037589 | 0.70993 | 0.026685 | 18.8865 | 13.0027 | 3.2105 | 0.042562 | -0.0081 | 1.1499 | 1.1743 | 0.9185
2012-10-30 03:41 | 24.9958 | 12.9783 | 14.1323 | 14.1607 | 13.913 | 46.5838 | 0.022458 | 0.70751 | 0.015889 | 31.5038 | 0.037596 | 0.70993 | 0.02669 | 18.883 | 12.9946 | 3.2115 | 0.042579 | -0.0163 | 1.1377 | 1.1661 | 0.9184
2012-10-30 03:41 | 25.0025 | 12.9606 | 14.1187 | 14.1958 | 13.9807 | 46.5566 | 0.022455 | 0.70754 | 0.015888 | 31.5093 | 0.037583 | 0.70987 | 0.026679 | 18.888 | 13.0054 | 3.2096 | 0.042567 | -0.0448 | 1.1133 | 1.1904 | 0.9753
2012-10-30 03:41 | 25.0077 | 12.9902 | 14.132 | 14.1889 | 13.9574 | 46.552 | 0.022445 | 0.70747 | 0.015879 | 31.5201 | 0.037592 | 0.70994 | 0.026688 | 18.8854 | 13.0065 | 3.2093 | 0.042567 | -0.0163 | 1.1255 | 1.1824 | 0.9509
2012-10-30 03:41 | 25.0115 | 12.9534 | 14.0952 | 14.1561 | 13.9531 | 46.563 | 0.022474 | 0.70769 | 0.015905 | 31.4891 | 0.03758 | 0.70964 | 0.026668 | 18.8833 | 13.0225 | 3.2101 | 0.042573 | -0.0691 | 1.0727 | 1.1336 | 0.9306
2012-10-30 03:41 | 25.0197 | 12.9819 | 14.1237 | 14.1806 | 13.9369 | 46.5699 | 0.02247 | 0.70774 | 0.015903 | 31.4972 | 0.037574 | 0.70962 | 0.026663 | 18.886 | 12.9941 | 3.2105 | 0.042566 | -0.0122 | 1.1296 | 1.1865 | 0.9428
2012-10-30 03:41 | 25.0212 | 12.9183 | 14.1699 | 14.1861 | 13.9466 | 46.5543 | 0.022464 | 0.7074 | 0.015891 | 31.4903 | 0.037589 | 0.7099 | 0.026684 | 18.8857 | 12.9468 | 3.2095 | 0.042575 | -0.0285 | 1.2231 | 1.2393 | 0.9998
2012-10-30 03:41 | 25.0137 | 12.9515 | 14.138 | 14.1462 | 13.9797 | 46.5601 | 0.022467 | 0.70747 | 0.015895 | 31.4893 | 0.037592 | 0.70975 | 0.026681 | 18.8804 | 12.9718 | 3.2099 | 0.042576 | -0.0203 | 1.1662 | 1.1744 | 1.0079
2012-10-30 03:41 | 25.0235 | 12.9654 | 14.1235 | 14.1235 | 13.9326 | 46.5583 | 0.022452 | 0.70738 | 0.015882 | 31.5064 | 0.037596 | 0.70982 | 0.026687 | 18.8803 | 12.9532 | 3.2097 | 0.042569 | 0.0122 | 1.1703 | 1.1703 | 0.9794
2012-10-30 03:41 | 25.0339 | 12.9351 | 14.1339 | 14.2151 | 13.9268 | 46.5549 | 0.022455 | 0.70737 | 0.015884 | 31.5016 | 0.037599 | 0.7099 | 0.026691 | 18.8807 | 12.988 | 3.2095 | 0.042575 | -0.0529 | 1.1459 | 1.2271 | 0.9388
2012-10-30 03:41 | 25.0384 | 12.9437 | 14.0937 | 14.2236 | 13.9069 | 46.5607 | 0.022455 | 0.70747 | 0.015886 | 31.5061 | 0.037589 | 0.7099 | 0.026684 | 18.8857 | 12.9396 | 3.2099 | 0.04257 | 0.0041 | 1.1541 | 1.284 | 0.9673
2012-10-30 03:41 | 25.0422 | 12.9515 | 14.1219 | 14.1787 | 13.9513 | 46.5607 | 0.022458 | 0.70728 | 0.015884 | 31.4935 | 0.037589 | 0.71 | 0.026688 | 18.8884 | 13.0003 | 3.2099 | 0.042572 | -0.0488 | 1.1216 | 1.1784 | 0.951
2012-10-30 03:41 | 25.0466 | 12.9397 | 14.1547 | 14.1791 | 13.9476 | 46.5578 | 0.022461 | 0.7074 | 0.015889 | 31.4945 | 0.037586 | 0.70974 | 0.026676 | 18.883 | 13.0007 | 3.2097 | 0.042565 | -0.061 | 1.154 | 1.1784 | 0.9469
2012-10-30 03:41 | 25.0444 | 12.9985 | 14.1078 | 14.1647 | 13.9332 | 46.5555 | 0.022442 | 0.7074 | 0.015875 | 31.5211 | 0.03758 | 0.70968 | 0.02667 | 18.8845 | 13.0025 | 3.2095 | 0.042545 | -0.004 | 1.1053 | 1.1622 | 0.9307
2012-10-30 03:41 | 25.0594 | 12.9687 | 14.0903 | 14.1756 | 13.932 | 46.5416 | 0.022458 | 0.70738 | 0.015886 | 31.498 | 0.037586 | 0.70977 | 0.026677 | 18.8838 | 12.9809 | 3.2086 | 0.042563 | -0.0122 | 1.1094 | 1.1947 | 0.9511
2012-10-30 03:41 | 25.0586 | 12.9354 | 14.0855 | 14.1505 | 13.9352 | 46.5289 | 0.022464 | 0.70738 | 0.015891 | 31.4896 | 0.03757 | 0.70984 | 0.026669 | 18.8937 | 12.9883 | 3.2077 | 0.04256 | -0.0529 | 1.0972 | 1.1622 | 0.9469
2012-10-30 03:41 | 25.0616 | 12.918 | 14.121 | 14.2062 | 13.922 | 46.5688 | 0.022458 | 0.70741 | 0.015887 | 31.4993 | 0.037586 | 0.70979 | 0.026678 | 18.8846 | 13.0035 | 3.2105 | 0.042565 | -0.0855 | 1.1175 | 1.2027 | 0.9185
2012-10-30 03:41 | 25.0758 | 12.9892 | 14.123 | 14.188 | 13.924 | 46.5578 | 0.022458 | 0.70743 | 0.015887 | 31.5 | 0.037586 | 0.70985 | 0.026681 | 18.8861 | 12.9811 | 3.2097 | 0.042568 | 0.0081 | 1.1419 | 1.2069 | 0.9429
2012-10-30 03:41 | 25.0668 | 12.9436 | 14.0937 | 14.1059 | 13.9109 | 46.5422 | 0.022467 | 0.70747 | 0.015895 | 31.4893 | 0.037599 | 0.70984 | 0.026689 | 18.8792 | 13.0046 | 3.2086 | 0.042584 | -0.061 | 1.0891 | 1.1013 | 0.9063
2012-10-30 03:41 | 25.0811 | 12.9538 | 14.0917 | 14.1851 | 13.909 | 46.5595 | 0.022445 | 0.70754 | 0.015881 | 31.5233 | 0.037583 | 0.70972 | 0.026674 | 18.8842 | 13.0392 | 3.2098 | 0.042555 | -0.0854 | 1.0525 | 1.1459 | 0.8698
2012-10-30 03:41 | 25.0856 | 12.9909 | 14.0841 | 14.1653 | 13.9338 | 46.5624 | 0.022445 | 0.70741 | 0.015878 | 31.5175 | 0.03758 | 0.70971 | 0.026671 | 18.8853 | 13.0356 | 3.21 | 0.042549 | -0.0447 | 1.0485 | 1.1297 | 0.8982
2012-10-30 03:41 | 25.0885 | 12.9531 | 14.1398 | 14.1763 | 13.9286 | 46.5445 | 0.022445 | 0.70737 | 0.015877 | 31.5156 | 0.03757 | 0.70981 | 0.026668 | 18.893 | 13.0629 | 3.2088 | 0.042545 | -0.1098 | 1.0769 | 1.1134 | 0.8657
2012-10-30 03:41 | 25.0923 | 12.9854 | 14.1395 | 14.1557 | 13.9324 | 46.5497 | 0.022464 | 0.70734 | 0.01589 | 31.4877 | 0.037583 | 0.70982 | 0.026677 | 18.8868 | 13.0179 | 3.2091 | 0.042567 | -0.0325 | 1.1216 | 1.1378 | 0.9145
2012-10-30 03:41 | 25.1005 | 12.9651 | 14.103 | 14.1721 | 13.9325 | 46.5456 | 0.022461 | 0.70737 | 0.015888 | 31.4932 | 0.03758 | 0.70975 | 0.026672 | 18.8864 | 13.0342 | 3.2089 | 0.04256 | -0.0691 | 1.0688 | 1.1379 | 0.8983
2012-10-30 03:41 | 25.1043 | 12.9526 | 14.1028 | 14.2043 | 13.9322 | 46.5491 | 0.022467 | 0.7074 | 0.015893 | 31.486 | 0.037596 | 0.70981 | 0.026686 | 18.8799 | 13.0502 | 3.2091 | 0.042579 | -0.0976 | 1.0526 | 1.1541 | 0.882
2012-10-30 03:42 | 25.1005 | 12.9569 | 14.099 | 14.1355 | 13.9528 | 46.5497 | 0.022445 | 0.70743 | 0.015878 | 31.5182 | 0.037596 | 0.70982 | 0.026687 | 18.8803 | 13.0464 | 3.2091 | 0.042565 | -0.0895 | 1.0526 | 1.0891 | 0.9064
2012-10-30 03:42 | 25.1155 | 12.9719 | 14.1261 | 14.1749 | 13.919 | 46.5277 | 0.022452 | 0.70743 | 0.015883 | 31.5084 | 0.037589 | 0.70979 | 0.02668 | 18.8831 | 13.0574 | 3.2076 | 0.042563 | -0.0855 | 1.0687 | 1.1175 | 0.8616
2012-10-30 03:42 | 25.1155 | 12.9353 | 14.1586 | 14.1708 | 13.915 | 46.5508 | 0.022464 | 0.70744 | 0.015892 | 31.4922 | 0.03758 | 0.7099 | 0.026678 | 18.8903 | 13.0492 | 3.2092 | 0.04257 | -0.1139 | 1.1094 | 1.1216 | 0.8658
2012-10-30 03:42 | 25.1319 | 12.9355 | 14.106 | 14.1385 | 13.9314 | 46.5433 | 0.022455 | 0.70731 | 0.015883 | 31.499 | 0.037577 | 0.70968 | 0.026668 | 18.886 | 13.0005 | 3.2087 | 0.042551 | -0.065 | 1.1055 | 1.138 | 0.9309
2012-10-30 03:42 | 25.1491 | 12.9161 | 14.0988 | 14.1435 | 13.8958 | 46.5456 | 0.022439 | 0.70738 | 0.015873 | 31.5247 | 0.037586 | 0.70974 | 0.026676 | 18.883 | 13.0096 | 3.2089 | 0.042549 | -0.0935 | 1.0892 | 1.1339 | 0.8862
2012-10-30 03:42 | 25.1536 | 12.9409 | 14.1155 | 14.1318 | 13.9003 | 46.5485 | 0.022445 | 0.70731 | 0.015876 | 31.513 | 0.03758 | 0.70979 | 0.026674 | 18.8876 | 13.0385 | 3.2091 | 0.04255 | -0.0976 | 1.077 | 1.0933 | 0.8618
2012-10-30 03:42 | 25.1529 | 12.9076 | 14.1229 | 14.1351 | 13.9239 | 46.5271 | 0.022455 | 0.70734 | 0.015883 | 31.5003 | 0.037599 | 0.70974 | 0.026685 | 18.8765 | 13.0215 | 3.2076 | 0.042568 | -0.1139 | 1.1014 | 1.1136 | 0.9024
2012-10-30 03:42 | 25.1604 | 12.9233 | 14.0614 | 14.1548 | 13.8908 | 46.5468 | 0.022442 | 0.70744 | 0.015876 | 31.523 | 0.037583 | 0.70979 | 0.026676 | 18.8861 | 12.9843 | 3.2089 | 0.042552 | -0.061 | 1.0771 | 1.1705 | 0.9065
2012-10-30 03:42 | 25.1604 | 12.9111 | 14.0614 | 14.1142 | 13.8867 | 46.5537 | 0.022445 | 0.70735 | 0.015877 | 31.5149 | 0.037586 | 0.70968 | 0.026674 | 18.8815 | 13.025 | 3.2094 | 0.042551 | -0.1139 | 1.0364 | 1.0892 | 0.8617
2012-10-30 03:42 | 25.1686 | 12.8827 | 14.0777 | 14.1224 | 13.9153 | 46.5474 | 0.022448 | 0.7073 | 0.015877 | 31.5082 | 0.037583 | 0.70961 | 0.026669 | 18.8811 | 13.0332 | 3.209 | 0.042546 | -0.1505 | 1.0445 | 1.0892 | 0.8821
2012-10-30 03:42 | 25.1746 | 12.9049 | 14.0634 | 14.0959 | 13.8888 | 46.5358 | 0.022455 | 0.70722 | 0.015881 | 31.4951 | 0.037592 | 0.70964 | 0.026677 | 18.8773 | 13.0473 | 3.2082 | 0.042558 | -0.1424 | 1.0161 | 1.0486 | 0.8415
2012-10-30 03:42 | 25.1738 | 12.9041 | 14.022 | 14.1682 | 13.8676 | 46.5601 | 0.022442 | 0.70719 | 0.015871 | 31.5121 | 0.037567 | 0.70958 | 0.026657 | 18.8883 | 13.0384 | 3.2099 | 0.042528 | -0.1343 | 0.9836 | 1.1298 | 0.8292
2012-10-30 03:42 | 25.1836 | 12.861 | 14.0481 | 14.1333 | 13.8774 | 46.5491 | 0.022445 | 0.70721 | 0.015873 | 31.5085 | 0.037577 | 0.70961 | 0.026665 | 18.8841 | 13.0563 | 3.2091 | 0.042538 | -0.1953 | 0.9918 | 1.077 | 0.8211
2012-10-30 03:42 | 25.1851 | 12.9073 | 14.078 | 14.1105 | 13.8952 | 46.5485 | 0.022452 | 0.70722 | 0.015879 | 31.4993 | 0.037577 | 0.70969 | 0.026668 | 18.8864 | 12.9887 | 3.2091 | 0.042547 | -0.0814 | 1.0893 | 1.1218 | 0.9065
2012-10-30 03:42 | 25.1963 | 12.9104 | 14.0729 | 14.0811 | 13.9267 | 46.5347 | 0.022442 | 0.70719 | 0.015871 | 31.5121 | 0.037577 | 0.70961 | 0.026665 | 18.8841 | 13.0405 | 3.2081 | 0.042536 | -0.1301 | 1.0324 | 1.0406 | 0.8862
2012-10-30 03:42 | 25.1993 | 12.9174 | 14.0881 | 14.1165 | 13.8931 | 46.5404 | 0.022455 | 0.7073 | 0.015882 | 31.4984 | 0.037599 | 0.70979 | 0.026688 | 18.878 | 13.0191 | 3.2085 | 0.04257 | -0.1017 | 1.069 | 1.0974 | 0.874
2012-10-30 03:42 | 25.209 | 12.9353 | 14.0978 | 14.0938 | 13.9028 | 46.5485 | 0.022461 | 0.7073 | 0.015887 | 31.49 | 0.037589 | 0.70968 | 0.026676 | 18.88 | 13.037 | 3.2091 | 0.042563 | -0.1017 | 1.0608 | 1.0568 | 0.8658
2012-10-30 03:42 | 25.2165 | 12.9428 | 14.1337 | 14.15 | 13.9063 | 46.526 | 0.022442 | 0.70724 | 0.015872 | 31.514 | 0.037589 | 0.70968 | 0.026676 | 18.88 | 13.0201 | 3.2075 | 0.042548 | -0.0773 | 1.1136 | 1.1299 | 0.8862
2012-10-30 03:42 | 25.2165 | 12.8939 | 14.0769 | 14.1459 | 13.8657 | 46.5202 | 0.022455 | 0.70722 | 0.015881 | 31.4951 | 0.037599 | 0.70979 | 0.026688 | 18.878 | 13.0282 | 3.2071 | 0.042569 | -0.1343 | 1.0487 | 1.1177 | 0.8375
2012-10-30 03:42 | 25.2255 | 12.9152 | 14.0778 | 14.159 | 13.8584 | 46.5404 | 0.022455 | 0.70711 | 0.015878 | 31.49 | 0.037574 | 0.70956 | 0.026661 | 18.8844 | 13.0291 | 3.2085 | 0.042539 | -0.1139 | 1.0487 | 1.1299 | 0.8293
2012-10-30 03:42 | 25.2202 | 12.9139 | 14.0765 | 14.0806 | 13.8937 | 46.5202 | 0.022458 | 0.70719 | 0.015882 | 31.4896 | 0.037589 | 0.70952 | 0.02667 | 18.8758 | 13.0441 | 3.2071 | 0.042552 | -0.1302 | 1.0324 | 1.0365 | 0.8496
2012-10-30 03:42 | 25.2284 | 12.8733 | 14.105 | 14.101 | 13.8532 | 46.5399 | 0.022445 | 0.70717 | 0.015872 | 31.5066 | 0.037574 | 0.70959 | 0.026662 | 18.8852 | 13.0279 | 3.2085 | 0.042534 | -0.1546 | 1.0771 | 1.0731 | 0.8253
2012-10-30 03:43 | 25.2299 | 12.8829 | 14.0659 | 14.0984 | 13.8506 | 46.541 | 0.022445 | 0.70708 | 0.01587 | 31.5027 | 0.037577 | 0.70951 | 0.026661 | 18.8814 | 13.0335 | 3.2085 | 0.042531 | -0.1506 | 1.0324 | 1.0649 | 0.8171
2012-10-30 03:43 | 25.2374 | 12.866 | 14.045 | 14.114 | 13.8825 | 46.5323 | 0.022442 | 0.70709 | 0.015869 | 31.5076 | 0.03758 | 0.70956 | 0.026665 | 18.8814 | 13.0084 | 3.2079 | 0.042534 | -0.1424 | 1.0366 | 1.1056 | 0.8741
2012-10-30 03:43 | 25.2442 | 12.9176 | 14.0558 | 14.1087 | 13.9259 | 46.5243 | 0.022448 | 0.70718 | 0.015875 | 31.503 | 0.037567 | 0.70962 | 0.026658 | 18.8895 | 13.0803 | 3.2074 | 0.042533 | -0.1627 | 0.9755 | 1.0284 | 0.8456
2012-10-30 03:43 | 25.2434 | 12.8435 | 14.1119 | 14.12 | 13.8357 | 46.5451 | 0.022448 | 0.70717 | 0.015874 | 31.5024 | 0.037586 | 0.70961 | 0.026671 | 18.8796 | 13.0348 | 3.2088 | 0.042545 | -0.1913 | 1.0771 | 1.0852 | 0.8009
2012-10-30 03:43 | 25.2449 | 12.8654 | 14.0687 | 14.0972 | 13.8656 | 46.5323 | 0.022439 | 0.70715 | 0.015868 | 31.5144 | 0.037583 | 0.70952 | 0.026666 | 18.8788 | 13.02 | 3.2079 | 0.042534 | -0.1546 | 1.0487 | 1.0772 | 0.8456
2012-10-30 03:43 | 25.2472 | 12.8921 | 14.0588 | 14.0913 | 13.8395 | 46.5271 | 0.022436 | 0.70712 | 0.015865 | 31.5173 | 0.037574 | 0.70949 | 0.026658 | 18.8825 | 12.9979 | 3.2076 | 0.042523 | -0.1058 | 1.0609 | 1.0934 | 0.8416
2012-10-30 03:43 | 25.2494 | 12.8902 | 14.0935 | 14.0895 | 13.8864 | 46.5323 | 0.022436 | 0.70717 | 0.015866 | 31.5192 | 0.037567 | 0.70955 | 0.026656 | 18.8876 | 13.0286 | 3.2079 | 0.042522 | -0.1384 | 1.0649 | 1.0609 | 0.8578
2012-10-30 03:43 | 25.2554 | 12.8922 | 14.067 | 14.1077 | 13.9046 | 46.5323 | 0.022442 | 0.70714 | 0.01587 | 31.5095 | 0.03757 | 0.70955 | 0.026658 | 18.8861 | 13.0508 | 3.2079 | 0.042528 | -0.1586 | 1.0162 | 1.0569 | 0.8538
2012-10-30 03:43 | 25.2621 | 12.9111 | 14.1144 | 14.1022 | 13.8503 | 46.5243 | 0.022442 | 0.70714 | 0.01587 | 31.5095 | 0.037589 | 0.70956 | 0.026672 | 18.8769 | 13.0575 | 3.2074 | 0.042542 | -0.1464 | 1.0569 | 1.0447 | 0.7928
2012-10-30 03:43 | 25.2696 | 12.8819 | 14.0609 | 14.0691 | 13.8619 | 46.5433 | 0.022461 | 0.70721 | 0.015885 | 31.4861 | 0.037583 | 0.70954 | 0.026666 | 18.8791 | 13.0366 | 3.2087 | 0.042551 | -0.1547 | 1.0243 | 1.0325 | 0.8253
2012-10-30 03:43 | 25.2703 | 12.8745 | 14.0738 | 14.1023 | 13.8951 | 46.5306 | 0.022442 | 0.70715 | 0.01587 | 31.5102 | 0.037574 | 0.70959 | 0.026662 | 18.8852 | 13.025 | 3.2078 | 0.042532 | -0.1505 | 1.0488 | 1.0773 | 0.8701
2012-10-30 03:43 | 25.2748 | 12.8912 | 14.0011 | 14.058 | 13.8915 | 46.53 | 0.022442 | 0.70721 | 0.015871 | 31.5127 | 0.03757 | 0.70962 | 0.02666 | 18.888 | 13.0011 | 3.2078 | 0.042531 | -0.1099 | 1 | 1.0569 | 0.8904
2012-10-30 03:43 | 25.2771 | 12.865 | 14.04 | 14.1131 | 13.8409 | 46.53 | 0.022439 | 0.70718 | 0.015868 | 31.5157 | 0.037567 | 0.70955 | 0.026656 | 18.8876 | 13.0034 | 3.2078 | 0.042524 | -0.1384 | 1.0366 | 1.1097 | 0.8375
2012-10-30 03:43 | 25.289 | 12.8403 | 14.0803 | 14.0925 | 13.8691 | 46.5167 | 0.022458 | 0.70717 | 0.015882 | 31.4883 | 0.037561 | 0.70958 | 0.026652 | 18.8913 | 12.9583 | 3.2069 | 0.042534 | -0.118 | 1.122 | 1.1342 | 0.9108
2012-10-30 03:43 | 25.2898 | 12.9103 | 14.0446 | 14.0893 | 13.8536 | 46.5399 | 0.022436 | 0.7073 | 0.015869 | 31.525 | 0.03758 | 0.70952 | 0.026664 | 18.8803 | 12.9876 | 3.2085 | 0.042533 | -0.0773 | 1.057 | 1.1017 | 0.866
2012-10-30 03:43 | 25.2988 | 12.9274 | 14.0455 | 14.082 | 13.8464 | 46.5295 | 0.022445 | 0.70724 | 0.015874 | 31.5098 | 0.037574 | 0.70954 | 0.02666 | 18.8837 | 12.9803 | 3.2077 | 0.042534 | -0.0529 | 1.0652 | 1.1017 | 0.8661
2012-10-30 03:43 | 25.2883 | 12.9006 | 14.0512 | 14.0918 | 13.905 | 46.5243 | 0.022448 | 0.70718 | 0.015875 | 31.503 | 0.037583 | 0.70962 | 0.02667 | 18.8815 | 13.0227 | 3.2074 | 0.042545 | -0.1221 | 1.0285 | 1.0691 | 0.8823
2012-10-30 03:43 | 25.2988 | 12.8786 | 14.0414 | 14.1064 | 13.887 | 46.5306 | 0.022445 | 0.70719 | 0.015873 | 31.5079 | 0.03758 | 0.70965 | 0.026669 | 18.8837 | 12.9925 | 3.2078 | 0.042542 | -0.1139 | 1.0489 | 1.1139 | 0.8945
2012-10-30 03:43 | 25.295 | 12.8544 | 14.0498 | 14.0538 | 13.8548 | 46.5364 | 0.022455 | 0.70724 | 0.015881 | 31.4958 | 0.037574 | 0.70961 | 0.026663 | 18.8856 | 12.9765 | 3.2082 | 0.042544 | -0.1221 | 1.0733 | 1.0773 | 0.8783
2012-10-30 03:43 | 25.3048 | 12.8724 | 14.0352 | 14.084 | 13.8605 | 46.5248 | 0.022445 | 0.70719 | 0.015873 | 31.5079 | 0.03758 | 0.70964 | 0.026668 | 18.8833 | 12.9863 | 3.2074 | 0.042541 | -0.1139 | 1.0489 | 1.0977 | 0.8742
2012-10-30 03:43 | 25.3078 | 12.8957 | 14.0504 | 14.087 | 13.831 | 46.5393 | 0.022461 | 0.70727 | 0.015886 | 31.4887 | 0.037586 | 0.70964 | 0.026672 | 18.8803 | 13.0178 | 3.2084 | 0.042558 | -0.1221 | 1.0326 | 1.0692 | 0.8132
2012-10-30 03:43 | 25.3122 | 12.8554 | 14.0304 | 14.0467 | 13.8801 | 46.5237 | 0.022445 | 0.70721 | 0.015873 | 31.5085 | 0.037583 | 0.70968 | 0.026672 | 18.883 | 12.9978 | 3.2073 | 0.042545 | -0.1424 | 1.0326 | 1.0489 | 0.8823
2012-10-30 03:43 | 25.325 | 12.8967 | 14.0067 | 14.0757 | 13.8279 | 46.5283 | 0.022448 | 0.70724 | 0.015876 | 31.5056 | 0.037558 | 0.70959 | 0.026651 | 18.8933 | 13.0147 | 3.2077 | 0.042527 | -0.118 | 0.992 | 1.061 | 0.8132
2012-10-30 03:43 | 25.3152 | 12.8746 | 14.0497 | 14.1065 | 13.8262 | 46.5381 | 0.022445 | 0.70724 | 0.015874 | 31.5098 | 0.037574 | 0.70955 | 0.026661 | 18.884 | 13.0211 | 3.2083 | 0.042535 | -0.1465 | 1.0286 | 1.0854 | 0.8051
2012-10-30 03:43 | 25.3257 | 12.8729 | 14.0439 | 14.0927 | 13.8448 | 46.5231 | 0.022445 | 0.70712 | 0.015871 | 31.5047 | 0.037564 | 0.70961 | 0.026656 | 18.8906 | 12.9706 | 3.2073 | 0.042527 | -0.0977 | 1.0733 | 1.1221 | 0.8742
2012-10-30 03:43 | 25.3212 | 12.8847 | 14.0557 | 14.0922 | 13.8038 | 46.5381 | 0.022445 | 0.70722 | 0.015874 | 31.5092 | 0.03758 | 0.70955 | 0.026665 | 18.881 | 13.0068 | 3.2083 | 0.042539 | -0.1221 | 1.0489 | 1.0854 | 0.797
2012-10-30 03:43 | 25.3377 | 12.8361 | 14.0519 | 14.1169 | 13.8528 | 46.5306 | 0.022445 | 0.70719 | 0.015873 | 31.5079 | 0.037574 | 0.70962 | 0.026663 | 18.886 | 12.9948 | 3.2078 | 0.042536 | -0.1587 | 1.0571 | 1.1221 | 0.858
2012-10-30 03:43 | 25.3392 | 12.8661 | 14.0005 | 14.0209 | 13.838 | 46.5104 | 0.022436 | 0.70717 | 0.015866 | 31.5192 | 0.037574 | 0.70955 | 0.026661 | 18.884 | 13.0248 | 3.2064 | 0.042527 | -0.1587 | 0.9757 | 0.9961 | 0.8132
2012-10-30 03:44 | 25.3392 | 12.8742 | 14.0452 | 14.0452 | 13.8705 | 46.5271 | 0.022448 | 0.70721 | 0.015875 | 31.5043 | 0.037558 | 0.70946 | 0.026646 | 18.8898 | 13.0329 | 3.2076 | 0.042521 | -0.1587 | 1.0123 | 1.0123 | 0.8376
2012-10-30 03:44 | 25.3482 | 12.8792 | 14.0095 | 14.0745 | 13.8145 | 46.5243 | 0.022445 | 0.70706 | 0.01587 | 31.5021 | 0.03758 | 0.70952 | 0.026664 | 18.8803 | 13.0094 | 3.2074 | 0.042534 | -0.1302 | 1.0001 | 1.0651 | 0.8051
2012-10-30 03:44 | 25.3504 | 12.8488 | 13.9874 | 14.093 | 13.833 | 46.5254 | 0.022452 | 0.70727 | 0.01588 | 31.5013 | 0.037577 | 0.70956 | 0.026663 | 18.8829 | 12.9872 | 3.2075 | 0.042543 | -0.1384 | 1.0002 | 1.1058 | 0.8458
2012-10-30 03:44 | 25.3601 | 12.887 | 14.0418 | 14.1027 | 13.867 | 46.5248 | 0.022452 | 0.70715 | 0.015877 | 31.4961 | 0.03757 | 0.70955 | 0.026658 | 18.8861 | 12.9969 | 3.2074 | 0.042535 | -0.1099 | 1.0449 | 1.1058 | 0.8701
2012-10-30 03:44 | 25.3609 | 12.8796 | 14.0182 | 14.0547 | 13.8191 | 46.5179 | 0.022445 | 0.70718 | 0.015873 | 31.5072 | 0.037567 | 0.70949 | 0.026653 | 18.886 | 13.0302 | 3.2069 | 0.042526 | -0.1506 | 0.988 | 1.0245 | 0.7889
2012-10-30 03:44 | 25.3594 | 12.8252 | 14.0492 | 14.0329 | 13.8135 | 46.5098 | 0.022455 | 0.70718 | 0.01588 | 31.4932 | 0.037564 | 0.70958 | 0.026655 | 18.8898 | 12.9921 | 3.2064 | 0.042535 | -0.1669 | 1.0571 | 1.0408 | 0.8214
2012-10-30 03:44 | 25.3616 | 12.86 | 14.0392 | 14.0839 | 13.8523 | 46.5358 | 0.022455 | 0.70719 | 0.01588 | 31.4938 | 0.037577 | 0.70959 | 0.026664 | 18.8837 | 13.0431 | 3.2082 | 0.042544 | -0.1831 | 0.9961 | 1.0408 | 0.8092
2012-10-30 03:44 | 25.3721 | 12.8583 | 14.0131 | 14.0497 | 13.8587 | 46.5104 | 0.022442 | 0.70719 | 0.015871 | 31.5121 | 0.037577 | 0.70964 | 0.026666 | 18.8849 | 13.0129 | 3.2064 | 0.042537 | -0.1546 | 1.0002 | 1.0368 | 0.8458
2012-10-30 03:44 | 25.3818 | 12.8639 | 14.0147 | 14.0594 | 13.844 | 46.5214 | 0.022442 | 0.70725 | 0.015872 | 31.5147 | 0.037586 | 0.70965 | 0.026673 | 18.8807 | 13.0023 | 3.2072 | 0.042545 | -0.1384 | 1.0124 | 1.0571 | 0.8417
2012-10-30 03:44 | 25.3826 | 12.8688 | 14.0277 | 14.048 | 13.8692 | 46.5225 | 0.022455 | 0.70719 | 0.01588 | 31.4938 | 0.037592 | 0.70965 | 0.026677 | 18.8777 | 13.0194 | 3.2073 | 0.042557 | -0.1506 | 1.0083 | 1.0286 | 0.8498
2012-10-30 03:44 | 25.3841 | 12.8499 | 14.0211 | 14.0373 | 13.8463 | 46.5092 | 0.022452 | 0.70725 | 0.015879 | 31.5006 | 0.037586 | 0.70969 | 0.026675 | 18.8819 | 13.0249 | 3.2063 | 0.042554 | -0.175 | 0.9962 | 1.0124 | 0.8214
2012-10-30 03:44 | 25.3826 | 12.8322 | 14.048 | 14.0399 | 13.8042 | 46.5341 | 0.022445 | 0.70725 | 0.015874 | 31.5104 | 0.03757 | 0.70964 | 0.026661 | 18.8884 | 13.0031 | 3.2081 | 0.042535 | -0.1709 | 1.0449 | 1.0368 | 0.8011
2012-10-30 03:44 | 25.3908 | 12.8485 | 14.0196 | 14.0521 | 13.8368 | 46.5115 | 0.022445 | 0.70722 | 0.015874 | 31.5092 | 0.037577 | 0.70954 | 0.026662 | 18.8822 | 13.0357 | 3.2065 | 0.042536 | -0.1872 | 0.9839 | 1.0164 | 0.8011
2012-10-30 03:44 | 25.3923 | 12.8419 | 13.9846 | 14.074 | 13.8464 | 46.5225 | 0.022445 | 0.70709 | 0.015871 | 31.5034 | 0.037567 | 0.70952 | 0.026655 | 18.8868 | 13.0006 | 3.2073 | 0.042526 | -0.1587 | 0.984 | 1.0734 | 0.8458
2012-10-30 03:44 | 25.399 | 12.8689 | 13.9994 | 14.0563 | 13.8206 | 46.5243 | 0.022439 | 0.70717 | 0.015868 | 31.515 | 0.037564 | 0.70955 | 0.026653 | 18.8891 | 13.0073 | 3.2074 | 0.042521 | -0.1384 | 0.9921 | 1.049 | 0.8133
2012-10-30 03:44 | 25.4118 | 12.841 | 14.0285 | 14.065 | 13.809 | 46.4988 | 0.022436 | 0.70717 | 0.015866 | 31.5192 | 0.037577 | 0.70952 | 0.026662 | 18.8818 | 13.0201 | 3.2056 | 0.042528 | -0.1791 | 1.0084 | 1.0449 | 0.7889
2012-10-30 03:44 | 25.4028 | 12.8442 | 14.0316 | 13.9788 | 13.8447 | 46.515 | 0.022439 | 0.70712 | 0.015867 | 31.5131 | 0.037574 | 0.70969 | 0.026666 | 18.8879 | 13.0314 | 3.2067 | 0.042533 | -0.1872 | 1.0002 | 0.9474 | 0.8133
2012-10-30 03:44 | 25.405 | 12.8261 | 13.9892 | 14.0704 | 13.8388 | 46.5133 | 0.022464 | 0.70705 | 0.015883 | 31.4748 | 0.03757 | 0.7094 | 0.026652 | 18.8822 | 12.9807 | 3.2066 | 0.042535 | -0.1546 | 1.0085 | 1.0897 | 0.8581
2012-10-30 03:44 | 25.417 | 12.8503 | 14.0052 | 14.0621 | 13.8589 | 46.5219 | 0.022436 | 0.70715 | 0.015866 | 31.5186 | 0.037558 | 0.70952 | 0.026648 | 18.8913 | 13.0212 | 3.2072 | 0.042514 | -0.1709 | 0.984 | 1.0409 | 0.8377
2012-10-30 03:44 | 25.417 | 12.8259 | 14.0337 | 14.0296 | 13.8061 | 46.5271 | 0.022461 | 0.70724 | 0.015885 | 31.4874 | 0.037577 | 0.70956 | 0.026663 | 18.8829 | 12.9846 | 3.2076 | 0.042548 | -0.1587 | 1.0491 | 1.045 | 0.8215
2012-10-30 03:44 | 25.4207 | 12.858 | 14.0292 | 14.0861 | 13.822 | 46.5162 | 0.022445 | 0.70714 | 0.015872 | 31.5053 | 0.03758 | 0.70956 | 0.026665 | 18.8814 | 13.0127 | 3.2068 | 0.042537 | -0.1547 | 1.0165 | 1.0734 | 0.8093
2012-10-30 03:44 | 25.4215 | 12.8385 | 14.03 | 14.0544 | 13.7943 | 46.5086 | 0.022445 | 0.70706 | 0.01587 | 31.5021 | 0.037567 | 0.70952 | 0.026655 | 18.8868 | 12.9891 | 3.2063 | 0.042525 | -0.1506 | 1.0409 | 1.0653 | 0.8052
2012-10-30 03:44 | 25.4349 | 12.856 | 14.0556 | 14.0516 | 13.8402 | 46.5185 | 0.022442 | 0.70708 | 0.015868 | 31.507 | 0.037577 | 0.70951 | 0.026661 | 18.8814 | 13.0391 | 3.207 | 0.042529 | -0.1831 | 1.0165 | 1.0125 | 0.8011
2012-10-30 03:44 | 25.4305 | 12.8109 | 13.9903 | 13.9984 | 13.7668 | 46.5283 | 0.022436 | 0.70709 | 0.015864 | 31.516 | 0.037555 | 0.70943 | 0.026643 | 18.8905 | 13.0591 | 3.2077 | 0.042507 | -0.2482 | 0.9312 | 0.9393 | 0.7077
2012-10-30 03:44 | 25.4342 | 12.8186 | 13.9858 | 14.0468 | 13.8355 | 46.5046 | 0.022442 | 0.70705 | 0.015868 | 31.5057 | 0.037567 | 0.70945 | 0.026652 | 18.8849 | 12.9977 | 3.206 | 0.04252 | -0.1791 | 0.9881 | 1.0491 | 0.8378
2012-10-30 03:44 | 25.4379 | 12.8508 | 14.0139 | 14.0383 | 13.7985 | 46.5017 | 0.022426 | 0.70709 | 0.015857 | 31.5301 | 0.03757 | 0.70943 | 0.026653 | 18.883 | 13.0014 | 3.2058 | 0.04251 | -0.1506 | 1.0125 | 1.0369 | 0.7971
2012-10-30 03:44 | 25.4462 | 12.794 | 13.9491 | 14.0222 | 13.815 | 46.5196 | 0.022436 | 0.70698 | 0.015862 | 31.5109 | 0.03758 | 0.70945 | 0.026661 | 18.8783 | 13.0056 | 3.2071 | 0.042523 | -0.2116 | 0.9435 | 1.0166 | 0.8094
2012-10-30 03:45 | 25.4469 | 12.8191 | 14.0026 | 14.0595 | 13.8238 | 46.5173 | 0.022433 | 0.70705 | 0.015861 | 31.5183 | 0.037555 | 0.70939 | 0.026641 | 18.8894 | 13.0145 | 3.2069 | 0.042502 | -0.1954 | 0.9881 | 1.045 | 0.8093
2012-10-30 03:45 | 25.4544 | 12.8307 | 14.0426 | 14.0223 | 13.7703 | 46.5029 | 0.022436 | 0.70705 | 0.015863 | 31.5141 | 0.037567 | 0.70932 | 0.026647 | 18.8814 | 13.0016 | 3.2059 | 0.04251 | -0.1709 | 1.041 | 1.0207 | 0.7687
2012-10-30 03:45 | 25.4596 | 12.84 | 13.995 | 14.0194 | 13.8162 | 46.5029 | 0.022439 | 0.70689 | 0.015862 | 31.5028 | 0.03757 | 0.70932 | 0.026649 | 18.8799 | 13.0028 | 3.2059 | 0.042511 | -0.1628 | 0.9922 | 1.0166 | 0.8134
2012-10-30 03:45 | 25.4566 | 12.8329 | 14.0245 | 14.0408 | 13.7969 | 46.511 | 0.022439 | 0.70708 | 0.015866 | 31.5112 | 0.037564 | 0.70949 | 0.026651 | 18.8876 | 12.9794 | 3.2065 | 0.042517 | -0.1465 | 1.0451 | 1.0614 | 0.8175
2012-10-30 03:45 | 25.4671 | 12.8149 | 13.9862 | 14.0472 | 13.8155 | 46.504 | 0.022448 | 0.70706 | 0.015872 | 31.4979 | 0.037558 | 0.70946 | 0.026646 | 18.8898 | 13.0103 | 3.206 | 0.042518 | -0.1954 | 0.9759 | 1.0369 | 0.8052
2012-10-30 03:45 | 25.4679 | 12.8238 | 14.0155 | 14.0317 | 13.8001 | 46.4982 | 0.022442 | 0.70701 | 0.015867 | 31.5037 | 0.037574 | 0.70942 | 0.026656 | 18.8806 | 13.0029 | 3.2056 | 0.042523 | -0.1791 | 1.0126 | 1.0288 | 0.7972
2012-10-30 03:45 | 25.4596 | 12.783 | 13.9909 | 13.995 | 13.7511 | 46.5104 | 0.022439 | 0.70712 | 0.015867 | 31.5131 | 0.037558 | 0.70936 | 0.026642 | 18.8871 | 12.9783 | 3.2064 | 0.042509 | -0.1953 | 1.0126 | 1.0167 | 0.7728
2012-10-30 03:45 | 25.4731 | 12.8494 | 14.0004 | 14.0288 | 13.7809 | 46.5133 | 0.022433 | 0.70692 | 0.015858 | 31.5125 | 0.037548 | 0.70927 | 0.026632 | 18.8898 | 13.0163 | 3.2066 | 0.04249 | -0.1669 | 0.9841 | 1.0125 | 0.7646
2012-10-30 03:45 | 25.4813 | 12.821 | 13.9517 | 13.9842 | 13.781 | 46.4994 | 0.022433 | 0.70693 | 0.015859 | 31.5131 | 0.037542 | 0.70927 | 0.026628 | 18.8928 | 13.0326 | 3.2057 | 0.042487 | -0.2116 | 0.9191 | 0.9516 | 0.7484
2012-10-30 03:45 | 25.4888 | 12.8 | 13.9795 | 14.0364 | 13.7966 | 46.5046 | 0.022445 | 0.70688 | 0.015866 | 31.4937 | 0.037561 | 0.70935 | 0.026644 | 18.8852 | 13.0116 | 3.206 | 0.04251 | -0.2116 | 0.9679 | 1.0248 | 0.785
2012-10-30 03:45 | 25.4776 | 12.7725 | 13.9805 | 14.0252 | 13.7651 | 46.4977 | 0.022439 | 0.70693 | 0.015863 | 31.5047 | 0.037555 | 0.70922 | 0.026635 | 18.8848 | 13.0045 | 3.2056 | 0.042498 | -0.232 | 0.976 | 1.0207 | 0.7606
2012-10-30 03:45 | 25.4926 | 12.8078 | 13.963 | 14.0483 | 13.7516 | 46.5121 | 0.022423 | 0.70683 | 0.015849 | 31.5227 | 0.037542 | 0.70932 | 0.026629 | 18.894 | 13.0439 | 3.2065 | 0.042478 | -0.2361 | 0.9191 | 1.0044 | 0.7077
2012-10-30 03:45 | 25.4896 | 12.7763 | 13.96 | 13.9884 | 13.8015 | 46.5086 | 0.022439 | 0.70696 | 0.015864 | 31.506 | 0.037564 | 0.70935 | 0.026646 | 18.8837 | 12.9961 | 3.2063 | 0.04251 | -0.2198 | 0.9639 | 0.9923 | 0.8054
2012-10-30 03:45 | 25.4963 | 12.8115 | 13.9911 | 14.0439 | 13.7716 | 46.4977 | 0.022426 | 0.70672 | 0.015849 | 31.5133 | 0.037533 | 0.70901 | 0.026611 | 18.8904 | 13.0028 | 3.2056 | 0.04246 | -0.1913 | 0.9883 | 1.0411 | 0.7688
2012-10-30 03:45 | 25.509 | 12.8161 | 13.9712 | 14.0119 | 13.7924 | 46.493 | 0.022442 | 0.70711 | 0.015869 | 31.5082 | 0.037558 | 0.7094 | 0.026644 | 18.8883 | 13.0115 | 3.2052 | 0.042513 | -0.1954 | 0.9597 | 1.0004 | 0.7809
2012-10-30 03:45 | 25.5068 | 12.7895 | 13.9894 | 14.0178 | 13.7902 | 46.5023 | 0.022445 | 0.70705 | 0.01587 | 31.5014 | 0.037561 | 0.70942 | 0.026646 | 18.8871 | 13.0581 | 3.2059 | 0.042516 | -0.2686 | 0.9313 | 0.9597 | 0.7321
2012-10-30 03:45 | 25.512 | 12.7947 | 14.0068 | 14.023 | 13.7791 | 46.4855 | 0.022445 | 0.70685 | 0.015865 | 31.4924 | 0.037551 | 0.70927 | 0.026634 | 18.8883 | 13.0185 | 3.2047 | 0.042499 | -0.2238 | 0.9883 | 1.0045 | 0.7606
2012-10-30 03:45 | 25.5195 | 12.7777 | 13.9492 | 13.9817 | 13.7582 | 46.4982 | 0.022423 | 0.70667 | 0.015846 | 31.5156 | 0.037542 | 0.70916 | 0.026623 | 18.8897 | 13.0545 | 3.2056 | 0.042469 | -0.2768 | 0.8947 | 0.9272 | 0.7037
2012-10-30 03:45 | 25.5217 | 12.8003 | 13.9839 | 13.988 | 13.7888 | 46.4907 | 0.02242 | 0.70679 | 0.015846 | 31.525 | 0.037548 | 0.7091 | 0.026625 | 18.8852 | 13.016 | 3.2051 | 0.042471 | -0.2157 | 0.9679 | 0.972 | 0.7728
2012-10-30 03:45 | 25.5247 | 12.8033 | 13.991 | 14.0438 | 13.7715 | 46.5098 | 0.022436 | 0.70717 | 0.015866 | 31.5192 | 0.037567 | 0.70949 | 0.026653 | 18.886 | 13.0312 | 3.2064 | 0.042519 | -0.2279 | 0.9598 | 1.0126 | 0.7403
2012-10-30 03:45 | 25.53 | 12.8289 | 14.0329 | 13.9882 | 13.7565 | 46.4907 | 0.022448 | 0.70686 | 0.015868 | 31.4889 | 0.037539 | 0.70906 | 0.026617 | 18.8886 | 13.0121 | 3.2051 | 0.042485 | -0.1832 | 1.0208 | 0.9761 | 0.7444
2012-10-30 03:45 | 25.5285 | 12.746 | 13.9704 | 13.9704 | 13.7713 | 46.4786 | 0.022445 | 0.70698 | 0.015868 | 31.4982 | 0.037564 | 0.70951 | 0.026652 | 18.8879 | 13.035 | 3.2042 | 0.04252 | -0.289 | 0.9354 | 0.9354 | 0.7363
2012-10-30 03:45 | 25.5382 | 12.7517 | 13.9761 | 13.972 | 13.785 | 46.4942 | 0.022436 | 0.70679 | 0.015858 | 31.5025 | 0.037539 | 0.70917 | 0.026622 | 18.8917 | 13.0122 | 3.2053 | 0.04248 | -0.2605 | 0.9639 | 0.9598 | 0.7728
2012-10-30 03:45 | 25.5434 | 12.8139 | 13.9569 | 14.026 | 13.778 | 46.4913 | 0.022439 | 0.70683 | 0.015861 | 31.5002 | 0.037551 | 0.70927 | 0.026634 | 18.8883 | 13.0214 | 3.2051 | 0.042495 | -0.2075 | 0.9355 | 1.0046 | 0.7566
2012-10-30 03:45 | 25.5464 | 12.7232 | 13.9558 | 14.0005 | 13.7932 | 46.4867 | 0.022455 | 0.70698 | 0.015875 | 31.4842 | 0.037564 | 0.70943 | 0.026649 | 18.886 | 13.0448 | 3.2048 | 0.042524 | -0.3216 | 0.911 | 0.9557 | 0.7484
2012-10-30 03:45 | 25.5502 | 12.7677 | 13.9637 | 13.9799 | 13.8255 | 46.4965 | 0.022433 | 0.70689 | 0.015858 | 31.5112 | 0.037555 | 0.70927 | 0.026637 | 18.8863 | 13.0486 | 3.2055 | 0.042495 | -0.2809 | 0.9151 | 0.9313 | 0.7769
2012-10-30 03:45 | 25.5449 | 12.7665 | 13.9787 | 14.0275 | 13.7795 | 46.4844 | 0.022442 | 0.70695 | 0.015865 | 31.5012 | 0.037555 | 0.70917 | 0.026633 | 18.8836 | 13.0636 | 3.2046 | 0.042498 | -0.2971 | 0.9151 | 0.9639 | 0.7159
2012-10-30 03:46 | 25.5502 | 12.7677 | 14.0043 | 13.9962 | 13.7442 | 46.4959 | 0.022445 | 0.70693 | 0.015867 | 31.4963 | 0.037555 | 0.7092 | 0.026634 | 18.8844 | 13.0079 | 3.2054 | 0.042501 | -0.2402 | 0.9964 | 0.9883 | 0.7363
2012-10-30 03:46 | 25.5554 | 12.7648 | 13.9607 | 13.9729 | 13.7453 | 46.4803 | 0.02243 | 0.70686 | 0.015855 | 31.5141 | 0.037567 | 0.70936 | 0.026649 | 18.8826 | 13.0375 | 3.2044 | 0.042504 | -0.2727 | 0.9232 | 0.9354 | 0.7078
2012-10-30 03:46 | 25.5614 | 12.7952 | 13.9789 | 13.983 | 13.7879 | 46.485 | 0.02243 | 0.70679 | 0.015853 | 31.5109 | 0.037542 | 0.7091 | 0.026621 | 18.8882 | 13.0232 | 3.2047 | 0.042474 | -0.228 | 0.9557 | 0.9598 | 0.7647
2012-10-30 03:46 | 25.5681 | 12.7653 | 13.9897 | 13.9816 | 13.7377 | 46.5011 | 0.022436 | 0.70688 | 0.015859 | 31.5064 | 0.037555 | 0.70929 | 0.026637 | 18.8867 | 13.0339 | 3.2058 | 0.042496 | -0.2686 | 0.9558 | 0.9477 | 0.7038
2012-10-30 03:46 | 25.5651 | 12.7134 | 13.9704 | 13.9989 | 13.7753 | 46.4751 | 0.022445 | 0.70683 | 0.015865 | 31.4918 | 0.037555 | 0.70919 | 0.026634 | 18.884 | 13.0269 | 3.204 | 0.042499 | -0.3135 | 0.9435 | 0.972 | 0.7484
2012-10-30 03:46 | 25.5696 | 12.7464 | 13.9668 | 13.9749 | 13.7676 | 46.4838 | 0.022433 | 0.70686 | 0.015857 | 31.5099 | 0.037545 | 0.70919 | 0.026626 | 18.889 | 13.0314 | 3.2046 | 0.042483 | -0.285 | 0.9354 | 0.9435 | 0.7362
2012-10-30 03:46 | 25.5689 | 12.7742 | 13.9295 | 13.9783 | 13.7425 | 46.4821 | 0.022439 | 0.70693 | 0.015863 | 31.5047 | 0.03757 | 0.70927 | 0.026647 | 18.8788 | 12.9778 | 3.2045 | 0.04251 | -0.2036 | 0.9517 | 1.0005 | 0.7647
2012-10-30 03:46 | 25.5794 | 12.7399 | 13.9726 | 14.0091 | 13.7408 | 46.4809 | 0.022436 | 0.70689 | 0.01586 | 31.507 | 0.037558 | 0.70925 | 0.026638 | 18.884 | 13.0127 | 3.2044 | 0.042498 | -0.2728 | 0.9599 | 0.9964 | 0.7281
2012-10-30 03:46 | 25.5801 | 12.7691 | 13.9245 | 13.9895 | 13.7415 | 46.4705 | 0.022436 | 0.70678 | 0.015857 | 31.5018 | 0.037558 | 0.70938 | 0.026643 | 18.8875 | 13.0052 | 3.2037 | 0.0425 | -0.2361 | 0.9193 | 0.9843 | 0.7363
2012-10-30 03:46 | 25.5823 | 12.7632 | 13.9633 | 13.9958 | 13.7641 | 46.474 | 0.022448 | 0.70686 | 0.015868 | 31.4889 | 0.037558 | 0.7092 | 0.026636 | 18.8829 | 12.983 | 3.2039 | 0.042504 | -0.2198 | 0.9803 | 1.0128 | 0.7811
2012-10-30 03:46 | 25.5906 | 12.7267 | 13.9634 | 13.9675 | 13.7317 | 46.4809 | 0.022426 | 0.70691 | 0.015853 | 31.5217 | 0.037564 | 0.70929 | 0.026644 | 18.8821 | 13.0239 | 3.2044 | 0.042497 | -0.2972 | 0.9395 | 0.9436 | 0.7078
2012-10-30 03:46 | 25.5936 | 12.7663 | 13.9339 | 13.9989 | 13.7388 | 46.489 | 0.022445 | 0.70688 | 0.015866 | 31.4937 | 0.037558 | 0.70932 | 0.026641 | 18.8859 | 13.0269 | 3.205 | 0.042507 | -0.2606 | 0.907 | 0.972 | 0.7119
2012-10-30 03:46 | 25.6018 | 12.7786 | 13.9462 | 13.9665 | 13.7795 | 46.4734 | 0.022433 | 0.70691 | 0.015858 | 31.5118 | 0.037558 | 0.70933 | 0.026641 | 18.8863 | 13.0269 | 3.2039 | 0.042499 | -0.2483 | 0.9193 | 0.9396 | 0.7526
2012-10-30 03:46 | 25.604 | 12.7523 | 13.9281 | 13.9606 | 13.7329 | 46.4717 | 0.022423 | 0.70691 | 0.015851 | 31.5259 | 0.037548 | 0.70935 | 0.026635 | 18.8917 | 13.0129 | 3.2038 | 0.042486 | -0.2606 | 0.9152 | 0.9477 | 0.72
2012-10-30 03:46 | 25.6018 | 12.7542 | 13.9381 | 13.9584 | 13.7836 | 46.485 | 0.022439 | 0.70695 | 0.015863 | 31.5054 | 0.037555 | 0.70936 | 0.02664 | 18.8886 | 13.0107 | 3.2047 | 0.042503 | -0.2565 | 0.9274 | 0.9477 | 0.7729
2012-10-30 03:46 | 25.613 | 12.7247 | 13.9818 | 13.994 | 13.7338 | 46.4855 | 0.022436 | 0.70693 | 0.015861 | 31.5089 | 0.037545 | 0.70933 | 0.026632 | 18.8929 | 13.0422 | 3.2047 | 0.042493 | -0.3175 | 0.9396 | 0.9518 | 0.6916
2012-10-30 03:46 | 25.6108 | 12.7591 | 13.943 | 13.9511 | 13.7682 | 46.4826 | 0.022445 | 0.70703 | 0.015869 | 31.5008 | 0.037561 | 0.70932 | 0.026643 | 18.8844 | 12.983 | 3.2045 | 0.042512 | -0.2239 | 0.96 | 0.9681 | 0.7852
2012-10-30 03:46 | 25.6175 | 12.7169 | 13.9375 | 14.0107 | 13.7302 | 46.4676 | 0.022433 | 0.70703 | 0.015861 | 31.5176 | 0.037564 | 0.70942 | 0.026649 | 18.8856 | 13.0223 | 3.2035 | 0.04251 | -0.3054 | 0.9152 | 0.9884 | 0.7079
2012-10-30 03:46 | 25.6242 | 12.8051 | 13.9889 | 13.9726 | 13.7856 | 46.478 | 0.02243 | 0.70693 | 0.015857 | 31.5173 | 0.037545 | 0.70942 | 0.026635 | 18.8952 | 13.0371 | 3.2042 | 0.042492 | -0.232 | 0.9518 | 0.9355 | 0.7485
2012-10-30 03:46 | 25.628 | 12.7234 | 13.9805 | 13.9439 | 13.7366 | 46.4797 | 0.022439 | 0.70698 | 0.015864 | 31.5067 | 0.037555 | 0.70933 | 0.026639 | 18.8878 | 13.0409 | 3.2043 | 0.042503 | -0.3175 | 0.9396 | 0.903 | 0.6957
2012-10-30 03:46 | 25.6235 | 12.6944 | 13.9313 | 14.0004 | 13.7606 | 46.4711 | 0.022433 | 0.70676 | 0.015855 | 31.5054 | 0.037555 | 0.70923 | 0.026635 | 18.8852 | 13.012 | 3.2037 | 0.04249 | -0.3176 | 0.9193 | 0.9884 | 0.7486
2012-10-30 03:46 | 25.6362 | 13.041 | 14.6103 | 13.9684 | 13.7407 | 48.1997 | 1.8155 | 45.2056 | 82.0686 | 24.9005 | 2.7276 | 43.8185 | 119.5201 | 16.0648 | 13.0288 | 3.3229 | 201.5887 | 0.0122 | 1.5815 | 0.9396 | 0.7119
2012-10-30 03:46 | 25.6407 | 18.0732 | 19.9757 | 14.1639 | 13.8713 | 52.8375 | 2.0666 | 51.7929 | 107.0371 | 25.0614 | 3.2159 | 51.6452 | 166.084 | 16.0595 | 13.0536 | 3.6426 | 273.1212 | 5.0196 | 6.9221 | 1.1103 | 0.8177
2012-10-30 03:46 | 25.6392 |  | 23.6279 | 14.7554 | 14.5158 | 53.4599 | 2.0649 | 51.7827 | 106.9242 | 25.078 | 3.2152 | 51.6348 | 166.0145 | 16.0597 | 13.0074 | 3.6855 | 272.9388 | -13.0074 | 10.6205 | 1.748 | 1.5084
2012-10-30 03:46 | 25.6489 | 30.0944 |  | 15.7631 | 15.8563 | 53.9252 | 2.0635 | 51.7764 | 106.8424 | 25.0911 | 3.2148 | 51.6283 | 165.9743 | 16.0596 | 13.0822 | 3.7176 | 272.8167 | 17.0122 | -13.0822 | 2.6809 | 2.7741
2012-10-30 03:46 | 25.6504 | 35.9779 | 28.5522 | 17.2504 | 17.392 | 54.3812 | 2.0624 | 51.7723 | 106.7755 | 25.1029 | 3.2145 | 51.6242 | 165.9478 | 16.0596 | 13.0511 | 3.749 | 272.7233 | 22.9268 | 15.5011 | 4.1993 | 4.3409
2012-10-30 03:47 | 25.6534 | 41.5322 | 31.2872 | 19.1347 | 19.038 | 54.8406 | 2.0613 | 51.7704 | 106.7167 | 25.1149 | 3.2144 | 51.6223 | 165.9364 | 16.0596 | 13.0582 | 3.7807 | 272.6531 | 28.474 | 18.229 | 6.0765 | 5.9798
2012-10-30 03:47 | 25.6602 | 46.7867 | 34.1112 | 21.357 | 20.9838 | 55.2879 | 2.0603 | 51.7699 | 106.6624 | 25.1271 | 3.2144 | 51.6215 | 165.9329 | 16.0594 | 13.0568 | 3.8115 | 272.5954 | 33.7299 | 21.0544 | 8.3002 | 7.927
2012-10-30 03:47 | 25.6579 | 51.5025 | 36.5252 | 24.073 | 23.088 | 55.7202 | 2.0594 | 51.77 | 106.6176 | 25.1378 | 3.2144 | 51.6215 | 165.9342 | 16.0592 | 13.0545 | 3.8414 | 272.5518 | 38.448 | 23.4707 | 11.0185 | 10.0335
2012-10-30 03:47 | 25.6654 | 56.0137 | 38.9146 |  |  | 56.1115 | 2.0587 | 51.7712 | 106.5798 | 25.1479 | 3.2146 | 51.6227 | 165.9453 | 16.0589 | 13.0254 | 3.8683 | 272.5251 | 42.9883 | 25.8892 | -13.0254 | -13.0254
2012-10-30 03:47 | 25.6699 | 60.7692 | 41.3563 | 29.4166 | 27.0279 | 56.501 | 2.0579 | 51.7726 | 106.5441 | 25.1577 | 3.2147 | 51.6239 | 165.9552 | 16.0587 | 13.1072 | 3.8952 | 272.4993 | 47.662 | 28.2491 | 16.3094 | 13.9207
2012-10-30 03:47 | 25.6706 | 65.4578 | 44.0604 | 33.0487 | 29.3068 | 56.9558 | 2.0571 | 51.7743 | 106.5072 | 25.168 | 3.2149 | 51.6254 | 165.9699 | 16.0582 | 13.0591 | 3.9265 | 272.4771 | 52.3987 | 31.0013 | 19.9896 | 16.2477
2012-10-30 03:47 | 25.6766 | 69.9986 | 46.9009 | 36.9221 | 31.8156 | 57.3818 | 2.0564 | 51.7759 | 106.4712 | 25.1781 | 3.2151 | 51.6269 | 165.9854 | 16.0577 | 13.0773 | 3.9559 | 272.4566 | 56.9213 | 33.8236 | 23.8448 | 18.7383
2012-10-30 03:47 | 25.6774 | 73.9987 | 49.3391 | 40.3618 | 33.8247 | 57.7256 | 2.0557 | 51.7772 | 106.4405 | 25.1867 | 3.2153 | 51.6282 | 165.9983 | 16.0572 | 13.1229 | 3.9796 | 272.4388 | 60.8758 | 36.2162 | 27.2389 | 20.7018
2012-10-30 03:47 | 25.6834 | 77.9431 | 51.7421 | 43.9085 | 35.8375 | 58.136 | 2.0551 | 51.779 | 106.4103 | 25.1956 | 3.2155 | 51.6298 | 166.0149 | 16.0566 | 13.0923 | 4.0079 | 272.4252 | 64.8508 | 38.6498 | 30.8162 | 22.7452
2012-10-30 03:47 | 25.6699 | 82.1433 | 54.3495 | 47.7228 | 38.0887 | 58.5445 | 2.0543 | 51.7802 | 106.3747 | 25.2052 | 3.2157 | 51.6309 | 166.0295 | 16.0559 | 13.0869 | 4.0361 | 272.4041 | 69.0564 | 41.2626 | 34.6359 | 25.0018
2012-10-30 03:47 | 25.6751 | 86.2887 | 56.9497 | 51.6482 | 40.4028 | 58.937 | 2.0537 | 51.7819 | 106.3457 | 25.2137 | 3.2159 | 51.6323 | 166.0449 | 16.0553 | 13.0921 | 4.0631 | 272.3906 | 73.1966 | 43.8576 | 38.5561 | 27.3107
2012-10-30 03:47 | 25.6879 | 90.3289 | 59.5666 | 55.6638 | 42.6855 | 59.2646 | 2.0532 | 51.7834 | 106.3196 | 25.2213 | 3.2161 | 51.6338 | 166.0616 | 16.0546 | 13.1049 | 4.0857 | 272.3813 | 77.224 | 46.4617 | 42.5589 | 29.5806
2012-10-30 03:47 | 25.6916 | 93.9694 | 61.7092 | 59.0137 | 44.583 | 59.6027 | 2.0526 | 51.7838 | 106.2908 | 25.2286 | 3.2163 | 51.634 | 166.0692 | 16.054 | 13.0842 | 4.109 | 272.36 | 80.8852 | 48.625 | 45.9295 | 31.4988
2012-10-30 03:47 | 25.6976 | 97.2695 | 63.8576 | 62.3823 | 46.5503 | 59.9362 | 2.052 | 51.7847 | 106.2646 | 25.2357 | 3.2165 | 51.6349 | 166.0811 | 16.0534 | 13.0902 | 4.132 | 272.3456 | 84.1793 | 50.7674 | 49.2921 | 33.4601
2012-10-30 03:47 | 25.6991 | 100.753 | 66.178 | 66.0282 | 48.6426 | 60.2876 | 2.0515 | 51.7861 | 106.2375 | 25.2434 | 3.2167 | 51.6361 | 166.0961 | 16.0527 | 13.0835 | 4.1562 | 272.3336 | 87.6695 | 53.0945 | 52.9447 | 35.5591
2012-10-30 03:47 | 25.6976 | 104.5284 | 68.5827 | 69.74 | 50.7475 | 60.6673 | 2.0509 | 51.7874 | 106.2092 | 25.2514 | 3.2169 | 51.6372 | 166.1105 | 16.052 | 13.0576 | 4.1824 | 272.3197 | 91.4708 | 55.5251 | 56.6824 | 37.6899
2012-10-30 03:47 | 25.7028 | 107.9409 | 71.0621 | 73.8065 | 53.0728 | 61.0158 | 2.0503 | 51.788 | 106.1809 | 25.2588 | 3.217 | 51.638 | 166.1211 | 16.0514 | 13.0954 | 4.2064 | 272.302 | 94.8455 | 57.9667 | 60.7111 | 39.9774
2012-10-30 03:47 | 25.7118 | 111.3063 | 73.3839 | 77.4879 | 55.1251 | 61.3469 | 2.0498 | 51.7889 | 106.1553 | 25.2657 | 3.2172 | 51.6385 | 166.1317 | 16.0507 | 13.0759 | 4.2293 | 272.287 | 98.2304 | 60.308 | 64.412 | 42.0492
2012-10-30 03:47 | 25.7066 | 114.6298 | 75.6381 | 81.1655 | 57.1826 | 61.7203 | 2.0492 | 51.7902 | 106.1276 | 25.2735 | 3.2174 | 51.6398 | 166.1463 | 16.0501 | 13.0951 | 4.255 | 272.274 | 101.5347 | 62.543 | 68.0704 | 44.0875
2012-10-30 03:47 | 25.7125 | 117.8615 | 78.1245 | 85.1505 | 59.4366 | 62.089 | 2.0485 | 51.791 | 106.0934 | 25.2826 | 3.2176 | 51.6405 | 166.1587 | 16.0494 | 13.1295 | 4.2804 | 272.2521 | 104.732 | 64.995 | 72.021 | 46.3071
2012-10-30 03:47 | 25.7133 | 121.9389 | 81.0278 | 89.7674 | 62.0664 | 62.497 | 2.0479 | 51.7922 | 106.0634 | 25.2909 | 3.2178 | 51.6414 | 166.1699 | 16.0488 | 13.0937 | 4.3085 | 272.2333 | 108.8452 | 67.9341 | 76.6737 | 48.9727
2012-10-30 03:47 | 25.7148 | 125.9517 | 83.6843 | 94.072 | 64.4329 | 62.8981 | 2.0473 | 51.7932 | 106.0387 | 25.2977 | 3.2179 | 51.6424 | 166.1804 | 16.0484 | 13.0952 | 4.3362 | 272.2191 | 112.8565 | 70.5891 | 80.9768 | 51.3377
2012-10-30 03:47 | 25.7096 | 129.0123 | 85.8568 | 97.5386 | 66.4689 | 63.2512 | 2.0468 | 51.7939 | 106.0117 | 25.3048 | 3.218 | 51.6429 | 166.1877 | 16.0481 | 13.1144 | 4.3605 | 272.1994 | 115.8979 | 72.7424 | 84.4242 | 53.3545
2012-10-30 03:47 | 25.7185 | 132.1033 | 88.2537 | 101.4945 | 68.734 | 63.5662 | 2.0463 | 51.7945 | 105.986 | 25.3115 | 3.2181 | 51.6434 | 166.1934 | 16.0478 | 13.1151 | 4.3823 | 272.1794 | 118.9882 | 75.1386 | 88.3794 | 55.6189
2012-10-30 03:47 | 25.7283 | 135.6056 | 90.8425 | 105.2548 | 70.9192 | 63.9453 | 2.0457 | 51.7951 | 105.9584 | 25.3188 | 3.2182 | 51.6441 | 166.2011 | 16.0475 | 13.0924 | 4.4084 | 272.1595 | 122.5132 | 77.7501 | 92.1624 | 57.8268
2012-10-30 03:47 | 25.7238 | 138.8207 | 93.2412 | 108.7374 | 73.0827 | 64.303 | 2.0452 | 51.7952 | 105.9314 | 25.3253 | 3.2182 | 51.644 | 166.2014 | 16.0474 | 13.0431 | 4.433 | 272.1329 | 125.7776 | 80.1981 | 95.6943 | 60.0396
2012-10-30 03:47 | 25.726 | 141.488 | 95.1877 | 112.1003 | 75.0853 | 64.6076 | 2.0447 | 51.7956 | 105.9062 | 25.3317 | 3.2183 | 51.6443 | 166.2048 | 16.0473 | 13.0616 | 4.454 | 272.1111 | 128.4264 | 82.1261 | 99.0387 | 62.0237
2012-10-30 03:48 | 25.726 | 144.3494 | 97.3357 | 115.7291 | 76.9116 | 64.9335 | 2.0442 | 51.7959 | 105.8787 | 25.3386 | 3.2183 | 51.6446 | 166.2061 | 16.0473 | 13.0779 | 4.4765 | 272.0848 | 131.2715 | 84.2578 | 102.6512 | 63.8337
2012-10-30 03:48 | 25.7417 | 147.4276 | 99.635 | 119.006 | 78.9264 | 65.2508 | 2.0437 | 51.7964 | 105.854 | 25.3449 | 3.2183 | 51.6449 | 166.2084 | 16.0473 | 13.0691 | 4.4984 | 272.0625 | 134.3585 | 86.5659 | 105.9369 | 65.8573
2012-10-30 03:48 | 25.7425 | 150.4965 | 101.9343 | 122.5532 | 81.0456 | 65.6109 | 2.0431 | 51.7966 | 105.8277 | 25.3515 | 3.2183 | 51.6452 | 166.2089 | 16.0474 | 13.017 | 4.5232 | 272.0366 | 137.4795 | 88.9173 | 109.5362 | 68.0286
2012-10-30 03:48 | 25.7372 | 153.5454 | 104.1936 | 126.1339 | 83.1373 | 65.9501 | 2.0426 | 51.7972 | 105.8013 | 25.3584 | 3.2183 | 51.6456 | 166.2097 | 16.0476 | 13.0931 | 4.5466 | 272.0109 | 140.4523 | 91.1005 | 113.0408 | 70.0442
2012-10-30 03:48 | 25.7395 | 156.2627 | 106.3135 | 129.2747 | 85.0789 | 66.2749 | 2.0421 | 51.7973 | 105.7776 | 25.3642 | 3.2182 | 51.6456 | 166.2069 | 16.0479 | 13.0466 | 4.569 | 271.9845 | 143.2161 | 93.2669 | 116.2281 | 72.0323
2012-10-30 03:48 | 25.7335 | 158.9875 | 108.3045 | 132.6872 | 87.1811 | 66.6037 | 2.0416 | 51.7978 | 105.7525 | 25.3707 | 3.2182 | 51.646 | 166.2075 | 16.0481 | 13.0162 | 4.5917 | 271.96 | 145.9713 | 95.2883 | 119.671 | 74.1649
2012-10-30 03:48 | 25.738 | 161.8395 | 110.549 | 136.0437 | 89.205 | 66.9436 | 2.0411 | 51.7979 | 105.7263 | 25.3771 | 3.2181 | 51.6461 | 166.2034 | 16.0485 | 13.0532 | 4.6151 | 271.9298 | 148.7863 | 97.4958 | 122.9905 | 76.1518
2012-10-30 03:48 | 25.7515 | 164.7941 | 112.7695 | 139.3696 | 91.2443 | 67.2452 | 2.0406 | 51.7982 | 105.7013 | 25.3833 | 3.218 | 51.6463 | 166.2 | 16.049 | 13.0504 | 4.6359 | 271.9012 | 151.7437 | 99.7191 | 126.3192 | 78.1939
2012-10-30 03:48 | 25.7522 | 167.6119 | 114.7549 | 142.3433 | 93.1561 | 67.5342 | 2.0402 | 51.7982 | 105.6769 | 25.3892 | 3.2179 | 51.6462 | 166.1935 | 16.0495 | 13.1 | 4.6558 | 271.8704 | 154.5119 | 101.6549 | 129.2433 | 80.0561
2012-10-30 03:48 | 25.7559 | 170.4951 | 116.9682 | 145.3098 | 95.0172 | 67.8451 | 2.0397 | 51.7983 | 105.6525 | 25.3952 | 3.2178 | 51.6464 | 166.1891 | 16.0501 | 13.0548 | 4.6772 | 271.8416 | 157.4403 | 103.9134 | 132.255 | 81.9624
2012-10-30 03:48 | 25.7604 | 173.0393 | 119.1347 | 148.4517 | 96.9844 | 68.167 | 2.0392 | 51.7987 | 105.6296 | 25.4011 | 3.2177 | 51.6465 | 166.1832 | 16.0507 | 13.0553 | 4.6994 | 271.8128 | 159.984 | 106.0794 | 135.3964 | 83.9291
2012-10-30 03:48 | 25.7589 | 175.5452 | 121.1033 | 151.5741 | 98.9452 | 68.4722 | 2.0388 | 51.7986 | 105.6056 | 25.4067 | 3.2175 | 51.6464 | 166.1743 | 16.0515 | 13.0986 | 4.7205 | 271.7799 | 162.4466 | 108.0047 | 138.4755 | 85.8466
2012-10-30 03:48 | 25.7634 | 178.5292 | 123.4254 | 154.9402 | 101.0782 | 68.7923 | 2.0382 | 51.7988 | 105.5787 | 25.4134 | 3.2174 | 51.6465 | 166.1654 | 16.0524 | 13.0827 | 4.7425 | 271.7441 | 165.4465 | 110.3427 | 141.8575 | 87.9955
2012-10-30 03:48 | 25.7679 | 181.4774 | 125.6775 | 158.2042 | 103.1393 | 69.0888 | 2.0378 | 51.7989 | 105.5535 | 25.4195 | 3.2172 | 51.6466 | 166.1568 | 16.0533 | 13.0628 | 4.763 | 271.7103 | 168.4146 | 112.6147 | 145.1414 | 90.0765
2012-10-30 03:48 | 25.7724 | 184.1425 | 127.5512 | 161.0238 | 104.9397 | 69.3639 | 2.0373 | 51.7994 | 105.5321 | 25.4252 | 3.217 | 51.6472 | 166.1509 | 16.0543 | 13.0876 | 4.7819 | 271.683 | 171.0549 | 114.4636 | 147.9362 | 91.8521
2012-10-30 03:48 | 25.7702 | 186.5914 | 129.4617 | 163.6972 | 106.7715 | 69.6558 | 2.0369 | 51.7998 | 105.5126 | 25.4303 | 3.2169 | 51.6475 | 166.1436 | 16.0552 | 13.1099 | 4.8021 | 271.6562 | 173.4815 | 116.3518 | 150.5873 | 93.6616
2012-10-30 03:48 | 25.7784 | 188.9896 | 131.5601 | 166.4159 | 108.5651 | 69.917 | 2.0365 | 51.7996 | 105.4908 | 25.4354 | 3.2167 | 51.6473 | 166.1326 | 16.0561 | 13.1018 | 4.8201 | 271.6234 | 175.8878 | 118.4583 | 153.3141 | 95.4633
2012-10-30 03:48 | 25.7814 | 191.6002 | 133.6098 | 169.094 | 110.3123 | 70.1545 | 2.0361 | 51.7993 | 105.468 | 25.4406 | 3.2165 | 51.6469 | 166.1204 | 16.0571 | 13.0885 | 4.8365 | 271.5884 | 178.5117 | 120.5213 | 156.0055 | 97.2238
2012-10-30 03:48 | 25.7829 | 193.9358 | 135.3887 | 171.9652 | 112.1723 | 70.4238 | 2.0357 | 51.7997 | 105.4467 | 25.4461 | 3.2162 | 51.6472 | 166.1103 | 16.0582 | 13.09 | 4.855 | 271.557 | 180.8458 | 122.2987 | 158.8752 | 99.0823
2012-10-30 03:48 | 25.7896 | 196.2817 | 137.297 | 174.6187 | 113.8415 | 70.6648 | 2.0353 | 51.7996 | 105.4261 | 25.451 | 3.216 | 51.6472 | 166.0992 | 16.0593 | 13.0926 | 4.8716 | 271.5253 | 183.1891 | 124.2044 | 161.5261 | 100.7489
2012-10-30 03:48 | 25.7806 | 198.6129 | 139.0269 | 177.2009 | 115.5828 | 70.937 | 2.0349 | 51.7997 | 105.4066 | 25.4558 | 3.2158 | 51.6471 | 166.0871 | 16.0604 | 13.0551 | 4.8904 | 271.4937 | 185.5578 | 125.9718 | 164.1458 | 102.5277
2012-10-30 03:48 | 25.7814 | 200.9154 | 140.8757 | 179.7329 | 117.225 | 71.1763 | 2.0345 | 51.7995 | 105.3853 | 25.4608 | 3.2156 | 51.647 | 166.0738 | 16.0616 | 13.0234 | 4.9069 | 271.4591 | 187.892 | 127.8523 | 166.7095 | 104.2016
2012-10-30 03:48 | 25.7866 | 203.6048 | 143.0267 | 182.7693 | 119.2482 | 71.441 | 2.034 | 51.7995 | 105.3626 | 25.4662 | 3.2153 | 51.6469 | 166.0583 | 16.063 | 13.0408 | 4.9251 | 271.4209 | 190.564 | 129.9859 | 169.7285 | 106.2074
2012-10-30 03:48 | 25.7919 | 206.0273 | 144.8881 | 185.7226 | 121.1857 | 71.7172 | 2.0336 | 51.799 | 105.3376 | 25.4718 | 3.215 | 51.6464 | 166.0413 | 16.0644 | 13.0664 | 4.9442 | 271.3789 | 192.9609 | 131.8217 | 172.6562 | 108.1193
2012-10-30 03:48 | 25.7896 | 208.2575 | 146.6747 | 187.8978 | 122.8397 | 71.9438 | 2.0332 | 51.7981 | 105.315 | 25.4764 | 3.2146 | 51.6456 | 166.0223 | 16.0657 | 13.0723 | 4.9598 | 271.3373 | 195.1852 | 133.6024 | 174.8255 | 109.7674
2012-10-30 03:48 | 25.8046 | 210.6864 | 148.4304 | 190.4136 | 124.5236 | 72.1749 | 2.0328 | 51.7981 | 105.2966 | 25.4808 | 3.2144 | 51.6455 | 166.0081 | 16.067 | 13.071 | 4.9757 | 271.3047 | 197.6154 | 135.3594 | 177.3426 | 111.4526
2012-10-30 03:49 | 25.8016 | 212.8296 | 149.9963 | 192.8345 | 125.9624 | 72.3992 | 2.0325 | 51.798 | 105.2787 | 25.485 | 3.2141 | 51.6453 | 165.9942 | 16.0682 | 13.0721 | 4.9912 | 271.2729 | 199.7575 | 136.9242 | 179.7624 | 112.8903
2012-10-30 03:49 | 25.7986 | 214.9925 | 151.8648 | 195.406 | 127.6535 | 72.6379 | 2.032 | 51.7975 | 105.2543 | 25.4904 | 3.2138 | 51.6448 | 165.9738 | 16.0699 | 13.0324 | 5.0077 | 271.2281 | 201.9601 | 138.8324 | 182.3736 | 114.6211
2012-10-30 03:49 | 25.8008 | 217.499 | 154.0496 | 198.2476 | 129.6944 | 72.8974 | 2.0316 | 51.7974 | 105.231 | 25.496 | 3.2134 | 51.6448 | 165.9554 | 16.0717 | 13.0509 | 5.0255 | 271.1864 | 204.4481 | 140.9987 | 185.1967 | 116.6435
2012-10-30 03:49 | 25.8053 | 219.9918 | 156.0421 | 200.8173 | 131.6023 | 73.1418 | 2.0312 | 51.7971 | 105.211 | 25.5006 | 3.2131 | 51.6443 | 165.9372 | 16.0731 | 13.1002 | 5.0424 | 271.1482 | 206.8916 | 142.9419 | 187.7171 | 118.5021
2012-10-30 03:49 | 25.8121 | 222.1281 | 157.7138 | 203.3863 | 133.0675 | 73.358 | 2.0309 | 51.7974 | 105.1938 | 25.505 | 3.2128 | 51.6446 | 165.9232 | 16.0747 | 13.0581 | 5.0573 | 271.1171 | 209.07 | 144.6557 | 190.3282 | 120.0094
2012-10-30 03:49 | 25.8008 | 224.1435 | 159.3999 | 205.922 | 134.5922 | 73.5782 | 2.0305 | 51.7967 | 105.171 | 25.5099 | 3.2124 | 51.6439 | 165.9021 | 16.0763 | 13.0916 | 5.0725 | 271.073 | 211.0519 | 146.3083 | 192.8304 | 121.5006
2012-10-30 03:49 | 25.8136 | 226.5681 | 161.3095 | 208.38 | 136.3718 | 73.8186 | 2.0301 | 51.7966 | 105.1505 | 25.5147 | 3.2121 | 51.6439 | 165.8843 | 16.078 | 13.0881 | 5.0891 | 271.0348 | 213.48 | 148.2214 | 195.2919 | 123.2837
2012-10-30 03:49 | 25.8128 | 228.9349 | 163.2269 | 210.8049 | 138.0745 | 74.0497 | 2.0297 | 51.7965 | 105.1314 | 25.5193 | 3.2117 | 51.6436 | 165.8655 | 16.0797 | 13.0873 | 5.105 | 270.997 | 215.8476 | 150.1396 | 197.7176 | 124.9872
2012-10-30 03:49 | 25.8143 | 230.9257 | 164.8492 | 213.2676 | 139.5473 | 74.2566 | 2.0294 | 51.7962 | 105.1128 | 25.5236 | 3.2114 | 51.6434 | 165.8484 | 16.0812 | 13.0603 | 5.1193 | 270.9611 | 217.8654 | 151.7889 | 200.2073 | 126.487
2012-10-30 03:49 | 25.8113 | 232.754 | 166.2892 | 215.4777 | 141.0437 | 74.4635 | 2.029 | 51.7957 | 105.0948 | 25.5274 | 3.2111 | 51.6429 | 165.8297 | 16.0827 | 13.0614 | 5.1335 | 270.9245 | 219.6926 | 153.2278 | 202.4163 | 127.9823
2012-10-30 03:49 | 25.8188 | 235.0478 | 168.1887 | 218.0131 | 142.7286 | 74.6941 | 2.0286 | 51.7955 | 105.0741 | 25.5323 | 3.2107 | 51.6427 | 165.809 | 16.0846 | 13.0526 | 5.1494 | 270.883 | 221.9952 | 155.1361 | 204.9605 | 129.676
2012-10-30 03:49 | 25.8218 | 237.5282 | 170.2101 | 220.7601 | 144.5413 | 74.9449 | 2.0282 | 51.7953 | 105.0526 | 25.5373 | 3.2103 | 51.6423 | 165.7868 | 16.0865 | 13.0719 | 5.1667 | 270.8394 | 224.4563 | 157.1382 | 207.6882 | 131.4694
2012-10-30 03:49 | 25.8225 | 239.5577 | 171.8956 | 223.0831 | 146.0919 | 75.1455 | 2.0279 | 51.795 | 105.0343 | 25.5414 | 3.2099 | 51.642 | 165.7666 | 16.0883 | 13.0482 | 5.1805 | 270.8009 | 226.5095 | 158.8474 | 210.0349 | 133.0437
2012-10-30 03:49 | 25.8218 | 241.4827 | 173.4521 | 225.1522 | 147.5347 | 75.3524 | 2.0276 | 51.7946 | 105.0168 | 25.5453 | 3.2096 | 51.6417 | 165.747 | 16.09 | 13.0475 | 5.1948 | 270.7638 | 228.4352 | 160.4046 | 212.1047 | 134.4872
2012-10-30 03:49 | 25.8278 | 243.3624 | 174.9382 | 227.251 | 148.9045 | 75.5396 | 2.0272 | 51.7942 | 104.9985 | 25.5493 | 3.2092 | 51.6413 | 165.7275 | 16.0916 | 13.086 | 5.2077 | 270.726 | 230.2764 | 161.8522 | 214.165 | 135.8185
2012-10-30 03:49 | 25.8345 | 245.7118 | 176.7621 | 229.6604 | 150.6039 | 75.7749 | 2.0269 | 51.7941 | 104.9791 | 25.554 | 3.2088 | 51.6412 | 165.7069 | 16.0935 | 13.052 | 5.2239 | 270.686 | 232.6598 | 163.7101 | 216.6084 | 137.5519
2012-10-30 03:49 | 25.8465 | 247.8535 | 178.5536 | 232.0623 | 152.1598 | 75.9933 | 2.0265 | 51.7933 | 104.9579 | 25.5583 | 3.2084 | 51.6404 | 165.6818 | 16.0955 | 13.1129 | 5.239 | 270.6397 | 234.7406 | 165.4407 | 218.9494 | 139.0469
2012-10-30 03:49 | 25.8413 | 249.4838 | 180.0472 | 234.0687 | 153.5452 | 76.1713 | 2.0262 | 51.7931 | 104.9424 | 25.5619 | 3.208 | 51.6401 | 165.6632 | 16.0971 | 13.0344 | 5.2512 | 270.6055 | 236.4494 | 167.0128 | 221.0343 | 140.5108
2012-10-30 03:49 | 25.8442 | 251.2777 | 181.5381 | 235.89 | 154.847 | 76.3187 | 2.0259 | 51.793 | 104.9263 | 25.5657 | 3.2077 | 51.64 | 165.6461 | 16.0987 | 13.0536 | 5.2614 | 270.5724 | 238.2241 | 168.4845 | 222.8364 | 141.7934
2012-10-30 03:49 | 25.8465 | 253.0503 | 182.982 | 237.8316 | 156.1414 | 76.5244 | 2.0255 | 51.7919 | 104.9069 | 25.5694 | 3.2073 | 51.6389 | 165.6218 | 16.1004 | 13.0477 | 5.2756 | 270.5287 | 240.0026 | 169.9343 | 224.7839 | 143.0937
2012-10-30 03:49 | 25.8465 | 254.9719 | 184.6087 | 240.1465 | 157.7637 | 76.7174 | 2.0252 | 51.7914 | 104.8884 | 25.5733 | 3.2069 | 51.6383 | 165.5979 | 16.1023 | 13.0233 | 5.2889 | 270.4862 | 241.9486 | 171.5854 | 227.1232 | 144.7404
2012-10-30 03:49 | 25.8517 | 256.8388 | 186.1982 | 242.1673 | 159.129 | 76.9249 | 2.0249 | 51.7908 | 104.872 | 25.5768 | 3.2065 | 51.6379 | 165.5778 | 16.1041 | 12.9674 | 5.3032 | 270.4498 | 243.8714 | 173.2308 | 229.1999 | 146.1616
2012-10-30 03:49 | 25.857 | 258.5604 | 187.6042 | 243.9716 | 160.3286 | 77.07 | 2.0246 | 51.7906 | 104.8567 | 25.5803 | 3.2062 | 51.6375 | 165.5592 | 16.1056 | 13.0379 | 5.3132 | 270.4158 | 245.5225 | 174.5663 | 230.9337 | 147.2907
2012-10-30 03:49 | 25.8607 | 260.3221 | 189.0051 | 245.8007 | 161.6089 | 77.2607 | 2.0243 | 51.79 | 104.8385 | 25.5842 | 3.2058 | 51.6371 | 165.537 | 16.1075 | 13.0619 | 5.3264 | 270.3756 | 247.2602 | 175.9432 | 232.7388 | 148.547
2012-10-30 03:49 | 25.8644 | 261.923 | 190.281 | 247.617 | 162.8979 | 77.4179 | 2.024 | 51.7895 | 104.8224 | 25.5876 | 3.2054 | 51.6365 | 165.5169 | 16.1091 | 13.0453 | 5.3372 | 270.3393 | 248.8777 | 177.2357 | 234.5717 | 149.8526
2012-10-30 03:49 | 25.8637 | 263.9038 | 191.7925 | 249.6588 | 164.2493 | 77.6127 | 2.0237 | 51.789 | 104.8038 | 25.5916 | 3.205 | 51.636 | 165.4932 | 16.1111 | 13.0364 | 5.3506 | 270.297 | 250.8674 | 178.7561 | 236.6224 | 151.2129
2012-10-30 03:49 | 25.8689 | 265.6557 | 193.4519 | 251.7988 | 165.8566 | 77.7883 | 2.0234 | 51.7884 | 104.7861 | 25.5954 | 3.2046 | 51.6354 | 165.4696 | 16.1131 | 13.0498 | 5.3627 | 270.2557 | 252.6059 | 180.4021 | 238.749 | 152.8068
2012-10-30 03:50 | 25.8659 | 267.2738 | 194.9347 | 253.5463 | 167.2818 | 77.9635 | 2.0231 | 51.7878 | 104.7708 | 25.5985 | 3.2042 | 51.6348 | 165.4477 | 16.1148 | 13.0346 | 5.3748 | 270.2184 | 254.2392 | 181.9001 | 240.5117 | 154.2472
2012-10-30 03:50 | 25.86 | 268.8648 | 196.297 | 255.3407 | 168.4753 | 78.1467 | 2.0228 | 51.7873 | 104.7539 | 25.6022 | 3.2038 | 51.6343 | 165.4275 | 16.1165 | 13.0572 | 5.3874 | 270.1814 | 255.8076 | 183.2398 | 242.2835 | 155.4181
2012-10-30 03:50 | 25.863 | 270.4869 | 197.5506 | 257.0569 | 169.7485 | 78.2929 | 2.0225 | 51.7872 | 104.7401 | 25.6054 | 3.2035 | 51.6343 | 165.4092 | 16.1182 | 13.0927 | 5.3975 | 270.1493 | 257.3942 | 184.4579 | 243.9642 | 156.6558
2012-10-30 03:50 | 25.8637 | 272.1484 | 199.0068 | 258.7386 | 171.0201 | 78.4686 | 2.0222 | 51.7867 | 104.7238 | 25.6089 | 3.2031 | 51.6337 | 165.3869 | 16.12 | 13.0487 | 5.4096 | 270.1107 | 259.0997 | 185.9581 | 245.6899 | 157.9714
2012-10-30 03:50 | 25.8734 | 273.9108 | 200.5745 | 260.8062 | 172.4966 | 78.6552 | 2.0219 | 51.7863 | 104.7071 | 25.6126 | 3.2027 | 51.6333 | 165.3645 | 16.1219 | 13.0298 | 5.4225 | 270.0716 | 260.881 | 187.5447 | 247.7764 | 159.4668
2012-10-30 03:50 | 25.8772 | 275.5537 | 202.0305 | 262.7768 | 173.7378 | 78.828 | 2.0216 | 51.7856 | 104.689 | 25.6163 | 3.2022 | 51.6326 | 165.3394 | 16.1239 | 13.0988 | 5.4344 | 270.0284 | 262.4549 | 188.9317 | 249.678 | 160.639
2012-10-30 03:50 | 25.8757 | 277.2019 | 203.4184 | 264.5812 | 174.9978 | 79.0153 | 2.0213 | 51.7849 | 104.6726 | 25.6197 | 3.2018 | 51.632 | 165.3165 | 16.1258 | 13.081 | 5.4473 | 269.9892 | 264.1209 | 190.3374 | 251.5002 | 161.9168
2012-10-30 03:50 | 25.8862 | 278.8301 | 204.8933 | 266.3807 | 176.2744 | 79.1794 | 2.021 | 51.7841 | 104.6546 | 25.6232 | 3.2014 | 51.6311 | 165.2906 | 16.1278 | 13.0793 | 5.4586 | 269.9452 | 265.7508 | 191.814 | 253.3014 | 163.1951
2012-10-30 03:50 | 25.8929 | 282.3472 | 207.8293 | 270.2218 | 179.0617 | 79.5221 | 2.0203 | 51.7818 | 104.6163 | 25.6304 | 3.2004 | 51.6289 | 165.2347 | 16.1319 | 13.0941 | 5.4823 | 269.8511 | 269.2531 | 194.7352 | 257.1277 | 165.9676
2012-10-30 03:50 | 25.9079 | 289.022 | 213.6172 | 277.5283 | 184.278 | 80.1995 | 2.0191 | 51.7791 | 104.5488 | 25.6442 | 3.1987 | 51.6261 | 165.1368 | 16.1397 | 13.0806 | 5.529 | 269.6856 | 275.9414 | 200.5366 | 264.4477 | 171.1974
2012-10-30 03:50 | 25.9101 | 293.5983 | 217.5847 | 282.5228 | 187.8809 | 80.6826 | 2.0183 | 51.7772 | 104.5019 | 25.6538 | 3.1975 | 51.6242 | 165.0662 | 16.1454 | 13.0665 | 5.5623 | 269.5681 | 280.5318 | 204.5182 | 269.4563 | 174.8144
2012-10-30 03:50 | 25.9131 | 294.9282 | 218.7098 | 283.9548 | 189.089 | 80.8254 | 2.018 | 51.7764 | 104.4872 | 25.6567 | 3.1971 | 51.6234 | 165.0432 | 16.1471 | 13.041 | 5.5721 | 269.5304 | 281.8872 | 205.6688 | 270.9138 | 176.048
2012-10-30 03:50 | 25.9168 | 296.4042 | 219.9734 | 285.433 | 190.302 | 80.9612 | 2.0178 | 51.7755 | 104.4722 | 25.6595 | 3.1967 | 51.6226 | 165.0199 | 16.1489 | 13.0855 | 5.5815 | 269.4921 | 283.3187 | 206.8879 | 272.3475 | 177.2165
2012-10-30 03:50 | 25.9266 | 297.8241 | 221.235 | 286.9975 | 191.4351 | 81.108 | 2.0175 | 51.7747 | 104.4573 | 25.6624 | 3.1962 | 51.6218 | 164.995 | 16.1508 | 13.0505 | 5.5916 | 269.4523 | 284.7736 | 208.1845 | 273.947 | 178.3846
2012-10-30 03:50 | 25.9281 | 299.2043 | 222.3386 | 288.4795 | 192.5287 | 81.2501 | 2.0173 | 51.7735 | 104.4401 | 25.6654 | 3.1958 | 51.6205 | 164.9695 | 16.1525 | 13.0642 | 5.6014 | 269.4096 | 286.1401 | 209.2744 | 275.4153 | 179.4645
2012-10-30 03:50 | 25.9325 | 300.6327 | 223.4803 | 289.8903 | 193.7042 | 81.3981 | 2.017 | 51.7726 | 104.4257 | 25.668 | 3.1954 | 51.6197 | 164.9478 | 16.1541 | 13.0686 | 5.6116 | 269.3736 | 287.5641 | 210.4117 | 276.8217 | 180.6356
2012-10-30 03:50 | 25.9311 | 301.8514 | 224.7185 | 291.364 | 194.6893 | 81.5379 | 2.0167 | 51.7717 | 104.4103 | 25.6709 | 3.195 | 51.6187 | 164.9231 | 16.156 | 13.0224 | 5.6212 | 269.3334 | 288.829 | 211.6961 | 278.3416 | 181.6669
2012-10-30 03:50 | 25.9408 | 303.2341 | 226.0344 | 292.9672 | 195.7998 | 81.6714 | 2.0165 | 51.7709 | 104.3953 | 25.6738 | 3.1946 | 51.6181 | 164.8991 | 16.1579 | 13.1095 | 5.6304 | 269.2944 | 290.1246 | 212.9249 | 279.8577 | 182.6903
2012-10-30 03:50 | 25.9438 | 304.6632 | 227.3158 | 294.432 | 196.8606 | 81.8119 | 2.0162 | 51.7704 | 104.3821 | 25.6766 | 3.1943 | 51.6175 | 164.8793 | 16.1595 | 13.1084 | 5.6401 | 269.2614 | 291.5548 | 214.2074 | 281.3236 | 183.7522
2012-10-30 03:50 | 25.9483 | 305.7291 | 228.2406 | 295.69 | 197.6753 | 81.9165 | 2.0161 | 51.7698 | 104.3708 | 25.6788 | 3.1939 | 51.6171 | 164.8602 | 16.1611 | 13.1455 | 5.6473 | 269.231 | 292.5836 | 215.0951 | 282.5445 | 184.5298
2012-10-30 03:51 | 25.9505 | 306.846 | 229.222 | 296.949 | 198.5861 | 82.0321 | 2.0158 | 51.7685 | 104.3548 | 25.6814 | 3.1935 | 51.6157 | 164.8362 | 16.1626 | 13.0947 | 5.6553 | 269.1911 | 293.7513 | 216.1273 | 283.8543 | 185.4914
2012-10-30 03:51 | 25.9513 | 308.1871 | 230.363 | 298.3558 | 199.657 | 82.1783 | 2.0155 | 51.7673 | 104.3387 | 25.6842 | 3.1931 | 51.6145 | 164.8097 | 16.1644 | 13.1077 | 5.6654 | 269.1484 | 295.0794 | 217.2553 | 285.2481 | 186.5493
2012-10-30 03:51 | 25.955 | 309.542 | 231.538 | 299.6995 | 200.6719 | 82.3089 | 2.0153 | 51.7664 | 104.3238 | 25.6869 | 3.1927 | 51.6136 | 164.7858 | 16.1662 | 13.14 | 5.6744 | 269.1096 | 296.402 | 218.398 | 286.5595 | 187.5319
2012-10-30 03:51 | 25.9572 | 310.8145 | 232.6524 | 301.1371 | 201.7327 | 82.4603 | 2.0151 | 51.7654 | 104.3109 | 25.6891 | 3.1923 | 51.6127 | 164.764 | 16.1678 | 13.0974 | 5.6848 | 269.0749 | 297.7171 | 219.555 | 288.0397 | 188.6353
2012-10-30 03:51 | 25.9602 | 311.9841 | 233.6416 | 302.4674 | 202.8062 | 82.5655 | 2.0148 | 51.7647 | 104.2973 | 25.6918 | 3.1919 | 51.6119 | 164.7407 | 16.1696 | 13.1126 | 5.6921 | 269.038 | 298.8715 | 220.529 | 289.3548 | 189.6936
2012-10-30 03:51 | 25.9737 | 313.3129 | 234.8571 | 303.8115 | 203.7565 | 82.7094 | 2.0146 | 51.7645 | 104.2853 | 25.6946 | 3.1916 | 51.6117 | 164.7219 | 16.1713 | 13.1016 | 5.702 | 269.0072 | 300.2113 | 221.7555 | 290.7099 | 190.6549
2012-10-30 03:51 | 25.9745 | 314.6553 | 236.04 | 305.2266 | 204.8202 | 82.8556 | [8 empty] | 13.1147 | 5.7121 | 0 | 301.5406 | 222.9253 | 292.1119 | 191.7055
2012-10-30 03:51
2012-10-30 03:54
2012-10-30 03:56
2012-10-30 03:58
2012-10-30 04:01
2012-10-30 04:03
2012-10-30 04:05 | [5 empty] | 51.7528 | 0.032695 | 0.70741 | 0.023129 | 21.6367 | 0.040727 | 0.76041 | 0.030969 | 18.6708 | -11.8386 | 3.5678 | 0.054098 | 11.8386 | 11.8386 | 11.8386 | 11.8386
2012-10-30 04:07 | 24.2754 | 56.2045 | 49.3993 | 56.7503 | 47.051 | 51.5002 | 0.0227 | 0.71651 | 0.016265 | 31.5644 | 0.038134 | 0.71883 | 0.027412 | 18.85 | 12.4125 | 3.5504 | 0.043677 | 43.792 | 36.9868 | 44.3378 | 34.6385
2012-10-30 04:07 | 24.2725 | 56.008 | 49.1899 | 56.5771 | 46.6732 | 51.4881 | 0.022694 | 0.71607 | 0.01625 | 31.5531 | 0.038102 | 0.71835 | 0.02737 | 18.8533 | 12.438 | 3.5496 | 0.04362 | 43.57 | 36.7519 | 44.1391 | 34.2352
2012-10-30 04:07 | 24.2784 | 55.828 | 49.0788 | 56.4166 | 46.5736 | 51.4725 | 0.022688 | 0.71563 | 0.016236 | 31.5423 | 0.038087 | 0.71777 | 0.027338 | 18.8455 | 12.4642 | 3.5485 | 0.043574 | 43.3638 | 36.6146 | 43.9524 | 34.1094
2012-10-30 04:07 | 24.2897 | 55.6689 | 48.9381 | 56.3311 | 46.6044 | 51.4459 | 0.029479 | 0.88124 | 0.025978 | 29.8937 | 0.046146 | 0.84777 | 0.039121 | 18.3716 | 12.4552 | 3.5467 | 0.065099 | 43.2137 | 36.4829 | 43.8759 | 34.1492
2012-10-30 04:07 | 24.2807 | 55.4894 | 48.8005 | 56.1827 | 46.5134 | 51.4459 | 0.076875 | 2.0341 | 0.15637 | 26.4593 | 0.14606 | 2.4549 | 0.35857 | 16.8076 | 12.4584 | 3.5467 | 0.51494 | 43.031 | 36.3421 | 43.7243 | 34.055
2012-10-30 04:07 | 24.2814 | 55.3933 | 48.8557 | 56.0789 | 46.4242 | 51.903 | 0.91144 | 22.4474 | 20.4596 | 24.6284 | 1.3937 | 22.496 | 31.3535 | 16.1408 | 12.4469 | 3.5782 | 51.813 | 42.9464 | 36.4088 | 43.632 | 33.9773
2012-10-30 04:07 | 24.2874 | 56.0888 | 49.6061 | 55.961 | 46.3208 | 52.7722 | 0.90457 | 22.3161 | 20.1864 | 24.6705 | 1.3795 | 22.2625 | 30.7118 | 16.1377 | 12.4529 | 3.6381 | 50.8982 | 43.6359 | 37.1532 | 43.5081 | 33.8679
2012-10-30 04:07 | 24.2919 | 57.0299 | 50.226 | 55.9848 | 46.3917 | 53.8934 | 1.5252 | 37.9263 | 57.846 | 24.8661 | 2.3817 | 38.301 | 91.2232 | 16.081 | 12.4696 | 3.7154 | 149.0692 | 44.5603 | 37.7564 | 43.5152 | 33.9221
2012-10-30 04:07 | 24.2941 | 59.6636 | 52.7834 | 56.0993 | 46.4252 | 55.1406 | 1.3323 | 33.0685 | 44.0579 | 24.8203 | 2.0495 | 32.9705 | 67.5725 | 16.0873 | 12.4759 | 3.8014 | 111.6303 | 47.1877 | 40.3075 | 43.6234 | 33.9493
2012-10-30 04:07 | 24.2926 | 62.2293 | 53.5196 | 56.4966 | 46.8534 | 54.8516 | 1.3482 | 33.459 | 45.1102 | 24.8171 | 2.0748 | 33.3746 | 69.245 | 16.0859 | 12.4378 | 3.7815 | 114.3552 | 49.7915 | 41.0818 | 44.0588 | 34.4156
2012-10-30 04:07 | 24.3016 | 64.3732 | 54.359 | 57.0861 | 47.534 | 55.0296 | 1.3459 | 33.4095 | 44.9671 | 24.8225 | 2.0718 | 33.3258 | 69.0437 | 16.0856 | 12.463 | 3.7937 | 114.0108 | 51.9102 | 41.896 | 44.6231 | 35.071
2012-10-30 04:07 | 24.3046 | 66.1925 | 55.0987 | 57.7775 | 48.0482 | 55.2041 | 1.3451 | 33.3947 | 44.9185 | 24.8273 | 2.0709 | 33.3114 | 68.9851 | 16.0853 | 12.4904 | 3.8058 | 113.9037 | 53.7021 | 42.6083 | 45.2871 | 35.5578
2012-10-30 04:07 | 24.3076 | 67.9593 | 55.8727 | 58.5457 | 48.6363 | 55.3307 | 1.345 | 33.3982 | 44.9196 | 24.8319 | 2.0712 | 33.315 | 69.0009 | 16.0852 | 12.526 | 3.8145 | 113.9205 | 55.4333 | 43.3467 | 46.0197 | 36.1103
2012-10-30 04:07 | 24.3076 | 69.7013 | 56.7167 | 59.4185 | 49.1509 | 55.4798 | 1.3448 | 33.3996 | 44.9152 | 24.8364 | 2.0713 | 33.3158 | 69.006 | 16.0847 | 12.5056 | 3.8248 | 113.9212 | 57.1957 | 44.2111 | 46.9129 | 36.6453
2012-10-30 04:07 | 24.3121 | 71.5143 | 57.5491 | 60.349 | 49.8061 | 55.652 | 1.344 | 33.3865 | 44.8712 | 24.8413 | 2.0705 | 33.3028 | 68.9527 | 16.0846 | 12.5223 | 3.8366 | 113.8239 | 58.992 | 45.0268 | 47.8267 | 37.2838
2012-10-30 04:07 | 24.3188 | 73.4225 | 58.5105 | 61.6124 | 50.6383 | 55.8104 | 1.3438 | 33.3878 | 44.8651 | 24.8466 | 2.0706 | 33.3042 | 68.9604 | 16.0841 | 12.5453 | 3.8476 | 113.8255 | 60.8772 | 45.9652 | 49.0671 | 38.093
2012-10-30 04:07 | 24.3218 | 75.172 | 59.479 | 62.9822 | 51.4701 | 55.9826 | 1.3437 | 33.3938 | 44.8723 | 24.8516 | 2.071 | 33.3099 | 68.9852 | 16.0839 | 12.5198 | 3.8594 | 113.8575 | 62.6522 | 46.9592 | 50.4624 | 38.9503
2012-10-30 04:07 | 24.3256 | 76.8747 | 60.4049 | 64.2704 | 52.2126 | 56.1646 | 1.3435 | 33.3942 | 44.8655 | 24.856 | 2.0711 | 33.3104 | 68.9887 | 16.0835 | 12.544 | 3.872 | 113.8542 | 64.3307 | 47.8609 | 51.7264 | 39.6686
2012-10-30 04:07 | 24.3293 | 78.4271 | 61.2337 | 65.5456 | 52.8535 | 56.3039 | 1.3433 | 33.3958 | 44.8619 | 24.8603 | 2.0712 | 33.3118 | 68.9954 | 16.0834 | 12.5558 | 3.8816 | 113.8572 | 65.8713 | 48.6779 | 52.9898 | 40.2977
2012-10-30 04:08 | 24.3323 | 80.0725 | 62.1689 | 66.9912 | 53.6641 | 56.4675 | 1.3432 | 33.3999 | 44.8643 | 24.8651 | 2.0715 | 33.3157 | 69.013 | 16.083 | 12.5873 | 3.8929 | 113.8773 | 67.4852 | 49.5816 | 54.4039 | 41.0768
2012-10-30 04:08 | 24.3443 | 81.8165 | 63.1855 | 68.6012 | 54.4987 | 56.6397 | 1.3431 | 33.4028 | 44.8647 | 24.8691 | 2.0717 | 33.3189 | 69.0277 | 16.0827 | 12.5749 | 3.9047 | 113.8923 | 69.2416 | 50.6106 | 56.0263 | 41.9238
2012-10-30 04:08 | 24.3353 | 83.3674 | 64.1263 | 70.1154 | 55.2922 | 56.8894 | 1.4719 | 36.6579 | 53.9567 | 24.9052 | 2.2735 | 36.5474 | 83.0915 | 16.0752 | 12.5903 | 3.922 | 137.0482 | 70.7771 | 51.536 | 57.5251 | 42.7019
2012-10-30 04:08 | 24.3472 | 85.3096 | 65.4636 | 71.5151 | 56.0285 | 57.7285 | 1.5574 | 38.8756 | 60.5459 | 24.9615 | 2.4129 | 38.7756 | 93.5605 | 16.0703 | 12.6143 | 3.9798 | 154.1064 | 72.6953 | 52.8493 | 58.9008 | 43.4142
2012-10-30 04:08 | 24.3435 | 87.9166 | 67.2018 | 73.2488 | 56.9074 | 58.0157 | 1.5569 | 38.8757 | 60.5246 | 24.9703 | 2.4129 | 38.7759 | 93.5631 | 16.0701 | 12.5903 | 3.9996 | 154.0877 | 75.3263 | 54.6115 | 60.6585 | 44.3171
2012-10-30 04:08 | 24.357 | 90.3229 | 68.6828 | 75.0357 | 57.9381 | 58.2527 | 1.5564 | 38.8751 | 60.5065 | 24.977 | 2.4129 | 38.7749 | 93.5593 | 16.07 | 12.616 | 4.0159 | 154.0658 | 77.7069 | 56.0668 | 62.4197 | 45.3221
2012-10-30 04:08 | 24.3637 | 92.5507 | 69.9334 | 76.6958 | 58.849 | 58.4492 | 1.5562 | 38.8765 | 60.4978 | 24.9825 | 2.413 | 38.7764 | 93.567 | 16.0699 | 12.6105 | 4.0295 | 154.0648 | 79.9402 | 57.3229 | 64.0853 | 46.2385
2012-10-30 04:08 | 24.3742 | 94.8986 | 71.3166 | 78.4986 | 59.9057 | 58.6682 | 1.5557 | 38.8742 | 60.4776 | 24.9878 | 2.4129 | 38.7738 | 93.5553 | 16.0697 | 12.6292 | 4.0446 | 154.033 | 82.2694 | 58.6874 | 65.8694 | 47.2765
2012-10-30 04:08 | 24.3764 | 96.7234 | 72.3655 | 80.1888 | 60.9917 | 58.8682 | 1.5555 | 38.8761 | 60.4734 | 24.9921 | 2.4129 | 38.7757 | 93.5636 | 16.0698 | 12.6557 | 4.0584 | 154.037 | 84.0677 | 59.7098 | 67.5331 | 48.336
2012-10-30 04:08 | 24.3816 | 98.6416 | 73.5426 | 82.1082 | 62.0873 | 59.0404 | 1.5553 | 38.8776 | 60.4654 | 24.9973 | 2.4131 | 38.7773 | 93.5732 | 16.0696 | 12.6691 | 4.0702 | 154.0386 | 85.9725 | 60.8735 | 69.4391 | 49.4182
2012-10-30 04:08 | 24.3914 | 100.9083 | 74.9253 | 84.2125 | 63.2287 | 59.2675 | 1.5544 | 38.8648 | 60.4126 | 25.0026 | 2.4123 | 38.7643 | 93.5104 | 16.0695 | 12.6667 | 4.0859 | 153.923 | 88.2416 | 62.2586 | 71.5458 | 50.562
2012-10-30 04:08 | 24.3899 | 102.8955 | 76.2041 | 86.1596 | 64.2655 | 59.501 | 1.5543 | 38.8706 | 60.4179 | 25.0079 | 2.4127 | 38.7701 | 93.5387 | 16.0695 | 12.6652 | 4.102 | 153.9566 | 90.2303 | 63.5389 | 73.4944 | 51.6003
2012-10-30 04:08 | 24.3936 | 104.9525 | 77.3462 | 87.944 | 65.2451 | 59.6935 | 1.554 | 38.87 | 60.4059 | 25.0121 | 2.4126 | 38.7692 | 93.5355 | 16.0693 | 12.6689 | 4.1153 | 153.9414 | 92.2836 | 64.6773 | 75.2751 | 52.5762
2012-10-30 04:08 | 24.4063 | 106.7693 | 78.4585 | 89.7969 | 66.2866 | 59.8668 | 1.5538 | 38.8696 | 60.3936 | 25.0166 | 2.4126 | 38.7688 | 93.5335 | 16.0693 | 12.6734 | 4.1272 | 153.9272 | 94.0959 | 65.7851 | 77.1235 | 53.6132
2012-10-30 04:08 | 24.4108 | 108.692 | 79.8887 | 91.9622 | 67.4684 | 60.0807 | 1.5536 | 38.8733 | 60.3926 | 25.0218 | 2.4128 | 38.7724 | 93.5518 | 16.0691 | 12.6617 | 4.142 | 153.9444 | 96.0303 | 67.227 | 79.3005 | 54.8067
2012-10-30 04:08 | 24.4205 | 110.5979 | 81.0952 | 93.9881 | 68.5931 | 60.2864 | 1.5538 | 38.8847 | 60.4179 | 25.026 | 2.4136 | 38.7838 | 93.6071 | 16.0691 | 12.647 | 4.1561 | 154.025 | 97.9509 | 68.4482 | 81.3411 | 55.9461
2012-10-30 04:08 | 24.4273 | 112.2633 | 82.2716 | 95.8215 | 69.6297 | 60.4505 | 1.5535 | 38.8856 | 60.4107 | 25.0302 | 2.4136 | 38.7847 | 93.6119 | 16.069 | 12.6904 | 4.1675 | 154.0226 | 99.5729 | 69.5812 | 83.1311 | 56.9393
2012-10-30 04:08 | 24.4325 | 114.0528 | 83.4342 | 97.6604 | 70.6601 | 60.595 | 1.5533 | 38.885 | 60.399 | 25.0342 | 2.4136 | 38.7839 | 93.608 | 16.069 | 12.7118 | 4.1774 | 154.007 | 101.341 | 70.7224 | 84.9486 | 57.9483
2012-10-30 04:08 | 24.4385 | 115.6736 | 84.5516 | 99.4697 | 71.6484 | 60.7881 | 1.5531 | 38.8866 | 60.3936 | 25.0385 | 2.4137 | 38.7855 | 93.6161 | 16.069 | 12.7382 | 4.1907 | 154.0097 | 102.9354 | 71.8134 | 86.7315 | 58.9102
2012-10-30 04:08 | 24.4452 | 117.5008 | 85.7829 | 101.4196 | 72.755 | 60.995 | 1.5529 | 38.8885 | 60.3887 | 25.043 | 2.4138 | 38.7874 | 93.6259 | 16.0689 | 12.6961 | 4.205 | 154.0145 | 104.8047 | 73.0868 | 88.7235 | 60.0589
2012-10-30 04:08 | 24.4467 | 119.2861 | 86.9478 | 103.2623 | 73.703 | 61.1967 | 1.5527 | 38.8903 | 60.3841 | 25.0473 | 2.4139 | 38.7889 | 93.633 | 16.0689 | 12.7301 | 4.2189 | 154.0172 | 106.556 | 74.2177 | 90.5322 | 60.9729
2012-10-30 04:08 | 24.4549 | 120.9122 | 88.0621 | 104.8965 | 74.7334 | 61.3261 | 1.5525 | 38.8911 | 60.3769 | 25.0513 | 2.414 | 38.7899 | 93.6376 | 16.0689 | 12.7099 | 4.2278 | 154.0145 | 108.2023 | 75.3522 | 92.1866 | 62.0235
2012-10-30 04:08 | 24.4632 | 122.4977 | 89.2177 | 106.6787 | 75.6869 | 61.4908 | 1.5523 | 38.8931 | 60.3739 | 25.0551 | 2.4141 | 38.7917 | 93.6461 | 16.0689 | 12.7385 | 4.2392 | 154.02 | 109.7592 | 76.4792 | 93.9402 | 62.9484
2012-10-30 04:08 | 24.4677 | 124.2023 | 90.4033 | 108.6051 | 76.7008 | 61.6665 | 1.5521 | 38.8939 | 60.3663 | 25.0593 | 2.4141 | 38.7924 | 93.6502 | 16.0688 | 12.6901 | 4.2513 | 154.0165 | 111.5122 | 77.7132 | 95.915 | 64.0107
2012-10-30 04:08 | 24.4789 | 125.9035 | 91.6822 | 110.6866 | 77.8767 | 61.8572 | 1.5519 | 38.8952 | 60.3604 | 25.0634 | 2.4142 | 38.7936 | 93.6553 | 16.069 | 12.7095 | 4.2644 | 154.0157 | 113.194 | 78.9727 | 97.9771 | 65.1672
2012-10-30 04:08 | 24.4864 | 127.6062 | 92.9042 | 112.5913 | 79.09 | 62.03 | 1.5515 | 38.8937 | 60.3448 | 25.0679 | 2.4141 | 38.7919 | 93.6472 | 16.069 | 12.7332 | 4.2763 | 153.992 | 114.873 | 80.171 | 99.8581 | 66.3568
2012-10-30 04:09 | 24.4968 | 129.2675 | 94.0607 | 114.3218 | 80.1192 | 62.1965 | 1.5513 | 38.8934 | 60.3355 | 25.0715 | 2.4141 | 38.7917 | 93.6464 | 16.0689 | 12.7274 | 4.2878 | 153.9819 | 116.5401 | 81.3333 | 101.5944 | 67.3918
2012-10-30 04:09 | 24.5013 | 130.8937 | 95.298 | 116.3621 | 81.233 | 62.3802 | 1.551 | 38.8931 | 60.325 | 25.0754 | 2.4141 | 38.7913 | 93.6447 | 16.0689 | 12.6749 | 4.3005 | 153.9697 | 118.2188 | 82.6231 | 103.6872 | 68.5581
2012-10-30 04:09 | 24.4998 | 132.3855 | 96.412 | 118.1967 | 82.3782 | 62.5403 | 1.5509 | 38.8952 | 60.3221 | 25.0793 | 2.4142 | 38.7935 | 93.6545 | 16.069 | 12.6775 | 4.3115 | 153.9766 | 119.708 | 83.7345 | 105.5192 | 69.7007
2012-10-30 04:09 | 24.5155 | 133.9645 | 97.5922 | 119.9549 | 83.4907 | 62.6946 | 1.5507 | 38.8978 | 60.3201 | 25.0835 | 2.4143 | 38.7961 | 93.6671 | 16.069 | 12.7339 | 4.3222 | 153.9872 | 121.2306 | 84.8583 | 107.221 | 70.7568
2012-10-30 04:09 | 24.5215 | 135.7188 | 98.9176 | 121.8908 | 84.5891 | 62.8796 | 1.5506 | 38.8998 | 60.3176 | 25.0871 | 2.4145 | 38.7978 | 93.6755 | 16.0689 | 12.7033 | 4.3349 | 153.9931 | 123.0155 | 86.2143 | 109.1875 | 71.8858
2012-10-30 04:09 | 24.5335 | 137.1361 | 99.9881 | 123.4959 | 85.5717 | 63.0639 | 1.5504 | 38.8999 | 60.3087 | 25.0909 | 2.4145 | 38.7979 | 93.676 | 16.069 | 12.7519 | 4.3476 | 153.9847 | 124.3842 | 87.2362 | 110.744 | 72.8198
2012-10-30 04:09 | 24.535 | 138.5364 | 101.0708 | 125.1182 | 86.5473 | 63.2263 | 1.5502 | 38.902 | 60.3073 | 25.0942 | 2.4146 | 38.7999 | 93.6854 | 16.069 | 12.7615 | 4.3588 | 153.9928 | 125.7749 | 88.3093 | 112.3567 | 73.7858
2012-10-30 04:09 | 24.5425 | 139.9591 | 102.1218 | 126.7669 | 87.5361 | 63.3691 | 1.55 | 38.902 | 60.2979 | 25.0982 | 2.4146 | 38.7999 | 93.6856 | 16.069 | 12.7527 | 4.3687 | 153.9834 | 127.2064 | 89.3691 | 114.0142 | 74.7834
2012-10-30 04:09 | 24.5552 | 141.4535 | 103.2952 | 128.3954 | 88.579 | 63.5286 | 1.5497 | 38.9018 | 60.2876 | 25.1022 | 2.4145 | 38.7996 | 93.6828 | 16.0692 | 12.7614 | 4.3797 | 153.9703 | 128.6921 | 90.5338 | 115.634 | 75.8176
2012-10-30 04:09 | 24.5671 | 142.9522 | 104.4188 | 130.1199 | 89.5526 | 63.6962 | 1.5495 | 38.9008 | 60.2774 | 25.1052 | 2.4145 | 38.7986 | 93.6782 | 16.0692 | 12.7448 | 4.3912 | 153.9556 | 130.2074 | 91.674 | 117.3751 | 76.8078
2012-10-30 04:09 | 24.5642 | 144.345 | 105.4672 | 131.6324 | 90.4544 | 63.8597 | 1.5493 | 38.9011 | 60.2704 | 25.1084 | 2.4145 | 38.7989 | 93.6788 | 16.0693 | 12.7785 | 4.4025 | 153.9492 | 131.5665 | 92.6887 | 118.8539 | 77.6759
2012-10-30 04:09 | 24.5701 | 145.8787 | 106.5699 | 133.2774 | 91.4934 | 64.0164 | 1.5491 | 38.9008 | 60.2609 | 25.112 | 2.4145 | 38.7985 | 93.6772 | 16.0692 | 12.7559 | 4.4133 | 153.9381 | 133.1228 | 93.814 | 120.5215 | 78.7375
2012-10-30 04:09 | 24.5806 | 147.322 | 107.6775 | 134.7333 | 92.4461 | 64.1527 | 1.5489 | 38.9013 | 60.2537 | 25.1157 | 2.4145 | 38.7988 | 93.679 | 16.0692 | 12.7258 | 4.4227 | 153.9327 | 134.5962 | 94.9517 | 122.0075 | 79.7203
2012-10-30 04:09 | 24.5911 | 148.6894 | 108.8761 | 136.2248 | 93.4401 | 64.2978 | 1.5487 | 38.9025 | 60.2499 | 25.1188 | 2.4145 | 38.8001 | 93.6846 | 16.0693 | 12.7647 | 4.4327 | 153.9345 | 135.9247 | 96.1114 | 123.4601 | 80.6754
2012-10-30 04:09 | 24.5956 | 149.9824 | 109.8835 | 137.8112 | 94.496 | 64.4735 | 1.5486 | 38.9028 | 60.2434 | 25.1219 | 2.4145 | 38.8002 | 93.6846 | 16.0694 | 12.7083 | 4.4448 | 153.928 | 137.2741 | 97.1752 | 125.1029 | 81.7877
2012-10-30 04:09 | 24.6083 | 151.2421 | 110.8539 | 139.3766 | 95.3483 | 64.6272 | 1.5484 | 38.9027 | 60.235 | 25.1252 | 2.4145 | 38.8 | 93.6834 | 16.0695 | 12.725 | 4.4554 | 153.9183 | 138.5171 | 98.1289 | 126.6516 | 82.6233
2012-10-30 04:09 | 24.6023 | 152.5265 | 111.8399 | 140.8056 | 96.2498 | 64.7619 | 1.5482 | 38.9032 | 60.2294 | 25.1282 | 2.4146 | 38.8007 | 93.6863 | 16.0695 | 12.7272 | 4.4647 | 153.9158 | 139.7993 | 99.1127 | 128.0784 | 83.5226
2012-10-30 04:09 | 24.6143 | 153.8143 | 112.8215 | 142.1117 | 97.1768 | 64.8925 | 1.548 | 38.9037 | 60.2233 | 25.1315 | 2.4146 | 38.8011 | 93.689 | 16.0694 | 12.727 | 4.4737 | 153.9123 | 141.0873 | 100.0945 | 129.3847 | 84.4498
2012-10-30 04:09 | 24.6262 | 155.3081 | 113.9813 | 143.5917 | 98.1944 | 65.0329 | 1.5478 | 38.9047 | 60.2171 | 25.1352 | 2.4146 | 38.802 | 93.6922 | 16.0696 | 12.7185 | 4.4834 | 153.9094 | 142.5896 | 101.2628 | 130.8732 | 85.4759
2012-10-30 04:09 | 24.63 | 156.7525 | 115.1751 | 145.2072 | 99.2 | 65.1809 | 1.5477 | 38.906 | 60.2133 | 25.1386 | 2.4147 | 38.8032 | 93.6976 | 16.0697 | 12.7345 | 4.4936 | 153.9109 | 144.018 | 102.4406 | 132.4727 | 86.4655
2012-10-30 04:09 | 24.6382 | 157.9775 | 116.1198 | 146.5899 | 100.1987 | 65.3225 | 1.5475 | 38.907 | 60.2094 | 25.1415 | 2.4147 | 38.8041 | 93.7018 | 16.0697 | 12.7671 | 4.5033 | 153.9112 | 145.2104 | 103.3527 | 133.8228 | 87.4316
2012-10-30 04:09 | 24.6434 | 159.2624 | 117.0314 | 147.9205 | 101.0999 | 65.4265 | 1.5474 | 38.9092 | 60.2093 | 25.1444 | 2.4149 | 38.8064 | 93.713 | 16.0696 | 12.7642 | 4.5105 | 153.9223 | 146.4982 | 104.2672 | 135.1563 | 88.3357
2012-10-30 04:09 | 24.6472 | 160.3954 | 117.9912 | 149.3781 | 101.9959 | 65.5577 | 1.5473 | 38.9104 | 60.2061 | 25.1472 | 2.4149 | 38.8074 | 93.7172 | 16.0698 | 12.7029 | 4.5195 | 153.9233 | 147.6925 | 105.2883 | 136.6752 | 89.293
2012-10-30 04:09 | 24.6569 | 161.6476 | 119.0408 | 150.8196 | 102.9281 | 65.6952 | 1.5471 | 38.9091 | 60.1946 | 25.1504 | 2.4149 | 38.8062 | 93.7112 | 16.0698 | 12.7248 | 4.529 | 153.9058 | 148.9228 | 106.316 | 138.0948 | 90.2033
2012-10-30 04:09 | 24.6554 | 162.9478 | 120.1062 | 152.0579 | 103.8378 | 65.8426 | 1.5469 | 38.9092 | 60.1882 | 25.1532 | 2.4149 | 38.8063 | 93.7114 | 16.0698 | 12.7355 | 4.5392 | 153.8996 | 150.2123 | 107.3707 | 139.3224 | 91.1023
2012-10-30 04:09 | 24.6749 | 164.1336 | 121.0145 | 153.4379 | 104.7421 | 65.9709 | 1.5467 | 38.9099 | 60.1827 | 25.1564 | 2.4149 | 38.8067 | 93.7131 | 16.0699 | 12.7591 | 4.548 | 153.8958 | 151.3745 | 108.2554 | 140.6788 | 91.983
2012-10-30 04:09 | 24.6786 | 165.3628 | 121.9681 | 154.7955 | 105.6005 | 66.1194 | 1.5465 | 38.9098 | 60.1753 | 25.1594 | 2.4149 | 38.8067 | 93.713 | 16.0699 | 12.7425 | 4.5583 | 153.8882 | 152.6203 | 109.2256 | 142.053 | 92.858
2012-10-30 04:10 | 24.6921 | 166.6725 | 123.0726 | 156.2839 | 106.6163 | 66.2558 | 1.5463 | 38.9097 | 60.1666 | 25.1629 | 2.4148 | 38.8065 | 93.7112 | 16.0701 | 12.7194 | 4.5677 | 153.8777 | 153.9531 | 110.3532 | 143.5645 | 93.8969
2012-10-30 04:10 | 24.701 | 168.0409 | 124.1579 | 157.6604 | 107.6012 | 66.3859 | 1.5461 | 38.9097 | 60.1596 | 25.1659 | 2.4148 | 38.8066 | 93.7111 | 16.0702 | 12.7201 | 4.5766 | 153.8707 | 155.3208 | 111.4378 | 144.9403 | 94.8811
2012-10-30 04:10 | 24.7123 | 169.1747 | 125.0672 | 158.8542 | 108.3956 | 66.5014 | 1.546 | 38.9111 | 60.1567 | 25.1689 | 2.4149 | 38.8079 | 93.7172 | 16.0702 | 12.7396 | 4.5846 | 153.874 | 156.4351 | 112.3276 | 146.1146 | 95.656
2012-10-30 04:10 | 24.7182 | 170.3192 | 125.9714 | 160.1249 | 109.128 | 66.643 | 1.5459 | 38.9117 | 60.1519 | 25.1716 | 2.4149 | 38.8083 | 93.7191 | 16.0702 | 12.7211 | 4.5944 | 153.871 | 157.5981 | 113.2503 | 147.4038 | 96.4069
2012-10-30 04:10 | 24.7242 | 171.4372 | 126.8953 | 161.2955 | 109.9402 | 66.765 | 1.5457 | 38.913 | 60.1494 | 25.1743 | 2.415 | 38.8095 | 93.7249 | 16.0702 | 12.7271 | 4.6028 | 153.8744 | 158.7101 | 114.1682 | 148.5684 | 97.2131
2012-10-30 04:10 | 24.7362 | 172.6007 | 127.8331 | 162.7604 | 111.0386 | 66.9077 | 1.5456 | 38.9133 | 60.143 | 25.1774 | 2.415 | 38.8099 | 93.7264 | 16.0703 | 12.7065 | 4.6126 | 153.8694 | 159.8942 | 115.1266 | 150.0539 | 98.3321
2012-10-30 04:10 | 24.7519 | 173.999 | 128.9009 | 164.1251 | 111.9933 | 67.0389 | 1.5454 | 38.9135 | 60.1362 | 25.1805 | 2.415 | 38.81 | 93.7256 | 16.0705 | 12.7222 | 4.6217 | 153.8618 | 161.2768 | 116.1787 | 151.4029 | 99.2711
2012-10-30 04:10 | 24.7586 | 175.1624 | 129.8611 | 165.4428 | 112.9052 | 67.1753 | 1.5452 | 38.9139 | 60.1306 | 25.1834 | 2.415 | 38.8103 | 93.7269 | 16.0705 | 12.7208 | 4.6311 | 153.8576 | 162.4416 | 117.1403 | 152.722 | 100.1844
2012-10-30 04:10 | 24.7601 | 176.3016 | 130.8013 | 166.577 | 113.7628 | 67.2961 | 1.5451 | 38.9144 | 60.1249 | 25.1864 | 2.415 | 38.8107 | 93.7293 | 16.0705 | 12.7467 | 4.6394 | 153.8542 | 163.5549 | 118.0546 | 153.8303 | 101.0161
2012-10-30 04:10 | 24.7818 | 178.7294 | 132.7706 | 169.2286 | 115.4748 | 67.5527 | 1.5442 | 38.9022 | 60.0739 | 25.1919 | 2.4143 | 38.7987 | 93.6704 | 16.0706 | 12.7277 | 4.6571 | 153.7443 | 166.0017 | 120.0429 | 156.5009 | 102.7471
2012-10-30 04:10 | 24.8207 | 183.6299 | 136.7186 | 174.6148 | 119.2273 | 68.0931 | 1.5438 | 38.91 | 60.0692 | 25.2041 | 2.4147 | 38.8062 | 93.7043 | 16.071 | 12.7382 | 4.6943 | 153.7734 | 170.8917 | 123.9804 | 161.8766 | 106.4891
2012-10-30 04:10 | 24.8521 | 187.1958 | 139.6818 | 178.4318 | 122.0542 | 68.4722 | 1.5435 | 38.9155 | 60.0658 | 25.2126 | 2.415 | 38.8116 | 93.7294 | 16.0711 | 12.7736 | 4.7205 | 153.7952 | 174.4222 | 126.9082 | 165.6582 | 109.2806
2012-10-30 04:10 | 24.8611 | 188.2058 | 140.4427 | 179.482 | 122.755 | 68.5901 | 1.5433 | 38.9145 | 60.0571 | 25.215 | 2.4149 | 38.8105 | 93.7236 | 16.0713 | 12.7786 | 4.7286 | 153.7807 | 175.4272 | 127.6641 | 166.7034 | 109.9764
2012-10-30 04:10 | 24.8701 | 189.1493 | 141.2807 | 180.5209 | 123.5206 | 68.6796 | 1.5432 | 38.9151 | 60.0525 | 25.2177 | 2.4149 | 38.8111 | 93.7261 | 16.0713 | 12.7469 | 4.7348 | 153.7785 | 176.4024 | 128.5338 | 167.774 | 110.7737
2012-10-30 04:10 | 24.8806 | 190.1848 | 142.04 | 181.5694 | 124.2424 | 68.8022 | 1.543 | 38.9156 | 60.0486 | 25.2199 | 2.415 | 38.8114 | 93.7285 | 16.0712 | 12.7818 | 4.7432 | 153.7771 | 177.403 | 129.2582 | 168.7876 | 111.4606
2012-10-30 04:10 | 24.8851 | 191.2736 | 142.9396 | 182.7339 | 125.0763 | 68.9264 | 1.5429 | 38.916 | 60.0429 | 25.2229 | 2.415 | 38.812 | 93.73 | 16.0714 | 12.77 | 4.7518 | 153.7729 | 178.5036 | 130.1696 | 169.9639 | 112.3063
2012-10-30 04:10 | 24.9045 | 192.5072 | 143.9085 | 183.9254 | 125.9065 | 69.0426 | 1.5428 | 38.9167 | 60.0394 | 25.2252 | 2.415 | 38.8124 | 93.7328 | 16.0712 | 12.8016 | 4.7598 | 153.7722 | 179.7056 | 131.1069 | 171.1238 | 113.1049
2012-10-30 04:10 | 24.9083 | 193.4739 | 144.6891 | 185.043 | 126.6299 | 69.1368 | 1.5425 | 38.9139 | 60.0246 | 25.2279 | 2.4148 | 38.8097 | 93.7189 | 16.0713 | 12.7932 | 4.7663 | 153.7435 | 180.6807 | 131.8959 | 172.2498 | 113.8367
2012-10-30 04:10 | 24.9135 | 194.4069 | 145.4946 | 186.0956 | 127.4008 | 69.2564 | 1.5424 | 38.9136 | 60.019 | 25.2298 | 2.4148 | 38.8094 | 93.718 | 16.0713 | 12.8025 | 4.7745 | 153.737 | 181.6044 | 132.6921 | 173.2931 | 114.5983
2012-10-30 04:10 | 24.9359 | 195.3572 | 146.2599 | 187.1069 | 128.1624 | 69.3518 | 1.5423 | 38.9146 | 60.0161 | 25.2323 | 2.4149 | 38.8103 | 93.7225 | 16.0712 | 12.8167 | 4.7811 | 153.7386 | 182.5405 | 133.4432 | 174.2902 | 115.3457
2012-10-30 04:11 | 24.9352 | 196.3633 | 146.987 | 188.1307 | 128.905 | 69.4518 | 1.5421 | 38.9139 | 60.009 | 25.2344 | 2.4148 | 38.8095 | 93.7184 | 16.0713 | 12.7916 | 4.788 | 153.7274 | 183.5717 | 134.1954 | 175.3391 | 116.1134
2012-10-30 04:11 | 24.9299 | 197.3688 | 147.8409 | 189.1816 | 129.7082 | 69.576 | 1.542 | 38.9143 | 60.0043 | 25.2369 | 2.4148 | 38.8099 | 93.7196 | 16.0714 | 12.827 | 4.7966 | 153.7238 | 184.5418 | 135.0139 | 176.3546 | 116.8812
2012-10-30 04:11 | 24.9449 | 198.3634 | 148.7502 | 190.3277 | 130.5205 | 69.6875 | 1.5417 | 38.913 | 59.9934 | 25.2399 | 2.4148 | 38.8085 | 93.7141 | 16.0712 | 12.7851 | 4.8043 | 153.7074 | 185.5783 | 135.9651 | 177.5426 | 117.7354
2012-10-30 04:11 | 24.9531 | 199.4339 | 149.6222 | 191.4634 | 131.231 | 69.791 | 1.5416 | 38.9118 | 59.9849 | 25.2418 | 2.4147 | 38.8074 | 93.7085 | 16.0713 | 12.8217 | 4.8114 | 153.6933 | 186.6122 | 136.8005 | 178.6417 | 118.4093
2012-10-30 04:11 | 24.9614 | 200.5125 | 150.491 | 192.5719 | 132.1442 | 69.9043 | 1.5414 | 38.912 | 59.9791 | 25.2445 | 2.4147 | 38.8075 | 93.709 | 16.0713 | 12.8422 | 4.8192 | 153.6881 | 187.6703 | 137.6488 | 179.7297 | 119.302
2012-10-30 04:11 | 24.9771 | 201.5594 | 151.3599 | 193.6528 | 132.8933 | 70.0129 | 1.5412 | 38.9099 | 59.9667 | 25.247 | 2.4146 | 38.8054 | 93.6978 | 16.0714 | 12.8091 | 4.8267 | 153.6646 | 188.7503 | 138.5508 | 180.8437 | 120.0842
2012-10-30 04:11 | 24.9748 | 202.4939 | 152.0991 | 194.5429 | 133.5177 | 70.1169 | 1.541 | 38.909 | 59.9582 | 25.2494 | 2.4145 | 38.8043 | 93.694 | 16.0712 | 12.8312 | 4.8339 | 153.6522 | 189.6627 | 139.2679 | 181.7117 | 120.6865
2012-10-30 04:11 | 24.9868 | 203.3561 | 152.7911 | 195.5929 | 134.1872 | 70.2262 | 1.5409 | 38.9096 | 59.9557 | 25.2512 | 2.4146 | 38.805 | 93.6976 | 16.0712 | 12.8188 | 4.8414 | 153.6533 | 190.5373 | 139.9723 | 182.7741 | 121.3684
2012-10-30 04:11 | 25.0018 | 204.2411 | 153.5406 | 196.5085 | 134.8867 | 70.3158 | 1.5408 | 38.9111 | 59.9549 | 25.2536 | 2.4147 | 38.8064 | 93.7041 | 16.0712 | 12.7891 | 4.8476 | 153.6591 | 191.452 | 140.7515 | 183.7194 | 122.0976
2012-10-30 04:11 | 24.9988 | 205.1791 | 154.311 | 197.4535 | 135.5647 | 70.4053 | 1.5407 | 38.911 | 59.9498 | 25.2556 | 2.4147 | 38.8064 | 93.704 | 16.0712 | 12.843 | 4.8537 | 153.6538 | 192.3361 | 141.468 | 184.6105 | 122.7217
2012-10-30 04:11 | 25.0175 | 206.1034 | 155.0995 | 198.5776 | 136.322 | 70.4972 | 1.5405 | 38.9107 | 59.943 | 25.258 | 2.4146 | 38.8059 | 93.7024 | 16.0711 | 12.8333 | 4.8601 | 153.6455 | 193.2701 | 142.2662 | 185.7443 | 123.4887
2012-10-30 04:11 | 25.022 | 206.9938 | 155.8239 | 199.5264 | 137.0463 | 70.6076 | 1.5404 | 38.9105 | 59.9374 | 25.2602 | 2.4146 | 38.8058 | 93.7021 | 16.071 | 12.7971 | 4.8677 | 153.6395 | 194.1967 | 143.0268 | 186.7293 | 124.2492
2012-10-30 04:11 | 25.0332 | 207.9815 | 156.5824 | 200.4899 | 137.5745 | 70.7215 | 1.5403 | 38.9115 | 59.9355 | 25.2622 | 2.4147 | 38.8067 | 93.7063 | 16.071 | 12.792 | 4.8755 | 153.6418 | 195.1895 | 143.7904 | 187.6979 | 124.7825
2012-10-30 04:11 | 25.0414 | 208.8009 | 157.2607 | 201.3954 | 138.257 | 70.7897 | 1.5403 | 38.9135 | 59.9371 | 25.2641 | 2.4148 | 38.8086 | 93.7159 | 16.071 | 12.8124 | 4.8802 | 153.653 | 195.9885 | 144.4483 | 188.583 | 125.4446
2012-10-30 04:11 | 25.0526 | 209.6667 | 158.0004 | 202.3355 | 139.054 | 70.8943 | 1.5402 | 38.9144 | 59.9347 | 25.2663 | 2.4149 | 38.8096 | 93.7207 | 16.071 | 12.787 | 4.8875 | 153.6555 | 196.8797 | 145.2134 | 189.5485 | 126.267
2012-10-30 04:11 | 25.0624 | 210.6649 | 158.7545 | 203.3569 | 139.708 | 71.0058 | 1.54 | 38.9146 | 59.9298 | 25.2687 | 2.4149 | 38.8097 | 93.7223 | 16.0708 | 12.7724 | 4.8951 | 153.6521 | 197.8925 | 145.9821 | 190.5845 | 126.9356
2012-10-30 04:11 | 25.0676 | 211.5286 | 159.4888 | 204.1967 | 140.2962 | 71.0746 | 1.5399 | 38.9133 | 59.9207 | 25.2708 | 2.4148 | 38.8083 | 93.716 | 16.0708 | 12.7817 | 4.8999 | 153.6367 | 198.7469 | 146.7071 | 191.415 | 127.5145
2012-10-30 04:11 | 25.0773 | 212.4125 | 160.1619 | 205.0568 | 141.0004 | 71.1416 | 1.5397 | 38.913 | 59.9158 | 25.2725 | 2.4148 | 38.8081 | 93.7148 | 16.0708 | 12.8036 | 4.9045 | 153.6307 | 199.6089 | 147.3583 | 192.2532 | 128.1968
2012-10-30 04:11 | 25.0796 | 213.1788 | 160.7657 | 205.9332 | 141.5939 | 71.2312 | 1.5396 | 38.9129 | 59.9106 | 25.2745 | 2.4148 | 38.8079 | 93.7142 | 16.0707 | 12.7896 | 4.9107 | 153.6248 | 200.3892 | 147.9761 | 193.1436 | 128.8043
2012-10-30 04:11 | 25.0833 | 214.0488 | 161.5109 | 206.7952 | 142.2235 | 71.3358 | 1.5395 | 38.9129 | 59.9059 | 25.2766 | 2.4148 | 38.8079 | 93.7146 | 16.0707 | 12.8258 | 4.9179 | 153.6205 | 201.223 | 148.6851 | 193.9694 | 129.3977
2012-10-30 04:11 | 25.0945 | 214.9381 | 162.2561 | 207.8184 | 142.8186 | 71.4375 | 1.5393 | 38.912 | 59.8979 | 25.2787 | 2.4148 | 38.8068 | 93.7104 | 16.0705 | 12.8248 | 4.9249 | 153.6083 | 202.1133 | 149.4313 | 194.9936 | 129.9938
2012-10-30 04:11 | 25.0983 | 215.7059 | 162.9515 | 208.6412 | 143.4834 | 71.4988 | 1.5392 | 38.9115 | 59.892 | 25.2806 | 2.4148 | 38.8064 | 93.7088 | 16.0704 | 12.8205 | 4.9291 | 153.6008 | 202.8854 | 150.131 | 195.8207 | 130.6629
2012-10-30 04:11 | 25.1222 | 217.0803 | 164.0365 | 210.1261 | 144.5568 | 71.6513 | 1.5389 | 38.9097 | 59.8786 | 25.2839 | 2.4147 | 38.8045 | 93.7005 | 16.0702 | 12.8403 | 4.9396 | 153.5791 | 204.24 | 151.1962 | 197.2858 | 131.7165
2012-10-30 04:11 | 25.1334 | 218.438 | 165.1328 | 211.6574 | 145.7607 | 71.8137 | 1.5387 | 38.9087 | 59.8675 | 25.2873 | 2.4146 | 38.8034 | 93.6962 | 16.0701 | 12.819 | 4.9508 | 153.5637 | 205.619 | 152.3138 | 198.8384 | 132.9417
2012-10-30 04:11 | 25.1349 | 219.1915 | 165.8074 | 212.3992 | 146.3165 | 71.8773 | 1.5385 | 38.9085 | 59.8626 | 25.2891 | 2.4146 | 38.8033 | 93.6957 | 16.0701 | 12.8083 | 4.9552 | 153.5584 | 206.3832 | 152.9991 | 199.5909 | 133.5082
2012-10-30 04:12 | 25.1551 | 219.9517 | 166.4192 | 213.2031 | 146.8333 | 71.9588 | 1.5385 | 38.9089 | 59.8595 | 25.2909 | 2.4147 | 38.8036 | 93.6986 | 16.0698 | 12.8488 | 4.9608 | 153.5581 | 207.1029 | 153.5704 | 200.3543 | 133.9845
2012-10-30 04:12 | 25.1604 | 220.7719 | 167.0863 | 214.0747 | 147.4048 | 72.0559 | 1.5384 | 38.9096 | 59.857 | 25.2929 | 2.4147 | 38.8041 | 93.7019 | 16.0697 | 12.8012 | 4.9675 | 153.5589 | 207.9707 | 154.2851 | 201.2735 | 134.6036
2012-10-30 04:12 | 25.1596 | 221.6134 | 167.845 | 214.9206 | 148.0629 | 72.1345 | 1.5382 | 38.9093 | 59.8519 | 25.2947 | 2.4147 | 38.8039 | 93.7015 | 16.0696 | 12.8411 | 4.973 | 153.5535 | 208.7723 | 155.0039 | 202.0795 | 135.2218
2012-10-30 04:12 | 25.1716 | 222.4636 | 168.5037 | 215.8067 | 148.7417 | 72.2183 | 1.5381 | 38.9081 | 59.8433 | 25.2968 | 2.4147 | 38.8027 | 93.6966 | 16.0694 | 12.849 | 4.9787 | 153.5399 | 209.6146 | 155.6547 | 202.9577 | 135.8927
2012-10-30 04:12 | 25.1888 | 223.3466 | 169.16 | 216.6862 | 149.3758 | 72.3056 | 1.5379 | 38.9072 | 59.8361 | 25.2986 | 2.4147 | 38.8016 | 93.6926 | 16.0692 | 12.8418 | 4.9847 | 153.5287 | 210.5048 | 156.3182 | 203.8444 | 136.534
2012-10-30 04:12 | 25.1858 | 224.1148 | 169.8431 | 217.4786 | 149.9475 | 72.3986 | 1.5378 | 38.9069 | 59.831 | 25.3003 | 2.4147 | 38.8013 | 93.6917 | 16.0691 | 12.851 | 4.9912 | 153.5228 | 211.2638 | 156.9921 | 204.6276 | 137.0965
2012-10-30 04:12 | 25.1985 | 224.8397 | 170.4758 | 218.2669 | 150.5699 | 72.4743 | 1.5377 | 38.9063 | 59.8249 | 25.3022 | 2.4146 | 38.8008 | 93.6897 | 16.069 | 12.8719 | 4.9964 | 153.5146 | 211.9678 | 157.6039 | 205.395 | 137.698
2012-10-30 04:12 | 25.2105 | 225.5519 | 171.1236 | 219.0978 | 151.1493 | 72.546 | 1.5375 | 38.9053 | 59.8174 | 25.304 | 2.4146 | 38.7997 | 93.6855 | 16.0688 | 12.8676 | 5.0013 | 153.503 | 212.6843 | 158.256 | 206.2302 | 138.2817
2012-10-30 04:12 | 25.2142 | 226.2637 | 171.7047 | 219.8888 | 151.6587 | 72.6304 | 1.5374 | 38.9046 | 59.8114 | 25.3057 | 2.4146 | 38.799 | 93.6832 | 16.0687 | 12.8916 | 5.0071 | 153.4946 | 213.3721 | 158.8131 | 206.9972 | 138.7671
2012-10-30 04:12 | 25.224 | 226.9815 | 172.2999 | 220.6032 | 152.2015 | 72.7153 | 1.5372 | 38.9024 | 59.8002 | 25.3076 | 2.4145 | 38.7968 | 93.6733 | 16.0685 | 12.877 | 5.013 | 153.4735 | 214.1045 | 159.4229 | 207.7262 | 139.3245
2012-10-30 04:12 | 25.2292 | 227.7221 | 172.8868 | 221.4035 | 152.763 | 72.8003 | 1.537 | 38.9002 | 59.7891 | 25.3094 | 2.4143 | 38.7946 | 93.6626 | 16.0685 | 12.8619 | 5.0189 | 153.4517 | 214.8602 | 160.0249 | 208.5416 | 139.9011
2012-10-30 04:12 | 25.2412 | 228.4301 | 173.504 | 222.1987 | 153.2852 | 72.8818 | 1.5369 | 38.9008 | 59.7872 | 25.311 | 2.4144 | 38.7951 | 93.6666 | 16.0683 | 12.882 | 5.0245 | 153.4538 | 215.5481 | 160.622 | 209.3167 | 140.4032
2012-10-30 04:12 | 25.2442 | 229.2981 | 174.1358 | 223.0164 | 153.8874 | 72.9644 | 1.5368 | 38.9022 | 59.7858 | 25.3134 | 2.4145 | 38.7964 | 93.6735 | 16.0681 | 12.8972 | 5.0302 | 153.4593 | 216.4009 | 161.2386 | 210.1192 | 140.9902
2012-10-30 04:12 | 25.2569 | 230.1167 | 174.8869 | 223.9223 | 154.6503 | 73.0499 | 1.5367 | 38.9026 | 59.7834 | 25.315 | 2.4146 | 38.7969 | 93.6771 | 16.0679 | 12.9099 | 5.0361 | 153.4605 | 217.2068 | 161.977 | 211.0124 | 141.7404
2012-10-30 04:12 | 25.2614 | 230.8798 | 175.5128 | 224.7019 | 155.1771 | 73.1222 | 1.5367 | 38.9041 | 59.7834 | 25.3168 | 2.4147 | 38.7981 | 93.6848 | 16.0677 | 12.8859 | 5.041 | 153.4682 | 217.9939 | 162.6269 | 211.816 | 142.2912
2012-10-30 04:12 | 25.2674 | 231.5775 | 176.1286 | 225.4987 | 155.6977 | 73.2002 | 1.5366 | 38.9044 | 59.7809 | 25.3184 | 2.4147 | 38.7986 | 93.6877 | 16.0676 | 12.8879 | 5.0464 | 153.4686 | 218.6896 | 163.2407 | 212.6108 | 142.8098
2012-10-30 04:12 | 25.2748 | 232.3472 | 176.7968 | 226.2417 | 156.2819 | 73.2828 | 1.5365 | 38.9043 | 59.7762 | 25.3202 | 2.4147 | 38.7984 | 93.6882 | 16.0673 | 12.8993 | 5.0521 | 153.4644 | 219.4479 | 163.8975 | 213.3424 | 143.3826
2012-10-30 04:12 | 25.2905 | 233.1487 | 177.3874 | 227.0912 | 156.8668 | 73.3811 | 1.5364 | 38.9053 | 59.7748 | 25.3221 | 2.4148 | 38.7993 | 93.694 | 16.0671 | 12.9232 | 5.0589 | 153.4688 | 220.2255 | 164.4642 | 214.168 | 143.9436
2012-10-30 04:12 | 25.3003 | 233.8341 | 177.9879 | 227.8325 | 157.4111 | 73.4406 | 1.5363 | 38.9055 | 59.7721 | 25.3234 | 2.4149 | 38.7994 | 93.6951 | 16.067 | 12.9574 | 5.063 | 153.4672 | 220.8767 | 165.0305 | 214.8751 | 144.4537
2012-10-30 04:12 | 25.3063 | 234.5707 | 178.5886 | 228.6053 | 158.0293 | 73.5273 | 1.5362 | 38.9056 | 59.7677 | 25.3255 | 2.4149 | 38.7996 | 93.6977 | 16.0667 | 12.9105 | 5.069 | 153.4654 | 221.6602 | 165.6781 | 215.6948 | 145.1188
2012-10-30 04:12 | 25.3115 | 235.3024 | 179.2278 | 229.3575 | 158.6469 | 73.6157 | 1.5361 | 38.9058 | 59.7645 | 25.3271 | 2.415 | 38.7997 | 93.6998 | 16.0664 | 12.8954 | 5.0751 | 153.4643 | 222.407 | 166.3324 | 216.4621 | 145.7515
2012-10-30 04:12 | 25.3152 | 236.0481 | 179.7952 | 230.061 | 159.2555 | 73.6862 | 1.536 | 38.9059 | 59.7603 | 25.3289 | 2.415 | 38.7997 | 93.7011 | 16.0662 | 12.9113 | 5.0799 | 153.4614 | 223.1368 | 166.8839 | 217.1497 | 146.3442
2012-10-30 04:12 | 25.3422 | 236.7307 | 180.433 | 230.7877 | 159.7595 | 73.7469 | 1.536 | 38.9069 | 59.7601 | 25.3304 | 2.4151 | 38.8008 | 93.7075 | 16.066 | 12.9505 | 5.0841 | 153.4676 | 223.7802 | 167.4825 | 217.8372 | 146.809
2012-10-30 04:12 | 25.3369 | 237.3181 | 180.9798 | 231.4662 | 160.177 | 73.8284 | 1.5359 | 38.9072 | 59.7567 | 25.3322 | 2.4151 | 38.801 | 93.7102 | 16.0657 | 12.9208 | 5.0897 | 153.467 | 224.3973 | 168.059 | 218.5454 | 147.2562
2012-10-30 04:12 | 25.3377 | 237.9507 | 181.5564 | 232.194 | 160.725 | 73.9133 | 1.5358 | 38.9068 | 59.7527 | 25.3334 | 2.4151 | 38.8006 | 93.7091 | 16.0655 | 12.9053 | 5.0956 | 153.4619 | 225.0454 | 168.6511 | 219.2887 | 147.8197
2012-10-30 04:13 | 25.3497 | 238.6925 | 182.1443 | 232.8386 | 161.1717 | 73.9694 | 1.5357 | 38.9064 | 59.7475 | 25.335 | 2.4151 | 38.8 | 93.7072 | 16.0654 | 12.9498 | 5.0995 | 153.4547 | 225.7427 | 169.1945 | 219.8888 | 148.2219
2012-10-30 04:13 | 25.3646 | 239.4018 | 182.6412 | 233.5292 | 161.5942 | 74.0428 | 1.5356 | 38.9065 | 59.7436 | 25.3369 | 2.4152 | 38.8003 | 93.7099 | 16.0651 | 12.9525 | 5.1045 | 153.4534 | 226.4493 | 169.6887 | 220.5767 | 148.6417
2012-10-30 04:13 | 25.3743 | 239.992 | 183.1604 | 234.2381 | 162.0932 | 74.1098 | 1.5355 | 38.9067 | 59.7406 | 25.3383 | 2.4152 | 38.8003 | 93.7116 | 16.0649 | 12.9337 | 5.1091 | 153.4522 | 227.0583 | 170.2267 | 221.3044 | 149.1595
2012-10-30 04:13 | 25.3743 | 240.498 | 183.7013 | 234.8664 | 162.6874 | 74.1757 | 1.5354 | 38.9071 | 59.7391 | 25.3396 | 2.4153 | 38.8008 | 93.7149 | 16.0647 | 12.9256 | 5.1137 | 153.454 | 227.5724 | 170.7757 | 221.9408 | 149.7618
2012-10-30 04:13 | 25.3871 | 241.2127 | 184.2042 | 235.4918 | 163.2131 | 74.2289 | 1.5353 | 38.907 | 59.7357 | 25.3408 | 2.4153 | 38.8006 | 93.7157 | 16.0644 | 12.9872 | 5.1173 | 153.4514 | 228.2255 | 171.217 | 222.5046 | 150.2259
2012-10-30 04:13 | 25.3915 | 241.9462 | 184.8124 | 236.2618 | 163.7577 | 74.304 | 1.5352 | 38.9064 | 59.7295 | 25.3427 | 2.4153 | 38.8 | 93.7144 | 16.0641 | 12.9632 | 5.1225 | 153.4439 | 228.983 | 171.8492 | 223.2986 | 150.7945
2012-10-30 04:13 | 25.4043 | 242.6646 | 185.4213 | 237.0244 | 164.1981 | 74.3982 | 1.5351 | 38.906 | 59.7246 | 25.3443 | 2.4153 | 38.7996 | 93.7144 | 16.0638 | 13.0126 | 5.129 | 153.4389 | 229.652 | 172.4087 | 224.0118 | 151.1855
2012-10-30 04:13 | 25.4185 | 243.294 | 185.8748 | 237.5881 | 164.6167 | 74.4554 | 1.535 | 38.9053 | 59.7186 | 25.3458 | 2.4153 | 38.7989 | 93.7116 | 16.0637 | 12.9698 | 5.133 | 153.4302 | 230.3242 | 172.905 | 224.6183 | 151.6469
2012-10-30 04:13 | 25.4177 | 243.8966 | 186.3723 | 238.2858 | 165.0087 | 74.5288 | 1.5349 | 38.9053 | 59.7155 | 25.3472 | 2.4154 | 38.7988 | 93.7139 | 16.0632 | 12.9771 | 5.138 | 153.4294 | 230.9195 | 173.3952 | 225.3087 | 152.0316
2012-10-30 04:13 | 25.4237 | 244.4628 | 186.904 | 238.8842 | 165.4931 | 74.5826 | 1.5348 | 38.9066 | 59.7153 | 25.3489 | 2.4155 | 38.8 | 93.7202 | 16.0631 | 12.9669 | 5.1417 | 153.4356 | 231.4959 | 173.9371 | 225.9173 | 152.5262
2012-10-30 04:13 | 25.4275 | 245.1286 | 187.5082 | 239.5432 | 166.0066 | 74.6554 | 1.5348 | 38.9065 | 59.7121 | 25.3502 | 2.4155 | 38.7999 | 93.7216 | 16.0628 | 12.9544 | 5.1467 | 153.4337 | 232.1742 | 174.5538 | 226.5888 | 153.0522
2012-10-30 04:13 | 25.4387 | 245.782 | 187.959 | 240.2525 | 166.5004 | 74.7219 | 1.5346 | 38.9054 | 59.7057 | 25.3516 | 2.4155 | 38.7989 | 93.7189 | 16.0624 | 12.9493 | 5.1513 | 153.4246 | 232.8327 | 175.0097 | 227.3032 | 153.5511
2012-10-30 04:13 | 25.4447 | 246.3479 | 188.4086 | 240.8585 | 166.9501 | 74.7849 | 1.5345 | 38.9048 | 59.7004 | 25.353 | 2.4155 | 38.7983 | 93.717 | 16.0622 | 12.9512 | 5.1557 | 153.4173 | 233.3967 | 175.4574 | 227.9073 | 153.9989
2012-10-30 04:13 | 25.4514 | 246.8205 | 188.8981 | 241.4573 | 167.4318 | 74.8369 | 1.5345 | 38.905 | 59.6984 | 25.3541 | 2.4155 | 38.7985 | 93.7195 | 16.062 | 12.9783 | 5.1593 | 153.4179 | 233.8422 | 175.9198 | 228.479 | 154.4535
2012-10-30 04:13 | 25.4544 | 247.4419 | 189.3487 | 242.0013 | 167.836 | 74.9033 | 1.5344 | 38.9054 | 59.6959 | 25.3556 | 2.4156 | 38.7986 | 93.7222 | 16.0616 | 12.9406 | 5.1638 | 153.4181 | 234.5013 | 176.4081 | 229.0607 | 154.8954
2012-10-30 04:13 | 25.4589 | 247.9316 | 189.8598 | 242.6134 | 168.2494 | 74.9692 | 1.5342 | 38.9037 | 59.6877 | 25.3569 | 2.4155 | 38.7971 | 93.7155 | 16.0615 | 12.9044 | 5.1684 | 153.4032 | 235.0272 | 176.9554 | 229.709 | 155.345
2012-10-30 04:13 | 25.4739 | 248.4553 | 190.3068 | 243.1222 | 168.5761 | 75.0155 | 1.5341 | 38.9027 | 59.682 | 25.3581 | 2.4155 | 38.796 | 93.7128 | 16.0611 | 12.9397 | 5.1716 | 153.3948 | 235.5156 | 177.3671 | 230.1825 | 155.6364
2012-10-30 04:13 | 25.4828 | 249.0353 | 190.7281 | 243.617 | 168.9902 | 75.0756 | 1.5341 | 38.9026 | 59.6792 | 25.3592 | 2.4155 | 38.7957 | 93.7119 | 16.061 | 12.9608 | 5.1757 | 153.3911 | 236.0745 | 177.7673 | 230.6562 | 156.0294
2012-10-30 04:13 | 25.4828 | 249.4499 | 191.1013 | 244.2007 | 169.3449 | 75.1154 | 1.534 | 38.9025 | 59.6765 | 25.3601 | 2.4156 | 38.7956 | 93.7136 | 16.0606 | 12.9731 | 5.1785 | 153.3901 | 236.4768 | 178.1282 | 231.2276 | 156.3718
2012-10-30 04:13 | 25.4941 | 249.9188 | 191.4975 | 244.7174 | 169.6875 | 75.201 | 1.5339 | 38.9024 | 59.6731 | 25.3615 | 2.4156 | 38.7956 | 93.7148 | 16.0604 | 12.964 | 5.1844 | 153.3879 | 236.9548 | 178.5335 | 231.7534 | 156.7235
2012-10-30 04:13 | 25.509 | 250.4577 | 191.9171 | 245.2571 | 170.1195 | 75.2414 | 1.5338 | 38.9019 | 59.669 | 25.3625 | 2.4156 | 38.795 | 93.7141 | 16.06 | 12.9504 | 5.1871 | 153.3832 | 237.5073 | 178.9667 | 232.3067 | 157.1691
2012-10-30 04:13 | 25.512 | 251.0628 | 192.4348 | 245.8827 | 170.5475 | 75.2957 | 1.5337 | 38.9019 | 59.6658 | 25.364 | 2.4157 | 38.7951 | 93.7157 | 16.0598 | 12.9656 | 5.1909 | 153.3815 | 238.0972 | 179.4692 | 232.9171 | 157.5819
2012-10-30 04:13 | 25.5217 | 251.6001 | 192.8296 | 246.4249 | 170.9745 | 75.3611 | 1.5337 | 38.9023 | 59.664 | 25.3652 | 2.4157 | 38.7954 | 93.7194 | 16.0595 | 12.9753 | 5.1954 | 153.3835 | 238.6248 | 179.8543 | 233.4496 | 157.9992
2012-10-30 04:13 | 25.5255 | 252.1237 | 193.2814 | 246.9298 | 171.3801 | 75.4119 | 1.5336 | 38.9014 | 59.6584 | 25.3664 | 2.4157 | 38.7945 | 93.7162 | 16.0592 | 12.9466 | 5.1989 | 153.3746 | 239.1771 | 180.3348 | 233.9832 | 158.4335
2012-10-30 04:13 | 25.533 | 252.5454 | 193.6938 | 247.4539 | 171.7816 | 75.4645 | 1.5335 | 38.9009 | 59.6537 | 25.3678 | 2.4157 | 38.794 | 93.7156 | 16.059 | 12.9581 | 5.2025 | 153.3693 | 239.5873 | 180.7357 | 234.4958 | 158.8235
2012-10-30 04:13 | 25.5337 | 253.1048 | 194.1268 | 247.9634 | 172.1608 | 75.5148 | 1.5334 | 38.9 | 59.6479 | 25.3691 | 2.4157 | 38.793 | 93.712 | 16.0588 | 12.967 | 5.206 | 153.3599 | 240.1378 | 181.1598 | 234.9964 | 159.1938
2012-10-30 04:14 | 25.5375 | 253.6515 | 194.6456 | 248.5267 | 172.4885 | 75.5795 | 1.5333 | 38.8991 | 59.6425 | 25.3702 | 2.4157 | 38.7921 | 93.7085 | 16.0586 | 12.9667 | 5.2105 | 153.351 | 240.6848 | 181.6789 | 235.56 | 159.5218
2012-10-30 04:14 | 25.5509 | 254.1375 | 194.9853 | 249.0956 | 172.865 | 75.6489 | 1.5332 | 38.8989 | 59.6395 | 25.3712 | 2.4157 | 38.7918 | 93.7093 | 16.0582 | 12.976 | 5.2152 | 153.3488 | 241.1615 | 182.0093 | 236.1196 | 159.889
2012-10-30 04:14 | 25.5517 | 254.6029 | 195.3282 | 249.5853 | 173.221 | 75.6754 | 1.5331 | 38.8987 | 59.6359 | 25.3725 | 2.4157 | 38.7917 | 93.7102 | 16.058 | 12.989 | 5.2171 | 153.3461 | 241.6139 | 182.3392 | 236.5963 | 160.232
2012-10-30 04:14 | 25.5696 | 255.0542 | 195.755 | 250.0452 | 173.5708 | 75.739 | 1.5331 | 38.899 | 59.6343 | 25.3736 | 2.4158 | 38.7919 | 93.7124 | 16.0578 | 12.9622 | 5.2214 | 153.3467 | 242.092 | 182.7928 | 237.083 | 160.6086
2012-10-30 04:14 | 25.5719 | 255.5483 | 196.1741 | 250.5597 | 174.0495 | 75.7847 | 1.533 | 38.8995 | 59.6326 | 25.3749 | 2.4159 | 38.7923 | 93.7167 | 16.0573 | 12.9482 | 5.2246 | 153.3493 | 242.6001 | 183.2259 | 237.6115 | 161.1013
2012-10-30 04:14 | 25.5764 | 256.0251 | 196.6308 | 251.0216 | 174.4485 | 75.8332 | 1.5329 | 38.8995 | 59.631 | 25.3756 | 2.4159 | 38.7924 | 93.7189 | 16.057 | 12.9853 | 5.2279 | 153.3499 | 243.0398 | 183.6455 | 238.0363 | 161.4632
2012-10-30 04:14 | 25.5868 | 256.5388 | 196.9755 | 251.5402 | 174.787 | 75.8875 | 1.5329 | 38.9005 | 59.6313 | 25.3768 | 2.416 | 38.7933 | 93.7254 | 16.0567 | 13.0364 | 5.2317 | 153.3567 | 243.5024 | 183.9391 | 238.5038 | 161.7506
2012-10-30 04:14 | 25.5906 | 256.9874 | 197.3962 | 251.9934 | 175.1151 | 75.9182 | 1.5329 | 38.9013 | 59.6316 | 25.3776 | 2.4161 | 38.7941 | 93.7301 | 16.0565 | 12.975 | 5.2338 | 153.3617 | 244.0124 | 184.4212 | 239.0184 | 162.1401
2012-10-30 04:14 | 25.5846 | 257.3442 | 197.7796 | 252.4055 | 175.4217 | 75.9789 | 1.5328 | 38.9 | 59.6245 | 25.3789 | 2.4161 | 38.7927 | 93.7256 | 16.0562 | 12.9812 | 5.238 | 153.3502 | 244.363 | 184.7984 | 239.4243 | 162.4405
2012-10-30 04:14 | 25.5988 | 257.8459 | 198.1321 | 253.0058 | 175.772 | 76.032 | 1.5327 | 38.8992 | 59.6199 | 25.38 | 2.416 | 38.792 | 93.7234 | 16.056 | 12.9954 | 5.2416 | 153.3433 | 244.8505 | 185.1367 | 240.0104 | 162.7766
2012-10-30 04:14 | 25.613 | 258.332 | 198.5869 | 253.4927 | 176.1458 | 76.0892 | 1.5326 | 38.8996 | 59.6183 | 25.3811 | 2.4161 | 38.7924 | 93.7265 | 16.0557 | 13.0137 | 5.2456 | 153.3448 | 245.3183 | 185.5732 | 240.479 | 163.1321
2012-10-30 04:14 | 25.607 | 258.8016 | 198.9153 | 253.9163 | 176.4447 | 76.1332 | 1.5326 | 38.9001 | 59.6176 | 25.3821 | 2.4162 | 38.7928 | 93.7311 | 16.0553 | 13.0036 | 5.2486 | 153.3487 | 245.798 | 185.9117 | 240.9127 | 163.4411
2012-10-30 04:14 | 25.6198 | 259.251 | 199.2547 | 254.343 | 176.8368 | 76.1742 | 1.5325 | 38.9003 | 59.6159 | 25.3831 | 2.4162 | 38.7929 | 93.7327 | 16.0552 | 12.992 | 5.2514 | 153.3487 | 246.259 | 186.2627 | 241.351 | 163.8448
2012-10-30 04:14 | 25.6362 | 259.6767 | 199.6488 | 254.8475 | 177.2598 | 76.2308 | 1.5325 | 38.9001 | 59.613 | 25.384 | 2.4163 | 38.7927 | 93.7335 | 16.0548 | 13.0166 | 5.2554 | 153.3464 | 246.6601 | 186.6322 | 241.8309 | 164.2432
2012-10-30 04:14 | 25.6355 | 260.0813 | 200.0415 | 255.245 | 177.5251 | 76.2638 | 1.5324 | 38.8996 | 59.6082 | 25.3854 | 2.4163 | 38.7922 | 93.7323 | 16.0546 | 12.9955 | 5.2576 | 153.3405 | 247.0858 | 187.046 | 242.2495 | 164.5296
2012-10-30 04:14 | 25.64 | 260.4676 | 200.353 | 255.6711 | 177.8151 | 76.3378 | 1.5323 | 38.8993 | 59.6058 | 25.386 | 2.4163 | 38.7919 | 93.7319 | 16.0544 | 13.0082 | 5.2627 | 153.3377 | 247.4594 | 187.3448 | 242.6629 | 164.8069
2012-10-30 04:14 | 25.643 | 260.8486 | 200.6786 | 256.0917 | 178.221 | 76.3586 | 1.5323 | 38.8998 | 59.6054 | 25.3869 | 2.4163 | 38.7924 | 93.7358 | 16.0542 | 12.9949 | 5.2642 | 153.3412 | 247.8537 | 187.6837 | 243.0968 | 165.2261
2012-10-30 04:14 | 25.6415 | 261.2444 | 201.0116 | 256.5311 | 178.6381 | 76.4129 | 1.5323 | 38.9008 | 59.6066 | 25.3877 | 2.4165 | 38.7933 | 93.7423 | 16.0538 | 12.9974 | 5.2679 | 153.3489 | 248.247 | 188.0142 | 243.5337 | 165.6407
2012-10-30 04:14 | 25.6617 | 261.7787 | 201.4529 | 257.0078 | 178.9518 | 76.4597 | 1.5322 | 38.9013 | 59.6046 | 25.3892 | 2.4165 | 38.7939 | 93.7471 | 16.0535 | 12.9891 | 5.2711 | 153.3517 | 248.7896 | 188.4638 | 244.0187 | 165.9627
2012-10-30 04:14 | 25.6617 | 262.1642 | 201.8504 | 257.4564 | 179.1905 | 76.499 | 1.5322 | 38.9015 | 59.6031 | 25.39 | 2.4166 | 38.7938 | 93.7483 | 16.0532 | 12.981 | 5.2738 | 153.3514 | 249.1832 | 188.8694 | 244.4754 | 166.2095
2012-10-30 04:14 | 25.6706 | 262.6014 | 202.2056 | 257.8709 | 179.489 | 76.5603 | 1.5321 | 38.9023 | 59.6034 | 25.391 | 2.4167 | 38.7947 | 93.7539 | 16.053 | 12.994 | 5.2781 | 153.3573 | 249.6074 | 189.2116 | 244.8769 | 166.495
2012-10-30 04:14 | 25.6781 | 262.932 | 202.4887 | 258.2957 | 179.6883 | 76.5851 | 1.5321 | 38.9023 | 59.6017 | 25.3916 | 2.4167 | 38.7947 | 93.7562 | 16.0526 | 12.9893 | 5.2798 | 153.358 | 249.9427 | 189.4994 | 245.3064 | 166.699
2012-10-30 04:14 | 25.6736 | 263.344 | 202.7754 | 258.7357 | 179.9148 | 76.6342 | 1.5321 | 38.9023 | 59.6003 | 25.3923 | 2.4168 | 38.7948 | 93.758 | 16.0524 | 12.9807 | 5.2832 | 153.3582 | 250.3633 | 189.7947 | 245.755 | 166.9341
2012-10-30 04:14 | 25.6849 | 263.6822 | 203.1174 | 259.1446 | 180.2001 | 76.6851 | 1.532 | 38.9017 | 59.5956 | 25.3935 | 2.4168 | 38.7941 | 93.756 | 16.0522 | 13.0123 | 5.2867 | 153.3515 | 250.6699 | 190.1051 | 246.1323 | 167.1878
2012-10-30 04:14 | 25.6983 | 264.2326 | 203.5283 | 259.6219 | 180.6403 | 76.725 | 1.5319 | 38.9008 | 59.5903 | 25.3947 | 2.4168 | 38.7931 | 93.7534 | 16.0517 | 13.0461 | 5.2894 | 153.3437 | 251.1865 | 190.4822 | 246.5758 | 167.5942
2012-10-30 04:14 | 25.7073 | 264.6929 | 203.9231 | 260.0713 | 181.0095 | 76.7833 | 1.5318 | 38.9007 | 59.5873 | 25.3958 | 2.4168 | 38.7929 | 93.7543 | 16.0514 | 13.0022 | 5.2934 | 153.3416 | 251.6907 | 190.9209 | 247.0691 | 168.0073
2012-10-30 04:14 | 25.7125 | 265.0481 | 204.2275 | 260.5129 | 181.3711 | 76.8273 | 1.5317 | 38.9 | 59.5829 | 25.3967 | 2.4168 | 38.7923 | 93.753 | 16.0511 | 12.987 | 5.2965 | 153.3359 | 252.0611 | 191.2405 | 247.5259 | 168.3841
2012-10-30 04:14 | 25.7163 | 265.4875 | 204.5738 | 260.8985 | 181.5982 | 76.8631 | 1.5316 | 38.8999 | 59.5809 | 25.3974 | 2.4168 | 38.7922 | 93.754 | 16.0508 | 13.0193 | 5.2989 | 153.3349 | 252.4682 | 191.5545 | 247.8792 | 168.5789
2012-10-30 04:15 | 25.7215 | 265.8272 | 204.8585 | 261.3166 | 181.952 | 76.9111 | 1.5316 | 38.9006 | 59.5811 | 25.3982 | 2.4169 | 38.7929 | 93.7595 | 16.0505 | 12.9797 | 5.3023 | 153.3406 | 252.8475 | 191.8788 | 248.3369 | 168.9723
2012-10-30 04:15 | 25.7335 | 266.193 | 205.1815 | 261.7298 | 182.1873 | 76.9544 | 1.5316 | 38.9006 | 59.5784 | 25.3994 | 2.4169 | 38.7929 | 93.7601 | 16.0504 | 12.9877 | 5.3052 | 153.3385 | 253.2053 | 192.1938 | 248.7421 | 169.1996
2012-10-30 04:15 | 25.7395 | 266.6266 | 205.5221 | 262.1369 | 182.4127 | 76.9891 | 1.5315 | 38.9006 | 59.5767 | 25.4002 | 2.417 | 38.7928 | 93.7615 | 16.0501 | 12.9855 | 5.3076 | 153.3383 | 253.6411 | 192.5366 | 249.1514 | 169.4272
2012-10-30 04:15 | 25.7387 | 266.9756 | 205.8166 | 262.5138 | 182.5294 | 77.0451 | 1.5314 | 38.8999 | 59.5722 | 25.4012 | 2.417 | 38.7922 | 93.7604 | 16.0498 | 13.0295 | 5.3115 | 153.3326 | 253.9461 | 192.7871 | 249.4843 | 169.4999
2012-10-30 04:15 | 25.747 | 267.3376 | 206.0454 | 262.8763 | 182.8355 | 77.0578 | 1.5313 | 38.8982 | 59.5656 | 25.4017 | 2.4169 | 38.7903 | 93.7523 | 16.0496 | 13.05 | 5.3124 | 153.3179 | 254.2876 | 192.9954 | 249.8263 | 169.7855
2012-10-30 04:15 | 25.7492 | 267.7051 | 206.3468 | 263.2833 | 183.0807 | 77.1128 | 1.5313 | 38.8982 | 59.5636 | 25.4026 | 2.4169 | 38.7903 | 93.7531 | 16.0495 | 13.0237 | 5.3162 | 153.3167 | 254.6814 | 193.3231 | 250.2596 | 170.057
2012-10-30 04:15 | 25.7604 | 268.0854 | 206.7006 | 263.6954 | 183.4681 | 77.1486 | 1.5312 | 38.898 | 59.5611 | 25.4034 | 2.417 | 38.7902 | 93.7549 | 16.049 | 12.9698 | 5.3186 | 153.316 | 255.1156 | 193.7308 | 250.7256 | 170.4983
2012-10-30 04:15 | 25.7664 | 268.476 | 206.9979 | 264.11 | 183.5917 | 77.2006 | 1.5311 | 38.8978 | 59.5575 | 25.4047 | 2.417 | 38.7898 | 93.7552 | 16.0487 | 12.9473 | 5.3222 | 153.3126 | 255.5287 | 194.0506 | 251.1627 | 170.6444
2012-10-30 04:15 | 25.7687 | 268.8162 | 207.2444 | 264.4936 | 183.7116 | 77.2451 | 1.5311 | 38.8973 | 59.5547 | 25.4053 | 2.417 | 38.7893 | 93.7553 | 16.0483 | 13.0066 | 5.3253 | 153.31 | 255.8096 | 194.2378 | 251.487 | 170.705
2012-10-30 04:15 | 25.7747 | 269.2067 | 207.5417 | 264.8847 | 184.1174 | 77.2792 | 1.531 | 38.8974 | 59.5532 | 25.406 | 2.4171 | 38.7894 | 93.7565 | 16.0481 | 13.0329 | 5.3276 | 153.3097 | 256.1738 | 194.5088 | 251.8518 | 171.0845
2012-10-30 04:15 | 25.7836 | 269.5534 | 207.8184 | 265.1426 | 184.5458 | 77.3087 | 1.531 | 38.8975 | 59.5521 | 25.4067 | 2.4171 | 38.7895 | 93.7585 | 16.0479 | 13.0378 | 5.3297 | 153.3105 | 256.5156 | 194.7806 | 252.1048 | 171.508
2012-10-30 04:15 | 25.7934 | 269.8816 | 208.1471 | 265.5763 | 184.9633 | 77.3688 | 1.531 | 38.8979 | 59.5509 | 25.4076 | 2.4172 | 38.7899 | 93.7622 | 16.0475 | 13.0028 | 5.3338 | 153.3131 | 256.8788 | 195.1443 | 252.5735 | 171.9605
2012-10-30 04:15 | 25.8016 | 270.2974 | 208.5019 | 265.9967 | 185.1676 | 77.3977 | 1.5309 | 38.8983 | 59.5493 | 25.4088 | 2.4172 | 38.7902 | 93.7651 | 16.0474 | 13.068 | 5.3358 | 153.3144 | 257.2294 | 195.4339 | 252.9287 | 172.0996
2012-10-30 04:15 | 25.7993 | 270.5979 | 208.728 | 266.3288 | 185.2163 | 77.4352 | 1.5309 | 38.8992 | 59.5508 | 25.4094 | 2.4173 | 38.791 | 93.7702 | 16.0471 | 13.0413 | 5.3384 | 153.321 | 257.5566 | 195.6867 | 253.2875 | 172.175
2012-10-30 04:15 | 25.8068 | 270.9391 | 209.019 | 266.6823 | 185.5101 | 77.4838 | 1.5308 | 38.8991 | 59.5483 | 25.4103 | 2.4174 | 38.791 | 93.7716 | 16.0468 | 13.0284 | 5.3417 | 153.3198 | 257.9107 | 195.9906 | 253.6539 | 172.4817
2012-10-30 04:15 | 25.8121 | 271.251 | 209.2724 | 266.9907 | 185.9076 | 77.5156 | 1.5308 | 38.8988 | 59.5458 | 25.411 | 2.4174 | 38.7908 | 93.7721 | 16.0466 | 13.0541 | 5.3439 | 153.3179 | 258.1969 | 196.2183 | 253.9366 | 172.8535
2012-10-30 04:15 | 25.8218 | 271.6254 | 209.6128 | 267.4124 | 186.3409 | 77.5618 | 1.5307 | 38.8989 | 59.5442 | 25.4118 | 2.4174 | 38.7907 | 93.7737 | 16.0463 | 13.0556 | 5.3471 | 153.3179 | 258.5698 | 196.5572 | 254.3568 | 173.2853
2012-10-30 04:15 | 25.833 | 271.9974 | 209.9469 | 267.851 | 186.5718 | 77.6092 | 1.5307 | 38.8983 | 59.54 | 25.4129 | 2.4174 | 38.7902 | 93.7727 | 16.046 | 13.0383 | 5.3504 | 153.3127 | 258.9591 | 196.9086 | 254.8127 | 173.5335
2012-10-30 04:15 | 25.8233 | 272.3019 | 210.1853 | 268.1483 | 186.8053 | 77.63 | 1.5306 | 38.8977 | 59.5359 | 25.4138 | 2.4174 | 38.7895 | 93.7703 | 16.0459 | 13.0612 | 5.3518 | 153.3062 | 259.2407 | 197.1241 | 255.0871 | 173.7441
2012-10-30 04:15 | 25.8345 | 272.6117 | 210.5037 | 268.4586 | 186.997 | 77.66 | 1.5306 | 38.8982 | 59.5359 | 25.4145 | 2.4175 | 38.79 | 93.774 | 16.0456 | 13.0195 | 5.3539 | 153.3099 | 259.5922 | 197.4842 | 255.4391 | 173.9775
2012-10-30 04:15 | 25.8398 | 272.9389 | 210.698 | 268.7785 | 187.2848 | 77.6953 | 1.5305 | 38.8981 | 59.5339 | 25.4151 | 2.4175 | 38.7899 | 93.7749 | 16.0454 | 13.0125 | 5.3563 | 153.3089 | 259.9264 | 197.6855 | 255.766 | 174.2723
2012-10-30 04:15 | 25.8457 | 273.2899 | 210.9914 | 269.1068 | 187.5418 | 77.7225 | 1.5305 | 38.8982 | 59.5326 | 25.4158 | 2.4175 | 38.7898 | 93.7759 | 16.0452 | 13.0754 | 5.3582 | 153.3085 | 260.2145 | 197.916 | 256.0314 | 174.4664
2012-10-30 04:15 | 25.854 | 273.5774 | 211.3226 | 269.453 | 187.7581 | 77.771 | 1.5304 | 38.898 | 59.5297 | 25.4169 | 2.4176 | 38.7898 | 93.7775 | 16.0449 | 13.0675 | 5.3615 | 153.3071 | 260.5099 | 198.2551 | 256.3855 | 174.6906
2012-10-30 04:15 | 25.8577 | 273.8834 | 211.5665 | 269.7556 | 188.0522 | 77.819 | 1.5304 | 38.8976 | 59.5279 | 25.417 | 2.4176 | 38.7894 | 93.7773 | 16.0446 | 13.0752 | 5.3648 | 153.3052 | 260.8082 | 198.4913 | 256.6804 | 174.977
2012-10-30 04:15 | 25.8622 | 274.2213 | 211.8309 | 270.0513 | 188.1823 | 77.83 | 1.5303 | 38.8972 | 59.5245 | 25.4179 | 2.4176 | 38.7888 | 93.7753 | 16.0444 | 13.0309 | 5.3656 | 153.2999 | 261.1904 | 198.8 | 257.0204 | 175.1514
2012-10-30 04:15 | 25.8689 | 274.4838 | 212.0463 | 270.3647 | 188.3892 | 77.8727 | 1.5302 | 38.8968 | 59.5217 | 25.4187 | 2.4176 | 38.7885 | 93.7751 | 16.0442 | 13.0335 | 5.3685 | 153.2968 | 261.4503 | 199.0128 | 257.3312 | 175.3557
2012-10-30 04:15 | 25.8824 | 274.8228 | 212.3237 | 270.7198 | 188.7167 | 77.897 | 1.5302 | 38.8966 | 59.5191 | 25.4195 | 2.4176 | 38.7883 | 93.7759 | 16.0439 | 13.0592 | 5.3702 | 153.295 | 261.7636 | 199.2645 | 257.6606 | 175.6575
2012-10-30 04:16 | 25.8884 | 275.1078 | 212.5817 | 271.0828 | 188.9661 | 77.9346 | 1.5302 | 38.8973 | 59.5196 | 25.4202 | 2.4177 | 38.7889 | 93.7803 | 16.0437 | 13.0611 | 5.3728 | 153.2999 | 262.0467 | 199.5206 | 258.0217 | 175.905
2012-10-30 04:16 | 25.8996 | 275.4173 | 212.8056 | 271.3036 | 189.1226 | 77.9733 | 1.5302 | 38.8977 | 59.5198 | 25.4207 | 2.4178 | 38.7893 | 93.7842 | 16.0433 | 13.0642 | 5.3755 | 153.304 | 262.3531 | 199.7414 | 258.2394 | 176.0584
2012-10-30 04:16 | 25.8876 | 275.7074 | 213.0062 | 271.6447 | 189.2362 | 77.9958 | 1.5301 | 38.8978 | 59.5184 | 25.4214 | 2.4178 | 38.7895 | 93.7855 | 16.0432 | 13.0522 | 5.377 | 153.3039 | 262.6552 | 199.954 | 258.5925 | 176.184
2012-10-30 04:16 | 25.9004 | 276.0224 | 213.3301 | 271.9523 | 189.4846 | 78.0334 | 1.5301 | 38.8978 | 59.5158 | 25.4224 | 2.4178 | 38.7893 | 93.7861 | 16.043 | 13.0813 | 5.3796 | 153.3019 | 262.9411 | 200.2488 | 258.871 | 176.4033
2012-10-30 04:16 | 25.9011 | 276.3406 | 213.571 | 272.2866 | 189.7365 | 78.0658 | 1.53 | 38.8973 | 59.5133 | 25.4229 | 2.4179 | 38.7888 | 93.7866 | 16.0425 | 13.0087 | 5.3819 | 153.2999 | 263.3319 | 200.5623 | 259.2779 | 176.7278
2012-10-30 04:16 | 25.9093 | 276.7206 | 213.8864 | 272.6051 | 190.1296 | 78.1033 | 1.5299 | 38.8968 | 59.5094 | 25.4239 | 2.4179 | 38.7882 | 93.7842 | 16.0424 | 13.0576 | 5.3844 | 153.2936 | 263.663 | 200.8288 | 259.5475 | 177.072
2012-10-30 04:16 | 25.9123 | 277.0101 | 214.1414 | 272.9688 | 190.3996 | 78.1311 | 1.5299 | 38.896 | 59.5054 | 25.4245 | 2.4178 | 38.7874 | 93.782 | 16.0422 | 12.9995 | 5.3864 | 153.2875 | 264.0106 | 201.1419 | 259.9693 | 177.4001
2012-10-30 04:16 | 25.9288 | 277.2395 | 214.3745 | 273.2645 | 190.5614 | 78.1704 | 1.5298 | 38.8955 | 59.5023 | 25.4253 | 2.4179 | 38.7871 | 93.7821 | 16.0418 | 13.0201 | 5.3891 | 153.2843 | 264.2194 | 201.3544 | 260.2444 | 177.5413
2012-10-30 04:16 | 25.9198 | 277.4938 | 214.5978 | 273.5269 | 190.7998 | 78.2021 | 1.5298 | 38.8955 | 59.5007 | 25.426 | 2.4179 | 38.787 | 93.7825 | 16.0417 | 13.0396 | 5.3913 | 153.2831 | 264.4542 | 201.5582 | 260.4873 | 177.7602
2012-10-30 04:16 | 25.9408 | 277.867 | 214.9063 | 273.92 | 191.0525 | 78.2513 | 1.5297 | 38.8952 | 59.4983 | 25.4266 | 2.4179 | 38.7868 | 93.7838 | 16.0413 | 13.0158 | 5.3946 | 153.2821 | 264.8512 | 201.8905 | 260.9042 | 178.0367
2012-10-30 04:16 | 25.9498 | 278.1933 | 215.1437 | 274.2817 | 191.2383 | 78.2906 | 1.5296 | 38.8942 | 59.494 | 25.427 | 2.4179 | 38.7855 | 93.7793 | 16.041 | 13.0655 | 5.3974 | 153.2733 | 265.1278 | 202.0782 | 261.2162 | 178.1728
2012-10-30 04:16 | 25.9438 | 278.3924 | 215.3464 | 274.5122 | 191.4444 | 78.3241 | 1.5296 | 38.8934 | 59.4899 | 25.4278 | 2.4179 | 38.7848 | 93.7768 | 16.0409 | 13.031 | 5.3997 | 153.2667 | 265.3614 | 202.3154 | 261.4812 | 178.4134
2012-10-30 04:16 | 25.9587 | 278.6782 | 215.5582 | 274.7518 | 191.6714 | 78.3299 | 1.5295 | 38.8925 | 59.4858 | 25.4284 | 2.4179 | 38.7839 | 93.7739 | 16.0406 | 13.0459 | 5.4001 | 153.2596 | 265.6323 | 202.5123 | 261.7059 | 178.6255
2012-10-30 04:16 | 25.964 | 278.9388 | 215.7919 | 275.0361 | 191.9164 | 78.3877 | 1.5294 | 38.8917 | 59.4819 | 25.429 | 2.4178 | 38.7831 | 93.7708 | 16.0404 | 13.0553 | 5.404 | 153.2527 | 265.8835 | 202.7366 | 261.9808 | 178.8611
2012-10-30 04:16 | 25.9745 | 279.2627 | 216.0977 | 275.2869 | 192.1744 | 78.4114 | 1.5294 | 38.8921 | 59.4817 | 25.4295 | 2.4179 | 38.7834 | 93.7746 | 16.0401 | 13.0454 | 5.4057 | 153.2563 | 266.2173 | 203.0523 | 262.2415 | 179.129
2012-10-30 04:16 | 25.9625 | 279.506 | 216.3102 | 275.5732 | 192.4217 | 78.4281 | 1.5294 | 38.8921 | 59.4802 | 25.4303 | 2.4179 | 38.7834 | 93.7751 | 16.04 | 13.0416 | 5.4068 | 153.2554 | 266.4644 | 203.2686 | 262.5316 | 179.3801
2012-10-30 04:16 | 25.9737 | 279.7841 | 216.4986 | 275.8207 | 192.6136 | 78.4778 | 1.5293 | 38.8923 | 59.4792 | 25.4309 | 2.418 | 38.7836 | 93.7772 | 16.0398 | 13.0039 | 5.4103 | 153.2565 | 266.7802 | 203.4947 | 262.8168 | 179.6097
2012-10-30 04:16 | 25.9782 | 280.0477 | 216.6803 | 276.0537 | 192.7281 | 78.4992 | 1.5293 | 38.8927 | 59.4798 | 25.4312 | 2.418 | 38.7841 | 93.7815 | 16.0394 | 13.041 | 5.4117 | 153.2613 | 267.0067 | 203.6393 | 263.0127 | 179.6871
2012-10-30 04:16 | 25.9804 | 280.3321 | 216.9187 | 276.3386 | 192.9032 | 78.5281 | 1.5293 | 38.8926 | 59.4771 | 25.4322 | 2.4181 | 38.7839 | 93.7821 | 16.0392 | 13.0228 | 5.4137 | 153.2592 | 267.3093 | 203.8959 | 263.3158 | 179.8804
2012-10-30 04:16 | 25.9894 | 280.5963 | 217.227 | 276.7 | 193.0733 | 78.5732 | 1.5292 | 38.8927 | 59.4759 | 25.4328 | 2.4181 | 38.7838 | 93.7833 | 16.0389 | 13.044 | 5.4168 | 153.2592 | 267.5523 | 204.183 | 263.656 | 180.0293
2012-10-30 04:16 | 25.9917 | 280.8731 | 217.4458 | 276.9734 | 193.3232 | 78.6038 | 1.5292 | 38.8916 | 59.4712 | 25.4334 | 2.4181 | 38.7827 | 93.7797 | 16.0387 | 13.0301 | 5.4189 | 153.2509 | 267.843 | 204.4157 | 263.9433 | 180.2931
2012-10-30 04:16 | 26.0014 | 281.1921 | 217.6721 | 277.2502 | 193.5097 | 78.6269 | 1.5291 | 38.8913 | 59.4692 | 25.4339 | 2.4181 | 38.7824 | 93.7801 | 16.0384 | 13.0438 | 5.4205 | 153.2493 | 268.1483 | 204.6283 | 264.2064 | 180.4659
2012-10-30 04:16 | 26.0089 | 281.4006 | 217.8489 | 277.4861 | 193.6391 | 78.6674 | 1.5291 | 38.8915 | 59.4685 | 25.4345 | 2.4181 | 38.7827 | 93.7821 | 16.0382 | 13.0432 | 5.4233 | 153.2506 | 268.3574 | 204.8057 | 264.4429 | 180.5959
2012-10-30 04:16 | 26.0029 | 281.6419 | 218.0398 | 277.7317 | 193.9632 | 78.694 | 1.5291 | 38.892 | 59.469 | 25.4349 | 2.4182 | 38.7831 | 93.7857 | 16.0379 | 13.0087 | 5.4252 | 153.2547 | 268.6332 | 205.0311 | 264.723 | 180.9545
2012-10-30 04:16 | 26.0051 | 281.8567 | 218.2152 | 278.001 | 194.1737 | 78.7211 | 1.529 | 38.8917 | 59.466 | 25.4358 | 2.4182 | 38.7829 | 93.7857 | 16.0378 | 13.0231 | 5.427 | 153.2516 | 268.8336 | 205.1921 | 264.9779 | 181.1506
2012-10-30 04:16 | 26.0126 | 282.1308 | 218.4589 | 278.2716 | 194.4603 | 78.7593 | 1.529 | 38.8919 | 59.4658 | 25.4362 | 2.4183 | 38.7831 | 93.7881 | 16.0375 | 13.0102 | 5.4297 | 153.2538 | 269.1206 | 205.4487 | 265.2614 | 181.4501
2012-10-30 04:16 | 26.0149 | 282.392 | 218.6817 | 278.5525 | 194.6002 | 78.7887 | 1.529 | 38.8927 | 59.4661 | 25.437 | 2.4183 | 38.7838 | 93.7918 | 16.0374 | 13.0533 | 5.4317 | 153.2579 | 269.3387 | 205.6284 | 265.4992 | 181.5469
2012-10-30 04:17 | 26.0209 | 282.6104 | 218.8924 | 278.7636 | 194.9049 | 78.809 | 1.529 | 38.8936 | 59.468 | 25.4375 | 2.4184 | 38.7846 | 93.7975 | 16.0371 | 13.0878 | 5.4331 | 153.2655 | 269.5226 | 205.8046 | 265.6758 | 181.8171
2012-10-30 04:17 | 26.0268 | 282.8481 | 219.0558 | 278.9746 | 195.186 | 78.8332 | 1.529 | 38.8943 | 59.4692 | 25.4379 | 2.4185 | 38.7854 | 93.8017 | 16.0371 | 13.0285 | 5.4348 | 153.2709 | 269.8196 | 206.0273 | 265.9461 | 182.1575
2012-10-30 04:17 | 26.0366 | 283.0781 | 219.3137 | 279.2552 | 195.2862 | 78.8575 | 1.529 | 38.8954 | 59.4708 | 25.4386 | 2.4186 | 38.7864 | 93.8084 | 16.0368 | 13.075 | 5.4364 | 153.2791 | 270.0031 | 206.2387 | 266.1802 | 182.2112
2012-10-30 04:17 | 26.0381 | 283.2998 | 219.5277 | 279.4578 | 195.3703 | 78.8933 | 1.529 | 38.8962 | 59.4723 | 25.4389 | 2.4187 | 38.7872 | 93.8133 | 16.0366 | 13.0724 | 5.4389 | 153.2856 | 270.2274 | 206.4553 | 266.3854 | 182.2979
2012-10-30 04:17 | 26.0373 | 283.5191 | 219.6687 | 279.6388 | 195.4756 | 78.9355 | 1.529 | 38.8961 | 59.4709 | 25.4395 | 2.4187 | 38.7871 | 93.8139 | 16.0365 | 13.1123 | 5.4418 | 153.2848 | 270.4068 | 206.5564 | 266.5265 | 182.3633
2012-10-30 04:17 | 26.0493 | 283.7396 | 219.9051 | 279.9293 | 195.6999 | 78.9581 | 1.5289 | 38.8964 | 59.4702 | 25.4401 | 2.4188 | 38.7873 | 93.8173 | 16.036 | 13.0958 | 5.4434 | 153.2875 | 270.6438 | 206.8093 | 266.8335 | 182.6041
2012-10-30 04:17 | 26.0448 | 283.8741 | 220.1289 | 280.1104 | 196.01 | 78.9714 | 1.5289 | 38.8967 | 59.4704 | 25.4404 | 2.4188 | 38.7876 | 93.8184 | 16.0361 | 13.0709 | 5.4443 | 153.2888 | 270.8032 | 207.058 | 267.0395 | 182.9391
2012-10-30 04:17
2012-10-30 04:19
2012-10-30 04:21
2012-10-30 04:24
2012-10-30 04:26
2012-10-30 04:29
2012-10-30 04:31
2012-10-30 04:34
2012-10-30 04:36
2012-10-30 04:39
2012-10-30 04:42
2012-10-30 04:45
2012-10-30 04:47
2012-10-30 04:50
2012-10-30 04:53
2012-10-30 04:55
2012-10-30 04:58
2012-10-30 05:01
2012-10-30 05:03
2012-10-30 05:06
2012-10-30 05:09
2012-10-30 05:11
2012-10-30 05:14
2012-10-30 05:17
2012-10-30 05:20
2012-10-30 05:22
2012-10-30 05:25
2012-10-30 05:27
2012-10-30 05:30
2012-10-30 05:32
2012-10-30 05:35
2012-10-30 05:37
2012-10-30 05:39
2012-10-30 05:42
2012-10-30 05:44
2012-10-30 05:47
2012-10-30 05:49
2012-10-30 05:51
2012-10-30 05:54
2012-10-30 05:56
2012-10-30 05:59
2012-10-30 06:01
2012-10-30 06:03
2012-10-30 06:06
2012-10-30 06:08
2012-10-30 06:11
2012-10-30 06:13
2012-10-30 06:15
2012-10-30 06:18
2012-10-30 06:20
2012-10-30 06:23
2012-10-30 06:25
2012-10-30 06:27
2012-10-30 06:30
2012-10-30 06:32
2012-10-30 06:34
2012-10-30 06:37
2012-10-30 06:39
2012-10-30 06:42
2012-10-30 06:44
2012-10-30 06:46
2012-10-30 06:49
2012-10-30 06:51
2012-10-30 06:54
2012-10-30 06:56
2012-10-30 06:58
2012-10-30 07:01
2012-10-30 07:03
2012-10-30 07:06
2012-10-30 07:08
2012-10-30 07:10
2012-10-30 07:13
2012-10-30 07:15
2012-10-30 07:18
2012-10-30 07:20
2012-10-30 07:23
2012-10-30 07:25
2012-10-30 07:27
2012-10-30 07:29
2012-10-30 07:32
2012-10-30 07:34
2012-10-30 07:37
2012-10-30 07:39
2012-10-30 07:41
2012-10-30 07:44
2012-10-30 07:46
2012-10-30 07:49
2012-10-30 07:51
2012-10-30 07:54
2012-10-30 07:56
2012-10-30 07:58
2012-10-30 08:01
2012-10-30 08:03
2012-10-30 08:06
2012-10-30 08:08
2012-10-30 08:10
2012-10-30 08:13
2012-10-30 08:15
2012-10-30 08:18
2012-10-30 08:20
2012-10-30 08:22
2012-10-30 08:25
2012-10-30 08:27
2012-10-30 08:30
2012-10-30 08:32
2012-10-30 08:34
2012-10-30 08:37
2012-10-30 08:39
2012-10-30 08:42
2012-10-30 08:44
2012-10-30 08:46
2012-10-30 08:49
2012-10-30 08:51
2012-10-30 08:53
2012-10-30 08:56
2012-10-30 08:58
2012-10-30 09:01
2012-10-30 09:03
2012-10-30 09:05
2012-10-30 09:08
2012-10-30 09:10
2012-10-30 09:13
2012-10-30 09:15
2012-10-30 09:17
2012-10-30 09:20
2012-10-30 09:22
2012-10-30 09:25
2012-10-30 09:27
2012-10-30 09:29
2012-10-30 09:32
2012-10-30 09:34
2012-10-30 09:37
2012-10-30 09:39
2012-10-30 09:41
2012-10-30 09:44
2012-10-30 09:46
2012-10-30 09:48
2012-10-30 09:51
2012-10-30 09:53
2012-10-30 09:56
2012-10-30 09:58
2012-10-30 10:00
2012-10-30 10:03
2012-10-30 10:05
2012-10-30 10:08
2012-10-30 10:10
2012-10-30 10:12
2012-10-30 10:15
2012-10-30 10:17
2012-10-30 10:20
2012-10-30 10:22
2012-10-30 10:24
2012-10-30 10:27
2012-10-30 10:29
2012-10-30 10:32
2012-10-30 10:34
2012-10-30 10:37
2012-10-30 10:39
2012-10-30 10:41
2012-10-30 10:44
2012-10-30 10:46
2012-10-30 10:48
2012-10-30 10:51
2012-10-30 10:53
2012-10-30 10:56
2012-10-30 10:58
2012-10-30 11:00
2012-10-30 11:03
2012-10-30 11:05
2012-10-30 11:08
2012-10-30 11:10
2012-10-30 11:12
2012-10-30 11:15
2012-10-30 11:17
2012-10-30 11:20
2012-10-30 11:22
2012-10-30 11:24
2012-10-30 11:27
2012-10-30 11:29
2012-10-30 11:31
2012-10-30 11:34
2012-10-30 11:36
2012-10-30 11:39
2012-10-30 11:41
2012-10-30 11:44
2012-10-30 11:46
2012-10-30 11:48
2012-10-30 11:51
2012-10-30 11:53
2012-10-30 11:56
2012-10-30 11:58
2012-10-30 12:00
2012-10-30 12:03
2012-10-30 12:05
2012-10-30 12:07
2012-10-30 12:10
2012-10-30 12:12
2012-10-30 12:15
2012-10-30 12:17
2012-10-30 12:19
2012-10-30 12:22
2012-10-30 12:24
2012-10-30 12:27
2012-10-30 12:29
2012-10-30 12:31
2012-10-30 12:34
2012-10-30 12:36
2012-10-30 12:39
2012-10-30 12:41
2012-10-30 12:43
2012-10-30 12:46
2012-10-30 12:48
2012-10-30 12:50
2012-10-30 12:53
2012-10-30 12:55
2012-10-30 12:58
2012-10-30 13:00
2012-10-30 13:02
2012-10-30 13:05
2012-10-30 13:07
2012-10-30 13:10
2012-10-30 13:12
2012-10-30 13:14
2012-10-30 13:17
2012-10-30 13:19
2012-10-30 13:22
2012-10-30 13:24
2012-10-30 13:26
2012-10-30 13:29
2012-10-30 13:31
2012-10-30 13:33
2012-10-30 13:36
2012-10-30 13:38 |  | -8.4566 | -6.8533 | -9.667 | -9.7642 | 47.7425 | 1.5622 | 37.9083 | 59.2187 | 24.2666 | 2.4099 | 38.5746 | 92.9595 | 16.007 | -9.1769 | 3.2914 | 152.1782 | 0.7203 | 2.3236 | -0.4901 | -0.5873
2012-10-30 13:39 | 20.589 | 17.1594 | 16.5786 | 11.404 | 11.3838 | 48.247 | 1.5603 | 38.6801 | 60.3536 | 24.7898 | 2.4099 | 38.5726 | 92.9566 | 16.0058 | 11.3838 | 3.3261 | 153.3102 | 5.7756 | 5.1948 | 0.0202 | 0
2012-10-30 13:39 | 20.5942 |  | 18.188 | 12.1498 | 12.2185 | 48.5626 | 1.5598 | 38.6794 | 60.3316 | 24.7978 | 2.4099 | 38.5719 | 92.9543 | 16.0056 | 11.4295 | 3.3479 | 153.2859 | -11.4295 | 6.7585 | 0.7203 | 0.789
2012-10-30 13:39 | 20.6024 | 23.3936 | 19.336 | 13.0268 | 12.9137 | 48.803 | 1.5594 | 38.6797 | 60.3178 | 24.804 | 2.41 | 38.5723 | 92.9583 | 16.0053 | 11.4255 | 3.3645 | 153.2761 | 11.9681 | 7.9105 | 1.6013 | 1.4882
2012-10-30 13:39 | 20.5979 | 25.7935 |  | 13.9099 | 13.7123 | 48.9851 | 1.5591 | 38.6806 | 60.3072 | 24.8095 | 2.4101 | 38.5731 | 92.9637 | 16.005 | 11.4453 | 3.377 | 153.2709 | 14.3482 | -11.4453 | 2.4646 | 2.267
2012-10-30 13:39 | 20.6076 | 28.0804 | 21.4208 | 14.9184 | 14.5199 | 49.1694 | 1.5587 | 38.6813 | 60.2932 | 24.8161 | 2.4101 | 38.5738 | 92.9682 | 16.0048 | 11.455 | 3.3897 | 153.2613 | 16.6254 | 9.9658 | 3.4634 | 3.0649
2012-10-30 13:39 | 20.6129 | 30.8419 | 22.7605 | 16.1091 | 15.5228 | 49.415 | 1.5584 | 38.6825 | 60.2811 | 24.8226 | 2.4102 | 38.5751 | 92.9757 | 16.0046 | 11.4441 | 3.4067 | 153.2567 | 19.3978 | 11.3164 | 4.665 | 4.0787
2012-10-30 13:39 | 20.6099 | 33.4577 | 24.0131 | 17.4764 | 16.4792 | 49.6156 | 1.558 | 38.684 | 60.2709 | 24.8288 | 2.4104 | 38.5765 | 92.9842 | 16.0043 | 11.4776 | 3.4205 | 153.2551 | 21.9801 | 12.5355 | 5.9988 | 5.0016
2012-10-30 13:39 | 20.6159 | 36.043 | 25.4128 | 19.1226 | 17.5104 | 49.8317 | 1.5577 | 38.6857 | 60.2611 | 24.8351 | 2.4105 | 38.5781 | 92.9935 | 16.0041 | 11.4957 | 3.4354 | 153.2545 | 24.5473 | 13.9171 | 7.6269 | 6.0147
2012-10-30 13:39 | 20.6301 | 38.3808 | 26.6428 |  | 18.567 | 50.0323 | 1.5574 | 38.6868 | 60.2512 | 24.8405 | 2.4106 | 38.5793 | 93 | 16.0039 | 11.5342 | 3.4492 | 153.2512 | 26.8466 | 15.1086 | -11.5342 | 7.0328
2012-10-30 13:39 | 20.6316 | 40.4238 | 27.7887 | 21.9777 | 19.4805 | 50.223 | 1.5572 | 38.6881 | 60.2433 | 24.8454 | 2.4107 | 38.5804 | 93.0069 | 16.0036 | 11.5357 | 3.4624 | 153.2503 | 28.8881 | 16.253 | 10.442 | 7.9448
2012-10-30 13:39 | 20.633 | 42.643 | 29.1747 | 23.902 |  | 50.4397 | 1.5569 | 38.6899 | 60.2344 | 24.8514 | 2.4109 | 38.5821 | 93.0159 | 16.0035 | 11.4926 | 3.4773 | 153.2503 | 31.1504 | 17.6821 | 12.4094 | -11.4926
2012-10-30 13:39 | 20.636 | 44.8789 | 30.3802 | 25.9737 | 21.5913 | 50.6506 | 1.5566 | 38.6912 | 60.2251 | 24.8569 | 2.411 | 38.5834 | 93.0231 | 16.0034 | 11.4915 | 3.4919 | 153.2482 | 33.3874 | 18.8887 | 14.4822 | 10.0998
2012-10-30 13:39 | 20.6465 | 46.9972 | 31.6517 | 27.7799 | 22.6401 | 50.8367 | 1.5563 | 38.6927 | 60.2163 | 24.8624 | 2.4111 | 38.5848 | 93.0312 | 16.0031 | 11.5182 | 3.5047 | 153.2475 | 35.479 | 20.1335 | 16.2617 | 11.1219
2012-10-30 13:39 | 20.6472 | 49.3617 | 33.0227 | 29.5833 | 23.7464 | 51.0124 | 1.556 | 38.6941 | 60.2077 | 24.8678 | 2.4112 | 38.586 | 93.0377 | 16.003 | 11.5068 | 3.5168 | 153.2453 | 37.8549 | 21.5159 | 18.0765 | 12.2396
2012-10-30 13:39 | 20.6592 | 51.6277 | 34.4118 | 31.5541 | 25.0021 | 51.2072 | 1.5557 | 38.6956 | 60.1999 | 24.8729 | 2.4113 | 38.5876 | 93.0465 | 16.0028 | 11.5512 | 3.5302 | 153.2464 | 40.0765 | 22.8606 | 20.0029 | 13.4509
2012-10-30 13:39 | 20.6682 | 53.7544 | 35.6828 | 33.7206 | 26.223 | 51.4008 | 1.5555 | 38.6969 | 60.1923 | 24.8778 | 2.4114 | 38.5888 | 93.0541 | 16.0025 | 11.548 | 3.5436 | 153.2464 | 42.2064 | 24.1348 | 22.1726 | 14.675
2012-10-30 13:39 | 20.6749 | 56.0096 | 36.9581 | 35.9291 | 27.248 | 51.595 | 1.5552 | 38.6985 | 60.1849 | 24.8829 | 2.4115 | 38.5903 | 93.0615 | 16.0024 | 11.5345 | 3.557 | 153.2464 | 44.4751 | 25.4236 | 24.3946 | 15.7135
2012-10-30 13:39 | 20.6861 | 60.7577 | 39.9379 | 40.6153 | 29.8784 | 52.0296 | 1.5546 | 38.7009 | 60.1661 | 24.8938 | 2.4117 | 38.5928 | 93.0751 | 16.0021 | 11.5659 | 3.5869 | 153.2412 | 49.1918 | 28.372 | 29.0494 | 18.3125
2012-10-30 13:39 | 20.7085 | 65.2143 | 42.9334 | 45.2672 | 32.5447 | 52.4572 | 1.5541 | 38.7035 | 60.1483 | 24.9044 | 2.4119 | 38.595 | 93.0881 | 16.0018 | 11.5924 | 3.6164 | 153.2364 | 53.6219 | 31.341 | 33.6748 | 20.9523
2012-10-30 13:39 | 20.716 | 67.0965 | 44.1748 | 47.6076 | 33.6189 | 52.6335 | 1.5538 | 38.7045 | 60.1393 | 24.9094 | 2.412 | 38.5959 | 93.094 | 16.0015 | 11.5594 | 3.6286 | 153.2334 | 55.5371 | 32.6154 | 36.0482 | 22.0595
2012-10-30 13:39 | 20.7152 | 69.151 | 45.4026 | 49.8801 | 34.7353 | 52.8508 | 1.5536 | 38.7053 | 60.1308 | 24.914 | 2.4121 | 38.5968 | 93.098 | 16.0016 | 11.5788 | 3.6435 | 153.2288 | 57.5722 | 33.8238 | 38.3013 | 23.1565
2012-10-30 13:39 | 20.7264 | 71.0239 | 46.7189 | 51.9119 | 35.792 | 52.9791 | 1.5533 | 38.7059 | 60.1235 | 24.9178 | 2.4121 | 38.5973 | 93.1008 | 16.0015 | 11.5698 | 3.6524 | 153.2242 | 59.4541 | 35.1491 | 40.3421 | 24.2222
2012-10-30 13:39 | 20.7444 | 72.6804 | 47.8966 | 53.9234 | 36.9059 | 53.1207 | 1.5531 | 38.7064 | 60.1149 | 24.922 | 2.4122 | 38.5979 | 93.1044 | 16.0014 | 11.5473 | 3.6621 | 153.2193 | 61.1331 | 36.3493 | 42.3761 | 25.3586
2012-10-30 13:39 | 20.7451 | 74.5234 | 49.1644 | 55.9293 | 37.9898 | 53.3126 | 1.321 | 32.9477 | 43.5236 | 24.9417 | 2.1762 | 34.8309 | 75.7997 | 16.0053 | 11.5642 | 3.6754 | 119.3233 | 62.9592 | 37.6002 | 44.3651 | 26.4256
2012-10-30 13:39 | 20.7444 | 76.0399 | 49.6181 | 57.8993 | 39.0911 | 51.9857 | 0.007682 | 0.32825 | 0.002522 | 42.7295 | 0.007877 | 0.2147 | 0.001691 | 27.2566 | 11.5433 | 3.5839 | 0.004213 | 64.4966 | 38.0748 | 46.356 | 27.5478
2012-10-30 13:39 | 20.7593 | 75.7051 | 48.1448 | 59.7428 | 40.1325 | 50.275 | 0.007685 | 0.32748 | 0.002517 | 42.6131 | 0.007884 | 0.21494 | 0.001695 | 27.2634 | 11.5946 | 3.466 | 0.004212 | 64.1105 | 36.5502 | 48.1482 | 28.5379
2012-10-30 13:40 | 20.7675 | 74.9147 | 47.6707 | 61.1311 | 40.7515 | 50.1808 | 0.007695 | 0.32732 | 0.002519 | 42.5371 | 0.007884 | 0.21499 | 0.001695 | 27.269 | 11.5299 | 3.4595 | 0.004214 | 63.3848 | 36.1408 | 49.6012 | 29.2216
2012-10-30 13:40 | 20.7712 | 74.0358 | 47.5655 | 62.3061 | 41.2092 | 50.0831 | 0.007701 | 0.32741 | 0.002521 | 42.5153 | 0.007884 | 0.21502 | 0.001695 | 27.2727 | 11.6065 | 3.4527 | 0.004216 | 62.4293 | 35.959 | 50.6996 | 29.6027
2012-10-30 13:40 | 20.7802 | 73.1275 | 47.6134 | 63.7842 | 41.672 | 50.0427 | 0.007704 | 0.32741 | 0.002522 | 42.4987 | 0.007887 | 0.21503 | 0.001696 | 27.2641 | 11.6155 | 3.4499 | 0.004218 | 61.512 | 35.9979 | 52.1687 | 30.0565
2012-10-30 13:40 | 20.7854 | 72.4281 | 47.7081 | 65.0844 | 42.0488 | 49.9849 | 0.007701 | 0.32741 | 0.002521 | 42.5153 | 0.007915 | 0.21507 | 0.001702 | 27.1731 | 11.6328 | 3.446 | 0.004223 | 60.7953 | 36.0753 | 53.4516 | 30.416
2012-10-30 13:40 | 20.7929 | 71.8146 | 47.7778 | 66.0489 | 42.3456 | 49.9375 | 0.00771 | 0.32751 | 0.002525 | 42.4787 | 0.007902 | 0.2151 | 0.0017 | 27.2215 | 11.6768 | 3.4427 | 0.004225 | 60.1378 | 36.101 | 54.3721 | 30.6688
2012-10-30 13:40 | 20.7952 | 71.3138 | 47.8618 | 66.8201 | 42.5472 | 49.9149 | 0.007704 | 0.3276 | 0.002524 | 42.5231 | 0.007915 | 0.21522 | 0.001703 | 27.1914 | 11.6669 | 3.4411 | 0.004227 | 59.6469 | 36.1949 | 55.1532 | 30.8803
2012-10-30 13:40 | 20.8004 | 70.7853 | 47.9292 | 67.4638 | 42.7791 | 49.8641 | 0.007723 | 0.32754 | 0.00253 | 42.411 | 0.007909 | 0.21519 | 0.001702 | 27.2084 | 11.6235 | 3.4376 | 0.004232 | 59.1618 | 36.3057 | 55.8403 | 31.1556
2012-10-30 13:40 | 20.8086 | 70.229 | 48.0113 | 68.0185 | 42.9084 | 49.8577 | 0.007714 | 0.32757 | 0.002527 | 42.4642 | 0.007912 | 0.21522 | 0.001703 | 27.2017 | 11.6601 | 3.4372 | 0.00423 | 58.5689 | 36.3512 | 56.3584 | 31.2483
2012-10-30 13:40 | 20.831 | 69.7477 | 48.0843 | 68.5757 | 43.0793 | 49.8092 | 0.007704 | 0.32753 | 0.002523 | 42.5136 | 0.007921 | 0.21525 | 0.001705 | 27.1745 | 11.6744 | 3.4338 | 0.004228 | 58.0733 | 36.4099 | 56.9013 | 31.4049
2012-10-30 13:40 | 20.8362 | 69.2491 | 48.0973 | 69.001 | 43.2642 | 49.7757 | 0.007723 | 0.32763 | 0.00253 | 42.4222 | 0.007902 | 0.21522 | 0.001701 | 27.2361 | 11.6755 | 3.4315 | 0.004231 | 57.5736 | 36.4218 | 57.3255 | 31.5887
2012-10-30 13:40 | 20.8362 | 68.8559 | 48.1128 | 69.2911 | 43.3619 | 49.7306 | 0.007729 | 0.32761 | 0.002532 | 42.3874 | 0.007921 | 0.21515 | 0.001704 | 27.1617 | 11.7039 | 3.4284 | 0.004236 | 57.152 | 36.4089 | 57.5872 | 31.658
2012-10-30 13:40 | 20.8557 | 68.4285 | 48.1595 | 69.4862 | 43.5063 | 49.7312 | 0.007714 | 0.32761 | 0.002527 | 42.4698 | 0.007912 | 0.21526 | 0.001703 | 27.2072 | 11.691 | 3.4285 | 0.00423 | 56.7375 | 36.4685 | 57.7952 | 31.8153
2012-10-30 13:40 | 20.8616 | 67.9874 | 48.1693 | 69.6524 | 43.5942 | 49.6919 | 0.007739 | 0.32768 | 0.002536 | 42.3419 | 0.007924 | 0.21531 | 0.001706 | 27.1715 | 11.6726 | 3.4258 | 0.004242 | 56.3148 | 36.4967 | 57.9798 | 31.9216
2012-10-30 13:40 | 20.8654 | 67.5173 | 48.1537 | 69.6982 | 43.6566 | 49.641 | 0.007732 | 0.32767 | 0.002534 | 42.3784 | 0.007921 | 0.21526 | 0.001705 | 27.1762 | 11.644 | 3.4223 | 0.004239 | 55.8733 | 36.5097 | 58.0542 | 32.0126
2012-10-30 13:40 | 20.8811 | 67.1201 | 48.1616 | 69.7559 | 43.7426 | 49.6121 | 0.007742 | 0.32761 | 0.002536 | 42.3162 | 0.007928 | 0.21525 | 0.001706 | 27.1505 | 11.6961 | 3.4203 | 0.004242 | 55.424 | 36.4655 | 58.0598 | 32.0465
2012-10-30 13:40 | 20.887 | 66.7513 | 48.1908 | 69.716 | 43.8578 | 49.5572 | 0.007723 | 0.32761 | 0.00253 | 42.4203 | 0.007918 | 0.21526 | 0.001704 | 27.1865 | 11.6737 | 3.4165 | 0.004234 | 55.0776 | 36.5171 | 58.0423 | 32.1841
2012-10-30 13:40 | 20.899 | 66.4075 | 48.1678 | 69.6822 | 43.8854 | 49.5607 | 0.007732 | 0.3276 | 0.002533 | 42.3691 | 0.007924 | 0.21533 | 0.001706 | 27.175 | 11.7181 | 3.4167 | 0.004239 | 54.6894 | 36.4497 | 57.9641 | 32.1673
2012-10-30 13:40 | 20.9109 | 66.033 | 48.1681 | 69.5452 | 43.9286 | 49.5433 | 0.007742 | 0.32766 | 0.002537 | 42.3219 | 0.007928 | 0.21529 | 0.001707 | 27.1559 | 11.734 | 3.4155 | 0.004244 | 54.299 | 36.4341 | 57.8112 | 32.1946
2012-10-30 13:40 | 20.9139 | 65.6226 | 48.1166 | 69.4146 | 43.9316 | 49.4965 | 0.007739 | 0.3276 | 0.002535 | 42.3308 | 0.007909 | 0.21532 | 0.001703 | 27.2248 | 11.6763 | 3.4123 | 0.004238 | 53.9463 | 36.4403 | 57.7383 | 32.2553
2012-10-30 13:40 | 20.9319 | 65.2692 | 48.0763 | 69.2074 | 43.9457 | 49.4827 | 0.007723 | 0.32768 | 0.002531 | 42.4296 | 0.007928 | 0.21529 | 0.001707 | 27.1559 | 11.755 | 3.4113 | 0.004238 | 53.5142 | 36.3213 | 57.4524 | 32.1907
2012-10-30 13:40 | 20.9311 | 64.9773 | 48.0288 | 69.031 | 43.9527 | 49.4653 | 0.007736 | 0.32764 | 0.002535 | 42.3528 | 0.007924 | 0.21535 | 0.001706 | 27.1769 | 11.6895 | 3.4101 | 0.004241 | 53.2878 | 36.3393 | 57.3415 | 32.2632
2012-10-30 13:40 | 20.949 | 64.6811 | 47.9962 | 68.8579 | 44.0721 | 49.4318 | 0.007736 | 0.32764 | 0.002535 | 42.3528 | 0.007924 | 0.21532 | 0.001706 | 27.1733 | 11.7114 | 3.4078 | 0.004241 | 52.9697 | 36.2848 | 57.1465 | 32.3607
2012-10-30 13:40 | 20.9617 | 64.2838 | 48.005 | 68.6376 | 44.0575 | 49.4301 | 0.007723 | 0.32761 | 0.00253 | 42.4203 | 0.007928 | 0.21528 | 0.001707 | 27.1541 | 11.7282 | 3.4077 | 0.004237 | 52.5556 | 36.2768 | 56.9094 | 32.3293
2012-10-30 13:40 | 20.9632 | 63.9518 | 47.9909 | 68.3336 | 44.0473 | 49.3937 | 0.007729 | 0.32766 | 0.002532 | 42.3931 | 0.007918 | 0.21538 | 0.001705 | 27.2012 | 11.7175 | 3.4052 | 0.004237 | 52.2343 | 36.2734 | 56.6161 | 32.3298
2012-10-30 13:40 | 20.9722 | 63.6043 | 47.9143 | 68.0942 | 44.0367 | 49.363 | 0.007742 | 0.32763 | 0.002536 | 42.3181 | 0.007943 | 0.21533 | 0.00171 | 27.11 | 11.7549 | 3.4031 | 0.004246 | 51.8494 | 36.1594 | 56.3393 | 32.2818
2012-10-30 13:40 | 20.9939 | 63.3499 | 47.8816 | 67.8064 | 44.0467 | 49.3619 | 0.007748 | 0.3276 | 0.002538 | 42.2816 | 0.007918 | 0.21535 | 0.001705 | 27.1975 | 11.6956 | 3.403 | 0.004243 | 51.6543 | 36.186 | 56.1108 | 32.3511
2012-10-30 13:40 | 21.0036 | 62.9913 | 47.8368 | 67.5217 | 44.0525 | 49.3243 | 0.007742 | 0.32757 | 0.002536 | 42.3106 | 0.007924 | 0.21529 | 0.001706 | 27.1696 | 11.7498 | 3.4004 | 0.004242 | 51.2415 | 36.087 | 55.7719 | 32.3027
2012-10-30 13:40 | 21.0096 | 62.6596 | 47.8156 | 67.1798 | 44.0234 | 49.2948 | 0.007742 | 0.32757 | 0.002536 | 42.3106 | 0.007928 | 0.21531 | 0.001707 | 27.1578 | 11.772 | 3.3984 | 0.004243 | 50.8876 | 36.0436 | 55.4078 | 32.2514
2012-10-30 13:40 |  | 69.2125 | 54.6071 | 72.3587 | 51.1248 | 50.7477 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014363 | 0.23506 | 0.003376 | 16.3655 | 19.3289 | 3.4985 | 0.008282 | 49.8836 | 35.2782 | 53.0298 | 31.7959
2012-10-30 13:41 | 19.0647 | 68.3477 | 54.3878 | 71.431 | 50.9977 | 50.6691 | 0.014181 | 0.34722 | 0.004924 | 24.4847 | 0.014376 | 0.23506 | 0.003379 | 16.3507 | 19.3528 | 3.4931 | 0.008303 | 48.9949 | 35.035 | 52.0782 | 31.6449
2012-10-30 13:41 | 19.0924 | 68.0816 | 54.3577 | 71.0894 | 51.0215 | 50.6379 | 0.01419 | 0.34722 | 0.004927 | 24.4692 | 0.014373 | 0.23503 | 0.003378 | 16.352 | 19.3371 | 3.491 | 0.008305 | 48.7445 | 35.0206 | 51.7523 | 31.6844
2012-10-30 13:41 | 19.0931 | 67.7923 | 54.3043 | 70.7816 | 50.9448 | 50.6125 | 0.014184 | 0.34722 | 0.004925 | 24.4796 | 0.014376 | 0.23503 | 0.003379 | 16.3486 | 19.3734 | 3.4892 | 0.008304 | 48.4189 | 34.9309 | 51.4082 | 31.5714
2012-10-30 13:41 | 19.108 | 67.5362 | 54.1918 | 70.4803 | 50.9365 | 50.6142 | 0.014187 | 0.34719 | 0.004926 | 24.4723 | 0.01437 | 0.23507 | 0.003378 | 16.3585 | 19.3527 | 3.4893 | 0.008304 | 48.1835 | 34.8391 | 51.1276 | 31.5838
2012-10-30 13:41 | 19.1155 | 67.2841 | 54.1144 | 70.1791 | 50.8626 | 50.5744 | 0.01419 | 0.34722 | 0.004927 | 24.4692 | 0.014376 | 0.23498 | 0.003378 | 16.3456 | 19.3484 | 3.4866 | 0.008305 | 47.9357 | 34.766 | 50.8307 | 31.5142
2012-10-30 13:41 | 19.1245 | 67.0411 | 54.0346 | 69.8452 | 50.8058 | 50.5299 | 0.01419 | 0.34718 | 0.004926 | 24.4662 | 0.014379 | 0.23503 | 0.003379 | 16.3452 | 19.3258 | 3.4835 | 0.008305 | 47.7153 | 34.7088 | 50.5194 | 31.48
2012-10-30 13:41 | 19.1342 | 66.8025 | 53.9748 | 69.5461 | 50.7264 | 50.5403 | 0.0142 | 0.34712 | 0.004929 | 24.4449 | 0.014386 | 0.23507 | 0.003382 | 16.3403 | 19.3671 | 3.4842 | 0.008311 | 47.4354 | 34.6077 | 50.179 | 31.3593
2012-10-30 13:41 | 19.1596 | 66.5108 | 53.9037 | 69.2322 | 50.717 | 50.5119 | 0.014172 | 0.34715 | 0.00492 | 24.4952 | 0.014382 | 0.23503 | 0.00338 | 16.3418 | 19.3885 | 3.4823 | 0.0083 | 47.1223 | 34.5152 | 49.8437 | 31.3285
2012-10-30 13:41 | 19.1611 | 66.2755 | 53.7816 | 68.9286 | 50.6255 | 50.4819 | 0.014187 | 0.34722 | 0.004926 | 24.4744 | 0.014373 | 0.2351 | 0.003379 | 16.3571 | 19.3071 | 3.4802 | 0.008305 | 46.9684 | 34.4745 | 49.6215 | 31.3184
2012-10-30 13:41 | 19.1723 | 66.0536 | 53.7078 | 68.6346 | 50.6057 | 50.4692 | 0.014187 | 0.34716 | 0.004925 | 24.4704 | 0.014379 | 0.23507 | 0.00338 | 16.3483 | 19.3223 | 3.4793 | 0.008305 | 46.7313 | 34.3855 | 49.3123 | 31.2834
2012-10-30 13:41 | 19.1775 | 65.7797 | 53.6358 | 68.3499 | 50.5102 | 50.4565 | 0.014206 | 0.34712 | 0.004931 | 24.4345 | 0.014386 | 0.23504 | 0.003381 | 16.3383 | 19.3314 | 3.4785 | 0.008312 | 46.4483 | 34.3044 | 49.0185 | 31.1788
2012-10-30 13:41 | 19.1879 | 65.5837 | 53.538 | 68.0588 | 50.4354 | 50.4091 | 0.014181 | 0.34716 | 0.004923 | 24.4807 | 0.014401 | 0.23501 | 0.003384 | 16.3193 | 19.3537 | 3.4752 | 0.008307 | 46.23 | 34.1843 | 48.7051 | 31.0817
2012-10-30 13:41 | 19.2021 | 65.315 | 53.4595 | 67.7868 | 50.4031 | 50.4247 | 0.014175 | 0.3471 | 0.00492 | 24.487 | 0.014386 | 0.23509 | 0.003382 | 16.3413 | 19.3481 | 3.4763 | 0.008302 | 45.9669 | 34.1114 | 48.4387 | 31.055
2012-10-30 13:41 | 19.2081 | 65.0915 | 53.4307 | 67.5179 | 50.3316 | 50.3912 | 0.01419 | 0.34716 | 0.004926 | 24.4652 | 0.014382 | 0.23498 | 0.00338 | 16.3388 | 19.3897 | 3.474 | 0.008306 | 45.7018 | 34.041 | 48.1282 | 30.9419
2012-10-30 13:41 | 19.223 | 64.8692 | 53.3143 | 67.2732 | 50.3387 | 50.3836 | 0.01419 | 0.34719 | 0.004927 | 24.4672 | 0.014379 | 0.23504 | 0.00338 | 16.3463 | 19.3967 | 3.4734 | 0.008307 | 45.4725 | 33.9176 | 47.8765 | 30.942
2012-10-30 13:41 | 19.2357 | 64.6333 | 53.1724 | 66.9002 | 50.2701 | 50.3599 | 0.014175 | 0.34719 | 0.004921 | 24.4931 | 0.014389 | 0.235 | 0.003381 | 16.3319 | 19.3423 | 3.4718 | 0.008302 | 45.291 | 33.8301 | 47.5579 | 30.9278
2012-10-30 13:41 | 19.2492 | 64.333 | 53.0815 | 66.6387 | 50.1789 | 50.349 | 0.014194 | 0.3471 | 0.004927 | 24.4542 | 0.014382 | 0.23497 | 0.003379 | 16.3379 | 19.3795 | 3.4711 | 0.008306 | 44.9535 | 33.702 | 47.2592 | 30.7994
2012-10-30 13:41 | 19.2648 | 64.1113 | 52.9851 | 66.3601 | 50.1325 | 50.3062 | 0.014178 | 0.34715 | 0.004922 | 24.4848 | 0.014386 | 0.23501 | 0.003381 | 16.3363 | 19.4108 | 3.4681 | 0.008303 | 44.7005 | 33.5743 | 46.9493 | 30.7217
2012-10-30 13:41 | 19.2686 | 63.9581 | 52.9502 | 66.1346 | 50.0239 | 50.312 | 0.014194 | 0.34709 | 0.004927 | 24.4531 | 0.014376 | 0.23501 | 0.003379 | 16.3477 | 19.4225 | 3.4685 | 0.008306 | 44.5356 | 33.5277 | 46.7121 | 30.6014
2012-10-30 13:41 | 19.2828 | 63.7235 | 52.891 | 65.8774 | 49.9955 | 50.279 | 0.014197 | 0.34707 | 0.004927 | 24.447 | 0.014379 | 0.235 | 0.003379 | 16.3432 | 19.3894 | 3.4662 | 0.008306 | 44.3341 | 33.5016 | 46.488 | 30.6061
2012-10-30 13:42 | 19.3149 | 62.9896 | 52.5287 | 64.9726 | 49.7601 | 50.2126 | 0.014194 | 0.34707 | 0.004926 | 24.4522 | 0.014386 | 0.23503 | 0.003381 | 16.3373 |  | 3.4617 | 0.008307 | 62.9896 | 52.5287 | 64.9726 | 49.7601
2012-10-30 13:42 | 19.3574 | 62.3 | 52.2347 | 64.1312 | 49.5234 | 50.1351 | 0.014203 | 0.34704 | 0.004929 | 24.4346 | 0.014382 | 0.235 | 0.00338 | 16.3398 | 19.385 | 3.4563 | 0.008309 | 42.915 | 32.8497 | 44.7462 | 30.1384
2012-10-30 13:42 | 19.3761 | 62.1424 | 52.1064 | 63.9163 | 49.449 | 50.1409 | 0.014197 | 0.34704 | 0.004927 | 24.445 | 0.014392 | 0.23497 | 0.003382 | 16.3265 |  | 3.4567 | 0.008309 | 62.1424 | 52.1064 | 63.9163 | 49.449
2012-10-30 13:42 | 19.3783 | 61.9605 | 52.0506 | 63.6084 | 49.393 | 50.1334 | 0.01419 | 0.34697 | 0.004924 | 24.4519 | 0.014382 | 0.23501 | 0.00338 | 16.3408 |  | 3.4562 | 0.008304 | 61.9605 | 52.0506 | 63.6084 | 49.393
2012-10-30 13:42 | 19.3798 | 61.7396 | 51.9515 | 63.472 | 49.2898 | 50.0831 | 0.014181 | 0.34697 | 0.00492 | 24.4675 | 0.014382 | 0.23501 | 0.00338 | 16.3408 |  | 3.4527 | 0.0083 | 61.7396 | 51.9515 | 63.472 | 49.2898
2012-10-30 13:42 | 19.391 | 61.5053 | 51.8891 | 63.2534 | 49.204 | 50.0918 | 0.014197 | 0.347 | 0.004926 | 24.4419 | 0.014379 | 0.23497 | 0.003379 | 16.3413 |  | 3.4533 | 0.008305 | 61.5053 | 51.8891 | 63.2534 | 49.204
2012-10-30 13:42 | 19.4104 | 61.3367 | 51.8195 | 62.9586 | 49.1381 | 50.0623 | 0.014184 | 0.34704 | 0.004922 | 24.4674 | 0.014401 | 0.23506 | 0.003385 | 16.3223 |  | 3.4513 | 0.008307 | 61.3367 | 51.8195 | 62.9586 | 49.1381
2012-10-30 13:42 | 19.4172 | 61.1555 | 51.6754 | 62.7317 | 49.0789 | 50.0369 | 0.014181 | 0.34703 | 0.004921 | 24.4715 | 0.014382 | 0.23503 | 0.00338 | 16.3418 |  | 3.4495 | 0.008301 | 61.1555 | 51.6754 | 62.7317 | 49.0789
2012-10-30 13:42 | 19.4254 | 60.9181 | 51.5675 | 62.5022 | 48.928 | 50.0282 | 0.0142 | 0.34699 | 0.004927 | 24.4357 | 0.014373 | 0.23498 | 0.003377 | 16.3491 |  | 3.4489 | 0.008304 | 60.9181 | 51.5675 | 62.5022 | 48.928
2012-10-30 13:42 | 19.4396 | 60.7557 | 51.4617 | 62.3247 | 48.9228 | 50.0317 | 0.01419 | 0.34702 | 0.004924 | 24.455 | 0.014395 | 0.23504 | 0.003383 | 16.3281 |  | 3.4492 | 0.008307 | 60.7557 | 51.4617 | 62.3247 | 48.9228
2012-10-30 13:42 | 19.4515 | 60.5295 | 51.3845 | 62.0682 | 48.8493 | 49.9935 | 0.014187 | 0.34702 | 0.004923 | 24.4601 | 0.014392 | 0.23497 | 0.003382 | 16.3265 |  | 3.4466 | 0.008305 | 60.5295 | 51.3845 | 62.0682 | 48.8493
2012-10-30 13:42 | 19.4717 | 60.4038 | 51.304 | 61.843 | 48.7298 | 49.9675 | 0.014187 | 0.34696 | 0.004922 | 24.4561 | 0.014395 | 0.23501 | 0.003383 | 16.3261 | 19.4044 | 3.4448 | 0.008305 | 40.9994 | 31.8996 | 42.4386 | 29.3254
2012-10-30 13:42 | 19.4739 | 60.2063 | 51.2288 | 61.6572 | 48.6311 | 49.9508 | 0.014197 | 0.34696 | 0.004926 | 24.4388 | 0.014392 | 0.23501 | 0.003382 | 16.3295 | 19.4066 | 3.4436 | 0.008308 | 40.7997 | 31.8222 | 42.2506 | 29.2245
2012-10-30 13:42 | 19.4933 | 60.0145 | 51.0817 | 61.4349 | 48.6078 | 49.9207 | 0.014187 | 0.34699 | 0.004923 | 24.4581 | 0.014389 | 0.23501 | 0.003382 | 16.3329 |  | 3.4415 | 0.008305 | 60.0145 | 51.0817 | 61.4349 | 48.6078
2012-10-30 13:42 | 19.4978 | 59.8499 | 51.0513 | 61.2015 | 48.5579 | 49.8884 | 0.014194 | 0.34703 | 0.004926 | 24.4491 | 0.014395 | 0.23498 | 0.003383 | 16.3241 |  | 3.4393 | 0.008309 | 59.8499 | 51.0513 | 61.2015 | 48.5579
2012-10-30 13:42 | 19.506 | 59.6775 | 50.9123 | 61.0178 | 48.4691 | 49.8762 | 0.014178 | 0.34697 | 0.004919 | 24.4726 | 0.014395 | 0.235 | 0.003383 | 16.325 | 19.3634 | 3.4385 | 0.008302 | 40.3141 | 31.5489 | 41.6544 | 29.1057
2012-10-30 13:42 | 19.5232 | 59.4565 | 50.8249 | 60.8278 | 48.3542 | 49.8624 | 0.014187 | 0.34704 | 0.004924 | 24.4622 | 0.014379 | 0.235 | 0.003379 | 16.3432 | 19.3568 | 3.4375 | 0.008303 | 40.0997 | 31.4681 | 41.471 | 28.9974
2012-10-30 13:42 | 19.5389 | 59.2953 | 50.6973 | 60.6246 | 48.234 | 49.8462 | 0.014194 | 0.3469 | 0.004924 | 24.4399 | 0.014404 | 0.23501 | 0.003385 | 16.3159 | 19.3685 | 3.4364 | 0.008309 | 39.9268 | 31.3288 | 41.2561 | 28.8655
2012-10-30 13:42 | 19.5575 | 59.1217 | 50.6345 | 60.4551 | 48.2215 | 49.8086 | 0.014194 | 0.34693 | 0.004924 | 24.4419 | 0.014373 | 0.23501 | 0.003378 | 16.3511 | 19.3832 | 3.4338 | 0.008302 | 39.7385 | 31.2513 | 41.0719 | 28.8383
2012-10-30 13:42 | 19.5553 | 58.985 | 50.5044 | 60.2378 | 48.13 | 49.815 | 0.014181 | 0.34696 | 0.00492 | 24.4664 | 0.014389 | 0.23501 | 0.003382 | 16.3329 | 19.3968 | 3.4342 | 0.008302 | 39.5882 | 31.1076 | 40.841 | 28.7332
2012-10-30 13:42 | 19.5702 | 58.8114 | 50.4069 | 60.0107 | 48.0323 | 49.7861 | 0.014197 | 0.34693 | 0.004925 | 24.4368 | 0.014389 | 0.23498 | 0.003381 | 16.3309 | 19.4077 | 3.4323 | 0.008306 | 39.4037 | 30.9992 | 40.603 | 28.6246
2012-10-30 13:42 | 19.5852 | 58.638 | 50.3173 | 59.8259 | 47.9696 | 49.7566 | 0.014194 | 0.34691 | 0.004924 | 24.441 | 0.014408 | 0.23509 | 0.003387 | 16.3164 | 19.4267 | 3.4302 | 0.008311 | 39.2113 | 30.8906 | 40.3992 | 28.5429
2012-10-30 13:42 | 19.5889 | 58.4532 | 50.2473 | 59.6029 | 47.8723 | 49.7566 | 0.014197 | 0.34689 | 0.004925 | 24.4338 | 0.014376 | 0.23506 | 0.003379 | 16.3507 | 19.4502 | 3.4302 | 0.008304 | 39.003 | 30.7971 | 40.1527 | 28.4221
2012-10-30 13:42 | 19.6046 | 58.3535 | 50.1428 | 59.4456 | 47.7015 | 49.7624 | 0.014203 | 0.34694 | 0.004928 | 24.4275 | 0.014389 | 0.23497 | 0.003381 | 16.3299 | 19.4303 | 3.4306 | 0.008309 | 38.9232 | 30.7125 | 40.0153 | 28.2712
2012-10-30 13:42 | 19.6158 | 58.1954 | 50.0145 | 59.2723 | 47.6971 | 49.7115 | 0.014187 | 0.34691 | 0.004922 | 24.453 | 0.014395 | 0.235 | 0.003383 | 16.325 | 19.4216 | 3.4271 | 0.008305 | 38.7738 | 30.5929 | 39.8507 | 28.2755
2012-10-30 13:43 | 19.6315 | 57.9994 | 49.9487 | 59.1188 | 47.6273 | 49.6994 | 0.014181 | 0.34689 | 0.004919 | 24.4613 | 0.014395 | 0.235 | 0.003383 | 16.325 | 19.469 | 3.4263 | 0.008302 | 38.5304 | 30.4797 | 39.6498 | 28.1583
2012-10-30 13:43 | 19.6352 | 57.8722 | 49.8361 | 58.9187 | 47.5222 | 49.6907 | 0.0142 | 0.347 | 0.004927 | 24.4368 | 0.014401 | 0.23498 | 0.003384 | 16.3173 | 19.4727 | 3.4257 | 0.008311 | 38.3995 | 30.3634 | 39.446 | 28.0495
2012-10-30 13:43 | 19.6472 | 57.6532 | 49.7705 | 58.7346 | 47.4526 | 49.6514 | 0.014184 | 0.34686 | 0.00492 | 24.4541 | 0.014404 | 0.23504 | 0.003386 | 16.3179 | 19.4729 | 3.423 | 0.008306 | 38.1803 | 30.2976 | 39.2617 | 27.9797
2012-10-30 13:43 | 19.6584 | 57.5027 | 49.6499 | 58.5496 | 47.3316 | 49.6647 | 0.014187 | 0.34686 | 0.004921 | 24.4489 | 0.014392 | 0.23497 | 0.003382 | 16.3265 | 19.4444 | 3.4239 | 0.008303 | 38.0583 | 30.2055 | 39.1052 | 27.8872
2012-10-30 13:43 | 19.674 | 57.345 | 49.4948 | 58.3728 | 47.2889 | 49.6543 | 0.014184 | 0.34689 | 0.00492 | 24.4561 | 0.014382 | 0.23498 | 0.00338 | 16.3388 | 19.5036 | 3.4232 | 0.0083 | 37.8414 | 29.9912 | 38.8692 | 27.7853
2012-10-30 13:43 | 19.6808 | 57.2132 | 49.4007 | 58.2065 | 47.1946 | 49.6231 | 0.014178 | 0.34689 | 0.004918 | 24.4665 | 0.014395 | 0.23491 | 0.003382 | 16.3191 | 19.4906 | 3.421 | 0.0083 | 37.7226 | 29.9101 | 38.7159 | 27.704
2012-10-30 13:43 | 19.6979 | 57.0955 | 49.3829 | 58.0042 | 47.0795 | 49.6202 | 0.014187 | 0.34691 | 0.004922 | 24.453 | 0.014389 | 0.23504 | 0.003382 | 16.3349 | 19.5473 | 3.4208 | 0.008304 | 37.5482 | 29.8356 | 38.4569 | 27.5322
2012-10-30 13:43 | 19.7054 | 56.8718 | 49.2779 | 57.8115 | 46.9742 | 49.6081 | 0.014197 | 0.34681 | 0.004924 | 24.4287 | 0.014392 | 0.23496 | 0.003381 | 16.3255 | 19.5152 | 3.42 | 0.008305 | 37.3566 | 29.7627 | 38.2963 | 27.459
2012-10-30 13:43 | 19.7181 | 56.7265 | 49.1082 | 57.651 | 46.8858 | 49.5873 | 0.014187 | 0.34689 | 0.004921 | 24.451 | 0.014382 | 0.2349 | 0.003378 | 16.3328 | 19.5081 | 3.4185 | 0.008299 | 37.2184 | 29.6001 | 38.1429 | 27.3777
2012-10-30 13:43 | 19.7293 | 56.626 | 49.0689 | 57.4928 | 46.8387 | 49.5555 | 0.014184 | 0.34681 | 0.004919 | 24.4511 | 0.014386 | 0.23498 | 0.00338 | 16.3343 | 19.5193 | 3.4164 | 0.008299 | 37.1067 | 29.5496 | 37.9735 | 27.3194
2012-10-30 13:43 | 19.7435 | 56.4667 | 48.9783 | 57.326 | 46.7907 | 49.5561 | 0.014181 | 0.34694 | 0.00492 | 24.4654 | 0.014389 | 0.23497 | 0.003381 | 16.3299 | 19.5137 | 3.4164 | 0.008301 | 36.953 | 29.4646 | 37.8123 | 27.277
2012-10-30 13:43 | 19.7569 | 56.3028 | 48.8791 | 57.1776 | 46.6757 | 49.5295 | 0.014187 | 0.3469 | 0.004921 | 24.452 | 0.014376 | 0.23498 | 0.003378 | 16.3456 | 19.5389 | 3.4146 | 0.008299 | 36.7639 | 29.3402 | 37.6387 | 27.1368
2012-10-30 13:43 | 19.7674 | 56.1822 | 48.777 | 57.0263 | 46.6123 | 49.5133 | 0.014206 | 0.34684 | 0.004927 | 24.4152 | 0.014392 | 0.23494 | 0.003381 | 16.3245 | 19.4622 | 3.4134 | 0.008308 | 36.72 | 29.3148 | 37.5641 | 27.1501
2012-10-30 13:43 | 19.7801 | 56.0176 | 48.6888 | 56.8425 | 46.5316 | 49.5081 | 0.014181 | 0.34683 | 0.004918 | 24.4572 | 0.01437 | 0.23497 | 0.003377 | 16.3515 | 19.4987 | 3.4131 | 0.008295 | 36.5189 | 29.1901 | 37.3438 | 27.0329
2012-10-30 13:43 | 19.7875 | 55.8398 | 48.5913 | 56.6688 | 46.4573 | 49.4803 | 0.014175 | 0.34681 | 0.004916 | 24.4666 | 0.014386 | 0.2349 | 0.003379 | 16.3282 | 19.4982 | 3.4112 | 0.008295 | 36.3416 | 29.0931 | 37.1706 | 26.9591
2012-10-30 13:43 | 19.8032 | 55.7591 | 48.4983 | 56.5496 | 46.3485 | 49.4671 | 0.014194 | 0.34693 | 0.004924 | 24.4419 | 0.014392 | 0.23491 | 0.003381 | 16.3225 | 19.4861 | 3.4103 | 0.008305 | 36.273 | 29.0122 | 37.0635 | 26.8624
2012-10-30 13:43 | 19.8092 | 55.6185 | 48.4033 | 56.3744 | 46.265 | 49.4682 | 0.014187 | 0.34679 | 0.00492 | 24.4439 | 0.014386 | 0.23497 | 0.00338 | 16.3333 | 19.4961 | 3.4103 | 0.0083 | 36.1224 | 28.9072 | 36.8783 | 26.7689
2012-10-30 13:43 | 19.8249 | 55.4722 | 48.3064 | 56.2513 | 46.1717 | 49.4399 | 0.014197 | 0.34679 | 0.004923 | 24.4266 | 0.014395 | 0.235 | 0.003383 | 16.325 | 19.5039 | 3.4084 | 0.008306 | 35.9683 | 28.8025 | 36.7474 | 26.6678
2012-10-30 13:43 | 19.8391 | 55.305 | 48.2157 | 56.0804 | 46.1314 | 49.4208 | 0.014194 | 0.34686 | 0.004923 | 24.4369 | 0.014386 | 0.235 | 0.003381 | 16.3353 | 19.5062 | 3.4071 | 0.008304 | 35.7988 | 28.7095 | 36.5742 | 26.6252
2012-10-30 13:43 | 19.8548 | 55.1741 | 48.1071 | 55.9573 | 46.0147 | 49.4087 | 0.014181 | 0.34684 | 0.004919 | 24.4583 | 0.014395 | 0.23496 | 0.003382 | 16.3221 | 19.4941 | 3.4062 | 0.008301 | 35.68 | 28.613 | 36.4632 | 26.5206
2012-10-30 13:43 | 19.854 | 55.0498 | 48.0208 | 55.7443 | 45.9516 | 49.3827 | 0.014194 | 0.3468 | 0.004922 | 24.4328 | 0.014379 | 0.23494 | 0.003378 | 16.3392 | 19.5012 | 3.4044 | 0.0083 | 35.5486 | 28.5196 | 36.2431 | 26.4504
2012-10-30 13:43 | 19.8794 | 54.8937 | 47.9413 | 55.5961 | 45.8602 | 49.3659 | 0.014197 | 0.34679 | 0.004923 | 24.4266 | 0.014398 | 0.23488 | 0.003382 | 16.3137 | 19.5227 | 3.4033 | 0.008305 | 35.371 | 28.4186 | 36.0734 | 26.3375
2012-10-30 13:43 | 19.8772 | 54.768 | 47.7992 | 55.455 | 45.7568 | 49.3584 | 0.014194 | 0.34673 | 0.004921 | 24.4277 | 0.014382 | 0.23497 | 0.003379 | 16.3379 | 19.5006 | 3.4028 | 0.0083 | 35.2674 | 28.2986 | 35.9544 | 26.2562
2012-10-30 13:43 | 19.8906 | 54.5767 | 47.7387 | 55.2716 | 45.7001 | 49.344 | 0.014184 | 0.34676 | 0.004918 | 24.447 | 0.014395 | 0.23494 | 0.003382 | 16.3211 | 19.4744 | 3.4018 | 0.0083 | 35.1023 | 28.2643 | 35.7972 | 26.2257
2012-10-30 13:43 | 19.9011 | 54.5254 | 47.5976 | 55.1355 | 45.5548 | 49.3179 | 0.01419 | 0.3468 | 0.004921 | 24.4397 | 0.014379 | 0.23497 | 0.003379 | 16.3413 | 19.5166 | 3.4 | 0.0083 | 35.0088 | 28.081 | 35.6189 | 26.0382
2012-10-30 13:43 | 19.9167 | 54.3402 | 47.551 | 54.9928 | 45.5471 | 49.322 | 0.014178 | 0.34689 | 0.004918 | 24.4665 | 0.014398 | 0.23496 | 0.003383 | 16.3187 | 19.4965 | 3.4003 | 0.008301 | 34.8437 | 28.0545 | 35.4963 | 26.0506
2012-10-30 13:43 | 19.9212 | 54.225 | 47.4078 | 54.8236 | 45.4464 | 49.307 | 0.01419 | 0.34674 | 0.00492 | 24.4357 | 0.014401 | 0.23494 | 0.003383 | 16.3143 | 19.5288 | 3.3992 | 0.008303 | 34.6962 | 27.879 | 35.2948 | 25.9176
2012-10-30 13:43 | 19.9421 | 54.0913 | 47.3393 | 54.6784 | 45.3816 | 49.285 | 0.014187 | 0.34677 | 0.00492 | 24.4428 | 0.014386 | 0.23496 | 0.00338 | 16.3323 | 19.5338 | 3.3977 | 0.0083 | 34.5575 | 27.8055 | 35.1446 | 25.8478
2012-10-30 13:44 | 19.9504 | 53.9103 | 47.2854 | 54.5554 | 45.2691 | 49.2607 | 0.01419 | 0.34679 | 0.004921 | 24.4387 | 0.014404 | 0.235 | 0.003385 | 16.3148 | 19.6016 | 3.396 | 0.008306 | 34.3087 | 27.6838 | 34.9538 | 25.6675
2012-10-30 13:44 | 19.9489 | 53.7851 | 47.1789 | 54.4303 | 45.2092 | 49.2561 | 0.014184 | 0.34679 | 0.004919 | 24.449 | 0.014382 | 0.23494 | 0.003379 | 16.3358 | 19.5406 | 3.3957 | 0.008298 | 34.2445 | 27.6383 | 34.8897 | 25.6686
2012-10-30 13:44 | 19.9705 | 53.6598 | 47.0643 | 54.2897 | 45.1217 | 49.2417 | 0.014175 | 0.34673 | 0.004915 | 24.4605 | 0.014392 | 0.23498 | 0.003382 | 16.3275 | 19.582 | 3.3947 | 0.008297 | 34.0778 | 27.4823 | 34.7077 | 25.5397
2012-10-30 13:44 | 19.9765 | 53.5112 | 46.977 | 54.118 | 45.003 | 49.2249 | 0.01419 | 0.34676 | 0.00492 | 24.4366 | 0.014395 | 0.23498 | 0.003383 | 16.3241 | 19.5603 | 3.3936 | 0.008303 | 33.9509 | 27.4167 | 34.5577 | 25.4427
2012-10-30 13:44 | 19.9929 | 53.4271 | 46.8883 | 53.9566 | 44.9376 | 49.2099 | 0.0142 | 0.34667 | 0.004923 | 24.4133 | 0.014398 | 0.23501 | 0.003384 | 16.3227 | 19.5608 | 3.3925 | 0.008307 | 33.8663 | 27.3275 | 34.3958 | 25.3768
2012-10-30 13:44 | 20.0064 | 53.2898 | 46.7968 | 53.8387 | 44.8692 | 49.2006 | 0.014178 | 0.34674 | 0.004916 | 24.4563 | 0.014395 | 0.2349 | 0.003381 | 16.318 | 19.5545 | 3.3919 | 0.008297 | 33.7353 | 27.2423 | 34.2842 | 25.3147
2012-10-30 13:44 | 20.0079 | 53.0786 | 46.7244 | 53.6895 | 44.7811 | 49.1654 | 0.01419 | 0.3467 | 0.00492 | 24.4326 | 0.014382 | 0.23491 | 0.003379 | 16.3338 | 19.5679 | 3.3895 | 0.008299 | 33.5107 | 27.1565 | 34.1216 | 25.2132
2012-10-30 13:44 | 20.0176 | 53.0071 | 46.6173 | 53.5368 | 44.6543 | 49.1544 | 0.01419 | 0.34673 | 0.00492 | 24.4346 | 0.014392 | 0.2349 | 0.003381 | 16.3214 | 19.5458 | 3.3887 | 0.008301 | 33.4613 | 27.0715 | 33.991 | 25.1085
2012-10-30 13:44 | 20.0288 | 52.8519 | 46.5118 | 53.3663 | 44.5759 | 49.1596 | 0.014203 | 0.34676 | 0.004925 | 24.4143 | 0.014382 | 0.23496 | 0.003379 | 16.3368 | 19.5491 | 3.3891 | 0.008304 | 33.3028 | 26.9627 | 33.8172 | 25.0268
2012-10-30 13:44 | 20.0422 | 52.7415 | 46.4162 | 53.2831 | 44.523 | 49.1209 | 0.014203 | 0.3467 | 0.004924 | 24.4102 | 0.014398 | 0.23497 | 0.003383 | 16.3197 | 19.5585 | 3.3864 | 0.008307 | 33.183 | 26.8577 | 33.7246 | 24.9645
2012-10-30 13:44 | 20.0579 | 52.6412 | 46.3268 | 53.1247 | 44.4178 | 49.1174 | 0.0142 | 0.34668 | 0.004923 | 24.4144 | 0.014398 | 0.2349 | 0.003382 | 16.3146 | 19.5544 | 3.3862 | 0.008305 | 33.0868 | 26.7724 | 33.5703 | 24.8634
2012-10-30 13:44 | 20.0594 | 52.4995 | 46.2388 | 52.956 | 44.353 | 49.1035 | 0.014187 | 0.34665 | 0.004918 | 24.4347 | 0.014408 | 0.23497 | 0.003385 | 16.3084 | 19.5321 | 3.3852 | 0.008303 | 32.9674 | 26.7067 | 33.4239 | 24.8209
2012-10-30 13:44 | 20.0803 | 52.3849 | 46.1857 | 52.9034 | 44.2803 | 49.1001 | 0.014197 | 0.34665 | 0.004921 | 24.4175 | 0.014379 | 0.23493 | 0.003378 | 16.3382 | 19.5768 | 3.385 | 0.008299 | 32.8081 | 26.6089 | 33.3266 | 24.7035
2012-10-30 13:44 | 20.0863 | 52.2322 | 46.036 | 52.7005 | 44.1264 | 49.0833 | 0.014194 | 0.34673 | 0.004921 | 24.4277 | 0.014392 | 0.235 | 0.003382 | 16.3284 | 19.5788 | 3.3838 | 0.008303 | 32.6534 | 26.4572 | 33.1217 | 24.5476
2012-10-30 13:44 | 20.1042 | 52.1108 | 45.9371 | 52.5752 | 44.0585 | 49.0544 | 0.014181 | 0.34668 | 0.004916 | 24.4471 | 0.014389 | 0.2349 | 0.00338 | 16.3248 | 19.5292 | 3.3818 | 0.008296 | 32.5816 | 26.4079 | 33.046 | 24.5293
2012-10-30 13:44 | 20.1109 | 51.982 | 45.8737 | 52.4813 | 43.9638 | 49.0556 | 0.014181 | 0.34665 | 0.004916 | 24.445 | 0.014398 | 0.23494 | 0.003383 | 16.3177 | 19.5558 | 3.3819 | 0.008299 | 32.4262 | 26.3179 | 32.9255 | 24.408
2012-10-30 13:44 | 20.1184 | 51.9004 | 45.7916 | 52.3882 | 43.9127 | 49.0405 | 0.014194 | 0.34665 | 0.00492 | 24.4226 | 0.014392 | 0.23491 | 0.003381 | 16.3225 | 19.5633 | 3.3809 | 0.008301 | 32.3371 | 26.2283 | 32.8249 | 24.3494
2012-10-30 13:44 | 20.1311 | 51.7311 | 45.6368 | 52.1803 | 43.8864 | 49.0203 | 0.01419 | 0.34665 | 0.004919 | 24.4295 | 0.014386 | 0.23501 | 0.003381 | 16.3363 | 19.5522 | 3.3795 | 0.0083 | 32.1789 | 26.0846 | 32.6281 | 24.3342
2012-10-30 13:44 | 20.1438 | 51.6353 | 45.5989 | 52.0652 | 43.7977 | 49.0099 | 0.014206 | 0.34667 | 0.004925 | 24.403 | 0.014395 | 0.23487 | 0.003381 | 16.316 | 19.5649 | 3.3787 | 0.008306 | 32.0704 | 26.034 | 32.5003 | 24.2328
2012-10-30 13:44 | 20.1505 | 51.5065 | 45.477 | 51.9325 | 43.6951 | 48.992 | 0.014187 | 0.34667 | 0.004918 | 24.4357 | 0.014401 | 0.23488 | 0.003383 | 16.3103 | 19.5676 | 3.3775 | 0.008301 | 31.9389 | 25.9094 | 32.3649 | 24.1275
2012-10-30 13:44 | 20.161 | 51.3697 | 45.394 | 51.7803 | 43.6198 | 48.9799 | 0.014178 | 0.34661 | 0.004914 | 24.4471 | 0.014376 | 0.23486 | 0.003376 | 16.3366 | 19.5781 | 3.3767 | 0.00829 | 31.7916 | 25.8159 | 32.2022 | 24.0417
2012-10-30 13:44 | 20.1834 | 51.2526 | 45.3268 | 51.6323 | 43.5485 | 48.9596 | 0.014184 | 0.34661 | 0.004916 | 24.4368 | 0.014398 | 0.23488 | 0.003382 | 16.3137 | 19.6283 | 3.3753 | 0.008298 | 31.6243 | 25.6985 | 32.004 | 23.9202
2012-10-30 13:44 | 20.1804 | 51.1528 | 45.2225 | 51.4976 | 43.4948 | 48.9371 | 0.014187 | 0.34658 | 0.004917 | 24.4296 | 0.014414 | 0.23487 | 0.003385 | 16.2945 | 19.5578 | 3.3737 | 0.008302 | 31.595 | 25.6647 | 31.9398 | 23.937
2012-10-30 13:44 | 20.1886 | 50.9943 | 45.1488 | 51.4167 | 43.339 | 48.9365 | 0.01419 | 0.34655 | 0.004918 | 24.4224 | 0.014382 | 0.23493 | 0.003379 | 16.3348 | 19.5819 | 3.3737 | 0.008297 | 31.4124 | 25.5669 | 31.8348 | 23.7571
2012-10-30 13:44 | 20.208 | 50.8897 | 45.0708 | 51.2772 | 43.2999 | 48.9186 | 0.014194 | 0.34657 | 0.004919 | 24.4165 | 0.014386 | 0.23493 | 0.00338 | 16.3303 | 19.5814 | 3.3724 | 0.008299 | 31.3083 | 25.4894 | 31.6958 | 23.7185
2012-10-30 13:44 | 20.2162 | 50.7428 | 44.9543 | 51.1692 | 43.23 | 48.899 | 0.014184 | 0.34664 | 0.004917 | 24.4388 | 0.014395 | 0.23494 | 0.003382 | 16.3211 | 19.6095 | 3.3711 | 0.008299 | 31.1333 | 25.3448 | 31.5597 | 23.6205
2012-10-30 13:44 | 20.2252 | 50.6859 | 44.8814 | 51.027 | 43.1687 | 48.8857 | 0.014184 | 0.34658 | 0.004916 | 24.4348 | 0.014379 | 0.23491 | 0.003378 | 16.3372 | 19.5907 | 3.3702 | 0.008294 | 31.0952 | 25.2907 | 31.4363 | 23.578
2012-10-30 13:45 | 20.2334 | 50.5157 | 44.7844 | 50.8996 | 43.0441 | 48.8845 | 0.014187 | 0.3466 | 0.004917 | 24.4306 | 0.014398 | 0.23486 | 0.003381 | 16.3116 | 19.5751 | 3.3701 | 0.008298 | 30.9406 | 25.2093 | 31.3245 | 23.469
2012-10-30 13:45 | 20.2551 | 50.4289 | 44.6969 | 50.774 | 42.9956 | 48.8712 | 0.014184 | 0.3466 | 0.004916 | 24.4358 | 0.014392 | 0.23488 | 0.00338 | 16.3205 | 19.6047 | 3.3692 | 0.008296 | 30.8242 | 25.0922 | 31.1693 | 23.3909
2012-10-30 13:45 | 20.2573 | 50.2798 | 44.6445 | 50.6133 | 42.8728 | 48.8475 | 0.0142 | 0.34661 | 0.004922 | 24.4092 | 0.014395 | 0.23488 | 0.003381 | 16.3171 | 19.6387 | 3.3675 | 0.008303 | 30.6411 | 25.0058 | 30.9746 | 23.2341
2012-10-30 13:45 | 20.2678 | 50.2282 | 44.4912 | 50.5075 | 42.77 | 48.8267 | 0.01419 | 0.34663 | 0.004919 | 24.4275 | 0.014389 | 0.23494 | 0.003381 | 16.3279 | 19.6095 | 3.3661 | 0.0083 | 30.6187 | 24.8817 | 30.898 | 23.1605
2012-10-30 13:45 | 20.2827 | 50.0492 | 44.4593 | 50.3711 | 42.6873 | 48.825 | 0.014187 | 0.3466 | 0.004917 | 24.4306 | 0.014382 | 0.23494 | 0.003379 | 16.3358 | 19.6085 | 3.366 | 0.008296 | 30.4407 | 24.8508 | 30.7626 | 23.0788
2012-10-30 13:45 | 20.2954 | 49.9532 | 44.3433 | 50.2791 | 42.6648 | 48.8238 | 0.014187 | 0.34667 | 0.004918 | 24.4357 | 0.014389 | 0.23491 | 0.00338 | 16.3259 | 19.5974 | 3.3659 | 0.008298 | 30.3558 | 24.7459 | 30.6817 | 23.0674
2012-10-30 13:45 | 20.3006 | 49.8614 | 44.2744 | 50.1834 | 42.5645 | 48.7897 | 0.014168 | 0.34655 | 0.00491 | 24.4603 | 0.014389 | 0.23487 | 0.00338 | 16.3228 | 19.6304 | 3.3636 | 0.00829 | 30.231 | 24.644 | 30.553 | 22.9341
2012-10-30 13:45 | 20.3066 | 49.7665 | 44.179 | 50.042 | 42.4768 | 48.7932 | 0.014172 | 0.34657 | 0.004912 | 24.4544 | 0.014382 | 0.23486 | 0.003378 | 16.3298 | 19.5966 | 3.3638 | 0.00829 | 30.1699 | 24.5824 | 30.4454 | 22.8802
2012-10-30 13:45 | 20.3216 | 49.5913 | 44.1042 | 49.8824 | 42.4253 | 48.7759 | 0.0142 | 0.34653 | 0.004921 | 24.4032 | 0.014389 | 0.23487 | 0.00338 | 16.3228 | 19.5759 | 3.3626 | 0.008301 | 30.0154 | 24.5283 | 30.3065 | 22.8494
2012-10-30 13:45 | 20.3231 | 49.5346 | 43.9926 | 49.7985 | 42.3213 | 48.7608 | 0.014203 | 0.34658 | 0.004923 | 24.4021 | 0.014389 | 0.23488 | 0.00338 | 16.3239 | 19.5496 | 3.3616 | 0.008303 | 29.985 | 24.443 | 30.2489 | 22.7717
2012-10-30 13:45 | 20.3268 | 49.3403 | 43.9182 | 49.7052 | 42.2039 | 48.7481 | 0.014194 | 0.3466 | 0.00492 | 24.4186 | 0.014392 | 0.23484 | 0.00338 | 16.3174 | 19.5771 | 3.3607 | 0.0083 | 29.7632 | 24.3411 | 30.1281 | 22.6268
2012-10-30 13:45 | 20.3328 | 49.2764 | 43.8306 | 49.6258 | 42.1239 | 48.7134 | 0.014187 | 0.34651 | 0.004916 | 24.4245 | 0.014382 | 0.23481 | 0.003377 | 16.3268 | 19.607 | 3.3583 | 0.008293 | 29.6694 | 24.2236 | 30.0188 | 22.5169
2012-10-30 13:45 | 20.3387 | 49.2241 | 43.7389 | 49.4726 | 42.0516 | 48.7366 | 0.014184 | 0.34653 | 0.004915 | 24.4307 | 0.014392 | 0.23494 | 0.003381 | 16.3245 | 19.5453 | 3.3599 | 0.008296 | 29.6788 | 24.1936 | 29.9273 | 22.5063
2012-10-30 13:45 | 20.3626 | 49.1082 | 43.7199 | 49.4149 | 42.0247 | 48.7181 | 0.014187 | 0.34645 | 0.004915 | 24.4205 | 0.014386 | 0.23491 | 0.003379 | 16.3293 | 19.5533 | 3.3586 | 0.008294 | 29.5549 | 24.1666 | 29.8616 | 22.4714
2012-10-30 13:45 | 20.3746 | 49.0153 | 43.5953 | 49.2794 | 41.9585 | 48.6932 | 0.014197 | 0.34655 | 0.00492 | 24.4104 | 0.014386 | 0.2349 | 0.003379 | 16.3282 | 19.5971 | 3.3569 | 0.008299 | 29.4182 | 23.9982 | 29.6823 | 22.3614
2012-10-30 13:45 | 20.3694 | 48.8781 | 43.5081 | 49.1461 | 41.8869 | 48.6926 | 0.014175 | 0.34645 | 0.004911 | 24.4411 | 0.014386 | 0.23484 | 0.003378 | 16.3242 | 19.6038 | 3.3569 | 0.008289 | 29.2743 | 23.9043 | 29.5423 | 22.2831
2012-10-30 13:45 | 20.3865 | 48.7864 | 43.4276 | 49.0583 | 41.8375 | 48.6707 | 0.01419 | 0.34648 | 0.004917 | 24.4173 | 0.014395 | 0.23487 | 0.003381 | 16.316 | 19.6487 | 3.3554 | 0.008298 | 29.1377 | 23.7789 | 29.4096 | 22.1888
2012-10-30 13:45 | 20.3903 | 48.6154 | 43.3455 | 48.899 | 41.6263 | 48.651 | 0.014187 | 0.34653 | 0.004916 | 24.4255 | 0.014398 | 0.23483 | 0.003381 | 16.3096 | 19.5493 | 3.354 | 0.008297 | 29.0661 | 23.7962 | 29.3497 | 22.077
2012-10-30 13:45 | 20.4045 | 48.548 | 43.2621 | 48.8122 | 41.5701 | 48.6603 | 0.014162 | 0.34647 | 0.004907 | 24.4646 | 0.014392 | 0.23488 | 0.00338 | 16.3205 | 19.6469 | 3.3546 | 0.008287 | 28.9011 | 23.6152 | 29.1653 | 21.9232
2012-10-30 13:45 | 20.4142 | 48.4295 | 43.1898 | 48.682 | 41.5211 | 48.6493 | 0.014184 | 0.34654 | 0.004915 | 24.4317 | 0.014398 | 0.2348 | 0.003381 | 16.3076 | 19.6407 | 3.3539 | 0.008296 | 28.7888 | 23.5491 | 29.0413 | 21.8804
2012-10-30 13:45 | 20.4254 | 48.3513 | 43.0917 | 48.5767 | 41.4502 | 48.6302 | 0.014172 | 0.3465 | 0.004911 | 24.4493 | 0.014395 | 0.23481 | 0.00338 | 16.3121 | 19.636 | 3.3526 | 0.008291 | 28.7153 | 23.4557 | 28.9407 | 21.8142
2012-10-30 13:45 | 20.4396 | 48.2061 | 43.0004 | 48.4626 | 41.3979 | 48.6083 | 0.014178 | 0.34647 | 0.004912 | 24.4369 | 0.014376 | 0.23483 | 0.003376 | 16.3346 | 19.6502 | 3.3511 | 0.008288 | 28.5559 | 23.3502 | 28.8124 | 21.7477
2012-10-30 13:45 | 20.4433 | 48.1282 | 42.9143 | 48.3808 | 41.3312 | 48.6129 | 0.014181 | 0.34648 | 0.004913 | 24.4328 | 0.014379 | 0.23493 | 0.003378 | 16.3382 | 19.6221 | 3.3514 | 0.008291 | 28.5061 | 23.2922 | 28.7587 | 21.7091
2012-10-30 13:45 | 20.456 | 48.0903 | 42.8567 | 48.2925 | 41.25 | 48.6019 | 0.014184 | 0.34644 | 0.004914 | 24.4246 | 0.014392 | 0.23493 | 0.003381 | 16.3234 | 19.607 | 3.3506 | 0.008295 | 28.4833 | 23.2497 | 28.6855 | 21.643
2012-10-30 13:45 | 20.4553 | 47.9691 | 42.7309 | 48.1518 | 41.1633 | 48.5776 | 0.014194 | 0.34645 | 0.004918 | 24.4084 | 0.014389 | 0.2349 | 0.00338 | 16.3248 | 19.6182 | 3.3489 | 0.008298 | 28.3509 | 23.1127 | 28.5336 | 21.5451
2012-10-30 13:45 | 20.4642 | 47.8225 | 42.7008 | 48.0363 | 41.1174 | 48.5678 | 0.014181 | 0.34648 | 0.004913 | 24.4328 | 0.014395 | 0.23487 | 0.003381 | 16.316 | 19.6311 | 3.3483 | 0.008294 | 28.1914 | 23.0697 | 28.4052 | 21.4863
2012-10-30 13:46 | 20.4792 | 47.7014 | 42.5908 | 47.9308 | 41.0659 | 48.5551 | 0.014194 | 0.34651 | 0.004918 | 24.4124 | 0.014386 | 0.23483 | 0.003378 | 16.3232 | 19.6262 | 3.3474 | 0.008296 | 28.0752 | 22.9646 | 28.3046 | 21.4397
2012-10-30 13:46 | 20.4762 | 47.6245 | 42.5096 | 47.8034 | 40.9807 | 48.547 | 0.01419 | 0.34655 | 0.004918 | 24.4224 | 0.014386 | 0.23484 | 0.003378 | 16.3242 | 19.6113 | 3.3468 | 0.008296 | 28.0132 | 22.8983 | 28.1921 | 21.3694
2012-10-30 13:46 | 20.4941 | 47.5296 | 42.4572 | 47.7124 | 40.893 | 48.5106 | 0.01419 | 0.34645 | 0.004916 | 24.4153 | 0.014379 | 0.23486 | 0.003377 | 16.3332 | 19.657 | 3.3443 | 0.008293 | 27.8726 | 22.8002 | 28.0554 | 21.236
2012-10-30 13:46 | 20.5016 | 47.4554 | 42.3982 | 47.6149 | 40.7987 | 48.4961 | 0.014194 | 0.3464 | 0.004917 | 24.4043 | 0.014392 | 0.23483 | 0.00338 | 16.3164 | 19.6566 | 3.3433 | 0.008297 | 27.7988 | 22.7416 | 27.9583 | 21.1421
2012-10-30 13:46 | 20.5128 | 47.2837 | 42.2922 | 47.4822 | 40.7199 | 48.4979 | 0.01419 | 0.34645 | 0.004916 | 24.4153 | 0.014386 | 0.23486 | 0.003379 | 16.3252 | 19.6916 | 3.3434 | 0.008295 | 27.5921 | 22.6006 | 27.7906 | 21.0283
2012-10-30 13:46 | 20.524 | 47.2132 | 42.19 | 47.3377 | 40.6059 | 48.4811 | 0.01419 | 0.34647 | 0.004916 | 24.4163 | 0.014379 | 0.2349 | 0.003378 | 16.3362 | 19.679 | 3.3423 | 0.008294 | 27.5342 | 22.511 | 27.6587 | 20.9269
2012-10-30 13:46 | 20.5352 | 47.0882 | 42.0957 | 47.2478 | 40.5701 | 48.4528 | 0.014181 | 0.34645 | 0.004913 | 24.4308 | 0.014398 | 0.23488 | 0.003382 | 16.3137 | 19.6624 | 3.3403 | 0.008295 | 27.4258 | 22.4333 | 27.5854 | 20.9077
2012-10-30 13:46 | 20.5419 | 47.0521 | 42.0438 | 47.1922 | 40.4907 | 48.4412 | 0.014184 | 0.3465 | 0.004915 | 24.4287 | 0.014389 | 0.23491 | 0.00338 | 16.3259 | 19.6691 | 3.3395 | 0.008295 | 27.383 | 22.3747 | 27.5231 | 20.8216
2012-10-30 13:46 | 20.5501 | 46.9514 | 41.9542 | 47.107 | 40.4519 | 48.4505 | 0.01419 | 0.34648 | 0.004917 | 24.4173 | 0.014392 | 0.23487 | 0.00338 | 16.3194 | 19.7011 | 3.3402 | 0.008297 | 27.2503 | 22.2531 | 27.4059 | 20.7508
2012-10-30 13:46 | 20.5591 | 46.7696 | 41.8694 | 46.9798 | 40.367 | 48.4216 | 0.0142 | 0.34647 | 0.00492 | 24.3991 | 0.014386 | 0.23486 | 0.003379 | 16.3252 | 19.7101 | 3.3382 | 0.008299 | 27.0595 | 22.1593 | 27.2697 | 20.6569
2012-10-30 13:46 | 20.5733 | 46.6943 | 41.8054 | 46.8656 | 40.3068 | 48.406 | 0.01419 | 0.34637 | 0.004915 | 24.4092 | 0.014386 | 0.23488 | 0.003379 | 16.3273 | 19.7362 | 3.3371 | 0.008294 | 26.9581 | 22.0692 | 27.1294 | 20.5706
2012-10-30 13:46 | 20.586 | 46.6252 | 41.7516 | 46.8004 | 40.202 | 48.4094 | 0.014194 | 0.34644 | 0.004917 | 24.4074 | 0.014389 | 0.23487 | 0.00338 | 16.3228 | 19.6933 | 3.3373 | 0.008297 | 26.9319 | 22.0583 | 27.1071 | 20.5087
2012-10-30 13:46 | 20.5882 | 46.534 | 41.6599 | 46.682 | 40.0946 | 48.406 | 0.014197 | 0.34645 | 0.004919 | 24.4033 | 0.014395 | 0.23488 | 0.003381 | 16.3171 | 19.7233 | 3.3371 | 0.0083 | 26.8107 | 21.9366 | 26.9587 | 20.3713
2012-10-30 13:46 | 20.5897 | 46.4226 | 41.5676 | 46.5199 | 40.0961 | 48.3863 | 0.014184 | 0.34644 | 0.004914 | 24.4246 | 0.014398 | 0.23487 | 0.003382 | 16.3126 | 19.7328 | 3.3358 | 0.008296 | 26.6898 | 21.8348 | 26.7871 | 20.3633
2012-10-30 13:46 | 20.6054 | 46.3526 | 41.462 | 46.45 | 40.0099 | 48.373 | 0.01419 | 0.3464 | 0.004915 | 24.4112 | 0.014392 | 0.23487 | 0.00338 | 16.3194 | 19.6849 | 3.3348 | 0.008295 | 26.6677 | 21.7771 | 26.7651 | 20.325
2012-10-30 13:46 | 20.6151 | 46.2182 | 41.4091 | 46.3701 | 39.8943 | 48.3655 | 0.014203 | 0.3464 | 0.00492 | 24.3889 | 0.014398 | 0.23483 | 0.003381 | 16.3096 | 19.6787 | 3.3343 | 0.008301 | 26.5395 | 21.7304 | 26.6914 | 20.2156
2012-10-30 13:46 | 20.6203 | 46.126 | 41.3283 | 46.2701 | 39.8682 | 48.3574 | 0.014187 | 0.34645 | 0.004915 | 24.4205 | 0.014392 | 0.23486 | 0.00338 | 16.3184 | 19.6998 | 3.3338 | 0.008295 | 26.4262 | 21.6285 | 26.5703 | 20.1684
2012-10-30 13:46 | 20.6226 | 46.0504 | 41.2562 | 46.1322 | 39.7804 | 48.3401 | 0.014194 | 0.34641 | 0.004917 | 24.4053 | 0.014408 | 0.2349 | 0.003384 | 16.3033 | 19.71 | 3.3326 | 0.008301 | 26.3404 | 21.5462 | 26.4222 | 20.0704
2012-10-30 13:46 | 20.6405 | 45.9008 | 41.1842 | 46.0722 | 39.7199 | 48.321 | 0.014194 | 0.34647 | 0.004918 | 24.4094 | 0.014398 | 0.23486 | 0.003381 | 16.3116 | 19.6962 | 3.3312 | 0.008299 | 26.2046 | 21.488 | 26.376 | 20.0237
2012-10-30 13:46 | 20.648 | 45.8148 | 41.0977 | 45.9784 | 39.6491 | 48.3233 | 0.014184 | 0.34647 | 0.004914 | 24.4266 | 0.014382 | 0.23488 | 0.003378 | 16.3318 | 19.6639 | 3.3314 | 0.008292 | 26.1509 | 21.4338 | 26.3145 | 19.9852
2012-10-30 13:46 | 20.6607 | 45.7379 | 40.997 | 45.8742 | 39.5599 | 48.3135 | 0.014187 | 0.34645 | 0.004915 | 24.4205 | 0.014414 | 0.23488 | 0.003386 | 16.2955 | 19.7322 | 3.3307 | 0.008301 | 26.0057 | 21.2648 | 26.142 | 19.8277
2012-10-30 13:46 | 20.6749 | 45.6273 | 40.9172 | 45.7482 | 39.5114 | 48.2742 | 0.014194 | 0.3465 | 0.004918 | 24.4114 | 0.014404 | 0.23487 | 0.003383 | 16.3058 | 19.7385 | 3.328 | 0.008301 | 25.8888 | 21.1787 | 26.0097 | 19.7729
2012-10-30 13:46 | 20.6734 | 45.5284 | 40.8062 | 45.6219 | 39.4237 | 48.2811 | 0.014181 | 0.34648 | 0.004913 | 24.4328 | 0.014389 | 0.23487 | 0.00338 | 16.3228 | 19.741 | 3.3285 | 0.008293 | 25.7874 | 21.0652 | 25.8809 | 19.6827
2012-10-30 13:46 | 20.6823 | 45.4009 | 40.7642 | 45.5451 | 39.3464 | 48.2863 | 0.014184 | 0.34634 | 0.004912 | 24.4174 | 0.014389 | 0.23484 | 0.003379 | 16.3208 | 19.734 | 3.3289 | 0.008291 | 25.6669 | 21.0302 | 25.8111 | 19.6124
2012-10-30 13:46 | 20.6943 | 45.3583 | 40.6666 | 45.4519 | 39.2683 | 48.265 | 0.014197 | 0.34645 | 0.004919 | 24.4033 | 0.014398 | 0.23484 | 0.003381 | 16.3106 | 19.7301 | 3.3274 | 0.0083 | 25.6282 | 20.9365 | 25.7218 | 19.5382
2012-10-30 13:46 | 20.7033 | 45.2036 | 40.6208 | 45.3634 | 39.1597 | 48.2476 | 0.014175 | 0.34634 | 0.004909 | 24.4329 | 0.014395 | 0.2349 | 0.003381 | 16.318 | 19.7192 | 3.3262 | 0.00829 | 25.4844 | 20.9016 | 25.6442 | 19.4405
2012-10-30 13:46 | 20.7055 | 45.1473 | 40.5016 | 45.2331 | 39.0639 | 48.2505 | 0.01419 | 0.34637 | 0.004915 | 24.4092 | 0.014398 | 0.2349 | 0.003382 | 16.3146 | 19.7294 | 3.3264 | 0.008297 | 25.4179 | 20.7722 | 25.5037 | 19.3345
2012-10-30 13:47 | 20.7212 | 45.0889 | 40.4273 | 45.1786 | 39.0522 | 48.2309 | 0.014187 | 0.3464 | 0.004914 | 24.4164 | 0.014398 | 0.23487 | 0.003382 | 16.3126 | 19.6974 | 3.325 | 0.008296 | 25.3915 | 20.7299 | 25.4812 | 19.3548
2012-10-30 13:47 | 20.7302 | 44.9965 | 40.3854 | 45.0355 | 38.9789 | 48.2118 | 0.01419 | 0.34641 | 0.004916 | 24.4122 | 0.014395 | 0.23477 | 0.003379 | 16.309 | 19.7302 | 3.3237 | 0.008295 | 25.2663 | 20.6552 | 25.3053 | 19.2487
2012-10-30 13:47 | 20.7339 | 44.8637 | 40.3225 | 44.9417 | 38.9316 | 48.2158 | 0.014187 | 0.34635 | 0.004914 | 24.4133 | 0.014386 | 0.23483 | 0.003378 | 16.3232 | 19.6981 | 3.324 | 0.008292 | 25.1656 | 20.6244 | 25.2436 | 19.2335
2012-10-30 13:47 | 20.7503 | 44.7826 | 40.2527 | 44.8255 | 38.8186 | 48.1945 | 0.014178 | 0.34634 | 0.00491 | 24.4278 | 0.014389 | 0.23481 | 0.003379 | 16.3189 | 19.7583 | 3.3225 | 0.008289 | 25.0243 | 20.4944 | 25.0672 | 19.0603
2012-10-30 13:47 | 20.7526 | 44.6757 | 40.165 | 44.7537 | 38.7308 | 48.1841 | 0.01419 | 0.34634 | 0.004915 | 24.4071 | 0.014382 | 0.23481 | 0.003377 | 16.3268 | 19.7129 | 3.3218 | 0.008292 | 24.9628 | 20.4521 | 25.0408 | 19.0179
2012-10-30 13:47 | 20.7668 | 44.6275 | 40.093 | 44.6431 | 38.6469 | 48.1927 | 0.0142 | 0.34634 | 0.004918 | 24.3899 | 0.014401 | 0.23486 | 0.003382 | 16.3082 | 19.7191 | 3.3224 | 0.0083 | 24.9084 | 20.3739 | 24.924 | 18.9278
2012-10-30 13:47 | 20.7698 | 44.5368 | 39.9863 | 44.5602 | 38.6107 | 48.1632 | 0.014197 | 0.34628 | 0.004916 | 24.391 | 0.014392 | 0.2348 | 0.003379 | 16.3144 | 19.7301 | 3.3204 | 0.008295 | 24.8067 | 20.2562 | 24.8301 | 18.8806
2012-10-30 13:47 | 20.775 | 44.4406 | 39.9249 | 44.4874 | 38.5453 | 48.165 | 0.01419 | 0.34637 | 0.004915 | 24.4092 | 0.014395 | 0.2348 | 0.00338 | 16.311 | 19.7233 | 3.3205 | 0.008295 | 24.7173 | 20.2016 | 24.7641 | 18.822
2012-10-30 13:47 | 20.7937 | 44.3617 | 39.8731 | 44.4398 | 38.5091 | 48.165 | 0.0142 | 0.34637 | 0.004918 | 24.392 | 0.014395 | 0.23484 | 0.003381 | 16.314 | 19.7102 | 3.3205 | 0.008299 | 24.6515 | 20.1629 | 24.7296 | 18.7989
2012-10-30 13:47 | 20.7899 | 44.2643 | 39.7832 | 44.315 | 38.3328 | 48.1292 | 0.014197 | 0.34635 | 0.004917 | 24.3961 | 0.014398 | 0.2348 | 0.003381 | 16.3076 | 19.7064 | 3.318 | 0.008298 | 24.5579 | 20.0768 | 24.6086 | 18.6264
2012-10-30 13:47 | 20.7996 | 44.1647 | 39.7262 | 44.2037 | 38.315 | 48.1234 | 0.014184 | 0.34631 | 0.004912 | 24.4154 | 0.014382 | 0.23477 | 0.003376 | 16.3237 | 19.7042 | 3.3176 | 0.008288 | 24.4605 | 20.022 | 24.4995 | 18.6108
2012-10-30 13:47 | 20.8056 | 44.077 | 39.6343 | 44.12 | 38.2426 | 48.1205 | 0.014206 | 0.34631 | 0.00492 | 24.3776 | 0.014389 | 0.23472 | 0.003377 | 16.3128 | 19.7261 | 3.3174 | 0.008297 | 24.3509 | 19.9082 | 24.3939 | 18.5165
2012-10-30 13:47 | 20.822 | 43.9724 | 39.5723 | 44.0466 | 38.2629 | 48.1003 | 0.014197 | 0.34628 | 0.004916 | 24.391 | 0.014392 | 0.23475 | 0.003379 | 16.3114 | 19.6988 | 3.316 | 0.008295 | 24.2736 | 19.8735 | 24.3478 | 18.5641
2012-10-30 13:47 | 20.822 | 43.9021 | 39.4861 | 43.91 | 38.1962 | 48.1066 | 0.014194 | 0.34631 | 0.004915 | 24.3982 | 0.014395 | 0.23484 | 0.003381 | 16.314 | 19.7226 | 3.3165 | 0.008296 | 24.1795 | 19.7635 | 24.1874 | 18.4736
2012-10-30 13:47 | 20.8392 | 43.81 | 39.4171 | 43.8959 | 38.1349 | 48.0939 | 0.014203 | 0.34632 | 0.004919 | 24.3838 | 0.014392 | 0.23483 | 0.00338 | 16.3164 | 19.7836 | 3.3156 | 0.008299 | 24.0264 | 19.6335 | 24.1123 | 18.3513
2012-10-30 13:47 | 20.8422 | 43.7545 | 39.3496 | 43.7662 | 38.0477 | 48.0731 | 0.014184 | 0.34628 | 0.004912 | 24.4134 | 0.014404 | 0.23481 | 0.003382 | 16.3019 | 19.711 | 3.3142 | 0.008294 | 24.0435 | 19.6386 | 24.0552 | 18.3367
2012-10-30 13:47 | 20.8467 | 43.6535 | 39.2678 | 43.6379 | 37.9777 | 48.0852 | 0.014194 | 0.34632 | 0.004916 | 24.3993 | 0.014395 | 0.23474 | 0.003379 | 16.307 | 19.7354 | 3.315 | 0.008295 | 23.9181 | 19.5324 | 23.9025 | 18.2423
2012-10-30 13:47 | 20.8534 | 43.5626 | 39.2314 | 43.5743 | 37.9177 | 48.0812 | 0.014187 | 0.34629 | 0.004913 | 24.4092 | 0.014379 | 0.23484 | 0.003377 | 16.3322 | 19.7222 | 3.3147 | 0.00829 | 23.8404 | 19.5092 | 23.8521 | 18.1955
2012-10-30 13:47 | 20.8661 | 43.4933 | 39.1814 | 43.5011 | 37.8558 | 48.0477 | 0.014194 | 0.34626 | 0.004915 | 24.3952 | 0.014404 | 0.23478 | 0.003382 | 16.2998 | 19.6872 | 3.3124 | 0.008297 | 23.8061 | 19.4942 | 23.8139 | 18.1686
2012-10-30 13:47 | 20.8743 | 43.3921 | 39.0916 | 43.4078 | 37.8248 | 48.0604 | 0.014178 | 0.34635 | 0.004911 | 24.4288 | 0.014395 | 0.2348 | 0.00338 | 16.311 | 19.7033 | 3.3133 | 0.008291 | 23.6888 | 19.3883 | 23.7045 | 18.1215
2012-10-30 13:47 | 20.8796 | 43.2568 | 39.042 | 43.3193 | 37.732 | 48.0315 | 0.014178 | 0.34632 | 0.00491 | 24.4268 | 0.014386 | 0.23475 | 0.003377 | 16.3182 | 19.6927 | 3.3113 | 0.008287 | 23.5641 | 19.3493 | 23.6266 | 18.0393
2012-10-30 13:47 | 20.8728 | 43.1797 | 38.9098 | 43.1953 | 37.6428 | 48.0303 | 0.014181 | 0.34624 | 0.00491 | 24.4155 | 0.014398 | 0.23481 | 0.003381 | 16.3087 | 19.7217 | 3.3112 | 0.008291 | 23.458 | 19.1881 | 23.4736 | 17.9211
2012-10-30 13:47 | 20.896 | 43.0935 | 38.9251 | 43.0935 | 37.5796 | 48.0176 | 0.01419 | 0.34626 | 0.004914 | 24.402 | 0.014389 | 0.23486 | 0.003379 | 16.3218 | 19.7171 | 3.3103 | 0.008293 | 23.3764 | 19.208 | 23.3764 | 17.8625
2012-10-30 13:47 | 20.8938 | 43.0014 | 38.8327 | 42.9936 | 37.5107 | 48.0078 | 0.014172 | 0.34626 | 0.004907 | 24.433 | 0.014392 | 0.23483 | 0.00338 | 16.3164 | 19.7069 | 3.3097 | 0.008287 | 23.2945 | 19.1258 | 23.2867 | 17.8038
2012-10-30 13:47 | 20.9132 | 42.9153 | 38.7463 | 42.9505 | 37.4398 | 47.9864 | 0.014184 | 0.34629 | 0.004912 | 24.4144 | 0.014392 | 0.23475 | 0.003379 | 16.3114 | 19.7383 | 3.3082 | 0.008291 | 23.177 | 19.008 | 23.2122 | 17.7015
2012-10-30 13:47 | 20.9177 | 42.8573 | 38.6802 | 42.8847 | 37.3933 | 47.9887 | 0.0142 | 0.34626 | 0.004917 | 24.3849 | 0.014392 | 0.23477 | 0.003379 | 16.3124 | 19.7189 | 3.3083 | 0.008296 | 23.1384 | 18.9613 | 23.1658 | 17.6744
2012-10-30 13:48 | 20.9266 | 42.7841 | 38.6028 | 42.8428 | 37.3472 | 47.9835 | 0.014194 | 0.34628 | 0.004915 | 24.3962 | 0.014379 | 0.23475 | 0.003376 | 16.3262 | 19.7079 | 3.308 | 0.008291 | 23.0762 | 18.8949 | 23.1349 | 17.6393
2012-10-30 13:48 | 20.9378 | 42.7133 | 38.567 | 42.7211 | 37.2328 | 47.9564 | 0.014197 | 0.34629 | 0.004916 | 24.3921 | 0.014392 | 0.23477 | 0.003379 | 16.3124 | 19.7072 | 3.3061 | 0.008295 | 23.0061 | 18.8598 | 23.0139 | 17.5256
2012-10-30 13:48 | 20.9401 | 42.6374 | 38.4908 | 42.6491 | 37.2116 | 47.9691 | 0.014194 | 0.34625 | 0.004915 | 24.3941 | 0.014398 | 0.23483 | 0.003381 | 16.3096 | 19.7055 | 3.307 | 0.008296 | 22.9319 | 18.7853 | 22.9436 | 17.5061
2012-10-30 13:48 | 20.9401 | 42.5631 | 38.4124 | 42.571 | 37.1213 | 47.9575 | 0.014181 | 0.34631 | 0.004911 | 24.4206 | 0.014392 | 0.23475 | 0.003379 | 16.3114 | 19.7333 | 3.3062 | 0.00829 | 22.8298 | 18.6791 | 22.8377 | 17.388
2012-10-30 13:48 | 20.9446 | 42.4582 | 38.3423 | 42.4777 | 37.0551 | 47.9292 | 0.014181 | 0.34624 | 0.00491 | 24.4155 | 0.014395 | 0.23481 | 0.00338 | 16.3121 | 19.7736 | 3.3042 | 0.00829 | 22.6846 | 18.5687 | 22.7041 | 17.2815
2012-10-30 13:48 | 20.9632 | 42.3869 | 38.2982 | 42.4142 | 37.0069 | 47.939 | 0.014181 | 0.34631 | 0.004911 | 24.4206 | 0.014382 | 0.23486 | 0.003378 | 16.3298 | 19.7366 | 3.3049 | 0.008289 | 22.6503 | 18.5616 | 22.6776 | 17.2703
2012-10-30 13:48 | 20.9707 | 42.3201 | 38.2115 | 42.324 | 37.0301 | 47.9263 | 0.014184 | 0.34632 | 0.004912 | 24.4165 | 0.014401 | 0.23477 | 0.003381 | 16.3022 | 19.7361 | 3.304 | 0.008293 | 22.584 | 18.4754 | 22.5879 | 17.294
2012-10-30 13:48 | 20.9782 | 42.1946 | 38.1602 | 42.2181 | 36.967 | 47.9257 | 0.014175 | 0.34625 | 0.004908 | 24.4268 | 0.014398 | 0.23474 | 0.00338 | 16.3036 | 19.7635 | 3.304 | 0.008288 | 22.4311 | 18.3967 | 22.4546 | 17.2035
2012-10-30 13:48 | 20.9864 | 42.1794 | 38.1291 | 42.0933 | 36.8534 | 47.924 | 0.014197 | 0.34625 | 0.004916 | 24.389 | 0.014408 | 0.2348 | 0.003383 | 16.2963 | 19.7836 | 3.3039 | 0.008299 | 22.3958 | 18.3455 | 22.3097 | 17.0698
2012-10-30 13:48 | 20.9901 | 42.0814 | 38.074 | 42.0345 | 36.7982 | 47.8991 | 0.014184 | 0.34631 | 0.004912 | 24.4154 | 0.014401 | 0.23478 | 0.003381 | 16.3032 | 19.7515 | 3.3022 | 0.008293 | 22.3299 | 18.3225 | 22.283 | 17.0467
2012-10-30 13:48 | 21.0028 | 41.9768 | 37.969 | 41.9611 | 36.7441 | 47.8876 | 0.014175 | 0.34629 | 0.004909 | 24.4299 | 0.014386 | 0.23478 | 0.003378 | 16.3202 | 19.7483 | 3.3014 | 0.008287 | 22.2285 | 18.2207 | 22.2128 | 16.9958
2012-10-30 13:48 | 21.0051 | 41.9047 | 37.9124 | 41.8891 | 36.715 | 47.8905 | 0.014184 | 0.34628 | 0.004912 | 24.4134 | 0.014392 | 0.23477 | 0.003379 | 16.3124 | 19.7864 | 3.3016 | 0.008291 | 22.1183 | 18.126 | 22.1027 | 16.9286
2012-10-30 13:48 | 21.0081 | 41.8178 | 37.833 | 41.7708 | 36.6866 | 47.8697 | 0.014184 | 0.34621 | 0.004911 | 24.4083 | 0.014398 | 0.23477 | 0.00338 | 16.3056 | 19.7775 | 3.3001 | 0.008291 | 22.0403 | 18.0555 | 21.9933 | 16.9091
2012-10-30 13:48 | 21.0245 | 41.7872 | 37.7709 | 41.7325 | 36.6127 | 47.8812 | 0.014194 | 0.34628 | 0.004915 | 24.3962 | 0.014404 | 0.23478 | 0.003382 | 16.2998 | 19.766 | 3.3009 | 0.008297 | 22.0212 | 18.0049 | 21.9665 | 16.8467
2012-10-30 13:48 | 21.0335 | 41.6711 | 37.6661 | 41.6632 | 36.6059 | 47.8674 | 0.014194 | 0.34631 | 0.004915 | 24.3982 | 0.014404 | 0.2349 | 0.003383 | 16.3078 | 19.771 | 3.3 | 0.008298 | 21.9001 | 17.8951 | 21.8922 | 16.8349
2012-10-30 13:48 | 21.0327 | 41.6116 | 37.6378 | 41.5646 | 36.4991 | 47.839 | 0.014203 | 0.34625 | 0.004918 | 24.3787 | 0.014389 | 0.23481 | 0.003379 | 16.3189 | 19.7583 | 3.298 | 0.008297 | 21.8533 | 17.8795 | 21.8063 | 16.7408
2012-10-30 13:48 | 21.0409 | 41.5102 | 37.6028 | 41.5063 | 36.413 | 47.8546 | 0.014206 | 0.34619 | 0.004918 | 24.3695 | 0.014398 | 0.23484 | 0.003381 | 16.3106 | 19.7864 | 3.2991 | 0.008299 | 21.7238 | 17.8164 | 21.7199 | 16.6266
2012-10-30 13:48 | 21.0566 | 41.475 | 37.5283 | 41.4242 | 36.3776 | 47.8275 | 0.0142 | 0.34622 | 0.004916 | 24.3818 | 0.014404 | 0.23481 | 0.003382 | 16.3019 | 19.7623 | 3.2972 | 0.008298 | 21.7127 | 17.766 | 21.6619 | 16.6153
2012-10-30 13:48 | 21.0529 | 41.3853 | 37.4735 | 41.3305 | 36.3268 | 47.8188 | 0.014194 | 0.34624 | 0.004914 | 24.3931 | 0.014395 | 0.23477 | 0.003379 | 16.309 | 19.7785 | 3.2966 | 0.008293 | 21.6068 | 17.695 | 21.552 | 16.5483
2012-10-30 13:48 | 21.0499 | 41.2805 | 37.4234 | 41.257 | 36.2452 | 47.8252 | 0.014181 | 0.34621 | 0.00491 | 24.4134 | 0.014398 | 0.2348 | 0.003381 | 16.3076 | 19.7477 | 3.2971 | 0.008291 | 21.5328 | 17.6757 | 21.5093 | 16.4975
2012-10-30 13:48 | 21.0693 | 41.2334 | 37.3604 | 41.1668 | 36.2096 | 47.8049 | 0.01419 | 0.34618 | 0.004912 | 24.3959 | 0.014389 | 0.23472 | 0.003377 | 16.3128 | 19.787 | 3.2957 | 0.008289 | 21.4464 | 17.5734 | 21.3798 | 16.4226
2012-10-30 13:48 | 21.0798 | 41.1264 | 37.2806 | 41.1186 | 36.1612 | 47.8015 | 0.014172 | 0.34625 | 0.004907 | 24.432 | 0.014408 | 0.2348 | 0.003383 | 16.2963 | 19.7497 | 3.2954 | 0.00829 | 21.3767 | 17.5309 | 21.3689 | 16.4115
2012-10-30 13:48 | 21.088 | 41.0642 | 37.2574 | 41.0133 | 36.0986 | 47.8026 | 0.014194 | 0.34624 | 0.004914 | 24.3931 | 0.014379 | 0.23475 | 0.003376 | 16.3262 | 19.7738 | 3.2955 | 0.00829 | 21.2904 | 17.4836 | 21.2395 | 16.3248
2012-10-30 13:48 | 21.0843 | 41.0018 | 37.2066 | 40.9548 | 36.0517 | 47.7731 | 0.014197 | 0.34618 | 0.004915 | 24.3839 | 0.014398 | 0.23475 | 0.00338 | 16.3046 | 19.7821 | 3.2935 | 0.008295 | 21.2197 | 17.4245 | 21.1727 | 16.2696
2012-10-30 13:48 | 21.088 | 40.8724 | 37.1317 | 40.8567 | 35.9807 | 47.7905 | 0.014187 | 0.34618 | 0.004911 | 24.4011 | 0.014386 | 0.23475 | 0.003377 | 16.3182 | 19.746 | 3.2947 | 0.008288 | 21.1264 | 17.3857 | 21.1107 | 16.2347
2012-10-30 13:48 | 21.103 | 40.8404 | 37.0603 | 40.7895 | 35.8778 | 47.7795 | 0.014184 | 0.34619 | 0.00491 | 24.4073 | 0.014386 | 0.23474 | 0.003377 | 16.3172 | 19.7928 | 3.2939 | 0.008287 | 21.0476 | 17.2675 | 20.9967 | 16.085
2012-10-30 13:48 | 21.1104 | 40.7695 | 36.9539 | 40.7381 | 35.901 | 47.7443 | 0.014187 | 0.34618 | 0.004911 | 24.4011 | 0.014408 | 0.23478 | 0.003383 | 16.2953 | 19.7883 | 3.2915 | 0.008294 | 20.9812 | 17.1656 | 20.9498 | 16.1127
2012-10-30 13:49 | 21.1246 | 40.7054 | 36.9641 | 40.6701 | 35.8208 | 47.7454 | 0.014203 | 0.34616 | 0.004917 | 24.3726 | 0.014392 | 0.23474 | 0.003378 | 16.3104 | 19.7587 | 3.2916 | 0.008295 | 20.9467 | 17.2054 | 20.9114 | 16.0621
2012-10-30 13:49 | 21.1306 | 40.6565 | 36.8877 | 40.5704 | 35.7757 | 47.746 | 0.014181 | 0.34622 | 0.00491 | 24.4144 | 0.014401 | 0.23471 | 0.00338 | 16.2982 | 19.7647 | 3.2916 | 0.00829 | 20.8918 | 17.123 | 20.8057 | 16.011
2012-10-30 13:49 | 21.1373 | 40.5497 | 36.8315 | 40.4753 | 35.6802 | 47.7339 | 0.014187 | 0.34618 | 0.004911 | 24.4011 | 0.014392 | 0.23477 | 0.003379 | 16.3124 | 19.7913 | 3.2908 | 0.00829 | 20.7584 | 17.0402 | 20.684 | 15.8889
2012-10-30 13:49 | 21.1642 | 40.2084 | 36.5481 | 40.1183 | 35.4162 | 47.6998 | 0.01419 | 0.34618 | 0.004912 | 24.3959 | 0.014392 | 0.23475 | 0.003379 | 16.3114 | 19.7784 | 3.2884 | 0.008291 | 20.43 | 16.7697 | 20.3399 | 15.6378
2012-10-30 13:49 | 21.1934 | 39.8732 | 36.2866 | 39.7831 | 35.1977 | 47.6593 | 0.0142 | 0.34622 | 0.004916 | 24.3818 | 0.014398 | 0.23477 | 0.00338 | 16.3056 | 19.7758 | 3.2856 | 0.008296 | 20.0974 | 16.5108 | 20.0073 | 15.4219
2012-10-30 13:49 | 21.2008 | 39.7983 | 36.2429 | 39.7003 | 35.1265 | 47.6385 | 0.014187 | 0.34609 | 0.00491 | 24.3949 | 0.014398 | 0.23475 | 0.00338 | 16.3046 | 19.8269 | 3.2842 | 0.00829 | 19.9714 | 16.416 | 19.8734 | 15.2996
2012-10-30 13:49 | 21.2076 | 39.6797 | 36.1514 | 39.6131 | 35.0586 | 47.6437 | 0.014197 | 0.34615 | 0.004914 | 24.3818 | 0.014386 | 0.23477 | 0.003377 | 16.3192 | 19.778 | 3.2846 | 0.008291 | 19.9017 | 16.3734 | 19.8351 | 15.2806
2012-10-30 13:49 | 21.2143 | 39.6511 | 36.0953 | 39.5688 | 35.0456 | 47.646 | 0.01419 | 0.34614 | 0.004912 | 24.3929 | 0.014395 | 0.23474 | 0.003379 | 16.307 | 19.7967 | 3.2847 | 0.008291 | 19.8544 | 16.2986 | 19.7721 | 15.2489
2012-10-30 13:49 | 21.2188 | 39.5811 | 36.0526 | 39.5027 | 34.9872 | 47.6223 | 0.014194 | 0.34614 | 0.004913 | 24.386 | 0.014389 | 0.23474 | 0.003378 | 16.3138 | 19.7773 | 3.2831 | 0.008291 | 19.8038 | 16.2753 | 19.7254 | 15.2099
2012-10-30 13:49 | 21.2285 | 39.5281 | 36.0269 | 39.4144 | 34.9143 | 47.6379 | 0.014194 | 0.34619 | 0.004914 | 24.3901 | 0.014398 | 0.23478 | 0.00338 | 16.3066 | 19.771 | 3.2842 | 0.008294 | 19.7571 | 16.2559 | 19.6434 | 15.1433
2012-10-30 13:49 | 21.2367 | 39.4697 | 35.9408 | 39.3678 | 34.8242 | 47.6188 | 0.014184 | 0.34615 | 0.00491 | 24.4042 | 0.014386 | 0.23475 | 0.003377 | 16.3182 | 19.8071 | 3.2828 | 0.008287 | 19.6626 | 16.1337 | 19.5607 | 15.0171
2012-10-30 13:49 | 21.2359 | 39.3748 | 35.9046 | 39.265 | 34.8391 | 47.5934 | 0.014194 | 0.34618 | 0.004914 | 24.389 | 0.014401 | 0.23474 | 0.00338 | 16.3002 | 19.7585 | 3.2811 | 0.008294 | 19.6163 | 16.1461 | 19.5065 | 15.0806
2012-10-30 13:49 | 21.2412 | 39.2939 | 35.8313 | 39.1841 | 34.7697 | 47.5957 | 0.014194 | 0.34616 | 0.004913 | 24.3881 | 0.014386 | 0.23477 | 0.003377 | 16.3192 | 19.8116 | 3.2812 | 0.00829 | 19.4823 | 16.0197 | 19.3725 | 14.9581
2012-10-30 13:49 | 21.2524 | 39.211 | 35.7796 | 39.1247 | 34.7455 | 47.5911 | 0.014187 | 0.34618 | 0.004911 | 24.4011 | 0.014382 | 0.23474 | 0.003376 | 16.3217 | 19.8188 | 3.2809 | 0.008287 | 19.3922 | 15.9608 | 19.3059 | 14.9267
2012-10-30 13:49 | 21.2569 | 39.1841 | 35.7409 | 39.0468 | 34.6909 | 47.5738 | 0.0142 | 0.34611 | 0.004915 | 24.3737 | 0.014389 | 0.23477 | 0.003378 | 16.3158 | 19.8233 | 3.2797 | 0.008293 | 19.3608 | 15.9176 | 19.2235 | 14.8676
2012-10-30 13:49 | 21.2621 | 39.0952 | 35.6793 | 38.9854 | 34.6214 | 47.5813 | 0.014184 | 0.34609 | 0.004909 | 24.4001 | 0.014401 | 0.23474 | 0.00338 | 16.3002 | 19.7927 | 3.2803 | 0.008289 | 19.3025 | 15.8866 | 19.1927 | 14.8287
2012-10-30 13:49 | 21.2815 | 39.0518 | 35.6279 | 38.9656 | 34.5385 | 47.5807 | 0.01419 | 0.34615 | 0.004912 | 24.3939 | 0.014398 | 0.23474 | 0.00338 | 16.3036 | 19.8161 | 3.2802 | 0.008292 | 19.2357 | 15.8118 | 19.1495 | 14.7224
2012-10-30 13:49 | 21.2838 | 38.9443 | 35.5791 | 38.8149 | 34.4857 | 47.5512 | 0.014203 | 0.34611 | 0.004916 | 24.3685 | 0.014386 | 0.23472 | 0.003377 | 16.3162 | 19.7826 | 3.2782 | 0.008293 | 19.1617 | 15.7965 | 19.0323 | 14.7031
2012-10-30 13:49 | 21.2942 | 38.8802 | 35.5226 | 38.739 | 34.3545 | 47.5449 | 0.014175 | 0.34616 | 0.004907 | 24.4207 | 0.014392 | 0.2348 | 0.003379 | 16.3144 | 19.8208 | 3.2777 | 0.008286 | 19.0594 | 15.7018 | 18.9182 | 14.5337
2012-10-30 13:50 | 21.2853 | 38.8086 | 35.4508 | 38.6948 | 34.3181 | 47.5524 | 0.014184 | 0.34611 | 0.004909 | 24.4012 | 0.014398 | 0.23465 | 0.003379 | 16.2976 | 19.8199 | 3.2783 | 0.008288 | 18.9887 | 15.6309 | 18.8749 | 14.4982
2012-10-30 13:50 | 21.2972 | 38.7342 | 35.4077 | 38.6557 | 34.2591 | 47.5322 | 0.014194 | 0.34615 | 0.004913 | 24.387 | 0.014395 | 0.2347 | 0.003378 | 16.304 | 19.8159 | 3.2769 | 0.008291 | 18.9183 | 15.5918 | 18.8398 | 14.4432
2012-10-30 13:50 | 21.3039 | 38.6349 | 35.32 | 38.5643 | 34.258 | 47.5183 | 0.014197 | 0.34609 | 0.004913 | 24.3778 | 0.014392 | 0.23474 | 0.003378 | 16.3104 | 19.8186 | 3.2759 | 0.008291 | 18.8163 | 15.5014 | 18.7457 | 14.4394
2012-10-30 13:50 | 21.3062 | 38.5823 | 35.2948 | 38.4842 | 34.154 | 47.5356 | 0.014184 | 0.34605 | 0.004908 | 24.3971 | 0.014398 | 0.23471 | 0.003379 | 16.3016 | 19.8169 | 3.2771 | 0.008287 | 18.7654 | 15.4779 | 18.6673 | 14.3371
2012-10-30 13:50 | 21.3144 | 38.5513 | 35.189 | 38.4218 | 34.115 | 47.5142 | 0.014194 | 0.34611 | 0.004913 | 24.384 | 0.014401 | 0.2347 | 0.00338 | 16.2972 | 19.8211 | 3.2756 | 0.008293 | 18.7302 | 15.3679 | 18.6007 | 14.2939
2012-10-30 13:50 | 21.3196 | 38.4898 | 35.1273 | 38.3603 | 34.0494 | 47.5061 | 0.01419 | 0.34609 | 0.004911 | 24.3898 | 0.014389 | 0.23471 | 0.003377 | 16.3118 | 19.7984 | 3.2751 | 0.008288 | 18.6914 | 15.3289 | 18.5619 | 14.251
2012-10-30 13:50 | 21.3204 | 38.3925 | 35.073 | 38.2709 | 33.9715 | 47.505 | 0.014187 | 0.34609 | 0.00491 | 24.3949 | 0.014386 | 0.23477 | 0.003377 | 16.3192 | 19.8629 | 3.275 | 0.008287 | 18.5296 | 15.2101 | 18.408 | 14.1086
2012-10-30 13:50 | 21.3338 | 38.351 | 35.0353 | 38.1901 | 33.977 | 47.4934 | 0.014203 | 0.34618 | 0.004917 | 24.3736 | 0.014389 | 0.23465 | 0.003376 | 16.3078 | 19.7688 | 3.2742 | 0.008293 | 18.5822 | 15.2665 | 18.4213 | 14.2082
2012-10-30 13:50 | 21.3473 | 38.2546 | 34.9623 | 38.1251 | 33.9 | 47.4749 | 0.01419 | 0.34612 | 0.004911 | 24.3918 | 0.014386 | 0.23471 | 0.003377 | 16.3152 | 19.8301 | 3.2729 | 0.008288 | 18.4245 | 15.1322 | 18.295 | 14.0699
2012-10-30 13:50 | 21.3346 | 38.1791 | 34.8945 | 38.0535 | 33.7889 | 47.49 | 0.014197 | 0.34606 | 0.004913 | 24.3757 | 0.014389 | 0.23471 | 0.003377 | 16.3118 | 19.8214 | 3.274 | 0.00829 | 18.3577 | 15.0731 | 18.2321 | 13.9675
2012-10-30 13:50 | 21.354 | 38.1632 | 34.8471 | 38.0062 | 33.7807 | 47.4738 | 0.014203 | 0.34609 | 0.004916 | 24.3675 | 0.014392 | 0.23474 | 0.003378 | 16.3104 | 19.8528 | 3.2728 | 0.008294 | 18.3104 | 14.9943 | 18.1534 | 13.9279
2012-10-30 13:50 | 21.3592 | 38.0232 | 34.7775 | 37.9172 | 33.7426 | 47.4628 | 0.014184 | 0.34614 | 0.00491 | 24.4032 | 0.014386 | 0.23472 | 0.003377 | 16.3162 | 19.8341 | 3.2721 | 0.008287 | 18.1891 | 14.9434 | 18.0831 | 13.9085
2012-10-30 13:50 | 21.3637 | 38.0238 | 34.7348 | 37.8589 | 33.6684 | 47.4599 | 0.014203 | 0.34611 | 0.004916 | 24.3685 | 0.014392 | 0.23471 | 0.003378 | 16.3084 | 19.8505 | 3.2719 | 0.008294 | 18.1733 | 14.8843 | 18.0084 | 13.8179
2012-10-30 13:50 | 21.3727 | 37.9268 | 34.6612 | 37.8012 | 33.6026 | 47.4432 | 0.014181 | 0.34606 | 0.004908 | 24.4032 | 0.014395 | 0.23474 | 0.003379 | 16.307 | 19.8555 | 3.2707 | 0.008287 | 18.0713 | 14.8057 | 17.9457 | 13.7471
2012-10-30 13:50 | 21.3757 | 37.8905 | 34.6052 | 37.7571 | 33.6095 | 47.4241 | 0.014209 | 0.34614 | 0.004918 | 24.3603 | 0.014398 | 0.23475 | 0.00338 | 16.3046 | 19.8506 | 3.2694 | 0.008298 | 18.0399 | 14.7546 | 17.9065 | 13.7589
2012-10-30 13:50 | 21.3757 | 37.8238 | 34.5776 | 37.6746 | 33.5387 | 47.4611 | 0.014197 | 0.34608 | 0.004913 | 24.3768 | 0.014382 | 0.23475 | 0.003376 | 16.3228 | 19.8426 | 3.272 | 0.008289 | 17.9812 | 14.735 | 17.832 | 13.6961
2012-10-30 13:50 | 21.3891 | 37.7666 | 34.5084 | 37.6252 | 33.4812 | 47.4235 | 0.014194 | 0.34608 | 0.004912 | 24.3819 | 0.014398 | 0.23475 | 0.00338 | 16.3046 | 19.8361 | 3.2694 | 0.008292 | 17.9305 | 14.6723 | 17.7891 | 13.6451
2012-10-30 13:50 | 21.3973 | 37.653 | 34.4065 | 37.5588 | 33.4225 | 47.4223 | 0.014187 | 0.34606 | 0.00491 | 24.3929 | 0.014401 | 0.23478 | 0.003381 | 16.3032 | 19.8483 | 3.2693 | 0.008291 | 17.8047 | 14.5582 | 17.7105 | 13.5742
2012-10-30 13:50 | 21.4041 | 37.6324 | 34.4015 | 37.491 | 33.3663 | 47.4299 | 0.01419 | 0.34615 | 0.004912 | 24.3939 | 0.014392 | 0.2347 | 0.003378 | 16.3074 | 19.8272 | 3.2698 | 0.00829 | 17.8052 | 14.5743 | 17.6638 | 13.5391
2012-10-30 13:50 | 21.4041 | 37.5342 | 34.3346 | 37.44 | 33.3466 | 47.3987 | 0.01419 | 0.34608 | 0.004911 | 24.3888 | 0.014389 | 0.23472 | 0.003377 | 16.3128 | 19.8272 | 3.2677 | 0.008288 | 17.707 | 14.5074 | 17.6128 | 13.5194
2012-10-30 13:50 | 21.4115 | 37.4317 | 34.2947 | 37.3335 | 33.2674 | 47.4015 | 0.014203 | 0.34603 | 0.004915 | 24.3634 | 0.014392 | 0.23475 | 0.003379 | 16.3114 | 19.8147 | 3.2679 | 0.008294 | 17.617 | 14.48 | 17.5188 | 13.4527
2012-10-30 13:50 | 21.419 | 37.3803 | 34.2629 | 37.2507 | 33.2119 | 47.3998 | 0.014194 | 0.34609 | 0.004912 | 24.3829 | 0.014382 | 0.23472 | 0.003376 | 16.3207 | 19.8302 | 3.2677 | 0.008288 | 17.5501 | 14.4327 | 17.4205 | 13.3817
2012-10-30 13:50 | 21.4183 | 37.3403 | 34.2228 | 37.2107 | 33.1206 | 47.4044 | 0.014178 | 0.34601 | 0.004906 | 24.4044 | 0.014398 | 0.23474 | 0.00338 | 16.3036 | 19.8135 | 3.2681 | 0.008286 | 17.5268 | 14.4093 | 17.3972 | 13.3071
2012-10-30 13:50 | 21.4295 | 37.2612 | 34.1278 | 37.1512 | 33.0964 | 47.3605 | 0.014194 | 0.34611 | 0.004913 | 24.384 | 0.014386 | 0.23475 | 0.003377 | 16.3182 | 19.8287 | 3.265 | 0.00829 | 17.4325 | 14.2991 | 17.3225 | 13.2677
2012-10-30 13:50 | 21.4317 | 37.2162 | 34.0906 | 37.1298 | 33.0592 | 47.3813 | 0.014194 | 0.34609 | 0.004912 | 24.3829 | 0.014386 | 0.23477 | 0.003377 | 16.3192 | 19.8349 | 3.2665 | 0.008289 | 17.3813 | 14.2557 | 17.2949 | 13.2243
2012-10-30 13:50 | 21.4496 | 37.1438 | 33.9983 | 37.0456 | 32.9865 | 47.3657 | 0.014187 | 0.34609 | 0.00491 | 24.3949 | 0.014395 | 0.2347 | 0.003378 | 16.304 | 19.8169 | 3.2654 | 0.008288 | 17.3269 | 14.1814 | 17.2287 | 13.1696
2012-10-30 13:51 | 21.4556 | 37.1145 | 33.9335 | 36.9141 | 32.9098 | 47.3484 | 0.014187 | 0.34609 | 0.00491 | 24.3949 | 0.014392 | 0.23471 | 0.003378 | 16.3084 | 19.8468 | 3.2642 | 0.008288 | 17.2677 | 14.0867 | 17.0673 | 13.063
2012-10-30 13:51 | 21.4579 | 37.0303 | 33.8846 | 36.9321 | 32.8649 | 47.353 | 0.01419 | 0.34603 | 0.00491 | 24.3858 | 0.014401 | 0.23472 | 0.00338 | 16.2992 | 19.8133 | 3.2645 | 0.00829 | 17.217 | 14.0713 | 17.1188 | 13.0516
2012-10-30 13:51 | 21.4564 | 36.9581 | 33.8241 | 36.7892 | 32.8279 | 47.3628 | 0.01419 | 0.34605 | 0.00491 | 24.3868 | 0.014392 | 0.23468 | 0.003378 | 16.3064 | 19.8357 | 3.2652 | 0.008288 | 17.1224 | 13.9884 | 16.9535 | 12.9922
2012-10-30 13:51 | 21.4616 | 36.8691 | 33.7545 | 36.7355 | 32.7898 | 47.3195 | 0.014184 | 0.34609 | 0.004909 | 24.4001 | 0.014392 | 0.23471 | 0.003378 | 16.3084 | 19.825 | 3.2622 | 0.008287 | 17.0441 | 13.9295 | 16.9105 | 12.9648
2012-10-30 13:51 | 21.4758 | 36.844 | 33.7136 | 36.6633 | 32.7252 | 47.3316 | 0.014194 | 0.34601 | 0.004911 | 24.3768 | 0.014398 | 0.23471 | 0.003379 | 16.3016 | 19.8192 | 3.263 | 0.00829 | 17.0248 | 13.8944 | 16.8441 | 12.906
2012-10-30 13:51 | 21.4863 | 36.7366 | 33.5902 | 36.5991 | 32.5939 | 47.3056 | 0.014187 | 0.34598 | 0.004908 | 24.3868 | 0.014404 | 0.23468 | 0.00338 | 16.2928 | 19.8218 | 3.2612 | 0.008288 | 16.9148 | 13.7684 | 16.7773 | 12.7721
2012-10-30 13:51 | 21.4863 | 36.6541 | 33.5548 | 36.5441 | 32.5033 | 47.3305 | 0.014178 | 0.34603 | 0.004906 | 24.4064 | 0.014392 | 0.23475 | 0.003379 | 16.3114 | 19.8616 | 3.263 | 0.008285 | 16.7925 | 13.6932 | 16.6825 | 12.6417
2012-10-30 13:51 | 21.4863 | 36.6109 | 33.5036 | 36.5087 | 32.5388 | 47.29 | 0.01419 | 0.34606 | 0.004911 | 24.3877 | 0.014389 | 0.2347 | 0.003377 | 16.3108 | 19.8297 | 3.2602 | 0.008288 | 16.7812 | 13.6739 | 16.679 | 12.7091
2012-10-30 13:51 | 21.4975 | 36.5553 | 33.4754 | 36.4178 | 32.4988 | 47.305 | 0.014187 | 0.34601 | 0.004909 | 24.3889 | 0.014398 | 0.23472 | 0.00338 | 16.3026 | 19.8529 | 3.2612 | 0.008289 | 16.7024 | 13.6225 | 16.5649 | 12.6459
2012-10-30 13:51 | 21.5049 | 36.5037 | 33.4238 | 36.4055 | 32.4904 | 47.3166 | 0.014194 | 0.34601 | 0.004911 | 24.3768 | 0.014389 | 0.23472 | 0.003377 | 16.3128 | 19.8244 | 3.262 | 0.008288 | 16.6793 | 13.5994 | 16.5811 | 12.666
2012-10-30 13:51 | 21.5094 | 36.3982 | 33.385 | 36.3353 | 32.4398 | 47.2784 | 0.014187 | 0.34605 | 0.004909 | 24.3919 | 0.014411 | 0.23484 | 0.003384 | 16.2959 | 19.825 | 3.2594 | 0.008293 | 16.5732 | 13.56 | 16.5103 | 12.6148
2012-10-30 13:51 | 21.5094 | 36.4139 | 33.3653 | 36.2725 | 32.4082 | 47.294 | 0.014175 | 0.34599 | 0.004904 | 24.4085 | 0.014386 | 0.23477 | 0.003377 | 16.3192 | 19.8369 | 3.2604 | 0.008281 | 16.577 | 13.5284 | 16.4356 | 12.5713
2012-10-30 13:51 | 21.5251 | 36.351 | 33.3022 | 36.241 | 32.3885 | 47.2721 | 0.014184 | 0.34602 | 0.004908 | 24.395 | 0.014398 | 0.23471 | 0.003379 | 16.3016 | 19.8207 | 3.2589 | 0.008287 | 16.5303 | 13.4815 | 16.4203 | 12.5678
2012-10-30 13:51 | 21.5259 | 36.2339 | 33.244 | 36.1632 | 32.3144 | 47.2669 | 0.0142 | 0.34601 | 0.004913 | 24.3665 | 0.014389 | 0.23468 | 0.003377 | 16.3098 | 19.8295 | 3.2586 | 0.00829 | 16.4044 | 13.4145 | 16.3337 | 12.4849
2012-10-30 13:51 | 21.5281 | 36.1968 | 33.2147 | 36.0632 | 32.2812 | 47.2808 | 0.014194 | 0.34602 | 0.004911 | 24.3778 | 0.014395 | 0.2347 | 0.003378 | 16.304 | 19.7998 | 3.2595 | 0.008289 | 16.397 | 13.4149 | 16.2634 | 12.4814
2012-10-30 13:51 | 21.5303 | 36.1479 | 33.1105 | 36.0457 | 32.2006 | 47.2571 | 0.01419 | 0.34608 | 0.004911 | 24.3888 | 0.014401 | 0.23467 | 0.003379 | 16.2952 | 19.7941 | 3.2579 | 0.00829 | 16.3538 | 13.3164 | 16.2516 | 12.4065
2012-10-30 13:51 | 21.5401 | 36.0948 | 33.1203 | 35.9887 | 32.1868 | 47.2721 | 0.01419 | 0.34601 | 0.00491 | 24.3837 | 0.014411 | 0.23468 | 0.003382 | 16.2849 | 19.8596 | 3.2589 | 0.008292 | 16.2352 | 13.2607 | 16.1291 | 12.3272
2012-10-30 13:51 | 21.549 | 36.0487 | 33.0899 | 35.9622 | 32.1169 | 47.2634 | 0.014206 | 0.34603 | 0.004916 | 24.3583 | 0.014398 | 0.23471 | 0.003379 | 16.3016 | 19.8287 | 3.2583 | 0.008295 | 16.22 | 13.2612 | 16.1335 | 12.2882
2012-10-30 13:51 | 21.5513 | 35.9763 | 33.0607 | 35.8426 | 32.0916 | 47.2553 | 0.014181 | 0.34599 | 0.004906 | 24.3981 | 0.014395 | 0.2347 | 0.003378 | 16.304 | 19.823 | 3.2578 | 0.008284 | 16.1533 | 13.2377 | 16.0196 | 12.2686
2012-10-30 13:51 | 21.5647 | 35.9386 | 32.952 | 35.8167 | 32.0656 | 47.2461 | 0.0142 | 0.34603 | 0.004914 | 24.3686 | 0.014382 | 0.23474 | 0.003376 | 16.3217 | 19.8564 | 3.2571 | 0.00829 | 16.0822 | 13.0956 | 15.9603 | 12.2092
2012-10-30 13:51 | 21.5595 | 35.8784 | 32.9389 | 35.7408 | 32.0328 | 47.2311 | 0.014181 | 0.34599 | 0.004906 | 24.3981 | 0.014389 | 0.23465 | 0.003376 | 16.3078 | 19.8352 | 3.2561 | 0.008282 | 16.0432 | 13.1037 | 15.9056 | 12.1976
2012-10-30 13:51 | 21.5707 | 35.7873 | 32.8871 | 35.7362 | 32.0086 | 47.2299 | 0.014187 | 0.34603 | 0.004909 | 24.3909 | 0.014389 | 0.23468 | 0.003377 | 16.3098 | 19.8424 | 3.256 | 0.008286 | 15.9449 | 13.0447 | 15.8938 | 12.1662
2012-10-30 13:51 | 21.5834 | 35.7765 | 32.8289 | 35.6585 | 31.9582 | 47.2415 | 0.0142 | 0.34605 | 0.004914 | 24.3696 | 0.014398 | 0.23468 | 0.003379 | 16.2996 | 19.8273 | 3.2568 | 0.008293 | 15.9492 | 13.0016 | 15.8312 | 12.1309
2012-10-30 13:51 | 21.5856 | 35.7197 | 32.7878 | 35.6175 | 31.9368 | 47.2212 | 0.014175 | 0.34596 | 0.004904 | 24.4064 | 0.014395 | 0.2347 | 0.003378 | 16.304 | 19.8295 | 3.2554 | 0.008282 | 15.8902 | 12.9583 | 15.788 | 12.1073
2012-10-30 13:51 | 21.5968 | 35.6483 | 32.799 | 35.5343 | 31.8889 | 47.1993 | 0.014194 | 0.34602 | 0.004911 | 24.3778 | 0.014404 | 0.2347 | 0.003381 | 16.2938 | 19.7968 | 3.2539 | 0.008292 | 15.8515 | 13.0022 | 15.7375 | 12.0921
2012-10-30 13:51 | 21.5894 | 35.6213 | 32.7128 | 35.4797 | 31.7987 | 47.2235 | 0.014184 | 0.34595 | 0.004907 | 24.3899 | 0.014382 | 0.23471 | 0.003376 | 16.3197 | 19.8452 | 3.2556 | 0.008283 | 15.7761 | 12.8676 | 15.6345 | 11.9535
2012-10-30 13:52 | 21.6043 | 35.5496 | 32.6765 | 35.4317 | 31.786 | 47.2178 | 0.014175 | 0.34601 | 0.004905 | 24.4095 | 0.014389 | 0.2347 | 0.003377 | 16.3108 | 19.7963 | 3.2552 | 0.008282 | 15.7533 | 12.8802 | 15.6354 | 11.9897
2012-10-30 13:52 | 21.6021 | 35.4963 | 32.6467 | 35.3429 | 31.7444 | 47.1941 | 0.01419 | 0.34606 | 0.004911 | 24.3877 | 0.014401 | 0.23472 | 0.00338 | 16.2992 | 19.838 | 3.2536 | 0.008291 | 15.6583 | 12.8087 | 15.5049 | 11.9064
2012-10-30 13:52 | 21.6013 | 35.3972 | 32.5514 | 35.2556 | 31.7554 | 47.2068 | 0.014181 | 0.34598 | 0.004906 | 24.3972 | 0.014389 | 0.23474 | 0.003378 | 16.3138 | 19.8252 | 3.2544 | 0.008284 | 15.572 | 12.7262 | 15.4304 | 11.9302
2012-10-30 13:52 | 21.6043 | 35.3923 | 32.4874 | 35.3019 | 31.6836 | 47.1877 | 0.014184 | 0.34595 | 0.004907 | 24.3899 | 0.014389 | 0.23472 | 0.003377 | 16.3128 | 19.8402 | 3.2531 | 0.008284 | 15.5521 | 12.6472 | 15.4617 | 11.8434
2012-10-30 13:52 | 21.6208 | 35.342 | 32.5118 | 35.1925 | 31.6567 | 47.1704 | 0.014203 | 0.34601 | 0.004914 | 24.3614 | 0.014376 | 0.23461 | 0.003373 | 16.3195 | 19.8208 | 3.2519 | 0.008287 | 15.5212 | 12.691 | 15.3717 | 11.8359
2012-10-30 13:52 | 21.6252 | 35.2323 | 32.4216 | 35.1576 | 31.6217 | 47.1779 | 0.014194 | 0.34602 | 0.004911 | 24.3778 | 0.014392 | 0.23471 | 0.003378 | 16.3084 | 19.8491 | 3.2524 | 0.008289 | 15.3832 | 12.5725 | 15.3085 | 11.7726
2012-10-30 13:52 | 21.6327 | 35.2202 | 32.4055 | 35.1021 | 31.6016 | 47.1594 | 0.014181 | 0.34599 | 0.004906 | 24.3981 | 0.014398 | 0.23465 | 0.003379 | 16.2976 | 19.8208 | 3.2512 | 0.008285 | 15.3994 | 12.5847 | 15.2813 | 11.7808
2012-10-30 13:52 | 21.6417 | 35.1505 | 32.3869 | 35.0718 | 31.5673 | 47.16 | 0.014203 | 0.34595 | 0.004913 | 24.3573 | 0.014382 | 0.23471 | 0.003376 | 16.3197 | 19.8616 | 3.2512 | 0.008289 | 15.2889 | 12.5253 | 15.2102 | 11.7057
2012-10-30 13:52 | 21.6379 | 35.1349 | 32.3398 | 34.9972 | 31.5595 | 47.1767 | 0.01419 | 0.34598 | 0.004909 | 24.3817 | 0.014404 | 0.23467 | 0.00338 | 16.2918 | 19.8698 | 3.2524 | 0.008289 | 15.2651 | 12.47 | 15.1274 | 11.6897
2012-10-30 13:52 | 21.6417 | 35.0364 | 32.3002 | 34.9263 | 31.4884 | 47.1449 | 0.014194 | 0.34592 | 0.00491 | 24.3707 | 0.014389 | 0.23471 | 0.003377 | 16.3118 | 19.8218 | 3.2502 | 0.008287 | 15.2146 | 12.4784 | 15.1045 | 11.6666
2012-10-30 13:52 | 21.6506 | 35.006 | 32.2343 | 34.8761 | 31.4816 | 47.1507 | 0.014181 | 0.34601 | 0.004907 | 24.3992 | 0.014392 | 0.23467 | 0.003377 | 16.3054 | 19.8546 | 3.2506 | 0.008284 | 15.1514 | 12.3797 | 15.0215 | 11.627
2012-10-30 13:52 | 21.6559 | 34.9601 | 32.1844 | 34.8303 | 31.4356 | 47.1565 | 0.0142 | 0.34599 | 0.004913 | 24.3655 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.824 | 3.251 | 0.008291 | 15.1361 | 12.3604 | 15.0063 | 11.6116
2012-10-30 13:52 | 21.6693 | 34.9342 | 32.2214 | 34.8162 | 31.4254 | 47.1397 | 0.014187 | 0.34593 | 0.004908 | 24.3837 | 0.014395 | 0.23468 | 0.003378 | 16.303 | 19.8056 | 3.2498 | 0.008286 | 15.1286 | 12.4158 | 15.0106 | 11.6198
2012-10-30 13:52 | 21.6678 | 34.8815 | 32.1608 | 34.7124 | 31.3805 | 47.1334 | 0.014187 | 0.34592 | 0.004908 | 24.3827 | 0.014389 | 0.23472 | 0.003377 | 16.3128 | 19.7722 | 3.2494 | 0.008285 | 15.1093 | 12.3886 | 14.9402 | 11.6083
2012-10-30 13:52 | 21.6716 | 34.7713 | 32.1331 | 34.6611 | 31.3331 | 47.1438 | 0.014206 | 0.34595 | 0.004915 | 24.3522 | 0.014382 | 0.23465 | 0.003375 | 16.3157 | 19.7919 | 3.2501 | 0.00829 | 14.9794 | 12.3412 | 14.8692 | 11.5412
2012-10-30 13:52 | 21.6753 | 34.7277 | 32.0462 | 34.5703 | 31.3131 | 47.1108 | 0.014194 | 0.34603 | 0.004912 | 24.3789 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 19.8116 | 3.2478 | 0.008289 | 14.9161 | 12.2346 | 14.7587 | 11.5015
2012-10-30 13:52 | 21.697 | 34.6865 | 32.0088 | 34.5133 | 31.2836 | 47.1195 | 0.014194 | 0.34592 | 0.00491 | 24.3707 | 0.014398 | 0.23464 | 0.003378 | 16.2966 | 19.8014 | 3.2484 | 0.008288 | 14.8851 | 12.2074 | 14.7119 | 11.4822
2012-10-30 13:52 | 21.6962 | 34.6581 | 31.9686 | 34.4653 | 31.1961 | 47.1004 | 0.01419 | 0.34598 | 0.004909 | 24.3817 | 0.014389 | 0.23478 | 0.003378 | 16.3168 | 19.8125 | 3.2471 | 0.008287 | 14.8456 | 12.1561 | 14.6528 | 11.3836
2012-10-30 13:52 | 21.6977 | 34.5377 | 31.9267 | 34.4393 | 31.233 | 47.1039 | 0.014194 | 0.34593 | 0.00491 | 24.3717 | 0.014386 | 0.23465 | 0.003376 | 16.3112 | 19.81 | 3.2473 | 0.008286 | 14.7277 | 12.1167 | 14.6293 | 11.423
2012-10-30 13:52 | 21.703 | 34.5311 | 31.8532 | 34.3501 | 31.1358 | 47.0854 | 0.01419 | 0.34593 | 0.004909 | 24.3786 | 0.014408 | 0.23465 | 0.003381 | 16.2863 | 19.8472 | 3.2461 | 0.00829 | 14.6839 | 12.006 | 14.5029 | 11.2886
2012-10-30 13:52 | 21.7082 | 34.4813 | 31.8702 | 34.2923 | 31.0977 | 47.097 | 0.014181 | 0.34599 | 0.004906 | 24.3981 | 0.014404 | 0.23465 | 0.00338 | 16.2908 | 19.8325 | 3.2469 | 0.008286 | 14.6488 | 12.0377 | 14.4598 | 11.2652
2012-10-30 13:52 | 21.7112 | 34.4174 | 31.8062 | 34.2402 | 31.077 | 47.0918 | 0.014181 | 0.34595 | 0.004906 | 24.3951 | 0.014395 | 0.23462 | 0.003377 | 16.299 | 19.8315 | 3.2465 | 0.008283 | 14.5859 | 11.9747 | 14.4087 | 11.2455
2012-10-30 13:52 | 21.7201 | 34.383 | 31.7245 | 34.2019 | 31.0386 | 47.0698 | 0.014184 | 0.34593 | 0.004907 | 24.3889 | 0.014386 | 0.23465 | 0.003376 | 16.3112 | 19.8125 | 3.245 | 0.008283 | 14.5705 | 11.912 | 14.3894 | 11.2261
2012-10-30 13:52 | 21.7299 | 34.3259 | 31.7658 | 34.1527 | 31.0248 | 47.0692 | 0.014184 | 0.34595 | 0.004907 | 24.3899 | 0.014395 | 0.23472 | 0.003379 | 16.306 | 19.8741 | 3.245 | 0.008286 | 14.4518 | 11.8917 | 14.2786 | 11.1507
2012-10-30 13:52 | 21.7351 | 34.2838 | 31.7001 | 34.091 | 30.9708 | 47.0635 | 0.01419 | 0.34595 | 0.004909 | 24.3796 | 0.014392 | 0.23467 | 0.003377 | 16.3054 | 19.8474 | 3.2446 | 0.008286 | 14.4364 | 11.8527 | 14.2436 | 11.1234
2012-10-30 13:52 | 21.7321 | 34.2375 | 31.6577 | 34.0683 | 30.9284 | 47.0611 | 0.014187 | 0.34593 | 0.004908 | 24.3837 | 0.014398 | 0.23465 | 0.003379 | 16.2976 | 19.8325 | 3.2444 | 0.008287 | 14.405 | 11.8252 | 14.2358 | 11.0959
2012-10-30 13:52 | 21.7448 | 34.1833 | 31.6073 | 33.9944 | 30.9017 | 47.06 | 0.014175 | 0.34589 | 0.004903 | 24.4013 | 0.014395 | 0.2347 | 0.003378 | 16.304 | 19.8173 | 3.2443 | 0.008281 | 14.366 | 11.79 | 14.1771 | 11.0844
2012-10-30 13:52 | 21.747 | 34.0792 | 31.5858 | 33.9651 | 30.8723 | 47.0646 | 0.014197 | 0.34592 | 0.004911 | 24.3656 | 0.014404 | 0.23464 | 0.00338 | 16.2898 | 19.8155 | 3.2446 | 0.008291 | 14.2637 | 11.7703 | 14.1496 | 11.0568
2012-10-30 13:53 | 21.7523 | 34.0727 | 31.5478 | 33.8917 | 30.7593 | 47.0548 | 0.01419 | 0.34598 | 0.004909 | 24.3817 | 0.014395 | 0.23474 | 0.003379 | 16.307 | 19.8009 | 3.244 | 0.008288 | 14.2718 | 11.7469 | 14.0908 | 10.9584
2012-10-30 13:53 | 21.7635 | 34.0721 | 31.5669 | 33.8556 | 30.7469 | 47.0675 | 0.014197 | 0.34598 | 0.004912 | 24.3697 | 0.014379 | 0.23461 | 0.003373 | 16.3161 | 19.8838 | 3.2448 | 0.008285 | 14.1883 | 11.6831 | 13.9718 | 10.8631
2012-10-30 13:53 | 21.765 | 34.0264 | 31.5092 | 33.7823 | 30.7444 | 47.019 | 0.01419 | 0.34596 | 0.004909 | 24.3806 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 19.8255 | 3.2415 | 0.008285 | 14.2009 | 11.6837 | 13.9568 | 10.9189
2012-10-30 13:53 | 21.7672 | 33.9184 | 31.4641 | 33.7609 | 30.6993 | 47.0496 | 0.014175 | 0.34592 | 0.004903 | 24.4034 | 0.014398 | 0.23468 | 0.003379 | 16.2996 | 19.8357 | 3.2436 | 0.008282 | 14.0827 | 11.6284 | 13.9252 | 10.8636
2012-10-30 13:53 | 21.7642 | 33.8878 | 31.4296 | 33.7264 | 30.6963 | 47.0288 | 0.014181 | 0.3459 | 0.004905 | 24.3921 | 0.014395 | 0.23468 | 0.003378 | 16.303 | 19.8247 | 3.2422 | 0.008283 | 14.0631 | 11.6049 | 13.9017 | 10.8716
2012-10-30 13:53 | 21.7732 | 33.8417 | 31.3795 | 33.637 | 30.6541 | 47.0184 | 0.01419 | 0.34593 | 0.004909 | 24.3786 | 0.014401 | 0.23468 | 0.00338 | 16.2962 | 19.8178 | 3.2414 | 0.008289 | 14.0239 | 11.5617 | 13.8192 | 10.8363
2012-10-30 13:53 | 21.7867 | 33.7961 | 31.3536 | 33.6268 | 30.636 | 47.0265 | 0.014178 | 0.34589 | 0.004904 | 24.3962 | 0.014389 | 0.23465 | 0.003376 | 16.3078 | 19.8073 | 3.242 | 0.00828 | 13.9888 | 11.5463 | 13.8195 | 10.8287
2012-10-30 13:53 | 21.7814 | 33.7042 | 31.3088 | 33.5507 | 30.5952 | 46.9947 | 0.014194 | 0.34592 | 0.00491 | 24.3707 | 0.014404 | 0.23462 | 0.00338 | 16.2888 | 19.7861 | 3.2398 | 0.00829 | 13.9181 | 11.5227 | 13.7646 | 10.8091
2012-10-30 13:53 | 21.7881 | 33.6715 | 31.2998 | 33.514 | 30.5822 | 47.0294 | 0.014194 | 0.3459 | 0.00491 | 24.3697 | 0.014392 | 0.23464 | 0.003377 | 16.3034 | 19.8207 | 3.2422 | 0.008287 | 13.8508 | 11.4791 | 13.6933 | 10.7615
2012-10-30 13:53 | 21.7941 | 33.6697 | 31.2585 | 33.5043 | 30.5606 | 47.0086 | 0.014187 | 0.34585 | 0.004907 | 24.3777 | 0.014382 | 0.23462 | 0.003374 | 16.3137 | 19.8427 | 3.2408 | 0.008281 | 13.827 | 11.4158 | 13.6616 | 10.7179
2012-10-30 13:53 | 21.7941 | 33.6539 | 31.2309 | 33.4334 | 30.5093 | 47.0034 | 0.014184 | 0.3459 | 0.004906 | 24.3869 | 0.014389 | 0.23464 | 0.003376 | 16.3068 | 19.8227 | 3.2404 | 0.008282 | 13.8312 | 11.4082 | 13.6107 | 10.6866
2012-10-30 13:53 | 21.8068 | 33.5564 | 31.2002 | 33.387 | 30.4668 | 47.0201 | 0.01419 | 0.34589 | 0.004908 | 24.3755 | 0.014389 | 0.23468 | 0.003377 | 16.3098 | 19.8833 | 3.2416 | 0.008285 | 13.6731 | 11.3169 | 13.5037 | 10.5835
2012-10-30 13:53 | 21.8068 | 33.5091 | 31.1805 | 33.328 | 30.4392 | 46.9785 | 0.014184 | 0.34592 | 0.004907 | 24.3879 | 0.014379 | 0.23461 | 0.003373 | 16.3161 | 19.8673 | 3.2387 | 0.00828 | 13.6418 | 11.3132 | 13.4607 | 10.5719
2012-10-30 13:53 | 21.8136 | 33.4962 | 31.1242 | 33.3033 | 30.446 | 47.0028 | 0.014203 | 0.34599 | 0.004914 | 24.3603 | 0.014398 | 0.23467 | 0.003379 | 16.2986 | 19.8661 | 3.2404 | 0.008293 | 13.6301 | 11.2581 | 13.4372 | 10.5799
2012-10-30 13:53 | 21.8218 | 33.4257 | 31.1127 | 33.2012 | 30.3832 | 46.9941 | 0.014194 | 0.34593 | 0.00491 | 24.3717 | 0.014373 | 0.23467 | 0.003373 | 16.3269 | 19.8704 | 3.2398 | 0.008283 | 13.5553 | 11.2423 | 13.3308 | 10.5128
2012-10-30 13:53 | 21.8188 | 33.3912 | 31.0742 | 33.1903 | 30.3369 | 46.9675 | 0.014184 | 0.34593 | 0.004907 | 24.3889 | 0.014392 | 0.23465 | 0.003377 | 16.3044 | 19.8474 | 3.2379 | 0.008284 | 13.5438 | 11.2268 | 13.3429 | 10.4895
2012-10-30 13:53 | 21.8278 | 33.3726 | 31.028 | 33.0969 | 30.3104 | 46.9918 | 0.0142 | 0.34601 | 0.004913 | 24.3665 | 0.014389 | 0.23464 | 0.003376 | 16.3068 | 19.8524 | 3.2396 | 0.008289 | 13.5202 | 11.1756 | 13.2445 | 10.458
2012-10-30 13:53 | 21.8345 | 33.2966 | 31.015 | 33.0918 | 30.2855 | 46.9589 | 0.014181 | 0.34592 | 0.004905 | 24.3931 | 0.014382 | 0.23462 | 0.003374 | 16.3137 | 19.8671 | 3.2373 | 0.008279 | 13.4295 | 11.1479 | 13.2247 | 10.4184
2012-10-30 13:53 | 21.8367 | 33.2476 | 30.9817 | 33.0467 | 30.2759 | 46.9843 | 0.014197 | 0.34585 | 0.00491 | 24.3605 | 0.014408 | 0.23462 | 0.00338 | 16.2843 | 19.8813 | 3.2391 | 0.00829 | 13.3663 | 11.1004 | 13.1654 | 10.3946
2012-10-30 13:53 | 21.8412 | 33.1812 | 30.9626 | 32.9685 | 30.1423 | 46.9669 | 0.014187 | 0.34593 | 0.004908 | 24.3837 | 0.014389 | 0.23464 | 0.003376 | 16.3068 | 19.8419 | 3.2379 | 0.008284 | 13.3393 | 11.1207 | 13.1266 | 10.3004
2012-10-30 13:53 | 21.8487 | 33.1808 | 30.903 | 32.9248 | 30.1301 | 46.9617 | 0.014197 | 0.34585 | 0.00491 | 24.3605 | 0.014379 | 0.23468 | 0.003374 | 16.3212 | 19.8494 | 3.2375 | 0.008284 | 13.3314 | 11.0536 | 13.0754 | 10.2807
2012-10-30 13:53 | 21.8524 | 33.1491 | 30.8437 | 32.8812 | 30.1102 | 46.9652 | 0.014197 | 0.3459 | 0.004911 | 24.3646 | 0.014408 | 0.23461 | 0.00338 | 16.2832 | 19.8531 | 3.2378 | 0.008291 | 13.296 | 10.9906 | 13.0281 | 10.2571
2012-10-30 13:53 | 21.8569 | 33.0669 | 30.8087 | 32.8542 | 30.0792 | 46.9508 | 0.014194 | 0.34585 | 0.004909 | 24.3656 | 0.014389 | 0.23461 | 0.003376 | 16.3047 | 19.8337 | 3.2368 | 0.008285 | 13.2332 | 10.975 | 13.0205 | 10.2455
2012-10-30 13:53 | 21.8599 | 33.0305 | 30.8157 | 32.8375 | 30.0151 | 46.9438 | 0.014187 | 0.3459 | 0.004907 | 24.3818 | 0.014392 | 0.23464 | 0.003377 | 16.3034 | 19.8167 | 3.2363 | 0.008284 | 13.2138 | 10.999 | 13.0208 | 10.1984
2012-10-30 13:53 | 21.8591 | 32.9904 | 30.7833 | 32.7422 | 29.9828 | 46.9473 | 0.014187 | 0.34585 | 0.004907 | 24.3777 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 19.8239 | 3.2365 | 0.008285 | 13.1665 | 10.9594 | 12.9183 | 10.1589
2012-10-30 13:53 | 21.8689 | 32.9135 | 30.6788 | 32.7126 | 29.9768 | 46.9357 | 0.014194 | 0.3459 | 0.00491 | 24.3697 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 19.8377 | 3.2357 | 0.008287 | 13.0758 | 10.8411 | 12.8749 | 10.1391
2012-10-30 13:53 | 21.8696 | 32.9418 | 30.6676 | 32.6503 | 29.9301 | 46.9155 | 0.014181 | 0.34593 | 0.004906 | 24.394 | 0.014392 | 0.23464 | 0.003377 | 16.3034 | 19.8424 | 3.2344 | 0.008283 | 13.0994 | 10.8252 | 12.8079 | 10.0877
2012-10-30 13:54 | 21.8816 | 32.875 | 30.6599 | 32.6268 | 29.9343 | 46.9178 | 0.014197 | 0.34589 | 0.004911 | 24.3635 | 0.014408 | 0.23464 | 0.003381 | 16.2853 | 19.8703 | 3.2345 | 0.008292 | 13.0047 | 10.7896 | 12.7565 | 10.064
2012-10-30 13:54 | 21.8816 | 32.7962 | 30.6442 | 32.5835 | 29.8633 | 46.9363 | 0.014184 | 0.34586 | 0.004906 | 24.3838 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.8504 | 3.2358 | 0.008282 | 12.9458 | 10.7938 | 12.7331 | 10.0129
2012-10-30 13:54 | 21.8838 | 32.7708 | 30.5557 | 32.5305 | 29.8536 | 46.9057 | 0.014187 | 0.34587 | 0.004907 | 24.3797 | 0.014389 | 0.23467 | 0.003377 | 16.3088 | 19.8566 | 3.2337 | 0.008284 | 12.9142 | 10.6991 | 12.6739 | 9.997
2012-10-30 13:54 | 21.886 | 32.7021 | 30.5184 | 32.4657 | 29.8479 | 46.9091 | 0.014197 | 0.34585 | 0.00491 | 24.3605 | 0.014382 | 0.2347 | 0.003375 | 16.3187 | 19.8508 | 3.2339 | 0.008285 | 12.8513 | 10.6676 | 12.6149 | 9.9971
2012-10-30 13:54 | 21.9047 | 32.6854 | 30.4859 | 32.4884 | 29.7956 | 46.9126 | 0.014187 | 0.34589 | 0.004907 | 24.3807 | 0.014398 | 0.23467 | 0.003379 | 16.2986 | 19.8576 | 3.2342 | 0.008286 | 12.8278 | 10.6283 | 12.6308 | 9.938
2012-10-30 13:54 | 21.8958 | 32.6331 | 30.477 | 32.3967 | 29.7473 | 46.8837 | 0.014197 | 0.34589 | 0.004911 | 24.3635 | 0.014404 | 0.23464 | 0.00338 | 16.2898 | 19.8806 | 3.2322 | 0.008291 | 12.7525 | 10.5964 | 12.5161 | 9.8667
2012-10-30 13:54 | 21.9047 | 32.6144 | 30.4425 | 32.3741 | 29.6891 | 46.8959 | 0.014187 | 0.34592 | 0.004908 | 24.3827 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 19.8576 | 3.233 | 0.008284 | 12.7568 | 10.5849 | 12.5165 | 9.8315
2012-10-30 13:54 | 21.9115 | 32.5543 | 30.3822 | 32.3257 | 29.6486 | 46.8987 | 0.014184 | 0.3458 | 0.004905 | 24.3797 | 0.014386 | 0.23457 | 0.003374 | 16.3052 | 19.8484 | 3.2332 | 0.008279 | 12.7059 | 10.5338 | 12.4773 | 9.8002
2012-10-30 13:54 | 21.9107 | 32.5298 | 30.3775 | 32.2698 | 29.6596 | 46.886 | 0.01419 | 0.34587 | 0.004908 | 24.3746 | 0.014392 | 0.23462 | 0.003377 | 16.3024 | 19.8636 | 3.2323 | 0.008285 | 12.6662 | 10.5139 | 12.4062 | 9.796
2012-10-30 13:54 | 21.9152 | 32.5107 | 30.3662 | 32.2388 | 29.5734 | 46.8791 | 0.014197 | 0.34589 | 0.004911 | 24.3635 | 0.014398 | 0.23464 | 0.003378 | 16.2966 | 19.8401 | 3.2318 | 0.008289 | 12.6706 | 10.5261 | 12.3987 | 9.7333
2012-10-30 13:54 | 21.9286 | 32.4532 | 30.3323 | 32.1734 | 29.571 | 46.8768 | 0.014187 | 0.34579 | 0.004906 | 24.3736 | 0.014379 | 0.23459 | 0.003373 | 16.3151 | 19.8815 | 3.2317 | 0.008279 | 12.5717 | 10.4508 | 12.2919 | 9.6895
2012-10-30 13:54 | 21.9204 | 32.3583 | 30.2965 | 32.1494 | 29.5825 | 46.8623 | 0.014184 | 0.34587 | 0.004906 | 24.3849 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.8733 | 3.2307 | 0.008284 | 12.485 | 10.4232 | 12.2761 | 9.7092
2012-10-30 13:54 | 21.9271 | 32.3532 | 30.2401 | 32.1364 | 29.5419 | 46.8438 | 0.014194 | 0.34582 | 0.004909 | 24.3636 | 0.014379 | 0.23467 | 0.003374 | 16.3201 | 19.852 | 3.2294 | 0.008283 | 12.5012 | 10.3881 | 12.2844 | 9.6899
2012-10-30 13:54 | 21.9316 | 32.3104 | 30.1854 | 32.0779 | 29.5109 | 46.8473 | 0.014184 | 0.34579 | 0.004905 | 24.3787 | 0.014382 | 0.23464 | 0.003375 | 16.3147 | 19.8685 | 3.2297 | 0.00828 | 12.4419 | 10.3169 | 12.2094 | 9.6424
2012-10-30 13:54 | 21.9346 | 32.2346 | 30.2042 | 32.0375 | 29.4508 | 46.8352 | 0.014206 | 0.34595 | 0.004915 | 24.3522 | 0.014386 | 0.23465 | 0.003376 | 16.3112 | 19.8914 | 3.2288 | 0.008291 | 12.3432 | 10.3128 | 12.1461 | 9.5594
2012-10-30 13:54 | 21.9369 | 32.2605 | 30.1394 | 32.0043 | 29.4136 | 46.845 | 0.014194 | 0.3459 | 0.00491 | 24.3697 | 0.014373 | 0.23465 | 0.003373 | 16.326 | 19.8379 | 3.2295 | 0.008283 | 12.4226 | 10.3015 | 12.1664 | 9.5757
2012-10-30 13:54 | 21.9443 | 32.2324 | 30.1429 | 31.9605 | 29.3737 | 46.83 | 0.014197 | 0.3458 | 0.004909 | 24.3574 | 0.014386 | 0.23471 | 0.003377 | 16.3152 | 19.8852 | 3.2285 | 0.008286 | 12.3472 | 10.2577 | 12.0753 | 9.4885
2012-10-30 13:54 | 21.9436 | 32.1923 | 30.0909 | 31.9204 | 29.3572 | 46.8184 | 0.01419 | 0.34585 | 0.004908 | 24.3725 | 0.014392 | 0.23462 | 0.003377 | 16.3024 | 19.8805 | 3.2277 | 0.008285 | 12.3118 | 10.2104 | 12.0399 | 9.4767
2012-10-30 13:54 | 21.9466 | 32.1244 | 30.0348 | 31.884 | 29.3405 | 46.8381 | 0.014181 | 0.34586 | 0.004905 | 24.389 | 0.014386 | 0.23467 | 0.003376 | 16.3122 | 19.8436 | 3.229 | 0.008281 | 12.2808 | 10.1912 | 12.0404 | 9.4969
2012-10-30 13:54 | 21.9496 | 32.0762 | 30.0259 | 31.887 | 29.3119 | 46.8253 | 0.014197 | 0.34586 | 0.00491 | 24.3615 | 0.014379 | 0.23465 | 0.003374 | 16.3191 | 19.8107 | 3.2281 | 0.008284 | 12.2655 | 10.2152 | 12.0763 | 9.5012
2012-10-30 13:54 | 21.9458 | 32.0211 | 30.0142 | 31.8359 | 29.3041 | 46.808 | 0.014206 | 0.3458 | 0.004912 | 24.342 | 0.014398 | 0.23462 | 0.003378 | 16.2956 | 19.8428 | 3.2269 | 0.00829 | 12.1783 | 10.1714 | 11.9931 | 9.4613
2012-10-30 13:54 | 21.9541 | 31.9821 | 29.9436 | 31.8205 | 29.2414 | 46.8294 | 0.01419 | 0.34589 | 0.004908 | 24.3755 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 19.8511 | 3.2284 | 0.008287 | 12.131 | 10.0925 | 11.9694 | 9.3903
2012-10-30 13:54 | 21.9563 | 31.9646 | 29.9498 | 31.74 | 29.2318 | 46.8115 | 0.014194 | 0.34579 | 0.004908 | 24.3616 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.8533 | 3.2272 | 0.008284 | 12.1113 | 10.0965 | 11.8867 | 9.3785
2012-10-30 13:54 | 21.969 | 31.9419 | 29.8875 | 31.7408 | 29.2011 | 46.8132 | 0.014172 | 0.34586 | 0.004902 | 24.4045 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 19.878 | 3.2273 | 0.008279 | 12.0639 | 10.0095 | 11.8628 | 9.3231
2012-10-30 13:54 | 21.963 | 31.8649 | 29.8894 | 31.6797 | 29.1675 | 46.804 | 0.014203 | 0.34579 | 0.004911 | 24.3461 | 0.014389 | 0.23459 | 0.003376 | 16.3038 | 19.8201 | 3.2267 | 0.008287 | 12.0448 | 10.0693 | 11.8596 | 9.3474
2012-10-30 13:54 | 21.9668 | 31.8214 | 29.8538 | 31.6204 | 29.1239 | 46.8132 | 0.014197 | 0.34585 | 0.00491 | 24.3605 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.8319 | 3.2273 | 0.008288 | 11.9895 | 10.0219 | 11.7885 | 9.292
2012-10-30 13:54 | 21.9705 | 31.7896 | 29.8023 | 31.6477 | 29.1118 | 46.7976 | 0.014181 | 0.34587 | 0.004905 | 24.39 | 0.014398 | 0.23467 | 0.003379 | 16.2986 | 19.8316 | 3.2262 | 0.008284 | 11.958 | 9.9707 | 11.8161 | 9.2802
2012-10-30 13:54 | 21.972 | 31.7833 | 29.7367 | 31.5192 | 29.0384 | 46.7832 | 0.01419 | 0.3458 | 0.004907 | 24.3694 | 0.014382 | 0.23467 | 0.003375 | 16.3167 | 19.8451 | 3.2252 | 0.008282 | 11.9382 | 9.8916 | 11.6741 | 9.1933
2012-10-30 13:55 | 21.9735 | 31.7611 | 29.7106 | 31.5285 | 28.9965 | 46.7843 | 0.0142 | 0.34587 | 0.004911 | 24.3574 | 0.014386 | 0.23468 | 0.003376 | 16.3132 | 19.8466 | 3.2253 | 0.008287 | 11.9145 | 9.864 | 11.6819 | 9.1499
2012-10-30 13:55 | 21.9824 | 31.6872 | 29.7116 | 31.5059 | 28.9738 | 46.7878 | 0.014203 | 0.34582 | 0.004912 | 24.3482 | 0.014382 | 0.23467 | 0.003375 | 16.3167 | 19.8515 | 3.2256 | 0.008287 | 11.8357 | 9.8601 | 11.6544 | 9.1223
2012-10-30 13:55 | 21.9922 | 31.6773 | 29.6543 | 31.4526 | 29.0388 | 46.7548 | 0.014194 | 0.34583 | 0.004909 | 24.3646 | 0.014392 | 0.23465 | 0.003377 | 16.3044 | 19.8573 | 3.2233 | 0.008286 | 11.82 | 9.797 | 11.5953 | 9.1815
2012-10-30 13:55 | 21.9847 | 31.5989 | 29.6074 | 31.4372 | 28.9209 | 46.7722 | 0.014178 | 0.3459 | 0.004904 | 24.3972 | 0.014392 | 0.23467 | 0.003377 | 16.3054 | 19.8179 | 3.2245 | 0.008281 | 11.781 | 9.7895 | 11.6193 | 9.103
2012-10-30 13:55 | 21.9937 | 31.596 | 29.6282 | 31.4265 | 28.9417 | 46.7722 | 0.014187 | 0.34587 | 0.004907 | 24.3797 | 0.014382 | 0.23462 | 0.003374 | 16.3137 | 19.8748 | 3.2245 | 0.008281 | 11.7212 | 9.7534 | 11.5517 | 9.0669
2012-10-30 13:55 | 21.9892 | 31.5285 | 29.5724 | 31.3156 | 28.8977 | 46.7548 | 0.014181 | 0.3458 | 0.004904 | 24.3849 | 0.014395 | 0.23462 | 0.003377 | 16.299 | 19.8782 | 3.2233 | 0.008281 | 11.6503 | 9.6942 | 11.4374 | 9.0195
2012-10-30 13:55 | 21.9974 | 31.4775 | 29.5412 | 31.308 | 28.8191 | 46.7537 | 0.01419 | 0.34586 | 0.004908 | 24.3735 | 0.014389 | 0.23465 | 0.003376 | 16.3078 | 19.8944 | 3.2232 | 0.008284 | 11.5831 | 9.6468 | 11.4136 | 8.9247
2012-10-30 13:55 | 22.0064 | 31.4708 | 29.5028 | 31.2776 | 28.8597 | 46.7843 | 0.014203 | 0.34586 | 0.004912 | 24.3512 | 0.014386 | 0.23461 | 0.003375 | 16.3081 | 19.8595 | 3.2253 | 0.008287 | 11.6113 | 9.6433 | 11.4181 | 9.0002
2012-10-30 13:55 | 22.0101 | 31.4272 | 29.5223 | 31.2182 | 28.816 | 46.7554 | 0.014197 | 0.34579 | 0.004909 | 24.3564 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 19.8832 | 3.2233 | 0.008287 | 11.544 | 9.6391 | 11.335 | 8.9328
2012-10-30 13:55 | 22.0101 | 31.3996 | 29.4316 | 31.1867 | 28.7213 | 46.7554 | 0.01419 | 0.34579 | 0.004907 | 24.3684 | 0.014392 | 0.23468 | 0.003378 | 16.3064 | 19.8832 | 3.2233 | 0.008285 | 11.5164 | 9.5484 | 11.3035 | 8.8381
2012-10-30 13:55 | 22.0108 | 31.349 | 29.4402 | 31.1637 | 28.651 | 46.7583 | 0.014184 | 0.3459 | 0.004906 | 24.3869 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 19.8799 | 3.2235 | 0.008286 | 11.4691 | 9.5603 | 11.2838 | 8.7711
2012-10-30 13:55 | 22.0265 | 31.3292 | 29.3848 | 31.1558 | 28.7022 | 46.7265 | 0.014197 | 0.34585 | 0.00491 | 24.3605 | 0.014392 | 0.23462 | 0.003377 | 16.3024 | 19.8876 | 3.2213 | 0.008287 | 11.4416 | 9.4972 | 11.2682 | 8.8146
2012-10-30 13:55 | 22.028 | 31.311 | 29.414 | 31.1178 | 28.6445 | 46.7473 | 0.0142 | 0.3458 | 0.00491 | 24.3523 | 0.014382 | 0.23465 | 0.003375 | 16.3157 | 19.8293 | 3.2228 | 0.008285 | 11.4817 | 9.5847 | 11.2885 | 8.8152
2012-10-30 13:55 | 22.0228 | 31.2664 | 29.3259 | 31.1008 | 28.6393 | 46.7566 | 0.014203 | 0.34583 | 0.004912 | 24.3492 | 0.014386 | 0.23465 | 0.003376 | 16.3112 | 19.852 | 3.2234 | 0.008288 | 11.4144 | 9.4739 | 11.2488 | 8.7873
2012-10-30 13:55 | 22.0333 | 31.2335 | 29.2693 | 31.0482 | 28.5867 | 46.7398 | 0.014178 | 0.34582 | 0.004903 | 24.3911 | 0.014398 | 0.23465 | 0.003379 | 16.2976 | 19.9104 | 3.2222 | 0.008282 | 11.3231 | 9.3589 | 11.1378 | 8.6763
2012-10-30 13:55 | 22.0325 | 31.1894 | 29.2449 | 31.0001 | 28.5346 | 46.7427 | 0.014187 | 0.34583 | 0.004906 | 24.3766 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 19.9056 | 3.2224 | 0.008284 | 11.2838 | 9.3393 | 11.0945 | 8.629
2012-10-30 13:55 | 22.0467 | 31.1444 | 29.2354 | 30.9788 | 28.533 | 46.7213 | 0.01419 | 0.34585 | 0.004908 | 24.3725 | 0.014386 | 0.23464 | 0.003376 | 16.3102 | 19.8998 | 3.221 | 0.008284 | 11.2446 | 9.3356 | 11.079 | 8.6332
2012-10-30 13:55 | 22.0355 | 31.1608 | 29.1887 | 30.8966 | 28.5178 | 46.741 | 0.014194 | 0.34579 | 0.004908 | 24.3616 | 0.014401 | 0.23464 | 0.003379 | 16.2932 | 19.8567 | 3.2223 | 0.008287 | 11.3041 | 9.332 | 11.0399 | 8.6611
2012-10-30 13:55 | 22.043 | 31.0776 | 29.1646 | 30.9317 | 28.474 | 46.73 | 0.014194 | 0.34592 | 0.00491 | 24.3707 | 0.014376 | 0.23459 | 0.003373 | 16.3185 | 19.8563 | 3.2216 | 0.008283 | 11.2213 | 9.3083 | 11.0754 | 8.6177
2012-10-30 13:55 | 22.049 | 31.1073 | 29.1627 | 30.9377 | 28.5037 | 46.7005 | 0.014203 | 0.34589 | 0.004913 | 24.3532 | 0.014392 | 0.23467 | 0.003377 | 16.3054 | 19.8543 | 3.2195 | 0.00829 | 11.253 | 9.3084 | 11.0834 | 8.6494
2012-10-30 13:55 | 22.0475 | 31.0151 | 29.1375 | 30.881 | 28.443 | 46.719 | 0.014203 | 0.34586 | 0.004912 | 24.3512 | 0.014382 | 0.23464 | 0.003375 | 16.3147 | 19.9006 | 3.2208 | 0.008287 | 11.1145 | 9.2369 | 10.9804 | 8.5424
2012-10-30 13:55 | 22.0572 | 30.9972 | 29.0881 | 30.8355 | 28.4054 | 46.7046 | 0.014187 | 0.34583 | 0.004906 | 24.3766 | 0.014386 | 0.23459 | 0.003375 | 16.3072 | 19.8625 | 3.2198 | 0.008281 | 11.1347 | 9.2256 | 10.973 | 8.5429
2012-10-30 13:55 | 22.052 | 30.9644 | 29.0671 | 30.8185 | 28.3725 | 46.7219 | 0.01419 | 0.34582 | 0.004907 | 24.3705 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 19.8653 | 3.221 | 0.008283 | 11.0991 | 9.2018 | 10.9532 | 8.5072
2012-10-30 13:55 | 22.0714 | 30.9759 | 29.0786 | 30.7945 | 28.4314 | 46.7023 | 0.0142 | 0.34585 | 0.004911 | 24.3554 | 0.014386 | 0.23459 | 0.003375 | 16.3072 | 19.8088 | 3.2197 | 0.008286 | 11.1671 | 9.2698 | 10.9857 | 8.6226
2012-10-30 13:55 | 22.0609 | 30.9102 | 29.0405 | 30.7367 | 28.267 | 46.7127 | 0.014187 | 0.34586 | 0.004907 | 24.3787 | 0.014395 | 0.23464 | 0.003378 | 16.3 | 19.8422 | 3.2204 | 0.008285 | 11.068 | 9.1983 | 10.8945 | 8.4248
2012-10-30 13:55 | 22.0684 | 30.8586 | 29.044 | 30.7048 | 28.3179 | 46.7109 | 0.0142 | 0.34582 | 0.004911 | 24.3533 | 0.014382 | 0.23467 | 0.003375 | 16.3167 | 19.8418 | 3.2202 | 0.008286 | 11.0168 | 9.2022 | 10.863 | 8.4761
2012-10-30 13:55 | 22.0714 | 30.8576 | 28.9365 | 30.688 | 28.2064 | 46.6895 | 0.014184 | 0.34585 | 0.004905 | 24.3828 | 0.014386 | 0.23464 | 0.003376 | 16.3102 | 19.8886 | 3.2188 | 0.008281 | 10.969 | 9.0479 | 10.7994 | 8.3178
2012-10-30 13:55 | 22.0796 | 30.8224 | 28.9487 | 30.6923 | 28.2225 | 46.7086 | 0.01419 | 0.34577 | 0.004907 | 24.3674 | 0.014389 | 0.23462 | 0.003376 | 16.3058 | 19.8849 | 3.2201 | 0.008283 | 10.9375 | 9.0638 | 10.8074 | 8.3376
2012-10-30 13:55 | 22.0766 | 30.8194 | 28.8746 | 30.6223 | 28.2314 | 46.6947 | 0.014197 | 0.34582 | 0.00491 | 24.3585 | 0.014386 | 0.23465 | 0.003376 | 16.3112 | 19.8938 | 3.2191 | 0.008286 | 10.9256 | 8.9808 | 10.7285 | 8.3376
2012-10-30 13:55 | 22.0871 | 30.7353 | 28.8694 | 30.6012 | 28.1393 | 46.6855 | 0.014194 | 0.34585 | 0.004909 | 24.3656 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 19.9323 | 3.2185 | 0.008288 | 10.803 | 8.9371 | 10.6689 | 8.207
2012-10-30 13:55 | 22.0818 | 30.73 | 28.8641 | 30.5644 | 28.1458 | 46.6988 | 0.014194 | 0.34579 | 0.004908 | 24.3616 | 0.014401 | 0.23465 | 0.003379 | 16.2942 | 19.8512 | 3.2194 | 0.008287 | 10.8788 | 9.0129 | 10.7132 | 8.2946
2012-10-30 13:55 | 22.0848 | 30.6778 | 28.8276 | 30.5121 | 28.1251 | 46.6739 | 0.01419 | 0.34576 | 0.004906 | 24.3664 | 0.014386 | 0.23459 | 0.003375 | 16.3072 | 19.8462 | 3.2177 | 0.008281 | 10.8316 | 8.9814 | 10.6659 | 8.2789
2012-10-30 13:56 | 22.0923 | 30.6774 | 28.7956 | 30.4999 | 28.1089 | 46.689 | 0.01419 | 0.3458 | 0.004907 | 24.3694 | 0.014389 | 0.23458 | 0.003375 | 16.3027 | 19.8657 | 3.2187 | 0.008282 | 10.8117 | 8.9299 | 10.6342 | 8.2432
2012-10-30 13:56 | 22.0946 | 30.6363 | 28.794 | 30.5219 | 28.0797 | 46.6647 | 0.0142 | 0.34577 | 0.00491 | 24.3503 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.8999 | 3.2171 | 0.008287 | 10.7364 | 8.8941 | 10.622 | 8.1798
2012-10-30 13:56 | 22.1035 | 30.5821 | 28.7595 | 30.4638 | 28.1044 | 46.6757 | 0.014168 | 0.34573 | 0.004898 | 24.4022 | 0.014389 | 0.23461 | 0.003376 | 16.3047 | 19.8848 | 3.2178 | 0.008274 | 10.6973 | 8.8747 | 10.579 | 8.2196
2012-10-30 13:56 | 22.1028 | 30.542 | 28.7035 | 30.4237 | 28.0366 | 46.6907 | 0.014194 | 0.34589 | 0.00491 | 24.3687 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 19.8363 | 3.2189 | 0.008284 | 10.7057 | 8.8672 | 10.5874 | 8.2003
2012-10-30 13:56 | 22.1005 | 30.5278 | 28.6933 | 30.3661 | 28.0106 | 46.6543 | 0.01419 | 0.34572 | 0.004906 | 24.3634 | 0.014389 | 0.23461 | 0.003376 | 16.3047 | 19.8539 | 3.2163 | 0.008282 | 10.6739 | 8.8394 | 10.5122 | 8.1567
2012-10-30 13:56 | 22.111 | 30.5107 | 28.6565 | 30.3411 | 28.0171 | 46.6664 | 0.01419 | 0.3458 | 0.004907 | 24.3694 | 0.014382 | 0.23465 | 0.003375 | 16.3157 | 19.9003 | 3.2172 | 0.008282 | 10.6104 | 8.7562 | 10.4408 | 8.1168
2012-10-30 13:56 | 22.1125 | 30.4767 | 28.6698 | 30.3466 | 27.9594 | 46.6676 | 0.01419 | 0.34582 | 0.004907 | 24.3705 | 0.014386 | 0.23458 | 0.003375 | 16.3061 | 19.842 | 3.2173 | 0.008282 | 10.6347 | 8.8278 | 10.5046 | 8.1174
2012-10-30 13:56 | 22.117 | 30.4221 | 28.5915 | 30.2485 | 27.9718 | 46.6525 | 0.0142 | 0.34567 | 0.004909 | 24.3431 | 0.014389 | 0.23465 | 0.003376 | 16.3078 | 19.8943 | 3.2162 | 0.008285 | 10.5278 | 8.6972 | 10.3542 | 8.0775
2012-10-30 13:56 | 22.1147 | 30.3567 | 28.5655 | 30.2738 | 27.9143 | 46.6647 | 0.014197 | 0.34583 | 0.00491 | 24.3594 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.8122 | 3.2171 | 0.008288 | 10.5445 | 8.7533 | 10.4616 | 8.1021
2012-10-30 13:56 | 22.1222 | 30.3523 | 28.4743 | 30.2616 | 27.9139 | 46.6352 | 0.014212 | 0.34569 | 0.004913 | 24.3236 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.8836 | 3.215 | 0.008287 | 10.4687 | 8.5907 | 10.378 | 8.0303
2012-10-30 13:56 | 22.1364 | 30.3508 | 28.5319 | 30.1891 | 27.8373 | 46.6433 | 0.014194 | 0.34577 | 0.004908 | 24.3606 | 0.014389 | 0.23465 | 0.003376 | 16.3078 | 19.8978 | 3.2156 | 0.008284 | 10.453 | 8.6341 | 10.2913 | 7.9395
2012-10-30 13:56 | 22.1342 | 30.2539 | 28.4982 | 30.1632 | 27.8272 | 46.6358 | 0.0142 | 0.34586 | 0.004911 | 24.3563 | 0.014386 | 0.23462 | 0.003375 | 16.3092 | 19.8876 | 3.2151 | 0.008286 | 10.3663 | 8.6106 | 10.2756 | 7.9396
2012-10-30 13:56 | 22.1372 | 30.2293 | 28.4696 | 30.0952 | 27.8262 | 46.6485 | 0.014194 | 0.34577 | 0.004908 | 24.3606 | 0.014398 | 0.23459 | 0.003378 | 16.2936 | 19.8746 | 3.2159 | 0.008286 | 10.3547 | 8.595 | 10.2206 | 7.9516
2012-10-30 13:56 | 22.1401 | 30.2559 | 28.4291 | 30.0665 | 27.7936 | 46.6375 | 0.014194 | 0.34582 | 0.004909 | 24.3636 | 0.014376 | 0.23455 | 0.003372 | 16.3155 | 19.8855 | 3.2152 | 0.008281 | 10.3704 | 8.5436 | 10.181 | 7.9081
2012-10-30 13:56 | 22.1491 | 30.1899 | 28.4499 | 30.0597 | 27.7592 | 46.6335 | 0.014184 | 0.34582 | 0.004905 | 24.3808 | 0.014379 | 0.23455 | 0.003373 | 16.3121 | 19.8905 | 3.2149 | 0.008278 | 10.2994 | 8.5594 | 10.1692 | 7.8687
2012-10-30 13:56 | 22.1499 | 30.1552 | 28.3915 | 30.0211 | 27.76 | 46.6364 | 0.01419 | 0.34579 | 0.004907 | 24.3684 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9073 | 3.2151 | 0.008281 | 10.2479 | 8.4842 | 10.1138 | 7.8527
2012-10-30 13:56 | 22.1514 | 30.1054 | 28.3693 | 29.9674 | 27.7773 | 46.6271 | 0.014181 | 0.34575 | 0.004903 | 24.3809 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.8888 | 3.2145 | 0.008279 | 10.2166 | 8.4805 | 10.0786 | 7.8885
2012-10-30 13:56 | 22.1603 | 30.0828 | 28.319 | 29.9289 | 27.6796 | 46.6011 | 0.014187 | 0.34579 | 0.004906 | 24.3736 | 0.014386 | 0.23464 | 0.003376 | 16.3102 | 19.8778 | 3.2127 | 0.008282 | 10.205 | 8.4412 | 10.0511 | 7.8018
2012-10-30 13:56 | 22.1543 | 30.0255 | 28.3407 | 29.9111 | 27.642 | 46.6271 | 0.014181 | 0.34576 | 0.004903 | 24.3818 | 0.014382 | 0.23464 | 0.003375 | 16.3147 | 19.8997 | 3.2145 | 0.008278 | 10.1258 | 8.441 | 10.0114 | 7.7423
2012-10-30 13:56 | 22.1641 | 29.9998 | 28.2834 | 29.8657 | 27.6518 | 46.6242 | 0.014197 | 0.3457 | 0.004908 | 24.3503 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 19.8896 | 3.2143 | 0.008288 | 10.1102 | 8.3938 | 9.9761 | 7.7622
2012-10-30 13:56 | 22.1528 | 29.9806 | 28.2129 | 29.8544 | 27.6168 | 46.6138 | 0.014184 | 0.3458 | 0.004905 | 24.3797 | 0.014398 | 0.23465 | 0.003379 | 16.2976 | 19.8783 | 3.2136 | 0.008284 | 10.1023 | 8.3346 | 9.9761 | 7.7385
2012-10-30 13:56 | 22.1678 | 29.9601 | 28.216 | 29.7826 | 27.5568 | 46.6161 | 0.014181 | 0.34579 | 0.004904 | 24.3839 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 19.8653 | 3.2137 | 0.008284 | 10.0948 | 8.3507 | 9.9173 | 7.6915
2012-10-30 13:56 | 22.173 | 29.9022 | 28.1778 | 29.7562 | 27.5423 | 46.6133 | 0.014203 | 0.34573 | 0.00491 | 24.342 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 19.8825 | 3.2135 | 0.008285 | 10.0197 | 8.2953 | 9.8737 | 7.6598
2012-10-30 13:56 | 22.1753 | 29.8729 | 28.1248 | 29.7309 | 27.517 | 46.6005 | 0.014194 | 0.34575 | 0.004908 | 24.3585 | 0.014389 | 0.23458 | 0.003375 | 16.3027 | 19.9407 | 3.2126 | 0.008283 | 9.9322 | 8.1841 | 9.7902 | 7.5763
2012-10-30 13:56 | 22.1783 | 29.8325 | 28.1278 | 29.6945 | 27.5121 | 46.6104 | 0.014184 | 0.34582 | 0.004905 | 24.3808 | 0.01437 | 0.23461 | 0.003371 | 16.3263 | 19.8958 | 3.2133 | 0.008276 | 9.9367 | 8.232 | 9.7987 | 7.6163
2012-10-30 13:56 | 22.1827 | 29.8606 | 28.0928 | 29.691 | 27.4888 | 46.5913 | 0.014194 | 0.34579 | 0.004908 | 24.3616 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.8683 | 3.212 | 0.008284 | 9.9923 | 8.2245 | 9.8227 | 7.6205
2012-10-30 13:56 | 22.1813 | 29.7527 | 28.0756 | 29.6698 | 27.448 | 46.6196 | 0.014203 | 0.34577 | 0.004911 | 24.3451 | 0.014386 | 0.23459 | 0.003375 | 16.3072 | 19.8868 | 3.214 | 0.008286 | 9.8659 | 8.1888 | 9.783 | 7.5612
2012-10-30 13:56 | 22.1857 | 29.761 | 28.0366 | 29.5677 | 27.4484 | 46.6028 | 0.014184 | 0.34579 | 0.004905 | 24.3787 | 0.014382 | 0.23471 | 0.003376 | 16.3197 | 19.8633 | 3.2128 | 0.008281 | 9.8977 | 8.1733 | 9.7044 | 7.5851
2012-10-30 13:56 | 22.1835 | 29.676 | 28.0344 | 29.5931 | 27.4146 | 46.5815 | 0.014187 | 0.34575 | 0.004905 | 24.3706 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 19.881 | 3.2113 | 0.008282 | 9.795 | 8.1534 | 9.7121 | 7.5336
2012-10-30 13:57 | 22.1872 | 29.7191 | 27.9552 | 29.5455 | 27.3828 | 46.5988 | 0.014194 | 0.34582 | 0.004909 | 24.3636 | 0.014395 | 0.23461 | 0.003377 | 16.298 | 19.8728 | 3.2125 | 0.008286 | 9.8463 | 8.0824 | 9.6727 | 7.51
2012-10-30 13:57 | 22.2014 | 29.6702 | 27.9773 | 29.5361 | 27.3694 | 46.574 | 0.014187 | 0.34582 | 0.004906 | 24.3756 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 19.9069 | 3.2108 | 0.008284 | 9.7633 | 8.0704 | 9.6292 | 7.4625
2012-10-30 13:57 | 22.1969 | 29.6263 | 27.9294 | 29.4645 | 27.3373 | 46.5728 | 0.014203 | 0.34576 | 0.004911 | 24.3441 | 0.014398 | 0.23459 | 0.003378 | 16.2936 | 19.8545 | 3.2107 | 0.008289 | 9.7718 | 8.0749 | 9.61 | 7.4828
2012-10-30 13:57 | 22.2037 | 29.6015 | 27.9401 | 29.4476 | 27.3085 | 46.578 | 0.014216 | 0.34577 | 0.004916 | 24.3229 | 0.014389 | 0.23464 | 0.003376 | 16.3068 | 19.9052 | 3.2111 | 0.008292 | 9.6963 | 8.0349 | 9.5424 | 7.4033
2012-10-30 13:57 | 22.1969 | 29.5592 | 27.8583 | 29.425 | 27.2741 | 46.563 | 0.014187 | 0.34569 | 0.004904 | 24.3665 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.8665 | 3.2101 | 0.00828 | 9.6927 | 7.9918 | 9.5585 | 7.4076
2012-10-30 13:57 | 22.2014 | 29.5361 | 27.851 | 29.4177 | 27.2273 | 46.5728 | 0.014187 | 0.3458 | 0.004906 | 24.3746 | 0.014398 | 0.23462 | 0.003378 | 16.2956 | 19.9109 | 3.2107 | 0.008284 | 9.6252 | 7.9401 | 9.5068 | 7.3164
2012-10-30 13:57 | 22.2164 | 29.4919 | 27.8581 | 29.4011 | 27.2423 | 46.5734 | 0.014197 | 0.34577 | 0.004909 | 24.3554 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.9538 | 3.2108 | 0.008285 | 9.5381 | 7.9043 | 9.4473 | 7.2885
2012-10-30 13:57 | 22.2164 | 29.5077 | 27.866 | 29.338 | 27.2067 | 46.5647 | 0.014194 | 0.34576 | 0.004908 | 24.3595 | 0.014401 | 0.23461 | 0.003379 | 16.2912 | 19.9499 | 3.2102 | 0.008287 | 9.5578 | 7.9161 | 9.3881 | 7.2568
2012-10-30 13:57 | 22.2104 | 29.4425 | 27.7889 | 29.332 | 27.1731 | 46.5728 | 0.014184 | 0.3457 | 0.004903 | 24.3726 | 0.014411 | 0.23462 | 0.003381 | 16.2809 | 19.9119 | 3.2107 | 0.008284 | 9.5306 | 7.877 | 9.4201 | 7.2612
2012-10-30 13:57 | 22.2186 | 29.3954 | 27.7458 | 29.2731 | 27.1853 | 46.5514 | 0.014194 | 0.34577 | 0.004908 | 24.3606 | 0.014395 | 0.23461 | 0.003377 | 16.298 | 19.9321 | 3.2093 | 0.008285 | 9.4633 | 7.8137 | 9.341 | 7.2532
2012-10-30 13:57 | 22.2201 | 29.3733 | 27.7789 | 29.2746 | 27.167 | 46.5485 | 0.014197 | 0.34572 | 0.004908 | 24.3513 | 0.014395 | 0.23465 | 0.003378 | 16.301 | 19.8937 | 3.2091 | 0.008286 | 9.4796 | 7.8852 | 9.3809 | 7.2733
2012-10-30 13:57 | 22.2253 | 29.3587 | 27.7249 | 29.2088 | 27.1367 | 46.5786 | 0.014197 | 0.34576 | 0.004909 | 24.3544 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 19.9189 | 3.2111 | 0.008289 | 9.4398 | 7.806 | 9.2899 | 7.2178
2012-10-30 13:57 | 22.2246 | 29.2949 | 27.7084 | 29.1805 | 27.1084 | 46.5243 | 0.0142 | 0.34577 | 0.00491 | 24.3503 | 0.014392 | 0.23471 | 0.003378 | 16.3084 | 19.9142 | 3.2074 | 0.008288 | 9.3807 | 7.7942 | 9.2663 | 7.1942
2012-10-30 13:57 | 22.2358 | 29.2903 | 27.6683 | 29.1641 | 27.1314 | 46.537 | 0.01419 | 0.34579 | 0.004907 | 24.3684 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.9453 | 3.2083 | 0.008285 | 9.345 | 7.723 | 9.2188 | 7.1861
2012-10-30 13:57 | 22.2381 | 29.2532 | 27.6745 | 29.1387 | 27.0982 | 46.5589 | 0.0142 | 0.34576 | 0.00491 | 24.3492 | 0.014389 | 0.23462 | 0.003376 | 16.3058 | 19.8478 | 3.2098 | 0.008286 | 9.4054 | 7.8267 | 9.2909 | 7.2504
2012-10-30 13:57 | 22.2493 | 29.2644 | 27.7015 | 29.0947 | 27.058 | 46.5277 | 0.014209 | 0.34576 | 0.004913 | 24.3338 | 0.014398 | 0.23464 | 0.003378 | 16.2966 | 19.9349 | 3.2076 | 0.008291 | 9.3295 | 7.7666 | 9.1598 | 7.1231
2012-10-30 13:57 | 22.25 | 29.2019 | 27.6588 | 29.0678 | 27.043 | 46.5387 | 0.014203 | 0.34577 | 0.004911 | 24.3451 | 0.014398 | 0.23462 | 0.003378 | 16.2956 | 19.9276 | 3.2084 | 0.008289 | 9.2743 | 7.7312 | 9.1402 | 7.1154
2012-10-30 13:57 | 22.25 | 29.194 | 27.5759 | 29.0362 | 27.0074 | 46.5243 | 0.014203 | 0.34577 | 0.004911 | 24.3451 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 19.9515 | 3.2074 | 0.00829 | 9.2425 | 7.6244 | 9.0847 | 7.0559
2012-10-30 13:57 | 22.2545 | 29.1196 | 27.5804 | 29.0052 | 26.9961 | 46.5375 | 0.014181 | 0.34576 | 0.004903 | 24.3818 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9361 | 3.2083 | 0.008281 | 9.1835 | 7.6443 | 9.0691 | 7.06
2012-10-30 13:57 | 22.256 | 29.1448 | 27.5464 | 28.9791 | 26.97 | 46.5277 | 0.014197 | 0.34577 | 0.004909 | 24.3554 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9256 | 3.2076 | 0.008286 | 9.2192 | 7.6208 | 9.0535 | 7.0444
2012-10-30 13:57 | 22.265 | 29.067 | 27.5317 | 28.921 | 26.9316 | 46.5208 | 0.014194 | 0.34573 | 0.004907 | 24.3575 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.9266 | 3.2071 | 0.008285 | 9.1404 | 7.6051 | 8.9944 | 7.005
2012-10-30 13:57 | 22.2754 | 29.0419 | 27.4868 | 28.9195 | 26.8907 | 46.5144 | 0.014203 | 0.34573 | 0.00491 | 24.342 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 19.9171 | 3.2067 | 0.008288 | 9.1248 | 7.5697 | 9.0024 | 6.9736
2012-10-30 13:57 | 22.2724 | 29.0152 | 27.4483 | 28.8771 | 26.8916 | 46.5162 | 0.014181 | 0.34575 | 0.004903 | 24.3809 | 0.014389 | 0.23462 | 0.003376 | 16.3058 | 19.958 | 3.2068 | 0.008279 | 9.0572 | 7.4903 | 8.9191 | 6.9336
2012-10-30 13:57 | 22.2694 | 28.9885 | 27.3821 | 28.807 | 26.8413 | 46.5104 | 0.014187 | 0.34576 | 0.004905 | 24.3715 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 19.9151 | 3.2064 | 0.008282 | 9.0734 | 7.467 | 8.8919 | 6.9262
2012-10-30 13:57 | 22.2777 | 28.9574 | 27.422 | 28.8074 | 26.8338 | 46.5202 | 0.014178 | 0.34582 | 0.004903 | 24.3911 | 0.014401 | 0.23464 | 0.003379 | 16.2932 | 19.9034 | 3.2071 | 0.008282 | 9.054 | 7.5186 | 8.904 | 6.9304
2012-10-30 13:57 | 22.2829 | 28.9626 | 27.4233 | 28.8284 | 26.8429 | 46.5179 | 0.01419 | 0.34575 | 0.004906 | 24.3654 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9046 | 3.2069 | 0.008284 | 9.058 | 7.5187 | 8.9238 | 6.9383
2012-10-30 13:57 | 22.2859 | 28.9103 | 27.3828 | 28.7367 | 26.8064 | 46.5086 | 0.014187 | 0.34575 | 0.004905 | 24.3706 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 19.9276 | 3.2063 | 0.008285 | 8.9827 | 7.4552 | 8.8091 | 6.8788
2012-10-30 13:57 | 22.2941 | 28.8514 | 27.3792 | 28.6975 | 26.7673 | 46.5179 | 0.014194 | 0.3457 | 0.004907 | 24.3554 | 0.014398 | 0.23462 | 0.003378 | 16.2956 | 19.9318 | 3.2069 | 0.008285 | 8.9196 | 7.4474 | 8.7657 | 6.8355
2012-10-30 13:57 | 22.2971 | 28.8702 | 27.3348 | 28.6847 | 26.7466 | 46.5023 | 0.014184 | 0.34576 | 0.004904 | 24.3767 | 0.014376 | 0.23461 | 0.003373 | 16.3195 | 19.9428 | 3.2059 | 0.008277 | 8.9274 | 7.392 | 8.7419 | 6.8038
2012-10-30 13:57 | 22.2896 | 28.8311 | 27.2642 | 28.6457 | 26.6957 | 46.5 | 0.014184 | 0.34567 | 0.004903 | 24.3706 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9033 | 3.2057 | 0.008279 | 8.9278 | 7.3609 | 8.7424 | 6.7924
2012-10-30 13:58 | 22.3061 | 28.7371 | 27.257 | 28.6464 | 26.7161 | 46.5086 | 0.014181 | 0.34577 | 0.004903 | 24.3829 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9238 | 3.2063 | 0.008279 | 8.8133 | 7.3332 | 8.7226 | 6.7923
2012-10-30 13:58 | 22.2978 | 28.7565 | 27.2289 | 28.5986 | 26.6802 | 46.4769 | 0.014203 | 0.34577 | 0.004911 | 24.3451 | 0.014386 | 0.23462 | 0.003375 | 16.3092 | 19.8796 | 3.2041 | 0.008286 | 8.8769 | 7.3493 | 8.719 | 6.8006
2012-10-30 13:58 | 22.3016 | 28.7011 | 27.1459 | 28.5629 | 26.6524 | 46.4884 | 0.01419 | 0.34573 | 0.004906 | 24.3643 | 0.014376 | 0.23458 | 0.003372 | 16.3175 | 19.9193 | 3.2049 | 0.008278 | 8.7818 | 7.2266 | 8.6436 | 6.7331
2012-10-30 13:58 | 22.3023 | 28.7215 | 27.1782 | 28.5321 | 26.6057 | 46.4982 | 0.014197 | 0.3458 | 0.004909 | 24.3574 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.9041 | 3.2056 | 0.008287 | 8.8174 | 7.2741 | 8.628 | 6.7016
2012-10-30 13:58 | 22.3061 | 28.6898 | 27.1385 | 28.5043 | 26.5977 | 46.4757 | 0.014194 | 0.34572 | 0.004907 | 24.3565 | 0.014382 | 0.23457 | 0.003374 | 16.3097 | 19.8879 | 3.204 | 0.008281 | 8.8019 | 7.2506 | 8.6164 | 6.7098
2012-10-30 13:58 | 22.3106 | 28.643 | 27.1312 | 28.493 | 26.5824 | 46.4948 | 0.014197 | 0.34577 | 0.004909 | 24.3554 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 19.9203 | 3.2054 | 0.008286 | 8.7227 | 7.2109 | 8.5727 | 6.6621
2012-10-30 13:58 | 22.3143 | 28.5993 | 27.1033 | 28.4888 | 26.523 | 46.474 | 0.014187 | 0.34576 | 0.004905 | 24.3715 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 19.9001 | 3.2039 | 0.008281 | 8.6992 | 7.2032 | 8.5887 | 6.6229
2012-10-30 13:58 | 22.3188 | 28.5801 | 27.092 | 28.4341 | 26.563 | 46.4763 | 0.01419 | 0.34575 | 0.004906 | 24.3654 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 19.8726 | 3.2041 | 0.008282 | 8.7075 | 7.2194 | 8.5615 | 6.6904
2012-10-30 13:58 | 22.3255 | 28.5513 | 27.1027 | 28.4132 | 26.5302 | 46.4803 | 0.014197 | 0.34573 | 0.004908 | 24.3523 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.8594 | 3.2044 | 0.008284 | 8.6919 | 7.2433 | 8.5538 | 6.6708
2012-10-30 13:58 | 22.3292 | 28.5629 | 27.0432 | 28.3774 | 26.5102 | 46.4624 | 0.014194 | 0.3457 | 0.004907 | 24.3554 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.9749 | 3.2031 | 0.008283 | 8.588 | 7.0683 | 8.4025 | 6.5353
2012-10-30 13:58 | 22.333 | 28.4878 | 27.0588 | 28.3378 | 26.4903 | 46.474 | 0.0142 | 0.34573 | 0.004909 | 24.3472 | 0.014382 | 0.23464 | 0.003375 | 16.3147 | 19.9587 | 3.2039 | 0.008284 | 8.5291 | 7.1001 | 8.3791 | 6.5316
2012-10-30 13:58 | 22.3322 | 28.4554 | 27.0028 | 28.3567 | 26.5211 | 46.4659 | 0.01419 | 0.34575 | 0.004906 | 24.3654 | 0.014389 | 0.23459 | 0.003376 | 16.3038 | 19.894 | 3.2034 | 0.008282 | 8.5614 | 7.1088 | 8.4627 | 6.6271
2012-10-30 13:58 | 22.3524 | 28.3335 | 26.9164 | 28.1954 | 26.3913 | 46.4561 | 0.01419 | 0.34575 | 0.004906 | 24.3654 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9062 | 3.2027 | 0.008283 | 8.4273 | 7.0102 | 8.2892 | 6.4851
2012-10-30 13:58 | 22.3666 | 28.2451 | 26.7608 | 28.0675 | 26.2989 | 46.437 | 0.014175 | 0.34576 | 0.004901 | 24.3922 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.9005 | 3.2014 | 0.008279 | 8.3446 | 6.8603 | 8.167 | 6.3984
2012-10-30 13:58 | 22.3674 | 28.1946 | 26.7814 | 28.0328 | 26.2681 | 46.4341 | 0.014187 | 0.34572 | 0.004905 | 24.3685 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 19.8733 | 3.2012 | 0.008283 | 8.3213 | 6.9081 | 8.1595 | 6.3948
2012-10-30 13:58 | 22.3801 | 28.1797 | 26.7585 | 27.9941 | 26.2295 | 46.4549 | 0.0142 | 0.34572 | 0.004909 | 24.3462 | 0.014389 | 0.23458 | 0.003375 | 16.3027 | 19.9339 | 3.2026 | 0.008284 | 8.2458 | 6.8246 | 8.0602 | 6.2956
2012-10-30 13:58 | 22.3816 | 28.1496 | 26.7363 | 27.9601 | 26.1994 | 46.4295 | 0.014194 | 0.34576 | 0.004908 | 24.3595 | 0.014376 | 0.23457 | 0.003372 | 16.3165 | 19.9354 | 3.2008 | 0.00828 | 8.2142 | 6.8009 | 8.0247 | 6.264
2012-10-30 13:58 | 22.3875 | 28.1042 | 26.7304 | 27.895 | 26.1619 | 46.4335 | 0.014206 | 0.3456 | 0.00491 | 24.3277 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 19.9334 | 3.2011 | 0.008287 | 8.1708 | 6.797 | 7.9616 | 6.2285
2012-10-30 13:58 | 22.3838 | 28.0728 | 26.7109 | 27.9031 | 26.1226 | 46.4283 | 0.01419 | 0.34573 | 0.004906 | 24.3643 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9576 | 3.2008 | 0.008284 | 8.1152 | 6.7533 | 7.9455 | 6.165
2012-10-30 13:58 | 22.3845 | 28.0459 | 26.684 | 27.8643 | 26.1825 | 46.4144 | 0.014203 | 0.3457 | 0.00491 | 24.34 | 0.014382 | 0.23461 | 0.003374 | 16.3127 | 19.9144 | 3.1998 | 0.008284 | 8.1315 | 6.7696 | 7.9499 | 6.2681
2012-10-30 13:58 | 22.3868 | 28.0601 | 26.5955 | 27.8351 | 26.1414 | 46.422 | 0.014178 | 0.34569 | 0.004901 | 24.3819 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 19.9606 | 3.2003 | 0.008278 | 8.0995 | 6.6349 | 7.8745 | 6.1808
2012-10-30 13:59 | 22.3875 | 27.9858 | 26.6356 | 27.8279 | 26.1421 | 46.4347 | 0.0142 | 0.34562 | 0.004908 | 24.3391 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.9453 | 3.2012 | 0.008284 | 8.0405 | 6.6903 | 7.8826 | 6.1968
2012-10-30 13:59 | 22.3973 | 27.964 | 26.5981 | 27.7627 | 26.1361 | 46.4075 | 0.014203 | 0.34566 | 0.004909 | 24.337 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.9671 | 3.1993 | 0.008285 | 7.9969 | 6.631 | 7.7956 | 6.169
2012-10-30 13:59 | 22.401 | 27.944 | 26.586 | 27.794 | 26.0806 | 46.422 | 0.01419 | 0.3457 | 0.004905 | 24.3623 | 0.014404 | 0.23461 | 0.003379 | 16.2878 | 19.9149 | 3.2003 | 0.008284 | 8.0291 | 6.6711 | 7.8791 | 6.1657
2012-10-30 13:59 | 22.4077 | 27.931 | 26.565 | 27.7415 | 26.111 | 46.422 | 0.01419 | 0.34573 | 0.004906 | 24.3643 | 0.014395 | 0.23467 | 0.003378 | 16.302 | 19.9256 | 3.2003 | 0.008284 | 8.0054 | 6.6394 | 7.8159 | 6.1854
2012-10-30 13:59 | 22.3988 | 27.8313 | 26.5482 | 27.6852 | 26.0547 | 46.411 | 0.014206 | 0.34575 | 0.004912 | 24.338 | 0.014379 | 0.23454 | 0.003372 | 16.3111 | 19.9407 | 3.1996 | 0.008284 | 7.8906 | 6.6075 | 7.7445 | 6.114
2012-10-30 13:59 | 22.4017 | 27.8776 | 26.5393 | 27.7118 | 26.0576 | 46.4098 | 0.014194 | 0.34564 | 0.004906 | 24.3513 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.9156 | 3.1995 | 0.008283 | 7.962 | 6.6237 | 7.7962 | 6.142
2012-10-30 13:59 | 22.4144 | 27.8706 | 26.5717 | 27.6258 | 26.0348 | 46.3798 | 0.014187 | 0.34567 | 0.004904 | 24.3654 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9243 | 3.1974 | 0.00828 | 7.9463 | 6.6474 | 7.7015 | 6.1105
2012-10-30 13:59 | 22.4122 | 27.8131 | 26.534 | 27.6315 | 25.9891 | 46.4023 | 0.014181 | 0.34572 | 0.004903 | 24.3788 | 0.014386 | 0.23457 | 0.003374 | 16.3052 | 19.9141 | 3.199 | 0.008277 | 7.899 | 6.6199 | 7.7174 | 6.075
2012-10-30 13:59 | 22.407 | 27.7921 | 26.4419 | 27.5829 | 25.9523 | 46.396 | 0.014197 | 0.34567 | 0.004908 | 24.3482 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9489 | 3.1985 | 0.008284 | 7.8432 | 6.493 | 7.634 | 6.0034
2012-10-30 13:59 | 22.4182 | 27.7836 | 26.465 | 27.598 | 25.9399 | 46.4121 | 0.0142 | 0.3457 | 0.004909 | 24.3451 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9082 | 3.1997 | 0.008286 | 7.8754 | 6.5568 | 7.6898 | 6.0317
2012-10-30 13:59 | 22.4249 | 27.7389 | 26.4559 | 27.5534 | 25.9624 | 46.385 | 0.014194 | 0.34573 | 0.004907 | 24.3575 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.9228 | 3.1978 | 0.008284 | 7.8161 | 6.5331 | 7.6306 | 6.0396
2012-10-30 13:59 | 22.4227 | 27.6933 | 26.4379 | 27.5473 | 25.9167 | 46.3832 | 0.01419 | 0.34562 | 0.004904 | 24.3562 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 19.9206 | 3.1977 | 0.008279 | 7.7727 | 6.5173 | 7.6267 | 5.9961
2012-10-30 13:59 | 22.4279 | 27.7262 | 26.3997 | 27.513 | 25.9338 | 46.3983 | 0.014184 | 0.34567 | 0.004903 | 24.3706 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 19.9538 | 3.1987 | 0.008278 | 7.7724 | 6.4459 | 7.5592 | 5.98
2012-10-30 13:59 | 22.4301 | 27.677 | 26.3466 | 27.4836 | 25.8649 | 46.3757 | 0.014197 | 0.3457 | 0.004908 | 24.3503 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 19.944 | 3.1971 | 0.008284 | 7.733 | 6.4026 | 7.5396 | 5.9209
2012-10-30 13:59 | 22.4414 | 27.6804 | 26.3342 | 27.491 | 25.8644 | 46.3861 | 0.0142 | 0.34575 | 0.00491 | 24.3482 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 19.9673 | 3.1979 | 0.008285 | 7.7131 | 6.3669 | 7.5237 | 5.8971
2012-10-30 13:59 | 22.4414 | 27.6054 | 26.3737 | 27.4396 | 25.8841 | 46.385 | 0.014181 | 0.34569 | 0.004902 | 24.3768 | 0.014382 | 0.23455 | 0.003373 | 16.3087 | 19.9473 | 3.1978 | 0.008275 | 7.6581 | 6.4264 | 7.4923 | 5.9368
2012-10-30 13:59 | 22.4421 | 27.5785 | 26.3152 | 27.3851 | 25.8335 | 46.3723 | 0.014197 | 0.3457 | 0.004908 | 24.3503 | 0.014395 | 0.23464 | 0.003378 | 16.3 | 19.9161 | 3.1969 | 0.008286 | 7.6624 | 6.3991 | 7.469 | 5.9174
2012-10-30 13:59 | 22.4354 | 27.5481 | 26.3006 | 27.4178 | 25.8189 | 46.3907 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014398 | 0.23459 | 0.003378 | 16.2936 | 19.9453 | 3.1982 | 0.008282 | 7.6028 | 6.3553 | 7.4725 | 5.8736
2012-10-30 13:59 | 22.4503 | 27.5354 | 26.3155 | 27.3459 | 25.8259 | 46.363 | 0.014194 | 0.34569 | 0.004907 | 24.3544 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 19.9522 | 3.1963 | 0.008285 | 7.5832 | 6.3633 | 7.3937 | 5.8737
2012-10-30 13:59 | 22.4391 | 27.5084 | 26.2885 | 27.3623 | 25.7949 | 46.3902 | 0.01419 | 0.34566 | 0.004905 | 24.3593 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.989 | 3.1981 | 0.008282 | 7.5194 | 6.2995 | 7.3733 | 5.8059
2012-10-30 13:59 | 22.4503 | 27.4722 | 26.2365 | 27.3143 | 25.8496 | 46.3659 | 0.01419 | 0.34573 | 0.004906 | 24.3643 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.9842 | 3.1965 | 0.008282 | 7.488 | 6.2523 | 7.3301 | 5.8654
2012-10-30 13:59 | 22.4548 | 27.4412 | 26.2292 | 27.2793 | 25.7198 | 46.3671 | 0.014197 | 0.3457 | 0.004908 | 24.3503 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.9607 | 3.1965 | 0.008284 | 7.4805 | 6.2685 | 7.3186 | 5.7591
2012-10-30 13:59 | 22.4556 | 27.3788 | 26.1707 | 27.2249 | 25.7048 | 46.374 | 0.014184 | 0.34572 | 0.004904 | 24.3737 | 0.014386 | 0.23455 | 0.003374 | 16.3042 | 19.9655 | 3.197 | 0.008278 | 7.4133 | 6.2052 | 7.2594 | 5.7393
2012-10-30 13:59 | 22.463 | 27.402 | 26.1347 | 27.1849 | 25.6648 | 46.3457 | 0.014187 | 0.3457 | 0.004904 | 24.3674 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.917 | 3.1951 | 0.008281 | 7.485 | 6.2177 | 7.2679 | 5.7478
2012-10-30 13:59 | 22.4608 | 27.3485 | 26.1483 | 27.1669 | 25.6705 | 46.3601 | 0.01419 | 0.34569 | 0.004905 | 24.3613 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 19.8988 | 3.1961 | 0.008285 | 7.4497 | 6.2495 | 7.2681 | 5.7717
2012-10-30 13:59 | 22.4608 | 27.309 | 26.1364 | 27.1353 | 25.6153 | 46.3439 | 0.014181 | 0.34573 | 0.004903 | 24.3798 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9508 | 3.1949 | 0.008277 | 7.3582 | 6.1856 | 7.1845 | 5.6645
2012-10-30 13:59 | 22.4653 | 27.2859 | 26.1252 | 27.1201 | 25.5605 | 46.3468 | 0.014203 | 0.34567 | 0.00491 | 24.338 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 19.9473 | 3.1951 | 0.008287 | 7.3386 | 6.1779 | 7.1728 | 5.6132
2012-10-30 14:00 | 22.4713 | 27.2721 | 26.0956 | 27.0787 | 25.5705 | 46.3474 | 0.014187 | 0.34569 | 0.004904 | 24.3665 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 19.9812 | 3.1952 | 0.008282 | 7.2909 | 6.1144 | 7.0975 | 5.5893
2012-10-30 14:00 | 22.4765 | 27.2616 | 26.0692 | 27.0602 | 25.5757 | 46.3503 | 0.014197 | 0.34569 | 0.004908 | 24.3493 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 19.9784 | 3.1954 | 0.008283 | 7.2832 | 6.0908 | 7.0818 | 5.5973
2012-10-30 14:00 | 22.4683 | 27.202 | 26.0373 | 27.052 | 25.5635 | 46.3653 | 0.014187 | 0.3457 | 0.004904 | 24.3674 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9423 | 3.1964 | 0.00828 | 7.2597 | 6.095 | 7.1097 | 5.6212
2012-10-30 14:00 | 22.4713 | 27.1695 | 26.0009 | 26.9642 | 25.5152 | 46.3312 | 0.01419 | 0.34577 | 0.004907 | 24.3674 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 19.9852 | 3.1941 | 0.008284 | 7.1843 | 6.0157 | 6.979 | 5.53
2012-10-30 14:00 | 22.478 | 27.1328 | 25.9878 | 27.038 | 25.4785 | 46.333 | 0.014175 | 0.34567 | 0.0049 | 24.386 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.96 | 3.1942 | 0.008278 | 7.1728 | 6.0278 | 7.078 | 5.5185
2012-10-30 14:00 | 22.478 | 27.1051 | 25.9957 | 26.9788 | 25.4548 | 46.3347 | 0.014181 | 0.34567 | 0.004902 | 24.3757 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 19.968 | 3.1943 | 0.008278 | 7.1371 | 6.0277 | 7.0108 | 5.4868
2012-10-30 14:00 | 22.481 | 27.1437 | 25.9829 | 26.9384 | 25.4578 | 46.3191 | 0.014181 | 0.34569 | 0.004902 | 24.3768 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 19.931 | 3.1932 | 0.008277 | 7.2127 | 6.0519 | 7.0074 | 5.5268
2012-10-30 14:00 | 22.4787 | 27.0743 | 25.9925 | 26.9558 | 25.4081 | 46.3428 | 0.01419 | 0.34566 | 0.004905 | 24.3593 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9287 | 3.1949 | 0.008279 | 7.1456 | 6.0638 | 7.0271 | 5.4794
2012-10-30 14:00 | 22.484 | 27.0914 | 25.9306 | 26.9572 | 25.4371 | 46.3174 | 0.014187 | 0.34564 | 0.004904 | 24.3634 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 19.93 | 3.1931 | 0.00828 | 7.1614 | 6.0006 | 7.0272 | 5.5071
2012-10-30 14:00 | 22.4899 | 27.0499 | 25.9207 | 26.8446 | 25.368 | 46.3208 | 0.014203 | 0.3457 | 0.00491 | 24.34 | 0.014389 | 0.23448 | 0.003374 | 16.2957 | 19.9479 | 3.1934 | 0.008284 | 7.102 | 5.9728 | 6.8967 | 5.4201
2012-10-30 14:00 | 22.4944 | 27.0702 | 25.8897 | 26.9162 | 25.3804 | 46.3162 | 0.014197 | 0.34566 | 0.004907 | 24.3473 | 0.014386 | 0.23458 | 0.003375 | 16.3061 | 19.9684 | 3.193 | 0.008282 | 7.1018 | 5.9213 | 6.9478 | 5.412
2012-10-30 14:00 | 22.487 | 26.9957 | 25.8626 | 26.818 | 25.3769 | 46.3168 | 0.014197 | 0.34562 | 0.004907 | 24.3442 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 19.941 | 3.1931 | 0.008283 | 7.0547 | 5.9216 | 6.877 | 5.4359
2012-10-30 14:00 | 22.4847 | 26.9737 | 25.8484 | 26.8078 | 25.3944 | 46.3197 | 0.014194 | 0.34563 | 0.004906 | 24.3504 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9028 | 3.1933 | 0.008282 | 7.0709 | 5.9456 | 6.905 | 5.4916
2012-10-30 14:00 | 22.4959 | 26.9572 | 25.8596 | 26.8072 | 25.299 | 46.3116 | 0.01419 | 0.34566 | 0.004905 | 24.3593 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 19.9539 | 3.1927 | 0.00828 | 7.0033 | 5.9057 | 6.8533 | 5.3451
2012-10-30 14:00 | 22.5019 | 26.9514 | 25.8301 | 26.7737 | 25.3286 | 46.3208 | 0.0142 | 0.34573 | 0.004909 | 24.3472 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9799 | 3.1934 | 0.008286 | 6.9715 | 5.8502 | 6.7938 | 5.3487
2012-10-30 14:00 | 22.4944 | 26.932 | 25.7989 | 26.7781 | 25.3488 | 46.3122 | 0.01419 | 0.34564 | 0.004905 | 24.3582 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9245 | 3.1928 | 0.008282 | 7.0075 | 5.8744 | 6.8536 | 5.4243
2012-10-30 14:00 | 22.5026 | 26.8968 | 25.7795 | 26.7507 | 25.2741 | 46.3243 | 0.014197 | 0.34569 | 0.004908 | 24.3493 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9327 | 3.1936 | 0.008285 | 6.9641 | 5.8468 | 6.818 | 5.3414
2012-10-30 14:00 | 22.5056 | 26.8998 | 25.7785 | 26.7221 | 25.2178 | 46.3023 | 0.014203 | 0.34567 | 0.00491 | 24.338 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 19.9436 | 3.1921 | 0.008285 | 6.9562 | 5.8349 | 6.7785 | 5.2742
2012-10-30 14:00 | 22.5101 | 26.8846 | 25.7475 | 26.6832 | 25.2539 | 46.296 | 0.014194 | 0.34569 | 0.004907 | 24.3544 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.9761 | 3.1916 | 0.008283 | 6.9085 | 5.7714 | 6.7071 | 5.2778
2012-10-30 14:00 | 22.4974 | 26.8205 | 25.7466 | 26.6587 | 25.2807 | 46.3237 | 0.0142 | 0.34562 | 0.004908 | 24.3391 | 0.014389 | 0.23459 | 0.003376 | 16.3038 | 19.9035 | 3.1936 | 0.008284 | 6.917 | 5.8431 | 6.7552 | 5.3772
2012-10-30 14:00 | 22.5116 | 26.8032 | 25.7332 | 26.6689 | 25.208 | 46.2913 | 0.014178 | 0.34572 | 0.004902 | 24.384 | 0.014386 | 0.23455 | 0.003374 | 16.3042 | 19.9616 | 3.1913 | 0.008276 | 6.8416 | 5.7716 | 6.7073 | 5.2464
2012-10-30 14:00 | 22.5056 | 26.813 | 25.6601 | 26.6234 | 25.1784 | 46.3029 | 0.014197 | 0.34573 | 0.004908 | 24.3523 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 19.9157 | 3.1921 | 0.008284 | 6.8973 | 5.7444 | 6.7077 | 5.2627
2012-10-30 14:00 | 22.5191 | 26.783 | 25.713 | 26.6488 | 25.2155 | 46.2844 | 0.014194 | 0.34564 | 0.004906 | 24.3513 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 19.9172 | 3.1908 | 0.008286 | 6.8658 | 5.7958 | 6.7316 | 5.2983
2012-10-30 14:00 | 22.5124 | 26.7408 | 25.6511 | 26.6263 | 25.1575 | 46.3168 | 0.014203 | 0.34564 | 0.004909 | 24.3359 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 19.9544 | 3.1931 | 0.008285 | 6.7864 | 5.6967 | 6.6719 | 5.2031
2012-10-30 14:00 | 22.5206 | 26.7529 | 25.6553 | 26.6108 | 25.1736 | 46.2971 | 0.014209 | 0.34556 | 0.00491 | 24.3196 | 0.014382 | 0.23448 | 0.003372 | 16.3036 | 19.9347 | 3.1917 | 0.008282 | 6.8182 | 5.7206 | 6.6761 | 5.2389
2012-10-30 14:00 | 22.5258 | 26.7345 | 25.6052 | 26.5923 | 25.1788 | 46.2989 | 0.014209 | 0.34567 | 0.004912 | 24.3277 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9599 | 3.1918 | 0.008289 | 6.7746 | 5.6453 | 6.6324 | 5.2189
2012-10-30 14:00 | 22.5258 | 26.7187 | 25.6052 | 26.5173 | 25.1235 | 46.289 | 0.01419 | 0.34562 | 0.004904 | 24.3562 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.9718 | 3.1912 | 0.008281 | 6.7469 | 5.6334 | 6.5455 | 5.1517
2012-10-30 14:00 | 22.5251 | 26.714 | 25.6203 | 26.489 | 25.0912 | 46.2965 | 0.014203 | 0.34569 | 0.00491 | 24.339 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.9671 | 3.1917 | 0.008288 | 6.7469 | 5.6532 | 6.5219 | 5.1241
2012-10-30 14:00 | 22.5243 | 26.65 | 25.5603 | 26.5158 | 25.0707 | 46.2977 | 0.014172 | 0.3457 | 0.004899 | 24.3932 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 19.9504 | 3.1918 | 0.008275 | 6.6996 | 5.6099 | 6.5654 | 5.1203
2012-10-30 14:01 | 22.543 | 26.649 | 25.5869 | 26.4634 | 25.0539 | 46.285 | 0.014197 | 0.34563 | 0.004907 | 24.3452 | 0.014386 | 0.23462 | 0.003375 | 16.3092 | 19.8932 | 3.1909 | 0.008282 | 6.7558 | 5.6937 | 6.5702 | 5.1607
2012-10-30 14:01 | 22.5363 | 26.6147 | 25.5881 | 26.4725 | 25.0788 | 46.2908 | 0.0142 | 0.34573 | 0.004909 | 24.3472 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9504 | 3.1913 | 0.008285 | 6.6643 | 5.6377 | 6.5221 | 5.1284
2012-10-30 14:01 | 22.5348 | 26.5816 | 25.5274 | 26.3921 | 25.0536 | 46.2729 | 0.0142 | 0.34562 | 0.004908 | 24.3391 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 19.9569 | 3.1901 | 0.008283 | 6.6247 | 5.5705 | 6.4352 | 5.0967
2012-10-30 14:01 | 22.5415 | 26.5883 | 25.4867 | 26.383 | 25.0682 | 46.289 | 0.014197 | 0.34566 | 0.004907 | 24.3473 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.9356 | 3.1912 | 0.008283 | 6.6527 | 5.5511 | 6.4474 | 5.1326
2012-10-30 14:01 | 22.5325 | 26.5122 | 25.4777 | 26.3819 | 25.0236 | 46.2607 | 0.014187 | 0.34569 | 0.004904 | 24.3665 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 19.9745 | 3.1892 | 0.008278 | 6.5377 | 5.5032 | 6.4074 | 5.0491
2012-10-30 14:01 | 22.5438 | 26.5353 | 25.4495 | 26.3339 | 24.956 | 46.2931 | 0.014194 | 0.34564 | 0.004906 | 24.3513 | 0.014382 | 0.23458 | 0.003374 | 16.3107 | 19.9339 | 3.1914 | 0.00828 | 6.6014 | 5.5156 | 6.4 | 5.0221
2012-10-30 14:01 | 22.5505 | 26.5104 | 25.4681 | 26.3683 | 24.9864 | 46.2578 | 0.014197 | 0.34564 | 0.004907 | 24.3462 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9206 | 3.189 | 0.008285 | 6.5898 | 5.5475 | 6.4477 | 5.0658
2012-10-30 14:01 | 22.5468 | 26.4633 | 25.4091 | 26.2935 | 24.9471 | 46.2867 | 0.014187 | 0.3457 | 0.004904 | 24.3674 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 19.9169 | 3.191 | 0.008279 | 6.5464 | 5.4922 | 6.3766 | 5.0302
2012-10-30 14:01 | 22.5385 | 26.4274 | 25.4561 | 26.3129 | 24.9586 | 46.259 | 0.014194 | 0.34564 | 0.004906 | 24.3513 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 19.9446 | 3.1891 | 0.008281 | 6.4828 | 5.5115 | 6.3683 | 5.014
2012-10-30 14:01 | 22.5535 | 26.4345 | 25.4039 | 26.2449 | 24.9104 | 46.2619 | 0.014187 | 0.34569 | 0.004904 | 24.3665 | 0.014392 | 0.23462 | 0.003377 | 16.3024 | 19.9516 | 3.1893 | 0.008281 | 6.4829 | 5.4523 | 6.2933 | 4.9588
2012-10-30 14:01 | 22.5505 | 26.4394 | 25.3496 | 26.2696 | 24.8995 | 46.2757 | 0.014187 | 0.34566 | 0.004904 | 24.3644 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9486 | 3.1902 | 0.008281 | 6.4908 | 5.401 | 6.321 | 4.9509
2012-10-30 14:01 | 22.5572 | 26.4303 | 25.3958 | 26.221 | 24.9338 | 46.2659 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014382 | 0.23452 | 0.003373 | 16.3067 | 19.9753 | 3.1896 | 0.008278 | 6.455 | 5.4205 | 6.2457 | 4.9585
2012-10-30 14:01 | 22.5505 | 26.3643 | 25.3496 | 26.2222 | 24.8718 | 46.259 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9606 | 3.1891 | 0.008281 | 6.4037 | 5.389 | 6.2616 | 4.9112
2012-10-30 14:01 | 22.5639 | 26.3896 | 25.3275 | 26.1882 | 24.8655 | 46.2549 | 0.014187 | 0.34569 | 0.004904 | 24.3665 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 19.934 | 3.1888 | 0.00828 | 6.4556 | 5.3935 | 6.2542 | 4.9315
2012-10-30 14:01 | 22.5572 | 26.3631 | 25.3168 | 26.1736 | 24.8272 | 46.2665 | 0.014184 | 0.3457 | 0.004903 | 24.3726 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.9194 | 3.1896 | 0.008279 | 6.4437 | 5.3974 | 6.2542 | 4.9078
2012-10-30 14:01 | 22.5699 | 26.3521 | 25.29 | 26.1666 | 24.836 | 46.2463 | 0.014197 | 0.34564 | 0.004907 | 24.3462 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.952 | 3.1882 | 0.008284 | 6.4001 | 5.338 | 6.2146 | 4.884
2012-10-30 14:01 | 22.5654 | 26.3003 | 25.2816 | 26.1305 | 24.8394 | 46.2659 | 0.014197 | 0.34566 | 0.004907 | 24.3473 | 0.014398 | 0.23461 | 0.003378 | 16.2946 | 19.9635 | 3.1896 | 0.008285 | 6.3368 | 5.3181 | 6.167 | 4.8759
2012-10-30 14:01 | 22.5632 | 26.2704 | 25.2715 | 26.1204 | 24.7858 | 46.2399 | 0.0142 | 0.34567 | 0.004909 | 24.3431 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 19.9134 | 3.1878 | 0.008286 | 6.357 | 5.3581 | 6.207 | 4.8724
2012-10-30 14:01 | 22.5722 | 26.2755 | 25.2923 | 26.1215 | 24.7909 | 46.2619 | 0.01419 | 0.3456 | 0.004904 | 24.3552 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9703 | 3.1893 | 0.008278 | 6.3052 | 5.322 | 6.1512 | 4.8206
2012-10-30 14:01 | 22.5752 | 26.2232 | 25.2519 | 26.1008 | 24.7978 | 46.2416 | 0.014184 | 0.3456 | 0.004902 | 24.3655 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 19.9493 | 3.1879 | 0.008282 | 6.2739 | 5.3026 | 6.1515 | 4.8485
2012-10-30 14:01 | 22.5879 | 26.232 | 25.2409 | 26.0582 | 24.7868 | 46.2584 | 0.014197 | 0.34567 | 0.004908 | 24.3482 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0339 | 3.1891 | 0.008283 | 6.1981 | 5.207 | 6.0243 | 4.7529
2012-10-30 14:01 | 22.5864 | 26.1831 | 25.2039 | 26.037 | 24.7498 | 46.2445 | 0.0142 | 0.34566 | 0.004908 | 24.3421 | 0.014382 | 0.23457 | 0.003374 | 16.3097 | 19.9565 | 3.1881 | 0.008282 | 6.2266 | 5.2474 | 6.0805 | 4.7933
2012-10-30 14:01 | 22.5781 | 26.1787 | 25.1877 | 26.0287 | 24.7178 | 46.2538 | 0.014194 | 0.34572 | 0.004907 | 24.3565 | 0.014398 | 0.23459 | 0.003378 | 16.2936 | 19.9522 | 3.1887 | 0.008285 | 6.2265 | 5.2355 | 6.0765 | 4.7656
2012-10-30 14:01 | 22.5849 | 26.1776 | 25.1629 | 26.0157 | 24.7286 | 46.2237 | 0.014187 | 0.34564 | 0.004904 | 24.3634 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 19.963 | 3.1867 | 0.008282 | 6.2146 | 5.1999 | 6.0527 | 4.7656
2012-10-30 14:01 | 22.5886 | 26.1339 | 25.1587 | 25.9721 | 24.677 | 46.2411 | 0.014194 | 0.34564 | 0.004906 | 24.3513 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9508 | 3.1879 | 0.008284 | 6.1831 | 5.2079 | 6.0213 | 4.7262
2012-10-30 14:01 | 22.5886 | 26.1458 | 25.1942 | 25.9286 | 24.6928 | 46.2353 | 0.014187 | 0.34566 | 0.004904 | 24.3644 | 0.014392 | 0.23462 | 0.003377 | 16.3024 | 19.9787 | 3.1875 | 0.008281 | 6.1671 | 5.2155 | 5.9499 | 4.7141
2012-10-30 14:01 | 22.5938 | 26.1194 | 25.1205 | 25.9812 | 24.623 | 46.2399 | 0.014194 | 0.34566 | 0.004906 | 24.3524 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.0199 | 3.1878 | 0.008283 | 6.0995 | 5.1006 | 5.9613 | 4.6031
2012-10-30 14:01 | 22.5976 | 26.0877 | 25.0769 | 25.9732 | 24.6031 | 46.2232 | 0.014181 | 0.34563 | 0.004901 | 24.3727 | 0.014401 | 0.23464 | 0.003379 | 16.2932 | 19.9957 | 3.1866 | 0.00828 | 6.092 | 5.0812 | 5.9775 | 4.6074
2012-10-30 14:02 | 22.6065 | 26.0137 | 25.0345 | 25.8597 | 24.5646 | 46.2237 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9567 | 3.1867 | 0.008281 | 6.057 | 5.0778 | 5.903 | 4.6079
2012-10-30 14:02 | 22.6208 | 25.9569 | 24.9777 | 25.7753 | 24.5631 | 46.2023 | 0.014203 | 0.34573 | 0.00491 | 24.342 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.967 | 3.1852 | 0.008287 | 5.9899 | 5.0107 | 5.8083 | 4.5961
2012-10-30 14:02 | 22.6178 | 25.942 | 24.9786 | 25.7644 | 24.4811 | 46.2203 | 0.01419 | 0.34559 | 0.004904 | 24.3542 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 19.924 | 3.1864 | 0.008282 | 6.018 | 5.0546 | 5.8404 | 4.5571
2012-10-30 14:02 | 22.6185 | 25.9033 | 24.9477 | 25.7453 | 24.5095 | 46.2029 | 0.0142 | 0.34563 | 0.004908 | 24.3401 | 0.014382 | 0.23457 | 0.003374 | 16.3097 | 19.9886 | 3.1852 | 0.008282 | 5.9147 | 4.9591 | 5.7567 | 4.5209
2012-10-30 14:02 | 22.6163 | 25.9169 | 24.8824 | 25.7273 | 24.4994 | 46.2174 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9665 | 3.1862 | 0.008283 | 5.9504 | 4.9159 | 5.7608 | 4.5329
2012-10-30 14:02 | 22.6208 | 25.8463 | 24.875 | 25.7476 | 24.4605 | 46.1827 | 0.01419 | 0.34564 | 0.004905 | 24.3582 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0109 | 3.1838 | 0.008282 | 5.8354 | 4.8641 | 5.7367 | 4.4496
2012-10-30 14:02 | 22.623 | 25.8209 | 24.9088 | 25.7459 | 24.4192 | 46.2364 | 0.014187 | 0.3456 | 0.004903 | 24.3603 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9852 | 3.1875 | 0.008279 | 5.8357 | 4.9236 | 5.7607 | 4.434
2012-10-30 14:02 | 22.6245 | 25.8184 | 24.8313 | 25.6803 | 24.4405 | 46.2081 | 0.014194 | 0.34566 | 0.004906 | 24.3524 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0066 | 3.1856 | 0.008282 | 5.8118 | 4.8247 | 5.6737 | 4.4339
2012-10-30 14:02 | 22.6305 | 25.8205 | 24.8294 | 25.6507 | 24.4504 | 46.2151 | 0.014184 | 0.3456 | 0.004902 | 24.3655 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 19.9527 | 3.1861 | 0.008281 | 5.8678 | 4.8767 | 5.698 | 4.4977
2012-10-30 14:02 | 22.6275 | 25.7938 | 24.8264 | 25.6082 | 24.3803 | 46.1891 | 0.014206 | 0.34563 | 0.00491 | 24.3298 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 19.9098 | 3.1843 | 0.008287 | 5.884 | 4.9166 | 5.6984 | 4.4705
2012-10-30 14:02 | 22.6342 | 25.7847 | 24.8331 | 25.6426 | 24.387 | 46.2197 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014386 | 0.23458 | 0.003375 | 16.3061 | 19.9964 | 3.1864 | 0.008279 | 5.7883 | 4.8367 | 5.6462 | 4.3906
2012-10-30 14:02 | 22.6417 | 25.8001 | 24.8485 | 25.6264 | 24.3708 | 46.1925 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0158 | 3.1845 | 0.008284 | 5.7843 | 4.8327 | 5.6106 | 4.355
2012-10-30 14:02 | 22.6305 | 25.7376 | 24.7663 | 25.6112 | 24.3438 | 46.2151 | 0.014203 | 0.34566 | 0.004909 | 24.337 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0126 | 3.1861 | 0.008287 | 5.725 | 4.7537 | 5.5986 | 4.3312
2012-10-30 14:02 | 22.6387 | 25.7221 | 24.8337 | 25.5326 | 24.3718 | 46.1891 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 19.9369 | 3.1843 | 0.008279 | 5.7852 | 4.8968 | 5.5957 | 4.4349
2012-10-30 14:02 | 22.6454 | 25.7485 | 24.7812 | 25.5945 | 24.3271 | 46.2226 | 0.014187 | 0.34563 | 0.004903 | 24.3624 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 19.9436 | 3.1866 | 0.008277 | 5.8049 | 4.8376 | 5.6509 | 4.3835
2012-10-30 14:02 | 22.6379 | 25.6858 | 24.7618 | 25.5199 | 24.3078 | 46.1862 | 0.014172 | 0.3456 | 0.004898 | 24.3861 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9921 | 3.1841 | 0.008275 | 5.6937 | 4.7697 | 5.5278 | 4.3157
2012-10-30 14:02 | 22.6394 | 25.6794 | 24.7436 | 25.5372 | 24.329 | 46.1908 | 0.014206 | 0.34554 | 0.004909 | 24.3237 | 0.014389 | 0.23457 | 0.003375 | 16.3018 | 19.9896 | 3.1844 | 0.008284 | 5.6898 | 4.754 | 5.5476 | 4.3394
2012-10-30 14:02 | 22.6507 | 25.6867 | 24.7667 | 25.4893 | 24.3443 | 46.1815 | 0.014197 | 0.34566 | 0.004907 | 24.3473 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9489 | 3.1838 | 0.008284 | 5.7378 | 4.8178 | 5.5404 | 4.3954
2012-10-30 14:02 | 22.6536 | 25.6343 | 24.7223 | 25.508 | 24.2958 | 46.1971 | 0.014194 | 0.34566 | 0.004906 | 24.3524 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 19.9918 | 3.1848 | 0.008284 | 5.6425 | 4.7305 | 5.5162 | 4.304
2012-10-30 14:02 | 22.6469 | 25.5881 | 24.7156 | 25.4342 | 24.2694 | 46.1815 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.9691 | 3.1838 | 0.008284 | 5.619 | 4.7465 | 5.4651 | 4.3003
2012-10-30 14:02 | 22.6641 | 25.6527 | 24.6854 | 25.4672 | 24.2669 | 46.2133 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.0023 | 3.1859 | 0.008285 | 5.6504 | 4.6831 | 5.4649 | 4.2646
2012-10-30 14:02 | 22.6521 | 25.5933 | 24.6971 | 25.4394 | 24.2509 | 46.1729 | 0.014194 | 0.3457 | 0.004907 | 24.3554 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 19.9663 | 3.1832 | 0.008282 | 5.627 | 4.7308 | 5.4731 | 4.2846
2012-10-30 14:02 | 22.6581 | 25.5599 | 24.6952 | 25.4296 | 24.253 | 46.1943 | 0.014197 | 0.34563 | 0.004907 | 24.3452 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 19.9723 | 3.1846 | 0.008285 | 5.5876 | 4.7229 | 5.4573 | 4.2807
2012-10-30 14:02 | 22.6574 | 25.5947 | 24.6313 | 25.417 | 24.2128 | 46.1613 | 0.014187 | 0.34563 | 0.004903 | 24.3624 | 0.014386 | 0.23457 | 0.003374 | 16.3052 | 19.9916 | 3.1824 | 0.008277 | 5.6031 | 4.6397 | 5.4254 | 4.2212
2012-10-30 14:02 | 22.6641 | 25.554 | 24.638 | 25.4079 | 24.2274 | 46.1896 | 0.0142 | 0.3456 | 0.004908 | 24.338 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 19.9863 | 3.1843 | 0.008286 | 5.5677 | 4.6517 | 5.4216 | 4.2411
2012-10-30 14:02 | 22.6708 | 25.5331 | 24.6171 | 25.3436 | 24.2104 | 46.1682 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.001 | 3.1828 | 0.00828 | 5.5321 | 4.6161 | 5.3426 | 4.2094
2012-10-30 14:02 | 22.6671 | 25.5254 | 24.6134 | 25.3478 | 24.1988 | 46.1804 | 0.014181 | 0.3456 | 0.004901 | 24.3706 | 0.014401 | 0.23444 | 0.003376 | 16.2791 | 20.0133 | 3.1837 | 0.008277 | 5.5121 | 4.6001 | 5.3345 | 4.1855
2012-10-30 14:03 | 22.6791 | 25.4822 | 24.6254 | 25.3163 | 24.1832 | 46.1786 | 0.0142 | 0.34566 | 0.004908 | 24.3421 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 19.9933 | 3.1836 | 0.008285 | 5.4889 | 4.6321 | 5.323 | 4.1899
2012-10-30 14:03 | 22.6701 | 25.5047 | 24.6124 | 25.3665 | 24.1821 | 46.1839 | 0.01419 | 0.34557 | 0.004904 | 24.3531 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.0003 | 3.1839 | 0.00828 | 5.5044 | 4.6121 | 5.3662 | 4.1818
2012-10-30 14:03 | 22.6723 | 25.499 | 24.6265 | 25.3214 | 24.1369 | 46.1694 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014382 | 0.23448 | 0.003372 | 16.3036 | 19.9865 | 3.1829 | 0.008277 | 5.5125 | 4.64 | 5.3349 | 4.1504
2012-10-30 14:03 | 22.6761 | 25.416 | 24.5513 | 25.2857 | 24.1407 | 46.1862 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 19.9943 | 3.1841 | 0.008281 | 5.4217 | 4.557 | 5.2914 | 4.1464
2012-10-30 14:03 | 22.6865 | 25.3948 | 24.538 | 25.2684 | 24.1511 | 46.1573 | 0.014209 | 0.34564 | 0.004911 | 24.3256 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.9528 | 3.1821 | 0.008287 | 5.442 | 4.5852 | 5.3156 | 4.1983
2012-10-30 14:03 | 22.6828 | 25.3792 | 24.5698 | 25.2489 | 24.1395 | 46.1723 | 0.0142 | 0.34557 | 0.004907 | 24.336 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9651 | 3.1831 | 0.008283 | 5.4141 | 4.6047 | 5.2838 | 4.1744
2012-10-30 14:03 | 22.6835 | 25.376 | 24.4916 | 25.2575 | 24.1086 | 46.1642 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.0217 | 3.1826 | 0.008281 | 5.3543 | 4.4699 | 5.2358 | 4.0869
2012-10-30 14:03 | 22.6828 | 25.3516 | 24.4909 | 25.2332 | 24.0684 | 46.1515 | 0.014197 | 0.34557 | 0.004906 | 24.3411 | 0.014386 | 0.23449 | 0.003373 | 16.3002 | 19.981 | 3.1817 | 0.008279 | 5.3706 | 4.5099 | 5.2522 | 4.0874
2012-10-30 14:03 | 22.6858 | 25.3665 | 24.5255 | 25.2006 | 24.0872 | 46.1526 | 0.0142 | 0.3456 | 0.004908 | 24.338 | 0.014382 | 0.23458 | 0.003374 | 16.3107 | 19.9721 | 3.1818 | 0.008282 | 5.3944 | 4.5534 | 5.2285 | 4.1151
2012-10-30 14:03 | 22.685 | 25.3578 | 24.4931 | 25.1761 | 24.1022 | 46.1538 | 0.014187 | 0.34563 | 0.004903 | 24.3624 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9992 | 3.1818 | 0.008279 | 5.3586 | 4.4939 | 5.1769 | 4.103
2012-10-30 14:03 | 22.6955 | 25.3169 | 24.4601 | 25.1866 | 24.0653 | 46.1567 | 0.014222 | 0.34559 | 0.004915 | 24.2994 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.0137 | 3.182 | 0.008294 | 5.3032 | 4.4464 | 5.1729 | 4.0516
2012-10-30 14:03 | 22.7 | 25.3175 | 24.4488 | 25.199 | 24.0264 | 46.1446 | 0.0142 | 0.34553 | 0.004906 | 24.333 | 0.014386 | 0.23461 | 0.003375 | 16.3081 | 20.0142 | 3.1812 | 0.008281 | 5.3033 | 4.4346 | 5.1848 | 4.0122
2012-10-30 14:03 | 22.7015 | 25.3032 | 24.4819 | 25.1452 | 24.0397 | 46.1711 | 0.014197 | 0.34563 | 0.004907 | 24.3452 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.0077 | 3.183 | 0.008285 | 5.2955 | 4.4742 | 5.1375 | 4.032
2012-10-30 14:03 | 22.691 | 25.3164 | 24.4517 | 25.1229 | 24.0095 | 46.1295 | 0.014197 | 0.34563 | 0.004907 | 24.3452 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 19.9573 | 3.1802 | 0.008284 | 5.3591 | 4.4944 | 5.1656 | 4.0522
2012-10-30 14:03 | 22.6925 | 25.2231 | 24.465 | 25.1362 | 24.0386 | 46.1665 | 0.014184 | 0.34557 | 0.004902 | 24.3634 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 19.9748 | 3.1827 | 0.00828 | 5.2483 | 4.4902 | 5.1614 | 4.0638
2012-10-30 14:03 | 22.7045 | 25.2667 | 24.4139 | 25.1246 | 24.0151 | 46.1417 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 19.9987 | 3.181 | 0.008284 | 5.268 | 4.4152 | 5.1259 | 4.0164
2012-10-30 14:03 | 22.7075 | 25.246 | 24.4327 | 25.0762 | 24.0142 | 46.1596 | 0.014194 | 0.34564 | 0.004906 | 24.3513 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 19.9698 | 3.1822 | 0.008281 | 5.2762 | 4.4629 | 5.1064 | 4.0444
2012-10-30 14:03 | 22.7022 | 25.221 | 24.376 | 25.0591 | 23.9575 | 46.1417 | 0.01419 | 0.3456 | 0.004904 | 24.3552 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 19.9405 | 3.181 | 0.008279 | 5.2805 | 4.4355 | 5.1186 | 4.017
2012-10-30 14:03 | 22.7082 | 25.2151 | 24.4057 | 25.0927 | 23.9793 | 46.148 | 0.014187 | 0.34559 | 0.004903 | 24.3593 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9425 | 3.1814 | 0.00828 | 5.2726 | 4.4632 | 5.1502 | 4.0368
2012-10-30 14:03 | 22.706 | 25.1497 | 24.3759 | 25.051 | 23.9811 | 46.1324 | 0.014203 | 0.34554 | 0.004908 | 24.3288 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 19.9363 | 3.1804 | 0.008285 | 5.2134 | 4.4396 | 5.1147 | 4.0448
2012-10-30 14:03 | 22.7172 | 25.1531 | 24.3436 | 25.0228 | 23.9449 | 46.1469 | 0.014209 | 0.34566 | 0.004911 | 24.3267 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 19.9595 | 3.1814 | 0.008287 | 5.1936 | 4.3841 | 5.0633 | 3.9854
2012-10-30 14:03 | 22.7194 | 25.1671 | 24.334 | 24.9973 | 23.9866 | 46.1405 | 0.014194 | 0.3456 | 0.004905 | 24.3483 | 0.014382 | 0.23452 | 0.003373 | 16.3067 | 19.9617 | 3.1809 | 0.008278 | 5.2054 | 4.3723 | 5.0356 | 4.0249
2012-10-30 14:03 | 22.7269 | 25.1509 | 24.3336 | 25.0206 | 23.9388 | 46.1336 | 0.01419 | 0.34559 | 0.004904 | 24.3542 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 19.9652 | 3.1805 | 0.008278 | 5.1857 | 4.3684 | 5.0554 | 3.9736
2012-10-30 14:03 | 22.7187 | 25.1388 | 24.2543 | 24.9927 | 23.899 | 46.144 | 0.01419 | 0.34562 | 0.004904 | 24.3562 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 19.957 | 3.1812 | 0.008283 | 5.1818 | 4.2973 | 5.0357 | 3.942
2012-10-30 14:03 | 22.7262 | 25.1107 | 24.3211 | 24.9646 | 23.9144 | 46.107 | 0.014206 | 0.34562 | 0.00491 | 24.3288 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 19.9765 | 3.1786 | 0.008285 | 5.1342 | 4.3446 | 4.9881 | 3.9379
2012-10-30 14:03 | 22.7269 | 25.0719 | 24.3257 | 24.9495 | 23.8796 | 46.1336 | 0.0142 | 0.34556 | 0.004907 | 24.335 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 19.9812 | 3.1805 | 0.008283 | 5.0907 | 4.3445 | 4.9683 | 3.8984
2012-10-30 14:03 | 22.7336 | 25.0707 | 24.2969 | 24.9168 | 23.8508 | 46.1422 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9599 | 3.181 | 0.008284 | 5.1108 | 4.337 | 4.9569 | 3.8909
2012-10-30 14:03 | 22.7232 | 24.9972 | 24.2904 | 24.9379 | 23.8325 | 46.1289 | 0.014197 | 0.34562 | 0.004907 | 24.3442 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 19.9695 | 3.1801 | 0.008284 | 5.0277 | 4.3209 | 4.9684 | 3.863
2012-10-30 14:03 | 22.7419 | 25.0672 | 24.2341 | 24.9014 | 23.863 | 46.1266 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014392 | 0.23445 | 0.003374 | 16.2903 | 19.9283 | 3.18 | 0.00828 | 5.1389 | 4.3058 | 4.9731 | 3.9347
2012-10-30 14:03 | 22.7329 | 25.0464 | 24.2172 | 24.8805 | 23.8066 | 46.1353 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014389 | 0.23447 | 0.003374 | 16.2947 | 19.9432 | 3.1806 | 0.008275 | 5.1032 | 4.274 | 4.9373 | 3.8634
2012-10-30 14:04 | 22.7344 | 25.0163 | 24.2148 | 24.8781 | 23.7884 | 46.1266 | 0.01419 | 0.34559 | 0.004904 | 24.3542 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 19.9407 | 3.18 | 0.008278 | 5.0756 | 4.2741 | 4.9374 | 3.8477
2012-10-30 14:04 | 22.7344 | 24.9689 | 24.1792 | 24.807 | 23.8239 | 46.1226 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 19.9327 | 3.1797 | 0.008279 | 5.0362 | 4.2465 | 4.8743 | 3.8912
2012-10-30 14:04 | 22.7299 | 24.9762 | 24.1747 | 24.8104 | 23.7444 | 46.1382 | 0.014194 | 0.34556 | 0.004905 | 24.3453 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9402 | 3.1808 | 0.008282 | 5.036 | 4.2345 | 4.8702 | 3.8042
2012-10-30 14:04 | 22.7336 | 24.9365 | 24.1626 | 24.8141 | 23.7639 | 46.1093 | 0.014194 | 0.34563 | 0.004906 | 24.3504 | 0.014382 | 0.23452 | 0.003373 | 16.3067 | 19.9759 | 3.1788 | 0.008279 | 4.9606 | 4.1867 | 4.8382 | 3.788
2012-10-30 14:04 | 22.7441 | 24.9707 | 24.181 | 24.8009 | 23.7231 | 46.1336 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 19.9464 | 3.1805 | 0.008279 | 5.0243 | 4.2346 | 4.8545 | 3.7767
2012-10-30 14:04 | 22.7426 | 24.9455 | 24.1756 | 24.7915 | 23.7532 | 46.111 | 0.014187 | 0.34556 | 0.004902 | 24.3573 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 19.925 | 3.1789 | 0.008281 | 5.0205 | 4.2506 | 4.8665 | 3.8282
2012-10-30 14:04 | 22.7411 | 24.9322 | 24.1228 | 24.79 | 23.7556 | 46.1197 | 0.01419 | 0.3456 | 0.004904 | 24.3552 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 19.9474 | 3.1795 | 0.008279 | 4.9848 | 4.1754 | 4.8426 | 3.8082
2012-10-30 14:04 | 22.7478 | 24.8638 | 24.1492 | 24.7533 | 23.6952 | 46.1012 | 0.014168 | 0.34554 | 0.004896 | 24.3889 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9182 | 3.1782 | 0.008274 | 4.9456 | 4.231 | 4.8351 | 3.777
2012-10-30 14:04 | 22.7441 | 24.9154 | 24.0863 | 24.7496 | 23.6994 | 46.1232 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9345 | 3.1797 | 0.008283 | 4.9809 | 4.1518 | 4.8151 | 3.7649
2012-10-30 14:04 | 22.7448 | 24.8095 | 24.1028 | 24.7226 | 23.6646 | 46.0989 | 0.014184 | 0.34557 | 0.004902 | 24.3634 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 19.9511 | 3.1781 | 0.008277 | 4.8584 | 4.1517 | 4.7715 | 3.7135
2012-10-30 14:04 | 22.7508 | 24.8866 | 24.1009 | 24.6931 | 23.6429 | 46.1336 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014386 | 0.23457 | 0.003374 | 16.3052 | 19.9491 | 3.1805 | 0.008278 | 4.9375 | 4.1518 | 4.744 | 3.6938
2012-10-30 14:04 | 22.7568 | 24.8531 | 24.0556 | 24.711 | 23.6015 | 46.0995 | 0.014175 | 0.34559 | 0.004899 | 24.38 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 19.9432 | 3.1781 | 0.008273 | 4.9099 | 4.1124 | 4.7678 | 3.6583
2012-10-30 14:04 | 22.7546 | 24.8233 | 24.077 | 24.6851 | 23.6112 | 46.1185 | 0.014197 | 0.34557 | 0.004906 | 24.3411 | 0.014398 | 0.23459 | 0.003378 | 16.2936 | 20.0009 | 3.1794 | 0.008284 | 4.8224 | 4.0761 | 4.6842 | 3.6103
2012-10-30 14:04 | 22.7583 | 24.7914 | 24.0413 | 24.6611 | 23.6267 | 46.1053 | 0.014206 | 0.3456 | 0.00491 | 24.3277 | 0.014382 | 0.23454 | 0.003373 | 16.3077 | 19.9646 | 3.1785 | 0.008283 | 4.8268 | 4.0767 | 4.6965 | 3.6621
2012-10-30 14:04 | 22.7665 | 24.8075 | 24.0376 | 24.6772 | 23.627 | 46.1105 | 0.014184 | 0.34556 | 0.004901 | 24.3625 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 19.9569 | 3.1789 | 0.008277 | 4.8506 | 4.0807 | 4.7203 | 3.6701
2012-10-30 14:04 | 22.774 | 24.8071 | 23.9898 | 24.6413 | 23.6464 | 46.0983 | 0.014181 | 0.34557 | 0.004901 | 24.3686 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 19.9644 | 3.178 | 0.008276 | 4.8427 | 4.0254 | 4.6769 | 3.682
2012-10-30 14:04 | 22.7605 | 24.762 | 23.9566 | 24.6317 | 23.5737 | 46.1058 | 0.014175 | 0.3456 | 0.004899 | 24.381 | 0.014386 | 0.23452 | 0.003374 | 16.3022 | 19.8949 | 3.1785 | 0.008273 | 4.8671 | 4.0617 | 4.7368 | 3.6788
2012-10-30 14:04 | 22.774 | 24.6729 | 23.9977 | 24.5821 | 23.6227 | 46.1047 | 0.01419 | 0.34557 | 0.004904 | 24.3531 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 19.9723 | 3.1785 | 0.008281 | 4.7006 | 4.0254 | 4.6098 | 3.6504
2012-10-30 14:04 | 22.7785 | 24.78 | 24.0101 | 24.5708 | 23.564 | 46.1116 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 19.9529 | 3.1789 | 0.008281 | 4.8271 | 4.0572 | 4.6179 | 3.6111
2012-10-30 14:04 | 22.7732 | 24.7195 | 23.9417 | 24.5655 | 23.5745 | 46.0949 | 0.01419 | 0.34557 | 0.004904 | 24.3531 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 19.9316 | 3.1778 | 0.008278 | 4.7879 | 4.0101 | 4.6339 | 3.6429
2012-10-30 14:04 | 22.783 | 24.6977 | 23.9949 | 24.5555 | 23.5962 | 46.1151 | 0.014187 | 0.34567 | 0.004904 | 24.3654 | 0.014386 | 0.23451 | 0.003374 | 16.3011 | 19.9694 | 3.1792 | 0.008278 | 4.7283 | 4.0255 | 4.5861 | 3.6268
2012-10-30 14:04 | 22.7792 | 24.6702 | 23.92 | 24.5596 | 23.5805 | 46.0821 | 0.014203 | 0.34559 | 0.004908 | 24.3319 | 0.014379 | 0.23454 | 0.003372 | 16.3111 | 19.9416 | 3.1769 | 0.00828 | 4.7286 | 3.9784 | 4.618 | 3.6389
2012-10-30 14:04 | 22.7822 | 24.6692 | 23.9625 | 24.5192 | 23.5796 | 46.107 | 0.014194 | 0.34562 | 0.004906 | 24.3494 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.9206 | 3.1786 | 0.008283 | 4.7486 | 4.0419 | 4.5986 | 3.659
2012-10-30 14:04 | 22.7889 | 24.6443 | 23.91 | 24.5298 | 23.5468 | 46.0769 | 0.0142 | 0.34562 | 0.004908 | 24.3391 | 0.014386 | 0.23455 | 0.003374 | 16.3042 | 19.9633 | 3.1765 | 0.008282 | 4.681 | 3.9467 | 4.5665 | 3.5835
2012-10-30 14:04 | 22.7904 | 24.634 | 23.8997 | 24.5313 | 23.5286 | 46.1139 | 0.014181 | 0.34557 | 0.004901 | 24.3686 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 19.9688 | 3.1791 | 0.008278 | 4.6652 | 3.9309 | 4.5625 | 3.5598
2012-10-30 14:04 | 22.7904 | 24.634 | 23.8957 | 24.4958 | 23.5483 | 46.0804 | 0.014175 | 0.34556 | 0.004898 | 24.3779 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 19.9248 | 3.1768 | 0.008275 | 4.7092 | 3.9709 | 4.571 | 3.6235
2012-10-30 14:04 | 22.7957 | 24.5959 | 23.905 | 24.5169 | 23.5181 | 46.0885 | 0.01419 | 0.34557 | 0.004904 | 24.3531 | 0.014386 | 0.23448 | 0.003373 | 16.2991 | 19.9421 | 3.1773 | 0.008277 | 4.6538 | 3.9629 | 4.5748 | 3.576
2012-10-30 14:04 | 22.7875 | 24.564 | 23.9007 | 24.4692 | 23.5099 | 46.0694 | 0.014184 | 0.3456 | 0.004902 | 24.3655 | 0.014401 | 0.23461 | 0.003379 | 16.2912 | 19.9499 | 3.176 | 0.008281 | 4.6141 | 3.9508 | 4.5193 | 3.56
2012-10-30 14:04 | 22.8024 | 24.6223 | 23.9038 | 24.492 | 23.4892 | 46.0954 | 0.014197 | 0.34562 | 0.004907 | 24.3442 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9888 | 3.1778 | 0.008281 | 4.6335 | 3.915 | 4.5032 | 3.5004
2012-10-30 14:04 | 22.7994 | 24.5996 | 23.8731 | 24.4693 | 23.4626 | 46.0816 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9937 | 3.1769 | 0.008282 | 4.6059 | 3.8794 | 4.4756 | 3.4689
2012-10-30 14:04 | 22.7979 | 24.5704 | 23.8637 | 24.4322 | 23.4611 | 46.0844 | 0.014181 | 0.34557 | 0.004901 | 24.3686 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9523 | 3.1771 | 0.008279 | 4.6181 | 3.9114 | 4.4799 | 3.5088
2012-10-30 14:05 | 22.8129 | 24.5854 | 23.855 | 24.4078 | 23.4168 | 46.0868 | 0.014209 | 0.34564 | 0.004911 | 24.3256 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.0192 | 3.1772 | 0.008286 | 4.5662 | 3.8358 | 4.3886 | 3.3976
2012-10-30 14:05 | 22.8091 | 24.55 | 23.8315 | 24.404 | 23.4367 | 46.0677 | 0.014187 | 0.34559 | 0.004903 | 24.3593 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.0034 | 3.1759 | 0.008278 | 4.5466 | 3.8281 | 4.4006 | 3.4333
2012-10-30 14:05 | 22.8054 | 24.5424 | 23.8633 | 24.4042 | 23.4212 | 46.0914 | 0.01419 | 0.34557 | 0.004904 | 24.3531 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0117 | 3.1775 | 0.008281 | 4.5307 | 3.8516 | 4.3925 | 3.4095
2012-10-30 14:05 | 22.8069 | 24.5242 | 23.8056 | 24.4057 | 23.4187 | 46.0677 | 0.014187 | 0.3456 | 0.004903 | 24.3603 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 19.9933 | 3.1759 | 0.008281 | 4.5309 | 3.8123 | 4.4124 | 3.4254
2012-10-30 14:05 | 22.8046 | 24.4666 | 23.7362 | 24.3718 | 23.3849 | 46.0937 | 0.014181 | 0.34557 | 0.004901 | 24.3686 | 0.014382 | 0.23459 | 0.003374 | 16.3117 | 20.0309 | 3.1777 | 0.008275 | 4.4357 | 3.7053 | 4.3409 | 3.354
2012-10-30 14:05 | 22.8091 | 24.4711 | 23.7841 | 24.325 | 23.3696 | 46.0683 | 0.014194 | 0.34551 | 0.004904 | 24.3422 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 19.9955 | 3.1759 | 0.008281 | 4.4756 | 3.7886 | 4.3295 | 3.3741
2012-10-30 14:05 | 22.8061 | 24.4523 | 23.789 | 24.3536 | 23.3706 | 46.0787 | 0.014175 | 0.34554 | 0.004898 | 24.3769 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0004 | 3.1767 | 0.008275 | 4.4519 | 3.7886 | 4.3532 | 3.3702
2012-10-30 14:05 | 22.8121 | 24.4385 | 23.7911 | 24.3162 | 23.3687 | 46.0642 | 0.014194 | 0.3456 | 0.004905 | 24.3483 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 19.9665 | 3.1757 | 0.008281 | 4.472 | 3.8246 | 4.3497 | 3.4022
2012-10-30 14:05 | 22.8129 | 24.4393 | 23.6932 | 24.317 | 23.3537 | 46.0873 | 0.0142 | 0.34557 | 0.004907 | 24.336 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 19.9713 | 3.1773 | 0.008285 | 4.468 | 3.7219 | 4.3457 | 3.3824
2012-10-30 14:05 | 22.8166 | 24.4075 | 23.7443 | 24.2733 | 23.3534 | 46.0752 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 19.963 | 3.1764 | 0.008283 | 4.4445 | 3.7813 | 4.3103 | 3.3904
2012-10-30 14:05 | 22.8196 | 24.4026 | 23.7196 | 24.2921 | 23.3288 | 46.0891 | 0.01419 | 0.34562 | 0.004904 | 24.3562 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.978 | 3.1774 | 0.008281 | 4.4246 | 3.7416 | 4.3141 | 3.3508
2012-10-30 14:05 | 22.8375 | 24.381 | 23.7573 | 24.2981 | 23.327 | 46.0654 | 0.014178 | 0.34564 | 0.004901 | 24.3788 | 0.014376 | 0.23454 | 0.003372 | 16.3145 | 19.9999 | 3.1757 | 0.008273 | 4.3811 | 3.7574 | 4.2982 | 3.3271
2012-10-30 14:05 | 22.8263 | 24.3935 | 23.6947 | 24.2317 | 23.3039 | 46.0902 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.0526 | 3.1775 | 0.008288 | 4.3409 | 3.6421 | 4.1791 | 3.2513
2012-10-30 14:05 | 22.8278 | 24.3674 | 23.6528 | 24.2292 | 23.3054 | 46.0579 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 20.0062 | 3.1752 | 0.008282 | 4.3612 | 3.6466 | 4.223 | 3.2992
2012-10-30 14:05 | 22.8316 | 24.3317 | 23.6566 | 24.237 | 23.2697 | 46.0764 | 0.014184 | 0.34559 | 0.004902 | 24.3645 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.022 | 3.1765 | 0.00828 | 4.3097 | 3.6346 | 4.215 | 3.2477
2012-10-30 14:05 | 22.8286 | 24.3524 | 23.6457 | 24.2221 | 23.2904 | 46.0642 | 0.014184 | 0.34557 | 0.004902 | 24.3634 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 19.943 | 3.1757 | 0.008278 | 4.4094 | 3.7027 | 4.2791 | 3.3474
2012-10-30 14:05 | 22.8338 | 24.33 | 23.6983 | 24.1957 | 23.2759 | 46.0579 | 0.014175 | 0.34557 | 0.004898 | 24.3789 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9842 | 3.1752 | 0.008274 | 4.3458 | 3.7141 | 4.2115 | 3.2917
2012-10-30 14:05 | 22.8316 | 24.3238 | 23.6053 | 24.2172 | 23.2342 | 46.0706 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.998 | 3.1761 | 0.008284 | 4.3258 | 3.6073 | 4.2192 | 3.2362
2012-10-30 14:05 | 22.839 | 24.2917 | 23.6087 | 24.1338 | 23.2653 | 46.0527 | 0.01419 | 0.34559 | 0.004904 | 24.3542 | 0.014386 | 0.23452 | 0.003374 | 16.3022 | 20.0214 | 3.1749 | 0.008278 | 4.2703 | 3.5873 | 4.1124 | 3.2439
2012-10-30 14:05 | 22.8398 | 24.2846 | 23.5938 | 24.1386 | 23.2266 | 46.0688 | 0.014194 | 0.34556 | 0.004905 | 24.3453 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 19.9942 | 3.176 | 0.008282 | 4.2904 | 3.5996 | 4.1444 | 3.2324
2012-10-30 14:05 | 22.8338 | 24.2707 | 23.6035 | 24.097 | 23.2285 | 46.0411 | 0.014184 | 0.34557 | 0.004902 | 24.3634 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.9882 | 3.1741 | 0.008278 | 4.2825 | 3.6153 | 4.1088 | 3.2403
2012-10-30 14:05 | 22.8443 | 24.2654 | 23.6061 | 24.1194 | 23.1838 | 46.0521 | 0.014194 | 0.3456 | 0.004905 | 24.3483 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 19.9947 | 3.1748 | 0.00828 | 4.2707 | 3.6114 | 4.1247 | 3.1891
2012-10-30 14:05 | 22.839 | 24.2167 | 23.5495 | 24.0864 | 23.2298 | 46.0451 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 20.0014 | 3.1743 | 0.008282 | 4.2153 | 3.5481 | 4.085 | 3.2284
2012-10-30 14:05 | 22.848 | 24.2968 | 23.5664 | 24.0875 | 23.2309 | 46.0428 | 0.014187 | 0.34557 | 0.004903 | 24.3583 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9944 | 3.1742 | 0.008279 | 4.3024 | 3.572 | 4.0931 | 3.2365
2012-10-30 14:05 | 22.8502 | 24.1884 | 23.5726 | 24.0897 | 23.1976 | 46.0446 | 0.014194 | 0.34557 | 0.004905 | 24.3463 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.9846 | 3.1743 | 0.008282 | 4.2038 | 3.588 | 4.1051 | 3.213
2012-10-30 14:05 | 22.8443 | 24.2418 | 23.539 | 24.0483 | 23.1838 | 46.0371 | 0.014212 | 0.34553 | 0.004911 | 24.3124 | 0.014404 | 0.23445 | 0.003377 | 16.2767 | 19.9627 | 3.1738 | 0.008288 | 4.2791 | 3.5763 | 4.0856 | 3.2211
2012-10-30 14:05 | 22.854 | 24.2159 | 23.5566 | 24.0659 | 23.2132 | 46.0613 | 0.014209 | 0.3456 | 0.004911 | 24.3226 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9884 | 3.1755 | 0.008285 | 4.2275 | 3.5682 | 4.0775 | 3.2248
2012-10-30 14:05 | 22.8502 | 24.1253 | 23.4936 | 24.0384 | 23.1778 | 46.0446 | 0.014184 | 0.34556 | 0.004901 | 24.3625 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0006 | 3.1743 | 0.008278 | 4.1247 | 3.493 | 4.0378 | 3.1772
2012-10-30 14:06 | 22.8547 | 24.1811 | 23.4942 | 24.0034 | 23.1705 | 46.0313 | 0.014197 | 0.34562 | 0.004907 | 24.3442 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.0251 | 3.1734 | 0.008283 | 4.156 | 3.4691 | 3.9783 | 3.1454
2012-10-30 14:06 | 22.86 | 24.1588 | 23.5034 | 23.9693 | 23.1876 | 46.0428 | 0.01419 | 0.34559 | 0.004904 | 24.3542 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.0224 | 3.1742 | 0.008278 | 4.1364 | 3.481 | 3.9469 | 3.1652
2012-10-30 14:06 | 22.8622 | 24.1373 | 23.4819 | 23.9438 | 23.1464 | 46.0243 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 19.9846 | 3.1729 | 0.008282 | 4.1527 | 3.4973 | 3.9592 | 3.1618
2012-10-30 14:06 | 22.8585 | 24.1178 | 23.4901 | 24.0151 | 23.1111 | 46.0347 | 0.014206 | 0.34559 | 0.004909 | 24.3268 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 19.9849 | 3.1736 | 0.008284 | 4.1329 | 3.5052 | 4.0302 | 3.1262
2012-10-30 14:06 | 22.8652 | 24.0653 | 23.4573 | 23.9587 | 23.1139 | 46.0417 | 0.014194 | 0.34562 | 0.004906 | 24.3494 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 20.0156 | 3.1741 | 0.008282 | 4.0497 | 3.4417 | 3.9431 | 3.0983
2012-10-30 14:06 | 22.8689 | 24.0926 | 23.4728 | 23.9505 | 23.0544 | 46.0347 | 0.014203 | 0.34559 | 0.004908 | 24.3319 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.9953 | 3.1736 | 0.008284 | 4.0973 | 3.4775 | 3.9552 | 3.0591
2012-10-30 14:06 | 22.8644 | 24.0645 | 23.4683 | 23.8868 | 23.0697 | 46.0342 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9828 | 3.1736 | 0.008285 | 4.0817 | 3.4855 | 3.904 | 3.0869
2012-10-30 14:06 | 22.8637 | 24.0361 | 23.4676 | 23.9177 | 23.069 | 46.0261 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 19.9661 | 3.173 | 0.008282 | 4.07 | 3.5015 | 3.9516 | 3.1029
2012-10-30 14:06 | 22.8659 | 24.0502 | 23.4106 | 23.8686 | 23.0277 | 46.0434 | 0.014212 | 0.34559 | 0.004911 | 24.3165 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 19.9483 | 3.1742 | 0.008287 | 4.1019 | 3.4623 | 3.9203 | 3.0794
2012-10-30 14:06 | 22.8757 | 24.0442 | 23.4283 | 23.9218 | 23.0928 | 46.0249 | 0.014203 | 0.3456 | 0.004909 | 24.3329 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9661 | 3.173 | 0.008286 | 4.0781 | 3.4622 | 3.9557 | 3.1267
2012-10-30 14:06 | 22.8704 | 24.027 | 23.4151 | 23.8652 | 23.0559 | 46.0423 | 0.0142 | 0.3456 | 0.004908 | 24.338 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9648 | 3.1742 | 0.008284 | 4.0622 | 3.4503 | 3.9004 | 3.0911
2012-10-30 14:06 | 22.8772 | 23.9944 | 23.4259 | 23.8641 | 23.0232 | 46.0336 | 0.01419 | 0.34559 | 0.004904 | 24.3542 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9756 | 3.1736 | 0.008281 | 4.0188 | 3.4503 | 3.8885 | 3.0476
2012-10-30 14:06 | 22.8772 | 23.9825 | 23.3943 | 23.8562 | 23.0311 | 46.0128 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 19.9796 | 3.1721 | 0.008282 | 4.0029 | 3.4147 | 3.8766 | 3.0515
2012-10-30 14:06 | 22.8764 | 23.9304 | 23.3619 | 23.8317 | 22.9988 | 46.0261 | 0.014206 | 0.3456 | 0.00491 | 24.3277 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9668 | 3.173 | 0.008287 | 3.9636 | 3.3951 | 3.8649 | 3.032
2012-10-30 14:06 | 22.8854 | 23.9947 | 23.367 | 23.8209 | 23.0157 | 46.0145 | 0.014187 | 0.34559 | 0.004903 | 24.3593 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9558 | 3.1722 | 0.008279 | 4.0389 | 3.4112 | 3.8651 | 3.0599
2012-10-30 14:06 | 22.8816 | 23.9198 | 23.3355 | 23.8093 | 22.9605 | 46.0001 | 0.014203 | 0.34559 | 0.004908 | 24.3319 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 19.972 | 3.1712 | 0.008283 | 3.9478 | 3.3635 | 3.8373 | 2.9885
2012-10-30 14:06 | 22.8854 | 23.9552 | 23.3354 | 23.8052 | 22.9446 | 46.0255 | 0.014187 | 0.34559 | 0.004903 | 24.3593 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 19.9678 | 3.173 | 0.008278 | 3.9874 | 3.3676 | 3.8374 | 2.9768
2012-10-30 14:06 | 22.8831 | 23.945 | 23.3134 | 23.7871 |  | 46.0191 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0295 | 3.1726 | 0.008282 | 3.9155 | 3.2839 | 3.7576 | -20.0295
2012-10-30 14:06 | 22.8861 | 23.9243 | 23.3124 | 23.794 | 22.965 | 46.0168 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9286 | 3.1724 | 0.008282 | 3.9957 | 3.3838 | 3.8654 | 3.0364
2012-10-30 14:06 | 22.8973 | 23.9434 | 23.3078 | 23.7855 |  | 46.0099 | 0.0142 | 0.3456 | 0.004908 | 24.338 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9478 | 3.1719 | 0.008285 | 3.9956 | 3.36 | 3.8377 | -19.9478
2012-10-30 14:06 | 22.8944 | 23.8497 | 23.2773 | 23.7707 |  | 46.0215 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014414 | 0.23461 | 0.003382 | 16.2765 | 19.9808 | 3.1727 | 0.008286 | 3.8689 | 3.2965 | 3.7899 | -19.9808
2012-10-30 14:06 | 22.8936 | 23.8647 | 23.2252 | 23.7265 | 22.941 | 46.0134 | 0.014209 | 0.34564 | 0.004911 | 24.3256 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 19.948 | 3.1722 | 0.008288 | 3.9167 | 3.2772 | 3.7785 | 2.993
2012-10-30 14:06 | 22.8973 | 23.8723 | 23.217 | 23.7026 |  | 46.0139 | 0.014216 | 0.34559 | 0.004913 | 24.3097 | 0.014401 | 0.23461 | 0.003379 | 16.2912 | 19.9717 | 3.1722 | 0.008292 | 3.9006 | 3.2453 | 3.7309 | -19.9717
2012-10-30 14:06 | 22.9011 | 23.8761 | 23.2208 | 23.6985 |  | 46.0116 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 19.9476 | 3.172 | 0.008286 | 3.9285 | 3.2732 | 3.7509 | -19.9476
2012-10-30 14:06 | 22.9026 | 23.8934 | 23.2065 | 23.6842 |  | 45.9978 | 0.014178 | 0.34559 | 0.0049 | 24.3748 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.013 | 3.1711 | 0.008278 | 3.8804 | 3.1935 | 3.6712 | -20.013
2012-10-30 14:06 | 22.9056 | 23.8254 | 23.2253 | 23.6832 |  | 46.022 | 0.014197 | 0.34563 | 0.004907 | 24.3452 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 19.984 | 3.1728 | 0.008284 | 3.8414 | 3.2413 | 3.6992 | -19.984
2012-10-30 14:06 | 22.9033 | 23.8231 | 23.2072 | 23.677 |  | 46.0082 | 0.014197 | 0.3456 | 0.004906 | 24.3432 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 19.9577 | 3.1718 | 0.008283 | 3.8654 | 3.2495 | 3.7193 | -19.9577
2012-10-30 14:06 | 22.9108 | 23.7753 | 23.2345 | 23.6292 |  | 45.9897 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 19.9333 | 3.1705 | 0.008282 | 3.842 | 3.3012 | 3.6959 | -19.9333
2012-10-30 14:07 | 22.9071 | 23.7992 | 23.1952 | 23.6018 |  | 46.0116 | 0.014212 | 0.34557 | 0.004911 | 24.3154 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9695 | 3.172 | 0.008288 | 3.8297 | 3.2257 | 3.6323 | -19.9695
2012-10-30 14:07 | 22.9086 | 23.7731 | 23.1731 | 23.6428 | 22.8413 | 45.9816 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 19.9431 | 3.17 | 0.008283 | 3.83 | 3.23 | 3.6997 | 2.8982
2012-10-30 14:07 | 22.9108 | 23.7358 | 23.1555 | 23.6213 | 22.8038 | 46.022 | 0.014187 | 0.34559 | 0.004903 | 24.3593 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 19.9772 | 3.1728 | 0.008277 | 3.7586 | 3.1783 | 3.6441 | 2.8266
2012-10-30 14:07 | 22.919 | 23.7401 | 23.2032 | 23.5979 | 22.7764 | 45.9995 | 0.014197 | 0.34562 | 0.004907 | 24.3442 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9734 | 3.1712 | 0.008284 | 3.7667 | 3.2298 | 3.6245 | 2.803
2012-10-30 14:07 | 22.9123 | 23.7136 | 23.157 | 23.5834 | 22.7697 | 45.9752 | 0.014197 | 0.3456 | 0.004906 | 24.3432 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 19.9667 | 3.1695 | 0.008283 | 3.7469 | 3.1903 | 3.6167 | 2.803
2012-10-30 14:07 | 22.9168 | 23.7458 | 23.1023 | 23.5879 | 22.8019 | 46.0139 | 0.014194 | 0.34556 | 0.004905 | 24.3453 | 0.014386 | 0.23452 | 0.003374 | 16.3022 | 19.9752 | 3.1722 | 0.008279 | 3.7706 | 3.1271 | 3.6127 | 2.8267
2012-10-30 14:07 | 22.9153 | 23.7324 | 23.1284 | 23.5548 |  | 46.0001 | 0.01419 | 0.34553 | 0.004903 | 24.3501 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 19.9777 | 3.1712 | 0.008283 | 3.7547 | 3.1507 | 3.5771 | -19.9777
2012-10-30 14:07 | 22.913 | 23.6827 | 23.1183 | 23.5525 | 22.7505 | 45.9868 | 0.01419 | 0.34556 | 0.004903 | 24.3521 | 0.014395 | 0.23461 | 0.003377 | 16.298 | 19.9754 | 3.1703 | 0.00828 | 3.7073 | 3.1429 | 3.5771 | 2.7751
2012-10-30 14:07 | 22.9258 | 23.6245 | 23.135 | 23.5179 | 22.7752 | 46.0012 | 0.014184 | 0.34556 | 0.004901 | 24.3625 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 19.9643 | 3.1713 | 0.008279 | 3.6602 | 3.1707 | 3.5536 | 2.8109
2012-10-30 14:07 | 22.9213 | 23.6397 | 23.0989 | 23.5095 | 22.743 | 45.9926 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9438 | 3.1707 | 0.008285 | 3.6959 | 3.1551 | 3.5657 | 2.7992
2012-10-30 14:07 | 22.9258 | 23.6363 | 23.0561 | 23.5179 | 22.7158 | 45.9839 | 0.0142 | 0.34556 | 0.004907 | 24.335 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9763 | 3.1701 | 0.008283 | 3.66 | 3.0798 | 3.5416 | 2.7395
2012-10-30 14:07 | 22.9317 | 23.6462 | 23.0935 | 23.4843 | 22.7058 | 45.9856 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 19.9702 | 3.1702 | 0.008281 | 3.676 | 3.1233 | 3.5141 | 2.7356
2012-10-30 14:07 | 22.9362 | 23.623 | 23.0822 | 23.4849 | 22.7064 | 45.9891 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 19.9507 | 3.1705 | 0.008283 | 3.6723 | 3.1315 | 3.5342 | 2.7557
2012-10-30 14:07 | 22.9325 | 23.5878 | 23.0193 | 23.4654 | 22.7145 | 45.9827 | 0.014194 | 0.34557 | 0.004905 | 24.3463 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.955 | 3.17 | 0.008281 | 3.6328 | 3.0643 | 3.5104 | 2.7595
2012-10-30 14:07 | 22.9422 | 23.6054 | 23.0685 | 23.4475 | 22.6846 | 45.9845 | 0.014181 | 0.34557 | 0.004901 | 24.3686 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 19.9527 | 3.1702 | 0.008275 | 3.6527 | 3.1158 | 3.4948 | 2.7319
2012-10-30 14:07 | 22.9429 | 23.614 |  | 23.4008 | 22.6694 | 45.9995 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 19.9294 | 3.1712 | 0.008282 | 3.6846 | -19.9294 | 3.4714 | 2.74
2012-10-30 14:07 | 22.9474 | 23.5395 | 22.9987 | 23.4369 | 22.67 | 45.9758 | 0.014209 | 0.34556 | 0.00491 | 24.3196 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 19.9259 | 3.1696 | 0.008286 | 3.6136 | 3.0728 | 3.511 | 2.7441
2012-10-30 14:07 | 22.9422 | 23.5461 | 22.9896 | 23.4119 | 22.7084 | 45.977 | 0.014203 | 0.3456 | 0.004909 | 24.3329 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 19.9487 | 3.1697 | 0.008285 | 3.5974 | 3.0409 | 3.4632 | 2.7597
2012-10-30 14:07 | 22.9497 | 23.5694 |  | 23.4036 | 22.6287 | 46.0041 | 0.014203 | 0.34564 | 0.004909 | 24.3359 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9442 | 3.1715 | 0.008287 | 3.6252 | -19.9442 | 3.4594 | 2.6845
2012-10-30 14:07 | 22.9497 | 23.5457 |  | 23.3997 | 22.6247 | 45.9885 | 0.014203 | 0.34562 | 0.004909 | 24.3339 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9242 | 3.1704 | 0.008286 | 3.6215 | -19.9242 | 3.4755 | 2.7005
2012-10-30 14:07 | 22.9444 | 23.5562 |  | 23.3786 | 22.5916 | 45.9666 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 19.9429 | 3.1689 | 0.008285 | 3.6133 | -19.9429 | 3.4357 | 2.6487
2012-10-30 14:07 | 22.9527 | 23.5093 |  | 23.3711 | 22.592 | 45.9735 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014379 | 0.23448 | 0.003372 | 16.307 | 19.9672 | 3.1694 | 0.008279 | 3.5421 | -19.9672 | 3.4039 | 2.6248
2012-10-30 14:07 | 22.9512 | 23.4999 |  | 23.3657 | 22.6103 | 45.9874 | 0.01419 | 0.34563 | 0.004904 | 24.3572 | 0.014389 | 0.23448 | 0.003374 | 16.2957 | 19.9537 | 3.1704 | 0.008278 | 3.5462 | -19.9537 | 3.412 | 2.6566
2012-10-30 14:07 | 22.9497 | 23.4707 |  | 23.3365 | 22.5771 | 45.966 | 0.014203 | 0.34559 | 0.004908 | 24.3319 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 19.9482 | 3.1689 | 0.008285 | 3.5225 | -19.9482 | 3.3883 | 2.6289
2012-10-30 14:07 | 22.9594 | 23.4962 |  | 23.3304 | 22.5947 | 45.9596 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014382 | 0.23448 | 0.003372 | 16.3036 | 19.9539 | 3.1685 | 0.008277 | 3.5423 | -19.9539 | 3.3765 | 2.6408
2012-10-30 14:07 | 22.9579 | 23.4632 |  | 23.3171 | 22.5536 | 45.9723 | 0.01419 | 0.34554 | 0.004903 | 24.3511 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 19.9564 | 3.1693 | 0.008281 | 3.5068 | -19.9564 | 3.3607 | 2.5972
2012-10-30 14:08 | 22.9684 | 23.4263 | 22.8773 | 23.3276 | 22.5244 | 45.9613 | 0.0142 | 0.34562 | 0.004908 | 24.3391 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 19.9589 | 3.1686 | 0.008288 | 3.4674 | 2.9184 | 3.3687 | 2.5655
2012-10-30 14:08 | 22.9616 | 23.455 | 22.8705 | 23.3445 | 22.5097 | 45.9602 | 0.014197 | 0.3456 | 0.004906 | 24.3432 | 0.014386 | 0.23445 | 0.003373 | 16.2971 | 19.9201 | 3.1685 | 0.008279 | 3.5349 | 2.9504 | 3.4244 | 2.5896
2012-10-30 14:08 | 22.9669 | 23.4327 | 22.848 | 23.3024 | 22.511 | 45.9758 | 0.014194 | 0.34551 | 0.004904 | 24.3422 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0054 | 3.1696 | 0.008279 | 3.4273 | 2.8426 | 3.297 | 2.5056
2012-10-30 14:08 | 22.9676 | 23.4215 | 22.8448 | 23.24 | 22.4839 | 45.9683 | 0.0142 | 0.34557 | 0.004907 | 24.336 | 0.014404 | 0.23458 | 0.003379 | 16.2858 | 19.9701 | 3.1691 | 0.008286 | 3.4514 | 2.8747 | 3.2699 | 2.5138
2012-10-30 14:08 | 22.9616 | 23.3919 | 22.7991 | 23.2379 | 22.5295 | 45.9648 | 0.014184 | 0.34559 | 0.004902 | 24.3645 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 19.9801 | 3.1688 | 0.008279 | 3.4118 | 2.819 | 3.2578 | 2.5494
2012-10-30 14:08 | 22.9684 | 23.4144 | 22.8218 | 23.2526 | 22.5323 | 45.9625 | 0.01419 | 0.3455 | 0.004903 | 24.3481 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 19.9869 | 3.1687 | 0.008279 | 3.4275 | 2.8349 | 3.2657 | 2.5454
2012-10-30 14:08 | 22.9691 | 23.3322 | 22.8225 | 23.2257 | 22.541 | 45.9619 | 0.014187 | 0.34562 | 0.004903 | 24.3614 | 0.014386 | 0.23457 | 0.003374 | 16.3052 | 19.9676 | 3.1686 | 0.008277 | 3.3646 | 2.8549 | 3.2581 | 2.5734
2012-10-30 14:08 | 22.9699 | 23.3291 | 22.8391 | 23.2383 | 22.518 | 45.9729 | 0.01419 | 0.34553 | 0.004903 | 24.3501 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 19.9404 | 3.1694 | 0.008281 | 3.3887 | 2.8987 | 3.2979 | 2.5776
2012-10-30 14:08 | 22.9766 | 23.3634 | 22.7943 | 23.1819 | 22.4969 | 45.955 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9911 | 3.1681 | 0.008281 | 3.3723 | 2.8032 | 3.1908 | 2.5058
2012-10-30 14:08 | 22.9766 | 23.3161 | 22.7864 | 23.1819 | 22.4969 | 45.9613 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 19.9791 | 3.1686 | 0.008283 | 3.337 | 2.8073 | 3.2028 | 2.5178
2012-10-30 14:08 | 22.9781 | 23.3057 | 22.8038 | 23.1794 | 22.4904 | 45.9648 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 19.9806 | 3.1688 | 0.008282 | 3.3251 | 2.8232 | 3.1988 | 2.5098
2012-10-30 14:08 | 22.9773 | 23.2773 | 22.7871 | 23.1904 | 22.4817 | 45.9475 | 0.01419 | 0.34553 | 0.004903 | 24.3501 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 19.9958 | 3.1676 | 0.008283 | 3.2815 | 2.7913 | 3.1946 | 2.4859
2012-10-30 14:08 | 22.9751 | 23.2869 | 22.769 | 23.1606 | 22.4597 | 45.9619 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 19.9576 | 3.1686 | 0.008281 | 3.3293 | 2.8114 | 3.203 | 2.5021
2012-10-30 14:08 | 22.9833 | 23.307 | 22.7574 | 23.1057 | 22.4441 | 45.9394 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 19.9938 | 3.1671 | 0.008286 | 3.3132 | 2.7636 | 3.1119 | 2.4503
2012-10-30 14:08 | 22.9863 | 23.2744 | 22.7763 | 23.0889 | 22.4748 | 45.9654 | 0.014184 | 0.34554 | 0.004901 | 24.3614 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 19.9968 | 3.1689 | 0.008276 | 3.2776 | 2.7795 | 3.0921 | 2.478
2012-10-30 14:08 | 22.9878 | 23.2878 | 22.754 | 23.0904 | 22.4605 | 45.9457 | 0.014194 | 0.34556 | 0.004905 | 24.3453 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 19.9783 | 3.1675 | 0.008284 | 3.3095 | 2.7757 | 3.1121 | 2.4822
2012-10-30 14:08 | 22.9848 | 23.2532 | 22.7232 | 23.1545 | 22.4138 | 45.9423 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.0033 | 3.1673 | 0.00828 | 3.2499 | 2.7199 | 3.1512 | 2.4105
2012-10-30 14:08 | 22.9811 | 23.2377 | 22.7235 | 23.0956 | 22.4022 | 45.944 | 0.014194 | 0.34556 | 0.004905 | 24.3453 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 19.9956 | 3.1674 | 0.008284 | 3.2421 | 2.7279 | 3.1 | 2.4066
2012-10-30 14:08 | 22.9953 | 23.2519 | 22.7694 | 23.0782 | 22.4362 | 45.9457 | 0.014181 | 0.34559 | 0.004901 | 24.3697 | 0.014382 | 0.23452 | 0.003373 | 16.3067 | 19.9938 | 3.1675 | 0.008274 | 3.2581 | 2.7756 | 3.0844 | 2.4424
2012-10-30 14:08 | 22.993 | 23.2219 | 22.7473 | 23.1035 | 22.422 | 45.9486 | 0.0142 | 0.34559 | 0.004907 | 24.337 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 19.9555 | 3.1677 | 0.008283 | 3.2664 | 2.7918 | 3.148 | 2.4665
2012-10-30 14:08 | 22.996 | 23.221 | 22.7265 |  | 22.3694 | 45.9382 | 0.014187 | 0.3456 | 0.004903 | 24.3603 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 19.9785 | 3.167 | 0.008281 | 3.2425 | 2.748 | -19.9785 | 2.3909
2012-10-30 14:08 | 22.9953 | 23.1729 | 22.7258 | 23.0466 | 22.3925 | 45.9527 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 19.9178 | 3.168 | 0.008282 | 3.2551 | 2.808 | 3.1288 | 2.4747
2012-10-30 14:08 | 22.9945 | 23.1682 | 22.6893 |  | 22.3362 | 45.9452 | 0.014184 | 0.34556 | 0.004901 | 24.3625 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 19.989 | 3.1675 | 0.00828 | 3.1792 | 2.7003 | -19.989 | 2.3472
2012-10-30 14:08 | 23.0005 | 23.1663 | 22.6993 |  | 22.366 | 45.9457 | 0.014187 | 0.34562 | 0.004903 | 24.3614 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 19.951 | 3.1675 | 0.008278 | 3.2153 | 2.7483 | -19.951 | 2.415
2012-10-30 14:08 | 22.9945 | 23.1405 | 22.6655 |  | 22.364 | 45.9296 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 19.973 | 3.1664 | 0.008281 | 3.1675 | 2.6925 | -19.973 | 2.391
2012-10-30 14:08 | 23.0087 | 23.139 | 22.6678 |  | 22.3623 | 45.9377 | 0.014187 | 0.34557 | 0.004903 | 24.3583 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 19.9792 | 3.1669 | 0.008279 | 3.1598 | 2.6886 | -19.9792 | 2.3831
2012-10-30 14:09 | 22.9983 | 23.1325 | 22.6574 |  | 22.336 | 45.9481 | 0.014187 | 0.3455 | 0.004902 | 24.3532 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 19.9808 | 3.1677 | 0.008279 | 3.1517 | 2.6766 | -19.9808 | 2.3552
2012-10-30 14:09 | 22.9998 | 23.0827 | 22.5796 |  | 22.266 | 45.9509 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 19.9943 | 3.1679 | 0.008283 | 3.0884 | 2.5853 | -19.9943 | 2.2717
2012-10-30 14:09 | 23.0057 | 23.0965 | 22.6291 |  | 22.3315 | 45.9365 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9602 | 3.1669 | 0.00828 | 3.1363 | 2.6689 | -19.9602 | 2.3713
2012-10-30 14:09 | 23.0035 | 23.0903 | 22.5436 |  | 22.2697 | 45.9319 | 0.01419 | 0.34551 | 0.004903 | 24.349 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 19.93 | 3.1665 | 0.008279 | 3.1603 | 2.6136 | -19.93 | 2.3397
2012-10-30 14:09 | 23.0147 |  | 22.5905 |  | 22.2452 | 45.9273 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 19.9652 | 3.1662 | 0.008283 | -19.9652 | 2.6253 | -19.9652 | 2.28
2012-10-30 14:09 | 23.0229 |  | 22.5551 | 22.8644 | 22.2136 | 45.9203 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 19.9854 | 3.1657 | 0.008281 | -19.9854 | 2.5697 | 2.879 | 2.2282
2012-10-30 14:09 | 23.0274 |  | 22.5556 | 22.8848 | 22.1705 | 45.9359 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0219 | 3.1668 | 0.008283 | -20.0219 | 2.5337 | 2.8629 | 2.1486
2012-10-30 14:09 | 23.0252 |  | 22.4899 | 22.839 | 22.2398 | 45.9192 | 0.014194 | 0.3456 | 0.004905 | 24.3483 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9797 | 3.1657 | 0.008282 | -19.9797 | 2.5102 | 2.8593 | 2.2601
2012-10-30 14:09 | 23.0267 |  | 22.547 | 22.8008 | 22.1857 | 45.9053 | 0.014203 | 0.34556 | 0.004908 | 24.3299 | 0.014389 | 0.23457 | 0.003375 | 16.3018 | 20.0212 | 3.1647 | 0.008283 | -20.0212 | 2.5258 | 2.7796 | 2.1645
2012-10-30 14:09 | 23.0356 | 22.9484 | 22.5162 | 22.8137 | 22.2144 | 45.9076 | 0.014187 | 0.34557 | 0.004903 | 24.3583 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0061 | 3.1649 | 0.00828 | 2.9423 | 2.5101 | 2.8076 | 2.2083
2012-10-30 14:09 | 23.0297 | 22.9267 | 22.4746 | 22.8078 | 22.1529 | 45.9168 | 0.0142 | 0.34553 | 0.004906 | 24.333 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 19.9882 | 3.1655 | 0.008284 | 2.9385 | 2.4864 | 2.8196 | 2.1647
2012-10-30 14:09 | 23.0371 | 22.9182 | 22.474 | 22.8429 | 22.1086 | 45.9197 | 0.014187 | 0.3456 | 0.004903 | 24.3603 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0476 | 3.1657 | 0.008279 | 2.8706 | 2.4264 | 2.7953 | 2.061
2012-10-30 14:09 | 23.0334 | 22.8908 | 22.51 | 22.7837 | 22.1765 | 45.9168 | 0.014184 | 0.34554 | 0.004901 | 24.3614 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.0399 | 3.1655 | 0.008278 | 2.8509 | 2.4701 | 2.7438 | 2.1366
2012-10-30 14:09 | 23.0356 | 22.9167 | 22.4606 | 22.7939 | 22.1469 | 45.9163 | 0.014206 | 0.34556 | 0.004909 | 24.3247 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.0461 | 3.1655 | 0.008284 | 2.8706 | 2.4145 | 2.7478 | 2.1008
2012-10-30 14:09 | 23.0356 | 22.9009 | 22.4844 | 22.7899 | 22.131 | 45.9163 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.0301 | 3.1655 | 0.008279 | 2.8708 | 2.4543 | 2.7598 | 2.1009
2012-10-30 14:09 | 23.0342 | 22.8916 | 22.4394 | 22.7806 | 22.1495 | 45.9267 | 0.01419 | 0.34551 | 0.004903 | 24.349 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0327 | 3.1662 | 0.00828 | 2.8589 | 2.4067 | 2.7479 | 2.1168
2012-10-30 14:09 | 23.0409 | 22.9062 | 22.4342 | 22.7555 | 22.1363 | 45.9163 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0314 | 3.1655 | 0.00828 | 2.8748 | 2.4028 | 2.7241 | 2.1049
2012-10-30 14:09 | 23.0461 | 22.8718 | 22.4235 | 22.7013 | 22.1256 | 45.9197 | 0.0142 | 0.34562 | 0.004908 | 24.3391 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.0246 | 3.1657 | 0.008287 | 2.8472 | 2.3989 | 2.6767 | 2.101
2012-10-30 14:09 | 23.0499 | 22.8597 | 22.4551 | 22.7368 | 22.1413 | 45.9238 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0084 | 3.166 | 0.008283 | 2.8513 | 2.4467 | 2.7284 | 2.1329
2012-10-30 14:10 | 23.0469 | 22.8091 | 22.4521 | 22.6981 | 22.1145 | 45.907 | 0.014194 | 0.34557 | 0.004905 | 24.3463 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.0614 | 3.1648 | 0.008284 | 2.7477 | 2.3907 | 2.6367 | 2.0531
2012-10-30 14:10 | 23.0506 | 22.8287 | 22.428 | 22.7177 | 22.1301 | 45.9082 | 0.01419 | 0.34554 | 0.004903 | 24.3511 | 0.014382 | 0.23454 | 0.003373 | 16.3077 | 20.0211 | 3.1649 | 0.008276 | 2.8076 | 2.4069 | 2.6966 | 2.109
2012-10-30 14:10 | 23.0409 | 22.7992 | 22.4024 | 22.6881 | 22.0727 | 45.8984 | 0.014194 | 0.34557 | 0.004905 | 24.3463 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.0554 | 3.1642 | 0.00828 | 2.7438 | 2.347 | 2.6327 | 2.0173
2012-10-30 14:10 | 23.0469 | 22.8012 | 22.3965 | 22.6743 | 22.0588 | 45.9036 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0294 | 3.1646 | 0.008284 | 2.7718 | 2.3671 | 2.6449 | 2.0294
2012-10-30 14:10 | 23.0521 | 22.7231 | 22.3898 | 22.6835 | 22.0799 | 45.8787 | 0.014187 | 0.34554 | 0.004902 | 24.3562 | 0.014386 | 0.23458 | 0.003375 | 16.3061 | 20.0426 | 3.1629 | 0.008277 | 2.6805 | 2.3472 | 2.6409 | 2.0373
2012-10-30 14:10 | 23.0536 | 22.7881 | 22.3794 | 22.685 | 22.0695 | 45.8764 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.0321 | 3.1627 | 0.008285 | 2.756 | 2.3473 | 2.6529 | 2.0374
2012-10-30 14:10 | 23.0626 | 22.7852 | 22.3924 | 22.6583 | 22.0944 | 45.8741 | 0.014203 | 0.34559 | 0.004908 | 24.3319 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.0211 | 3.1626 | 0.008284 | 2.7641 | 2.3713 | 2.6372 | 2.0733
2012-10-30 14:10 | 23.0611 | 22.7638 | 22.4028 | 22.6488 | 22.0293 | 45.866 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0556 | 3.162 | 0.008282 | 2.7082 | 2.3472 | 2.5932 | 1.9737
2012-10-30 14:10 | 23.0656 | 22.7802 | 22.3556 | 22.6613 | 22.0457 | 45.8891 | 0.01419 | 0.34553 | 0.004903 | 24.3501 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.0721 | 3.1636 | 0.008277 | 2.7081 | 2.2835 | 2.5892 | 1.9736
2012-10-30 14:10 | 23.0626 | 22.7059 | 22.3169 | 22.6067 | 22.0268 | 45.8752 | 0.0142 | 0.34554 | 0.004907 | 24.3339 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0771 | 3.1626 | 0.008283 | 2.6288 | 2.2398 | 2.5296 | 1.9497
2012-10-30 14:10 | 23.0738 | 22.725 | 22.3559 | 22.6298 | 22.0221 | 45.8671 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0523 | 3.1621 | 0.008281 | 2.6727 | 2.3036 | 2.5775 | 1.9698
2012-10-30 14:10 | 23.0626 | 22.7019 | 22.3169 | 22.6107 | 22.0229 | 45.888 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0571 | 3.1635 | 0.008281 | 2.6448 | 2.2598 | 2.5536 | 1.9658
2012-10-30 14:10 | 23.076 | 22.6915 | 22.3621 | 22.5923 | 22.0164 | 45.8614 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0505 | 3.1617 | 0.008277 | 2.641 | 2.3116 | 2.5418 | 1.9659
2012-10-30 14:10 | 23.0663 | 22.662 | 22.3365 | 22.5628 | 22.0226 | 45.8677 | 0.014187 | 0.34557 | 0.004903 | 24.3583 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 20.0408 | 3.1621 | 0.008279 | 2.6212 | 2.2957 | 2.522 | 1.9818
2012-10-30 14:10 | 23.0693 | 22.6808 | 22.3157 | 22.5896 | 22.0097 | 45.8654 | 0.01419 | 0.34551 | 0.004903 | 24.349 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 20.0438 | 3.162 | 0.008281 | 2.637 | 2.2719 | 2.5458 | 1.9659
2012-10-30 14:10 | 23.0745 | 22.694 | 22.305 | 22.5749 | 21.999 | 45.8683 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014386 | 0.23445 | 0.003373 | 16.2971 | 20.049 | 3.1622 | 0.008278 | 2.645 | 2.256 | 2.5259 | 1.95
2012-10-30 14:10 | 23.0685 | 22.6642 | 22.2791 | 22.5531 | 21.985 | 45.8591 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.035 | 3.1615 | 0.008282 | 2.6292 | 2.2441 | 2.5181 | 1.95
2012-10-30 14:10 | 23.0775 | 22.6732 | 22.3 | 22.4946 | 21.9741 | 45.8562 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.012 | 3.1613 | 0.008282 | 2.6612 | 2.288 | 2.4826 | 1.9621
2012-10-30 14:10 | 23.085 | 22.6807 | 22.2599 | 22.5775 | 21.9498 | 45.8619 | 0.014197 | 0.34559 | 0.004906 | 24.3422 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.0435 | 3.1617 | 0.008286 | 2.6372 | 2.2164 | 2.534 | 1.9063
2012-10-30 14:10 | 23.0842 | 22.664 | 22.2551 | 22.545 | 21.9411 | 45.8556 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.0867 | 3.1613 | 0.008282 | 2.5773 | 2.1684 | 2.4583 | 1.8544
2012-10-30 14:10 | 23.0887 | 22.6368 | 22.2596 | 22.5256 | 21.9416 | 45.8446 | 0.014206 | 0.34556 | 0.004909 | 24.3247 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.0872 | 3.1605 | 0.008287 | 2.5496 | 2.1724 | 2.4384 | 1.8544
2012-10-30 14:10 | 23.0887 | 22.6129 | 22.3112 | 22.5018 | 21.9138 | 45.8654 | 0.0142 | 0.3456 | 0.004908 | 24.338 | 0.014382 | 0.23449 | 0.003372 | 16.3047 | 20.0992 | 3.162 | 0.00828 | 2.5137 | 2.212 | 2.4026 | 1.8146
2012-10-30 14:10 | 23.0887 | 22.6169 | 22.204 | 22.4621 | 21.9257 | 45.8585 | 0.014194 | 0.34557 | 0.004905 | 24.3463 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.0192 | 3.1615 | 0.008281 | 2.5977 | 2.1848 | 2.4429 | 1.9065
2012-10-30 14:10 | 23.0932 | 22.5857 | 22.2204 | 22.4825 | 21.9103 | 45.8446 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014386 | 0.23448 | 0.003373 | 16.2991 | 20.0677 | 3.1605 | 0.008278 | 2.518 | 2.1527 | 2.4148 | 1.8426
2012-10-30 14:10 | 23.0962 | 22.6165 | 22.2671 | 22.5212 | 21.9213 | 45.8533 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.0467 | 3.1611 | 0.008279 | 2.5698 | 2.2204 | 2.4745 | 1.8746
2012-10-30 14:10 | 23.0932 | 22.5738 | 22.2283 | 22.4587 | 21.9381 | 45.8585 | 0.0142 | 0.34556 | 0.004907 | 24.335 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.0517 | 3.1615 | 0.008284 | 2.5221 | 2.1766 | 2.407 | 1.8864
2012-10-30 14:11 | 23.0962 | 22.5173 | 22.2154 | 22.4498 | 21.8576 | 45.8481 | 0.014203 | 0.34556 | 0.004908 | 24.3299 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 20.0387 | 3.1608 | 0.008282 | 2.4786 | 2.1767 | 2.4111 | 1.8189
2012-10-30 14:11 | 23.0977 | 22.5426 | 22.2408 | 22.4513 | 21.9108 | 45.8487 | 0.014187 | 0.34551 | 0.004902 | 24.3542 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.0522 | 3.1608 | 0.008277 | 2.4904 | 2.1886 | 2.3991 | 1.8586
2012-10-30 14:11 | 23.1037 | 22.5645 | 22.1911 | 22.4811 | 21.9407 | 45.8458 | 0.014194 | 0.34551 | 0.004904 | 24.3422 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 20.0822 | 3.1606 | 0.008279 | 2.4823 | 2.1089 | 2.3989 | 1.8585
2012-10-30 14:11 | 23.1022 | 22.5272 | 22.1936 | 22.4121 | 21.8278 | 45.8521 | 0.014187 | 0.3455 | 0.004902 | 24.3532 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.0447 | 3.161 | 0.008277 | 2.4825 | 2.1489 | 2.3674 | 1.7831
2012-10-30 14:11 | 23.1082 | 22.5332 | 22.1519 | 22.4419 | 21.8617 | 45.8313 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0627 | 3.1596 | 0.008285 | 2.4705 | 2.0892 | 2.3792 | 1.799
2012-10-30 14:11 | 23.1059 | 22.5031 | 22.1496 | 22.4158 | 21.8594 | 45.8533 | 0.014209 | 0.34557 | 0.00491 | 24.3206 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0164 | 3.1611 | 0.008287 | 2.4867 | 2.1332 | 2.3994 | 1.843
2012-10-30 14:11 | 23.1029 | 22.5319 | 22.1466 | 22.4088 | 21.8564 | 45.8527 | 0.01419 | 0.34551 | 0.004903 | 24.349 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.0014 | 3.1611 | 0.008278 | 2.5305 | 2.1452 | 2.4074 | 1.855
2012-10-30 14:11 | 23.1127 | 22.506 | 22.1962 | 22.363 | 21.8304 | 45.8371 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0272 | 3.16 | 0.008285 | 2.4788 | 2.169 | 2.3358 | 1.8032
2012-10-30 14:11 | 23.1186 | 22.4761 | 22.1703 | 22.3769 | 21.8641 | 45.8319 | 0.014203 | 0.34554 | 0.004908 | 24.3288 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.0091 | 3.1597 | 0.008287 | 2.467 | 2.1612 | 2.3678 | 1.855
2012-10-30 14:11 | 23.1201 | 22.4499 | 22.1559 | 22.3744 | 21.8417 | 45.8417 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.0226 | 3.1603 | 0.008285 | 2.4273 | 2.1333 | 2.3518 | 1.8191
2012-10-30 14:11 | 23.1209 | 22.4507 | 22.1328 | 22.3355 | 21.8465 | 45.8377 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 19.9754 | 3.1601 | 0.008281 | 2.4753 | 2.1574 | 2.3601 | 1.8711
2012-10-30 14:11 | 23.1291 | 22.4549 | 22.1291 | 22.3675 | 21.8468 | 45.8267 | 0.014181 | 0.34548 | 0.004899 | 24.3625 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.0356 | 3.1593 | 0.008275 | 2.4193 | 2.0935 | 2.3319 | 1.8112
2012-10-30 14:11 | 23.1224 | 22.4402 | 22.1304 | 22.3052 | 21.8003 | 45.8307 | 0.0142 | 0.34554 | 0.004907 | 24.3339 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 19.9689 | 3.1596 | 0.008281 | 2.4713 | 2.1615 | 2.3363 | 1.8314
2012-10-30 14:11 | 23.1284 | 22.4582 | 22.1165 | 22.3072 | 21.8222 | 45.851 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014395 | 0.23462 | 0.003377 | 16.299 | 19.9949 | 3.161 | 0.008281 | 2.4633 | 2.1216 | 2.3123 | 1.8273
2012-10-30 14:11 | 23.1269 | 22.4368 | 22.0951 | 22.3216 | 21.7809 | 45.8313 | 0.014184 | 0.3456 | 0.004902 | 24.3655 | 0.014404 | 0.23461 | 0.003379 | 16.2878 | 20.0014 | 3.1596 | 0.008281 | 2.4354 | 2.0937 | 2.3202 | 1.7795
2012-10-30 14:11 | 23.1351 | 22.4291 | 22.0834 | 22.31 | 21.7891 | 45.8203 | 0.014209 | 0.3456 | 0.004911 | 24.3226 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0576 | 3.1589 | 0.008288 | 2.3715 | 2.0258 | 2.2524 | 1.7315
2012-10-30 14:11 | 23.1269 | 22.4368 | 22.0792 | 22.27 | 21.765 | 45.8238 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9934 | 3.1591 | 0.008278 | 2.4434 | 2.0858 | 2.2766 | 1.7716
2012-10-30 14:11 | 23.1306 | 22.3809 | 22.0551 | 22.3094 | 21.7965 | 45.8365 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0291 | 3.16 | 0.008279 | 2.3518 | 2.026 | 2.2803 | 1.7674
2012-10-30 14:11 | 23.1366 | 22.379 | 22.0889 | 22.2956 | 21.7508 | 45.8267 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.0111 | 3.1593 | 0.008278 | 2.3679 | 2.0778 | 2.2845 | 1.7397
2012-10-30 14:11 | 23.1343 | 22.3648 | 22.0548 | 22.2575 | 21.7405 | 45.8244 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0448 | 3.1591 | 0.008282 | 2.32 | 2.01 | 2.2127 | 1.6957
2012-10-30 14:11 | 23.1313 | 22.3816 | 21.9922 | 22.2505 | 21.7375 | 45.8232 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.0418 | 3.1591 | 0.008282 | 2.3398 | 1.9504 | 2.2087 | 1.6957
2012-10-30 14:11 | 23.1411 | 22.3438 | 22.0497 | 22.2842 | 21.7672 | 45.8255 | 0.014206 | 0.34553 | 0.004909 | 24.3227 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.0676 | 3.1592 | 0.008285 | 2.2762 | 1.9821 | 2.2166 | 1.6996
2012-10-30 14:11 | 23.1373 | 22.336 | 22.0538 | 22.2764 | 21.6838 | 45.8134 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 20.0118 | 3.1584 | 0.008283 | 2.3242 | 2.042 | 2.2646 | 1.672
2012-10-30 14:11 | 23.1478 | 22.3425 | 22.0643 | 22.2392 | 21.7501 | 45.8319 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0823 | 3.1597 | 0.008285 | 2.2602 | 1.982 | 2.1569 | 1.6678
2012-10-30 14:12 | 23.1448 | 22.3038 | 22.0017 | 22.2442 | 21.7033 | 45.829 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.0433 | 3.1595 | 0.008285 | 2.2605 | 1.9584 | 2.2009 | 1.66
2012-10-30 14:12 | 23.1441 | 22.3428 | 22.0288 | 22.2077 | 21.7066 | 45.8163 | 0.01419 | 0.34551 | 0.004903 | 24.349 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0506 | 3.1586 | 0.008279 | 2.2922 | 1.9782 | 2.1571 | 1.656
2012-10-30 14:12 | 23.147 | 22.3298 | 21.9999 | 22.2066 | 21.7214 | 45.8186 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.1095 | 3.1587 | 0.008282 | 2.2203 | 1.8904 | 2.0971 | 1.6119
2012-10-30 14:12 | 23.1485 | 22.2558 | 22.0412 | 22.2002 | 21.707 | 45.8348 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.047 | 3.1599 | 0.00828 | 2.2088 | 1.9942 | 2.1532 | 1.66
2012-10-30 14:12 | 23.1515 | 22.2946 | 22.0362 | 22.2072 | 21.6861 | 45.8169 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.094 | 3.1586 | 0.008282 | 2.2006 | 1.9422 | 2.1132 | 1.5921
2012-10-30 14:12 | 23.1478 | 22.2511 | 22.0166 | 22.1558 | 21.7262 | 45.825 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.1023 | 3.1592 | 0.008282 | 2.1488 | 1.9143 | 2.0535 | 1.6239
2012-10-30 14:12 | 23.15 | 22.2494 | 22.0029 | 22.158 | 21.7005 | 45.8284 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 20.0765 | 3.1594 | 0.008286 | 2.1729 | 1.9264 | 2.0815 | 1.624
2012-10-30 14:12 | 23.147 | 22.2702 | 21.9681 | 22.1669 | 21.7134 | 45.8047 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0495 | 3.1578 | 0.008285 | 2.2207 | 1.9186 | 2.1174 | 1.6639
2012-10-30 14:12 | 23.1485 | 22.228 | 21.9656 | 22.1406 | 21.6791 | 45.8024 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.055 | 3.1576 | 0.008281 | 2.173 | 1.9106 | 2.0856 | 1.6241
2012-10-30 14:12 | 23.156 | 22.2395 | 21.9453 | 22.16 | 21.6667 | 45.8215 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0505 | 3.1589 | 0.008279 | 2.189 | 1.8948 | 2.1095 | 1.6162
2012-10-30 14:12 | 23.1575 | 22.2171 | 21.9507 | 22.1257 | 21.6722 | 45.8198 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.052 | 3.1588 | 0.008289 | 2.1651 | 1.8987 | 2.0737 | 1.6202
2012-10-30 14:12 | 23.165 | 22.2604 | 21.9463 | 22.1451 | 21.6598 | 45.8053 | 0.014178 | 0.34554 | 0.004899 | 24.3717 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0435 | 3.1578 | 0.008275 | 2.2169 | 1.9028 | 2.1016 | 1.6163
2012-10-30 14:12 | 23.1627 | 22.2144 | 21.948 | 22.1389 | 21.6415 | 45.8226 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0292 | 3.159 | 0.008283 | 2.1852 | 1.9188 | 2.1097 | 1.6123
2012-10-30 14:12 | 23.165 | 22.2087 | 21.9463 | 22.1173 | 21.5722 | 45.8047 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0195 | 3.1578 | 0.008279 | 2.1892 | 1.9268 | 2.0978 | 1.5527
2012-10-30 14:12 | 23.1695 | 22.2132 | 21.9429 | 22.0781 | 21.6324 | 45.8111 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.04 | 3.1582 | 0.008284 | 2.1732 | 1.9029 | 2.0381 | 1.5924
2012-10-30 14:12 | 23.171 | 22.2147 | 21.9205 | 22.1472 | 21.5941 | 45.8117 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0575 | 3.1583 | 0.008282 | 2.1572 | 1.863 | 2.0897 | 1.5366
2012-10-30 14:12 | 23.168 | 22.172 | 21.9294 | 22.1124 | 21.5831 | 45.8018 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014376 | 0.23458 | 0.003372 | 16.3175 | 20.0265 | 3.1576 | 0.008276 | 2.1455 | 1.9029 | 2.0859 | 1.5566
2012-10-30 14:12 | 23.1657 | 22.1657 | 21.8873 | 22.0544 | 21.5888 | 45.8076 | 0.01419 | 0.34553 | 0.004903 | 24.3501 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 19.9802 | 3.158 | 0.008284 | 2.1855 | 1.9071 | 2.0742 | 1.6086
2012-10-30 14:12 | 23.1747 | 22.1906 | 21.9162 | 22.0634 | 21.5699 | 45.8099 | 0.01419 | 0.34554 | 0.004903 | 24.3511 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.0452 | 3.1581 | 0.008278 | 2.1454 | 1.871 | 2.0182 | 1.5247
2012-10-30 14:12 | 23.1732 | 22.1573 | 21.9107 | 22.0778 | 21.5844 | 45.8094 | 0.014203 | 0.34556 | 0.004908 | 24.3299 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.0757 | 3.1581 | 0.008286 | 2.0816 | 1.835 | 2.0021 | 1.5087
2012-10-30 14:12 | 23.1792 | 22.1593 | 21.9088 | 22.0599 | 21.5943 | 45.8163 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0897 | 3.1586 | 0.008285 | 2.0696 | 1.8191 | 1.9702 | 1.5046
2012-10-30 14:12 | 23.168 | 22.1322 | 21.8737 | 22.0448 | 21.5592 | 45.8007 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.0625 | 3.1575 | 0.008285 | 2.0697 | 1.8112 | 1.9823 | 1.4967
2012-10-30 14:12 | 23.1755 | 22.1874 | 21.8892 | 22.0562 | 21.5986 | 45.7943 | 0.014194 | 0.34556 | 0.004905 | 24.3453 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.054 | 3.1571 | 0.00828 | 2.1334 | 1.8352 | 2.0022 | 1.5446
2012-10-30 14:12 | 23.1822 | 22.1067 | 21.8919 | 22.0391 | 21.5376 | 45.8111 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0767 | 3.1582 | 0.008285 | 2.03 | 1.8152 | 1.9624 | 1.4609
2012-10-30 14:12 | 23.1822 | 22.1146 | 21.8879 | 22.0232 | 21.5734 | 45.7845 | 0.014212 | 0.34554 | 0.004911 | 24.3134 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0767 | 3.1564 | 0.008288 | 2.0379 | 1.8112 | 1.9465 | 1.4967
2012-10-30 14:12 | 23.1844 | 22.1009 | 21.914 | 22.0214 | 21.5796 | 45.8111 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0629 | 3.1582 | 0.008285 | 2.038 | 1.8511 | 1.9585 | 1.5167
2012-10-30 14:13 | 23.1822 | 22.0987 | 21.868 | 22.0271 | 21.5535 | 45.8007 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.0727 | 3.1575 | 0.008283 | 2.026 | 1.7953 | 1.9544 | 1.4808
2012-10-30 14:13 | 23.1852 | 22.0977 | 21.8471 | 22.0063 | 21.5605 | 45.7938 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0677 | 3.157 | 0.008281 | 2.03 | 1.7794 | 1.9386 | 1.4928
2012-10-30 14:13 | 23.1755 | 22.0761 | 21.8812 | 21.9807 | 21.5269 | 45.8042 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.07 | 3.1577 | 0.008279 | 2.0061 | 1.8112 | 1.9107 | 1.4569
2012-10-30 14:13 | 23.1732 | 22.0659 | 21.8431 | 22.0062 | 21.5525 | 45.7978 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0517 | 3.1573 | 0.008285 | 2.0142 | 1.7914 | 1.9545 | 1.5008
2012-10-30 14:13 | 23.1829 | 22.0756 | 21.8369 | 21.9881 | 21.5502 | 45.7874 | 0.014197 | 0.34556 | 0.004906 | 24.3401 | 0.014392 | 0.23442 | 0.003374 | 16.2883 | 20.0214 | 3.1566 | 0.00828 | 2.0542 | 1.8155 | 1.9667 | 1.5288
2012-10-30 14:13 | 23.1867 | 22.0674 | 21.8367 | 21.9601 | 21.4983 | 45.8007 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0372 | 3.1575 | 0.008286 | 2.0302 | 1.7995 | 1.9229 | 1.4611
2012-10-30 14:13 | 23.1874 | 22.0363 | 21.8016 | 21.9568 | 21.4671 | 45.8053 | 0.014187 | 0.34547 | 0.004901 | 24.3512 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0419 | 3.1578 | 0.008277 | 1.9944 | 1.7597 | 1.9149 | 1.4252
2012-10-30 14:13 | 23.1904 | 22.0353 | 21.8126 | 21.9558 | 21.506 | 45.7949 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0449 | 3.1571 | 0.008277 | 1.9904 | 1.7677 | 1.9109 | 1.4611
2012-10-30 14:13 | 23.1859 | 22.0428 | 21.8081 | 21.9553 | 21.5333 | 45.788 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014386 | 0.23445 | 0.003373 | 16.2971 | 20.0364 | 3.1566 | 0.008279 | 2.0064 | 1.7717 | 1.9189 | 1.4969
2012-10-30 14:13 | 23.1889 | 22.0458 | 21.815 | 21.9384 | 21.5443 | 45.807 | 0.014184 | 0.34548 | 0.0049 | 24.3574 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0714 | 3.1579 | 0.008277 | 1.9744 | 1.7436 | 1.867 | 1.4729
2012-10-30 14:13 | 23.1919 | 22.009 | 21.7942 | 21.9294 | 21.4796 | 45.7914 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014386 | 0.23449 | 0.003373 | 16.3002 | 20.0664 | 3.1569 | 0.008278 | 1.9426 | 1.7278 | 1.863 | 1.4132
2012-10-30 14:13 | 23.1949 | 22.02 | 21.7414 | 21.9046 | 21.4587 | 45.766 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 20.0334 | 3.1551 | 0.008279 | 1.9866 | 1.708 | 1.8712 | 1.4253
2012-10-30 14:13 | 23.1949 | 22.0239 | 21.8091 | 21.9285 | 21.4467 | 45.7897 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0734 | 3.1567 | 0.008279 | 1.9505 | 1.7357 | 1.8551 | 1.3733
2012-10-30 14:13 | 23.1919 | 21.9812 | 21.7703 | 21.9454 | 21.4995 | 45.7753 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0584 | 3.1557 | 0.008282 | 1.9228 | 1.7119 | 1.887 | 1.4411
2012-10-30 14:13 | 23.1964 | 22.0016 | 21.7907 | 21.8623 | 21.508 | 45.7753 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014379 | 0.23451 | 0.003372 | 16.3091 | 20.0789 | 3.1557 | 0.008279 | 1.9227 | 1.7118 | 1.7834 | 1.4291
2012-10-30 14:13 | 23.2001 | 21.9695 | 21.7626 | 21.9058 | 21.4639 | 45.7995 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0786 | 3.1574 | 0.008281 | 1.8909 | 1.684 | 1.8272 | 1.3853
2012-10-30 14:13 | 23.2001 | 21.9774 | 21.8063 | 21.862 | 21.4519 | 45.7735 | 0.014187 | 0.34556 | 0.004902 | 24.3573 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.0706 | 3.1556 | 0.008279 | 1.9068 | 1.7357 | 1.7914 | 1.3813
2012-10-30 14:13 | 23.1956 | 21.9729 | 21.7421 | 21.8456 | 21.4634 | 45.7741 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0461 | 3.1557 | 0.008279 | 1.9268 | 1.696 | 1.7995 | 1.4173
2012-10-30 14:13 | 23.1956 | 21.961 | 21.7541 | 21.8615 | 21.4514 | 45.8007 | 0.014175 | 0.34548 | 0.004897 | 24.3728 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0541 | 3.1575 | 0.008274 | 1.9069 | 1.7 | 1.8074 | 1.3973
2012-10-30 14:13 | 23.2106 | 21.972 | 21.777 | 21.8765 | 21.4505 | 45.7718 | 0.01419 | 0.34554 | 0.004903 | 24.3511 | 0.014386 | 0.23452 | 0.003374 | 16.3022 | 20.0531 | 3.1555 | 0.008277 | 1.9189 | 1.7239 | 1.8234 | 1.3974
2012-10-30 14:13 | 23.2188 | 21.9484 | 21.7534 | 21.8887 | 21.4228 | 45.7781 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.0813 | 3.1559 | 0.008283 | 1.8671 | 1.6721 | 1.8074 | 1.3415
2012-10-30 14:13 | 23.2009 | 21.9146 | 21.6758 | 21.839 | 21.4208 | 45.7891 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0594 | 3.1567 | 0.008278 | 1.8552 | 1.6164 | 1.7796 | 1.3614
2012-10-30 14:13 | 23.2173 | 21.9509 | 21.7121 | 21.8434 | 21.4691 | 45.7833 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0758 | 3.1563 | 0.008283 | 1.8751 | 1.6363 | 1.7676 | 1.3933
2012-10-30 14:13 | 23.2121 | 21.9059 | 21.683 | 21.874 | 21.444 | 45.7735 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014379 | 0.23448 | 0.003372 | 16.307 | 20.0746 | 3.1556 | 0.008274 | 1.8313 | 1.6084 | 1.7994 | 1.3694
2012-10-30 14:13 | 23.2173 | 21.927 | 21.7041 | 21.8156 | 21.4253 | 45.7787 | 0.014209 | 0.34553 | 0.00491 | 24.3175 | 0.014392 | 0.23445 | 0.003374 | 16.2903 | 20.0598 | 3.156 | 0.008284 | 1.8672 | 1.6443 | 1.7558 | 1.3655
2012-10-30 14:13 | 23.2143 | 21.8922 | 21.717 | 21.8404 | 21.4303 | 45.7729 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.0368 | 3.1556 | 0.008283 | 1.8554 | 1.6802 | 1.8036 | 1.3935
2012-10-30 14:13 | 23.2248 | 21.9226 | 21.6877 | 21.7912 | 21.4089 | 45.7579 | 0.014184 | 0.34547 | 0.0049 | 24.3563 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.0113 | 3.1545 | 0.008279 | 1.9113 | 1.6764 | 1.7799 | 1.3976
2012-10-30 14:14 | 23.2128 | 21.8946 | 21.6757 | 21.8031 | 21.365 | 45.7833 | 0.014206 | 0.34554 | 0.004909 | 24.3237 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0233 | 3.1563 | 0.008286 | 1.8713 | 1.6524 | 1.7798 | 1.3417
2012-10-30 14:14 | 23.2263 | 21.9201 | 21.6932 | 21.7768 | 21.3825 | 45.7625 | 0.014184 | 0.34547 | 0.0049 | 24.3563 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0688 | 3.1549 | 0.008276 | 1.8513 | 1.6244 | 1.708 | 1.3137
2012-10-30 14:14 | 23.2158 | 21.8738 | 21.6429 | 21.7942 | 21.36 | 45.7874 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014373 | 0.23449 | 0.00337 | 16.3149 | 20.0263 | 3.1566 | 0.00828 | 1.8475 | 1.6166 | 1.7679 | 1.3337
2012-10-30 14:14 | 23.2136 | 21.8755 | 21.6526 | 21.7601 | 21.3618 | 45.7695 | 0.014184 | 0.34546 | 0.0049 | 24.3553 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0281 | 3.1553 | 0.008276 | 1.8474 | 1.6245 | 1.732 | 1.3337
2012-10-30 14:14 | 23.2233 | 21.8375 | 21.6544 | 21.7937 | 21.3835 | 45.7475 | 0.014184 | 0.34541 | 0.004899 | 24.3522 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0418 | 3.1538 | 0.008275 | 1.7957 | 1.6126 | 1.7519 | 1.3417
2012-10-30 14:14 | 23.2151 | 21.8691 | 21.6581 | 21.7656 | 21.3514 | 45.7706 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.0296 | 3.1554 | 0.008285 | 1.8395 | 1.6285 | 1.736 | 1.3218
2012-10-30 14:14 | 23.2226 | 21.8328 | 21.6377 | 21.7691 | 21.3748 | 45.7545 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0371 | 3.1543 | 0.00828 | 1.7957 | 1.6006 | 1.732 | 1.3377
2012-10-30 14:14 | 23.2263 | 21.8684 | 21.6096 | 21.7489 | 21.3626 | 45.7654 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0328 | 3.1551 | 0.008283 | 1.8356 | 1.5768 | 1.7161 | 1.3298
2012-10-30 14:14 | 23.233 | 21.8671 | 21.6123 | 21.7755 | 21.3772 | 45.7677 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.0355 | 3.1552 | 0.008279 | 1.8316 | 1.5768 | 1.74 | 1.3417
2012-10-30 14:14 | 23.2293 | 21.8515 | 21.5847 | 21.7758 | 21.3656 | 45.7643 | 0.014209 | 0.34554 | 0.00491 | 24.3185 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 19.9838 | 3.155 | 0.008288 | 1.8677 | 1.6009 | 1.792 | 1.3818
2012-10-30 14:14 | 23.2308 | 21.8331 | 21.6181 | 21.7096 | 21.3471 | 45.7805 | 0.01419 | 0.34553 | 0.004903 | 24.3501 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.0093 | 3.1561 | 0.00828 | 1.8238 | 1.6088 | 1.7003 | 1.3378
2012-10-30 14:14 | 23.2285 | 21.8308 | 21.6038 | 21.7153 | 21.3568 | 45.7729 | 0.014206 | 0.34554 | 0.004909 | 24.3237 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.059 | 3.1556 | 0.008286 | 1.7718 | 1.5448 | 1.6563 | 1.2978
2012-10-30 14:14 | 23.2427 | 21.8529 | 21.626 | 21.7295 | 21.375 | 45.7625 | 0.01419 | 0.34551 | 0.004903 | 24.349 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.0172 | 3.1549 | 0.008282 | 1.8357 | 1.6088 | 1.7123 | 1.3578
2012-10-30 14:14 | 23.2338 | 21.8321 | 21.5772 | 21.7286 | 21.2544 | 45.7573 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014417 | 0.23457 | 0.003382 | 16.2701 | 20.0323 | 3.1545 | 0.008288 | 1.7998 | 1.5449 | 1.6963 | 1.2221
2012-10-30 14:14 | 23.2353 | 21.8336 | 21.5787 | 21.7102 | 21.3397 | 45.7573 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.0218 | 3.1545 | 0.008282 | 1.8118 | 1.5569 | 1.6884 | 1.3179
2012-10-30 14:14 | 23.2405 | 21.7592 | 21.572 | 21.7114 | 21.2891 | 45.7608 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.003 | 3.1547 | 0.008287 | 1.7562 | 1.569 | 1.7084 | 1.2861
2012-10-30 14:14 | 23.236 | 21.8383 | 21.5715 | 21.691 | 21.3045 | 45.777 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0105 | 3.1559 | 0.008283 | 1.8278 | 1.561 | 1.6805 | 1.294
2012-10-30 14:14 | 23.236 | 21.7905 | 21.5715 | 21.675 | 21.3125 | 45.7533 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 19.9945 | 3.1542 | 0.00828 | 1.796 | 1.577 | 1.6805 | 1.318
2012-10-30 14:14 | 23.2353 | 21.7818 | 21.5708 | 21.6903 | 21.2799 | 45.7498 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 19.9778 | 3.154 | 0.008282 | 1.804 | 1.593 | 1.7125 | 1.3021
2012-10-30 14:14 | 23.242 | 21.7885 | 21.5894 | 21.677 | 21.2427 | 45.7683 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 20.0205 | 3.1553 | 0.008279 | 1.768 | 1.5689 | 1.6565 | 1.2222
2012-10-30 14:14 | 23.2375 | 21.7761 | 21.5331 | 21.6566 | 21.298 | 45.7545 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 20.02 | 3.1543 | 0.008283 | 1.7561 | 1.5131 | 1.6366 | 1.278
2012-10-30 14:14 | 23.2427 | 21.7693 | 21.5702 | 21.6738 | 21.2873 | 45.755 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0212 | 3.1543 | 0.008281 | 1.7481 | 1.549 | 1.6526 | 1.2661
2012-10-30 14:14 | 23.2442 | 21.7589 | 21.5438 | 21.6633 | 21.2609 | 45.7521 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0187 | 3.1541 | 0.00828 | 1.7402 | 1.5251 | 1.6446 | 1.2422
2012-10-30 14:14 | 23.2353 | 21.75 | 21.5469 | 21.6305 | 21.248 | 45.755 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.0098 | 3.1543 | 0.008282 | 1.7402 | 1.5371 | 1.6207 | 1.2382
2012-10-30 14:14 | 23.2435 | 21.6945 | 21.5192 | 21.6108 | 21.2482 | 45.7591 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.034 | 3.1546 | 0.008282 | 1.6605 | 1.4852 | 1.5768 | 1.2142
2012-10-30 14:14 | 23.2517 | 21.7584 | 21.5234 | 21.6429 | 21.2763 | 45.7394 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.0342 | 3.1533 | 0.008282 | 1.7242 | 1.4892 | 1.6087 | 1.2421
2012-10-30 14:15 | 23.2457 | 21.7086 | 21.5493 | 21.6369 | 21.2384 | 45.7487 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.0362 | 3.1539 | 0.008284 | 1.6724 | 1.5131 | 1.6007 | 1.2022
2012-10-30 14:15 | 23.2487 | 21.7236 | 21.5005 | 21.5961 | 21.2893 | 45.7545 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.0392 | 3.1543 | 0.008277 | 1.6844 | 1.4613 | 1.5569 | 1.2501
2012-10-30 14:15 | 23.2502 | 21.7012 | 21.506 | 21.6215 | 21.219 | 45.7516 | 0.014187 | 0.3454 | 0.0049 | 24.3461 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0447 | 3.1541 | 0.008276 | 1.6565 | 1.4613 | 1.5768 | 1.1743
2012-10-30 14:15 | 23.248 | 21.7069 | 21.4958 | 21.5795 | 21.2487 | 45.7539 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 19.9745 | 3.1543 | 0.008284 | 1.7324 | 1.5213 | 1.605 | 1.2742
2012-10-30 14:15 | 23.251 | 21.7259 | 21.459 | 21.5984 | 21.2158 | 45.7481 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0575 | 3.1539 | 0.008278 | 1.6684 | 1.4015 | 1.5409 | 1.1583
2012-10-30 14:15 | 23.2472 | 21.6783 | 21.4831 | 21.5827 | 21.204 | 45.7469 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.0617 | 3.1538 | 0.008285 | 1.6166 | 1.4214 | 1.521 | 1.1423
2012-10-30 14:15 | 23.2517 | 21.7067 | 21.4477 | 21.5912 | 21.1886 | 45.7423 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.0062 | 3.1535 | 0.008282 | 1.7005 | 1.4415 | 1.585 | 1.1824
2012-10-30 14:15 | 23.2569 | 21.68 | 21.4529 | 21.5685 | 21.2417 | 45.7533 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0314 | 3.1542 | 0.008282 | 1.6486 | 1.4215 | 1.5371 | 1.2103
2012-10-30 14:15 | 23.2502 | 21.6693 | 21.4821 | 21.5299 | 21.207 | 45.7498 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014382 | 0.23454 | 0.003373 | 16.3077 | 20.0207 | 3.154 | 0.008278 | 1.6486 | 1.4614 | 1.5092 | 1.1863
2012-10-30 14:15 | 23.2502 | 21.6653 | 21.4741 | 21.5578 | 21.1831 | 45.7556 | 0.014184 | 0.34547 | 0.0049 | 24.3563 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.0407 | 3.1544 | 0.008274 | 1.6246 | 1.4334 | 1.5171 | 1.1424
2012-10-30 14:15 | 23.2502 | 21.6813 | 21.4462 | 21.5578 | 21.1991 | 45.7394 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0287 | 3.1533 | 0.008278 | 1.6526 | 1.4175 | 1.5291 | 1.1704
2012-10-30 14:15 | 23.2599 | 21.6711 | 21.436 | 21.5635 | 21.2008 | 45.7521 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.0064 | 3.1541 | 0.008281 | 1.6647 | 1.4296 | 1.5571 | 1.1944
2012-10-30 14:15 | 23.2517 | 21.631 | 21.4358 | 21.5354 | 21.1886 | 45.7579 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.0022 | 3.1545 | 0.008283 | 1.6288 | 1.4336 | 1.5332 | 1.1864
2012-10-30 14:15 | 23.2674 | 21.6347 | 21.4037 | 21.5511 | 21.2203 | 45.7365 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0339 | 3.1531 | 0.008278 | 1.6008 | 1.3698 | 1.5172 | 1.1864
2012-10-30 14:15 | 23.2682 | 21.6355 | 21.4244 | 21.532 | 21.1692 | 45.7539 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014379 | 0.23448 | 0.003372 | 16.307 | 20.0027 | 3.1543 | 0.008274 | 1.6328 | 1.4217 | 1.5293 | 1.1665
2012-10-30 14:15 | 23.2704 | 21.6577 | 21.4585 | 21.5143 | 21.1475 | 45.7469 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0129 | 3.1538 | 0.008282 | 1.6448 | 1.4456 | 1.5014 | 1.1346
2012-10-30 14:15 | 23.2682 | 21.6196 | 21.4045 | 21.5001 | 21.1852 | 45.7394 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014389 | 0.23457 | 0.003375 | 16.3018 | 20.0347 | 3.1533 | 0.008282 | 1.5849 | 1.3698 | 1.4654 | 1.1505
2012-10-30 14:15 | 23.2689 | 21.6243 | 21.453 | 21.5247 | 21.1181 | 45.7296 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.0354 | 3.1526 | 0.008279 | 1.5889 | 1.4176 | 1.4893 | 1.0827
2012-10-30 14:15 | 23.2704 | 21.6019 | 21.4186 | 21.5063 | 21.1475 | 45.7516 | 0.014206 | 0.34556 | 0.004909 | 24.3247 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0329 | 3.1541 | 0.008285 | 1.569 | 1.3857 | 1.4734 | 1.1146
2012-10-30 14:15 | 23.2689 | 21.5884 | 21.4131 | 21.5128 | 21.15 | 45.725 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.0114 | 3.1523 | 0.008279 | 1.577 | 1.4017 | 1.5014 | 1.1386
2012-10-30 14:15 | 23.2741 | 21.6335 | 21.3944 | 21.4542 | 21.1392 | 45.7406 | 0.014197 | 0.34557 | 0.004906 | 24.3411 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.0206 | 3.1534 | 0.008282 | 1.6129 | 1.3738 | 1.4336 | 1.1186
2012-10-30 14:15 | 23.2712 | 21.5947 | 21.3556 | 21.4951 | 21.0924 | 45.7342 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.0217 | 3.1529 | 0.008278 | 1.573 | 1.3339 | 1.4734 | 1.0707
2012-10-30 14:15 | 23.2756 | 21.6071 | 21.376 | 21.5115 | 21.1567 | 45.7336 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.0501 | 3.1529 | 0.008287 | 1.557 | 1.3259 | 1.4614 | 1.1066
2012-10-30 14:15 | 23.2816 | 21.6051 | 21.366 | 21.4657 | 21.1108 | 45.7365 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014386 | 0.23455 | 0.003374 | 16.3042 | 19.9961 | 3.1531 | 0.008275 | 1.609 | 1.3699 | 1.4696 | 1.1147
2012-10-30 14:15 | 23.2741 | 21.5458 | 21.3545 | 21.4781 | 21.1113 | 45.7383 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 19.9926 | 3.1532 | 0.008283 | 1.5532 | 1.3619 | 1.4855 | 1.1187
2012-10-30 14:15 | 23.2839 | 21.5716 | 21.3962 | 21.4281 | 21.0932 | 45.7204 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 19.9704 | 3.152 | 0.008283 | 1.6012 | 1.4258 | 1.4577 | 1.1228
2012-10-30 14:16 | 23.2809 | 21.5765 | 21.3414 | 21.4331 | 21.1221 | 45.7331 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0034 | 3.1528 | 0.008283 | 1.5731 | 1.338 | 1.4297 | 1.1187
2012-10-30 14:16 | 23.2861 | 21.5499 | 21.3466 | 21.4582 | 21.0794 | 45.7383 | 0.014209 | 0.34554 | 0.00491 | 24.3185 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 19.9726 | 3.1532 | 0.008286 | 1.5773 | 1.374 | 1.4856 | 1.1068
2012-10-30 14:16 | 23.2816 | 21.5215 | 21.3022 | 21.4418 | 21.0669 | 45.7227 | 0.014184 | 0.34546 | 0.0049 | 24.3553 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 19.9721 | 3.1521 | 0.008275 | 1.5494 | 1.3301 | 1.4697 | 1.0948
2012-10-30 14:16 | 23.2801 | 21.524 | 21.3446 | 21.3924 | 21.0335 | 45.744 | 0.014194 | 0.34551 | 0.004904 | 24.3422 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 19.9826 | 3.1536 | 0.00828 | 1.5414 | 1.362 | 1.4098 | 1.0509
2012-10-30 14:16 | 23.2884 | 21.5801 | 21.3449 | 21.4286 | 21.0857 | 45.729 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 19.9829 | 3.1526 | 0.008283 | 1.5972 | 1.362 | 1.4457 | 1.1028
2012-10-30 14:16 | 23.2884 | 21.5601 | 21.3409 | 21.4366 | 21.0817 | 45.7256 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9589 | 3.1523 | 0.008283 | 1.6012 | 1.382 | 1.4777 | 1.1228
2012-10-30 14:16 | 23.2839 | 21.5118 | 21.2487 | 21.4241 | 21.0892 | 45.74 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 19.9704 | 3.1533 | 0.008278 | 1.5414 | 1.2783 | 1.4537 | 1.1188
2012-10-30 14:16 | 23.3018 | 21.5217 | 21.3543 | 21.454 | 21.0592 | 45.7394 | 0.014175 | 0.34548 | 0.004897 | 24.3728 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 19.9883 | 3.1533 | 0.008275 | 1.5334 | 1.366 | 1.4657 | 1.0709
2012-10-30 14:16 | 23.2921 | 21.524 | 21.3486 | 21.4204 | 21.0176 | 45.7342 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.0066 | 3.1529 | 0.008284 | 1.5174 | 1.342 | 1.4138 | 1.011
2012-10-30 14:16 | 23.2943 | 21.4744 | 21.3189 | 21.3987 | 21.0278 | 45.718 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 19.9808 | 3.1518 | 0.008276 | 1.4936 | 1.3381 | 1.4179 | 1.047
2012-10-30 14:16 | 23.2891 | 21.4811 | 21.2858 | 21.3935 | 21.0425 | 45.736 | 0.014181 | 0.34551 | 0.0049 | 24.3645 | 0.014382 | 0.23445 | 0.003372 | 16.3016 | 19.9796 | 3.153 | 0.008272 | 1.5015 | 1.3062 | 1.4139 | 1.0629
2012-10-30 14:16 | 23.2928 | 21.4809 | 21.2696 | 21.3653 | 21.0223 | 45.7204 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 19.9833 | 3.152 | 0.008283 | 1.4976 | 1.2863 | 1.382 | 1.039
2012-10-30 14:16 | 23.2973 | 21.5133 | 21.29 | 21.3817 | 21.0108 | 45.7365 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 19.9718 | 3.1531 | 0.008287 | 1.5415 | 1.3182 | 1.4099 | 1.039
2012-10-30 14:16 | 23.2973 | 21.4575 | 21.2422 | 21.3578 | 21.0268 | 45.7215 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.0238 | 3.152 | 0.008283 | 1.4337 | 1.2184 | 1.334 | 1.003
2012-10-30 14:16 | 23.2973 | 21.4854 | 21.2701 | 21.3498 | 21.0228 | 45.7279 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 19.9598 | 3.1525 | 0.008279 | 1.5256 | 1.3103 | 1.39 | 1.063
2012-10-30 14:16 | 23.3063 | 21.4824 | 21.2911 | 21.3748 | 21.0437 | 45.7377 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014386 | 0.23451 | 0.003374 | 16.3011 | 19.9888 | 3.1532 | 0.008276 | 1.4936 | 1.3023 | 1.386 | 1.0549
2012-10-30 14:16 | 23.3055 | 21.4736 | 21.2743 | 21.3461 | 20.9911 | 45.7076 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.008 | 3.1511 | 0.008284 | 1.4656 | 1.2663 | 1.3381 | 0.9831
2012-10-30 14:16 | 23.3093 | 21.4615 | 21.2821 | 21.4097 | 20.9869 | 45.7319 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0278 | 3.1528 | 0.008285 | 1.4337 | 1.2543 | 1.3819 | 0.9591
2012-10-30 14:16 | 23.3063 | 21.4744 | 21.2831 | 21.3349 | 21.0078 | 45.718 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.0448 | 3.1518 | 0.008282 | 1.4296 | 1.2383 | 1.2901 | 0.963
2012-10-30 14:16 | 23.307 | 21.4512 | 21.2758 | 21.3436 | 21.0245 | 45.71 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.0575 | 3.1512 | 0.008281 | 1.3937 | 1.2183 | 1.2861 | 0.967
2012-10-30 14:16 | 23.3123 | 21.4207 | 21.3329 | 21.3569 | 20.9659 | 45.7371 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0428 | 3.1531 | 0.008282 | 1.3779 | 1.2901 | 1.3141 | 0.9231
2012-10-30 14:16 | 23.3168 | 21.4411 | 21.2417 | 21.3614 | 21.0223 | 45.7175 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.0593 | 3.1518 | 0.008286 | 1.3818 | 1.1824 | 1.3021 | 0.963
2012-10-30 14:16 | 23.31 | 21.4184 | 21.2429 | 21.3267 | 20.9756 | 45.7128 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014417 | 0.23459 | 0.003382 | 16.2721 | 20.0405 | 3.1514 | 0.008287 | 1.3779 | 1.2024 | 1.2862 | 0.9351
2012-10-30 14:16 | 23.313 | 21.4333 | 21.2499 | 21.3057 | 20.9826 | 45.729 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0435 | 3.1526 | 0.008285 | 1.3898 | 1.2064 | 1.2622 | 0.9391
2012-10-30 14:16 | 23.3227 | 21.4271 | 21.2277 | 21.3314 | 21.0083 | 45.7152 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0652 | 3.1516 | 0.008278 | 1.3619 | 1.1625 | 1.2662 | 0.9431
2012-10-30 14:16 | 23.325 | 21.4214 | 21.23 | 21.3137 | 20.9786 | 45.7111 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.0315 | 3.1513 | 0.008278 | 1.3899 | 1.1985 | 1.2822 | 0.9471
2012-10-30 14:17 | 23.3272 | 21.4037 | 21.2202 | 21.308 | 20.9689 | 45.7261 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.0057 | 3.1524 | 0.008282 | 1.398 | 1.2145 | 1.3023 | 0.9632
2012-10-30 14:17 | 23.3205 | 21.4289 | 21.2574 | 21.2614 | 20.9223 | 45.6995 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 19.991 | 3.1505 | 0.008281 | 1.4379 | 1.2664 | 1.2704 | 0.9313
2012-10-30 14:17 | 23.319 | 21.3875 | 21.1881 | 21.2719 | 20.9766 | 45.7221 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 19.9815 | 3.1521 | 0.008277 | 1.406 | 1.2066 | 1.2904 | 0.9951
2012-10-30 14:17 | 23.3257 | 21.3942 | 21.2267 | 21.2427 | 20.9594 | 45.7076 | 0.014194 | 0.3454 | 0.004903 | 24.3341 | 0.014389 | 0.23444 | 0.003373 | 16.2927 | 20.0242 | 3.1511 | 0.008276 | 1.37 | 1.2025 | 1.2185 | 0.9352
2012-10-30 14:17 | 23.3369 | 21.4014 | 21.1661 | 21.3017 | 20.9307 | 45.7059 | 0.014184 | 0.34556 | 0.004901 | 24.3625 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.0114 | 3.151 | 0.008276 | 1.39 | 1.1547 | 1.2903 | 0.9193
2012-10-30 14:17 | 23.3227 | 21.3752 | 21.1838 | 21.3154 | 20.9644 | 45.7094 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.0132 | 3.1512 | 0.008285 | 1.362 | 1.1706 | 1.3022 | 0.9512
2012-10-30 14:17 | 23.3265 | 21.3312 | 21.2076 | 21.2953 | 20.9402 | 45.6961 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.049 | 3.1503 | 0.008288 | 1.2822 | 1.1586 | 1.2463 | 0.8912
2012-10-30 14:17 | 23.3272 | 21.3718 | 21.1843 | 21.2601 | 20.901 | 45.7273 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.0217 | 3.1524 | 0.008284 | 1.3501 | 1.1626 | 1.2384 | 0.8793
2012-10-30 14:17 | 23.3369 | 21.3456 | 21.194 | 21.2459 | 20.9546 | 45.7071 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.0394 | 3.151 | 0.008285 | 1.3062 | 1.1546 | 1.2065 | 0.9152
2012-10-30 14:17 | 23.3272 | 21.3558 | 21.1883 | 21.2521 | 20.921 | 45.699 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.0177 | 3.1505 | 0.008279 | 1.3381 | 1.1706 | 1.2344 | 0.9033
2012-10-30 14:17 | 23.331 | 21.3636 | 21.1642 | 21.24 | 20.9128 | 45.703 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0135 | 3.1508 | 0.008284 | 1.3501 | 1.1507 | 1.2265 | 0.8993
2012-10-30 14:17 | 23.3377 | 21.3504 | 21.1908 | 21.2347 | 20.9275 | 45.6845 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 20.0322 | 3.1495 | 0.008283 | 1.3182 | 1.1586 | 1.2025 | 0.8953
2012-10-30 14:17 | 23.3325 | 21.3372 | 21.1577 | 21.2614 | 20.8943 | 45.7238 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.047 | 3.1522 | 0.008286 | 1.2902 | 1.1107 | 1.2144 | 0.8473
2012-10-30 14:17 | 23.3437 | 21.3643 | 21.1689 | 21.2527 | 20.9215 | 45.7007 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 20.0502 | 3.1506 | 0.008279 | 1.3141 | 1.1187 | 1.2025 | 0.8713
2012-10-30 14:17 | 23.3392 | 21.3519 | 21.1365 | 21.2203 | 20.905 | 45.6972 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.0177 | 3.1504 | 0.008279 | 1.3342 | 1.1188 | 1.2026 | 0.8873
2012-10-30 14:17 | 23.3407 | 21.3135 | 21.146 | 21.2497 | 20.8546 | 45.7221 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014386 | 0.23449 | 0.003373 | 16.3002 | 20.0072 | 3.1521 | 0.008279 | 1.3063 | 1.1388 | 1.2425 | 0.8474
2012-10-30 14:17 | 23.3362 | 21.309 | 21.1654 | 21.2412 | 20.902 | 45.6967 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 19.9987 | 3.1503 | 0.008284 | 1.3103 | 1.1667 | 1.2425 | 0.9033
2012-10-30 14:17 | 23.3482 | 21.309 | 21.1535 | 21.1974 | 20.8981 | 45.7204 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.0387 | 3.152 | 0.008288 | 1.2703 | 1.1148 | 1.1587 | 0.8594
2012-10-30 14:17 | 23.3512 | 21.316 | 21.1325 | 21.2402 | 20.917 | 45.6961 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.0697 | 3.1503 | 0.008282 | 1.2463 | 1.0628 | 1.1705 | 0.8473
2012-10-30 14:17 | 23.3444 | 21.2773 | 21.1098 | 21.2175 | 20.8943 | 45.6863 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0589 | 3.1496 | 0.008281 | 1.2184 | 1.0509 | 1.1586 | 0.8354
2012-10-30 14:17 | 23.3474 | 21.3401 | 21.1527 | 21.2085 | 20.8653 | 45.7221 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.0139 | 3.1521 | 0.008282 | 1.3262 | 1.1388 | 1.1946 | 0.8514
2012-10-30 14:17 | 23.3504 | 21.3072 | 21.1397 | 21.1916 | 20.8483 | 45.7048 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.0409 | 3.1509 | 0.008284 | 1.2663 | 1.0988 | 1.1507 | 0.8074
2012-10-30 14:17 | 23.3541 | 21.3109 | 21.1155 | 21.2033 | 20.896 | 45.6995 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014408 | 0.23442 | 0.003378 | 16.2702 | 19.9966 | 3.1505 | 0.008285 | 1.3143 | 1.1189 | 1.2067 | 0.8994
2012-10-30 14:17 | 23.3474 | 21.2723 | 21.1048 | 21.1447 | 20.8933 | 45.7013 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23441 | 0.003376 | 16.2771 | 20.0539 | 3.1506 | 0.008281 | 1.2184 | 1.0509 | 1.0908 | 0.8394
2012-10-30 14:17 | 23.3556 | 21.3005 | 21.1409 | 21.1808 | 20.8216 | 45.6863 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0581 | 3.1496 | 0.008278 | 1.2424 | 1.0828 | 1.1227 | 0.7635
2012-10-30 14:17 | 23.3497 | 21.2667 | 21.1071 | 21.1789 | 20.8676 | 45.7076 | 0.014184 | 0.34546 | 0.0049 | 24.3553 | 0.014386 | 0.23447 | 0.003373 | 16.2981 | 20.0442 | 3.1511 | 0.008273 | 1.2225 | 1.0629 | 1.1347 | 0.8234
2012-10-30 14:17 | 23.3497 | 21.2826 | 21.1031 | 21.1749 | 20.8556 | 45.6782 | 0.014197 | 0.34538 | 0.004903 | 24.3279 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0362 | 3.1491 | 0.008278 | 1.2464 | 1.0669 | 1.1387 | 0.8194
2012-10-30 14:18 | 23.3512 | 21.2961 | 21.0926 | 21.1645 | 20.8252 | 45.7013 | 0.014197 | 0.34538 | 0.004903 | 24.3279 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.0337 | 3.1506 | 0.008281 | 1.2624 | 1.0589 | 1.1308 | 0.7915
2012-10-30 14:18 | 23.3467 | 21.2876 | 21.0762 | 21.14 | 20.8167 | 45.7007 | 0.014184 | 0.3455 | 0.004901 | 24.3584 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.0412 | 3.1506 | 0.008277 | 1.2464 | 1.035 | 1.0988 | 0.7755
2012-10-30 14:18 | 23.3534 | 21.2704 | 21.1307 | 21.1467 | 20.8194 | 45.6845 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.0199 | 3.1495 | 0.008284 | 1.2505 | 1.1108 | 1.1268 | 0.7995
2012-10-30 14:18 | 23.3519 | 21.233 | 21.0933 | 21.1292 | 20.8099 | 45.703 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.0344 | 3.1508 | 0.008283 | 1.1986 | 1.0589 | 1.0948 | 0.7755
2012-10-30 14:18 | 23.3467 | 21.2397 | 21.0881 | 21.1839 | 20.8486 | 45.6949 | 0.014184 | 0.3455 | 0.004901 | 24.3584 | 0.014411 | 0.23442 | 0.003378 | 16.2668 | 20.0252 | 3.1502 | 0.008279 | 1.2145 | 1.0629 | 1.1587 | 0.8234
2012-10-30 14:18 | 23.3541 | 21.2192 | 21.0676 | 21.1713 | 20.8281 | 45.6799 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.0206 | 3.1492 | 0.008287 | 1.1986 | 1.047 | 1.1507 | 0.8075
2012-10-30 14:18 | 23.3549 | 21.2798 | 21.0564 | 21.1243 | 20.8409 | 45.6938 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0454 | 3.1501 | 0.008285 | 1.2344 | 1.011 | 1.0789 | 0.7955
2012-10-30 14:18 | 23.3549 | 21.236 | 21.0644 | 21.1322 | 20.8688 | 45.6759 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.0454 | 3.1489 | 0.008282 | 1.1906 | 1.019 | 1.0868 | 0.8234
2012-10-30 14:18 | 23.3512 | 21.2123 | 21.0727 | 21.1206 | 20.8092 | 45.6932 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014386 | 0.23449 | 0.003373 | 16.3002 | 19.9937 | 3.1501 | 0.008277 | 1.2186 | 1.079 | 1.1269 | 0.8155
2012-10-30 14:18 | 23.3609 | 21.218 | 21.1063 | 21.1183 | 20.8429 | 45.7019 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0394 | 3.1507 | 0.008278 | 1.1786 | 1.0669 | 1.0789 | 0.8035
2012-10-30 14:18 | 23.3556 | 21.2048 | 21.0811 | 21.0611 | 20.8216 | 45.6811 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0021 | 3.1493 | 0.008283 | 1.2027 | 1.079 | 1.059 | 0.8195
2012-10-30 14:18 | 23.3594 | 21.1926 | 21.0649 | 21.1208 | 20.7855 | 45.6822 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.0019 | 3.1493 | 0.008283 | 1.1907 | 1.063 | 1.1189 | 0.7836
2012-10-30 14:18 | 23.3571 | 21.2063 | 21.0586 | 21.1105 | 20.7672 | 45.6938 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.0476 | 3.1501 | 0.008282 | 1.1587 | 1.011 | 1.0629 | 0.7196
2012-10-30 14:18 | 23.3594 | 21.2046 | 21.0609 | 21.1008 | 20.8094 | 45.6845 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.0299 | 3.1495 | 0.008276 | 1.1747 | 1.031 | 1.0709 | 0.7795
2012-10-30 14:18 | 23.3571 | 21.2102 | 21.0746 | 21.1145 | 20.7951 | 45.6984 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.0396 | 3.1504 | 0.008284 | 1.1706 | 1.035 | 1.0749 | 0.7555
2012-10-30 14:18 | 23.3571 | 21.1983 | 21.0626 | 21.0786 | 20.8031 | 45.6683 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0196 | 3.1484 | 0.008282 | 1.1787 | 1.043 | 1.059 | 0.7835
2012-10-30 14:18 | 23.3564 | 21.1896 | 21.0659 | 21.0938 | 20.7865 | 45.6816 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0229 | 3.1493 | 0.008282 | 1.1667 | 1.043 | 1.0709 | 0.7636
2012-10-30 14:18 | 23.3579 | 21.215 | 21.0714 | 21.0874 | 20.8039 | 45.7053 | 0.014209 | 0.34554 | 0.00491 | 24.3185 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.0364 | 3.1509 | 0.008287 | 1.1786 | 1.035 | 1.051 | 0.7675
2012-10-30 14:18 | 23.3624 | 21.2116 | 21.0599 | 21.1038 | 20.7845 | 45.6741 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.0289 | 3.1488 | 0.008282 | 1.1827 | 1.031 | 1.0749 | 0.7556
2012-10-30 14:18 | 23.3624 | 21.1717 | 21.036 | 21.0599 | 20.7725 | 45.6868 | 0.0142 | 0.34553 | 0.004906 | 24.333 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 19.9728 | 3.1496 | 0.008284 | 1.1989 | 1.0632 | 1.0871 | 0.7997
2012-10-30 14:18 | 23.3661 | 21.1554 | 21.0317 | 21.0397 | 20.7762 | 45.6845 | 0.014209 | 0.34554 | 0.00491 | 24.3185 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.0286 | 3.1495 | 0.008291 | 1.1268 | 1.0031 | 1.0111 | 0.7476
2012-10-30 14:18 | 23.3631 | 21.1724 | 21.0367 | 21.0926 | 20.7532 | 45.6556 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0136 | 3.1475 | 0.008282 | 1.1588 | 1.0231 | 1.079 | 0.7396
2012-10-30 14:18 | 23.3556 | 21.105 | 21.0372 | 21.0651 | 20.7138 | 45.6874 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0221 | 3.1497 | 0.008282 | 1.0829 | 1.0151 | 1.043 | 0.6917
2012-10-30 14:18 | 23.3631 | 21.1165 | 21.0048 | 21.0567 | 20.7812 | 45.6868 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0056 | 3.1496 | 0.008279 | 1.1109 | 0.9992 | 1.0511 | 0.7756
2012-10-30 14:18 | 23.3631 | 21.1045 | 21.0088 | 21.0327 | 20.7572 | 45.6707 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.0416 | 3.1485 | 0.008284 | 1.0629 | 0.9672 | 0.9911 | 0.7156
2012-10-30 14:19 | 23.3631 | 21.1245 | 21.0127 | 21.0527 | 20.7492 | 45.6764 | 0.014184 | 0.34556 | 0.004901 | 24.3625 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0456 | 3.1489 | 0.008279 | 1.0789 | 0.9671 | 1.0071 | 0.7036
2012-10-30 14:19 | 23.3743 | 21.1317 | 21 | 21.0279 | 20.7444 | 45.6863 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.0368 | 3.1496 | 0.00828 | 1.0949 | 0.9632 | 0.9911 | 0.7076
2012-10-30 14:19 | 23.3713 | 21.1327 | 21.013 | 21.0569 | 20.7414 | 45.6759 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0498 | 3.1489 | 0.008282 | 1.0829 | 0.9632 | 1.0071 | 0.6916
2012-10-30 14:19 | 23.3654 | 21.0709 | 20.9951 | 20.9991 | 20.6876 | 45.6747 | 0.014187 | 0.34551 | 0.004902 | 24.3542 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0359 | 3.1488 | 0.008279 | 1.035 | 0.9592 | 0.9632 | 0.6517
2012-10-30 14:19 | 23.3736 | 21.115 | 20.9993 | 20.9714 | 20.7198 | 45.6787 | 0.014184 | 0.34551 | 0.004901 | 24.3593 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.0361 | 3.1491 | 0.008279 | 1.0789 | 0.9632 | 0.9353 | 0.6837
2012-10-30 14:19 | 23.3773 | 21.0828 | 20.9751 | 21.0309 | 20.7514 | 45.6579 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014386 | 0.23448 | 0.003373 | 16.2991 | 20.0638 | 3.1477 | 0.008281 | 1.019 | 0.9113 | 0.9671 | 0.6876
2012-10-30 14:19 | 23.3743 | 21.1157 | 21.008 | 21.008 | 20.7125 | 45.6897 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.0248 | 3.1498 | 0.00828 | 1.0909 | 0.9832 | 0.9832 | 0.6877
2012-10-30 14:19 | 23.3691 | 21.0746 | 20.9629 | 21.0028 | 20.7113 | 45.6683 | 0.014219 | 0.34548 | 0.004912 | 24.2974 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.0356 | 3.1484 | 0.008292 | 1.039 | 0.9273 | 0.9672 | 0.6757
2012-10-30 14:19 | 23.3706 | 21.0721 | 20.9963 | 21.0083 | 20.6968 | 45.6672 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.009 | 3.1483 | 0.008282 | 1.0631 | 0.9873 | 0.9993 | 0.6878
2012-10-30 14:19 | 23.3781 | 21.1036 | 20.9759 | 20.9719 | 20.7043 | 45.6759 | 0.014187 | 0.34548 | 0.004901 | 24.3522 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0406 | 3.1489 | 0.008276 | 1.063 | 0.9353 | 0.9313 | 0.6637
2012-10-30 14:19 | 23.3751 | 21.0567 | 20.9409 | 20.9689 | 20.6813 | 45.6707 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.0176 | 3.1485 | 0.008285 | 1.0391 | 0.9233 | 0.9513 | 0.6637
2012-10-30 14:19 | 23.3811 | 21.0747 | 20.9789 | 20.9749 | 20.7153 | 45.6498 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.0636 | 3.1471 | 0.008284 | 1.0111 | 0.9153 | 0.9113 | 0.6517
2012-10-30 14:19 | 23.3863 | 21.0479 | 20.9441 | 20.9601 | 20.7165 | 45.6672 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0368 | 3.1483 | 0.008284 | 1.0111 | 0.9073 | 0.9233 | 0.6797
2012-10-30 14:19 | 23.3855 | 21.0751 | 20.9194 | 20.9912 | 20.7117 | 45.6678 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0159 | 3.1483 | 0.008282 | 1.0592 | 0.9035 | 0.9753 | 0.6958
2012-10-30 14:19 | 23.3781 | 21.0557 | 20.9799 | 20.9719 | 20.7203 | 45.6602 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.0125 | 3.1478 | 0.008283 | 1.0432 | 0.9674 | 0.9594 | 0.7078
2012-10-30 14:19 | 23.3885 | 21.0741 | 20.9503 | 20.9783 | 20.7227 | 45.6678 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.0149 | 3.1483 | 0.008282 | 1.0592 | 0.9354 | 0.9634 | 0.7078
2012-10-30 14:19 | 23.3863 | 21.0439 | 20.9322 | 20.9841 | 20.6965 | 45.6857 | 0.014184 | 0.34548 | 0.0049 | 24.3574 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.0648 | 3.1496 | 0.008275 | 0.9791 | 0.8674 | 0.9193 | 0.6317
2012-10-30 14:19 | 23.3885 | 21.0701 | 20.9463 | 20.9384 | 20.7307 | 45.655 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.103 | 3.1475 | 0.008282 | 0.9671 | 0.8433 | 0.8354 | 0.6277
2012-10-30 14:19 | 23.3923 | 21.022 | 20.9302 | 20.9501 | 20.6546 | 45.666 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.0548 | 3.1482 | 0.008284 | 0.9672 | 0.8754 | 0.8953 | 0.5998
2012-10-30 14:19 | 23.3893 | 21.035 | 20.9831 | 20.9392 | 20.7035 | 45.6655 | 0.014175 | 0.3455 | 0.004897 | 24.3738 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.0718 | 3.1482 | 0.008276 | 0.9632 | 0.9113 | 0.8674 | 0.6317
2012-10-30 14:19 | 23.3855 | 21.0032 | 20.9354 | 20.9354 | 20.6678 | 45.6631 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.04 | 3.148 | 0.008281 | 0.9632 | 0.8954 | 0.8954 | 0.6278
2012-10-30 14:19 | 23.3915 | 21.0372 | 20.9533 | 20.9254 | 20.6937 | 45.6574 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.062 | 3.1476 | 0.008282 | 0.9752 | 0.8913 | 0.8634 | 0.6317
2012-10-30 14:19 | 23.3968 | 21.0145 | 20.9347 | 20.9347 | 20.707 | 45.6683 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.0833 | 3.1484 | 0.008285 | 0.9312 | 0.8514 | 0.8514 | 0.6237
2012-10-30 14:19 | 23.3938 | 21.0035 | 20.9397 | 20.9197 | 20.6801 | 45.6597 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1123 | 3.1478 | 0.008283 | 0.8912 | 0.8274 | 0.8074 | 0.5678
2012-10-30 14:19 | 23.393 | 20.9868 | 20.9628 | 20.9309 | 20.6273 | 45.6597 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1115 | 3.1478 | 0.00829 | 0.8753 | 0.8513 | 0.8194 | 0.5158
2012-10-30 14:20 | 23.3893 | 21.023 | 20.9551 | 20.9392 | 20.6396 | 45.6562 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1038 | 3.1475 | 0.008287 | 0.9192 | 0.8513 | 0.8354 | 0.5358
2012-10-30 14:20 | 23.3908 | 21.0125 | 20.9367 | 20.9407 | 20.6571 | 45.6602 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.1133 | 3.1478 | 0.008282 | 0.8992 | 0.8234 | 0.8274 | 0.5438
2012-10-30 14:20 | 23.3863 | 20.9761 | 20.9002 | 20.9162 | 20.6526 | 45.6666 | 0.014206 | 0.34556 | 0.004909 | 24.3247 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1168 | 3.1483 | 0.008286 | 0.8593 | 0.7834 | 0.7994 | 0.5358
2012-10-30 14:20 | 23.393 | 20.9748 | 20.9109 | 20.8949 | 20.6473 | 45.6539 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1115 | 3.1474 | 0.00828 | 0.8633 | 0.7994 | 0.7834 | 0.5358
2012-10-30 14:20 | 23.3923 | 20.9501 | 20.8942 | 20.9062 | 20.6546 | 45.6643 | 0.0142 | 0.34557 | 0.004907 | 24.336 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1148 | 3.1481 | 0.008283 | 0.8353 | 0.7794 | 0.7914 | 0.5398
2012-10-30 14:20 | 23.3945 | 21.0002 | 20.9444 | 20.8924 | 20.6408 | 45.6585 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.093 | 3.1477 | 0.008282 | 0.9072 | 0.8514 | 0.7994 | 0.5478
2012-10-30 14:20 | 23.3945 | 20.9483 | 20.9084 | 20.9084 | 20.6768 | 45.666 | 0.014212 | 0.34554 | 0.004911 | 24.3134 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.093 | 3.1482 | 0.008289 | 0.8553 | 0.8154 | 0.8154 | 0.5838
2012-10-30 14:20 | 23.3968 | 20.9467 | 20.8668 | 20.9107 | 20.6631 | 45.6579 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.0993 | 3.1477 | 0.008287 | 0.8474 | 0.7675 | 0.8114 | 0.5638
2012-10-30 14:20 | 23.399 | 20.9808 | 20.853 | 20.9169 | 20.6253 | 45.6527 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.0655 | 3.1473 | 0.00828 | 0.9153 | 0.7875 | 0.8514 | 0.5598
2012-10-30 14:20 | 23.3983 | 20.9681 | 20.8763 | 20.9082 | 20.6326 | 45.6741 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.0888 | 3.1488 | 0.008282 | 0.8793 | 0.7875 | 0.8194 | 0.5438
2012-10-30 14:20 | 23.3953 | 20.9412 | 20.8693 | 20.8453 | 20.6376 | 45.6522 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.0698 | 3.1473 | 0.008286 | 0.8714 | 0.7995 | 0.7755 | 0.5678
2012-10-30 14:20 | 23.3953 | 20.9611 | 20.8932 | 20.8493 | 20.6216 | 45.6562 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0898 | 3.1475 | 0.008282 | 0.8713 | 0.8034 | 0.7595 | 0.5318
2012-10-30 14:20 | 23.399 | 20.9688 | 20.869 | 20.861 | 20.6293 | 45.6493 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.0695 | 3.1471 | 0.00828 | 0.8993 | 0.7995 | 0.7915 | 0.5598
2012-10-30 14:20 | 23.396 | 20.9538 | 20.862 | 20.834 | 20.6343 | 45.6446 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.0505 | 3.1467 | 0.008282 | 0.9033 | 0.8115 | 0.7835 | 0.5838
2012-10-30 14:20 | 23.4005 | 20.9543 | 20.8385 | 20.8625 | 20.6668 | 45.6493 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.115 | 3.1471 | 0.008286 | 0.8393 | 0.7235 | 0.7475 | 0.5518
2012-10-30 14:20 | 23.3953 | 20.9252 | 20.8453 | 20.8094 | 20.5817 | 45.6533 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.0578 | 3.1473 | 0.008282 | 0.8674 | 0.7875 | 0.7516 | 0.5239
2012-10-30 14:20 | 23.396 | 20.9339 | 20.846 | 20.85 | 20.6463 | 45.6527 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 20.0385 | 3.1473 | 0.008286 | 0.8954 | 0.8075 | 0.8115 | 0.6078
2012-10-30 14:20 | 23.4042 | 20.9101 | 20.8103 | 20.7983 | 20.6385 | 45.6539 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.0266 | 3.1474 | 0.008281 | 0.8835 | 0.7837 | 0.7717 | 0.6119
2012-10-30 14:20 | 23.4005 | 20.9064 | 20.8425 | 20.8346 | 20.5909 | 45.6672 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.059 | 3.1483 | 0.008281 | 0.8474 | 0.7835 | 0.7756 | 0.5319
2012-10-30 14:20 | 23.4087 | 20.8987 | 20.8707 | 20.8348 | 20.615 | 45.6452 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.0471 | 3.1468 | 0.008285 | 0.8516 | 0.8236 | 0.7877 | 0.5679
2012-10-30 14:20 | 23.393 | 20.8909 | 20.8071 | 20.7871 | 20.5714 | 45.6435 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.0234 | 3.1467 | 0.008289 | 0.8675 | 0.7837 | 0.7637 | 0.548
2012-10-30 14:20 | 23.4005 | 20.9104 | 20.7906 | 20.8106 | 20.5749 | 45.6608 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.067 | 3.1479 | 0.008287 | 0.8434 | 0.7236 | 0.7436 | 0.5079
2012-10-30 14:20 | 23.4065 | 20.9124 | 20.8645 | 20.8206 | 20.5969 | 45.6464 | 0.014206 | 0.34556 | 0.004909 | 24.3247 | 0.014389 | 0.23459 | 0.003376 | 16.3038 | 20.097 | 3.1469 | 0.008285 | 0.8154 | 0.7675 | 0.7236 | 0.4999
2012-10-30 14:20 | 23.4005 | 20.8665 | 20.8266 | 20.8186 | 20.5988 | 45.6389 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.043 | 3.1463 | 0.008281 | 0.8235 | 0.7836 | 0.7756 | 0.5558
2012-10-30 14:20 | 23.4125 | 20.9104 | 20.8066 | 20.7986 | 20.5629 | 45.6579 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.0509 | 3.1477 | 0.008288 | 0.8595 | 0.7557 | 0.7477 | 0.512
2012-10-30 14:20 | 23.411 | 20.881 | 20.8211 | 20.8251 | 20.5934 | 45.6412 | 0.014203 | 0.3456 | 0.004909 | 24.3329 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.0735 | 3.1465 | 0.008284 | 0.8075 | 0.7476 | 0.7516 | 0.5199
2012-10-30 14:21 | 23.4125 | 20.8985 | 20.8226 | 20.8146 | 20.5429 | 45.6458 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.103 | 3.1468 | 0.008284 | 0.7955 | 0.7196 | 0.7116 | 0.4399
2012-10-30 14:21 | 23.4042 | 20.8462 | 20.7903 | 20.7863 | 20.5506 | 45.6446 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.0747 | 3.1467 | 0.008284 | 0.7715 | 0.7156 | 0.7116 | 0.4759
2012-10-30 14:21 | 23.4065 | 20.8366 | 20.8046 | 20.7926 | 20.5249 | 45.6458 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.0449 | 3.1468 | 0.008283 | 0.7917 | 0.7597 | 0.7477 | 0.48
2012-10-30 14:21 | 23.4102 | 20.8682 | 20.8083 | 20.7564 | 20.5606 | 45.6498 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.0967 | 3.1471 | 0.008286 | 0.7715 | 0.7116 | 0.6597 | 0.4639
2012-10-30 14:21 | 23.4155 | 20.8496 | 20.7896 | 20.7777 | 20.5779 | 45.6446 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.094 | 3.1467 | 0.008285 | 0.7556 | 0.6956 | 0.6837 | 0.4839
2012-10-30 14:21 | 23.4132 | 20.8712 | 20.7913 | 20.8233 | 20.5236 | 45.6556 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.0677 | 3.1475 | 0.008283 | 0.8035 | 0.7236 | 0.7556 | 0.4559
2012-10-30 14:21 | 23.417 | 20.8271 | 20.7752 | 20.7871 | 20.5394 | 45.6383 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0595 | 3.1463 | 0.008284 | 0.7676 | 0.7157 | 0.7276 | 0.4799
2012-10-30 14:21 | 23.414 | 20.884 | 20.7562 | 20.7881 | 20.5004 | 45.6244 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.0605 | 3.1453 | 0.008283 | 0.8235 | 0.6957 | 0.7276 | 0.4399
2012-10-30 14:21 | 23.4199 | 20.83 | 20.7821 | 20.802 | 20.5783 | 45.6516 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0704 | 3.1472 | 0.008284 | 0.7596 | 0.7117 | 0.7316 | 0.5079
2012-10-30 14:21 | 23.4214 | 20.8395 | 20.7796 | 20.7596 | 20.4998 | 45.636 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.0719 | 3.1461 | 0.008284 | 0.7676 | 0.7077 | 0.6877 | 0.4279
2012-10-30 14:21 | 23.4177 | 20.8438 | 20.7439 | 20.7279 | 20.5441 | 45.6446 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.0642 | 3.1467 | 0.00829 | 0.7796 | 0.6797 | 0.6637 | 0.4799
2012-10-30 14:21 | 23.4184 | 20.8325 | 20.7606 | 20.7606 | 20.5128 | 45.6418 | 0.01419 | 0.3455 | 0.004903 | 24.3481 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0488 | 3.1465 | 0.008279 | 0.7837 | 0.7118 | 0.7118 | 0.464
2012-10-30 14:21 | 23.4199 | 20.802 | 20.7781 | 20.7741 | 20.5063 | 45.6383 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0463 | 3.1463 | 0.008285 | 0.7557 | 0.7318 | 0.7278 | 0.46
2012-10-30 14:21 | 23.4199 | 20.834 | 20.7581 | 20.7501 | 20.5303 | 45.6412 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.0303 | 3.1465 | 0.008284 | 0.8037 | 0.7278 | 0.7198 | 0.5
2012-10-30 14:21 | 23.4327 | 20.8268 | 20.7549 | 20.7549 | 20.4951 | 45.6314 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0711 | 3.1458 | 0.008287 | 0.7557 | 0.6838 | 0.6838 | 0.424
2012-10-30 14:21 | 23.4229 | 20.7851 | 20.7131 | 20.7052 | 20.4773 | 45.64 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014411 | 0.23445 | 0.003379 | 16.2688 | 20.0373 | 3.1464 | 0.008286 | 0.7478 | 0.6758 | 0.6679 | 0.44
2012-10-30 14:21 | 23.4342 | 20.8163 | 20.7205 | 20.7524 | 20.4926 | 45.6435 | 0.014209 | 0.34554 | 0.00491 | 24.3185 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.0566 | 3.1467 | 0.008288 | 0.7597 | 0.6639 | 0.6958 | 0.436
2012-10-30 14:21 | 23.4356 | 20.8137 | 20.7498 | 20.7418 | 20.4701 | 45.6325 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014404 | 0.23458 | 0.003379 | 16.2858 | 20.062 | 3.1459 | 0.008284 | 0.7517 | 0.6878 | 0.6798 | 0.4081
2012-10-30 14:21 | 23.4379 | 20.82 | 20.7601 | 20.7162 | 20.4804 | 45.6435 | 0.014194 | 0.34559 | 0.004905 | 24.3473 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0323 | 3.1467 | 0.008282 | 0.7877 | 0.7278 | 0.6839 | 0.4481
2012-10-30 14:21 | 23.4356 | 20.7978 | 20.7578 | 20.6699 | 20.49 | 45.6389 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.034 | 3.1463 | 0.008285 | 0.7638 | 0.7238 | 0.6359 | 0.456
2012-10-30 14:21 | 23.4379 | 20.7641 | 20.7601 | 20.7002 | 20.4644 | 45.6412 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0363 | 3.1465 | 0.008284 | 0.7278 | 0.7238 | 0.6639 | 0.4281
2012-10-30 14:21 | 23.4364 | 20.7786 | 20.7027 | 20.7266 | 20.4749 | 45.6227 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.0668 | 3.1452 | 0.008285 | 0.7118 | 0.6359 | 0.6598 | 0.4081
2012-10-30 14:21 | 23.4454 | 20.8116 | 20.7436 | 20.7037 | 20.4679 | 45.6435 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 20.0398 | 3.1467 | 0.008282 | 0.7718 | 0.7038 | 0.6639 | 0.4281
2012-10-30 14:22 | 23.4499 | 20.7641 | 20.7242 | 20.6882 | 20.4764 | 45.6348 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0603 | 3.1461 | 0.008283 | 0.7038 | 0.6639 | 0.6279 | 0.4161
2012-10-30 14:22 | 23.4581 | 20.7683 | 20.7124 | 20.6724 | 20.4606 | 45.6094 | 0.01419 | 0.34548 | 0.004902 | 24.3471 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.0485 | 3.1443 | 0.00828 | 0.7198 | 0.6639 | 0.6239 | 0.4121
2012-10-30 14:22 | 23.4618 | 20.7081 | 20.7321 | 20.7201 | 20.4643 | 45.6337 | 0.014194 | 0.34554 | 0.004905 | 24.3442 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.0922 | 3.146 | 0.008279 | 0.6159 | 0.6399 | 0.6279 | 0.3721
2012-10-30 14:22 | 23.4626 | 20.7449 | 20.7289 | 20.7009 | 20.4811 | 45.6348 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1091 | 3.1461 | 0.008286 | 0.6358 | 0.6198 | 0.5918 | 0.372
2012-10-30 14:22 | 23.4596 | 20.7498 | 20.7139 | 20.6699 | 20.4621 | 45.6227 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.098 | 3.1452 | 0.008284 | 0.6518 | 0.6159 | 0.5719 | 0.3641
2012-10-30 14:22 | 23.4611 | 20.7553 | 20.7274 | 20.6515 | 20.4516 | 45.629 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.0995 | 3.1457 | 0.008282 | 0.6558 | 0.6279 | 0.552 | 0.3521
2012-10-30 14:22 | 23.4633 | 20.7336 | 20.6936 | 20.6537 | 20.4658 | 45.6394 | 0.014206 | 0.34553 | 0.004909 | 24.3227 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.0697 | 3.1464 | 0.008287 | 0.6639 | 0.6239 | 0.584 | 0.3961
2012-10-30 14:22 | 23.4648 | 20.7231 | 20.7391 | 20.5992 | 20.4393 | 45.625 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0952 | 3.1454 | 0.008283 | 0.6279 | 0.6439 | 0.504 | 0.3441
2012-10-30 14:22 | 23.4648 | 20.7431 | 20.6711 | 20.6312 | 20.4353 | 45.6198 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014404 | 0.23462 | 0.00338 | 16.2888 | 20.0712 | 3.145 | 0.008286 | 0.6719 | 0.5999 | 0.56 | 0.3641
2012-10-30 14:22 | 23.4768 | 20.779 | 20.7471 | 20.6991 | 20.4833 | 45.6262 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.0792 | 3.1455 | 0.008279 | 0.6998 | 0.6679 | 0.6199 | 0.4041
2012-10-30 14:22 | 23.4738 | 20.7081 | 20.7001 | 20.6642 | 20.4363 | 45.6181 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0922 | 3.1449 | 0.008279 | 0.6159 | 0.6079 | 0.572 | 0.3441
2012-10-30 14:22 | 23.4693 | 20.7476 | 20.7036 | 20.6477 | 20.4438 | 45.6186 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.1077 | 3.1449 | 0.008281 | 0.6399 | 0.5959 | 0.54 | 0.3361
2012-10-30 14:22 | 23.4745 | 20.7088 | 20.7008 | 20.6649 | 20.425 | 45.6366 | 0.014184 | 0.34546 | 0.0049 | 24.3553 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.0729 | 3.1462 | 0.008281 | 0.6359 | 0.6279 | 0.592 | 0.3521
2012-10-30 14:22 | 23.4783 | 20.7086 | 20.6726 | 20.6567 | 20.4208 | 45.6273 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1127 | 3.1455 | 0.008282 | 0.5959 | 0.5599 | 0.544 | 0.3081
2012-10-30 14:22 | 23.4745 | 20.6888 | 20.6888 | 20.6329 | 20.441 | 45.6233 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1129 | 3.1453 | 0.008285 | 0.5759 | 0.5759 | 0.52 | 0.3281
2012-10-30 14:22 | 23.4835 | 20.7178 | 20.7018 | 20.6739 | 20.406 | 45.6244 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1059 | 3.1453 | 0.008282 | 0.6119 | 0.5959 | 0.568 | 0.3001
2012-10-30 14:22 | 23.4828 | 20.6771 | 20.6771 | 20.6052 | 20.4693 | 45.6308 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1172 | 3.1458 | 0.008282 | 0.5599 | 0.5599 | 0.488 | 0.3521
2012-10-30 14:22 | 23.488 | 20.7303 | 20.6863 | 20.6344 | 20.4225 | 45.6204 | 0.014203 | 0.34554 | 0.004908 | 24.3288 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.0784 | 3.1451 | 0.008283 | 0.6519 | 0.6079 | 0.556 | 0.3441
2012-10-30 14:22 | 23.4813 | 20.6956 | 20.6677 | 20.6197 | 20.4118 | 45.6105 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0957 | 3.1444 | 0.008281 | 0.5999 | 0.572 | 0.524 | 0.3161
2012-10-30 14:22 | 23.4768 | 20.7311 | 20.6512 | 20.6352 | 20.3913 | 45.6169 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.0792 | 3.1448 | 0.008281 | 0.6519 | 0.572 | 0.556 | 0.3121
2012-10-30 14:22 | 23.4842 | 20.7185 | 20.6466 | 20.5946 | 20.3827 | 45.6273 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.1267 | 3.1455 | 0.008282 | 0.5918 | 0.5199 | 0.4679 | 0.256
2012-10-30 14:23 | 23.4887 | 20.695 | 20.6671 | 20.6071 | 20.4272 | 45.6256 | 0.01419 | 0.34557 | 0.004904 | 24.3531 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1432 | 3.1454 | 0.008284 | 0.5518 | 0.5239 | 0.4639 | 0.284
2012-10-30 14:23 | 23.4887 | 20.715 | 20.6511 | 20.6431 | 20.4232 | 45.6314 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.0951 | 3.1458 | 0.008286 | 0.6199 | 0.556 | 0.548 | 0.3281
2012-10-30 14:23 | 23.4857 | 20.6441 | 20.6521 | 20.5801 | 20.4082 | 45.6244 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014395 | 0.23444 | 0.003375 | 16.2859 | 20.1121 | 3.1453 | 0.008283 | 0.532 | 0.54 | 0.468 | 0.2961
2012-10-30 14:23 | 23.485 | 20.6634 | 20.6594 | 20.5954 | 20.4035 | 45.6215 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.0634 | 3.1451 | 0.008288 | 0.6 | 0.596 | 0.532 | 0.3401
2012-10-30 14:23 | 23.488 | 20.6784 | 20.6584 | 20.5784 | 20.4065 | 45.6088 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014417 | 0.23457 | 0.003382 | 16.2701 | 20.0704 | 3.1443 | 0.008286 | 0.608 | 0.588 | 0.508 | 0.3361
2012-10-30 14:23 | 23.4835 | 20.6858 | 20.6699 | 20.5579 | 20.33 | 45.6186 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.0939 | 3.1449 | 0.008281 | 0.5919 | 0.576 | 0.464 | 0.2361
2012-10-30 14:23 | 23.4887 | 20.6751 | 20.6551 | 20.6071 | 20.4152 | 45.6169 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014392 | 0.23444 | 0.003374 | 16.2893 | 20.0831 | 3.1448 | 0.008281 | 0.592 | 0.572 | 0.524 | 0.3321
2012-10-30 14:23 | 23.4842 | 20.6865 | 20.6666 | 20.5586 | 20.3627 | 45.6065 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.0626 | 3.1441 | 0.008285 | 0.6239 | 0.604 | 0.496 | 0.3001
2012-10-30 14:23 | 23.482 | 20.6644 | 20.6004 | 20.5604 | 20.3685 | 45.6152 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.0724 | 3.1447 | 0.008281 | 0.592 | 0.528 | 0.488 | 0.2961
2012-10-30 14:23 | 23.485 | 20.6554 | 20.5954 | 20.5674 | 20.3715 | 45.6273 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1034 | 3.1455 | 0.008284 | 0.552 | 0.492 | 0.464 | 0.2681
2012-10-30 14:23 | 23.4895 | 20.6639 | 20.6838 | 20.5839 | 20.412 | 45.6169 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1159 | 3.1448 | 0.008283 | 0.548 | 0.5679 | 0.468 | 0.2961
2012-10-30 14:23 | 23.4932 | 20.6156 | 20.6396 | 20.5556 | 20.3677 | 45.6152 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.0596 | 3.1447 | 0.008289 | 0.556 | 0.58 | 0.496 | 0.3081
2012-10-30 14:23 | 23.485 | 20.6594 | 20.6074 | 20.5794 | 20.3595 | 45.6262 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.0634 | 3.1455 | 0.008288 | 0.596 | 0.544 | 0.516 | 0.2961
2012-10-30 14:23 | 23.4887 | 20.6311 | 20.6271 | 20.5431 | 20.3792 | 45.6157 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.0511 | 3.1447 | 0.008284 | 0.58 | 0.576 | 0.492 | 0.3281
2012-10-30 14:23 | 23.4835 | 20.6459 | 20.5979 | 20.5699 | 20.382 | 45.6077 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0859 | 3.1442 | 0.008286 | 0.56 | 0.512 | 0.484 | 0.2961
2012-10-30 14:23 | 23.4902 | 20.6406 | 20.6086 | 20.5486 | 20.3487 | 45.6192 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.0486 | 3.145 | 0.008284 | 0.592 | 0.56 | 0.5 | 0.3001
2012-10-30 14:23 | 23.485 | 20.6394 | 20.6114 | 20.5235 | 20.3315 | 45.6366 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.1275 | 3.1462 | 0.008283 | 0.5119 | 0.4839 | 0.396 | 0.204
2012-10-30 14:23 | 23.4895 | 20.6119 | 20.6039 | 20.524 | 20.328 | 45.6094 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.0559 | 3.1443 | 0.008285 | 0.556 | 0.548 | 0.4681 | 0.2721
2012-10-30 14:23 | 23.4887 | 20.6151 | 20.5991 | 20.5152 | 20.3672 | 45.6123 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.0671 | 3.1445 | 0.008286 | 0.548 | 0.532 | 0.4481 | 0.3001
2012-10-30 14:23 | 23.4925 | 20.6109 | 20.6029 | 20.543 | 20.335 | 45.6134 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.0589 | 3.1446 | 0.008288 | 0.552 | 0.544 | 0.4841 | 0.2761
2012-10-30 14:23 | 23.4955 | 20.5819 | 20.5939 | 20.542 | 20.33 | 45.6048 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0339 | 3.144 | 0.008282 | 0.548 | 0.56 | 0.5081 | 0.2961
2012-10-30 14:23 | 23.491 | 20.6054 | 20.6054 | 20.5135 | 20.2855 | 45.6082 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23444 | 0.003378 | 16.2712 | 20.0574 | 3.1442 | 0.008286 | 0.548 | 0.548 | 0.4561 | 0.2281
2012-10-30 14:23 | 23.497 | 20.6594 | 20.5674 | 20.5355 | 20.3155 | 45.6181 | 0.014187 | 0.34553 | 0.004902 | 24.3553 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0714 | 3.1449 | 0.008279 | 0.588 | 0.496 | 0.4641 | 0.2441
2012-10-30 14:24 | 23.4955 | 20.6259 | 20.5699 | 20.514 | 20.37 | 45.6025 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0899 | 3.1438 | 0.008281 | 0.536 | 0.48 | 0.4241 | 0.2801
2012-10-30 14:24 | 23.4902 | 20.5926 | 20.5886 | 20.5087 | 20.3247 | 45.6013 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0846 | 3.1438 | 0.008284 | 0.508 | 0.504 | 0.4241 | 0.2401
2012-10-30 14:24 | 23.4977 | 20.6041 | 20.5761 | 20.5042 | 20.3202 | 45.6111 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1001 | 3.1444 | 0.008283 | 0.504 | 0.476 | 0.4041 | 0.2201
2012-10-30 14:24 | 23.4977 | 20.6401 | 20.5961 | 20.5242 | 20.3762 | 45.6215 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1321 | 3.1451 | 0.008288 | 0.508 | 0.464 | 0.3921 | 0.2441
2012-10-30 14:24 | 23.4985 | 20.6009 | 20.5609 | 20.513 | 20.341 | 45.5938 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1369 | 3.1432 | 0.008282 | 0.464 | 0.424 | 0.3761 | 0.2041
2012-10-30 14:24 | 23.4895 | 20.5719 | 20.5559 | 20.516 | 20.272 | 45.5967 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0959 | 3.1434 | 0.008287 | 0.476 | 0.46 | 0.4201 | 0.1761
2012-10-30 14:24 | 23.4925 | 20.5749 | 20.5989 | 20.503 | 20.311 | 45.6152 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1269 | 3.1447 | 0.008286 | 0.448 | 0.472 | 0.3761 | 0.1841
2012-10-30 14:24 | 23.494 | 20.5524 | 20.5405 | 20.5285 | 20.3245 | 45.5944 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1124 | 3.1433 | 0.00828 | 0.44 | 0.4281 | 0.4161 | 0.2121
2012-10-30 14:24 | 23.4962 | 20.6106 | 20.5746 | 20.5027 | 20.2947 | 45.6088 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0866 | 3.1443 | 0.008279 | 0.524 | 0.488 | 0.4161 | 0.2081
2012-10-30 14:24 | 23.4932 | 20.5756 | 20.5596 | 20.4597 | 20.2837 | 45.6169 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014414 | 0.23442 | 0.003379 | 16.2634 | 20.1036 | 3.1448 | 0.008283 | 0.472 | 0.456 | 0.3561 | 0.1801
2012-10-30 14:24 | 23.4917 | 20.5382 | 20.5461 | 20.5102 | 20.2942 | 45.6007 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1061 | 3.1437 | 0.008286 | 0.4321 | 0.44 | 0.4041 | 0.1881
2012-10-30 14:24 | 23.4955 | 20.5899 | 20.5739 | 20.482 | 20.286 | 45.599 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.1219 | 3.1436 | 0.008287 | 0.468 | 0.452 | 0.3601 | 0.1641
2012-10-30 14:24 | 23.4962 | 20.5946 | 20.5506 | 20.5027 | 20.2947 | 45.6088 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.0986 | 3.1443 | 0.008278 | 0.496 | 0.452 | 0.4041 | 0.1961
2012-10-30 14:24 | 23.5037 | 20.5502 | 20.5621 | 20.4822 | 20.2982 | 45.5984 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1141 | 3.1436 | 0.008279 | 0.4361 | 0.448 | 0.3681 | 0.1841
2012-10-30 14:24 | 23.4947 | 20.5651 | 20.5531 | 20.4852 | 20.3212 | 45.6025 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0771 | 3.1438 | 0.00828 | 0.488 | 0.476 | 0.4081 | 0.2441
2012-10-30 14:24 | 23.4999 | 20.5743 | 20.5304 | 20.4664 | 20.3024 | 45.5949 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014389 | 0.23455 | 0.003375 | 16.3008 | 20.0743 | 3.1433 | 0.008279 | 0.5 | 0.4561 | 0.3921 | 0.2281
2012-10-30 14:24 | 23.5059 | 20.5324 | 20.5044 | 20.4524 | 20.3084 | 45.6146 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.1323 | 3.1447 | 0.008287 | 0.4001 | 0.3721 | 0.3201 | 0.1761
2012-10-30 14:24 | 23.5044 | 20.5149 | 20.5708 | 20.4949 | 20.3069 | 45.5915 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0908 | 3.1431 | 0.008281 | 0.4241 | 0.48 | 0.4041 | 0.2161
2012-10-30 14:24 | 23.5089 | 20.5673 | 20.5514 | 20.4914 | 20.2634 | 45.5973 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.0913 | 3.1435 | 0.00828 | 0.476 | 0.4601 | 0.4001 | 0.1721
2012-10-30 14:24 | 23.5059 | 20.5164 | 20.5164 | 20.4484 | 20.3084 | 45.5961 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.1003 | 3.1434 | 0.008287 | 0.4161 | 0.4161 | 0.3481 | 0.2081
2012-10-30 14:24 | 23.5052 | 20.5916 | 20.5517 | 20.4797 | 20.2517 | 45.6181 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1036 | 3.1449 | 0.008286 | 0.488 | 0.4481 | 0.3761 | 0.1481
2012-10-30 14:24 | 23.5029 | 20.4774 | 20.5214 | 20.4414 | 20.2734 | 45.5909 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0813 | 3.143 | 0.00828 | 0.3961 | 0.4401 | 0.3601 | 0.1921
2012-10-30 14:24 | 23.4917 | 20.5302 | 20.5302 | 20.4182 | 20.2622 | 45.5955 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.1101 | 3.1434 | 0.008283 | 0.4201 | 0.4201 | 0.3081 | 0.1521
2012-10-30 14:24 | 23.5082 | 20.5387 | 20.5347 | 20.4587 | 20.2547 | 45.6053 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1026 | 3.144 | 0.008284 | 0.4361 | 0.4321 | 0.3561 | 0.1521
2012-10-30 14:25 | 23.5074 | 20.5339 | 20.5219 | 20.4619 | 20.2859 | 45.5863 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.0898 | 3.1427 | 0.008286 | 0.4441 | 0.4321 | 0.3721 | 0.1961
2012-10-30 14:25 | 23.5022 | 20.5047 | 20.4967 | 20.4207 | 20.2607 | 45.5845 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1607 | 3.1426 | 0.008283 | 0.344 | 0.336 | 0.26 | 0.1
2012-10-30 14:25 | 23.5097 | 20.5042 | 20.5322 | 20.4602 | 20.2362 | 45.5776 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1441 | 3.1421 | 0.00828 | 0.3601 | 0.3881 | 0.3161 | 0.0921
2012-10-30 14:25 | 23.5037 | 20.5382 | 20.5022 | 20.4582 | 20.2462 | 45.6019 | 0.014184 | 0.34544 | 0.0049 | 24.3543 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.1381 | 3.1438 | 0.008279 | 0.4001 | 0.3641 | 0.3201 | 0.1081
2012-10-30 14:25 | 23.5082 | 20.4867 | 20.5387 | 20.4467 | 20.2467 | 45.5828 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1026 | 3.1425 | 0.008284 | 0.3841 | 0.4361 | 0.3441 | 0.1441
2012-10-30 14:25 | 23.5014 | 20.4879 | 20.4799 | 20.3959 | 20.2519 | 45.5949 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1078 | 3.1433 | 0.008282 | 0.3801 | 0.3721 | 0.2881 | 0.1441
2012-10-30 14:25 | 23.5067 | 20.4972 | 20.5052 | 20.4292 | 20.2452 | 45.5967 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1131 | 3.1434 | 0.008283 | 0.3841 | 0.3921 | 0.3161 | 0.1321
2012-10-30 14:25 | 23.5074 | 20.5299 | 20.4939 | 20.4379 | 20.2499 | 45.6094 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.1539 | 3.1443 | 0.008279 | 0.376 | 0.34 | 0.284 | 0.096
2012-10-30 14:25 | 23.5082 | 20.5027 | 20.4787 | 20.4067 | 20.2587 | 45.5903 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.1587 | 3.143 | 0.008282 | 0.344 | 0.32 | 0.248 | 0.1
2012-10-30 14:25 | 23.5067 | 20.4932 | 20.5172 | 20.4092 | 20.2532 | 45.5799 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1251 | 3.1423 | 0.008282 | 0.3681 | 0.3921 | 0.2841 | 0.1281
2012-10-30 14:25 | 23.5089 | 20.5314 | 20.4434 | 20.4114 | 20.2234 | 45.6059 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1313 | 3.1441 | 0.008284 | 0.4001 | 0.3121 | 0.2801 | 0.0921
2012-10-30 14:25 | 23.5067 | 20.4612 | 20.4612 | 20.4172 | 20.2532 | 45.5955 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.1171 | 3.1434 | 0.008283 | 0.3441 | 0.3441 | 0.3001 | 0.1361
2012-10-30 14:25 | 23.5082 | 20.4787 | 20.4747 | 20.3947 | 20.2747 | 45.5955 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0706 | 3.1434 | 0.008285 | 0.4081 | 0.4041 | 0.3241 | 0.2041
2012-10-30 14:25 | 23.5112 | 20.4737 | 20.4817 | 20.4337 | 20.2177 | 45.5973 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.0896 | 3.1435 | 0.008283 | 0.3841 | 0.3921 | 0.3441 | 0.1281
2012-10-30 14:25 | 23.5216 | 20.5161 | 20.4641 | 20.4121 | 20.2241 | 45.5874 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.096 | 3.1428 | 0.008283 | 0.4201 | 0.3681 | 0.3161 | 0.1281
2012-10-30 14:25 | 23.5194 | 20.4459 | 20.4699 | 20.3859 | 20.2139 | 45.5765 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.0978 | 3.142 | 0.008282 | 0.3481 | 0.3721 | 0.2881 | 0.1161
2012-10-30 14:25 | 23.5164 | 20.4869 | 20.4269 | 20.4549 | 20.2509 | 45.5892 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.1388 | 3.1429 | 0.00828 | 0.3481 | 0.2881 | 0.3161 | 0.1121
2012-10-30 14:25 | 23.5216 | 20.4521 | 20.4801 | 20.3641 | 20.2041 | 45.5978 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.116 | 3.1435 | 0.008288 | 0.3361 | 0.3641 | 0.2481 | 0.0881
2012-10-30 14:25 | 23.5209 | 20.4274 | 20.4674 | 20.3754 | 20.2074 | 45.6088 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1113 | 3.1443 | 0.008283 | 0.3161 | 0.3561 | 0.2641 | 0.0961
2012-10-30 14:25 | 23.5082 | 20.4467 | 20.4267 | 20.3747 | 20.1987 | 45.5915 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1066 | 3.1431 | 0.008285 | 0.3401 | 0.3201 | 0.2681 | 0.0921
2012-10-30 14:25 | 23.5164 | 20.4549 | 20.4469 | 20.3789 | 20.1789 | 45.588 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0908 | 3.1428 | 0.008283 | 0.3641 | 0.3561 | 0.2881 | 0.0881
2012-10-30 14:25 | 23.5149 | 20.4454 | 20.4134 | 20.3654 | 20.1974 | 45.5967 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1053 | 3.1434 | 0.008285 | 0.3401 | 0.3081 | 0.2601 | 0.0921
2012-10-30 14:25 | 23.5261 | 20.4686 | 20.4166 | 20.3886 | 20.2406 | 45.5817 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1325 | 3.1424 | 0.008285 | 0.3361 | 0.2841 | 0.2561 | 0.1081
2012-10-30 14:26 | 23.5201 | 20.4466 | 20.4426 | 20.3746 | 20.2186 | 45.5834 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1585 | 3.1425 | 0.008283 | 0.2881 | 0.2841 | 0.2161 | 0.0601
2012-10-30 14:26 | 23.5239 | 20.4104 | 20.4424 | 20.3704 | 20.1944 | 45.5996 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.1383 | 3.1436 | 0.00828 | 0.2721 | 0.3041 | 0.2321 | 0.0561
2012-10-30 14:26 | 23.5291 | 20.4676 | 20.4196 | 20.3876 | 20.1956 | 45.5834 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1075 | 3.1425 | 0.008279 | 0.3601 | 0.3121 | 0.2801 | 0.0881
2012-10-30 14:26 | 23.5186 | 20.4331 | 20.4451 | 20.3651 | 20.2011 | 45.5845 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23444 | 0.003378 | 16.2712 | 20.133 | 3.1426 | 0.008284 | 0.3001 | 0.3121 | 0.2321 | 0.0681
2012-10-30 14:26 | 23.5321 | 20.4546 | 20.4186 | 20.3826 | 20.2026 | 45.5765 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1305 | 3.142 | 0.008287 | 0.3241 | 0.2881 | 0.2521 | 0.0721
2012-10-30 14:26 | 23.5209 | 20.4754 | 20.4274 | 20.3594 | 20.1954 | 45.5915 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.0953 | 3.1431 | 0.008282 | 0.3801 | 0.3321 | 0.2641 | 0.1001
2012-10-30 14:26 | 23.5254 | 20.4239 | 20.4199 | 20.3639 | 20.1679 | 45.5828 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.0998 | 3.1425 | 0.008283 | 0.3241 | 0.3201 | 0.2641 | 0.0681
2012-10-30 14:26 | 23.5284 | 20.3989 | 20.4589 | 20.3469 | 20.1829 | 45.588 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1228 | 3.1428 | 0.008283 | 0.2761 | 0.3361 | 0.2241 | 0.0601
2012-10-30 14:26 | 23.5343 | 20.4208 | 20.4248 | 20.3328 | 20.1808 | 45.5874 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1607 | 3.1428 | 0.008282 | 0.2601 | 0.2641 | 0.1721 | 0.0201
2012-10-30 14:26 | 23.5358 | 20.3943 | 20.4943 | 20.3703 | 20.1502 | 45.5822 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.1502 | 3.1424 | 0.008281 | 0.2441 | 0.3441 | 0.2201 | 0
2012-10-30 14:26 | 23.5366 | 20.4751 | 20.4551 | 20.3591 | 20.1951 | 45.5759 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.095 | 3.142 | 0.008283 | 0.3801 | 0.3601 | 0.2641 | 0.1001
2012-10-30 14:26 | 23.5284 | 20.4029 | 20.4109 | 20.3269 | 20.1749 | 45.5672 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1388 | 3.1414 | 0.008286 | 0.2641 | 0.2721 | 0.1881 | 0.0361
2012-10-30 14:26 | 23.5261 | 20.4046 | 20.4006 | 20.3366 | 20.1886 | 45.5892 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.1165 | 3.1429 | 0.008284 | 0.2881 | 0.2841 | 0.2201 | 0.0721
2012-10-30 14:26 | 23.5373 | 20.4398 | 20.4158 | 20.3678 | 20.1677 | 45.588 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.1397 | 3.1428 | 0.008285 | 0.3001 | 0.2761 | 0.2281 | 0.028
2012-10-30 14:26 | 23.5291 | 20.4236 | 20.4196 | 20.3076 | 20.1836 | 45.5817 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.1395 | 3.1424 | 0.008284 | 0.2841 | 0.2801 | 0.1681 | 0.0441
2012-10-30 14:26 | 23.5351 | 20.4096 | 20.4416 | 20.3296 | 20.1495 | 45.5765 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.1215 | 3.142 | 0.008279 | 0.2881 | 0.3201 | 0.2081 | 0.028
2012-10-30 14:26 | 23.5314 | 20.4099 | 20.4019 | 20.2979 | 20.1939 | 45.5817 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1218 | 3.1424 | 0.008284 | 0.2881 | 0.2801 | 0.1761 | 0.0721
2012-10-30 14:26 | 23.5343 | 20.3928 | 20.4168 | 20.3208 | 20.1808 | 45.5926 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1327 | 3.1432 | 0.008276 | 0.2601 | 0.2841 | 0.1881 | 0.0481
2012-10-30 14:26 | 23.5441 | 20.4026 | 20.4266 | 20.3106 | 20.1825 | 45.5712 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.1505 | 3.1417 | 0.008287 | 0.2521 | 0.2761 | 0.1601 | 0.032
2012-10-30 14:26 | 23.5291 | 20.3876 | 20.4356 | 20.2796 | 20.1635 | 45.5736 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.1195 | 3.1418 | 0.008281 | 0.2681 | 0.3161 | 0.1601 | 0.044
2012-10-30 14:26 | 23.5456 | 20.4281 | 20.4041 | 20.3321 | 20.2081 | 45.5845 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.14 | 3.1426 | 0.008279 | 0.2881 | 0.2641 | 0.1921 | 0.0681
2012-10-30 14:26 | 23.5388 | 20.3573 | 20.3813 | 20.3053 | 20.2013 | 45.5886 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 20.1052 | 3.1429 | 0.008287 | 0.2521 | 0.2761 | 0.2001 | 0.0961
2012-10-30 14:26 | 23.5433 | 20.3738 | 20.3778 | 20.2898 | 20.1817 | 45.573 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.1297 | 3.1418 | 0.008282 | 0.2441 | 0.2481 | 0.1601 | 0.052
2012-10-30 14:26 | 23.5426 | 20.3691 | 20.3971 | 20.2971 | 20.165 | 45.5684 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.129 | 3.1415 | 0.008279 | 0.2401 | 0.2681 | 0.1681 | 0.036
2012-10-30 14:27 | 23.5478 | 20.3903 | 20.4063 | 20.3383 | 20.1462 | 45.5926 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014382 | 0.23452 | 0.003373 | 16.3067 | 20.1462 | 3.1432 | 0.008277 | 0.2441 | 0.2601 | 0.1921 | 0
2012-10-30 14:27 | 23.5493 | 20.3998 | 20.4118 | 20.2998 | 20.1517 | 45.5736 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1197 | 3.1418 | 0.00828 | 0.2801 | 0.2921 | 0.1801 | 0.032
2012-10-30 14:27 | 23.5508 | 20.3533 | 20.3733 | 20.3093 | 20.1332 | 45.577 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0892 | 3.1421 | 0.008285 | 0.2641 | 0.2841 | 0.2201 | 0.044
2012-10-30 14:27 | 23.5493 | 20.4038 | 20.3918 | 20.2678 | 20.1437 | 45.5585 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1197 | 3.1408 | 0.008283 | 0.2841 | 0.2721 | 0.1481 | 0.024
2012-10-30 14:27 | 23.556 | 20.3625 | 20.3785 | 20.3305 | 20.1584 | 45.5886 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1064 | 3.1429 | 0.008284 | 0.2561 | 0.2721 | 0.2241 | 0.052
2012-10-30 14:27 | 23.556 | 20.3825 | 20.3465 | 20.2945 | 20.1544 | 45.5805 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.1104 | 3.1423 | 0.008282 | 0.2721 | 0.2361 | 0.1841 | 0.044
2012-10-30 14:27 | 23.553 | 20.3715 | 20.3915 | 20.2995 | 20.1394 | 45.5753 | 0.01419 | 0.3455 | 0.004903 | 24.3481 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1074 | 3.142 | 0.008281 | 0.2641 | 0.2841 | 0.1921 | 0.032
2012-10-30 14:27 | 23.5628 | 20.3373 | 20.3613 | 20.2573 | 20.1332 | 45.573 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1332 | 3.1418 | 0.008284 | 0.2041 | 0.2281 | 0.1241 | 0
2012-10-30 14:27 | 23.5628 | 20.3693 | 20.3733 | 20.2533 | 20.1372 | 45.5851 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014389 | 0.23447 | 0.003374 | 16.2947 | 20.1252 | 3.1426 | 0.008275 | 0.2441 | 0.2481 | 0.1281 | 0.012
2012-10-30 14:27 | 23.5628 | 20.3053 | 20.4173 | 20.3093 | 20.1572 | 45.5793 | 0.014187 | 0.34537 | 0.0049 | 24.344 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.1372 | 3.1422 | 0.00828 | 0.1681 | 0.2801 | 0.1721 | 0.02
2012-10-30 14:27 | 23.5568 | 20.3313 | 20.3713 | 20.2833 | 20.1312 | 45.5672 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1072 | 3.1414 | 0.008284 | 0.2241 | 0.2641 | 0.1761 | 0.024
2012-10-30 14:27 | 23.5628 | 20.3493 | 20.3693 | 20.2853 | 20.1692 | 45.566 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1612 | 3.1413 | 0.00828 | 0.1881 | 0.2081 | 0.1241 | 0.008
2012-10-30 14:27 | 23.5635 | 20.318 | 20.35 | 20.266 | 20.1459 | 45.5765 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1059 | 3.142 | 0.008281 | 0.2121 | 0.2441 | 0.1601 | 0.04
2012-10-30 14:27 | 23.5695 | 20.34 | 20.384 | 20.288 | 20.1639 | 45.5655 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1719 | 3.1413 | 0.008283 | 0.1681 | 0.2121 | 0.1161 | -0.008
2012-10-30 14:27 | 23.5658 | 20.3203 | 20.3683 | 20.2723 | 20.1282 | 45.5689 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1482 | 3.1415 | 0.008285 | 0.1721 | 0.2201 | 0.1241 | -0.02
2012-10-30 14:27 | 23.559 | 20.3495 | 20.3535 | 20.2535 | 20.1294 | 45.5655 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0974 | 3.1413 | 0.008284 | 0.2521 | 0.2561 | 0.1561 | 0.032
2012-10-30 14:27 | 23.5672 | 20.3017 | 20.3817 | 20.2457 | 20.1576 | 45.5753 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014395 | 0.23444 | 0.003375 | 16.2859 | 20.1336 | 3.142 | 0.008285 | 0.1681 | 0.2481 | 0.1121 | 0.024
2012-10-30 14:27 | 23.5747 | 20.3652 | 20.3652 | 20.2532 | 20.1371 | 45.5603 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1131 | 3.1409 | 0.008286 | 0.2521 | 0.2521 | 0.1401 | 0.024
2012-10-30 14:27 | 23.5755 | 20.31 | 20.394 | 20.29 | 20.1259 | 45.5707 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1459 | 3.1416 | 0.008282 | 0.1641 | 0.2481 | 0.1441 | -0.02
2012-10-30 14:27 | 23.5725 | 20.319 | 20.371 | 20.251 | 20.1229 | 45.5695 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1749 | 3.1416 | 0.008284 | 0.1441 | 0.1961 | 0.0761 | -0.052
2012-10-30 14:27 | 23.577 | 20.3475 | 20.3715 | 20.2715 | 20.1634 | 45.5736 | 0.014197 | 0.34538 | 0.004903 | 24.3279 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1634 | 3.1418 | 0.00828 | 0.1841 | 0.2081 | 0.1081 | 0
2012-10-30 14:27 | 23.5732 | 20.2837 | 20.3637 | 20.2877 | 20.1276 | 45.5568 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014401 | 0.23441 | 0.003376 | 16.2771 | 20.1516 | 3.1407 | 0.00828 | 0.1321 | 0.2121 | 0.1361 | -0.024
2012-10-30 14:27 | 23.5755 | 20.302 | 20.358 | 20.258 | 20.1299 | 45.5765 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1779 | 3.142 | 0.008282 | 0.1241 | 0.1801 | 0.0801 | -0.048
2012-10-30 14:28 | 23.5815 | 20.312 | 20.368 | 20.244 | 20.1479 | 45.5712 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1399 | 3.1417 | 0.008283 | 0.1721 | 0.2281 | 0.1041 | 0.008
2012-10-30 14:28 | 23.5785 | 20.309 | 20.357 | 20.285 | 20.1129 | 45.5608 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1249 | 3.141 | 0.008283 | 0.1841 | 0.2321 | 0.1601 | -0.012
2012-10-30 14:28 | 23.5687 | 20.3112 | 20.3352 | 20.2272 | 20.1191 | 45.562 | 0.01419 | 0.3455 | 0.004903 | 24.3481 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1111 | 3.141 | 0.00828 | 0.2001 | 0.2241 | 0.1161 | 0.008
2012-10-30 14:28 | 23.5859 | 20.3324 | 20.3284 | 20.2284 | 20.1443 | 45.5712 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1603 | 3.1417 | 0.008277 | 0.1721 | 0.1681 | 0.0681 | -0.016
2012-10-30 14:28 | 23.5822 | 20.2847 | 20.3567 | 20.2327 | 20.1046 | 45.5799 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1526 | 3.1423 | 0.008287 | 0.1321 | 0.2041 | 0.0801 | -0.048
2012-10-30 14:28 | 23.5859 | 20.2724 | 20.3884 | 20.2203 | 20.1163 | 45.5545 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.1443 | 3.1405 | 0.008284 | 0.1281 | 0.2441 | 0.076 | -0.028
2012-10-30 14:28 | 23.5829 | 20.2774 | 20.3734 | 20.2814 | 20.1093 | 45.5585 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.1533 | 3.1408 | 0.008286 | 0.1241 | 0.2201 | 0.1281 | -0.044
2012-10-30 14:28 | 23.5889 | 20.2794 | 20.3754 | 20.2354 | 20.1033 | 45.577 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.1113 | 3.1421 | 0.008283 | 0.1681 | 0.2641 | 0.1241 | -0.008
2012-10-30 14:28 | 23.5859 | 20.2724 | 20.3764 | 20.2364 | 20.0923 | 45.558 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.01442 | 0.23448 | 0.003381 | 16.2607 | 20.1523 | 3.1408 | 0.008288 | 0.1201 | 0.2241 | 0.0841 | -0.06
2012-10-30 14:28 | 23.5859 | 20.2724 | 20.3524 | 20.2364 | 20.1083 | 45.5533 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1563 | 3.1404 | 0.008281 | 0.1161 | 0.1961 | 0.0801 | -0.048
2012-10-30 14:28 | 23.5889 | 20.2714 | 20.3194 | 20.2314 | 20.0913 | 45.5655 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.1193 | 3.1413 | 0.008289 | 0.1521 | 0.2001 | 0.1121 | -0.028
2012-10-30 14:28 | 23.5867 | 20.2492 | 20.3412 | 20.2332 | 20.1091 | 45.562 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.1251 | 3.141 | 0.008278 | 0.1241 | 0.2161 | 0.1081 | -0.016
2012-10-30 14:28 | 23.5934 | 20.2879 | 20.3359 | 20.2519 | 20.1118 | 45.558 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1158 | 3.1408 | 0.008287 | 0.1721 | 0.2201 | 0.1361 | -0.004
2012-10-30 14:28 | 23.5912 | 20.2617 | 20.3057 | 20.2176 | 20.0976 | 45.5689 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1296 | 3.1415 | 0.008281 | 0.1321 | 0.1761 | 0.088 | -0.032
2012-10-30 14:28 | 23.5919 | 20.2624 | 20.3344 | 20.2384 | 20.0582 | 45.5516 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1143 | 3.1403 | 0.008282 | 0.1481 | 0.2201 | 0.1241 | -0.0561
2012-10-30 14:28 | 23.5919 | 20.2624 | 20.3184 | 20.2143 | 20.0902 | 45.5568 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1143 | 3.1407 | 0.008282 | 0.1481 | 0.2041 | 0.1 | -0.0241
2012-10-30 14:28 | 23.5822 | 20.2727 | 20.3527 | 20.2327 | 20.0805 | 45.5551 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1366 | 3.1406 | 0.008285 | 0.1361 | 0.2161 | 0.0961 | -0.0561
2012-10-30 14:28 | 23.5957 | 20.2582 | 20.3302 | 20.2021 | 20.1141 | 45.5597 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1661 | 3.1409 | 0.008283 | 0.0921 | 0.1641 | 0.036 | -0.052
2012-10-30 14:28 | 23.5942 | 20.2567 | 20.3567 | 20.2206 | 20.1006 | 45.5487 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1246 | 3.1401 | 0.008283 | 0.1321 | 0.2321 | 0.096 | -0.024
2012-10-30 14:28 | 23.5867 | 20.2332 | 20.3252 | 20.2091 | 20.081 | 45.5585 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.1291 | 3.1408 | 0.008288 | 0.1041 | 0.1961 | 0.08 | -0.0481
2012-10-30 14:28 | 23.5904 | 20.2689 | 20.3529 | 20.2088 | 20.0968 | 45.5637 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.1248 | 3.1412 | 0.008284 | 0.1441 | 0.2281 | 0.084 | -0.028
2012-10-30 14:28 | 23.5957 | 20.2341 | 20.3382 | 20.2261 | 20.1061 | 45.5562 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1501 | 3.1406 | 0.008282 | 0.084 | 0.1881 | 0.076 | -0.044
2012-10-30 14:28 | 23.5972 | 20.2997 | 20.3157 | 20.2116 | 20.1076 | 45.5406 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.1596 | 3.1396 | 0.00829 | 0.1401 | 0.1561 | 0.052 | -0.052
2012-10-30 14:29 | 23.5867 | 20.2532 | 20.3452 | 20.1771 | 20.065 | 45.5678 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1291 | 3.1414 | 0.008285 | 0.1241 | 0.2161 | 0.048 | -0.0641
2012-10-30 14:29 | 23.5987 | 20.2371 | 20.3132 | 20.2291 | 20.1051 | 45.5545 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.1331 | 3.1405 | 0.008285 | 0.104 | 0.1801 | 0.096 | -0.028
2012-10-30 14:29 | 23.5874 | 20.2419 | 20.3139 | 20.1458 | 20.0377 | 45.551 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1578 | 3.1403 | 0.008281 | 0.0841 | 0.1561 | -0.012 | -0.1201
2012-10-30 14:29 | 23.5889 | 20.2273 | 20.2714 | 20.1993 | 20.0752 | 45.5747 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014423 | 0.23444 | 0.003381 | 16.2543 | 20.1433 | 3.1419 | 0.008287 | 0.084 | 0.1281 | 0.056 | -0.0681
2012-10-30 14:29 | 23.5964 | 20.2348 | 20.2949 | 20.1868 | 20.0747 | 45.5539 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1468 | 3.1405 | 0.008284 | 0.088 | 0.1481 | 0.04 | -0.0721
2012-10-30 14:29 | 23.5897 | 20.2322 | 20.3042 | 20.1721 | 20.056 | 45.5556 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1401 | 3.1406 | 0.008286 | 0.0921 | 0.1641 | 0.032 | -0.0841
2012-10-30 14:29 | 23.5919 | 20.2063 | 20.3304 | 20.2063 | 20.0462 | 45.5464 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.1303 | 3.14 | 0.008281 | 0.076 | 0.2001 | 0.076 | -0.0841
2012-10-30 14:29 | 23.5994 | 20.2258 | 20.3059 | 20.1738 | 20.0817 | 45.5395 | 0.014184 | 0.34547 | 0.0049 | 24.3563 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1378 | 3.1395 | 0.008278 | 0.088 | 0.1681 | 0.036 | -0.0561
2012-10-30 14:29 | 23.5919 | 20.2464 | 20.3224 | 20.1583 | 20.0582 | 45.5603 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1103 | 3.1409 | 0.008284 | 0.1361 | 0.2121 | 0.048 | -0.0521
2012-10-30 14:29 | 23.5882 | 20.1946 | 20.2827 | 20.1866 | 20.0825 | 45.5377 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1586 | 3.1394 | 0.008287 | 0.036 | 0.1241 | 0.028 | -0.0761
2012-10-30 14:29 | 23.6031 | 20.2415 | 20.3176 | 20.1495 | 20.0494 | 45.5406 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.1615 | 3.1396 | 0.008285 | 0.08 | 0.1561 | -0.012 | -0.1121
2012-10-30 14:29 | 23.5987 | 20.2171 | 20.2972 | 20.1691 | 20.021 | 45.5637 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.1731 | 3.1412 | 0.008281 | 0.044 | 0.1241 | -0.004 | -0.1521
2012-10-30 14:29 | 23.6061 | 20.2486 | 20.2926 | 20.1645 | 20.0564 | 45.566 | 0.014194 | 0.34537 | 0.004902 | 24.332 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1565 | 3.1413 | 0.008281 | 0.0921 | 0.1361 | 0.008 | -0.1001
2012-10-30 14:29 | 23.6046 | 20.227 | 20.2911 | 20.187 | 20.0509 | 45.5452 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.203 | 3.1399 | 0.008284 | 0.024 | 0.0881 | -0.016 | -0.1521
2012-10-30 14:29 | 23.6069 | 20.1853 | 20.2574 | 20.1613 | 20.0212 | 45.5568 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1373 | 3.1407 | 0.008286 | 0.048 | 0.1201 | 0.024 | -0.1161
2012-10-30 14:29 | 23.5972 | 20.1876 | 20.2677 | 20.1156 | 20.0435 | 45.5424 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.1196 | 3.1397 | 0.008282 | 0.068 | 0.1481 | -0.004 | -0.0761
2012-10-30 14:29 | 23.5994 | 20.1778 | 20.2979 | 20.1578 | 20.0737 | 45.5418 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1418 | 3.1397 | 0.008283 | 0.036 | 0.1561 | 0.016 | -0.0681
2012-10-30 14:29 | 23.6069 | 20.1933 | 20.2574 | 20.1173 | 20.0011 | 45.5666 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1333 | 3.1414 | 0.008284 | 0.06 | 0.1241 | -0.016 | -0.1322
2012-10-30 14:30 | 23.6129 | 20.1833 | 20.2874 | 20.1393 | 20.0272 | 45.5412 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1553 | 3.1396 | 0.008286 | 0.028 | 0.1321 | -0.016 | -0.1281
2012-10-30 14:30 | 23.5987 | 20.1851 | 20.2211 | 20.1651 | 19.9889 | 45.5545 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.1371 | 3.1405 | 0.008287 | 0.048 | 0.084 | 0.028 | -0.1482
2012-10-30 14:30 | 23.6188 | 20.1532 | 20.2813 | 20.1652 | 20.0571 | 45.5626 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1292 | 3.1411 | 0.008284 | 0.024 | 0.1521 | 0.036 | -0.0721
2012-10-30 14:30 | 23.6203 | 20.1827 | 20.2467 | 20.1587 | 20.0185 | 45.5458 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1467 | 3.1399 | 0.008285 | 0.036 | 0.1 | 0.012 | -0.1282
2012-10-30 14:30 | 23.6203 | 20.1947 | 20.2467 | 20.1427 | 20.0266 | 45.5412 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.1146 | 3.1396 | 0.008285 | 0.0801 | 0.1321 | 0.0281 | -0.088
2012-10-30 14:30 | 23.6181 | 20.1965 | 20.2205 | 20.1485 | 20.0284 | 45.558 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1645 | 3.1408 | 0.008281 | 0.032 | 0.056 | -0.016 | -0.1361
2012-10-30 14:30 | 23.6196 | 20.142 | 20.2741 | 20.1179 | 20.0259 | 45.5516 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014404 | 0.23442 | 0.003377 | 16.2747 | 20.0939 | 3.1403 | 0.008284 | 0.0481 | 0.1802 | 0.024 | -0.068
2012-10-30 14:30 | 23.6218 | 20.1722 | 20.2162 | 20.1201 | 19.996 | 45.547 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1201 | 3.14 | 0.008282 | 0.0521 | 0.0961 | 0 | -0.1241
2012-10-30 14:30 | 23.6353 | 20.1817 | 20.2457 | 20.1377 | 20.0255 | 45.551 | 0.014206 | 0.34537 | 0.004906 | 24.3115 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1176 | 3.1403 | 0.008284 | 0.0641 | 0.1281 | 0.0201 | -0.0921
2012-10-30 14:30 | 23.6293 | 20.1797 | 20.2517 | 20.1677 | 19.9795 | 45.5481 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.1036 | 3.1401 | 0.008283 | 0.0761 | 0.1481 | 0.0641 | -0.1241
2012-10-30 14:30 | 23.6293 | 20.1597 | 20.2197 | 20.1437 | 20.0115 | 45.5366 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1076 | 3.1393 | 0.00828 | 0.0521 | 0.1121 | 0.0361 | -0.0961
2012-10-30 14:30 | 23.6301 | 20.1765 | 20.2605 | 20.1164 | 20.0083 | 45.5343 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.1325 | 3.1391 | 0.008284 | 0.044 | 0.128 | -0.0161 | -0.1242
2012-10-30 14:30 | 23.6338 | 20.1442 | 20.2682 | 20.1241 | 20 | 45.5568 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1081 | 3.1407 | 0.008282 | 0.0361 | 0.1601 | 0.016 | -0.1081
2012-10-30 14:30 | 23.6353 | 20.1577 | 20.2177 | 20.1497 | 20.0375 | 45.5412 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.0736 | 3.1396 | 0.008282 | 0.0841 | 0.1441 | 0.0761 | -0.0361
2012-10-30 14:30 | 23.6345 | 20.1649 | 20.2529 | 20.0808 | 20.0207 | 45.5412 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014389 | 0.23445 | 0.003374 | 16.2937 | 20.0648 | 3.1396 | 0.008277 | 0.1001 | 0.1881 | 0.016 | -0.0441
2012-10-30 14:30 | 23.633 | 20.1233 | 20.2074 | 20.0753 | 19.9712 | 45.5597 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.0673 | 3.1409 | 0.008284 | 0.056 | 0.1401 | 0.008 | -0.0961
2012-10-30 14:30 | 23.6383 | 20.1847 | 20.2487 | 20.1246 | 20.0005 | 45.5366 | 0.01419 | 0.3455 | 0.004903 | 24.3481 | 0.01442 | 0.23448 | 0.003381 | 16.2607 | 20.0806 | 3.1393 | 0.008284 | 0.1041 | 0.1681 | 0.044 | -0.0801
2012-10-30 14:30 | 23.642 | 20.1604 | 20.2164 | 20.1323 | 20.0162 | 45.5424 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.0843 | 3.1397 | 0.008287 | 0.0761 | 0.1321 | 0.048 | -0.0681
2012-10-30 14:30 | 23.639 | 20.1414 | 20.2254 | 20.0973 | 19.9732 | 45.5585 | 0.014216 | 0.34541 | 0.00491 | 24.2974 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0893 | 3.1408 | 0.008287 | 0.0521 | 0.1361 | 0.008 | -0.1161
2012-10-30 14:30 | 23.642 | 20.1444 | 20.2244 | 20.1283 | 19.9842 | 45.5291 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.1123 | 3.1388 | 0.008285 | 0.0321 | 0.1121 | 0.016 | -0.1281
2012-10-30 14:30 | 23.642 | 20.2244 | 20.2604 | 20.0843 | 19.9842 | 45.5296 | 0.014187 | 0.34543 | 0.004901 | 24.3481 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.0963 | 3.1388 | 0.00828 | 0.1281 | 0.1641 | -0.012 | -0.1121
2012-10-30 14:30 | 23.6473 | 20.1336 | 20.2297 | 20.0936 | 20.0055 | 45.5429 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1096 | 3.1397 | 0.008287 | 0.024 | 0.1201 | -0.016 | -0.1041
2012-10-30 14:30 | 23.6488 | 20.1191 | 20.2112 | 20.1111 | 19.971 | 45.54 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1231 | 3.1395 | 0.008281 | -0.004 | 0.0881 | -0.012 | -0.1521
2012-10-30 14:30 | 23.6465 | 20.1569 | 20.2369 | 20.1368 | 20.0327 | 45.5273 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1328 | 3.1387 | 0.008284 | 0.0241 | 0.1041 | 0.004 | -0.1001
2012-10-30 14:31 | 23.6473 | 20.1456 | 20.2537 | 20.0936 | 19.9775 | 45.54 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1016 | 3.1395 | 0.008282 | 0.044 | 0.1521 | -0.008 | -0.1241
2012-10-30 14:31 | 23.6532 | 20.1275 | 20.2036 | 20.0915 | 20.0234 | 45.5424 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.1435 | 3.1397 | 0.00828 | -0.016 | 0.0601 | -0.052 | -0.1201
2012-10-30 14:31 | 23.6577 | 20.144 | 20.2241 | 20.128 | 19.9999 | 45.5343 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.116 | 3.1391 | 0.008286 | 0.028 | 0.1081 | 0.012 | -0.1161
2012-10-30 14:31 | 23.6637 | 20.098 | 20.2061 | 20.098 | 20.0219 | 45.5574 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.122 | 3.1407 | 0.008285 | -0.024 | 0.0841 | -0.024 | -0.1001
2012-10-30 14:31 | 23.651 | 20.1333 | 20.2214 | 20.0733 | 20.0132 | 45.5343 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1053 | 3.1391 | 0.008283 | 0.028 | 0.1161 | -0.032 | -0.0921
2012-10-30 14:31 | 23.6652 | 20.1756 | 20.2516 | 20.1035 | 19.9914 | 45.5325 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.0955 | 3.139 | 0.008286 | 0.0801 | 0.1561 | 0.008 | -0.1041
2012-10-30 14:31 | 23.6585 | 20.1328 | 20.2209 | 20.0648 | 19.9927 | 45.5528 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.1368 | 3.1404 | 0.008287 | -0.004 | 0.0841 | -0.072 | -0.1441
2012-10-30 14:31 | 23.6652 | 20.1115 | 20.2316 | 20.0514 | 19.9793 | 45.5412 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1595 | 3.1396 | 0.008288 | -0.048 | 0.0721 | -0.1081 | -0.1802
2012-10-30 14:31 | 23.6592 | 20.1335 | 20.2256 | 20.0895 | 19.9974 | 45.5291 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.1095 | 3.1388 | 0.008277 | 0.024 | 0.1161 | -0.02 | -0.1121
2012-10-30 14:31 | 23.6532 | 20.1075 | 20.2556 | 20.0835 | 19.9713 | 45.5458 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0955 | 3.1399 | 0.008286 | 0.012 | 0.1601 | -0.012 | -0.1242
2012-10-30 14:31 | 23.6592 | 20.1736 | 20.2216 | 20.0775 | 20.0014 | 45.5285 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1455 | 3.1387 | 0.008284 | 0.0281 | 0.0761 | -0.068 | -0.1441
2012-10-30 14:31 | 23.6622 | 20.1285 | 20.1806 | 20.1085 | 19.9523 | 45.5273 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.1045 | 3.1387 | 0.00828 | 0.024 | 0.0761 | 0.004 | -0.1522
2012-10-30 14:31 | 23.6622 | 20.1485 | 20.2566 | 20.0564 | 19.9643 | 45.5464 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 20.1285 | 3.14 | 0.008281 | 0.02 | 0.1281 | -0.0721 | -0.1642
2012-10-30 14:31 | 23.6607 | 20.107 | 20.2151 | 20.075 | 19.9708 | 45.5187 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.01442 | 0.23457 | 0.003382 | 16.2667 | 20.079 | 3.1381 | 0.008288 | 0.028 | 0.1361 | -0.004 | -0.1082
2012-10-30 14:31 | 23.6562 | 20.0865 | 20.1946 | 20.0785 | 19.9623 | 45.5331 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.1065 | 3.1391 | 0.008287 | -0.02 | 0.0881 | -0.028 | -0.1442
2012-10-30 14:31 | 23.6547 | 20.101 | 20.2131 | 20.065 | 19.9488 | 45.5458 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.117 | 3.1399 | 0.008288 | -0.016 | 0.0961 | -0.052 | -0.1682
2012-10-30 14:31 | 23.6577 | 20.132 | 20.2041 | 20.08 | 19.9478 | 45.5198 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.136 | 3.1381 | 0.008284 | -0.004 | 0.0681 | -0.056 | -0.1882
2012-10-30 14:31 | 23.6532 | 20.1075 | 20.2436 | 20.0715 | 19.9633 | 45.5458 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.1315 | 3.1399 | 0.008279 | -0.024 | 0.1121 | -0.06 | -0.1682
2012-10-30 14:31 | 23.6592 | 20.1175 | 20.2536 | 20.0695 | 19.9573 | 45.5308 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1415 | 3.1389 | 0.008285 | -0.024 | 0.1121 | -0.072 | -0.1842
2012-10-30 14:31 | 23.6473 | 20.1096 | 20.1697 | 20.0696 | 19.9294 | 45.5319 | 0.014187 | 0.34543 | 0.004901 | 24.3481 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1256 | 3.139 | 0.008278 | -0.016 | 0.0441 | -0.056 | -0.1962
2012-10-30 14:31 | 23.6532 | 20.1515 | 20.1996 | 20.0434 | 19.9393 | 45.5377 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.1195 | 3.1394 | 0.008287 | 0.032 | 0.0801 | -0.0761 | -0.1802
2012-10-30 14:31 | 23.6577 | 20.12 | 20.2161 | 20.0479 | 19.9718 | 45.5348 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.156 | 3.1392 | 0.00829 | -0.036 | 0.0601 | -0.1081 | -0.1842
2012-10-30 14:31 | 23.6562 | 20.0945 | 20.1706 | 20.0464 | 19.9503 | 45.5366 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.01442 | 0.23448 | 0.003381 | 16.2607 | 20.1385 | 3.1393 | 0.008289 | -0.044 | 0.0321 | -0.0921 | -0.1882
2012-10-30 14:31 | 23.6532 | 20.0995 | 20.2116 | 20.0234 | 19.9754 | 45.5366 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.1075 | 3.1393 | 0.008285 | -0.008 | 0.1041 | -0.0841 | -0.1321
2012-10-30 14:31 | 23.66 | 20.0983 | 20.1784 | 20.0502 | 19.9661 | 45.5262 | 0.014209 | 0.34553 | 0.00491 | 24.3175 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1423 | 3.1386 | 0.008287 | -0.044 | 0.0361 | -0.0921 | -0.1762
2012-10-30 14:32 | 23.6637 | 20.122 | 20.162 | 20.0659 | 19.9738 | 45.5372 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.13 | 3.1393 | 0.008287 | -0.008 | 0.032 | -0.0641 | -0.1562
2012-10-30 14:32 | 23.651 | 20.0733 | 20.1734 | 20.0252 | 19.9651 | 45.5389 | 0.014212 | 0.34556 | 0.004911 | 24.3145 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1333 | 3.1395 | 0.008289 | -0.06 | 0.0401 | -0.1081 | -0.1682
2012-10-30 14:32 | 23.6592 | 20.1455 | 20.1816 | 20.0534 | 19.9173 | 45.5389 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1095 | 3.1395 | 0.008284 | 0.036 | 0.0721 | -0.0561 | -0.1922
2012-10-30 14:32 | 23.6667 | 20.133 | 20.161 | 20.0249 | 19.9248 | 45.5331 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014373 | 0.23448 | 0.00337 | 16.3139 | 20.121 | 3.1391 | 0.008275 | 0.012 | 0.04 | -0.0961 | -0.1962
2012-10-30 14:32 | 23.6585 | 20.1208 | 20.1889 | 20.0688 | 19.9526 | 45.5337 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1048 | 3.1391 | 0.008283 | 0.016 | 0.0841 | -0.036 | -0.1522
2012-10-30 14:32 | 23.6645 | 20.0868 | 20.1909 | 20.0427 | 19.9266 | 45.5042 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014414 | 0.23445 | 0.003379 | 16.2654 | 20.1148 | 3.1371 | 0.008286 | -0.028 | 0.0761 | -0.0721 | -0.1882
2012-10-30 14:32 | 23.6667 | 20.101 | 20.161 | 20.0529 | 19.9568 | 45.5337 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.141 | 3.1391 | 0.008283 | -0.04 | 0.02 | -0.0881 | -0.1842
2012-10-30 14:32 | 23.6652 | 20.1195 | 20.2116 | 20.0594 | 19.9553 | 45.5244 | 0.014194 | 0.34553 | 0.004904 | 24.3432 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1035 | 3.1385 | 0.008282 | 0.016 | 0.1081 | -0.0441 | -0.1482
2012-10-30 14:32 | 23.6637 | 20.082 | 20.1701 | 20.0339 | 19.9538 | 45.5308 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.102 | 3.1389 | 0.00828 | -0.02 | 0.0681 | -0.0681 | -0.1482
2012-10-30 14:32 | 23.663 | 20.0773 | 20.1453 | 20.0172 | 19.9331 | 45.5343 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1373 | 3.1391 | 0.008286 | -0.06 | 0.008 | -0.1201 | -0.2042
2012-10-30 14:32 | 23.6652 | 20.0875 | 20.1916 | 20.0034 | 19.9353 | 45.5181 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.0915 | 3.138 | 0.008279 | -0.004 | 0.1001 | -0.0881 | -0.1562
2012-10-30 14:32 | 23.6689 | 20.1192 | 20.1793 | 20.0111 | 19.9029 | 45.5383 | 0.014194 | 0.3454 | 0.004903 | 24.3341 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1472 | 3.1394 | 0.008283 | -0.028 | 0.0321 | -0.1361 | -0.2443
2012-10-30 14:32 | 23.663 | 20.0853 | 20.1453 | 20.0212 | 19.913 | 45.5337 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1013 | 3.1391 | 0.008286 | -0.016 | 0.044 | -0.0801 | -0.1883
2012-10-30 14:32 | 23.6674 | 20.0937 | 20.1738 | 20.0336 | 19.9375 | 45.5146 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1057 | 3.1378 | 0.008285 | -0.012 | 0.0681 | -0.0721 | -0.1682
2012-10-30 14:32 | 23.6704 | 20.0686 | 20.1607 | 20.0046 | 19.9365 | 45.5424 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.1407 | 3.1397 | 0.008287 | -0.0721 | 0.02 | -0.1361 | -0.2042
2012-10-30 14:32 | 23.6562 | 20.0665 | 20.1626 | 20.0024 | 19.9343 | 45.5331 | 0.014194 | 0.3454 | 0.004903 | 24.3341 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1065 | 3.1391 | 0.008282 | -0.04 | 0.0561 | -0.1041 | -0.1722
2012-10-30 14:32 | 23.6652 | 20.0434 | 20.1756 | 20.0354 | 19.9313 | 45.5273 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.1155 | 3.1387 | 0.00828 | -0.0721 | 0.0601 | -0.0801 | -0.1842
2012-10-30 14:32 | 23.6674 | 20.0817 | 20.1457 | 20.0056 | 19.9215 | 45.5441 | 0.014194 | 0.34538 | 0.004902 | 24.333 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.1177 | 3.1398 | 0.00828 | -0.036 | 0.028 | -0.1121 | -0.1962
2012-10-30 14:32 | 23.6674 | 20.0737 | 20.1097 | 20.0416 | 19.9054 | 45.5221 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1137 | 3.1383 | 0.008291 | -0.04 | -0.004 | -0.0721 | -0.2083
2012-10-30 14:32 | 23.6674 | 20.0416 | 20.1257 | 20.0176 | 19.8654 | 45.5267 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.1417 | 3.1386 | 0.00828 | -0.1001 | -0.016 | -0.1241 | -0.2763
2012-10-30 14:32 | 23.6749 | 20.0932 | 20.1372 | 20.0211 | 19.9169 | 45.5337 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1412 | 3.1391 | 0.008284 | -0.048 | -0.004 | -0.1201 | -0.2243
2012-10-30 14:32 | 23.6689 | 20.0711 | 20.1793 | 19.967 | 19.9149 | 45.5302 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1112 | 3.1389 | 0.008285 | -0.0401 | 0.0681 | -0.1442 | -0.1963
2012-10-30 14:32 | 23.6727 | 20.0509 | 20.143 | 20.0269 | 19.9147 | 45.5221 | 0.014209 | 0.34553 | 0.00491 | 24.3175 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.135 | 3.1383 | 0.008289 | -0.0841 | 0.008 | -0.1081 | -0.2203
2012-10-30 14:32 | 23.6645 | 20.0347 | 20.1669 | 19.9666 | 19.9065 | 45.525 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.0868 | 3.1385 | 0.008282 | -0.0521 | 0.0801 | -0.1202 | -0.1803
2012-10-30 14:33 | 23.6742 | 20.0644 | 20.1365 | 19.9763 | 19.9002 | 45.5239 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.1485 | 3.1384 | 0.008278 | -0.0841 | -0.012 | -0.1722 | -0.2483
2012-10-30 14:33 | 23.6809 | 20.0511 | 20.1592 | 20.0031 | 19.935 | 45.514 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0952 | 3.1377 | 0.008286 | -0.0441 | 0.064 | -0.0921 | -0.1602
2012-10-30 14:33 | 23.6794 | 20.0616 | 20.1617 | 20.0056 | 19.9214 | 45.5239 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.1377 | 3.1384 | 0.008285 | -0.0761 | 0.024 | -0.1321 | -0.2163
2012-10-30 14:33 | 23.6839 | 20.0621 | 20.1502 | 20.0221 | 19.9299 | 45.5187 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1102 | 3.1381 | 0.008287 | -0.0481 | 0.04 | -0.0881 | -0.1803
2012-10-30 14:33 | 23.6787 | 20.0329 | 20.097 | 19.9928 | 19.9167 | 45.5071 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.109 | 3.1373 | 0.008286 | -0.0761 | -0.012 | -0.1162 | -0.1923
2012-10-30 14:33 | 23.6846 | 20.0909 | 20.1549 | 19.9947 | 19.9106 | 45.5285 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1429 | 3.1387 | 0.008287 | -0.052 | 0.012 | -0.1482 | -0.2323
2012-10-30 14:33 | 23.6846 | 20.0628 | 20.1629 | 19.9907 | 19.9387 | 45.5158 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1189 | 3.1379 | 0.008285 | -0.0561 | 0.044 | -0.1282 | -0.1802
2012-10-30 14:33 | 23.6802 | 20.0264 | 20.1185 | 20.0024 | 19.8942 | 45.5181 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014401 | 0.23444 | 0.003376 | 16.2791 | 20.1345 | 3.138 | 0.008283 | -0.1081 | -0.016 | -0.1321 | -0.2403
2012-10-30 14:33 | 23.6839 | 20.0741 | 20.1542 | 20.0261 | 19.9019 | 45.525 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.1782 | 3.1385 | 0.00828 | -0.1041 | -0.024 | -0.1521 | -0.2763
2012-10-30 14:33 | 23.6891 | 20.0233 | 20.1234 | 20.0113 | 19.8991 | 45.5123 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1234 | 3.1376 | 0.008286 | -0.1001 | 0 | -0.1121 | -0.2243
2012-10-30 14:33 | 23.6929 | 20.0511 | 20.1512 | 19.951 | 19.9109 | 45.5106 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1232 | 3.1375 | 0.008289 | -0.0721 | 0.028 | -0.1722 | -0.2123
2012-10-30 14:33 | 23.6906 | 20.0488 | 20.1489 | 19.9967 | 19.8685 | 45.5187 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1369 | 3.1381 | 0.008281 | -0.0881 | 0.012 | -0.1402 | -0.2684
2012-10-30 14:33 | 23.6884 | 20.0266 | 20.1147 | 19.9905 | 19.9104 | 45.5262 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.1267 | 3.1386 | 0.008291 | -0.1001 | -0.012 | -0.1362 | -0.2163
2012-10-30 14:33 | 23.6981 | 20.0042 | 20.1484 | 20.0082 | 19.8881 | 45.5163 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1204 | 3.1379 | 0.008281 | -0.1162 | 0.028 | -0.1122 | -0.2323
2012-10-30 14:33 | 23.6936 | 20.0438 | 20.1519 | 19.9717 | 19.9036 | 45.5135 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1199 | 3.1377 | 0.008286 | -0.0761 | 0.032 | -0.1482 | -0.2163
2012-10-30 14:33 | 23.7063 | 20.0204 | 20.1045 | 19.9804 | 19.8602 | 45.5198 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.1246 | 3.1381 | 0.008288 | -0.1042 | -0.0201 | -0.1442 | -0.2644
2012-10-30 14:33 | 23.6996 | 20.0458 | 20.1299 | 19.9937 | 19.8896 | 45.5215 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1459 | 3.1383 | 0.00828 | -0.1001 | -0.016 | -0.1522 | -0.2563
2012-10-30 14:33 | 23.6974 | 20.0316 | 20.1157 | 19.9995 | 19.8874 | 45.51 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1197 | 3.1375 | 0.008287 | -0.0881 | -0.004 | -0.1202 | -0.2323
2012-10-30 14:33 | 23.7011 | 20.0313 | 20.1474 | 19.9752 | 19.9111 | 45.5071 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0833 | 3.1373 | 0.008284 | -0.052 | 0.0641 | -0.1081 | -0.1722
2012-10-30 14:33 | 23.6951 | 20.0453 | 20.1134 | 19.9532 | 19.869 | 45.5285 | 0.014194 | 0.34551 | 0.004904 | 24.3422 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.1214 | 3.1387 | 0.008285 | -0.0761 | -0.008 | -0.1682 | -0.2524
2012-10-30 14:33 | 23.6966 | 20.0067 | 20.1549 | 19.9827 | 19.9026 | 45.5215 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.0988 | 3.1383 | 0.008284 | -0.0921 | 0.0561 | -0.1161 | -0.1962
2012-10-30 14:33 | 23.7018 | 20.028 | 20.1161 | 19.9518 | 19.8797 | 45.5054 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1481 | 3.1371 | 0.008284 | -0.1201 | -0.032 | -0.1963 | -0.2684
2012-10-30 14:33 | 23.7093 | 20.0234 | 20.1676 | 19.9634 | 19.8952 | 45.5198 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1236 | 3.1381 | 0.008286 | -0.1002 | 0.044 | -0.1602 | -0.2284
2012-10-30 14:33 | 23.7116 | 20.0538 | 20.1179 | 19.9857 | 19.8695 | 45.5221 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1659 | 3.1383 | 0.008287 | -0.1121 | -0.048 | -0.1802 | -0.2964
2012-10-30 14:33 | 23.7078 | 20.05 | 20.09 | 19.9699 | 19.9138 | 45.5158 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1221 | 3.1379 | 0.008285 | -0.0721 | -0.0321 | -0.1522 | -0.2083
2012-10-30 14:34 | 23.719 | 20.0492 | 20.1653 | 20.0131 | 19.8689 | 45.5146 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1453 | 3.1378 | 0.008285 | -0.0961 | 0.02 | -0.1322 | -0.2764
2012-10-30 14:34 | 23.7123 | 20.0385 | 20.1105 | 19.9704 | 19.8942 | 45.5273 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1226 | 3.1387 | 0.008287 | -0.0841 | -0.0121 | -0.1522 | -0.2284
2012-10-30 14:34 | 23.7033 | 20.0094 | 20.1496 | 19.9654 | 19.8612 | 45.5071 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1096 | 3.1373 | 0.008283 | -0.1002 | 0.04 | -0.1442 | -0.2484
2012-10-30 14:34 | 23.7168 | 20.0149 | 20.111 | 19.9548 | 19.8987 | 45.5146 | 0.014216 | 0.34537 | 0.00491 | 24.2944 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.1551 | 3.1378 | 0.008291 | -0.1402 | -0.0441 | -0.2003 | -0.2564
2012-10-30 14:34 | 23.7146 | 20.0087 | 20.1369 | 19.9646 | 19.8485 | 45.5239 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014414 | 0.23458 | 0.003381 | 16.2745 | 20.1649 | 3.1384 | 0.008288 | -0.1562 | -0.028 | -0.2003 | -0.3164
2012-10-30 14:34 | 23.7138 | 19.9959 | 20.1201 | 19.9198 | 19.8917 | 45.5198 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.1481 | 3.1381 | 0.008288 | -0.1522 | -0.028 | -0.2283 | -0.2564
2012-10-30 14:34 | 23.7161 | 19.9702 | 20.1264 | 19.9982 | 19.8299 | 45.5013 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1264 | 3.1369 | 0.008281 | -0.1562 | 0 | -0.1282 | -0.2965
2012-10-30 14:34 | 23.7078 | 19.9979 | 20.1461 | 19.9699 | 19.8737 | 45.521 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1301 | 3.1382 | 0.008284 | -0.1322 | 0.016 | -0.1602 | -0.2564
2012-10-30 14:34 | 23.719 | 20.0011 | 20.0932 | 19.937 | 19.8729 | 45.5221 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1212 | 3.1383 | 0.008284 | -0.1201 | -0.028 | -0.1842 | -0.2483
2012-10-30 14:34 | 23.7198 | 19.9859 | 20.1341 | 19.9859 | 19.8697 | 45.5169 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1581 | 3.1379 | 0.008283 | -0.1722 | -0.024 | -0.1722 | -0.2884
2012-10-30 14:34 | 23.728 | 20.0141 | 20.1302 | 19.954 | 19.9019 | 45.5071 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014392 | 0.23444 | 0.003374 | 16.2893 | 20.1783 | 3.1373 | 0.008281 | -0.1642 | -0.0481 | -0.2243 | -0.2764
2012-10-30 14:34 | 23.7318 | 20.0099 | 20.1501 | 19.9818 | 19.8937 | 45.5083 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1621 | 3.1373 | 0.008281 | -0.1522 | -0.012 | -0.1803 | -0.2684
2012-10-30 14:34 | 23.7198 | 19.9698 | 20.066 | 19.9859 | 19.9057 | 45.5163 | 0.01419 | 0.34537 | 0.004901 | 24.3389 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.11 | 3.1379 | 0.00828 | -0.1402 | -0.044 | -0.1241 | -0.2043
2012-10-30 14:34 | 23.7213 | 19.9914 | 20.1075 | 19.9713 | 19.8992 | 45.5262 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.1636 | 3.1386 | 0.008283 | -0.1722 | -0.0561 | -0.1923 | -0.2644
2012-10-30 14:34 | 23.7228 | 20.0209 | 20.085 | 19.9128 | 19.8527 | 45.5215 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014414 | 0.23445 | 0.003379 | 16.2654 | 20.121 | 3.1383 | 0.00828 | -0.1001 | -0.036 | -0.2082 | -0.2683
2012-10-30 14:34 | 23.7235 | 19.9816 | 20.1177 | 19.9936 | 19.8654 | 45.5065 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1298 | 3.1372 | 0.008284 | -0.1482 | -0.0121 | -0.1362 | -0.2644
2012-10-30 14:34 | 23.7273 | 20.0014 | 20.0815 | 19.9733 | 19.8812 | 45.5291 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.1536 | 3.1388 | 0.008287 | -0.1522 | -0.0721 | -0.1803 | -0.2724
2012-10-30 14:34 | 23.7243 | 19.9784 | 20.1145 | 19.9383 | 19.8702 | 45.5244 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.1466 | 3.1385 | 0.008281 | -0.1682 | -0.0321 | -0.2083 | -0.2764
2012-10-30 14:34 | 23.7332 | 19.9953 | 20.0994 | 19.9512 | 19.8831 | 45.5187 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1875 | 3.1381 | 0.008282 | -0.1922 | -0.0881 | -0.2363 | -0.3044
2012-10-30 14:34 | 23.731 | 20.0011 | 20.1092 | 19.953 | 19.8528 | 45.4967 | 0.014194 | 0.34536 | 0.004902 | 24.3311 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.1853 | 3.1365 | 0.00828 | -0.1842 | -0.0761 | -0.2323 | -0.3325
2012-10-30 14:34 | 23.7288 | 19.9668 | 20.087 | 19.9228 | 19.8506 | 45.5129 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1751 | 3.1377 | 0.008287 | -0.2083 | -0.0881 | -0.2523 | -0.3245
2012-10-30 14:34 | 23.7332 | 19.9993 | 20.1114 | 19.9672 | 19.9031 | 45.514 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1835 | 3.1377 | 0.008281 | -0.1842 | -0.0721 | -0.2163 | -0.2804
2012-10-30 14:34 | 23.725 | 19.9951 | 20.1112 | 19.943 | 19.8629 | 45.5054 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1713 | 3.1371 | 0.008281 | -0.1762 | -0.0601 | -0.2283 | -0.3084
2012-10-30 14:35 | 23.728 | 20.0021 | 20.1062 | 19.95 | 19.8779 | 45.5031 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.1663 | 3.137 | 0.008283 | -0.1642 | -0.0601 | -0.2163 | -0.2884
2012-10-30 14:35 | 23.7295 | 19.9836 | 20.0917 | 19.9114 | 19.8553 | 45.5059 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1478 | 3.1372 | 0.00828 | -0.1642 | -0.0561 | -0.2364 | -0.2925
2012-10-30 14:35 | 23.7325 | 19.9906 | 20.1067 | 19.9465 | 19.8583 | 45.5187 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1548 | 3.1381 | 0.008281 | -0.1642 | -0.0481 | -0.2083 | -0.2965
2012-10-30 14:35 | 23.7295 | 19.9956 | 20.0917 | 19.9395 | 19.8634 | 45.4984 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1758 | 3.1367 | 0.008281 | -0.1802 | -0.0841 | -0.2363 | -0.3124
2012-10-30 14:35 | 23.734 | 19.956 | 20.0722 | 19.9199 | 19.8919 | 45.5019 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014395 | 0.23444 | 0.003375 | 16.2859 | 20.1643 | 3.1369 | 0.008279 | -0.2083 | -0.0921 | -0.2444 | -0.2724
2012-10-30 14:35 | 23.728 | 19.9861 | 20.0862 | 19.9139 | 19.8498 | 45.495 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.0982 | 3.1364 | 0.008287 | -0.1121 | -0.012 | -0.1843 | -0.2484
2012-10-30 14:35 | 23.731 | 19.965 | 20.1012 | 19.9009 | 19.8769 | 45.5048 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.1052 | 3.1371 | 0.008284 | -0.1402 | -0.004 | -0.2043 | -0.2283
2012-10-30 14:35 | 23.7273 | 19.9733 | 20.0935 | 19.9733 | 19.8692 | 45.5088 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1255 | 3.1374 | 0.008286 | -0.1522 | -0.032 | -0.1522 | -0.2563
2012-10-30 14:35 | 23.728 | 19.97 | 20.0702 | 19.922 | 19.8779 | 45.495 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1463 | 3.1364 | 0.008282 | -0.1763 | -0.0761 | -0.2243 | -0.2684
2012-10-30 14:35 | 23.7332 | 19.9993 | 20.0874 | 19.9352 | 19.8711 | 45.5019 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.0914 | 3.1369 | 0.008286 | -0.0921 | -0.004 | -0.1562 | -0.2203
2012-10-30 14:35 | 23.7318 | 19.9939 | 20.094 | 19.9418 | 19.8496 | 45.5013 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.106 | 3.1369 | 0.008289 | -0.1121 | -0.012 | -0.1642 | -0.2564
2012-10-30 14:35 | 23.7362 | 20.0223 | 20.0824 | 19.9903 | 19.842 | 45.5059 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.1224 | 3.1372 | 0.008285 | -0.1001 | -0.04 | -0.1321 | -0.2804
2012-10-30 14:35 | 23.7318 | 19.9698 | 20.074 | 19.9338 | 19.8617 | 45.5002 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014386 | 0.23452 | 0.003374 | 16.3022 | 20.082 | 3.1368 | 0.008282 | -0.1122 | -0.008 | -0.1482 | -0.2203
2012-10-30 14:35 | 23.7347 | 19.9968 | 20.0969 | 19.9206 | 19.8485 | 45.5013 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1049 | 3.1369 | 0.008281 | -0.1081 | -0.008 | -0.1843 | -0.2564
2012-10-30 14:35 | 23.734 | 19.984 | 20.0802 | 19.9119 | 19.8558 | 45.5129 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.0922 | 3.1377 | 0.008283 | -0.1082 | -0.012 | -0.1803 | -0.2364
2012-10-30 14:35 | 23.7332 | 19.9592 | 20.0954 | 19.8911 | 19.8791 | 45.5042 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1194 | 3.1371 | 0.00828 | -0.1602 | -0.024 | -0.2283 | -0.2403
2012-10-30 14:35 | 23.7362 | 19.9662 | 20.0463 | 19.9101 | 19.854 | 45.4996 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1344 | 3.1367 | 0.008283 | -0.1682 | -0.0881 | -0.2243 | -0.2804
2012-10-30 14:35 | 23.7325 | 19.9745 | 20.0547 | 19.9425 | 19.8784 | 45.499 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1067 | 3.1367 | 0.008281 | -0.1322 | -0.052 | -0.1642 | -0.2283
2012-10-30 14:35 | 23.731 | 19.965 | 20.0291 | 19.937 | 19.8448 | 45.5233 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.1092 | 3.1384 | 0.008289 | -0.1442 | -0.0801 | -0.1722 | -0.2644
2012-10-30 14:35 | 23.737 | 19.959 | 20.0551 | 19.9229 | 19.8468 | 45.4979 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1032 | 3.1366 | 0.008285 | -0.1442 | -0.0481 | -0.1803 | -0.2564
2012-10-30 14:35 | 23.7355 | 19.9815 | 20.0336 | 19.9054 | 19.8734 | 45.4938 | 0.014216 | 0.34553 | 0.004912 | 24.3056 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.1337 | 3.1363 | 0.008292 | -0.1522 | -0.1001 | -0.2283 | -0.2603
2012-10-30 14:35 | 23.7392 | 19.9572 | 20.0654 | 19.9011 | 19.857 | 45.5146 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1374 | 3.1378 | 0.008283 | -0.1802 | -0.072 | -0.2363 | -0.2804
2012-10-30 14:35 | 23.7332 | 19.9752 | 20.0313 | 19.9111 | 19.831 | 45.51 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.1435 | 3.1375 | 0.008287 | -0.1683 | -0.1122 | -0.2324 | -0.3125
2012-10-30 14:35 | 23.7347 | 20.0088 | 20.0088 | 19.9006 | 19.8165 | 45.4961 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.141 | 3.1365 | 0.008287 | -0.1322 | -0.1322 | -0.2404 | -0.3245
2012-10-30 14:35 | 23.7392 | 19.9572 | 20.0453 | 19.9292 | 19.9011 | 45.5007 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1855 | 3.1368 | 0.008284 | -0.2283 | -0.1402 | -0.2563 | -0.2844
2012-10-30 14:36 | 23.7407 | 20.0068 | 20.0709 | 19.8986 | 19.8144 | 45.5013 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014417 | 0.23445 | 0.00338 | 16.2621 | 20.1149 | 3.1369 | 0.008285 | -0.1081 | -0.044 | -0.2163 | -0.3005
2012-10-30 14:36 | 23.737 | 19.947 | 20.0632 | 19.9229 | 19.8348 | 45.4927 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.1232 | 3.1363 | 0.00828 | -0.1762 | -0.06 | -0.2003 | -0.2884
2012-10-30 14:36 | 23.7422 | 19.9882 | 20.0483 | 19.8801 | 19.8039 | 45.4932 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014411 | 0.23444 | 0.003378 | 16.2679 | 20.1525 | 3.1363 | 0.008286 | -0.1643 | -0.1042 | -0.2724 | -0.3486
2012-10-30 14:36 | 23.7467 | 19.9927 | 20.0528 | 19.8645 | 19.8525 | 45.5065 | 0.014184 | 0.34541 | 0.004899 | 24.3522 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.1249 | 3.1372 | 0.008275 | -0.1322 | -0.0721 | -0.2604 | -0.2724
2012-10-30 14:36 | 23.7504 | 19.9724 | 20.0085 | 19.9123 | 19.8522 | 45.4996 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1326 | 3.1367 | 0.008285 | -0.1602 | -0.1241 | -0.2203 | -0.2804
2012-10-30 14:36 | 23.7437 | 19.9457 | 20.0258 | 19.8856 | 19.8575 | 45.5054 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1139 | 3.1371 | 0.008285 | -0.1682 | -0.0881 | -0.2283 | -0.2564
2012-10-30 14:36 | 23.7475 | 19.9294 | 20.0256 | 19.8774 | 19.8212 | 45.4996 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1137 | 3.1367 | 0.008281 | -0.1843 | -0.0881 | -0.2363 | -0.2925
2012-10-30 14:36 | 23.7534 | 19.9994 | 20.0315 | 19.8873 | 19.8392 | 45.484 | 0.014187 | 0.34546 | 0.004901 | 24.3502 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1236 | 3.1357 | 0.008277 | -0.1242 | -0.0921 | -0.2363 | -0.2844
2012-10-30 14:36 | 23.7475 | 19.9615 | 20.0536 | 19.8974 | 19.8212 | 45.4886 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.1017 | 3.136 | 0.008288 | -0.1402 | -0.0481 | -0.2043 | -0.2805
2012-10-30 14:36 | 23.7504 | 19.9604 | 20.0285 | 19.8642 | 19.7881 | 45.5106 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1086 | 3.1375 | 0.008285 | -0.1482 | -0.0801 | -0.2444 | -0.3205
2012-10-30 14:36 | 23.7587 | 19.9647 | 20.0248 | 19.9046 | 19.8525 | 45.5013 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0969 | 3.1369 | 0.008281 | -0.1322 | -0.0721 | -0.1923 | -0.2444
2012-10-30 14:36 | 23.749 | 19.9309 | 20.0111 | 19.8789 | 19.8468 | 45.5002 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1272 | 3.1368 | 0.00828 | -0.1963 | -0.1161 | -0.2483 | -0.2804
2012-10-30 14:36 | 23.7572 | 19.9632 | 20.0433 | 19.9151 | 19.7748 | 45.4909 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1554 | 3.1361 | 0.008285 | -0.1922 | -0.1121 | -0.2403 | -0.3806
2012-10-30 14:36 | 23.7527 | 19.9386 | 20.0308 | 19.8946 | 19.8144 | 45.4822 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.0989 | 3.1355 | 0.008283 | -0.1603 | -0.0681 | -0.2043 | -0.2845
2012-10-30 14:36 | 23.7647 | 19.9547 | 20.0548 | 19.9186 | 19.7984 | 45.5007 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.1509 | 3.1368 | 0.008283 | -0.1962 | -0.0961 | -0.2323 | -0.3525
2012-10-30 14:36 | 23.7602 | 19.9461 | 19.9982 | 19.886 | 19.8219 | 45.5002 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.1464 | 3.1368 | 0.008279 | -0.2003 | -0.1482 | -0.2604 | -0.3245
2012-10-30 14:36 | 23.7617 | 19.9476 | 20.0117 | 19.8916 | 19.8234 | 45.4846 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.176 | 3.1357 | 0.008282 | -0.2284 | -0.1643 | -0.2844 | -0.3526
2012-10-30 14:36 | 23.7602 | 19.9582 | 20.0783 | 19.8941 | 19.7899 | 45.5065 | 0.01419 | 0.3454 | 0.004901 | 24.3409 | 0.014386 | 0.23448 | 0.003373 | 16.2991 | 20.1665 | 3.1372 | 0.008274 | -0.2083 | -0.0882 | -0.2724 | -0.3766
2012-10-30 14:36 | 23.7609 | 19.9669 | 20.055 | 19.8667 | 19.8266 | 45.5158 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.1151 | 3.1379 | 0.008286 | -0.1482 | -0.0601 | -0.2484 | -0.2885
2012-10-30 14:36 | 23.7594 | 19.9494 | 20.0655 | 19.8492 | 19.8171 | 45.495 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1616 | 3.1364 | 0.00828 | -0.2122 | -0.0961 | -0.3124 | -0.3445
2012-10-30 14:36 | 23.7691 | 19.911 | 19.9911 | 19.8428 | 19.8148 | 45.5059 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1313 | 3.1372 | 0.008282 | -0.2203 | -0.1402 | -0.2885 | -0.3165
2012-10-30 14:36 | 23.7669 | 19.9729 | 20.021 | 19.8807 | 19.8086 | 45.4927 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.1411 | 3.1363 | 0.00828 | -0.1682 | -0.1201 | -0.2604 | -0.3325
2012-10-30 14:36 | 23.7594 | 19.9293 | 20.0255 | 19.8652 | 19.8291 | 45.4892 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1576 | 3.136 | 0.008282 | -0.2283 | -0.1321 | -0.2924 | -0.3285
2012-10-30 14:36 | 23.7662 | 19.9081 | 20.0523 | 19.844 | 19.844 | 45.495 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014398 | 0.23442 | 0.003375 | 16.2815 | 20.1725 | 3.1364 | 0.008282 | -0.2644 | -0.1202 | -0.3285 | -0.3285
2012-10-30 14:37 | 23.7676 | 19.9255 | 20.0497 | 19.9055 | 19.8413 | 45.4909 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1378 | 3.1361 | 0.008286 | -0.2123 | -0.0881 | -0.2323 | -0.2965
2012-10-30 14:37 | 23.7654 | 19.8912 | 20.0154 | 19.8472 | 19.8271 | 45.4944 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.1797 | 3.1364 | 0.008284 | -0.2885 | -0.1643 | -0.3325 | -0.3526
2012-10-30 14:37 | 23.7647 | 19.9627 | 20.0147 | 19.8384 | 19.7944 | 45.5094 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1589 | 3.1374 | 0.008281 | -0.1962 | -0.1442 | -0.3205 | -0.3645
2012-10-30 14:37 | 23.7617 | 19.9677 | 20.0438 | 19.8475 | 19.8314 | 45.477 | 0.014187 | 0.3454 | 0.0049 | 24.3461 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.18 | 3.1352 | 0.008276 | -0.2123 | -0.1362 | -0.3325 | -0.3486
2012-10-30 14:37 | 23.7647 | 19.9547 | 20.0188 | 19.8545 | 19.7903 | 45.5019 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.175 | 3.1369 | 0.008286 | -0.2203 | -0.1562 | -0.3205 | -0.3847
2012-10-30 14:37 | 23.7662 | 19.872 | 20.0403 | 19.852 | 19.8159 | 45.4973 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1484 | 3.1366 | 0.008282 | -0.2764 | -0.1081 | -0.2964 | -0.3325
2012-10-30 14:37 | 23.7647 | 19.9106 | 20.0348 | 19.8986 | 19.8344 | 45.5013 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1669 | 3.1369 | 0.008286 | -0.2563 | -0.1321 | -0.2683 | -0.3325
2012-10-30 14:37 | 23.7714 | 19.9373 | 20.0655 | 19.8772 | 19.777 | 45.4955 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1696 | 3.1365 | 0.008283 | -0.2323 | -0.1041 | -0.2924 | -0.3926
2012-10-30 14:37 | 23.7594 | 19.9373 | 20.0215 | 19.8852 | 19.757 | 45.4961 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1576 | 3.1365 | 0.008282 | -0.2203 | -0.1361 | -0.2724 | -0.4006
2012-10-30 14:37 | 23.7744 | 19.9083 | 20.0605 | 19.8962 | 19.8321 | 45.4846 | 0.014219 | 0.34548 | 0.004912 | 24.2974 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 20.2207 | 3.1357 | 0.008286 | -0.3124 | -0.1602 | -0.3245 | -0.3886
2012-10-30 14:37 | 23.7647 | 19.9186 | 20.0388 | 19.9026 | 19.8224 | 45.4967 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.203 | 3.1365 | 0.008287 | -0.2844 | -0.1642 | -0.3004 | -0.3806
2012-10-30 14:37 | 23.7729 | 19.9148 | 20.047 | 19.8466 | 19.8346 | 45.495 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1792 | 3.1364 | 0.008287 | -0.2644 | -0.1322 | -0.3326 | -0.3446
2012-10-30 14:37 | 23.7699 | 19.9158 | 20.032 | 19.8356 | 19.8517 | 45.4909 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014401 | 0.23444 | 0.003376 | 16.2791 | 20.1561 | 3.1361 | 0.008282 | -0.2403 | -0.1241 | -0.3205 | -0.3044
2012-10-30 14:37 | 23.7654 | 19.8993 | 20.0034 | 19.8672 | 19.8271 | 45.4869 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.1797 | 3.1359 | 0.008282 | -0.2804 | -0.1763 | -0.3125 | -0.3526
2012-10-30 14:37 | 23.7669 | 19.9208 | 20.049 | 19.8647 | 19.7444 | 45.4863 | 0.014194 | 0.34538 | 0.004902 | 24.333 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1371 | 3.1358 | 0.00828 | -0.2163 | -0.0881 | -0.2724 | -0.3927
2012-10-30 14:37 | 23.7617 | 19.8916 | 20.0558 | 19.8555 | 19.7914 | 45.4869 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1439 | 3.1359 | 0.008282 | -0.2523 | -0.0881 | -0.2884 | -0.3525
2012-10-30 14:37 | 23.7699 | 19.8917 | 20.044 | 19.8557 | 19.8316 | 45.4955 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 20.1601 | 3.1365 | 0.008279 | -0.2684 | -0.1161 | -0.3044 | -0.3285
2012-10-30 14:37 | 23.7744 | 19.9123 | 20.0605 | 19.8882 | 19.788 | 45.4817 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.1646 | 3.1355 | 0.008283 | -0.2523 | -0.1041 | -0.2764 | -0.3766
2012-10-30 14:37 | 23.7736 | 19.9035 | 20.0397 | 19.8674 | 19.7992 | 45.4961 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.1678 | 3.1365 | 0.008279 | -0.2643 | -0.1281 | -0.3004 | -0.3686
2012-10-30 14:37 | 23.7744 | 19.9123 | 20.0685 | 19.8562 | 19.8 | 45.4979 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1606 | 3.1366 | 0.00828 | -0.2483 | -0.0921 | -0.3044 | -0.3606
2012-10-30 14:37 | 23.7721 | 19.91 | 20.0262 | 19.8138 | 19.7737 | 45.4921 | 0.014212 | 0.34537 | 0.004908 | 24.3012 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1583 | 3.1362 | 0.008288 | -0.2483 | -0.1321 | -0.3445 | -0.3846
2012-10-30 14:37 | 23.7706 | 19.9205 | 19.9926 | 19.8163 | 19.8003 | 45.4927 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1448 | 3.1363 | 0.008282 | -0.2243 | -0.1522 | -0.3285 | -0.3445
2012-10-30 14:37 | 23.7706 | 19.9045 | 20.0206 | 19.7842 | 19.8043 | 45.4898 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1809 | 3.1361 | 0.008283 | -0.2764 | -0.1603 | -0.3967 | -0.3766
2012-10-30 14:38 | 23.7684 | 19.8662 | 20.0265 | 19.8021 | 19.794 | 45.4961 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 20.1266 | 3.1365 | 0.008282 | -0.2604 | -0.1001 | -0.3245 | -0.3326
2012-10-30 14:38 | 23.7706 | 19.9085 | 20.0166 | 19.8243 | 19.7882 | 45.4961 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1168 | 3.1365 | 0.008287 | -0.2083 | -0.1002 | -0.2925 | -0.3286
2012-10-30 14:38 | 23.7759 | 19.8857 | 20.0059 | 19.8617 | 19.8136 | 45.4863 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.1741 | 3.1358 | 0.008288 | -0.2884 | -0.1682 | -0.3124 | -0.3605
2012-10-30 14:38 | 23.7789 | 19.9047 | 20.049 | 19.8366 | 19.7845 | 45.4927 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.1491 | 3.1363 | 0.008283 | -0.2444 | -0.1001 | -0.3125 | -0.3646
2012-10-30 14:38 | 23.7781 | 19.9039 | 20.0001 | 19.8358 | 19.7917 | 45.5013 | 0.014216 | 0.34541 | 0.00491 | 24.2974 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1363 | 3.1369 | 0.008288 | -0.2324 | -0.1362 | -0.3005 | -0.3446
2012-10-30 14:38 | 23.7804 | 19.8902 | 20.0345 | 19.8261 | 19.786 | 45.4828 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.1666 | 3.1356 | 0.008284 | -0.2764 | -0.1321 | -0.3405 | -0.3806
2012-10-30 14:38 | 23.7781 | 19.8799 | 20.0081 | 19.8358 | 19.7997 | 45.4782 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.1683 | 3.1353 | 0.008281 | -0.2884 | -0.1602 | -0.3325 | -0.3686
2012-10-30 14:38 | 23.7848 | 19.9147 | 19.9868 | 19.8345 | 19.7824 | 45.4909 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.171 | 3.1361 | 0.008282 | -0.2563 | -0.1842 | -0.3365 | -0.3886
2012-10-30 14:38 | 23.7804 | 19.8982 | 20.0304 | 19.8662 | 19.806 | 45.488 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1746 | 3.1359 | 0.008284 | -0.2764 | -0.1442 | -0.3084 | -0.3686
2012-10-30 14:38 | 23.7848 | 19.8986 | 20.0068 | 19.8265 | 19.7503 | 45.4921 | 0.01419 | 0.34536 | 0.004901 | 24.3379 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.171 | 3.1362 | 0.008278 | -0.2724 | -0.1642 | -0.3445 | -0.4207
2012-10-30 14:38 | 23.7759 | 19.8857 | 19.9979 | 19.8176 | 19.7855 | 45.4874 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1341 | 3.1359 | 0.008286 | -0.2484 | -0.1362 | -0.3165 | -0.3486
2012-10-30 14:38 | 23.7871 | 19.8849 | 19.9811 | 19.8168 | 19.7887 | 45.4927 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1813 | 3.1363 | 0.008284 | -0.2964 | -0.2002 | -0.3645 | -0.3926
2012-10-30 14:38 | 23.7804 | 19.9062 | 19.9944 | 19.8221 | 19.77 | 45.4886 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1746 | 3.136 | 0.008279 | -0.2684 | -0.1802 | -0.3525 | -0.4046
2012-10-30 14:38 | 23.7878 | 19.9257 | 20.0178 | 19.8375 | 19.7854 | 45.4898 | 0.014203 | 0.34534 | 0.004905 | 24.3146 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.138 | 3.1361 | 0.008282 | -0.2123 | -0.1202 | -0.3005 | -0.3526
2012-10-30 14:38 | 23.7796 | 19.9054 | 20.0176 | 19.8253 | 19.7652 | 45.4874 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.1138 | 3.1359 | 0.008289 | -0.2084 | -0.0962 | -0.2885 | -0.3486
2012-10-30 14:38 | 23.7863 | 19.8801 | 20.0083 | 19.8119 | 19.7879 | 45.4898 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.0844 | 3.1361 | 0.008286 | -0.2043 | -0.0761 | -0.2725 | -0.2965
2012-10-30 14:38 | 23.7766 | 19.8904 | 19.9906 | 19.8143 | 19.7822 | 45.4811 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1108 | 3.1355 | 0.008283 | -0.2204 | -0.1202 | -0.2965 | -0.3286
2012-10-30 14:38 | 23.7886 | 19.9104 | 19.9946 | 19.8463 | 19.7581 | 45.4851 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1148 | 3.1357 | 0.008282 | -0.2044 | -0.1202 | -0.2685 | -0.3567
2012-10-30 14:38 | 23.7856 | 19.9114 | 20.0116 | 19.8433 | 19.7912 | 45.4915 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1438 | 3.1362 | 0.00828 | -0.2324 | -0.1322 | -0.3005 | -0.3526
2012-10-30 14:38 | 23.7774 | 19.8552 | 19.9754 | 19.8271 | 19.787 | 45.4921 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1636 | 3.1362 | 0.008282 | -0.3084 | -0.1882 | -0.3365 | -0.3766
2012-10-30 14:38 | 23.7766 | 19.8664 | 20.0427 | 19.8864 | 19.7902 | 45.5059 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1308 | 3.1372 | 0.00828 | -0.2644 | -0.0881 | -0.2444 | -0.3406
2012-10-30 14:38 | 23.7736 | 19.8594 | 20.0277 | 19.7912 | 19.7632 | 45.4857 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 20.1518 | 3.1358 | 0.008284 | -0.2924 | -0.1241 | -0.3606 | -0.3886
2012-10-30 14:38 | 23.7863 | 19.8961 | 20.0003 | 19.816 | 19.7438 | 45.4846 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.1565 | 3.1357 | 0.008284 | -0.2604 | -0.1562 | -0.3405 | -0.4127
2012-10-30 14:39 | 23.7766 | 19.8263 | 20.0146 | 19.7702 | 19.7982 | 45.4979 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1628 | 3.1366 | 0.008286 | -0.3365 | -0.1482 | -0.3926 | -0.3646
2012-10-30 14:39 | 23.7826 | 19.9004 | 19.9886 | 19.8203 | 19.7481 | 45.4742 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1488 | 3.135 | 0.008284 | -0.2484 | -0.1602 | -0.3285 | -0.4007
2012-10-30 14:39 | 23.7774 | 19.8912 | 19.9754 | 19.8191 | 19.803 | 45.4874 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1516 | 3.1359 | 0.008282 | -0.2604 | -0.1762 | -0.3325 | -0.3486
2012-10-30 14:39 | 23.7819 | 19.8757 | 20.0279 | 19.8035 | 19.7835 | 45.5013 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1962 | 3.1369 | 0.008285 | -0.3205 | -0.1683 | -0.3927 | -0.4127
2012-10-30 14:39 | 23.7729 | 19.8907 | 20.0229 | 19.8306 | 19.7785 | 45.4799 | 0.014206 | 0.34537 | 0.004906 | 24.3115 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.1631 | 3.1354 | 0.008286 | -0.2724 | -0.1402 | -0.3325 | -0.3846
2012-10-30 14:39 | 23.7729 | 19.8947 | 19.9669 | 19.8026 | 19.7625 | 45.4944 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.1872 | 3.1364 | 0.008285 | -0.2925 | -0.2203 | -0.3846 | -0.4247
2012-10-30 14:39 | 23.7766 | 19.8904 | 19.9706 | 19.8183 | 19.7622 | 45.4932 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1909 | 3.1363 | 0.008284 | -0.3005 | -0.2203 | -0.3726 | -0.4287
2012-10-30 14:39 | 23.7744 | 19.8882 | 19.9804 | 19.8081 | 19.776 | 45.4799 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1646 | 3.1354 | 0.008283 | -0.2764 | -0.1842 | -0.3565 | -0.3886
2012-10-30 14:39 | 23.7766 | 19.9024 | 19.9746 | 19.8503 | 19.7702 | 45.4903 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.1468 | 3.1361 | 0.008285 | -0.2444 | -0.1722 | -0.2965 | -0.3766
2012-10-30 14:39 | 23.7684 | 19.9183 | 20.0225 | 19.8141 | 19.7259 | 45.4724 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1626 | 3.1349 | 0.008286 | -0.2443 | -0.1401 | -0.3485 | -0.4367
2012-10-30 14:39 | 23.7662 | 19.868 | 19.9682 | 19.848 | 19.7959 | 45.4713 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1765 | 3.1348 | 0.008283 | -0.3085 | -0.2083 | -0.3285 | -0.3806
2012-10-30 14:39 | 23.7721 | 19.8939 | 19.9981 | 19.8058 | 19.7537 | 45.4828 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1583 | 3.1356 | 0.008284 | -0.2644 | -0.1602 | -0.3525 | -0.4046
2012-10-30 14:39 | 23.7647 | 19.8865 | 20.0107 | 19.8304 | 19.7543 | 45.4886 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1589 | 3.136 | 0.008286 | -0.2724 | -0.1482 | -0.3285 | -0.4046
2012-10-30 14:39 | 23.7639 | 19.8817 | 19.9539 | 19.7535 | 19.7936 | 45.4776 | 0.01419 | 0.34538 | 0.004901 | 24.3399 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.1862 | 3.1352 | 0.008279 | -0.3045 | -0.2323 | -0.4327 | -0.3926
2012-10-30 14:39 | 23.7602 | 19.882 | 19.9421 | 19.7818 | 19.7257 | 45.4707 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014404 | 0.23445 | 0.003377 | 16.2767 | 20.1985 | 3.1348 | 0.008284 | -0.3165 | -0.2564 | -0.4167 | -0.4728
2012-10-30 14:39 | 23.7639 | 19.8577 | 19.9739 | 19.7936 | 19.7334 | 45.4979 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.1541 | 3.1366 | 0.008284 | -0.2964 | -0.1802 | -0.3605 | -0.4207
2012-10-30 14:39 | 23.7602 | 19.85 | 19.9942 | 19.8099 | 19.7618 | 45.4794 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.1504 | 3.1353 | 0.008286 | -0.3004 | -0.1562 | -0.3405 | -0.3886
2012-10-30 14:39 | 23.7587 | 19.8645 | 20.0248 | 19.7843 | 19.7322 | 45.4742 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.1449 | 3.135 | 0.00829 | -0.2804 | -0.1201 | -0.3606 | -0.4127
2012-10-30 14:39 | 23.7654 | 19.8872 | 19.9634 | 19.779 | 19.8111 | 45.4892 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.1316 | 3.136 | 0.008283 | -0.2444 | -0.1682 | -0.3526 | -0.3205
2012-10-30 14:39 | 23.7594 | 19.8331 | 19.9494 | 19.7369 | 19.761 | 45.4851 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1216 | 3.1357 | 0.008284 | -0.2885 | -0.1722 | -0.3847 | -0.3606
2012-10-30 14:39 | 23.7624 | 19.8762 | 19.9564 | 19.776 | 19.756 | 45.4753 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1526 | 3.1351 | 0.008283 | -0.2764 | -0.1962 | -0.3766 | -0.3966
2012-10-30 14:39 | 23.7699 | 19.8877 | 19.9639 | 19.7755 | 19.7314 | 45.4921 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.1561 | 3.1362 | 0.008284 | -0.2684 | -0.1922 | -0.3806 | -0.4247
2012-10-30 14:39 | 23.7654 | 19.8512 | 19.9273 | 19.8151 | 19.767 | 45.477 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1316 | 3.1352 | 0.008286 | -0.2804 | -0.2043 | -0.3165 | -0.3646
2012-10-30 14:40 | 23.7759 | 19.8416 | 19.9659 | 19.7935 | 19.7374 | 45.4747 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1381 | 3.135 | 0.00828 | -0.2965 | -0.1722 | -0.3446 | -0.4007
2012-10-30 14:40 | 23.7804 | 19.8501 | 19.9543 | 19.786 | 19.7299 | 45.4759 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.1867 | 3.1351 | 0.00829 | -0.3366 | -0.2324 | -0.4007 | -0.4568
2012-10-30 14:40 | 23.7871 | 19.8448 | 19.953 | 19.7967 | 19.7767 | 45.4811 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1693 | 3.1355 | 0.008283 | -0.3245 | -0.2163 | -0.3726 | -0.3926
2012-10-30 14:40 | 23.7789 | 19.9047 | 19.9488 | 19.8086 | 19.7083 | 45.4892 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1771 | 3.136 | 0.008283 | -0.2724 | -0.2283 | -0.3685 | -0.4688
2012-10-30 14:40 | 23.7781 | 19.8478 | 19.956 | 19.7757 | 19.7356 | 45.4799 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1924 | 3.1354 | 0.008281 | -0.3446 | -0.2364 | -0.4167 | -0.4568
2012-10-30 14:40 | 23.7729 | 19.8547 | 19.9869 | 19.7745 | 19.7344 | 45.4817 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1832 | 3.1355 | 0.008283 | -0.3285 | -0.1963 | -0.4087 | -0.4488
2012-10-30 14:40 | 23.7796 | 19.8453 | 19.9615 | 19.8293 | 19.7131 | 45.4655 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014417 | 0.23442 | 0.00338 | 16.26 | 20.1498 | 3.1344 | 0.008285 | -0.3045 | -0.1883 | -0.3205 | -0.4367
2012-10-30 14:40 | 23.7901 | 19.8558 | 19.972 | 19.8157 | 19.7356 | 45.4753 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1964 | 3.1351 | 0.008283 | -0.3406 | -0.2244 | -0.3807 | -0.4608
2012-10-30 14:40 | 23.7759 | 19.8577 | 20.0099 | 19.7775 | 19.7575 | 45.4869 | 0.014225 | 0.34543 | 0.004914 | 24.2831 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1902 | 3.1359 | 0.008291 | -0.3325 | -0.1803 | -0.4127 | -0.4327
2012-10-30 14:40 | 23.7781 | 19.8398 | 19.96 | 19.7597 | 19.7556 | 45.4776 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.2044 | 3.1352 | 0.008284 | -0.3646 | -0.2444 | -0.4447 | -0.4488
2012-10-30 14:40 | 23.7766 | 19.8744 | 19.9465 | 19.7541 | 19.7341 | 45.4678 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014395 | 0.23461 | 0.003377 | 16.298 | 20.1508 | 3.1346 | 0.00828 | -0.2764 | -0.2043 | -0.3967 | -0.4167
2012-10-30 14:40 | 23.7766 | 19.8544 | 19.9666 | 19.7622 | 19.7221 | 45.4805 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.1148 | 3.1354 | 0.008289 | -0.2604 | -0.1482 | -0.3526 | -0.3927
2012-10-30 14:40 | 23.7759 | 19.8216 | 19.9779 | 19.7895 | 19.7294 | 45.4834 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014417 | 0.23455 | 0.003382 | 16.2691 | 20.1501 | 3.1356 | 0.008288 | -0.3285 | -0.1722 | -0.3606 | -0.4207
2012-10-30 14:40 | 23.7826 | 19.8403 | 19.9445 | 19.7922 | 19.7722 | 45.4776 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1408 | 3.1352 | 0.008283 | -0.3005 | -0.1963 | -0.3486 | -0.3686
2012-10-30 14:40 | 23.7826 | 19.8483 | 19.9525 | 19.7642 | 19.7682 | 45.4794 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.1448 | 3.1353 | 0.008278 | -0.2965 | -0.1923 | -0.3806 | -0.3766
2012-10-30 14:40 | 23.7819 | 19.7955 | 19.9598 | 19.7795 | 19.6993 | 45.4898 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2162 | 3.1361 | 0.008287 | -0.4207 | -0.2564 | -0.4367 | -0.5169
2012-10-30 14:40 | 23.7826 | 19.8163 | 19.9605 | 19.7481 | 19.7161 | 45.4811 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1528 | 3.1355 | 0.008287 | -0.3365 | -0.1923 | -0.4047 | -0.4367
2012-10-30 14:40 | 23.7878 | 19.8615 | 19.9657 | 19.8215 | 19.7132 | 45.4661 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.162 | 3.1344 | 0.008284 | -0.3005 | -0.1963 | -0.3405 | -0.4488
2012-10-30 14:40 | 23.7848 | 19.8345 | 19.9707 | 19.7704 | 19.7463 | 45.473 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2071 | 3.1349 | 0.008284 | -0.3726 | -0.2364 | -0.4367 | -0.4608
2012-10-30 14:40 | 23.7856 | 19.8393 | 19.9555 | 19.7912 | 19.711 | 45.4742 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1959 | 3.135 | 0.008281 | -0.3566 | -0.2404 | -0.4047 | -0.4849
2012-10-30 14:40 | 23.7834 | 19.805 | 19.9493 | 19.793 | 19.7249 | 45.4927 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.1296 | 3.1363 | 0.008281 | -0.3246 | -0.1803 | -0.3366 | -0.4047
2012-10-30 14:40 | 23.7893 | 19.831 | 19.9552 | 19.7789 | 19.7388 | 45.4753 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1795 | 3.1351 | 0.008287 | -0.3485 | -0.2243 | -0.4006 | -0.4407
2012-10-30 14:40 | 23.7841 | 19.8619 | 19.934 | 19.7737 | 19.7376 | 45.4678 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.1423 | 3.1346 | 0.008291 | -0.2804 | -0.2083 | -0.3686 | -0.4047
2012-10-30 14:41 | 23.7908 | 19.8285 | 19.9888 | 19.8004 | 19.7122 | 45.4886 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.165 | 3.136 | 0.008282 | -0.3365 | -0.1762 | -0.3646 | -0.4528
2012-10-30 14:41 | 23.7916 | 19.8493 | 19.9134 | 19.7371 | 19.705 | 45.4765 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.01442 | 0.23449 | 0.003381 | 16.2617 | 20.1658 | 3.1351 | 0.008288 | -0.3165 | -0.2524 | -0.4287 | -0.4608
2012-10-30 14:41 | 23.7886 | 19.8623 | 19.9465 | 19.7702 | 19.7381 | 45.4661 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1748 | 3.1344 | 0.008285 | -0.3125 | -0.2283 | -0.4046 | -0.4367
2012-10-30 14:41 | 23.7938 | 19.8515 | 19.9317 | 19.7313 | 19.7232 | 45.4736 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.18 | 3.1349 | 0.008286 | -0.3285 | -0.2483 | -0.4487 | -0.4568
2012-10-30 14:41 | 23.7908 | 19.8245 | 19.9367 | 19.7323 | 19.7403 | 45.4718 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.181 | 3.1348 | 0.008286 | -0.3565 | -0.2443 | -0.4487 | -0.4407
2012-10-30 14:41 | 23.7908 | 19.8004 | 19.9607 | 19.7283 | 19.7162 | 45.4695 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.01442 | 0.23449 | 0.003381 | 16.2617 | 20.169 | 3.1347 | 0.008284 | -0.3686 | -0.2083 | -0.4407 | -0.4528
2012-10-30 14:41 | 23.7901 | 19.8478 | 19.96 | 19.7035 | 19.7195 | 45.4661 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1763 | 3.1344 | 0.008285 | -0.3285 | -0.2163 | -0.4728 | -0.4568
2012-10-30 14:41 | 23.7968 | 19.8144 | 19.9387 | 19.7864 | 19.7102 | 45.4834 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.167 | 3.1356 | 0.008286 | -0.3526 | -0.2283 | -0.3806 | -0.4568
2012-10-30 14:41 | 23.7961 | 19.7937 | 19.93 | 19.7696 | 19.7296 | 45.4724 | 0.014219 | 0.34546 | 0.004912 | 24.2954 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1783 | 3.1349 | 0.008291 | -0.3846 | -0.2483 | -0.4087 | -0.4487
2012-10-30 14:41 | 23.7998 | 19.8254 | 19.9377 | 19.7493 | 19.7453 | 45.473 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.154 | 3.1349 | 0.008284 | -0.3286 | -0.2163 | -0.4047 | -0.4087
2012-10-30 14:41 | 23.7968 | 19.8144 | 19.9267 | 19.7743 | 19.6982 | 45.484 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.183 | 3.1357 | 0.008287 | -0.3686 | -0.2563 | -0.4087 | -0.4848
2012-10-30 14:41 | 23.8058 | 19.7874 | 19.9797 | 19.7513 | 19.7753 | 45.4863 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.16 | 3.1358 | 0.008285 | -0.3726 | -0.1803 | -0.4087 | -0.3847
2012-10-30 14:41 | 23.7961 | 19.8258 | 19.962 | 19.7736 | 19.7255 | 45.4643 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1703 | 3.1343 | 0.008284 | -0.3445 | -0.2083 | -0.3967 | -0.4448
2012-10-30 14:41 | 23.805 | 19.8306 | 19.9389 | 19.7505 | 19.7264 | 45.469 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1952 | 3.1346 | 0.008287 | -0.3646 | -0.2563 | -0.4447 | -0.4688
2012-10-30 14:41 | 23.8013 | 19.8029 | 19.9752 | 19.7788 | 19.6947 | 45.4614 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1635 | 3.1341 | 0.008281 | -0.3606 | -0.1883 | -0.3847 | -0.4688
2012-10-30 14:41 | 23.805 | 19.8146 | 19.9188 | 19.7545 | 19.7144 | 45.4851 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.1792 | 3.1357 | 0.008286 | -0.3646 | -0.2604 | -0.4247 | -0.4648
2012-10-30 14:41 | 23.8073 | 19.8089 | 19.9412 | 19.7287 | 19.7167 | 45.477 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1535 | 3.1352 | 0.008285 | -0.3446 | -0.2123 | -0.4248 | -0.4368
2012-10-30 14:41 | 23.8035 | 19.8131 | 19.9334 | 19.777 | 19.7249 | 45.4707 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1817 | 3.1348 | 0.008283 | -0.3686 | -0.2483 | -0.4047 | -0.4568
2012-10-30 14:41 | 23.8035 | 19.8051 | 19.9654 | 19.769 | 19.7289 | 45.4695 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.1977 | 3.1347 | 0.008289 | -0.3926 | -0.2323 | -0.4287 | -0.4688
2012-10-30 14:41 | 23.8043 | 19.854 | 19.9622 | 19.7257 | 19.7418 | 45.4742 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1825 | 3.135 | 0.008285 | -0.3285 | -0.2203 | -0.4568 | -0.4407
2012-10-30 14:41 | 23.8065 | 19.7921 | 19.9203 | 19.76 | 19.7359 | 45.4759 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1367 | 3.1351 | 0.008283 | -0.3446 | -0.2164 | -0.3767 | -0.4008
2012-10-30 14:41 | 23.8065 | 19.8081 | 19.9644 | 19.768 | 19.7279 | 45.4718 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1687 | 3.1348 | 0.008281 | -0.3606 | -0.2043 | -0.4007 | -0.4408
2012-10-30 14:41 | 23.8043 | 19.7618 | 19.9181 | 19.7538 | 19.7017 | 45.4655 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1785 | 3.1344 | 0.008282 | -0.4167 | -0.2604 | -0.4247 | -0.4768
2012-10-30 14:42 | 23.808 | 19.7815 | 19.9379 | 19.7495 | 19.7254 | 45.4736 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1342 | 3.1349 | 0.008286 | -0.3527 | -0.1963 | -0.3847 | -0.4088
2012-10-30 14:42 | 23.8125 | 19.778 | 19.9504 | 19.786 | 19.7179 | 45.4557 | 0.014187 | 0.34547 | 0.004901 | 24.3512 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1226 | 3.1337 | 0.00828 | -0.3446 | -0.1722 | -0.3366 | -0.4047
2012-10-30 14:42 | 23.8155 | 19.8011 | 19.9454 | 19.7209 | 19.7289 | 45.484 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1817 | 3.1357 | 0.008283 | -0.3806 | -0.2363 | -0.4608 | -0.4528
2012-10-30 14:42 | 23.8148 | 19.8124 | 19.9527 | 19.7964 | 19.7442 | 45.4747 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.01442 | 0.23452 | 0.003382 | 16.2637 | 20.161 | 3.135 | 0.008293 | -0.3486 | -0.2083 | -0.3646 | -0.4168
2012-10-30 14:42 | 23.8133 | 19.8189 | 19.9992 | 19.7307 | 19.7026 | 45.4557 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.1515 | 3.1337 | 0.008286 | -0.3326 | -0.1523 | -0.4208 | -0.4489
2012-10-30 14:42 | 23.811 | 19.7805 | 19.9489 | 19.7164 | 19.7204 | 45.4678 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 20.1732 | 3.1346 | 0.00828 | -0.3927 | -0.2243 | -0.4568 | -0.4528
2012-10-30 14:42 | 23.8148 | 19.8044 | 19.9366 | 19.7603 | 19.7282 | 45.4591 | 0.014219 | 0.34541 | 0.004911 | 24.2923 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.209 | 3.134 | 0.008288 | -0.4046 | -0.2724 | -0.4487 | -0.4808
2012-10-30 14:42 | 23.8207 | 19.7742 | 19.9586 | 19.7582 | 19.7341 | 45.4736 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1549 | 3.1349 | 0.008291 | -0.3807 | -0.1963 | -0.3967 | -0.4208
2012-10-30 14:42 | 23.82 | 19.7414 | 19.9739 | 19.7374 | 19.7655 | 45.4666 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.1662 | 3.1345 | 0.008284 | -0.4248 | -0.1923 | -0.4288 | -0.4007
2012-10-30 14:42 | 23.8178 | 19.8234 | 19.9677 | 19.7513 | 19.6951 | 45.4661 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.156 | 3.1344 | 0.008287 | -0.3326 | -0.1883 | -0.4047 | -0.4609
2012-10-30 14:42 | 23.8133 | 19.8269 | 19.9592 | 19.7668 | 19.7107 | 45.4632 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.2075 | 3.1342 | 0.008288 | -0.3806 | -0.2483 | -0.4407 | -0.4968
2012-10-30 14:42 | 23.8155 | 19.785 | 19.9614 | 19.7329 | 19.7249 | 45.4695 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.1817 | 3.1347 | 0.008284 | -0.3967 | -0.2203 | -0.4488 | -0.4568
2012-10-30 14:42 | 23.8163 | 19.8419 | 19.9261 | 19.7738 | 19.7097 | 45.4609 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.01442 | 0.23452 | 0.003382 | 16.2637 | 20.1945 | 3.1341 | 0.008291 | -0.3526 | -0.2684 | -0.4207 | -0.4848
2012-10-30 14:42 | 23.8155 | 19.773 | 19.9494 | 19.753 | 19.7329 | 45.4678 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1937 | 3.1346 | 0.008288 | -0.4207 | -0.2443 | -0.4407 | -0.4608
2012-10-30 14:42 | 23.8185 | 19.8161 | 19.9524 | 19.76 | 19.7159 | 45.4609 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.2167 | 3.1341 | 0.008289 | -0.4006 | -0.2643 | -0.4567 | -0.5008
2012-10-30 14:42 | 23.8178 | 19.7994 | 19.9557 | 19.7873 | 19.7031 | 45.4695 | 0.014212 | 0.3455 | 0.00491 | 24.3104 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.2361 | 3.1347 | 0.008286 | -0.4367 | -0.2804 | -0.4488 | -0.533
2012-10-30 14:42 | 23.8275 | 19.785 | 19.9574 | 19.785 | 19.6968 | 45.4626 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.2418 | 3.1342 | 0.008287 | -0.4568 | -0.2844 | -0.4568 | -0.545
2012-10-30 14:42 | 23.8103 | 19.7678 | 19.9201 | 19.7157 | 19.7237 | 45.4522 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.2166 | 3.1335 | 0.008285 | -0.4488 | -0.2965 | -0.5009 | -0.4929
2012-10-30 14:42 | 23.8163 | 19.7618 | 19.9381 | 19.7297 | 19.7337 | 45.4626 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014417 | 0.23459 | 0.003382 | 16.2721 | 20.2105 | 3.1342 | 0.008287 | -0.4487 | -0.2724 | -0.4808 | -0.4768
2012-10-30 14:42 | 23.8222 | 19.8038 | 19.9721 | 19.7276 | 19.7276 | 45.4655 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2084 | 3.1344 | 0.008288 | -0.4046 | -0.2363 | -0.4808 | -0.4808
2012-10-30 14:43 | 23.82 | 19.8016 | 19.9539 | 19.7454 | 19.7214 | 45.4603 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014398 | 0.23459 | 0.003378 | 16.2936 | 20.2303 | 3.134 | 0.008282 | -0.4287 | -0.2764 | -0.4849 | -0.5089
2012-10-30 14:43 | 23.8207 | 19.7622 | 19.9425 | 19.7341 | 19.694 | 45.458 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1789 | 3.1339 | 0.008284 | -0.4167 | -0.2364 | -0.4448 | -0.4849
2012-10-30 14:43 | 23.8185 | 19.776 | 19.9163 | 19.7439 | 19.6958 | 45.451 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.1887 | 3.1334 | 0.008281 | -0.4127 | -0.2724 | -0.4448 | -0.4929
2012-10-30 14:43 | 23.82 | 19.7454 | 19.9258 | 19.7294 | 19.6933 | 45.4609 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1742 | 3.1341 | 0.008286 | -0.4288 | -0.2484 | -0.4448 | -0.4809
2012-10-30 14:43 | 23.8133 | 19.7508 | 19.9271 | 19.7548 | 19.7347 | 45.4597 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1835 | 3.134 | 0.008283 | -0.4327 | -0.2564 | -0.4287 | -0.4488
2012-10-30 14:43 | 23.8207 | 19.7902 | 19.9586 | 19.7221 | 19.694 | 45.462 | 0.014194 | 0.34538 | 0.004902 | 24.333 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.1709 | 3.1342 | 0.008283 | -0.3807 | -0.2123 | -0.4488 | -0.4769
2012-10-30 14:43 | 23.8118 | 19.7974 | 19.9296 | 19.7653 | 19.7172 | 45.4574 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.19 | 3.1338 | 0.008283 | -0.3926 | -0.2604 | -0.4247 | -0.4728
2012-10-30 14:43 | 23.814 | 19.7915 | 19.9879 | 19.7394 | 19.7114 | 45.4522 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2203 | 3.1335 | 0.008286 | -0.4288 | -0.2324 | -0.4809 | -0.5089
2012-10-30 14:43 | 23.823 | 19.7685 | 19.9729 | 19.7404 | 19.6763 | 45.4678 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.2373 | 3.1346 | 0.008286 | -0.4688 | -0.2644 | -0.4969 | -0.561
2012-10-30 14:43 | 23.8192 | 19.8248 | 19.901 | 19.7406 | 19.6845 | 45.4586 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1734 | 3.1339 | 0.008281 | -0.3486 | -0.2724 | -0.4328 | -0.4889
2012-10-30 14:43 | 23.8207 | 19.7822 | 19.9546 | 19.694 | 19.702 | 45.4614 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1749 | 3.1341 | 0.008285 | -0.3927 | -0.2203 | -0.4809 | -0.4729
2012-10-30 14:43 | 23.8245 | 19.798 | 19.9303 | 19.7379 | 19.7018 | 45.4539 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.2067 | 3.1336 | 0.008283 | -0.4087 | -0.2764 | -0.4688 | -0.5049
2012-10-30 14:43 | 23.823 | 19.7645 | 19.9168 | 19.7324 | 19.6763 | 45.4528 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1812 | 3.1335 | 0.008291 | -0.4167 | -0.2644 | -0.4488 | -0.5049
2012-10-30 14:43 | 23.8237 | 19.7411 | 19.9295 | 19.7291 | 19.689 | 45.462 | 0.014194 | 0.34537 | 0.004902 | 24.332 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1899 | 3.1342 | 0.008277 | -0.4488 | -0.2604 | -0.4608 | -0.5009
2012-10-30 14:43 | 23.8245 | 19.774 | 19.9103 | 19.6778 | 19.7098 | 45.4724 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1587 | 3.1349 | 0.008284 | -0.3847 | -0.2484 | -0.4809 | -0.4489
2012-10-30 14:43 | 23.8237 | 19.7692 | 19.9255 | 19.713 | 19.6449 | 45.4597 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014404 | 0.23445 | 0.003377 | 16.2767 | 20.23 | 3.134 | 0.008282 | -0.4608 | -0.3045 | -0.517 | -0.5851
2012-10-30 14:43 | 23.8215 | 19.7309 | 19.9313 | 19.7429 | 19.6788 | 45.458 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.1837 | 3.1339 | 0.008284 | -0.4528 | -0.2524 | -0.4408 | -0.5049
2012-10-30 14:43 | 23.82 | 19.7374 | 19.9298 | 19.7454 | 19.7093 | 45.4718 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1702 | 3.1348 | 0.008286 | -0.4328 | -0.2404 | -0.4248 | -0.4609
2012-10-30 14:43 | 23.823 | 19.7765 | 19.9328 | 19.7204 | 19.6723 | 45.4557 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.2012 | 3.1337 | 0.008289 | -0.4247 | -0.2684 | -0.4808 | -0.5289
2012-10-30 14:43 | 23.8185 | 19.7399 | 19.9323 | 19.7479 | 19.7038 | 45.4505 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.1847 | 3.1334 | 0.008283 | -0.4448 | -0.2524 | -0.4368 | -0.4809
2012-10-30 14:43 | 23.829 | 19.7785 | 19.9308 | 19.7584 | 19.6542 | 45.4539 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.1712 | 3.1336 | 0.008289 | -0.3927 | -0.2404 | -0.4128 | -0.517
2012-10-30 14:43 | 23.8192 | 19.7567 | 19.9571 | 19.7085 | 19.6845 | 45.4539 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.1894 | 3.1336 | 0.008283 | -0.4327 | -0.2323 | -0.4809 | -0.5049
2012-10-30 14:43 | 23.8327 | 19.7822 | 19.9265 | 19.7301 | 19.702 | 45.4609 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1589 | 3.1341 | 0.008288 | -0.3767 | -0.2324 | -0.4288 | -0.4569
2012-10-30 14:44 | 23.8222 | 19.7957 | 19.9 | 19.7717 | 19.6795 | 45.4528 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1924 | 3.1335 | 0.008285 | -0.3967 | -0.2924 | -0.4207 | -0.5129
2012-10-30 14:44 | 23.8267 | 19.7962 | 19.9325 | 19.716 | 19.7201 | 45.4568 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1849 | 3.1338 | 0.008286 | -0.3887 | -0.2524 | -0.4689 | -0.4648
2012-10-30 14:44 | 23.8282 | 19.7937 | 19.95 | 19.7216 | 19.7135 | 45.4516 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.2064 | 3.1334 | 0.008288 | -0.4127 | -0.2564 | -0.4848 | -0.4929
2012-10-30 14:44 | 23.817 | 19.7545 | 19.9228 | 19.7144 | 19.6703 | 45.4695 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1672 | 3.1347 | 0.008281 | -0.4127 | -0.2444 | -0.4528 | -0.4969
2012-10-30 14:44 | 23.8207 | 19.7501 | 19.9025 | 19.7181 | 19.702 | 45.4586 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1829 | 3.1339 | 0.00829 | -0.4328 | -0.2804 | -0.4648 | -0.4809
2012-10-30 14:44 | 23.8237 | 19.7251 | 19.9415 | 19.713 | 19.6689 | 45.4412 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1659 | 3.1327 | 0.008287 | -0.4408 | -0.2244 | -0.4529 | -0.497
2012-10-30 14:44 | 23.8275 | 19.789 | 19.9173 | 19.7289 | 19.6687 | 45.4586 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1817 | 3.1339 | 0.008284 | -0.3927 | -0.2644 | -0.4528 | -0.513
2012-10-30 14:44 | 23.82 | 19.7374 | 19.9218 | 19.6853 | 19.6733 | 45.4522 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1942 | 3.1335 | 0.008286 | -0.4568 | -0.2724 | -0.5089 | -0.5209
2012-10-30 14:44 | 23.8305 | 19.764 | 19.9042 | 19.6918 | 19.7399 | 45.4603 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1887 | 3.134 | 0.008282 | -0.4247 | -0.2845 | -0.4969 | -0.4488
2012-10-30 14:44 | 23.8237 | 19.7732 | 19.9375 | 19.693 | 19.681 | 45.4568 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1739 | 3.1338 | 0.008283 | -0.4007 | -0.2364 | -0.4809 | -0.4929
2012-10-30 14:44 | 23.8282 | 19.7617 | 19.8819 | 19.7015 | 19.6775 | 45.4464 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2104 | 3.1331 | 0.008286 | -0.4487 | -0.3285 | -0.5089 | -0.5329
2012-10-30 14:44 | 23.8275 | 19.7489 | 19.9373 | 19.7209 | 19.7008 | 45.4401 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1577 | 3.1326 | 0.008283 | -0.4088 | -0.2204 | -0.4368 | -0.4569
2012-10-30 14:44 | 23.8335 | 19.795 | 19.9313 | 19.7509 | 19.7148 | 45.4562 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.1717 | 3.1338 | 0.008285 | -0.3767 | -0.2404 | -0.4208 | -0.4569
2012-10-30 14:44 | 23.8252 | 19.7747 | 19.939 | 19.6825 | 19.7025 | 45.4574 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.1634 | 3.1338 | 0.008286 | -0.3887 | -0.2244 | -0.4809 | -0.4609
2012-10-30 14:44 | 23.8335 | 19.791 | 19.9353 | 19.7469 | 19.7108 | 45.4487 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.1637 | 3.1332 | 0.008288 | -0.3727 | -0.2284 | -0.4168 | -0.4529
2012-10-30 14:44 | 23.8387 | 19.7521 | 19.9485 | 19.712 | 19.72 | 45.4562 | 0.014187 | 0.34541 | 0.0049 | 24.3471 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.1689 | 3.1338 | 0.00828 | -0.4168 | -0.2204 | -0.4569 | -0.4489
2012-10-30 14:44 | 23.8335 | 19.7389 | 19.9113 | 19.6828 | 19.6747 | 45.4383 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.1516 | 3.1325 | 0.008285 | -0.4127 | -0.2403 | -0.4688 | -0.4769
2012-10-30 14:44 | 23.8267 | 19.7682 | 19.8844 | 19.6719 | 19.708 | 45.4678 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014423 | 0.23449 | 0.003382 | 16.2583 | 20.1569 | 3.1346 | 0.008288 | -0.3887 | -0.2725 | -0.485 | -0.4489
2012-10-30 14:44 | 23.835 | 19.7564 | 19.9128 | 19.7123 | 19.6923 | 45.4591 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.1772 | 3.134 | 0.008285 | -0.4208 | -0.2644 | -0.4649 | -0.4849
2012-10-30 14:44 | 23.8364 | 19.7819 | 19.9061 | 19.7137 | 19.6656 | 45.4574 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.1786 | 3.1338 | 0.008286 | -0.3967 | -0.2725 | -0.4649 | -0.513
2012-10-30 14:44 | 23.8394 | 19.7167 | 19.8971 | 19.6846 | 19.6846 | 45.4516 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1656 | 3.1334 | 0.008286 | -0.4489 | -0.2685 | -0.481 | -0.481
2012-10-30 14:44 | 23.8409 | 19.7503 | 19.9187 | 19.7503 | 19.6741 | 45.4539 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.1671 | 3.1336 | 0.008289 | -0.4168 | -0.2484 | -0.4168 | -0.493
2012-10-30 14:44 | 23.832 | 19.7454 | 19.8897 | 19.6893 | 19.6291 | 45.4597 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.01442 | 0.23455 | 0.003382 | 16.2657 | 20.1582 | 3.134 | 0.008288 | -0.4128 | -0.2685 | -0.4689 | -0.5291
2012-10-30 14:45 | 23.8432 | 19.7406 | 19.925 | 19.7205 | 19.6884 | 45.4562 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1894 | 3.1338 | 0.008288 | -0.4488 | -0.2644 | -0.4689 | -0.501
2012-10-30 14:45 | 23.8417 | 19.7471 | 19.8954 | 19.695 | 19.6509 | 45.4505 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.1959 | 3.1334 | 0.008283 | -0.4488 | -0.3005 | -0.5009 | -0.545
2012-10-30 14:45 | 23.8477 | 19.7852 | 19.8894 | 19.729 | 19.6889 | 45.4441 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2099 | 3.1329 | 0.008284 | -0.4247 | -0.3205 | -0.4809 | -0.521
2012-10-30 14:45 | 23.8477 | 19.7691 | 19.8974 | 19.7531 | 19.6849 | 45.4522 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.2339 | 3.1335 | 0.008286 | -0.4648 | -0.3365 | -0.4808 | -0.549
2012-10-30 14:45 | 23.8439 | 19.7934 | 19.9136 | 19.7092 | 19.625 | 45.458 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.2141 | 3.1339 | 0.00828 | -0.4207 | -0.3005 | -0.5049 | -0.5891
2012-10-30 14:45 | 23.8424 | 19.7678 | 19.8801 | 19.7157 | 19.6997 | 45.4499 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.2046 | 3.1333 | 0.008284 | -0.4368 | -0.3245 | -0.4889 | -0.5049
2012-10-30 14:45 | 23.8484 | 19.7778 | 19.8941 | 19.7257 | 19.6896 | 45.4516 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.2066 | 3.1334 | 0.008286 | -0.4288 | -0.3125 | -0.4809 | -0.517
2012-10-30 14:45 | 23.8439 | 19.7493 | 19.8736 | 19.7052 | 19.6691 | 45.4482 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1781 | 3.1332 | 0.008285 | -0.4288 | -0.3045 | -0.4729 | -0.509
2012-10-30 14:45 | 23.8507 | 19.7521 | 19.8964 | 19.6959 | 19.6879 | 45.4482 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2169 | 3.1332 | 0.00828 | -0.4648 | -0.3205 | -0.521 | -0.529
2012-10-30 14:45 | 23.8484 | 19.7418 | 19.9061 | 19.7097 | 19.6816 | 45.4534 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2066 | 3.1336 | 0.008283 | -0.4648 | -0.3005 | -0.4969 | -0.525
2012-10-30 14:45 | 23.8536 | 19.763 | 19.9153 | 19.7229 | 19.6628 | 45.4534 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014389 | 0.23457 | 0.003375 | 16.3018 | 20.2158 | 3.1336 | 0.00828 | -0.4528 | -0.3005 | -0.4929 | -0.553
2012-10-30 14:45 | 23.8514 | 19.7448 | 19.877 | 19.7327 | 19.7007 | 45.4406 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2256 | 3.1327 | 0.008279 | -0.4808 | -0.3486 | -0.4929 | -0.5249
2012-10-30 14:45 | 23.8529 | 19.7663 | 19.8986 | 19.7102 | 19.65 | 45.4377 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2031 | 3.1325 | 0.008283 | -0.4368 | -0.3045 | -0.4929 | -0.5531
2012-10-30 14:45 | 23.8454 | 19.7588 | 19.851 | 19.6866 | 19.7027 | 45.4551 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.1996 | 3.1337 | 0.008281 | -0.4408 | -0.3486 | -0.513 | -0.4969
2012-10-30 14:45 | 23.8521 | 19.7495 | 19.9058 | 19.7014 | 19.6733 | 45.4499 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.1903 | 3.1333 | 0.00828 | -0.4408 | -0.2845 | -0.4889 | -0.517
2012-10-30 14:45 | 23.8536 | 19.747 | 19.8873 | 19.6988 | 19.6788 | 45.4464 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2118 | 3.1331 | 0.008285 | -0.4648 | -0.3245 | -0.513 | -0.533
2012-10-30 14:45 | 23.8432 | 19.7646 | 19.8408 | 19.6844 | 19.6443 | 45.4447 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014395 | 0.23444 | 0.003375 | 16.2859 | 20.1814 | 3.133 | 0.008284 | -0.4168 | -0.3406 | -0.497 | -0.5371
2012-10-30 14:45 | 23.8551 | 19.7725 | 19.8848 | 19.7164 | 19.6763 | 45.4424 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1813 | 3.1328 | 0.00828 | -0.4088 | -0.2965 | -0.4649 | -0.505
2012-10-30 14:45 | 23.8544 | 19.7598 | 19.8841 | 19.7277 | 19.6796 | 45.4534 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.2046 | 3.1336 | 0.008282 | -0.4448 | -0.3205 | -0.4769 | -0.525
2012-10-30 14:45 | 23.8551 | 19.7404 | 19.8888 | 19.6843 | 19.6522 | 45.4464 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.01442 | 0.23449 | 0.003381 | 16.2617 | 20.1893 | 3.1331 | 0.008288 | -0.4489 | -0.3005 | -0.505 | -0.5371
2012-10-30 14:45 | 23.8536 | 19.7149 | 19.8993 | 19.6748 | 19.6708 | 45.451 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.1758 | 3.1334 | 0.008286 | -0.4609 | -0.2765 | -0.501 | -0.505
2012-10-30 14:45 | 23.8559 | 19.7813 | 19.9136 | 19.6971 | 19.6931 | 45.4406 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2141 | 3.1327 | 0.008284 | -0.4328 | -0.3005 | -0.517 | -0.521
2012-10-30 14:45 | 23.8559 | 19.7412 | 19.8896 | 19.661 | 19.645 | 45.4482 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.2261 | 3.1332 | 0.008282 | -0.4849 | -0.3365 | -0.5651 | -0.5811
2012-10-30 14:46 | 23.8589 | 19.7202 | 19.8886 | 19.7322 | 19.66 | 45.4568 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014411 | 0.23442 | 0.003378 | 16.2668 | 20.2211 | 3.1338 | 0.008283 | -0.5009 | -0.3325 | -0.4889 | -0.5611
2012-10-30 14:46 | 23.8499 | 19.7072 | 19.8876 | 19.7072 | 19.6671 | 45.4534 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 20.2241 | 3.1336 | 0.008288 | -0.5169 | -0.3365 | -0.5169 | -0.557
2012-10-30 14:46 | 23.8611 | 19.7384 | 19.8667 | 19.6903 | 19.6903 | 45.4297 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.2674 | 3.1319 | 0.008284 | -0.529 | -0.4007 | -0.5771 | -0.5771
2012-10-30 14:46 | 23.8574 | 19.7267 | 19.9031 | 19.6866 | 19.6866 | 45.4568 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2637 | 3.1338 | 0.008283 | -0.537 | -0.3606 | -0.5771 | -0.5771
2012-10-30 14:46 | 23.8611 | 19.7023 | 19.8908 | 19.6983 | 19.6743 | 45.447 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2714 | 3.1331 | 0.008284 | -0.5691 | -0.3806 | -0.5731 | -0.5971
2012-10-30 14:46 | 23.8581 | 19.7194 | 19.9158 | 19.7274 | 19.6873 | 45.4539 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.2323 | 3.1336 | 0.008283 | -0.5129 | -0.3165 | -0.5049 | -0.545
2012-10-30 14:46 | 23.8544 | 19.7317 | 19.9041 | 19.6676 | 19.6836 | 45.4424 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1966 | 3.1328 | 0.008287 | -0.4649 | -0.2925 | -0.529 | -0.513
2012-10-30 14:46 | 23.8566 | 19.7339 | 19.8863 | 19.7018 | 19.6417 | 45.4429 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.2228 | 3.1328 | 0.008286 | -0.4889 | -0.3365 | -0.521 | -0.5811
2012-10-30 14:46 | 23.8604 | 19.7297 | 19.9221 | 19.6816 | 19.6776 | 45.4395 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1946 | 3.1326 | 0.008281 | -0.4649 | -0.2725 | -0.513 | -0.517
2012-10-30 14:46 | 23.8626 | 19.7439 | 19.9043 | 19.6477 | 19.6316 | 45.4597 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 20.2008 | 3.134 | 0.008279 | -0.4569 | -0.2965 | -0.5531 | -0.5692
2012-10-30 14:46 | 23.8529 | 19.7382 | 19.8785 | 19.6661 | 19.6661 | 45.4487 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.131 | 3.1332 | 0.008285 | -0.3928 | -0.2525 | -0.4649 | -0.4649
2012-10-30 14:46 | 23.8641 | 19.7414 | 19.9098 | 19.6773 | 19.6692 | 45.4499 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1822 | 3.1333 | 0.008282 | -0.4408 | -0.2724 | -0.5049 | -0.513
2012-10-30 14:46 | 23.8679 | 19.7412 | 19.8655 | 19.6931 | 19.653 | 45.4401 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.2061 | 3.1326 | 0.008285 | -0.4649 | -0.3406 | -0.513 | -0.5531
2012-10-30 14:46 | 23.8656 | 19.7108 | 19.9153 | 19.6988 | 19.6828 | 45.4493 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.1637 | 3.1333 | 0.00828 | -0.4529 | -0.2484 | -0.4649 | -0.4809
2012-10-30 14:46 | 23.8596 | 19.7449 | 19.8412 | 19.6888 | 19.6647 | 45.4458 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1938 | 3.133 | 0.008286 | -0.4489 | -0.3526 | -0.505 | -0.5291
2012-10-30 14:46 | 23.8611 | 19.7224 | 19.8467 | 19.7063 | 19.6662 | 45.4487 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.2073 | 3.1332 | 0.008283 | -0.4849 | -0.3606 | -0.501 | -0.5411
2012-10-30 14:46 | 23.8656 | 19.7429 | 19.9113 | 19.6748 | 19.7108 | 45.4406 | 0.014194 | 0.3454 | 0.004903 | 24.3341 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.2078 | 3.1327 | 0.008281 | -0.4649 | -0.2965 | -0.533 | -0.497
2012-10-30 14:46 | 23.8671 | 19.7645 | 19.8968 | 19.7043 | 19.6763 | 45.4476 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.1852 | 3.1332 | 0.008284 | -0.4207 | -0.2884 | -0.4809 | -0.5089
2012-10-30 14:46 | 23.8693 | 19.7145 | 19.8629 | 19.6785 | 19.6985 | 45.4458 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.01442 | 0.23457 | 0.003382 | 16.2667 | 20.2115 | 3.133 | 0.008294 | -0.497 | -0.3486 | -0.533 | -0.513
2012-10-30 14:46 | 23.8649 | 19.7583 | 19.8905 | 19.7222 | 19.638 | 45.4539 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014411 | 0.23445 | 0.003379 | 16.2688 | 20.2311 | 3.1336 | 0.008287 | -0.4728 | -0.3406 | -0.5089 | -0.5931
2012-10-30 14:46 | 23.8656 | 19.6908 | 19.8792 | 19.6547 | 19.6186 | 45.4429 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2038 | 3.1328 | 0.008281 | -0.513 | -0.3246 | -0.5491 | -0.5852
2012-10-30 14:46 | 23.8701 | 19.7394 | 19.8717 | 19.6833 | 19.6833 | 45.4476 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2203 | 3.1332 | 0.008284 | -0.4809 | -0.3486 | -0.537 | -0.537
2012-10-30 14:46 | 23.8708 | 19.7521 | 19.8724 | 19.6599 | 19.704 | 45.4528 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1769 | 3.1335 | 0.008282 | -0.4248 | -0.3045 | -0.517 | -0.4729
2012-10-30 14:47 | 23.8723 | 19.7456 | 19.8699 | 19.7055 | 19.6654 | 45.4487 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014411 | 0.23444 | 0.003378 | 16.2679 | 20.2105 | 3.1332 | 0.008288 | -0.4649 | -0.3406 | -0.505 | -0.5451
2012-10-30 14:47 | 23.8641 | 19.7174 | 19.8416 | 19.7013 | 19.7134 | 45.4476 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.1742 | 3.1332 | 0.008285 | -0.4568 | -0.3326 | -0.4729 | -0.4608
2012-10-30 14:47 | 23.8619 | 19.7071 | 19.8555 | 19.6269 | 19.6751 | 45.4441 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 20.1961 | 3.1329 | 0.008288 | -0.489 | -0.3406 | -0.5692 | -0.521
2012-10-30 14:47 | 23.8716 | 19.7249 | 19.8972 | 19.7128 | 19.6687 | 45.4487 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.2658 | 3.1332 | 0.008285 | -0.5409 | -0.3686 | -0.553 | -0.5971
2012-10-30 14:47 | 23.8679 | 19.7172 | 19.8735 | 19.6811 | 19.661 | 45.4487 | 0.014187 | 0.34541 | 0.0049 | 24.3471 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.186 | 3.1332 | 0.008275 | -0.4688 | -0.3125 | -0.5049 | -0.525
2012-10-30 14:47 | 23.8693 | 19.7426 | 19.8549 | 19.6865 | 19.6825 | 45.4424 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.2235 | 3.1328 | 0.008282 | -0.4809 | -0.3686 | -0.537 | -0.541
2012-10-30 14:47 | 23.8671 | 19.7484 | 19.9088 | 19.7123 | 19.6562 | 45.4476 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.1893 | 3.1332 | 0.008284 | -0.4409 | -0.2805 | -0.477 | -0.5331
2012-10-30 14:47 | 23.8656 | 19.7229 | 19.8832 | 19.6707 | 19.6788 | 45.4372 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1918 | 3.1324 | 0.008283 | -0.4689 | -0.3086 | -0.5211 | -0.513
2012-10-30 14:47 | 23.8634 | 19.7487 | 19.881 | 19.6846 | 19.6645 | 45.4464 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1815 | 3.1331 | 0.008282 | -0.4328 | -0.3005 | -0.4969 | -0.517
2012-10-30 14:47 | 23.8649 | 19.7262 | 19.9146 | 19.6821 | 19.7021 | 45.436 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.2151 | 3.1324 | 0.008283 | -0.4889 | -0.3005 | -0.533 | -0.513
2012-10-30 14:47 | 23.8656 | 19.7189 | 19.8993 | 19.6788 | 19.6467 | 45.4534 | 0.014187 | 0.34543 | 0.004901 | 24.3481 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.1918 | 3.1336 | 0.00828 | -0.4729 | -0.2925 | -0.513 | -0.5451
2012-10-30 14:47 | 23.8723 | 19.7015 | 19.918 | 19.7175 | 19.6855 | 45.4453 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.2105 | 3.133 | 0.008285 | -0.509 | -0.2925 | -0.493 | -0.525
2012-10-30 14:47 | 23.8581 | 19.7033 | 19.8878 | 19.6913 | 19.6472 | 45.4453 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.1803 | 3.133 | 0.008284 | -0.477 | -0.2925 | -0.489 | -0.5331
2012-10-30 14:47 | 23.8679 | 19.7011 | 19.8895 | 19.645 | 19.669 | 45.4291 | 0.014231 | 0.3454 | 0.004915 | 24.2708 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.154 | 3.1319 | 0.008292 | -0.4529 | -0.2645 | -0.509 | -0.485
2012-10-30 14:47 | 23.8626 | 19.6878 | 19.8802 | 19.7078 | 19.6637 | 45.458 | 0.014187 | 0.3454 | 0.0049 | 24.3461 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1727 | 3.1339 | 0.008277 | -0.4849 | -0.2925 | -0.4649 | -0.509
2012-10-30 14:47 | 23.8566 | 19.7379 | 19.9384 | 19.6537 | 19.6497 | 45.4389 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1387 | 3.1326 | 0.008283 | -0.4008 | -0.2003 | -0.485 | -0.489
2012-10-30 14:47 | 23.8634 | 19.7367 | 19.9091 | 19.6966 | 19.6645 | 45.4447 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1815 | 3.133 | 0.008279 | -0.4448 | -0.2724 | -0.4849 | -0.517
2012-10-30 14:47 | 23.8634 | 19.7167 | 19.865 | 19.6846 | 19.6926 | 45.4424 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.1535 | 3.1328 | 0.00828 | -0.4368 | -0.2885 | -0.4689 | -0.4609
2012-10-30 14:47 | 23.8649 | 19.7342 | 19.8986 | 19.6781 | 19.6861 | 45.4482 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.163 | 3.1332 | 0.008286 | -0.4288 | -0.2644 | -0.4849 | -0.4769
2012-10-30 14:47 | 23.8679 | 19.7572 | 19.9016 | 19.673 | 19.7091 | 45.4389 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.2061 | 3.1326 | 0.008283 | -0.4489 | -0.3045 | -0.5331 | -0.497
2012-10-30 14:47 | 23.8664 | 19.7197 | 19.9281 | 19.6756 | 19.6796 | 45.4256 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1926 | 3.1316 | 0.00828 | -0.4729 | -0.2645 | -0.517 | -0.513
2012-10-30 14:47 | 23.8634 | 19.6926 | 19.889 | 19.6525 | 19.7086 | 45.4429 | 0.014194 | 0.34538 | 0.004902 | 24.333 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1936 | 3.1328 | 0.008281 | -0.501 | -0.3046 | -0.5411 | -0.485
2012-10-30 14:47 | 23.8641 | 19.7414 | 19.9138 | 19.6532 | 19.6612 | 45.4366 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1782 | 3.1324 | 0.008285 | -0.4368 | -0.2644 | -0.525 | -0.517
2012-10-30 14:47 | 23.8619 | 19.7352 | 19.8996 | 19.6911 | 19.659 | 45.4476 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.2161 | 3.1332 | 0.008289 | -0.4809 | -0.3165 | -0.525 | -0.5571
2012-10-30 14:48 | 23.8529 | 19.7142 | 19.9026 | 19.6621 | 19.646 | 45.4337 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.155 | 3.1322 | 0.008279 | -0.4408 | -0.2524 | -0.4929 | -0.509
2012-10-30 14:48 | 23.8611 | 19.7104 | 19.8908 | 19.6462 | 19.6703 | 45.4447 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1752 | 3.133 | 0.008281 | -0.4648 | -0.2844 | -0.529 | -0.5049
2012-10-30 14:48 | 23.8611 | 19.7184 | 19.8908 | 19.6783 | 19.6422 | 45.4401 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014411 | 0.23444 | 0.003378 | 16.2679 | 20.1752 | 3.1326 | 0.008283 | -0.4568 | -0.2844 | -0.4969 | -0.533
2012-10-30 14:48 | 23.8544 | 19.6916 | 19.9041 | 19.6636 | 19.6716 | 45.4418 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.1926 | 3.1328 | 0.008281 | -0.501 | -0.2885 | -0.529 | -0.521
2012-10-30 14:48 | 23.8604 | 19.6976 | 19.9101 | 19.6816 | 19.6575 | 45.4389 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.1906 | 3.1326 | 0.008287 | -0.493 | -0.2805 | -0.509 | -0.5331
2012-10-30 14:48 | 23.8656 | 19.7349 | 19.8832 | 19.7108 | 19.6748 | 45.4551 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.1998 | 3.1337 | 0.008287 | -0.4649 | -0.3166 | -0.489 | -0.525
2012-10-30 14:48 | 23.8596 | 19.7249 | 19.8893 | 19.6688 | 19.6688 | 45.4429 | 0.01419 | 0.34538 | 0.004901 | 24.3399 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.1898 | 3.1328 | 0.008281 | -0.4649 | -0.3005 | -0.521 | -0.521
2012-10-30 14:48 | 23.8514 | 19.6806 | 19.8891 | 19.6405 | 19.6485 | 45.432 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.2136 | 3.1321 | 0.008283 | -0.533 | -0.3245 | -0.5731 | -0.5651
2012-10-30 14:48 | 23.8634 | 19.6926 | 19.9051 | 19.6766 | 19.7046 | 45.4389 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23445 | 0.003379 | 16.2688 | 20.1735 | 3.1326 | 0.008285 | -0.4809 | -0.2684 | -0.4969 | -0.4689
2012-10-30 14:48 | 23.8596 | 19.7089 | 19.9293 | 19.6487 | 19.6487 | 45.4372 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.2138 | 3.1324 | 0.00828 | -0.5049 | -0.2845 | -0.5651 | -0.5651
2012-10-30 14:48 | 23.8604 | 19.7297 | 19.9301 | 19.6816 | 19.6816 | 45.4372 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1745 | 3.1324 | 0.008284 | -0.4448 | -0.2444 | -0.4929 | -0.4929
2012-10-30 14:48 | 23.8551 | 19.7124 | 19.9369 | 19.7003 | 19.6482 | 45.4377 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014404 | 0.23445 | 0.003377 | 16.2767 | 20.1933 | 3.1325 | 0.008283 | -0.4809 | -0.2564 | -0.493 | -0.5451
2012-10-30 14:48 | 23.8649 | 19.6941 | 19.8986 | 19.6821 | 19.658 | 45.4308 | 0.014187 | 0.34546 | 0.004901 | 24.3502 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2431 | 3.132 | 0.00828 | -0.549 | -0.3445 | -0.561 | -0.5851
2012-10-30 14:48 | 23.8619 | 19.7152 | 19.8835 | 19.655 | 19.663 | 45.4424 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014417 | 0.23455 | 0.003382 | 16.2691 | 20.2081 | 3.1328 | 0.008287 | -0.4929 | -0.3246 | -0.5531 | -0.5451
2012-10-30 14:48 | 23.8596 | 19.7129 | 19.9053 | 19.6567 | 19.6728 | 45.4308 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.1938 | 3.132 | 0.008288 | -0.4809 | -0.2885 | -0.5371 | -0.521
2012-10-30 14:48 | 23.8589 | 19.7082 | 19.8445 | 19.664 | 19.64 | 45.4314 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.173 | 3.132 | 0.008287 | -0.4648 | -0.3285 | -0.509 | -0.533
2012-10-30 14:48 | 23.8559 | 19.7533 | 19.8695 | 19.6851 | 19.6651 | 45.432 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014392 | 0.23441 | 0.003374 | 16.2873 | 20.174 | 3.1321 | 0.008283 | -0.4207 | -0.3045 | -0.4889 | -0.5089
2012-10-30 14:48 | 23.8611 | 19.7144 | 19.8507 | 19.6662 | 19.6582 | 45.4401 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1913 | 3.1326 | 0.008287 | -0.4769 | -0.3406 | -0.5251 | -0.5331
2012-10-30 14:48 | 23.8596 | 19.7289 | 19.8692 | 19.6728 | 19.6527 | 45.4424 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1858 | 3.1328 | 0.008284 | -0.4569 | -0.3166 | -0.513 | -0.5331
2012-10-30 14:48 | 23.8589 | 19.7442 | 19.8685 | 19.6801 | 19.64 | 45.4389 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.149 | 3.1326 | 0.008286 | -0.4048 | -0.2805 | -0.4689 | -0.509
2012-10-30 14:48 | 23.8611 | 19.7424 | 19.8867 | 19.6863 | 19.6703 | 45.4302 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.1552 | 3.132 | 0.00828 | -0.4128 | -0.2685 | -0.4689 | -0.4849
2012-10-30 14:48 | 23.8664 | 19.7197 | 19.856 | 19.6435 | 19.6034 | 45.4198 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014417 | 0.23442 | 0.00338 | 16.26 | 20.1765 | 3.1312 | 0.008287 | -0.4568 | -0.3205 | -0.533 | -0.5731
2012-10-30 14:48 | 23.8671 | 19.7043 | 19.8807 | 19.6682 | 19.6522 | 45.4268 | 0.014216 | 0.34537 | 0.00491 | 24.2944 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.1532 | 3.1317 | 0.008284 | -0.4489 | -0.2725 | -0.485 | -0.501
2012-10-30 14:49 | 23.8656 | 19.6988 | 19.8592 | 19.6788 | 19.6346 | 45.4291 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1677 | 3.1319 | 0.008284 | -0.4689 | -0.3085 | -0.4889 | -0.5331
2012-10-30 14:49 | 23.8693 | 19.7346 | 19.8709 | 19.6905 | 19.6343 | 45.4383 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014389 | 0.23444 | 0.003373 | 16.2927 | 20.2075 | 3.1325 | 0.008278 | -0.4729 | -0.3366 | -0.517 | -0.5732
2012-10-30 14:49 | 23.8753 | 19.6925 | 19.8889 | 19.6724 | 19.7246 | 45.4297 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2175 | 3.1319 | 0.008289 | -0.525 | -0.3286 | -0.5451 | -0.4929
2012-10-30 14:49 | 23.8686 | 19.7098 | 19.8461 | 19.6697 | 19.6457 | 45.4279 | 0.014219 | 0.34536 | 0.004911 | 24.2883 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1988 | 3.1318 | 0.008289 | -0.489 | -0.3527 | -0.5291 | -0.5531
2012-10-30 14:49 | 23.8679 | 19.6971 | 19.8815 | 19.6971 | 19.6329 | 45.4534 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.174 | 3.1336 | 0.008286 | -0.4769 | -0.2925 | -0.4769 | -0.5411
2012-10-30 14:49 | 23.8708 | 19.712 | 19.8724 | 19.6639 | 19.6559 | 45.4331 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.213 | 3.1322 | 0.008284 | -0.501 | -0.3406 | -0.5491 | -0.5571
2012-10-30 14:49 | 23.8716 | 19.7249 | 19.8972 | 19.6767 | 19.6767 | 45.4193 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.2138 | 3.1312 | 0.008283 | -0.4889 | -0.3166 | -0.5371 | -0.5371
2012-10-30 14:49 | 23.8679 | 19.7292 | 19.8535 | 19.673 | 19.673 | 45.4389 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.1981 | 3.1326 | 0.008287 | -0.4689 | -0.3446 | -0.5251 | -0.5251
2012-10-30 14:49 | 23.8708 | 19.7201 | 19.8764 | 19.6519 | 19.6398 | 45.436 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.229 | 3.1324 | 0.008281 | -0.5089 | -0.3526 | -0.5771 | -0.5892
2012-10-30 14:49 | 23.8693 | 19.7306 | 19.8749 | 19.7105 | 19.6544 | 45.4302 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1754 | 3.132 | 0.008285 | -0.4448 | -0.3005 | -0.4649 | -0.521
2012-10-30 14:49 | 23.8626 | 19.6998 | 19.8321 | 19.6557 | 19.6196 | 45.4273 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2008 | 3.1318 | 0.008284 | -0.501 | -0.3687 | -0.5451 | -0.5812
2012-10-30 14:49 | 23.8679 | 19.7332 | 19.8695 | 19.6851 | 19.6329 | 45.4273 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.166 | 3.1318 | 0.008282 | -0.4328 | -0.2965 | -0.4809 | -0.5331
2012-10-30 14:49 | 23.8753 | 19.7165 | 19.909 | 19.6804 | 19.6403 | 45.4302 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.1814 | 3.132 | 0.00828 | -0.4649 | -0.2724 | -0.501 | -0.5411
2012-10-30 14:49 | 23.8641 | 19.7174 | 19.8537 | 19.6933 | 19.6331 | 45.4406 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.1662 | 3.1327 | 0.008282 | -0.4488 | -0.3125 | -0.4729 | -0.5331
2012-10-30 14:49 | 23.8723 | 19.7376 | 19.8699 | 19.6454 | 19.6654 | 45.4221 | 0.014194 | 0.3454 | 0.004903 | 24.3341 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.1864 | 3.1314 | 0.00828 | -0.4488 | -0.3165 | -0.541 | -0.521
2012-10-30 14:49 | 23.8671 | 19.7043 | 19.9128 | 19.6803 | 19.6321 | 45.4424 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.1332 | 3.1328 | 0.008283 | -0.4289 | -0.2204 | -0.4529 | -0.5011
2012-10-30 14:49 | 23.8723 | 19.7216 | 19.8498 | 19.6293 | 19.6494 | 45.4395 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014417 | 0.23445 | 0.00338 | 16.2621 | 20.2385 | 3.1326 | 0.008289 | -0.5169 | -0.3887 | -0.6092 | -0.5891
2012-10-30 14:50 | 23.8701 | 19.7434 | 19.8877 | 19.6592 | 19.6672 | 45.436 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.1762 | 3.1324 | 0.008281 | -0.4328 | -0.2885 | -0.517 | -0.509
2012-10-30 14:50 | 23.8716 | 19.6928 | 19.8812 | 19.6808 | 19.6126 | 45.4204 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2018 | 3.1313 | 0.008285 | -0.509 | -0.3206 | -0.521 | -0.5892
2012-10-30 14:50 | 23.8656 | 19.6948 | 19.8351 | 19.6427 | 19.6667 | 45.4377 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.2038 | 3.1325 | 0.008288 | -0.509 | -0.3687 | -0.5611 | -0.5371
2012-10-30 14:50 | 23.8656 | 19.6788 | 19.8792 | 19.6507 | 19.6627 | 45.432 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.2238 | 3.1321 | 0.008288 | -0.545 | -0.3446 | -0.5731 | -0.5611
2012-10-30 14:50 | 23.8664 | 19.7157 | 19.8399 | 19.6675 | 19.6435 | 45.4297 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2126 | 3.1319 | 0.008287 | -0.4969 | -0.3727 | -0.5451 | -0.5691
2012-10-30 14:50 | 23.8641 | 19.7053 | 19.8857 | 19.6532 | 19.6331 | 45.4262 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.1903 | 3.1317 | 0.008287 | -0.485 | -0.3046 | -0.5371 | -0.5572
2012-10-30 14:50 | 23.8671 | 19.6883 | 19.8727 | 19.6642 | 19.6482 | 45.4112 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1732 | 3.1306 | 0.008288 | -0.4849 | -0.3005 | -0.509 | -0.525
2012-10-30 14:50 | 23.8693 | 19.7105 | 19.8949 | 19.6343 | 19.6825 | 45.4262 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.1674 | 3.1317 | 0.00828 | -0.4569 | -0.2725 | -0.5331 | -0.4849
2012-10-30 14:50 | 23.8679 | 19.7051 | 19.8976 | 19.653 | 19.661 | 45.4273 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.186 | 3.1318 | 0.008285 | -0.4809 | -0.2884 | -0.533 | -0.525
2012-10-30 14:50 | 23.8693 | 19.7386 | 19.8669 | 19.6504 | 19.6223 | 45.4198 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2195 | 3.1312 | 0.008287 | -0.4809 | -0.3526 | -0.5691 | -0.5972
2012-10-30 14:50 | 23.8731 | 19.7264 | 19.9068 | 19.6823 | 19.6381 | 45.4262 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.1672 | 3.1317 | 0.008284 | -0.4408 | -0.2604 | -0.4849 | -0.5291
2012-10-30 14:50 | 23.8753 | 19.7246 | 19.8809 | 19.6443 | 19.6604 | 45.4285 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014426 | 0.23452 | 0.003383 | 16.257 | 20.2215 | 3.1318 | 0.008289 | -0.4969 | -0.3406 | -0.5772 | -0.5611
2012-10-30 14:50 | 23.8776 | 19.7269 | 19.8592 | 19.6306 | 19.6707 | 45.4308 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014392 | 0.23442 | 0.003374 | 16.2883 | 20.2478 | 3.132 | 0.00828 | -0.5209 | -0.3886 | -0.6172 | -0.5771
2012-10-30 14:50 | 23.8716 | 19.7048 | 19.8972 | 19.6607 | 19.6848 | 45.4233 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.2338 | 3.1315 | 0.008287 | -0.529 | -0.3366 | -0.5731 | -0.549
2012-10-30 14:50 | 23.8798 | 19.693 | 19.8934 | 19.6849 | 19.6809 | 45.425 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.222 | 3.1316 | 0.008284 | -0.529 | -0.3286 | -0.5371 | -0.5411
2012-10-30 14:50 | 23.8746 | 19.7078 | 19.8682 | 19.6717 | 19.6396 | 45.4227 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2448 | 3.1314 | 0.008279 | -0.537 | -0.3766 | -0.5731 | -0.6052
2012-10-30 14:50 | 23.8776 | 19.6988 | 19.8792 | 19.6507 | 19.6707 | 45.4314 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014389 | 0.23448 | 0.003374 | 16.2957 | 20.2118 | 3.132 | 0.008284 | -0.513 | -0.3326 | -0.5611 | -0.5411
2012-10-30 14:50 | 23.8753 | 19.7646 | 19.8609 | 19.6524 | 19.6484 | 45.4129 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.2175 | 3.1308 | 0.008291 | -0.4529 | -0.3566 | -0.5651 | -0.5691
2012-10-30 14:50 | 23.8768 | 19.706 | 19.8944 | 19.686 | 19.69 | 45.4216 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.215 | 3.1314 | 0.008282 | -0.509 | -0.3206 | -0.529 | -0.525
2012-10-30 14:50 | 23.8679 | 19.6971 | 19.8535 | 19.6129 | 19.6329 | 45.4216 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2141 | 3.1314 | 0.008284 | -0.517 | -0.3606 | -0.6012 | -0.5812
2012-10-30 14:50 | 23.8716 | 19.6848 | 19.8491 | 19.6527 | 19.6366 | 45.4193 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014408 | 0.23444 | 0.003378 | 16.2712 | 20.1857 | 3.1312 | 0.008282 | -0.5009 | -0.3366 | -0.533 | -0.5491
2012-10-30 14:50 | 23.8776 | 19.7509 | 19.8792 | 19.6747 | 19.6306 | 45.4343 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1917 | 3.1322 | 0.00829 | -0.4408 | -0.3125 | -0.517 | -0.5611
2012-10-30 14:50 | 23.8791 | 19.7364 | 19.8967 | 19.6481 | 19.6602 | 45.4152 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.2213 | 3.1309 | 0.008287 | -0.4849 | -0.3246 | -0.5732 | -0.5611
2012-10-30 14:51 | 23.8813 | 19.7065 | 19.8548 | 19.6544 | 19.6704 | 45.4158 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.1914 | 3.131 | 0.008287 | -0.4849 | -0.3366 | -0.537 | -0.521
2012-10-30 14:51 | 23.8776 | 19.6948 | 19.8752 | 19.6787 | 19.6707 | 45.4297 | 0.014206 | 0.34536 | 0.004906 | 24.3105 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.2158 | 3.1319 | 0.008283 | -0.521 | -0.3406 | -0.5371 | -0.5451
2012-10-30 14:51 | 23.8768 | 19.706 | 19.8704 | 19.6178 | 19.6579 | 45.4279 | 0.014197 | 0.34538 | 0.004903 | 24.3279 | 0.014401 | 0.23441 | 0.003376 | 16.2771 | 20.219 | 3.1318 | 0.008279 | -0.513 | -0.3486 | -0.6012 | -0.5611
2012-10-30 14:51 | 23.8753 | 19.7085 | 19.8488 | 19.6203 | 19.6323 | 45.4164 | 0.014203 | 0.34534 | 0.004905 | 24.3146 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.1894 | 3.131 | 0.008284 | -0.4809 | -0.3406 | -0.5691 | -0.5571
2012-10-30 14:51 | 23.8843 | 19.7175 | 19.8739 | 19.6694 | 19.6814 | 45.4245 | 0.0142 | 0.34533 | 0.004904 | 24.3187 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.1824 | 3.1316 | 0.008282 | -0.4649 | -0.3085 | -0.513 | -0.501
2012-10-30 14:51 | 23.8798 | 19.697 | 19.9135 | 19.6529 | 19.6408 | 45.4221 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1979 | 3.1314 | 0.008282 | -0.5009 | -0.2844 | -0.545 | -0.5571
2012-10-30 14:51 | 23.8783 | 19.7075 | 19.8679 | 19.6634 | 19.6112 | 45.4325 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.2205 | 3.1321 | 0.008283 | -0.513 | -0.3526 | -0.5571 | -0.6093
2012-10-30 14:51 | 23.8776 | 19.7228 | 19.8632 | 19.6266 | 19.6587 | 45.4291 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23441 | 0.003378 | 16.2658 | 20.1998 | 3.1319 | 0.008284 | -0.477 | -0.3366 | -0.5732 | -0.5411
2012-10-30 14:51 | 23.8791 | 19.7123 | 19.8887 | 19.6481 | 19.6682 | 45.4146 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.2373 | 3.1309 | 0.008284 | -0.525 | -0.3486 | -0.5892 | -0.5691
2012-10-30 14:51 | 23.8813 | 19.6985 | 19.8669 | 19.6624 | 19.6584 | 45.4112 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2235 | 3.1306 | 0.008283 | -0.525 | -0.3566 | -0.5611 | -0.5651
2012-10-30 14:51 | 23.8858 | 19.6909 | 19.8754 | 19.6669 | 19.6508 | 45.4291 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.248 | 3.1319 | 0.00828 | -0.5571 | -0.3726 | -0.5811 | -0.5972
2012-10-30 14:51 | 23.8798 | 19.693 | 19.8734 | 19.6448 | 19.693 | 45.421 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.2941 | 3.1313 | 0.008282 | -0.6011 | -0.4207 | -0.6493 | -0.6011
2012-10-30 14:51 | 23.8791 | 19.6762 | 19.8727 | 19.6241 | 19.6762 | 45.4193 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.2613 | 3.1312 | 0.008282 | -0.5851 | -0.3886 | -0.6372 | -0.5851
2012-10-30 14:51 | 23.8836 | 19.6847 | 19.8652 | 19.6526 | 19.6526 | 45.4239 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.2338 | 3.1315 | 0.00828 | -0.5491 | -0.3686 | -0.5812 | -0.5812
2012-10-30 14:51 | 23.8851 | 19.6862 | 19.9107 | 19.6301 | 19.6501 | 45.4256 | 0.014209 | 0.34537 | 0.004907 | 24.3064 | 0.014404 | 0.23444 | 0.003377 | 16.2758 | 20.2433 | 3.1316 | 0.008284 | -0.5571 | -0.3326 | -0.6132 | -0.5932
2012-10-30 14:51 | 23.8806 | 19.7018 | 19.8141 | 19.6777 | 19.6496 | 45.4297 | 0.014187 | 0.34538 | 0.0049 | 24.345 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2348 | 3.1319 | 0.008278 | -0.533 | -0.4207 | -0.5571 | -0.5852
2012-10-30 14:51 | 23.888 | 19.6651 | 19.8896 | 19.6891 | 19.641 | 45.4083 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014408 | 0.23441 | 0.003377 | 16.2692 | 20.2702 | 3.1304 | 0.008281 | -0.6051 | -0.3806 | -0.5811 | -0.6292
2012-10-30 14:51 | 23.8783 | 19.7075 | 19.916 | 19.6273 | 19.6393 | 45.4273 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.2325 | 3.1318 | 0.008284 | -0.525 | -0.3165 | -0.6052 | -0.5932
2012-10-30 14:51 | 23.8798 | 19.709 | 19.8814 | 19.6408 | 19.6649 | 45.4198 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.23 | 3.1312 | 0.008287 | -0.521 | -0.3486 | -0.5892 | -0.5651
2012-10-30 14:51 | 23.8828 | 19.708 | 19.8964 | 19.6799 | 19.6478 | 45.4198 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014404 | 0.23442 | 0.003377 | 16.2747 | 20.265 | 3.1312 | 0.008281 | -0.557 | -0.3686 | -0.5851 | -0.6172
2012-10-30 14:51 | 23.8791 | 19.7043 | 19.9368 | 19.6281 | 19.6401 | 45.4135 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2493 | 3.1308 | 0.008282 | -0.545 | -0.3125 | -0.6212 | -0.6092
2012-10-30 14:51 | 23.8821 | 19.6712 | 19.9358 | 19.6993 | 19.6471 | 45.406 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.2483 | 3.1303 | 0.008285 | -0.5771 | -0.3125 | -0.549 | -0.6012
2012-10-30 14:51 | 23.8798 | 19.7411 | 19.8814 | 19.6649 | 19.6649 | 45.4239 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 20.1899 | 3.1315 | 0.008279 | -0.4488 | -0.3085 | -0.525 | -0.525
2012-10-30 14:51 | 23.8806 | 19.7178 | 19.9223 | 19.6817 | 19.6697 | 45.4262 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 20.2428 | 3.1317 | 0.008277 | -0.525 | -0.3205 | -0.5611 | -0.5731
2012-10-30 14:52 | 23.8798 | 19.705 | 19.9215 | 19.6649 | 19.6769 | 45.4175 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.214 | 3.1311 | 0.008283 | -0.509 | -0.2925 | -0.5491 | -0.5371
2012-10-30 14:52 | 23.8851 | 19.7223 | 19.8747 | 19.6902 | 19.6622 | 45.4187 | 0.014206 | 0.34536 | 0.004906 | 24.3105 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2593 | 3.1312 | 0.008284 | -0.537 | -0.3846 | -0.5691 | -0.5971
2012-10-30 14:52 | 23.8776 | 19.7028 | 19.8912 | 19.6266 | 19.6747 | 45.4175 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2358 | 3.1311 | 0.008281 | -0.533 | -0.3446 | -0.6092 | -0.5611
2012-10-30 14:52 | 23.8843 | 19.6894 | 19.8819 | 19.6413 | 19.6654 | 45.4193 | 0.014194 | 0.34534 | 0.004902 | 24.33 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.2305 | 3.1312 | 0.008277 | -0.5411 | -0.3486 | -0.5892 | -0.5651
2012-10-30 14:52 | 23.8761 | 19.7053 | 19.8897 | 19.6812 | 19.6451 | 45.4262 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.2063 | 3.1317 | 0.008285 | -0.501 | -0.3166 | -0.5251 | -0.5612
2012-10-30 14:52 | 23.8783 | 19.7035 | 19.912 | 19.6674 | 19.6594 | 45.4262 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.1644 | 3.1317 | 0.008281 | -0.4609 | -0.2524 | -0.497 | -0.505
2012-10-30 14:52 | 23.8746 | 19.6797 | 19.9123 | 19.6396 | 19.6797 | 45.4054 | 0.014209 | 0.34537 | 0.004907 | 24.3064 | 0.014408 | 0.23442 | 0.003378 | 16.2702 | 20.1767 | 3.1302 | 0.008285 | -0.497 | -0.2644 | -0.5371 | -0.497
2012-10-30 14:52 | 23.8753 | 19.6724 | 19.8929 | 19.6644 | 19.6484 | 45.4152 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.1534 | 3.1309 | 0.00828 | -0.481 | -0.2605 | -0.489 | -0.505
2012-10-30 14:52 | 23.8738 | 19.683 | 19.8994 | 19.6589 | 19.6469 | 45.4227 | 0.014197 | 0.34531 | 0.004902 | 24.3228 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.1679 | 3.1314 | 0.008277 | -0.4849 | -0.2685 | -0.509 | -0.521
2012-10-30 14:52 | 23.8828 | 19.7401 | 19.9205 | 19.6719 | 19.6759 | 45.4308 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.1328 | 3.132 | 0.008283 | -0.3927 | -0.2123 | -0.4609 | -0.4569
2012-10-30 14:52 | 23.8746 | 19.7239 | 19.9002 | 19.6477 | 19.6757 | 45.4256 | 0.0142 | 0.34533 | 0.004904 | 24.3187 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.0886 | 3.1316 | 0.008282 | -0.3647 | -0.1884 | -0.4409 | -0.4129
2012-10-30 14:52 | 23.8746 | 19.6958 | 19.8962 | 19.6878 | 19.6316 | 45.4135 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.0726 | 3.1308 | 0.008279 | -0.3768 | -0.1764 | -0.3848 | -0.441
2012-10-30 14:52 | 23.8738 | 19.719 | 19.8754 | 19.6629 | 19.6589 | 45.4123 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.0638 | 3.1307 | 0.008287 | -0.3448 | -0.1884 | -0.4009 | -0.4049
2012-10-30 14:52 | 23.8828 | 19.704 | 19.8724 | 19.6518 | 19.6799 | 45.4204 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.0647 | 3.1313 | 0.008284 | -0.3607 | -0.1923 | -0.4129 | -0.3848
2012-10-30 14:52 | 23.8716 | 19.7048 | 19.9013 | 19.6607 | 19.6166 | 45.4233 | 0.014206 | 0.34536 | 0.004906 | 24.3105 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.0255 | 3.1315 | 0.008284 | -0.3207 | -0.1242 | -0.3648 | -0.4089
2012-10-30 14:52 | 23.8746 | 19.7399 | 19.8842 | 19.6477 | 19.6517 | 45.4164 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0285 | 3.131 | 0.008285 | -0.2886 | -0.1443 | -0.3808 | -0.3768
2012-10-30 14:52 | 23.8671 | 19.7244 | 19.8446 | 19.6682 | 19.6602 | 45.4233 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014395 | 0.23441 | 0.003374 | 16.2839 | 19.997 | 3.1315 | 0.008279 | -0.2726 | -0.1524 | -0.3288 | -0.3368
2012-10-30 14:52 | 23.8731 | 19.7023 | 19.8667 | 19.6502 | 19.6261 | 45.4135 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.047 | 3.1308 | 0.008287 | -0.3447 | -0.1803 | -0.3968 | -0.4209
2012-10-30 14:52 | 23.8746 | 19.7198 | 19.8722 | 19.6517 | 19.6838 | 45.425 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.0485 | 3.1316 | 0.00828 | -0.3287 | -0.1763 | -0.3968 | -0.3647
2012-10-30 14:52 | 23.8753 | 19.7005 | 19.8849 | 19.6363 | 19.6524 | 45.4187 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.0292 | 3.1312 | 0.008283 | -0.3287 | -0.1443 | -0.3929 | -0.3768
2012-10-30 14:52 | 23.8783 | 19.7235 | 19.8558 | 19.6433 | 19.6714 | 45.3932 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.0763 | 3.1294 | 0.008279 | -0.3528 | -0.2205 | -0.433 | -0.4049
2012-10-30 14:52 | 23.8641 | 19.7174 | 19.8416 | 19.6853 | 19.6492 | 45.4135 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.0781 | 3.1308 | 0.008286 | -0.3607 | -0.2365 | -0.3928 | -0.4289
2012-10-30 14:52 | 23.8731 | 19.7705 | 19.8587 | 19.6421 | 19.6702 | 45.4221 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.0671 | 3.1314 | 0.008282 | -0.2966 | -0.2084 | -0.425 | -0.3969
2012-10-30 14:52 | 23.8693 | 19.7426 | 19.8468 | 19.6664 | 19.6263 | 45.4169 | 0.014219 | 0.34537 | 0.004911 | 24.2893 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.0593 | 3.131 | 0.008288 | -0.3167 | -0.2125 | -0.3929 | -0.433
2012-10-30 14:53 | 23.8634 | 19.7086 | 19.8249 | 19.6685 | 19.6004 | 45.4013 | 0.014209 | 0.34534 | 0.004907 | 24.3043 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0694 | 3.13 | 0.008283 | -0.3608 | -0.2445 | -0.4009 | -0.469
2012-10-30 14:53 | 23.8574 | 19.7147 | 19.8309 | 19.6585 | 19.6505 | 45.4123 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.0834 | 3.1307 | 0.008283 | -0.3687 | -0.2525 | -0.4249 | -0.4329
2012-10-30 14:53 | 23.8671 | 19.7043 | 19.8046 | 19.6442 | 19.6642 | 45.4175 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.0931 | 3.1311 | 0.008282 | -0.3888 | -0.2885 | -0.4489 | -0.4289
2012-10-30 14:53 | 23.8693 | 19.7186 | 19.8268 | 19.6865 | 19.6544 | 45.4164 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.0232 | 3.131 | 0.008281 | -0.3046 | -0.1964 | -0.3367 | -0.3688
2012-10-30 14:53 | 23.8634 | 19.7327 | 19.8049 | 19.6525 | 19.6244 | 45.4065 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.0694 | 3.1303 | 0.008282 | -0.3367 | -0.2645 | -0.4169 | -0.445
2012-10-30 14:53 | 23.8641 | 19.7414 | 19.8416 | 19.6372 | 19.6171 | 45.4106 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.0901 | 3.1306 | 0.008285 | -0.3487 | -0.2485 | -0.4529 | -0.473
2012-10-30 14:53 | 23.8679 | 19.7332 | 19.8254 | 19.653 | 19.641 | 45.4204 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.0739 | 3.1313 | 0.008282 | -0.3407 | -0.2485 | -0.4209 | -0.4329
2012-10-30 14:53 | 23.8596 | 19.7129 | 19.785 | 19.6407 | 19.6367 | 45.4302 | 0.014187 | 0.3454 | 0.0049 | 24.3461 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.0576 | 3.132 | 0.008277 | -0.3447 | -0.2726 | -0.4169 | -0.4209
2012-10-30 14:53 | 23.8679 | 19.7172 | 19.8054 | 19.665 | 19.6209 | 45.4117 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.0699 | 3.1307 | 0.008288 | -0.3527 | -0.2645 | -0.4049 | -0.449
2012-10-30 14:53 | 23.8656 | 19.7108 | 19.8231 | 19.6547 | 19.6146 | 45.4141 | 0.0142 | 0.34533 | 0.004904 | 24.3187 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.1116 | 3.1308 | 0.00828 | -0.4008 | -0.2885 | -0.4569 | -0.497
2012-10-30 14:53 | 23.8649 | 19.7262 | 19.8104 | 19.6219 | 19.5978 | 45.4227 | 0.01419 | 0.34541 | 0.004901 | 24.3419 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.0869 | 3.1314 | 0.008279 | -0.3607 | -0.2765 | -0.465 | -0.4891
2012-10-30 14:53 | 23.8768 | 19.7381 | 19.8584 | 19.6458 | 19.6579 | 45.425 | 0.014187 | 0.34541 | 0.0049 | 24.3471 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.0868 | 3.1316 | 0.00828 | -0.3487 | -0.2284 | -0.441 | -0.4289
2012-10-30 14:53 | 23.8806 | 19.7339 | 19.79 | 19.6537 | 19.6336 | 45.41 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.0826 | 3.1306 | 0.008286 | -0.3487 | -0.2926 | -0.4289 | -0.449
2012-10-30 14:53 | 23.8761 | 19.7454 | 19.8376 | 19.6532 | 19.6171 | 45.4123 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 20.05 | 3.1307 | 0.008278 | -0.3046 | -0.2124 | -0.3968 | -0.4329
2012-10-30 14:53 | 23.8776 | 19.6827 | 19.8191 | 19.6426 | 19.6146 | 45.4262 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014423 | 0.23448 | 0.003382 | 16.2573 | 20.0836 | 3.1317 | 0.008287 | -0.4009 | -0.2645 | -0.441 | -0.469
2012-10-30 14:53 | 23.8753 | 19.7125 | 19.8088 | 19.6403 | 19.6243 | 45.4175 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1053 | 3.1311 | 0.00829 | -0.3928 | -0.2965 | -0.465 | -0.481
2012-10-30 14:53 | 23.8798 | 19.697 | 19.7812 | 19.6529 | 19.6328 | 45.4216 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.0778 | 3.1314 | 0.008282 | -0.3808 | -0.2966 | -0.4249 | -0.445
2012-10-30 14:53 | 23.8738 | 19.7672 | 19.8193 | 19.6549 | 19.6188 | 45.4008 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.0958 | 3.1299 | 0.008283 | -0.3286 | -0.2765 | -0.4409 | -0.477
2012-10-30 14:53 | 23.8806 | 19.7098 | 19.778 | 19.6296 | 19.6296 | 45.4302 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.0986 | 3.132 | 0.008287 | -0.3888 | -0.3206 | -0.469 | -0.469
2012-10-30 14:53 | 23.8858 | 19.7431 | 19.8273 | 19.6669 | 19.6468 | 45.4181 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1439 | 3.1311 | 0.008285 | -0.4008 | -0.3166 | -0.477 | -0.4971
2012-10-30 14:53 | 23.8768 | 19.698 | 19.8103 | 19.6378 | 19.6338 | 45.4019 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1549 | 3.13 | 0.008288 | -0.4569 | -0.3446 | -0.5171 | -0.5211
2012-10-30 14:53 | 23.8798 | 19.713 | 19.8413 | 19.6809 | 19.6328 | 45.4019 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.1459 | 3.13 | 0.008287 | -0.4329 | -0.3046 | -0.465 | -0.5131
2012-10-30 14:53 | 23.8738 | 19.6549 | 19.8353 | 19.6469 | 19.6148 | 45.4158 | 0.014238 | 0.34543 | 0.004918 | 24.2609 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.1359 | 3.131 | 0.008295 | -0.481 | -0.3006 | -0.489 | -0.5211
2012-10-30 14:53 | 23.8783 | 19.6995 | 19.8278 | 19.6233 | 19.6393 | 45.4025 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.1444 | 3.13 | 0.008283 | -0.4449 | -0.3166 | -0.5211 | -0.5051
2012-10-30 14:54 | 23.8798 | 19.709 | 19.8654 | 19.6328 | 19.6168 | 45.4002 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1939 | 3.1299 | 0.008288 | -0.4849 | -0.3285 | -0.5611 | -0.5771
2012-10-30 14:54 | 23.8753 | 19.7286 | 19.8248 | 19.6363 | 19.6484 | 45.4025 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.1414 | 3.13 | 0.008285 | -0.4128 | -0.3166 | -0.5051 | -0.493
2012-10-30 14:54 | 23.8761 | 19.7013 | 19.8136 | 19.6211 | 19.6532 | 45.4048 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1542 | 3.1302 | 0.008281 | -0.4529 | -0.3406 | -0.5331 | -0.501
2012-10-30 14:54 | 23.8738 | 19.703 | 19.8513 | 19.6188 | 19.6509 | 45.3961 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014423 | 0.23447 | 0.003382 | 16.2563 | 20.1118 | 3.1296 | 0.008286 | -0.4088 | -0.2605 | -0.493 | -0.4609
2012-10-30 14:54 | 23.8806 | 19.7258 | 19.8501 | 19.6657 | 19.6657 | 45.3932 | 0.014216 | 0.34534 | 0.004909 | 24.2923 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.1747 | 3.1294 | 0.008285 | -0.4489 | -0.3246 | -0.509 | -0.509
2012-10-30 14:54 | 23.8761 | 19.6933 | 19.8256 | 19.6652 | 19.6331 | 45.4048 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.1582 | 3.1302 | 0.00828 | -0.4649 | -0.3326 | -0.493 | -0.5251
2012-10-30 14:54 | 23.8813 | 19.7065 | 19.8428 | 19.6624 | 19.6584 | 45.388 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1994 | 3.129 | 0.008281 | -0.4929 | -0.3566 | -0.537 | -0.541
2012-10-30 14:54 | 23.8783 | 19.7396 | 19.8478 | 19.6594 | 19.6193 | 45.3967 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.1764 | 3.1296 | 0.008289 | -0.4368 | -0.3286 | -0.517 | -0.5571
2012-10-30 14:54 | 23.8813 | 19.7025 | 19.8508 | 19.6744 | 19.6383 | 45.4048 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1914 | 3.1302 | 0.008284 | -0.4889 | -0.3406 | -0.517 | -0.5531
2012-10-30 14:54 | 23.8686 | 19.6737 | 19.8461 | 19.6256 | 19.6256 | 45.4112 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.1667 | 3.1306 | 0.008289 | -0.493 | -0.3206 | -0.5411 | -0.5411
2012-10-30 14:54 | 23.8738 | 19.695 | 19.8674 | 19.6549 | 19.6469 | 45.3852 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.1238 | 3.1289 | 0.008285 | -0.4288 | -0.2564 | -0.4689 | -0.4769
2012-10-30 14:54 | 23.8753 | 19.6644 | 19.8569 | 19.6443 | 19.6403 | 45.3979 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.1934 | 3.1297 | 0.008284 | -0.529 | -0.3365 | -0.5491 | -0.5531
2012-10-30 14:54 | 23.8723 | 19.6574 | 19.8579 | 19.6133 | 19.6093 | 45.4129 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014417 | 0.23455 | 0.003382 | 16.2691 | 20.1424 | 3.1308 | 0.00829 | -0.485 | -0.2845 | -0.5291 | -0.5331
2012-10-30 14:54 | 23.8723 | 19.7216 | 19.8539 | 19.6413 | 19.6373 | 45.4089 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.2265 | 3.1305 | 0.008281 | -0.5049 | -0.3726 | -0.5852 | -0.5892
2012-10-30 14:54 | 23.8746 | 19.7118 | 19.8682 | 19.6316 | 19.6236 | 45.4083 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.2408 | 3.1304 | 0.00828 | -0.529 | -0.3726 | -0.6092 | -0.6172
2012-10-30 14:54 | 23.8753 | 19.7085 | 19.8649 | 19.6123 | 19.6524 | 45.4135 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.2175 | 3.1308 | 0.008286 | -0.509 | -0.3526 | -0.6052 | -0.5651
2012-10-30 14:54 | 23.8791 | 19.7163 | 19.8246 | 19.608 | 19.6722 | 45.4175 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2173 | 3.1311 | 0.008288 | -0.501 | -0.3927 | -0.6093 | -0.5451
2012-10-30 14:54 | 23.8791 | 19.7123 | 19.8647 | 19.6963 | 19.6161 | 45.3875 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.2173 | 3.129 | 0.008284 | -0.505 | -0.3526 | -0.521 | -0.6012
2012-10-30 14:54 | 23.8791 | 19.6883 | 19.8406 | 19.6241 | 19.6201 | 45.3967 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.1812 | 3.1296 | 0.008284 | -0.4929 | -0.3406 | -0.5571 | -0.5611
2012-10-30 14:54 | 23.8858 | 19.707 | 19.8233 | 19.6428 | 19.6428 | 45.41 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.1879 | 3.1306 | 0.008282 | -0.4809 | -0.3646 | -0.5451 | -0.5451
2012-10-30 14:54 | 23.8851 | 19.7103 | 19.8386 | 19.6381 | 19.6301 | 45.4089 | 0.014212 | 0.34538 | 0.004909 | 24.3022 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.1992 | 3.1305 | 0.008289 | -0.4889 | -0.3606 | -0.5611 | -0.5691
2012-10-30 14:54 | 23.8753 | 19.6644 | 19.8488 | 19.6203 | 19.6123 | 45.3973 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.1454 | 3.1297 | 0.008281 | -0.481 | -0.2966 | -0.5251 | -0.5331
2012-10-30 14:54 | 23.8836 | 19.6847 | 19.8852 | 19.6045 | 19.6607 | 45.3938 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.1537 | 3.1294 | 0.008284 | -0.469 | -0.2685 | -0.5492 | -0.493
2012-10-30 14:54 | 23.8821 | 19.6471 | 19.8877 | 19.6552 | 19.615 | 45.3967 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.1682 | 3.1296 | 0.008284 | -0.5211 | -0.2805 | -0.513 | -0.5532
2012-10-30 14:55 | 23.8948 | 19.712 | 19.8443 | 19.6358 | 19.6558 | 45.4169 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1609 | 3.131 | 0.008283 | -0.4489 | -0.3166 | -0.5251 | -0.5051
2012-10-30 14:55 | 23.8806 | 19.6737 | 19.8541 | 19.6135 | 19.6216 | 45.3984 | 0.014222 | 0.3454 | 0.004912 | 24.2862 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2228 | 3.1298 | 0.008289 | -0.5491 | -0.3687 | -0.6093 | -0.6012
2012-10-30 14:55 | 23.8865 | 19.6676 | 19.848 | 19.6395 | 19.6074 | 45.3857 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1806 | 3.1289 | 0.008288 | -0.513 | -0.3326 | -0.5411 | -0.5732
2012-10-30 14:55 | 23.8873 | 19.6684 | 19.8488 | 19.6122 | 19.6082 | 45.3961 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.1814 | 3.1296 | 0.008285 | -0.513 | -0.3326 | -0.5692 | -0.5732
2012-10-30 14:55 | 23.8903 | 19.7115 | 19.8719 | 19.6433 | 19.6393 | 45.3875 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1724 | 3.129 | 0.008282 | -0.4609 | -0.3005 | -0.5291 | -0.5331
2012-10-30 14:55 | 23.8918 | 19.7491 | 19.8373 | 19.6288 | 19.6488 | 45.3956 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.1979 | 3.1296 | 0.008288 | -0.4488 | -0.3606 | -0.5691 | -0.5491
2012-10-30 14:55 | 23.8798 | 19.701 | 19.8373 | 19.6408 | 19.6368 | 45.4117 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.1939 | 3.1307 | 0.008284 | -0.4929 | -0.3566 | -0.5531 | -0.5571
2012-10-30 14:55 | 23.8933 | 19.6984 | 19.8548 | 19.6182 | 19.6343 | 45.3996 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.2275 | 3.1298 | 0.008287 | -0.5291 | -0.3727 | -0.6093 | -0.5932
2012-10-30 14:55 | 23.8865 | 19.7157 | 19.848 | 19.6315 | 19.6114 | 45.3909 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2207 | 3.1292 | 0.008288 | -0.505 | -0.3727 | -0.5892 | -0.6093
2012-10-30 14:55 | 23.8873 | 19.6804 | 19.8328 | 19.6202 | 19.6563 | 45.4054 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.2054 | 3.1302 | 0.008284 | -0.525 | -0.3726 | -0.5852 | -0.5491
2012-10-30 14:55 | 23.8933 | 19.6744 | 19.8829 | 19.6222 | 19.6383 | 45.4042 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.1834 | 3.1302 | 0.008287 | -0.509 | -0.3005 | -0.5612 | -0.5451
2012-10-30 14:55 | 23.8858 | 19.6709 | 19.8834 | 19.6107 | 19.6348 | 45.3863 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.24 | 3.1289 | 0.008289 | -0.5691 | -0.3566 | -0.6293 | -0.6052
2012-10-30 14:55 | 23.8843 | 19.7255 | 19.8859 | 19.6293 | 19.6574 | 45.4071 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2305 | 3.1304 | 0.008285 | -0.505 | -0.3446 | -0.6012 | -0.5731
2012-10-30 14:55 | 23.8873 | 19.6924 | 19.8568 | 19.6122 | 19.6403 | 45.4019 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.01442 | 0.23455 | 0.003382 | 16.2657 | 20.1774 | 3.13 | 0.008288 | -0.485 | -0.3206 | -0.5652 | -0.5371
2012-10-30 14:55 | 23.8925 | 19.6936 | 19.8821 | 19.6014 | 19.6575 | 45.4008 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.2066 | 3.1299 | 0.008279 | -0.513 | -0.3245 | -0.6052 | -0.5491
2012-10-30 14:55 | 23.8888 | 19.6739 | 19.8944 | 19.6458 | 19.6378 | 45.3759 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1829 | 3.1282 | 0.008284 | -0.509 | -0.2885 | -0.5371 | -0.5451
2012-10-30 14:55 | 23.897 | 19.6901 | 19.8866 | 19.6099 | 19.6179 | 45.395 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.2392 | 3.1295 | 0.008278 | -0.5491 | -0.3526 | -0.6293 | -0.6213
2012-10-30 14:55 | 23.894 | 19.647 | 19.8435 | 19.6109 | 19.6109 | 45.4083 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.2081 | 3.1304 | 0.008285 | -0.5611 | -0.3646 | -0.5972 | -0.5972
2012-10-30 14:55 | 23.8873 | 19.6724 | 19.8769 | 19.6323 | 19.6082 | 45.388 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014389 | 0.23444 | 0.003373 | 16.2927 | 20.1974 | 3.129 | 0.008281 | -0.525 | -0.3205 | -0.5651 | -0.5892
2012-10-30 14:55 | 23.8903 | 19.6714 | 19.8438 | 19.6393 | 19.6393 | 45.3927 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.1684 | 3.1294 | 0.008285 | -0.497 | -0.3246 | -0.5291 | -0.5291
2012-10-30 14:55 | 23.8978 | 19.6668 | 19.8673 | 19.6107 | 19.6348 | 45.3869 | 0.01419 | 0.34547 | 0.004902 | 24.346 | 0.014417 | 0.23438 | 0.003379 | 16.2571 | 20.1799 | 3.129 | 0.008281 | -0.5131 | -0.3126 | -0.5692 | -0.5451
2012-10-30 14:55 | 23.9037 | 19.6567 | 19.8452 | 19.6046 | 19.6246 | 45.3765 | 0.014184 | 0.34543 | 0.0049 | 24.3533 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2339 | 3.1283 | 0.008278 | -0.5772 | -0.3887 | -0.6293 | -0.6093
2012-10-30 14:55 | 23.906 | 19.675 | 19.8435 | 19.6149 | 19.663 | 45.3823 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2081 | 3.1287 | 0.008288 | -0.5331 | -0.3646 | -0.5932 | -0.5451
2012-10-30 14:55 | 23.9 | 19.6731 | 19.8575 | 19.5968 | 19.6129 | 45.3909 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.2021 | 3.1292 | 0.008281 | -0.529 | -0.3446 | -0.6053 | -0.5892
2012-10-30 14:56 | 23.9008 | 19.6418 | 19.8142 | 19.6538 | 19.6538 | 45.3707 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2029 | 3.1279 | 0.008288 | -0.5611 | -0.3887 | -0.5491 | -0.5491
2012-10-30 14:56 | 23.9097 | 19.6948 | 19.8432 | 19.6065 | 19.6467 | 45.3765 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1878 | 3.1283 | 0.008286 | -0.493 | -0.3446 | -0.5813 | -0.5411
2012-10-30 14:56 | 23.9097 | 19.6948 | 19.8271 | 19.6106 | 19.6467 | 45.3944 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.2078 | 3.1295 | 0.008286 | -0.513 | -0.3807 | -0.5972 | -0.5611
2012-10-30 14:56 | 23.9052 | 19.6903 | 19.8547 | 19.6542 | 19.6221 | 45.3794 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.2033 | 3.1285 | 0.008283 | -0.513 | -0.3486 | -0.5491 | -0.5812
2012-10-30 14:56 | 23.9105 | 19.6475 | 19.868 | 19.6515 | 19.6234 | 45.3938 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014389 | 0.23454 | 0.003375 | 16.2997 | 20.2166 | 3.1294 | 0.008281 | -0.5691 | -0.3486 | -0.5651 | -0.5932
2012-10-30 14:56 | 23.9112 | 19.6642 | 19.8366 | 19.6161 | 19.6562 | 45.388 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.2454 | 3.129 | 0.008285 | -0.5812 | -0.4088 | -0.6293 | -0.5892
2012-10-30 14:56 | 23.912 | 19.681 | 19.8615 | 19.5888 | 19.6409 | 45.3724 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2862 | 3.128 | 0.008288 | -0.6052 | -0.4247 | -0.6974 | -0.6453
2012-10-30 14:56 | 23.9105 | 19.6675 | 19.8359 | 19.6434 | 19.6033 | 45.4019 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.2206 | 3.13 | 0.008285 | -0.5531 | -0.3847 | -0.5772 | -0.6173
2012-10-30 14:56 | 23.9165 | 19.6575 | 19.858 | 19.6374 | 19.6374 | 45.395 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.2387 | 3.1295 | 0.008281 | -0.5812 | -0.3807 | -0.6013 | -0.6013
2012-10-30 14:56 | 23.9209 | 19.6899 | 19.8704 | 19.5977 | 19.6258 | 45.3823 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.219 | 3.1287 | 0.008289 | -0.5291 | -0.3486 | -0.6213 | -0.5932
2012-10-30 14:56 | 23.9135 | 19.6825 | 19.851 | 19.6464 | 19.6505 | 45.3909 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.2076 | 3.1292 | 0.008287 | -0.5251 | -0.3566 | -0.5612 | -0.5571
2012-10-30 14:56 | 23.9217 | 19.6587 | 19.8632 | 19.6667 | 19.6386 | 45.3915 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.2158 | 3.1293 | 0.008285 | -0.5571 | -0.3526 | -0.5491 | -0.5772
2012-10-30 14:56 | 23.9142 | 19.6552 | 19.8797 | 19.6471 | 19.6431 | 45.3996 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014417 | 0.23444 | 0.00338 | 16.2611 | 20.2404 | 3.1298 | 0.008283 | -0.5852 | -0.3607 | -0.5933 | -0.5973
2012-10-30 14:56 | 23.9127 | 19.6777 | 19.8742 | 19.6376 | 19.6537 | 45.388 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.2148 | 3.129 | 0.00829 | -0.5371 | -0.3406 | -0.5772 | -0.5611
2012-10-30 14:56 | 23.9195 | 19.6645 | 19.861 | 19.6163 | 19.6645 | 45.3973 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.2296 | 3.1297 | 0.008286 | -0.5651 | -0.3686 | -0.6133 | -0.5651
2012-10-30 14:56 | 23.9262 | 19.6872 | 19.9118 | 19.6471 | 19.6391 | 45.3967 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.2323 | 3.1296 | 0.008286 | -0.5451 | -0.3205 | -0.5852 | -0.5932
2012-10-30 14:56 | 23.9209 | 19.6418 | 19.8303 | 19.6418 | 19.6057 | 45.3701 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.195 | 3.1278 | 0.008287 | -0.5532 | -0.3647 | -0.5532 | -0.5893
2012-10-30 14:56 | 23.9172 | 19.6501 | 19.8306 | 19.6461 | 19.6381 | 45.3869 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.1713 | 3.129 | 0.008288 | -0.5212 | -0.3407 | -0.5252 | -0.5332
2012-10-30 14:56 | 23.9262 | 19.6391 | 19.8797 | 19.6551 | 19.619 | 45.3904 | 0.014203 | 0.34536 | 0.004905 | 24.3156 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2283 | 3.1292 | 0.008285 | -0.5892 | -0.3486 | -0.5732 | -0.6093
2012-10-30 14:56 | 23.9254 | 19.6744 | 19.8749 | 19.6102 | 19.6664 | 45.3805 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2075 | 3.1285 | 0.008283 | -0.5331 | -0.3326 | -0.5973 | -0.5411
2012-10-30 14:56 | 23.9165 | 19.6494 | 19.874 | 19.6334 | 19.5853 | 45.3823 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.2266 | 3.1287 | 0.008289 | -0.5772 | -0.3526 | -0.5932 | -0.6413
2012-10-30 14:56 | 23.9195 | 19.6565 | 19.8369 | 19.6444 | 19.6444 | 45.388 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2016 | 3.129 | 0.008288 | -0.5451 | -0.3647 | -0.5572 | -0.5572
2012-10-30 14:56 | 23.9254 | 19.6824 | 19.8308 | 19.5942 | 19.6263 | 45.4019 | 0.014219 | 0.34547 | 0.004912 | 24.2964 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.2556 | 3.13 | 0.008292 | -0.5732 | -0.4248 | -0.6614 | -0.6293
2012-10-30 14:56 | 23.9209 | 19.6338 | 19.8784 | 19.6498 | 19.6538 | 45.3753 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2991 | 3.1282 | 0.008284 | -0.6653 | -0.4207 | -0.6493 | -0.6453
2012-10-30 14:57 | 23.9247 | 19.6376 | 19.8622 | 19.6175 | 19.6336 | 45.3805 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2388 | 3.1285 | 0.008286 | -0.6012 | -0.3766 | -0.6213 | -0.6052
2012-10-30 14:57 | 23.9299 | 19.6909 | 19.8914 | 19.6548 | 19.6428 | 45.3927 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.2521 | 3.1294 | 0.008283 | -0.5612 | -0.3607 | -0.5973 | -0.6093
2012-10-30 14:57 | 23.9232 | 19.6642 | 19.8727 | 19.5799 | 19.6441 | 45.3817 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.2133 | 3.1286 | 0.008289 | -0.5491 | -0.3406 | -0.6334 | -0.5692
2012-10-30 14:57 | 23.9269 | 19.6518 | 19.8844 | 19.6398 | 19.6237 | 45.3707 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.209 | 3.1279 | 0.008284 | -0.5572 | -0.3246 | -0.5692 | -0.5853
2012-10-30 14:57 | 23.9157 | 19.6366 | 19.8492 | 19.6166 | 19.6005 | 45.3788 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2258 | 3.1284 | 0.008287 | -0.5892 | -0.3766 | -0.6092 | -0.6253
2012-10-30 14:57 | 23.9359 | 19.6729 | 19.8694 | 19.6648 | 19.6568 | 45.3834 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.222 | 3.1287 | 0.008284 | -0.5491 | -0.3526 | -0.5572 | -0.5652
2012-10-30 14:57 | 23.9247 | 19.6978 | 19.8501 | 19.6256 | 19.6256 | 45.3823 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.2268 | 3.1287 | 0.008285 | -0.529 | -0.3767 | -0.6012 | -0.6012
2012-10-30 14:57 | 23.918 | 19.6189 | 19.8434 | 19.5948 | 19.6189 | 45.3776 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.1961 | 3.1283 | 0.008288 | -0.5772 | -0.3527 | -0.6013 | -0.5772
2012-10-30 14:57 | 23.9224 | 19.6553 | 19.8438 | 19.6433 | 19.6393 | 45.3846 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1965 | 3.1288 | 0.008286 | -0.5412 | -0.3527 | -0.5532 | -0.5572
2012-10-30 14:57 | 23.9322 | 19.6972 | 19.8256 | 19.6331 | 19.6371 | 45.3672 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23445 | 0.003377 | 16.2767 | 20.2343 | 3.1276 | 0.008285 | -0.5371 | -0.4087 | -0.6012 | -0.5972
2012-10-30 14:57 | 23.9247 | 19.6456 | 19.8982 | 19.5975 | 19.6536 | 45.3759 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.2388 | 3.1282 | 0.008286 | -0.5932 | -0.3406 | -0.6413 | -0.5852
2012-10-30 14:57 | 23.9202 | 19.6772 | 19.8416 | 19.6692 | 19.6251 | 45.3742 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2103 | 3.1281 | 0.008286 | -0.5331 | -0.3687 | -0.5411 | -0.5852
2012-10-30 14:57 | 23.9224 | 19.6353 | 19.8639 | 19.5912 | 19.6473 | 45.3742 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.2085 | 3.1281 | 0.008283 | -0.5732 | -0.3446 | -0.6173 | -0.5612
2012-10-30 14:57 | 23.9232 | 19.616 | 19.8526 | 19.604 | 19.6561 | 45.369 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.1973 | 3.1277 | 0.008292 | -0.5813 | -0.3447 | -0.5933 | -0.5412
2012-10-30 14:57 | 23.9269 | 19.6639 | 19.9165 | 19.5756 | 19.6438 | 45.3788 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2491 | 3.1284 | 0.008283 | -0.5852 | -0.3326 | -0.6735 | -0.6053
2012-10-30 14:57 | 23.9269 | 19.6278 | 19.8604 | 19.6318 | 19.6759 | 45.384 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.2531 | 3.1288 | 0.008283 | -0.6253 | -0.3927 | -0.6213 | -0.5772
2012-10-30 14:57 | 23.9284 | 19.7055 | 19.8619 | 19.6293 | 19.6453 | 45.3661 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.2626 | 3.1275 | 0.008282 | -0.5571 | -0.4007 | -0.6333 | -0.6173
2012-10-30 14:57 | 23.9262 | 19.6912 | 19.8276 | 19.619 | 19.6391 | 45.3719 | 0.014206 | 0.34536 | 0.004906 | 24.3105 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.2484 | 3.1279 | 0.008287 | -0.5572 | -0.4208 | -0.6294 | -0.6093
2012-10-30 14:57 | 23.9314 | 19.6563 | 19.8849 | 19.6443 | 19.6323 | 45.3724 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2576 | 3.128 | 0.008287 | -0.6013 | -0.3727 | -0.6133 | -0.6253
2012-10-30 14:57 | 23.9314 | 19.6523 | 19.913 | 19.6764 | 19.6523 | 45.3701 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.2335 | 3.1278 | 0.008284 | -0.5812 | -0.3205 | -0.5571 | -0.5812
2012-10-30 14:57 | 23.9352 | 19.632 | 19.9007 | 19.6361 | 19.628 | 45.3667 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014414 | 0.23445 | 0.003379 | 16.2654 | 20.2533 | 3.1276 | 0.008286 | -0.6213 | -0.3526 | -0.6172 | -0.6253
2012-10-30 14:57 | 23.9322 | 19.6611 | 19.8777 | 19.6571 | 19.6531 | 45.3707 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014423 | 0.23457 | 0.003383 | 16.2633 | 20.2063 | 3.1279 | 0.008291 | -0.5452 | -0.3286 | -0.5492 | -0.5532
2012-10-30 14:57 | 23.9389 | 19.6598 | 19.8683 | 19.6357 | 19.6277 | 45.369 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.209 | 3.1277 | 0.008287 | -0.5492 | -0.3407 | -0.5733 | -0.5813
2012-10-30 14:57 | 23.9322 | 19.6371 | 19.8616 | 19.605 | 19.597 | 45.3782 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.2223 | 3.1284 | 0.00829 | -0.5852 | -0.3607 | -0.6173 | -0.6253
2012-10-30 14:58 | 23.9359 | 19.6287 | 19.8814 | 19.6207 | 19.6408 | 45.38 | 0.014184 | 0.34544 | 0.0049 | 24.3543 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.254 | 3.1285 | 0.008279 | -0.6253 | -0.3726 | -0.6333 | -0.6132
2012-10-30 14:58 | 23.9314 | 19.6483 | 19.8649 | 19.6242 | 19.6483 | 45.3597 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.2736 | 3.1271 | 0.008289 | -0.6253 | -0.4087 | -0.6494 | -0.6253
2012-10-30 14:58 | 23.9314 | 19.6443 | 19.8087 | 19.6002 | 19.6443 | 45.3863 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.2135 | 3.1289 | 0.008285 | -0.5692 | -0.4048 | -0.6133 | -0.5692
2012-10-30 14:58 | 23.9441 | 19.6329 | 19.8695 | 19.657 | 19.645 | 45.3771 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2382 | 3.1283 | 0.008286 | -0.6053 | -0.3687 | -0.5812 | -0.5932
2012-10-30 14:58 | 23.9441 | 19.657 | 19.8976 | 19.6209 | 19.6289 | 45.3609 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2542 | 3.1272 | 0.008285 | -0.5972 | -0.3566 | -0.6333 | -0.6253
2012-10-30 14:58 | 23.9441 | 19.657 | 19.8455 | 19.6369 | 19.6249 | 45.3644 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.2663 | 3.1274 | 0.008288 | -0.6093 | -0.4208 | -0.6294 | -0.6414
2012-10-30 14:58 | 23.9501 | 19.667 | 19.8595 | 19.6269 | 19.655 | 45.3667 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014414 | 0.23444 | 0.003379 | 16.2645 | 20.2843 | 3.1276 | 0.008284 | -0.6173 | -0.4248 | -0.6574 | -0.6293
2012-10-30 14:58 | 23.9426 | 19.6555 | 19.868 | 19.6314 | 19.6675 | 45.3811 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.2607 | 3.1286 | 0.008286 | -0.6052 | -0.3927 | -0.6293 | -0.5932
2012-10-30 14:58 | 23.9449 | 19.6819 | 19.8463 | 19.6658 | 19.6257 | 45.3759 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.243 | 3.1282 | 0.008288 | -0.5611 | -0.3967 | -0.5772 | -0.6173
2012-10-30 14:58 | 23.9411 | 19.662 | 19.8425 | 19.6219 | 19.646 | 45.373 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2713 | 3.128 | 0.008289 | -0.6093 | -0.4288 | -0.6494 | -0.6253
2012-10-30 14:58 | 23.9434 | 19.7044 | 19.8809 | 19.6041 | 19.6282 | 45.3591 | 0.014197 | 0.34531 | 0.004902 | 24.3228 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2495 | 3.1271 | 0.008281 | -0.5451 | -0.3686 | -0.6454 | -0.6213
2012-10-30 14:58 | 23.9456 | 19.6545 | 19.871 | 19.6104 | 19.6384 | 45.38 | 0.014219 | 0.34538 | 0.004911 | 24.2902 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2477 | 3.1285 | 0.008289 | -0.5932 | -0.3767 | -0.6373 | -0.6093
2012-10-30 14:58 | 23.9509 | 19.6758 | 19.8322 | 19.6116 | 19.6598 | 45.3649 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.237 | 3.1275 | 0.008286 | -0.5612 | -0.4048 | -0.6254 | -0.5772
2012-10-30 14:58 | 23.9479 | 19.6247 | 19.8453 | 19.6367 | 19.6287 | 45.3661 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.222 | 3.1275 | 0.008283 | -0.5973 | -0.3767 | -0.5853 | -0.5933
2012-10-30 14:58 | 23.9486 | 19.6214 | 19.846 | 19.6174 | 19.6294 | 45.3678 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.2187 | 3.1277 | 0.008285 | -0.5973 | -0.3727 | -0.6013 | -0.5893
2012-10-30 14:58 | 23.9479 | 19.6367 | 19.8533 | 19.6086 | 19.6167 | 45.358 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.202 | 3.127 | 0.008286 | -0.5653 | -0.3487 | -0.5934 | -0.5853
2012-10-30 14:58 | 23.9539 | 19.6628 | 19.8593 | 19.6227 | 19.6267 | 45.373 | 0.014212 | 0.34537 | 0.004908 | 24.3012 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.2039 | 3.128 | 0.008288 | -0.5411 | -0.3446 | -0.5812 | -0.5772
2012-10-30 14:58 | 23.9568 | 19.6416 | 19.8261 | 19.6216 | 19.6296 | 45.3684 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.1788 | 3.1277 | 0.008288 | -0.5372 | -0.3527 | -0.5572 | -0.5492
2012-10-30 14:58 | 23.9606 | 19.6695 | 19.8339 | 19.6254 | 19.6374 | 45.3672 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.1946 | 3.1276 | 0.008291 | -0.5251 | -0.3607 | -0.5692 | -0.5572
2012-10-30 14:58 | 23.9568 | 19.6496 | 19.8381 | 19.6336 | 19.6456 | 45.373 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.1708 | 3.128 | 0.008288 | -0.5212 | -0.3327 | -0.5372 | -0.5252
2012-10-30 14:58 | 23.9636 | 19.6243 | 19.861 | 19.5842 | 19.6404 | 45.3748 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.1736 | 3.1281 | 0.00829 | -0.5493 | -0.3126 | -0.5894 | -0.5332
2012-10-30 14:58 | 23.9583 | 19.6271 | 19.8316 | 19.611 | 19.6431 | 45.3649 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.1723 | 3.1275 | 0.008287 | -0.5452 | -0.3407 | -0.5613 | -0.5292
2012-10-30 14:59 | 23.9696 | 19.6785 | 19.8549 | 19.6624 | 19.6384 | 45.3678 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.2196 | 3.1277 | 0.008286 | -0.5411 | -0.3647 | -0.5572 | -0.5812
2012-10-30 14:59 | 23.9666 | 19.6554 | 19.8038 | 19.6113 | 19.6354 | 45.3655 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.2046 | 3.1275 | 0.00829 | -0.5492 | -0.4008 | -0.5933 | -0.5692
2012-10-30 14:59 | 23.9666 | 19.6273 | 19.8439 | 19.6273 | 19.6273 | 45.3522 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.1966 | 3.1266 | 0.008288 | -0.5693 | -0.3527 | -0.5693 | -0.5693
2012-10-30 14:59 | 23.9673 | 19.6481 | 19.8807 | 19.6441 | 19.6321 | 45.3603 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.2334 | 3.1271 | 0.00829 | -0.5853 | -0.3527 | -0.5893 | -0.6013
2012-10-30 14:59 | 23.9673 | 19.616 | 19.8446 | 19.604 | 19.6321 | 45.3603 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.1693 | 3.1271 | 0.008283 | -0.5533 | -0.3247 | -0.5653 | -0.5372
2012-10-30 14:59 | 23.9733 | 19.6822 | 19.8626 | 19.622 | 19.6461 | 45.3817 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.2073 | 3.1286 | 0.008289 | -0.5251 | -0.3447 | -0.5853 | -0.5612
2012-10-30 14:59 | 23.9621 | 19.683 | 19.8434 | 19.6068 | 19.6469 | 45.3771 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2242 | 3.1283 | 0.008288 | -0.5412 | -0.3808 | -0.6174 | -0.5773
2012-10-30 14:59 | 23.98 | 19.6127 | 19.8774 | 19.6127 | 19.6448 | 45.3626 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.21 | 3.1273 | 0.008289 | -0.5973 | -0.3326 | -0.5973 | -0.5652
2012-10-30 14:59 | 23.9755 | 19.6362 | 19.8488 | 19.5801 | 19.6523 | 45.3626 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014392 | 0.23448 | 0.003375 | 16.2923 | 20.2015 | 3.1273 | 0.008284 | -0.5653 | -0.3527 | -0.6214 | -0.5492
2012-10-30 14:59 | 23.9793 | 19.6681 | 19.8245 | 19.6441 | 19.64 | 45.3713 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.2213 | 3.1279 | 0.008285 | -0.5532 | -0.3968 | -0.5772 | -0.5813
2012-10-30 14:59 | 23.9778 | 19.6626 | 19.8271 | 19.5864 | 19.6506 | 45.3649 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2559 | 3.1275 | 0.008287 | -0.5933 | -0.4288 | -0.6695 | -0.6053
2012-10-30 14:59 | 23.977 | 19.6618 | 19.8663 | 19.6297 | 19.6297 | 45.3586 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.2551 | 3.127 | 0.00828 | -0.5933 | -0.3888 | -0.6254 | -0.6254
2012-10-30 14:59 | 23.9778 | 19.6666 | 19.8712 | 19.6666 | 19.6506 | 45.3696 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.2519 | 3.1278 | 0.008285 | -0.5853 | -0.3807 | -0.5853 | -0.6013
2012-10-30 14:59 | 23.974 | 19.6869 | 19.8553 | 19.5706 | 19.5866 | 45.3661 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.2281 | 3.1275 | 0.008287 | -0.5412 | -0.3728 | -0.6575 | -0.6415
2012-10-30 14:59 | 23.9763 | 19.6651 | 19.8817 | 19.605 | 19.6571 | 45.3748 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014417 | 0.23455 | 0.003382 | 16.2691 | 20.2103 | 3.1281 | 0.008286 | -0.5452 | -0.3286 | -0.6053 | -0.5532
2012-10-30 14:59 | 23.974 | 19.6749 | 19.8633 | 19.5746 | 19.6267 | 45.3667 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.212 | 3.1276 | 0.008284 | -0.5371 | -0.3487 | -0.6374 | -0.5853
2012-10-30 14:59 | 23.9808 | 19.6736 | 19.8421 | 19.6255 | 19.6335 | 45.3591 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.2228 | 3.1271 | 0.008285 | -0.5492 | -0.3807 | -0.5973 | -0.5893
2012-10-30 14:59 | 23.9815 | 19.6262 | 19.8428 | 19.6102 | 19.6382 | 45.362 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.2596 | 3.1273 | 0.008287 | -0.6334 | -0.4168 | -0.6494 | -0.6214
2012-10-30 14:59 | 23.9838 | 19.6445 | 19.805 | 19.5964 | 19.6285 | 45.3707 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.2459 | 3.1279 | 0.008284 | -0.6014 | -0.4409 | -0.6495 | -0.6174
2012-10-30 14:59 | 23.9883 | 19.645 | 19.8656 | 19.6009 | 19.629 | 45.3655 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.2343 | 3.1275 | 0.008283 | -0.5893 | -0.3687 | -0.6334 | -0.6053
2012-10-30 14:59 | 23.9808 | 19.6295 | 19.826 | 19.5814 | 19.5934 | 45.3615 | 0.014197 | 0.34538 | 0.004903 | 24.3279 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.2709 | 3.1272 | 0.008277 | -0.6414 | -0.4449 | -0.6895 | -0.6775
2012-10-30 14:59 | 23.9883 | 19.6691 | 19.8416 | 19.6049 | 19.625 | 45.3591 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014423 | 0.23459 | 0.003384 | 16.2653 | 20.2464 | 3.1271 | 0.008292 | -0.5773 | -0.4048 | -0.6415 | -0.6214
2012-10-30 14:59 | 23.989 | 19.6778 | 19.8342 | 19.5936 | 19.6618 | 45.3574 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.227 | 3.1269 | 0.008284 | -0.5492 | -0.3928 | -0.6334 | -0.5652
2012-10-30 15:00 | 23.9875 | 19.6643 | 19.8528 | 19.6242 | 19.6121 | 45.369 | 0.014209 | 0.34534 | 0.004907 | 24.3043 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2536 | 3.1277 | 0.008286 | -0.5893 | -0.4008 | -0.6294 | -0.6415
2012-10-30 15:00 | 23.9912 | 19.6519 | 19.8445 | 19.6158 | 19.6279 | 45.3586 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.2533 | 3.127 | 0.008285 | -0.6014 | -0.4088 | -0.6375 | -0.6254
2012-10-30 15:00 | 23.9965 | 19.6773 | 19.8257 | 19.6171 | 19.6332 | 45.3545 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2626 | 3.1267 | 0.008288 | -0.5853 | -0.4369 | -0.6455 | -0.6294
2012-10-30 15:00 | 23.9927 | 19.6855 | 19.8339 | 19.6294 | 19.6133 | 45.3615 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014417 | 0.23457 | 0.003382 | 16.2701 | 20.3028 | 3.1272 | 0.008291 | -0.6173 | -0.4689 | -0.6734 | -0.6895
2012-10-30 15:00 | 23.9957 | 19.6725 | 19.845 | 19.6203 | 19.6203 | 45.3551 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2578 | 3.1268 | 0.00829 | -0.5853 | -0.4128 | -0.6375 | -0.6375
2012-10-30 15:00 | 24.001 | 19.6417 | 19.8583 | 19.6377 | 19.6216 | 45.3568 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.2631 | 3.1269 | 0.008289 | -0.6214 | -0.4048 | -0.6254 | -0.6415
2012-10-30 15:00 | 23.9957 | 19.6605 | 19.845 | 19.6083 | 19.6043 | 45.3667 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2578 | 3.1276 | 0.008288 | -0.5973 | -0.4128 | -0.6495 | -0.6535
2012-10-30 15:00 | 23.9965 | 19.6211 | 19.8738 | 19.6252 | 19.589 | 45.3719 | 0.014225 | 0.34543 | 0.004914 | 24.2831 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.2506 | 3.1279 | 0.008293 | -0.6295 | -0.3768 | -0.6254 | -0.6616
2012-10-30 15:00 | 24.0002 | 19.685 | 19.8174 | 19.6248 | 19.6289 | 45.3626 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.2903 | 3.1273 | 0.008285 | -0.6053 | -0.4729 | -0.6655 | -0.6614
2012-10-30 15:00 | 24.0047 | 19.6695 | 19.8379 | 19.6374 | 19.6334 | 45.3568 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2628 | 3.1269 | 0.008284 | -0.5933 | -0.4249 | -0.6254 | -0.6294
2012-10-30 15:00 | 24.0077 | 19.6805 | 19.8369 | 19.6243 | 19.6203 | 45.3609 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.2577 | 3.1272 | 0.008287 | -0.5772 | -0.4208 | -0.6334 | -0.6374
2012-10-30 15:00 | 24.0084 | 19.633 | 19.8697 | 19.621 | 19.621 | 45.3707 | 0.014209 | 0.34537 | 0.004907 | 24.3064 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.2705 | 3.1279 | 0.008284 | -0.6375 | -0.4008 | -0.6495 | -0.6495
2012-10-30 15:00 | 24.004 | 19.6447 | 19.8533 | 19.6006 | 19.6086 | 45.3591 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.2941 | 3.1271 | 0.008285 | -0.6494 | -0.4408 | -0.6935 | -0.6855
2012-10-30 15:00 | 24.0077 | 19.6644 | 19.8409 | 19.6725 | 19.6404 | 45.358 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.3218 | 3.127 | 0.008288 | -0.6574 | -0.4809 | -0.6493 | -0.6814
2012-10-30 15:00 | 24.0084 | 19.6651 | 19.8376 | 19.609 | 19.633 | 45.3632 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2825 | 3.1273 | 0.008284 | -0.6174 | -0.4449 | -0.6735 | -0.6495
2012-10-30 15:00 | 24.0084 | 19.6732 | 19.8336 | 19.617 | 19.6411 | 45.3678 | 0.014212 | 0.34537 | 0.004908 | 24.3012 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.2785 | 3.1277 | 0.008288 | -0.6053 | -0.4449 | -0.6615 | -0.6374
2012-10-30 15:00 | 23.9942 | 19.687 | 19.8354 | 19.6309 | 19.6309 | 45.3424 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2723 | 3.1259 | 0.00829 | -0.5853 | -0.4369 | -0.6414 | -0.6414
2012-10-30 15:00 | 24.004 | 19.6768 | 19.8773 | 19.6367 | 19.6046 | 45.3557 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2781 | 3.1268 | 0.008282 | -0.6013 | -0.4008 | -0.6414 | -0.6735
2012-10-30 15:00 | 24.0017 | 19.6825 | 19.867 | 19.6584 | 19.6183 | 45.3655 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.2958 | 3.1275 | 0.008286 | -0.6133 | -0.4288 | -0.6374 | -0.6775
2012-10-30 15:00 | 24.001 | 19.6497 | 19.8703 | 19.6136 | 19.6537 | 45.358 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2591 | 3.127 | 0.008285 | -0.6094 | -0.3888 | -0.6455 | -0.6054
2012-10-30 15:00 | 24.001 | 19.6377 | 19.8743 | 19.6417 | 19.6176 | 45.358 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2951 | 3.127 | 0.008286 | -0.6574 | -0.4208 | -0.6534 | -0.6775
2012-10-30 15:00 | 24.001 | 19.6417 | 19.8663 | 19.6457 | 19.5935 | 45.3609 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.2991 | 3.1272 | 0.008288 | -0.6574 | -0.4328 | -0.6534 | -0.7056
2012-10-30 15:01 | 24.0032 | 19.6479 | 19.8565 | 19.6399 | 19.6278 | 45.3516 | 0.014212 | 0.34538 | 0.004909 | 24.3022 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.3013 | 3.1265 | 0.008289 | -0.6534 | -0.4448 | -0.6614 | -0.6735
2012-10-30 15:01 | 23.9957 | 19.6524 | 19.877 | 19.6725 | 19.6524 | 45.3586 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2898 | 3.127 | 0.008282 | -0.6374 | -0.4128 | -0.6173 | -0.6374
2012-10-30 15:01 | 24.0017 | 19.6624 | 19.859 | 19.6183 | 19.6544 | 45.3609 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.3158 | 3.1272 | 0.008281 | -0.6534 | -0.4568 | -0.6975 | -0.6614
2012-10-30 15:01 | 24.0047 | 19.6735 | 19.9181 | 19.6534 | 19.6293 | 45.3638 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.3068 | 3.1274 | 0.008282 | -0.6333 | -0.3887 | -0.6534 | -0.6775
2012-10-30 15:01 | 24.0017 | 19.6504 | 19.8991 | 19.6183 | 19.6304 | 45.3591 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.3198 | 3.1271 | 0.008281 | -0.6694 | -0.4207 | -0.7015 | -0.6894
2012-10-30 15:01 | 23.9987 | 19.6314 | 19.888 | 19.6073 | 19.6474 | 45.3655 | 0.014209 | 0.34537 | 0.004907 | 24.3064 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.2808 | 3.1275 | 0.008286 | -0.6494 | -0.3928 | -0.6735 | -0.6334
2012-10-30 15:01 | 24.0032 | 19.6639 | 19.8685 | 19.6118 | 19.5998 | 45.3603 | 0.014203 | 0.34534 | 0.004905 | 24.3146 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.2933 | 3.1271 | 0.008282 | -0.6294 | -0.4248 | -0.6815 | -0.6935
2012-10-30 15:01 | 23.9957 | 19.6564 | 19.881 | 19.6244 | 19.6284 | 45.3667 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.3098 | 3.1276 | 0.008286 | -0.6534 | -0.4288 | -0.6854 | -0.6814
2012-10-30 15:01 | 24.0077 | 19.6805 | 19.881 | 19.6404 | 19.6323 | 45.3557 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.3258 | 3.1268 | 0.008282 | -0.6453 | -0.4448 | -0.6854 | -0.6935
2012-10-30 15:01 | 24.0002 | 19.6449 | 19.8294 | 19.6248 | 19.6569 | 45.3759 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.2943 | 3.1282 | 0.008284 | -0.6494 | -0.4649 | -0.6695 | -0.6374
2012-10-30 15:01 | 23.9965 | 19.6412 | 19.9019 | 19.6011 | 19.6693 | 45.3551 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.3106 | 3.1268 | 0.008283 | -0.6694 | -0.4087 | -0.7095 | -0.6413
2012-10-30 15:01 | 23.9957 | 19.6685 | 19.877 | 19.6284 | 19.6605 | 45.3482 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23445 | 0.003379 | 16.2688 | 20.2898 | 3.1263 | 0.008285 | -0.6213 | -0.4128 | -0.6614 | -0.6293
2012-10-30 15:01 | 23.9898 | 19.6465 | 19.8791 | 19.6385 | 19.6345 | 45.347 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.2919 | 3.1262 | 0.008284 | -0.6454 | -0.4128 | -0.6534 | -0.6574
2012-10-30 15:01 | 23.9942 | 19.6429 | 19.8916 | 19.5988 | 19.6469 | 45.369 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.3003 | 3.1277 | 0.008283 | -0.6574 | -0.4087 | -0.7015 | -0.6534
2012-10-30 15:01 | 23.9935 | 19.6542 | 19.8788 | 19.582 | 19.6382 | 45.3591 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.2756 | 3.1271 | 0.008287 | -0.6214 | -0.3968 | -0.6936 | -0.6374
2012-10-30 15:01 | 23.992 | 19.6327 | 19.8613 | 19.6287 | 19.6447 | 45.3407 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.2861 | 3.1258 | 0.008285 | -0.6534 | -0.4248 | -0.6574 | -0.6414
2012-10-30 15:01 | 23.992 | 19.6367 | 19.8693 | 19.5725 | 19.6487 | 45.358 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.2701 | 3.127 | 0.008285 | -0.6334 | -0.4008 | -0.6976 | -0.6214
2012-10-30 15:01 | 23.9868 | 19.6395 | 19.8481 | 19.6315 | 19.6556 | 45.3563 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.2769 | 3.1269 | 0.008287 | -0.6374 | -0.4288 | -0.6454 | -0.6213
2012-10-30 15:01 | 23.9957 | 19.6885 | 19.857 | 19.6203 | 19.6564 | 45.3395 | 0.014225 | 0.34544 | 0.004914 | 24.2841 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.2978 | 3.1257 | 0.008294 | -0.6093 | -0.4408 | -0.6775 | -0.6414
2012-10-30 15:01 | 23.989 | 19.6297 | 19.8423 | 19.6337 | 19.6538 | 45.3441 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.2511 | 3.126 | 0.008285 | -0.6214 | -0.4088 | -0.6174 | -0.5973
2012-10-30 15:01 | 23.9853 | 19.6541 | 19.8626 | 19.6019 | 19.6621 | 45.3568 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.2634 | 3.1269 | 0.008285 | -0.6093 | -0.4008 | -0.6615 | -0.6013
2012-10-30 15:01 | 23.992 | 19.6327 | 19.8533 | 19.6327 | 19.6487 | 45.3459 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.2741 | 3.1261 | 0.008284 | -0.6414 | -0.4208 | -0.6414 | -0.6254
2012-10-30 15:02 | 23.9927 | 19.7016 | 19.874 | 19.6173 | 19.6093 | 45.3487 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014408 | 0.23442 | 0.003378 | 16.2702 | 20.2828 | 3.1263 | 0.008285 | -0.5812 | -0.4088 | -0.6655 | -0.6735
2012-10-30 15:02 | 23.998 | 19.6628 | 19.8392 | 19.6267 | 19.6467 | 45.3511 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2641 | 3.1265 | 0.008284 | -0.6013 | -0.4249 | -0.6374 | -0.6174
2012-10-30 15:02 | 23.9965 | 19.6773 | 19.8377 | 19.5931 | 19.6733 | 45.3499 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.2866 | 3.1264 | 0.008291 | -0.6093 | -0.4489 | -0.6935 | -0.6133
2012-10-30 15:02 | 24.0047 | 19.6053 | 19.87 | 19.6414 | 19.6053 | 45.3511 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014417 | 0.23444 | 0.00338 | 16.2611 | 20.2668 | 3.1265 | 0.008287 | -0.6615 | -0.3968 | -0.6254 | -0.6615
2012-10-30 15:02 | 24.0032 | 19.6479 | 19.8364 | 19.6038 | 19.6639 | 45.3274 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 20.2893 | 3.1249 | 0.008288 | -0.6414 | -0.4529 | -0.6855 | -0.6254
2012-10-30 15:02 | 24.0002 | 19.6369 | 19.8013 | 19.5807 | 19.5887 | 45.3603 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.2863 | 3.1271 | 0.008284 | -0.6494 | -0.485 | -0.7056 | -0.6976
2012-10-30 15:02 | 24.0055 | 19.6542 | 19.8507 | 19.6662 | 19.6221 | 45.3516 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2876 | 3.1265 | 0.008289 | -0.6334 | -0.4369 | -0.6214 | -0.6655
2012-10-30 15:02 | 24.0107 | 19.6755 | 19.876 | 19.6554 | 19.6153 | 45.336 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.2648 | 3.1255 | 0.008283 | -0.5893 | -0.3888 | -0.6094 | -0.6495
2012-10-30 15:02 | 24.0055 | 19.6542 | 19.8628 | 19.598 | 19.6301 | 45.3655 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2996 | 3.1275 | 0.008291 | -0.6454 | -0.4368 | -0.7016 | -0.6695
2012-10-30 15:02 | 24.0114 | 19.6561 | 19.8727 | 19.6401 | 19.6401 | 45.3487 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.3095 | 3.1263 | 0.008283 | -0.6534 | -0.4368 | -0.6694 | -0.6694
2012-10-30 15:02 | 24.0084 | 19.6411 | 19.8737 | 19.617 | 19.617 | 45.3476 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.2825 | 3.1263 | 0.008284 | -0.6414 | -0.4088 | -0.6655 | -0.6655
2012-10-30 15:02 | 24.0062 | 19.6389 | 19.8394 | 19.5867 | 19.6349 | 45.3505 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.2683 | 3.1265 | 0.008286 | -0.6294 | -0.4289 | -0.6816 | -0.6334
2012-10-30 15:02 | 24.0122 | 19.677 | 19.8494 | 19.6007 | 19.6088 | 45.3505 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.2502 | 3.1265 | 0.008286 | -0.5732 | -0.4008 | -0.6495 | -0.6414
2012-10-30 15:02 | 24.0047 | 19.6574 | 19.8499 | 19.6093 | 19.6253 | 45.3418 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2748 | 3.1259 | 0.008288 | -0.6174 | -0.4249 | -0.6655 | -0.6495
2012-10-30 15:02 | 24.0077 | 19.6283 | 19.8409 | 19.6123 | 19.6163 | 45.3505 | 0.014209 | 0.34537 | 0.004907 | 24.3064 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.2858 | 3.1265 | 0.008285 | -0.6575 | -0.4449 | -0.6735 | -0.6695
2012-10-30 15:02 | 24.0107 | 19.6233 | 19.876 | 19.6113 | 19.6514 | 45.3453 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.3128 | 3.1261 | 0.008283 | -0.6895 | -0.4368 | -0.7015 | -0.6614
2012-10-30 15:02 | 24.007 | 19.6116 | 19.8683 | 19.6156 | 19.6557 | 45.3511 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.3051 | 3.1265 | 0.008288 | -0.6935 | -0.4368 | -0.6895 | -0.6494
2012-10-30 15:02 | 24.0062 | 19.6509 | 19.8514 | 19.5987 | 19.6108 | 45.3482 | 0.014219 | 0.34547 | 0.004912 | 24.2964 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.2643 | 3.1263 | 0.00829 | -0.6134 | -0.4129 | -0.6656 | -0.6535
2012-10-30 15:02 | 24.0107 | 19.6554 | 19.8439 | 19.6353 | 19.6032 | 45.3487 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.3128 | 3.1263 | 0.008287 | -0.6574 | -0.4689 | -0.6775 | -0.7096
2012-10-30 15:02 | 24.0099 | 19.6586 | 19.8391 | 19.6386 | 19.6426 | 45.3355 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.2599 | 3.1254 | 0.008289 | -0.6013 | -0.4208 | -0.6213 | -0.6173
2012-10-30 15:03 | 24.0092 | 19.6258 | 19.8625 | 19.6098 | 19.6298 | 45.3459 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.2552 | 3.1261 | 0.008286 | -0.6294 | -0.3927 | -0.6454 | -0.6254
2012-10-30 15:03 | 24.0092 | 19.6379 | 19.8625 | 19.6218 | 19.5857 | 45.3499 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014389 | 0.23449 | 0.003374 | 16.2968 | 20.2633 | 3.1264 | 0.008282 | -0.6254 | -0.4008 | -0.6415 | -0.6776
2012-10-30 15:03 | 24.0099 | 19.6185 | 19.8191 | 19.6386 | 19.6185 | 45.3459 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2399 | 3.1261 | 0.008284 | -0.6214 | -0.4208 | -0.6013 | -0.6214
2012-10-30 15:03 | 24.0077 | 19.6524 | 19.8369 | 19.6123 | 19.6083 | 45.332 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2818 | 3.1252 | 0.008284 | -0.6294 | -0.4449 | -0.6695 | -0.6735
2012-10-30 15:03 | 24.0144 | 19.6431 | 19.8155 | 19.635 | 19.631 | 45.3493 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.2644 | 3.1264 | 0.008283 | -0.6213 | -0.4489 | -0.6294 | -0.6334
2012-10-30 15:03 | 24.0182 | 19.683 | 19.8193 | 19.5987 | 19.6228 | 45.3568 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.3043 | 3.1269 | 0.008285 | -0.6213 | -0.485 | -0.7056 | -0.6815
2012-10-30 15:03 | 24.0242 | 19.6448 | 19.8855 | 19.6047 | 19.6248 | 45.3418 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.2943 | 3.1259 | 0.008283 | -0.6495 | -0.4088 | -0.6896 | -0.6695
2012-10-30 15:03 | 24.0182 | 19.6228 | 19.8675 | 19.5947 | 19.6388 | 45.3441 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.2642 | 3.126 | 0.008282 | -0.6414 | -0.3967 | -0.6695 | -0.6254
2012-10-30 15:03 | 24.0279 | 19.6606 | 19.8371 | 19.6164 | 19.6204 | 45.3395 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014411 | 0.23442 | 0.003378 | 16.2668 | 20.2659 | 3.1257 | 0.008285 | -0.6053 | -0.4288 | -0.6495 | -0.6455
2012-10-30 15:03 | 24.0204 | 19.625 | 19.8656 | 19.621 | 19.637 | 45.3591 | 0.014219 | 0.3455 | 0.004913 | 24.2984 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.2464 | 3.1271 | 0.00829 | -0.6214 | -0.3808 | -0.6254 | -0.6094
2012-10-30 15:03 | 24.0279 | 19.6325 | 19.8611 | 19.6445 | 19.6525 | 45.3522 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.2659 | 3.1266 | 0.008285 | -0.6334 | -0.4048 | -0.6214 | -0.6134
2012-10-30 15:03 | 24.0234 | 19.648 | 19.8085 | 19.6159 | 19.636 | 45.3383 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.2694 | 3.1256 | 0.008287 | -0.6214 | -0.4609 | -0.6535 | -0.6334
2012-10-30 15:03 | 24.0167 | 19.5932 | 19.8339 | 19.5691 | 19.6373 | 45.3378 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.2667 | 3.1256 | 0.008282 | -0.6735 | -0.4328 | -0.6976 | -0.6294
2012-10-30 15:03 | 24.0286 | 19.6532 | 19.8418 | 19.6091 | 19.6332 | 45.3464 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2586 | 3.1262 | 0.008289 | -0.6054 | -0.4168 | -0.6495 | -0.6254
2012-10-30 15:03 | 24.0271 | 19.6277 | 19.8162 | 19.6036 | 19.6156 | 45.3626 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.2771 | 3.1273 | 0.008288 | -0.6494 | -0.4609 | -0.6735 | -0.6615
2012-10-30 15:03 | 24.0249 | 19.5853 | 19.8541 | 19.6094 | 19.5813 | 45.3499 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.2629 | 3.1264 | 0.008284 | -0.6776 | -0.4088 | -0.6535 | -0.6816
2012-10-30 15:03 | 24.0309 | 19.6596 | 19.832 | 19.6234 | 19.6716 | 45.3349 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.2489 | 3.1254 | 0.008283 | -0.5893 | -0.4169 | -0.6255 | -0.5773
2012-10-30 15:03 | 24.0316 | 19.6763 | 19.8608 | 19.588 | 19.6161 | 45.3308 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2536 | 3.1251 | 0.008284 | -0.5773 | -0.3928 | -0.6656 | -0.6375
2012-10-30 15:03 | 24.0316 | 19.6001 | 19.8448 | 19.588 | 19.6522 | 45.3568 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.2416 | 3.1269 | 0.008286 | -0.6415 | -0.3968 | -0.6536 | -0.5894
2012-10-30 15:03 | 24.0316 | 19.6402 | 19.8408 | 19.584 | 19.6763 | 45.3447 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.2656 | 3.1261 | 0.008288 | -0.6254 | -0.4248 | -0.6816 | -0.5893
2012-10-30 15:03 | 24.0354 | 19.6641 | 19.8446 | 19.632 | 19.656 | 45.3383 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014423 | 0.23455 | 0.003383 | 16.2624 | 20.2574 | 3.1256 | 0.008293 | -0.5933 | -0.4128 | -0.6254 | -0.6014
2012-10-30 15:03 | 24.0309 | 19.6395 | 19.8441 | 19.5913 | 19.6194 | 45.3407 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014414 | 0.23459 | 0.003381 | 16.2755 | 20.2569 | 3.1258 | 0.008287 | -0.6174 | -0.4128 | -0.6656 | -0.6375
2012-10-30 15:04 | 24.0339 | 19.6585 | 19.8671 | 19.6425 | 19.6626 | 45.3349 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.2359 | 3.1254 | 0.008281 | -0.5774 | -0.3688 | -0.5934 | -0.5733
2012-10-30 15:04 | 24.0301 | 19.6226 | 19.8152 | 19.5785 | 19.6507 | 45.3539 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.2521 | 3.1267 | 0.008291 | -0.6295 | -0.4369 | -0.6736 | -0.6014
2012-10-30 15:04 | 24.0331 | 19.6577 | 19.8904 | 19.6176 | 19.6537 | 45.3453 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2511 | 3.1261 | 0.008288 | -0.5934 | -0.3607 | -0.6335 | -0.5974
2012-10-30 15:04 | 24.0324 | 19.645 | 19.8696 | 19.6129 | 19.649 | 45.3499 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.2384 | 3.1264 | 0.008285 | -0.5934 | -0.3688 | -0.6255 | -0.5894
2012-10-30 15:04 | 24.0316 | 19.6522 | 19.8608 | 19.6322 | 19.6241 | 45.3349 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014395 | 0.23442 | 0.003374 | 16.2849 | 20.2616 | 3.1254 | 0.008282 | -0.6094 | -0.4008 | -0.6294 | -0.6375
2012-10-30 15:04 | 24.0271 | 19.6838 | 19.8322 | 19.6156 | 19.6477 | 45.3314 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.2892 | 3.1251 | 0.008287 | -0.6054 | -0.457 | -0.6736 | -0.6415
2012-10-30 15:04 | 24.0354 | 19.6681 | 19.8726 | 19.6039 | 19.632 | 45.3424 | 0.014206 | 0.34536 | 0.004906 | 24.3105 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.2133 | 3.1259 | 0.008281 | -0.5452 | -0.3407 | -0.6094 | -0.5813
2012-10-30 15:04 | 24.0294 | 19.6219 | 19.8546 | 19.6099 | 19.6219 | 45.336 | 0.014222 | 0.34546 | 0.004913 | 24.2903 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.2995 | 3.1255 | 0.008291 | -0.6776 | -0.4449 | -0.6896 | -0.6776
2012-10-30 15:04 | 24.0242 | 19.6448 | 19.8574 | 19.6167 | 19.6087 | 45.3378 | 0.014216 | 0.34538 | 0.00491 | 24.2954 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.2903 | 3.1256 | 0.008286 | -0.6455 | -0.4329 | -0.6736 | -0.6816
2012-10-30 15:04 | 24.0301 | 19.6748 | 19.8874 | 19.6186 | 19.6668 | 45.332 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.2801 | 3.1252 | 0.00829 | -0.6053 | -0.3927 | -0.6615 | -0.6133
2012-10-30 15:04 | 24.0264 | 19.643 | 19.8676 | 19.6189 | 19.651 | 45.3505 | 0.014222 | 0.34544 | 0.004913 | 24.2892 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.2684 | 3.1265 | 0.008291 | -0.6254 | -0.4008 | -0.6495 | -0.6174
2012-10-30 15:04 | 24.0219 | 19.6506 | 19.8511 | 19.6064 | 19.6425 | 45.347 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.2639 | 3.1262 | 0.008289 | -0.6133 | -0.4128 | -0.6575 | -0.6214
2012-10-30 15:04 | 24.0256 | 19.6623 | 19.8307 | 19.6141 | 19.582 | 45.3516 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.2797 | 3.1265 | 0.008284 | -0.6174 | -0.449 | -0.6656 | -0.6977
2012-10-30 15:04 | 24.0271 | 19.6798 | 19.7921 | 19.6237 | 19.6437 | 45.3418 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.3132 | 3.1259 | 0.008286 | -0.6334 | -0.5211 | -0.6895 | -0.6695
2012-10-30 15:04 | 24.0316 | 19.6803 | 19.8207 | 19.6402 | 19.6201 | 45.3274 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.2977 | 3.1249 | 0.008289 | -0.6174 | -0.477 | -0.6575 | -0.6776
2012-10-30 15:04 | 24.0339 | 19.6585 | 19.835 | 19.5783 | 19.6465 | 45.3331 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23458 | 0.003379 | 16.2858 | 20.304 | 3.1253 | 0.008287 | -0.6455 | -0.469 | -0.7257 | -0.6575
2012-10-30 15:04 | 24.0234 | 19.632 | 19.8125 | 19.5798 | 19.64 | 45.3453 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.2895 | 3.1261 | 0.008282 | -0.6575 | -0.477 | -0.7097 | -0.6495
2012-10-30 15:04 | 24.0361 | 19.6487 | 19.8332 | 19.6728 | 19.6367 | 45.3326 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.2501 | 3.1252 | 0.008282 | -0.6014 | -0.4169 | -0.5773 | -0.6134
2012-10-30 15:04 | 24.0309 | 19.6957 | 19.7959 | 19.6154 | 19.6074 | 45.3343 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.2609 | 3.1253 | 0.008288 | -0.5652 | -0.465 | -0.6455 | -0.6535
2012-10-30 15:04 | 24.0391 | 19.6357 | 19.8041 | 19.6236 | 19.6517 | 45.3441 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.2291 | 3.126 | 0.008285 | -0.5934 | -0.425 | -0.6055 | -0.5774
2012-10-30 15:04 | 24.0451 | 19.6818 | 19.8101 | 19.6096 | 19.6296 | 45.3511 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2631 | 3.1265 | 0.008287 | -0.5813 | -0.453 | -0.6535 | -0.6335
2012-10-30 15:04 | 24.0346 | 19.6713 | 19.8077 | 19.6352 | 19.6271 | 45.3401 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.2646 | 3.1257 | 0.008278 | -0.5933 | -0.4569 | -0.6294 | -0.6375
2012-10-30 15:04 | 24.0354 | 19.6841 | 19.8325 | 19.5517 | 19.644 | 45.332 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.2614 | 3.1252 | 0.008291 | -0.5773 | -0.4289 | -0.7097 | -0.6174
2012-10-30 15:05 | 24.0428 | 19.6394 | 19.8319 | 19.6032 | 19.6273 | 45.3256 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014392 | 0.23454 | 0.003375 | 16.2963 | 20.2848 | 3.1247 | 0.008282 | -0.6454 | -0.4529 | -0.6816 | -0.6575
2012-10-30 15:05 | 24.0421 | 19.7069 | 19.8553 | 19.6306 | 19.6146 | 45.3343 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 20.2641 | 3.1253 | 0.008288 | -0.5572 | -0.4088 | -0.6335 | -0.6495
2012-10-30 15:05 | 24.0316 | 19.6964 | 19.8367 | 19.5961 | 19.5961 | 45.3435 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.2296 | 3.126 | 0.008283 | -0.5332 | -0.3929 | -0.6335 | -0.6335
2012-10-30 15:05 | 24.0286 | 19.6693 | 19.8177 | 19.6091 | 19.6252 | 45.3383 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.1985 | 3.1256 | 0.008282 | -0.5292 | -0.3808 | -0.5894 | -0.5733
2012-10-30 15:05 | 24.0361 | 19.7009 | 19.8172 | 19.6246 | 19.6367 | 45.3326 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.186 | 3.1252 | 0.008287 | -0.4851 | -0.3688 | -0.5614 | -0.5493
2012-10-30 15:05 | 24.0339 | 19.6585 | 19.8471 | 19.6184 | 19.5863 | 45.3464 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.1998 | 3.1262 | 0.008282 | -0.5413 | -0.3527 | -0.5814 | -0.6135
2012-10-30 15:05 | 24.0264 | 19.6912 | 19.8315 | 19.6029 | 19.5949 | 45.3326 | 0.014216 | 0.34541 | 0.00491 | 24.2974 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.1843 | 3.1252 | 0.008287 | -0.4931 | -0.3528 | -0.5814 | -0.5894
2012-10-30 15:05 | 24.0242 | 19.6689 | 19.7892 | 19.5927 | 19.5967 | 45.3251 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014389 | 0.23447 | 0.003374 | 16.2947 | 20.1661 | 3.1247 | 0.00828 | -0.4972 | -0.3769 | -0.5734 | -0.5694
2012-10-30 15:05 | 24.0301 | 19.7149 | 19.7871 | 19.6026 | 19.6267 | 45.3418 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.196 | 3.1259 | 0.008284 | -0.4811 | -0.4089 | -0.5934 | -0.5693
2012-10-30 15:05 | 24.0286 | 19.7054 | 19.8378 | 19.6332 | 19.6131 | 45.3418 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.1985 | 3.1259 | 0.008285 | -0.4931 | -0.3607 | -0.5653 | -0.5854
2012-10-30 15:05 | 24.0264 | 19.6992 | 19.8235 | 19.6189 | 19.635 | 45.3464 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.2003 | 3.1262 | 0.008285 | -0.5011 | -0.3768 | -0.5814 | -0.5653
2012-10-30 15:05 | 24.0294 | 19.6942 | 19.8105 | 19.6139 | 19.638 | 45.3308 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014404 | 0.23445 | 0.003377 | 16.2767 | 20.2554 | 3.1251 | 0.008283 | -0.5612 | -0.4449 | -0.6415 | -0.6174
2012-10-30 15:05 | 24.0227 | 19.6554 | 19.7998 | 19.5872 | 19.6233 | 45.3435 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.2727 | 3.126 | 0.008283 | -0.6173 | -0.4729 | -0.6855 | -0.6494
2012-10-30 15:05 | 24.0279 | 19.6566 | 19.829 | 19.5843 | 19.5924 | 45.3349 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014423 | 0.23451 | 0.003382 | 16.2593 | 20.302 | 3.1254 | 0.008285 | -0.6454 | -0.473 | -0.7177 | -0.7096
2012-10-30 15:05 | 24.0346 | 19.6873 | 19.8357 | 19.6472 | 19.6151 | 45.3372 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014401 | 0.23444 | 0.003376 | 16.2791 | 20.2566 | 3.1255 | 0.008284 | -0.5693 | -0.4209 | -0.6094 | -0.6415
2012-10-30 15:05 | 24.0346 | 19.6713 | 19.8317 | 19.575 | 19.6231 | 45.336 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.2446 | 3.1255 | 0.008289 | -0.5733 | -0.4129 | -0.6696 | -0.6215
2012-10-30 15:05 | 24.0301 | 19.6668 | 19.8352 | 19.6186 | 19.6226 | 45.3476 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.2321 | 3.1263 | 0.008288 | -0.5653 | -0.3969 | -0.6135 | -0.6095
2012-10-30 15:05 | 24.0294 | 19.6781 | 19.8185 | 19.5738 | 19.5939 | 45.3412 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.2514 | 3.1258 | 0.008285 | -0.5733 | -0.4329 | -0.6776 | -0.6575
2012-10-30 15:05 | 24.0271 | 19.6437 | 19.8322 | 19.5956 | 19.6357 | 45.347 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014398 | 0.23445 | 0.003376 | 16.2835 | 20.2771 | 3.1262 | 0.008288 | -0.6334 | -0.4449 | -0.6815 | -0.6414
2012-10-30 15:05 | 24.0301 | 19.7109 | 19.8473 | 19.5946 | 19.6106 | 45.3464 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.2521 | 3.1262 | 0.008286 | -0.5412 | -0.4048 | -0.6575 | -0.6415
2012-10-30 15:05 | 24.0242 | 19.6569 | 19.8213 | 19.6127 | 19.6248 | 45.3291 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2702 | 3.125 | 0.008287 | -0.6133 | -0.4489 | -0.6575 | -0.6454
2012-10-30 15:05 | 24.0271 | 19.6838 | 19.8002 | 19.5956 | 19.6036 | 45.3245 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.3092 | 3.1247 | 0.008292 | -0.6254 | -0.509 | -0.7136 | -0.7056
2012-10-30 15:05 | 24.0346 | 19.6833 | 19.8558 | 19.6151 | 19.6312 | 45.3493 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.3087 | 3.1264 | 0.008283 | -0.6254 | -0.4529 | -0.6936 | -0.6775
2012-10-30 15:05 | 24.0339 | 19.6626 | 19.831 | 19.6024 | 19.6264 | 45.3389 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.2519 | 3.1257 | 0.008285 | -0.5893 | -0.4209 | -0.6495 | -0.6255
2012-10-30 15:06 | 24.0376 | 19.6542 | 19.8387 | 19.582 | 19.59 | 45.332 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.2876 | 3.1252 | 0.008281 | -0.6334 | -0.4489 | -0.7056 | -0.6976
2012-10-30 15:06 | 24.0346 | 19.6472 | 19.8397 | 19.583 | 19.6432 | 45.3366 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.2887 | 3.1255 | 0.008281 | -0.6415 | -0.449 | -0.7057 | -0.6455
2012-10-30 15:06 | 24.0279 | 19.6245 | 19.805 | 19.5924 | 19.6164 | 45.336 | 0.014216 | 0.34534 | 0.004909 | 24.2923 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.282 | 3.1255 | 0.008288 | -0.6575 | -0.477 | -0.6896 | -0.6656
2012-10-30 15:06 | 24.0414 | 19.666 | 19.8265 | 19.5978 | 19.5938 | 45.343 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.2754 | 3.1259 | 0.008285 | -0.6094 | -0.4489 | -0.6776 | -0.6816
2012-10-30 15:06 | 24.0428 | 19.6634 | 19.8239 | 19.6474 | 19.6193 | 45.3314 | 0.014222 | 0.34543 | 0.004913 | 24.2882 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.2848 | 3.1251 | 0.008292 | -0.6214 | -0.4609 | -0.6374 | -0.6655
2012-10-30 15:06 | 24.0346 | 19.6753 | 19.8638 | 19.567 | 19.6472 | 45.3308 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2846 | 3.1251 | 0.008288 | -0.6093 | -0.4208 | -0.7176 | -0.6374
2012-10-30 15:06 | 24.0376 | 19.6502 | 19.8147 | 19.6542 | 19.6061 | 45.347 | 0.014216 | 0.34541 | 0.00491 | 24.2974 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.2957 | 3.1262 | 0.008286 | -0.6455 | -0.481 | -0.6415 | -0.6896
2012-10-30 15:06 | 24.0458 | 19.6464 | 19.8229 | 19.6062 | 19.6183 | 45.3314 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2958 | 3.1251 | 0.008283 | -0.6494 | -0.4729 | -0.6896 | -0.6775
2012-10-30 15:06 | 24.0414 | 19.6741 | 19.8225 | 19.6219 | 19.658 | 45.3476 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014439 | 0.23449 | 0.003386 | 16.2403 | 20.2354 | 3.1263 | 0.00829 | -0.5613 | -0.4129 | -0.6135 | -0.5774
2012-10-30 15:06 | 24.0391 | 19.6918 | 19.8202 | 19.5915 | 19.6196 | 45.3568 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.2731 | 3.1269 | 0.008287 | -0.5813 | -0.4529 | -0.6816 | -0.6535
2012-10-30 15:06 | 24.0399 | 19.6966 | 19.829 | 19.6244 | 19.6284 | 45.3343 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.2739 | 3.1253 | 0.008286 | -0.5773 | -0.4449 | -0.6495 | -0.6455
2012-10-30 15:06 | 24.0451 | 19.6818 | 19.8462 | 19.6136 | 19.5975 | 45.3349 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.3192 | 3.1254 | 0.008285 | -0.6374 | -0.473 | -0.7056 | -0.7217
2012-10-30 15:06 | 24.0466 | 19.6672 | 19.8277 | 19.6391 | 19.591 | 45.3366 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014404 | 0.23458 | 0.003379 | 16.2858 | 20.3087 | 3.1255 | 0.008285 | -0.6415 | -0.481 | -0.6696 | -0.7177
2012-10-30 15:06 | 24.0399 | 19.6525 | 19.8089 | 19.6003 | 19.6204 | 45.3511 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.314 | 3.1265 | 0.008287 | -0.6615 | -0.5051 | -0.7137 | -0.6936
2012-10-30 15:06 | 24.0421 | 19.6748 | 19.8392 | 19.6146 | 19.5945 | 45.3285 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.2881 | 3.1249 | 0.008285 | -0.6133 | -0.4489 | -0.6735 | -0.6936
2012-10-30 15:06 | 24.0421 | 19.6708 | 19.8232 | 19.6025 | 19.6186 | 45.3389 | 0.014231 | 0.3455 | 0.004917 | 24.2779 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 20.2601 | 3.1257 | 0.008297 | -0.5893 | -0.4369 | -0.6576 | -0.6415
2012-10-30 15:06 | 24.0384 | 19.647 | 19.8435 | 19.6309 | 19.6149 | 45.3355 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.2364 | 3.1254 | 0.008283 | -0.5894 | -0.3929 | -0.6055 | -0.6215
2012-10-30 15:06 | 24.0421 | 19.6547 | 19.8513 | 19.6146 | 19.6306 | 45.3447 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.2801 | 3.1261 | 0.008289 | -0.6254 | -0.4288 | -0.6655 | -0.6495
2012-10-30 15:06 | 24.0473 | 19.6559 | 19.8685 | 19.6278 | 19.6479 | 45.3412 | 0.014219 | 0.34541 | 0.004911 | 24.2923 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.2653 | 3.1258 | 0.00829 | -0.6094 | -0.3968 | -0.6375 | -0.6174
2012-10-30 15:06 | 24.0406 | 19.6492 | 19.8177 | 19.6091 | 19.593 | 45.3389 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.2947 | 3.1257 | 0.008284 | -0.6455 | -0.477 | -0.6856 | -0.7017
2012-10-30 15:06 | 24.0443 | 19.701 | 19.8294 | 19.6088 | 19.6368 | 45.3505 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.2863 | 3.1265 | 0.008284 | -0.5853 | -0.4569 | -0.6775 | -0.6495
2012-10-30 15:06 | 24.0473 | 19.6599 | 19.8364 | 19.6037 | 19.6519 | 45.3389 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.3054 | 3.1257 | 0.008285 | -0.6455 | -0.469 | -0.7017 | -0.6535
2012-10-30 15:06 | 24.0458 | 19.6745 | 19.8349 | 19.5822 | 19.6584 | 45.3389 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.2918 | 3.1257 | 0.00828 | -0.6173 | -0.4569 | -0.7096 | -0.6334
2012-10-30 15:07 | 24.0436 | 19.6361 | 19.8167 | 19.588 | 19.588 | 45.3499 | 0.014194 | 0.34538 | 0.004902 | 24.333 | 0.014395 | 0.23445 | 0.003375 | 16.2869 | 20.2776 | 3.1264 | 0.008277 | -0.6415 | -0.4609 | -0.6896 | -0.6896
2012-10-30 15:07 | 24.0466 | 19.6552 | 19.8237 | 19.591 | 19.6151 | 45.3401 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2646 | 3.1257 | 0.008289 | -0.6094 | -0.4409 | -0.6736 | -0.6495
2012-10-30 15:07 | 24.0458 | 19.7266 | 19.8269 | 19.5982 | 19.6303 | 45.3181 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.2718 | 3.1242 | 0.008286 | -0.5452 | -0.4449 | -0.6736 | -0.6415
2012-10-30 15:07 | 24.0518 | 19.6443 | 19.8208 | 19.6323 | 19.6203 | 45.3378 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.3139 | 3.1256 | 0.008281 | -0.6696 | -0.4931 | -0.6816 | -0.6936
2012-10-30 15:07 | 24.0526 | 19.6973 | 19.8136 | 19.601 | 19.589 | 45.3395 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.2906 | 3.1257 | 0.008282 | -0.5933 | -0.477 | -0.6896 | -0.7016
2012-10-30 15:07 | 24.0563 | 19.6488 | 19.8414 | 19.6127 | 19.6167 | 45.3297 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.3023 | 3.125 | 0.008282 | -0.6535 | -0.4609 | -0.6896 | -0.6856
2012-10-30 15:07 | 24.0503 | 19.679 | 19.7993 | 19.5987 | 19.6228 | 45.3383 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.3244 | 3.1256 | 0.008286 | -0.6454 | -0.5251 | -0.7257 | -0.7016
2012-10-30 15:07 | 24.0541 | 19.6466 | 19.8151 | 19.6386 | 19.6065 | 45.3459 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014433 | 0.23451 | 0.003385 | 16.248 | 20.3282 | 3.1261 | 0.008293 | -0.6816 | -0.5131 | -0.6896 | -0.7217
2012-10-30 15:07 | 24.0571 | 19.6577 | 19.8141 | 19.5814 | 19.6135 | 45.3401 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.3232 | 3.1257 | 0.008287 | -0.6655 | -0.5091 | -0.7418 | -0.7097
2012-10-30 15:07 | 24.0541 | 19.6386 | 19.8432 | 19.6145 | 19.6426 | 45.3245 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.3202 | 3.1247 | 0.008288 | -0.6816 | -0.477 | -0.7057 | -0.6776
2012-10-30 15:07 | 24.0623 | 19.6869 | 19.8193 | 19.5987 | 19.6428 | 45.3459 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.3164 | 3.1261 | 0.008288 | -0.6295 | -0.4971 | -0.7177 | -0.6736
2012-10-30 15:07 | 24.0556 | 19.6201 | 19.8206 | 19.604 | 19.6281 | 45.332 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.3177 | 3.1252 | 0.008284 | -0.6976 | -0.4971 | -0.7137 | -0.6896
2012-10-30 15:07 | 24.0638 | 19.6804 | 19.8529 | 19.6002 | 19.6403 | 45.3262 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014423 | 0.23451 | 0.003382 | 16.2593 | 20.3419 | 3.1248 | 0.008286 | -0.6615 | -0.489 | -0.7417 | -0.7016
2012-10-30 15:07 | 24.0615 | 19.6581 | 19.8867 | 19.5818 | 19.6059 | 45.336 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.3356 | 3.1255 | 0.008286 | -0.6775 | -0.4489 | -0.7538 | -0.7297
2012-10-30 15:07 | 24.0593 | 19.6679 | 19.8283 | 19.6117 | 19.5957 | 45.3424 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.3134 | 3.1259 | 0.00829 | -0.6455 | -0.4851 | -0.7017 | -0.7177
2012-10-30 15:07 | 24.0556 | 19.6642 | 19.8046 | 19.6 | 19.592 | 45.3482 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.2856 | 3.1263 | 0.008288 | -0.6214 | -0.481 | -0.6856 | -0.6936
2012-10-30 15:07 | 24.0608 | 19.6814 | 19.8379 | 19.6132 | 19.6333 | 45.3251 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.3309 | 3.1247 | 0.008287 | -0.6495 | -0.493 | -0.7177 | -0.6976
2012-10-30 15:07 | 24.0511 | 19.6757 | 19.8522 | 19.6236 | 19.5835 | 45.3401 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.3532 | 3.1257 | 0.008287 | -0.6775 | -0.501 | -0.7296 | -0.7697
2012-10-30 15:07 | 24.0541 | 19.6908 | 19.8071 | 19.6306 | 19.6145 | 45.3372 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.3722 | 3.1255 | 0.008281 | -0.6814 | -0.5651 | -0.7416 | -0.7577
2012-10-30 15:07 | 24.0586 | 19.6712 | 19.7956 | 19.6271 | 19.6311 | 45.3464 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 20.3687 | 3.1262 | 0.008287 | -0.6975 | -0.5731 | -0.7416 | -0.7376
2012-10-30 15:07 | 24.0481 | 19.6888 | 19.8292 | 19.5965 | 19.6246 | 45.3412 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.3462 | 3.1258 | 0.008288 | -0.6574 | -0.517 | -0.7497 | -0.7216
2012-10-30 15:07 | 24.063 | 19.6876 | 19.82 | 19.5913 | 19.6315 | 45.3274 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.3331 | 3.1249 | 0.008289 | -0.6455 | -0.5131 | -0.7418 | -0.7016
2012-10-30 15:07 | 24.0593 | 19.7 | 19.8083 | 19.6157 | 19.6398 | 45.3326 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.3414 | 3.1252 | 0.008288 | -0.6414 | -0.5331 | -0.7257 | -0.7016
2012-10-30 15:07 | 24.0615 | 19.7102 | 19.8105 | 19.63 | 19.63 | 45.3401 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 20.3236 | 3.1257 | 0.008287 | -0.6134 | -0.5131 | -0.6936 | -0.6936
2012-10-30 15:08 | 24.0675 | 19.7082 | 19.7764 | 19.644 | 19.636 | 45.3407 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.3696 | 3.1258 | 0.008284 | -0.6614 | -0.5932 | -0.7256 | -0.7336
2012-10-30 15:08 | 24.0608 | 19.7015 | 19.7456 | 19.6092 | 19.6212 | 45.3378 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.3589 | 3.1256 | 0.008286 | -0.6574 | -0.6133 | -0.7497 | -0.7377
2012-10-30 15:08 | 24.0623 | 19.691 | 19.7952 | 19.6227 | 19.6187 | 45.3366 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.3644 | 3.1255 | 0.008284 | -0.6734 | -0.5692 | -0.7417 | -0.7457
2012-10-30 15:08 | 24.0638 | 19.6804 | 19.7727 | 19.6042 | 19.6042 | 45.336 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.3459 | 3.1255 | 0.008286 | -0.6655 | -0.5732 | -0.7417 | -0.7417
2012-10-30 15:08 | 24.0713 | 19.6558 | 19.8203 | 19.6197 | 19.6197 | 45.3522 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014423 | 0.23451 | 0.003382 | 16.2593 | 20.3494 | 3.1266 | 0.008288 | -0.6936 | -0.5291 | -0.7297 | -0.7297
2012-10-30 15:08 | 24.0638 | 19.6965 | 19.7847 | 19.6122 | 19.6002 | 45.3308 | 0.014209 | 0.34536 | 0.004907 | 24.3054 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.3219 | 3.1251 | 0.008287 | -0.6254 | -0.5372 | -0.7097 | -0.7217
2012-10-30 15:08 | 24.075 | 19.6475 | 19.824 | 19.5672 | 19.6234 | 45.3326 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.3891 | 3.1252 | 0.008285 | -0.7416 | -0.5651 | -0.8219 | -0.7657
2012-10-30 15:08 | 24.0758 | 19.6924 | 19.8128 | 19.6282 | 19.6162 | 45.3349 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.3419 | 3.1254 | 0.00828 | -0.6495 | -0.5291 | -0.7137 | -0.7257
2012-10-30 15:08 | 24.0705 | 19.6831 | 19.7874 | 19.639 | 19.5828 | 45.3401 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014389 | 0.23445 | 0.003374 | 16.2937 | 20.3406 | 3.1257 | 0.008282 | -0.6575 | -0.5532 | -0.7016 | -0.7578
2012-10-30 15:08 | 24.0765 | 19.7092 | 19.7974 | 19.5808 | 19.6289 | 45.3383 | 0.014194 | 0.34541 | 0.004903 | 24.3351 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.3906 | 3.1256 | 0.008283 | -0.6814 | -0.5932 | -0.8098 | -0.7617
2012-10-30 15:08 | 24.0832 | 19.6798 | 19.8202 | 19.6115 | 19.6236 | 45.3314 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.3493 | 3.1251 | 0.008285 | -0.6695 | -0.5291 | -0.7378 | -0.7257
2012-10-30 15:08 | 24.0832 | 19.6958 | 19.8161 | 19.6276 | 19.6115 | 45.3297 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.3973 | 3.125 | 0.008291 | -0.7015 | -0.5812 | -0.7697 | -0.7858
2012-10-30 15:08 | 24.0862 | 19.6988 | 19.8352 | 19.6266 | 19.6186 | 45.3297 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.4204 | 3.125 | 0.008284 | -0.7216 | -0.5852 | -0.7938 | -0.8018
2012-10-30 15:08 | 24.0832 | 19.7078 | 19.7921 | 19.5915 | 19.6276 | 45.3285 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.3893 | 3.1249 | 0.008284 | -0.6815 | -0.5972 | -0.7978 | -0.7617
2012-10-30 15:08 | 24.093 | 19.6855 | 19.7979 | 19.6093 | 19.6173 | 45.3331 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.3911 | 3.1253 | 0.008283 | -0.7056 | -0.5932 | -0.7818 | -0.7738
2012-10-30 15:08 | 24.0832 | 19.6677 | 19.8242 | 19.6316 | 19.6115 | 45.3366 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4094 | 3.1255 | 0.008284 | -0.7417 | -0.5852 | -0.7778 | -0.7979
2012-10-30 15:08 | 24.0855 | 19.662 | 19.8104 | 19.5777 | 19.6058 | 45.3372 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4197 | 3.1255 | 0.008285 | -0.7577 | -0.6093 | -0.842 | -0.8139
2012-10-30 15:08 | 24.0937 | 19.7023 | 19.8266 | 19.622 | 19.618 | 45.3222 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014417 | 0.23449 | 0.003381 | 16.2651 | 20.3798 | 3.1245 | 0.008287 | -0.6775 | -0.5532 | -0.7578 | -0.7618
2012-10-30 15:08 | 24.0967 | 19.6692 | 19.7935 | 19.609 | 19.609 | 45.3256 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.3708 | 3.1247 | 0.008284 | -0.7016 | -0.5773 | -0.7618 | -0.7618
2012-10-30 15:08 | 24.0922 | 19.6807 | 19.8412 | 19.5925 | 19.6045 | 45.3378 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4144 | 3.1256 | 0.008283 | -0.7337 | -0.5732 | -0.8219 | -0.8099
2012-10-30 15:08 | 24.1012 | 19.6897 | 19.8101 | 19.6215 | 19.6015 | 45.3251 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.3993 | 3.1247 | 0.008284 | -0.7096 | -0.5892 | -0.7778 | -0.7978
2012-10-30 15:08 | 24.0997 | 19.6842 | 19.8166 | 19.608 | 19.62 | 45.3262 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.3858 | 3.1248 | 0.008285 | -0.7016 | -0.5692 | -0.7778 | -0.7658
2012-10-30 15:08 | 24.1019 | 19.6985 | 19.8148 | 19.6102 | 19.6062 | 45.3424 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.42 | 3.1259 | 0.00829 | -0.7215 | -0.6052 | -0.8098 | -0.8138
2012-10-30 15:09 | 24.1072 | 19.6436 | 19.792 | 19.6034 | 19.6115 | 45.3256 | 0.014194 | 0.34538 | 0.004902 | 24.333 | 0.014395 | 0.23458 | 0.003377 | 16.2959 | 20.4013 | 3.1247 | 0.008279 | -0.7577 | -0.6093 | -0.7979 | -0.7898
2012-10-30 15:09 | 24.1049 | 19.6613 | 19.8098 | 19.6052 | 19.6212 | 45.3372 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.4311 | 3.1255 | 0.008286 | -0.7698 | -0.6213 | -0.8259 | -0.8099
2012-10-30 15:09 | 24.1057 | 19.7023 | 19.8106 | 19.622 | 19.6421 | 45.3308 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.4479 | 3.1251 | 0.008285 | -0.7456 | -0.6373 | -0.8259 | -0.8058
2012-10-30 15:09 | 24.1034 | 19.6919 | 19.7802 | 19.6277 | 19.6358 | 45.3279 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4336 | 3.1249 | 0.008289 | -0.7417 | -0.6534 | -0.8059 | -0.7978
2012-10-30 15:09 | 24.1169 | 19.6934 | 19.8258 | 19.6172 | 19.6131 | 45.332 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.407 | 3.1252 | 0.008286 | -0.7136 | -0.5812 | -0.7898 | -0.7939
2012-10-30 15:09 | 24.1221 | 19.6384 | 19.823 | 19.6224 | 19.6063 | 45.3366 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4443 | 3.1255 | 0.008288 | -0.8059 | -0.6213 | -0.8219 | -0.838
2012-10-30 15:09 | 24.1274 | 19.6838 | 19.8082 | 19.6598 | 19.6477 | 45.3204 | 0.014212 | 0.34538 | 0.004909 | 24.3022 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4295 | 3.1244 | 0.008287 | -0.7457 | -0.6213 | -0.7697 | -0.7818
2012-10-30 15:09 | 24.1229 | 19.6713 | 19.7997 | 19.6071 | 19.6352 | 45.3308 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.425 | 3.1251 | 0.008284 | -0.7537 | -0.6253 | -0.8179 | -0.7898
2012-10-30 15:09 | 24.1281 | 19.6926 | 19.849 | 19.6003 | 19.6324 | 45.3331 | 0.014212 | 0.34538 | 0.004909 | 24.3022 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.4022 | 3.1253 | 0.008289 | -0.7096 | -0.5532 | -0.8019 | -0.7698
2012-10-30 15:09 | 24.1333 | 19.6817 | 19.8342 | 19.6376 | 19.6456 | 45.3222 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014423 | 0.23458 | 0.003383 | 16.2643 | 20.4154 | 3.1245 | 0.008292 | -0.7337 | -0.5812 | -0.7778 | -0.7698
2012-10-30 15:09 | 24.1333 | 19.6416 | 19.8141 | 19.6456 | 19.6536 | 45.3331 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.3874 | 3.1253 | 0.008284 | -0.7458 | -0.5733 | -0.7418 | -0.7338
2012-10-30 15:09 | 24.1356 | 19.7081 | 19.8284 | 19.6439 | 19.6399 | 45.3314 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.3536 | 3.1251 | 0.008286 | -0.6455 | -0.5252 | -0.7097 | -0.7137
2012-10-30 15:09 | 24.1326 | 19.6609 | 19.8254 | 19.6369 | 19.6248 | 45.3181 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014436 | 0.23449 | 0.003385 | 16.2437 | 20.3746 | 3.1242 | 0.008291 | -0.7137 | -0.5492 | -0.7377 | -0.7498
2012-10-30 15:09 | 24.1318 | 19.6642 | 19.8367 | 19.608 | 19.62 | 45.3227 | 0.014225 | 0.3454 | 0.004913 | 24.2811 | 0.01442 | 0.23448 | 0.003381 | 16.2607 | 20.3738 | 3.1245 | 0.008294 | -0.7096 | -0.5371 | -0.7658 | -0.7538
2012-10-30 15:09 | 24.1333 | 19.6737 | 19.786 | 19.5774 | 19.6456 | 45.3412 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4274 | 3.1258 | 0.008286 | -0.7537 | -0.6414 | -0.85 | -0.7818
2012-10-30 15:09 | 24.1318 | 19.6922 | 19.8407 | 19.5959 | 19.624 | 45.3181 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.4179 | 3.1242 | 0.008289 | -0.7257 | -0.5772 | -0.822 | -0.7939
2012-10-30 15:09 | 24.1363 | 19.6606 | 19.785 | 19.6165 | 19.5964 | 45.3401 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 20.3783 | 3.1257 | 0.008285 | -0.7177 | -0.5933 | -0.7618 | -0.7819
2012-10-30 15:09 | 24.1423 | 19.7108 | 19.8552 | 19.6305 | 19.6185 | 45.3297 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.4004 | 3.125 | 0.008285 | -0.6896 | -0.5452 | -0.7699 | -0.7819
2012-10-30 15:10 | 24.1333 | 19.6857 | 19.8221 | 19.6014 | 19.6095 | 45.3152 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.4114 | 3.124 | 0.008281 | -0.7257 | -0.5893 | -0.81 | -0.8019
2012-10-30 15:10 | 24.1386 | 19.675 | 19.8074 | 19.5827 | 19.6148 | 45.3199 | 0.014206 | 0.34537 | 0.004906 | 24.3115 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4888 | 3.1244 | 0.008283 | -0.8138 | -0.6814 | -0.9061 | -0.874
2012-10-30 15:10 | 24.1468 | 19.6912 | 19.8477 | 19.643 | 19.5828 | 45.3331 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.489 | 3.1253 | 0.008283 | -0.7978 | -0.6413 | -0.846 | -0.9062
2012-10-30 15:10 | 24.1468 | 19.6872 | 19.8156 | 19.5668 | 19.6711 | 45.3152 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4489 | 3.124 | 0.008284 | -0.7617 | -0.6333 | -0.8821 | -0.7778
2012-10-30 15:10 | 24.1423 | 19.6586 | 19.8151 | 19.5984 | 19.6265 | 45.3227 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.4484 | 3.1245 | 0.008289 | -0.7898 | -0.6333 | -0.85 | -0.8219
2012-10-30 15:10 | 24.1483 | 19.6646 | 19.8492 | 19.5763 | 19.6325 | 45.336 | 0.014187 | 0.34546 | 0.004901 | 24.3502 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4184 | 3.1255 | 0.00828 | -0.7538 | -0.5692 | -0.8421 | -0.7859
2012-10-30 15:10 | 24.1543 | 19.6626 | 19.811 | 19.6626 | 19.6184 | 45.3245 | 0.014219 | 0.3454 | 0.004911 | 24.2913 | 0.014395 | 0.23459 | 0.003377 | 16.297 | 20.4404 | 3.1247 | 0.008288 | -0.7778 | -0.6294 | -0.7778 | -0.822
2012-10-30 15:10 | 24.1595 | 19.736 | 19.8243 | 19.6276 | 19.6397 | 45.3268 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4176 | 3.1248 | 0.008282 | -0.6816 | -0.5933 | -0.79 | -0.7779
2012-10-30 15:10 | 24.155 | 19.6593 | 19.8117 | 19.6151 | 19.6191 | 45.3487 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.4171 | 3.1263 | 0.008281 | -0.7578 | -0.6054 | -0.802 | -0.798
2012-10-30 15:10 | 24.1573 | 19.6736 | 19.8301 | 19.6214 | 19.6375 | 45.3222 | 0.014194 | 0.34543 | 0.004903 | 24.3361 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.4714 | 3.1245 | 0.008282 | -0.7978 | -0.6413 | -0.85 | -0.8339
2012-10-30 15:10 | 24.155 | 19.6874 | 19.8037 | 19.6312 | 19.6352 | 45.3118 | 0.014212 | 0.34537 | 0.004908 | 24.3012 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.4852 | 3.1238 | 0.008288 | -0.7978 | -0.6815 | -0.854 | -0.85
2012-10-30 15:10 | 24.158 | 19.6904 | 19.8669 | 19.6101 | 19.6181 | 45.3303 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.4922 | 3.1251 | 0.008286 | -0.8018 | -0.6253 | -0.8821 | -0.8741
2012-10-30 15:10 | 24.1535 | 19.6658 | 19.8223 | 19.5735 | 19.6297 | 45.3279 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4837 | 3.1249 | 0.00828 | -0.8179 | -0.6614 | -0.9102 | -0.854
2012-10-30 15:10 | 24.1558 | 19.6801 | 19.8446 | 19.5878 | 19.6601 | 45.3193 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23444 | 0.003378 | 16.2712 | 20.506 | 3.1243 | 0.008286 | -0.8259 | -0.6614 | -0.9182 | -0.8459
2012-10-30 15:10 | 24.1565 | 19.6768 | 19.8293 | 19.6166 | 19.6206 | 45.3303 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.5107 | 3.1251 | 0.008288 | -0.8339 | -0.6814 | -0.8941 | -0.8901
2012-10-30 15:10 | 24.164 | 19.6723 | 19.8448 | 19.6281 | 19.6402 | 45.3314 | 0.014219 | 0.34548 | 0.004912 | 24.2974 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.4621 | 3.1251 | 0.008288 | -0.7898 | -0.6173 | -0.834 | -0.8219
2012-10-30 15:10 | 24.158 | 19.6823 | 19.8629 | 19.6342 | 19.6302 | 45.3222 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014389 | 0.23451 | 0.003374 | 16.2977 | 20.4842 | 3.1245 | 0.00828 | -0.8019 | -0.6213 | -0.85 | -0.854
2012-10-30 15:10 | 24.1558 | 19.64 | 19.8486 | 19.628 | 19.644 | 45.3274 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4339 | 3.1249 | 0.008288 | -0.7939 | -0.5853 | -0.8059 | -0.7899
2012-10-30 15:10 | 24.1618 | 19.7022 | 19.8346 | 19.634 | 19.6661 | 45.3233 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4679 | 3.1246 | 0.008287 | -0.7657 | -0.6333 | -0.8339 | -0.8018
2012-10-30 15:10 | 24.1633 | 19.6796 | 19.8682 | 19.6033 | 19.6395 | 45.3199 | 0.014212 | 0.34538 | 0.004909 | 24.3022 | 0.014417 | 0.23444 | 0.00338 | 16.2611 | 20.4254 | 3.1244 | 0.008289 | -0.7458 | -0.5572 | -0.8221 | -0.7859
2012-10-30 15:10 | 24.1543 | 19.6746 | 19.8431 | 19.6385 | 19.6184 | 45.3112 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4444 | 3.1238 | 0.008286 | -0.7698 | -0.6013 | -0.8059 | -0.826
2012-10-30 15:10 | 24.1543 | 19.6826 | 19.8632 | 19.6224 | 19.6345 | 45.3245 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4244 | 3.1247 | 0.008283 | -0.7418 | -0.5612 | -0.802 | -0.7899
2012-10-30 15:10 | 24.1588 | 19.6831 | 19.8436 | 19.5748 | 19.6229 | 45.3251 | 0.014187 | 0.34541 | 0.0049 | 24.3471 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4169 | 3.1247 | 0.008278 | -0.7338 | -0.5733 | -0.8421 | -0.794
2012-10-30 15:10 | 24.1573 | 19.6937 | 19.8501 | 19.5733 | 19.6616 | 45.3268 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.4835 | 3.1248 | 0.008285 | -0.7898 | -0.6334 | -0.9102 | -0.8219
2012-10-30 15:10 | 24.161 | 19.6813 | 19.906 | 19.601 | 19.6332 | 45.3285 | 0.014225 | 0.34544 | 0.004914 | 24.2841 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.5152 | 3.1249 | 0.008292 | -0.8339 | -0.6092 | -0.9142 | -0.882
2012-10-30 15:10 | 24.1588 | 19.6831 | 19.8877 | 19.5988 | 19.631 | 45.3071 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014417 | 0.23448 | 0.00338 | 16.2641 | 20.485 | 3.1235 | 0.008288 | -0.8019 | -0.5973 | -0.8862 | -0.854
2012-10-30 15:11 | 24.161 | 19.6893 | 19.8578 | 19.6051 | 19.6492 | 45.3314 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.4872 | 3.1251 | 0.008284 | -0.7979 | -0.6294 | -0.8821 | -0.838
2012-10-30 15:11 | 24.1677 | 19.6599 | 19.8726 | 19.6559 | 19.6599 | 45.3204 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.01442 | 0.23458 | 0.003383 | 16.2677 | 20.4618 | 3.1244 | 0.00829 | -0.8019 | -0.5892 | -0.8059 | -0.8019
2012-10-30 15:11 | 24.158 | 19.6783 | 19.8428 | 19.6342 | 19.6342 | 45.3071 | 0.014209 | 0.34536 | 0.004907 | 24.3054 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.4882 | 3.1235 | 0.008285 | -0.8099 | -0.6454 | -0.854 | -0.854
2012-10-30 15:11 | 24.1565 | 19.6728 | 19.8574 | 19.6006 | 19.6728 | 45.3389 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.4546 | 3.1257 | 0.008289 | -0.7818 | -0.5972 | -0.854 | -0.7818
2012-10-30 15:11 | 24.1588 | 19.7152 | 19.8637 | 19.6269 | 19.643 | 45.3106 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.4169 | 3.1237 | 0.008288 | -0.7017 | -0.5532 | -0.79 | -0.7739
2012-10-30 15:11 | 24.1573 | 19.6616 | 19.8582 | 19.6214 | 19.6415 | 45.3285 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.4274 | 3.1249 | 0.008286 | -0.7658 | -0.5692 | -0.806 | -0.7859
2012-10-30 15:11 | 24.1588 | 19.7112 | 19.8396 | 19.6029 | 19.635 | 45.3239 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.4169 | 3.1246 | 0.008286 | -0.7057 | -0.5773 | -0.814 | -0.7819
2012-10-30 15:11 | 24.1655 | 19.6738 | 19.8704 | 19.6457 | 19.6938 | 45.3077 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.3875 | 3.1235 | 0.00828 | -0.7137 | -0.5171 | -0.7418 | -0.6937
2012-10-30 15:11 | 24.161 | 19.6813 | 19.8578 | 19.6332 | 19.6251 | 45.3303 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.403 | 3.1251 | 0.008285 | -0.7217 | -0.5452 | -0.7698 | -0.7779
2012-10-30 15:11 | 24.1662 | 19.6945 | 19.839 | 19.5782 | 19.6945 | 45.3268 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014411 | 0.23444 | 0.003378 | 16.2679 | 20.4283 | 3.1248 | 0.008286 | -0.7338 | -0.5893 | -0.8501 | -0.7338
2012-10-30 15:11 | 24.1692 | 19.6895 | 19.8099 | 19.6173 | 19.6454 | 45.3251 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.4753 | 3.1247 | 0.008286 | -0.7858 | -0.6654 | -0.858 | -0.8299
2012-10-30 15:11 | 24.1535 | 19.6698 | 19.8142 | 19.5895 | 19.6216 | 45.3204 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.4476 | 3.1244 | 0.008281 | -0.7778 | -0.6334 | -0.8581 | -0.826
2012-10-30 15:11 | 24.164 | 19.7124 | 19.8608 | 19.6121 | 19.6281 | 45.3216 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.4821 | 3.1245 | 0.008282 | -0.7697 | -0.6213 | -0.87 | -0.854
2012-10-30 15:11 | 24.1655 | 19.73 | 19.8102 | 19.5855 | 19.6256 | 45.3129 | 0.014203 | 0.34536 | 0.004905 | 24.3156 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.4316 | 3.1239 | 0.008281 | -0.7016 | -0.6214 | -0.8461 | -0.806
2012-10-30 15:11 | 24.1655 | 19.6377 | 19.8423 | 19.6136 | 19.6377 | 45.3106 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014392 | 0.23451 | 0.003375 | 16.2943 | 20.4596 | 3.1237 | 0.008284 | -0.8219 | -0.6173 | -0.846 | -0.8219
2012-10-30 15:11 | 24.1573 | 19.6977 | 19.8501 | 19.6014 | 19.6174 | 45.3366 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.4354 | 3.1255 | 0.008284 | -0.7377 | -0.5853 | -0.834 | -0.818
2012-10-30 15:11 | 24.167 | 19.6913 | 19.8638 | 19.6111 | 19.6552 | 45.3164 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4811 | 3.1241 | 0.008282 | -0.7898 | -0.6173 | -0.87 | -0.8259
2012-10-30 15:11 | 24.1648 | 19.6891 | 19.8456 | 19.6289 | 19.6249 | 45.3239 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4429 | 3.1246 | 0.00829 | -0.7538 | -0.5973 | -0.814 | -0.818
2012-10-30 15:11 | 24.1588 | 19.6831 | 19.8556 | 19.5948 | 19.647 | 45.3262 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4249 | 3.1248 | 0.008282 | -0.7418 | -0.5693 | -0.8301 | -0.7779
2012-10-30 15:11 | 24.164 | 19.6883 | 19.8448 | 19.6362 | 19.6121 | 45.3135 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4421 | 3.1239 | 0.008284 | -0.7538 | -0.5973 | -0.8059 | -0.83
2012-10-30 15:11 | 24.1625 | 19.7029 | 19.8473 | 19.6226 | 19.6668 | 45.3146 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4406 | 3.124 | 0.008288 | -0.7377 | -0.5933 | -0.818 | -0.7738
2012-10-30 15:11 | 24.1633 | 19.6595 | 19.8682 | 19.6234 | 19.6716 | 45.3274 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.4494 | 3.1249 | 0.008286 | -0.7899 | -0.5812 | -0.826 | -0.7778
2012-10-30 15:11 | 24.17 | 19.6743 | 19.8709 | 19.6141 | 19.6582 | 45.3251 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4241 | 3.1247 | 0.008286 | -0.7498 | -0.5532 | -0.81 | -0.7659
2012-10-30 15:11 | 24.176 | 19.6763 | 19.8448 | 19.6482 | 19.6482 | 45.3135 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.422 | 3.1239 | 0.008284 | -0.7457 | -0.5772 | -0.7738 | -0.7738
2012-10-30 15:11 | 24.1692 | 19.7136 | 19.842 | 19.6253 | 19.6373 | 45.336 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.4313 | 3.1255 | 0.00829 | -0.7177 | -0.5893 | -0.806 | -0.794
2012-10-30 15:11 | 24.1662 | 19.6825 | 19.835 | 19.6183 | 19.6424 | 45.3175 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.4243 | 3.1242 | 0.008283 | -0.7418 | -0.5893 | -0.806 | -0.7819
2012-10-30 15:11 | 24.179 | 19.6993 | 19.8598 | 19.6351 | 19.6592 | 45.3331 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.4811 | 3.1253 | 0.00829 | -0.7818 | -0.6213 | -0.846 | -0.8219
2012-10-30 15:11 | 24.1752 | 19.6755 | 19.8199 | 19.6152 | 19.6594 | 45.3285 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.4693 | 3.1249 | 0.008283 | -0.7938 | -0.6494 | -0.8541 | -0.8099
2012-10-30 15:12 | 24.1737 | 19.682 | 19.8746 | 19.6137 | 19.6619 | 45.3314 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4798 | 3.1251 | 0.008286 | -0.7978 | -0.6052 | -0.8661 | -0.8179
2012-10-30 15:12 | 24.173 | 19.6773 | 19.8498 | 19.593 | 19.6492 | 45.3135 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.5032 | 3.1239 | 0.008283 | -0.8259 | -0.6534 | -0.9102 | -0.854
2012-10-30 15:12 | 24.1812 | 19.7015 | 19.854 | 19.6052 | 19.6534 | 45.3297 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.5394 | 3.125 | 0.008286 | -0.8379 | -0.6854 | -0.9342 | -0.886
2012-10-30 15:12 | 24.1775 | 19.6858 | 19.8463 | 19.6135 | 19.6657 | 45.317 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.5918 | 3.1242 | 0.008288 | -0.906 | -0.7455 | -0.9783 | -0.9261
2012-10-30 15:12 | 24.182 | 19.6903 | 19.8187 | 19.5739 | 19.6461 | 45.3274 | 0.014222 | 0.34543 | 0.004913 | 24.2882 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.5242 | 3.1249 | 0.008292 | -0.8339 | -0.7055 | -0.9503 | -0.8781
2012-10-30 15:12 | 24.1835 | 19.6998 | 19.8603 | 19.6356 | 19.6516 | 45.317 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.5057 | 3.1242 | 0.008285 | -0.8059 | -0.6454 | -0.8701 | -0.8541
2012-10-30 15:12 | 24.1827 | 19.687 | 19.8795 | 19.6669 | 19.683 | 45.3146 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.5249 | 3.124 | 0.008287 | -0.8379 | -0.6454 | -0.858 | -0.8419
2012-10-30 15:12 | 24.1857 | 19.7341 | 19.8384 | 19.6538 | 19.6418 | 45.3175 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.5279 | 3.1242 | 0.008286 | -0.7938 | -0.6895 | -0.8741 | -0.8861
2012-10-30 15:12 | 24.1812 | 19.6694 | 19.8419 | 19.6052 | 19.6774 | 45.3123 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.5274 | 3.1238 | 0.008284 | -0.858 | -0.6855 | -0.9222 | -0.85
2012-10-30 15:12 | 24.1797 | 19.6559 | 19.8244 | 19.6438 | 19.6278 | 45.3066 | 0.014222 | 0.34546 | 0.004913 | 24.2903 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.5339 | 3.1234 | 0.008293 | -0.878 | -0.7095 | -0.8901 | -0.9061
2012-10-30 15:12 | 24.1894 | 19.7057 | 19.8622 | 19.6335 | 19.6776 | 45.3199 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.5516 | 3.1244 | 0.008288 | -0.8459 | -0.6894 | -0.9181 | -0.874
2012-10-30 15:12 | 24.1879 | 19.6801 | 19.8206 | 19.632 | 19.6279 | 45.3239 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.5541 | 3.1246 | 0.008282 | -0.874 | -0.7335 | -0.9221 | -0.9262
2012-10-30 15:12 | 24.1917 | 19.716 | 19.8685 | 19.6518 | 19.6719 | 45.3083 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.4858 | 3.1236 | 0.008289 | -0.7698 | -0.6173 | -0.834 | -0.8139
2012-10-30 15:12 | 24.1894 | 19.6616 | 19.8341 | 19.6415 | 19.6535 | 45.3129 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.5276 | 3.1239 | 0.008285 | -0.866 | -0.6935 | -0.8861 | -0.8741
2012-10-30 15:12 | 24.1879 | 19.6681 | 19.8567 | 19.6279 | 19.648 | 45.3251 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.506 | 3.1247 | 0.00829 | -0.8379 | -0.6493 | -0.8781 | -0.858
2012-10-30 15:12 | 24.1872 | 19.6794 | 19.8439 | 19.6433 | 19.6513 | 45.3031 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.4973 | 3.1232 | 0.00828 | -0.8179 | -0.6534 | -0.854 | -0.846
2012-10-30 15:12 | 24.1917 | 19.692 | 19.8645 | 19.6077 | 19.6518 | 45.3129 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4858 | 3.1239 | 0.008288 | -0.7938 | -0.6213 | -0.8781 | -0.834
2012-10-30 15:12 | 24.1827 | 19.687 | 19.8434 | 19.6147 | 19.6268 | 45.3141 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4768 | 3.124 | 0.008283 | -0.7898 | -0.6334 | -0.8621 | -0.85
2012-10-30 15:12 | 24.1909 | 19.6992 | 19.8877 | 19.635 | 19.659 | 45.3106 | 0.014216 | 0.3454 | 0.00491 | 24.2964 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.505 | 3.1237 | 0.008288 | -0.8058 | -0.6173 | -0.87 | -0.846
2012-10-30 15:12 | 24.1954 | 19.6716 | 19.8481 | 19.6033 | 19.6475 | 45.3199 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.01442 | 0.23441 | 0.00338 | 16.2557 | 20.4775 | 3.1244 | 0.008285 | -0.8059 | -0.6294 | -0.8742 | -0.83
2012-10-30 15:12 | 24.1887 | 19.7291 | 19.8575 | 19.6047 | 19.6769 | 45.332 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.4387 | 3.1252 | 0.008286 | -0.7096 | -0.5812 | -0.834 | -0.7618
2012-10-30 15:12 | 24.1939 | 19.7022 | 19.8867 | 19.6219 | 19.6781 | 45.3175 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.5521 | 3.1242 | 0.008287 | -0.8499 | -0.6654 | -0.9302 | -0.874
2012-10-30 15:12 | 24.1969 | 19.7212 | 19.9018 | 19.665 | 19.6771 | 45.3066 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.5391 | 3.1234 | 0.008287 | -0.8179 | -0.6373 | -0.8741 | -0.862
2012-10-30 15:12 | 24.1917 | 19.6839 | 19.8966 | 19.6478 | 19.6879 | 45.3227 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.5259 | 3.1245 | 0.008286 | -0.842 | -0.6293 | -0.8781 | -0.838
2012-10-30 15:12 | 24.1902 | 19.6985 | 19.871 | 19.6664 | 19.6824 | 45.3314 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.5404 | 3.1251 | 0.008286 | -0.8419 | -0.6694 | -0.874 | -0.858
2012-10-30 15:12 | 24.1857 | 19.706 | 19.8545 | 19.6378 | 19.686 | 45.3135 | 0.014194 | 0.3454 | 0.004903 | 24.3341 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.5359 | 3.1239 | 0.008279 | -0.8299 | -0.6814 | -0.8981 | -0.8499
2012-10-30 15:13 | 24.1872 | 19.7236 | 19.876 | 19.6393 | 19.6714 | 45.3123 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.5053 | 3.1238 | 0.008288 | -0.7817 | -0.6293 | -0.866 | -0.8339
2012-10-30 15:13 | 24.1857 | 19.698 | 19.8825 | 19.6177 | 19.698 | 45.3216 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.5439 | 3.1245 | 0.008281 | -0.8459 | -0.6614 | -0.9262 | -0.8459
2012-10-30 15:13 | 24.1812 | 19.6975 | 19.8861 | 19.6534 | 19.6333 | 45.3181 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.5194 | 3.1242 | 0.008285 | -0.8219 | -0.6333 | -0.866 | -0.8861
2012-10-30 15:13 | 24.179 | 19.6752 | 19.8839 | 19.6311 | 19.6913 | 45.3158 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.5732 | 3.1241 | 0.008282 | -0.898 | -0.6893 | -0.9421 | -0.8819
2012-10-30 15:13 | 24.1835 | 19.6918 | 19.8803 | 19.6356 | 19.6757 | 45.3187 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.5337 | 3.1243 | 0.008285 | -0.8419 | -0.6534 | -0.8981 | -0.858
2012-10-30 15:13 | 24.1715 | 19.6958 | 19.8643 | 19.6236 | 19.6718 | 45.3222 | 0.014187 | 0.34544 | 0.004901 | 24.3491 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.5417 | 3.1245 | 0.008278 | -0.8459 | -0.6774 | -0.9181 | -0.8699
2012-10-30 15:13 | 24.1805 | 19.6727 | 19.8894 | 19.6286 | 19.6446 | 45.3175 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.5547 | 3.1242 | 0.008288 | -0.882 | -0.6653 | -0.9261 | -0.9101
2012-10-30 15:13 | 24.1835 | 19.6878 | 19.8964 | 19.6797 | 19.6717 | 45.3175 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014404 | 0.23444 | 0.003377 | 16.2758 | 20.5057 | 3.1242 | 0.008286 | -0.8179 | -0.6093 | -0.826 | -0.834
2012-10-30 15:13 | 24.1857 | 19.6619 | 19.8625 | 19.6538 | 19.6739 | 45.3083 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.5319 | 3.1236 | 0.008285 | -0.87 | -0.6694 | -0.8781 | -0.858
2012-10-30 15:13 | 24.1812 | 19.7136 | 19.8861 | 19.6574 | 19.6614 | 45.3112 | 0.0142 | 0.34536 | 0.004904 | 24.3208 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.5154 | 3.1238 | 0.008283 | -0.8018 | -0.6293 | -0.858 | -0.854
2012-10-30 15:13 | 24.1782 | 19.7547 | 19.8991 | 19.6544 | 19.6704 | 45.3303 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014392 | 0.23458 | 0.003376 | 16.2993 | 20.4883 | 3.1251 | 0.008282 | -0.7336 | -0.5892 | -0.8339 | -0.8179
2012-10-30 15:13 | 24.1782 | 19.6785 | 19.8791 | 19.6263 | 19.6865 | 45.3112 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.4963 | 3.1238 | 0.008285 | -0.8178 | -0.6172 | -0.87 | -0.8098
2012-10-30 15:13 | 24.173 | 19.6933 | 19.8859 | 19.6371 | 19.6813 | 45.3158 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014389 | 0.23448 | 0.003374 | 16.2957 | 20.4952 | 3.1241 | 0.00828 | -0.8019 | -0.6093 | -0.8581 | -0.8139
2012-10-30 15:13 | 24.17 | 19.7144 | 19.8668 | 19.6141 | 19.6261 | 45.317 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014417 | 0.23455 | 0.003382 | 16.2691 | 20.4761 | 3.1242 | 0.008292 | -0.7617 | -0.6093 | -0.862 | -0.85
2012-10-30 15:13 | 24.1722 | 19.6925 | 19.869 | 19.6203 | 19.7046 | 45.3239 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4303 | 3.1246 | 0.008288 | -0.7378 | -0.5613 | -0.81 | -0.7257
2012-10-30 15:13 | 24.1662 | 19.6745 | 19.855 | 19.6343 | 19.6825 | 45.299 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.4483 | 3.1229 | 0.008286 | -0.7738 | -0.5933 | -0.814 | -0.7658
2012-10-30 15:13 | 24.1633 | 19.6595 | 19.8601 | 19.6033 | 19.6555 | 45.3008 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.4814 | 3.123 | 0.008281 | -0.8219 | -0.6213 | -0.8781 | -0.8259
2012-10-30 15:13 | 24.1677 | 19.7322 | 19.9167 | 19.6278 | 19.688 | 45.3222 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014414 | 0.23458 | 0.003381 | 16.2745 | 20.5179 | 3.1245 | 0.00829 | -0.7857 | -0.6012 | -0.8901 | -0.8299
2012-10-30 15:13 | 24.1625 | 19.6828 | 19.8634 | 19.6347 | 19.6949 | 45.3227 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.4606 | 3.1245 | 0.008285 | -0.7778 | -0.5972 | -0.8259 | -0.7657
2012-10-30 15:13 | 24.1648 | 19.6931 | 19.8977 | 19.6329 | 19.649 | 45.3146 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.503 | 3.124 | 0.008285 | -0.8099 | -0.6053 | -0.8701 | -0.854
2012-10-30 15:13 | 24.1662 | 19.6865 | 19.8791 | 19.6303 | 19.6785 | 45.3227 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.4403 | 3.1245 | 0.008286 | -0.7538 | -0.5612 | -0.81 | -0.7618
2012-10-30 15:13 | 24.161 | 19.6653 | 19.8298 | 19.6532 | 19.6372 | 45.3129 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4191 | 3.1239 | 0.008282 | -0.7538 | -0.5893 | -0.7659 | -0.7819
2012-10-30 15:13 | 24.161 | 19.6733 | 19.8578 | 19.6532 | 19.6853 | 45.3326 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4591 | 3.1252 | 0.008286 | -0.7858 | -0.6013 | -0.8059 | -0.7738
2012-10-30 15:13 | 24.1677 | 19.7281 | 19.8766 | 19.6358 | 19.6599 | 45.3303 | 0.014209 | 0.34553 | 0.00491 | 24.3175 | 0.014417 | 0.23459 | 0.003382 | 16.2721 | 20.5139 | 3.1251 | 0.008292 | -0.7858 | -0.6373 | -0.8781 | -0.854
2012-10-30 15:13 | 24.1618 | 19.6861 | 19.8787 | 19.5978 | 19.6701 | 45.3141 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4439 | 3.124 | 0.008286 | -0.7578 | -0.5652 | -0.8461 | -0.7738
2012-10-30 15:14 | 24.1648 | 19.661 | 19.8937 | 19.637 | 19.7132 | 45.3025 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014414 | 0.23458 | 0.003381 | 16.2745 | 20.4669 | 3.1232 | 0.008289 | -0.8059 | -0.5732 | -0.8299 | -0.7537
2012-10-30 15:14 | 24.1648 | 19.7092 | 19.8456 | 19.641 | 19.657 | 45.3222 | 0.014222 | 0.34546 | 0.004913 | 24.2903 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.4669 | 3.1245 | 0.00829 | -0.7577 | -0.6213 | -0.8259 | -0.8099
2012-10-30 15:14 | 24.158 | 19.6823 | 19.8669 | 19.6181 | 19.6502 | 45.3152 | 0.014212 | 0.3455 | 0.00491 | 24.3104 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.4281 | 3.124 | 0.008287 | -0.7458 | -0.5612 | -0.81 | -0.7779
2012-10-30 15:14 | 24.1618 | 19.6901 | 19.8867 | 19.6299 | 19.646 | 45.3245 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4279 | 3.1247 | 0.008285 | -0.7378 | -0.5412 | -0.798 | -0.7819
2012-10-30 15:14 | 24.1633 | 19.7237 | 19.8561 | 19.6395 | 19.6555 | 45.3042 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4574 | 3.1233 | 0.008288 | -0.7337 | -0.6013 | -0.8179 | -0.8019
2012-10-30 15:14 | 24.1588 | 19.6711 | 19.8958 | 19.6269 | 19.6591 | 45.3199 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.4489 | 3.1244 | 0.008286 | -0.7778 | -0.5531 | -0.822 | -0.7898
2012-10-30 15:14 | 24.161 | 19.7094 | 19.8538 | 19.6332 | 19.6653 | 45.3407 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.379 | 3.1258 | 0.008284 | -0.6696 | -0.5252 | -0.7458 | -0.7137
2012-10-30 15:14 | 24.1648 | 19.7212 | 19.8616 | 19.641 | 19.657 | 45.3158 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.491 | 3.1241 | 0.008288 | -0.7698 | -0.6294 | -0.85 | -0.834
2012-10-30 15:14 | 24.1573 | 19.7177 | 19.8742 | 19.6335 | 19.6576 | 45.3002 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014417 | 0.23447 | 0.00338 | 16.2631 | 20.4674 | 3.123 | 0.008286 | -0.7497 | -0.5932 | -0.8339 | -0.8098
2012-10-30 15:14 | 24.158 | 19.6743 | 19.8548 | 19.6382 | 19.6462 | 45.3164 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014395 | 0.23454 | 0.003376 | 16.2929 | 20.4561 | 3.1241 | 0.008282 | -0.7818 | -0.6013 | -0.8179 | -0.8099
2012-10-30 15:14 | 24.1633 | 19.6957 | 19.8802 | 19.6154 | 19.6475 | 45.3245 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4814 | 3.1247 | 0.008284 | -0.7857 | -0.6012 | -0.866 | -0.8339
2012-10-30 15:14 | 24.1662 | 19.6785 | 19.8711 | 19.6464 | 19.6665 | 45.317 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4603 | 3.1242 | 0.008285 | -0.7818 | -0.5892 | -0.8139 | -0.7938
2012-10-30 15:14 | 24.1603 | 19.7328 | 19.8612 | 19.6525 | 19.6445 | 45.3181 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.4063 | 3.1242 | 0.008286 | -0.6735 | -0.5451 | -0.7538 | -0.7618
2012-10-30 15:14 | 24.1662 | 19.6905 | 19.851 | 19.6544 | 19.6464 | 45.3071 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.4563 | 3.1235 | 0.008289 | -0.7658 | -0.6053 | -0.8019 | -0.8099
2012-10-30 15:14 | 24.1603 | 19.7047 | 19.8812 | 19.6244 | 19.6284 | 45.321 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4264 | 3.1244 | 0.008288 | -0.7217 | -0.5452 | -0.802 | -0.798
2012-10-30 15:14 | 24.1633 | 19.7077 | 19.8722 | 19.6274 | 19.6957 | 45.3181 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.4133 | 3.1242 | 0.008284 | -0.7056 | -0.5411 | -0.7859 | -0.7176
2012-10-30 15:14 | 24.1685 | 19.6888 | 19.8774 | 19.6045 | 19.6447 | 45.3071 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.3665 | 3.1235 | 0.00829 | -0.6777 | -0.4891 | -0.762 | -0.7218
2012-10-30 15:14 | 24.1633 | 19.6796 | 19.9083 | 19.6114 | 19.6435 | 45.3008 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.3933 | 3.123 | 0.008286 | -0.7137 | -0.485 | -0.7819 | -0.7498
2012-10-30 15:14 | 24.1655 | 19.6938 | 19.8744 | 19.6698 | 19.7179 | 45.3129 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.4276 | 3.1239 | 0.008288 | -0.7338 | -0.5532 | -0.7578 | -0.7097
2012-10-30 15:14 | 24.1677 | 19.672 | 19.8806 | 19.5997 | 19.68 | 45.3187 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.4057 | 3.1243 | 0.008287 | -0.7337 | -0.5251 | -0.806 | -0.7257
2012-10-30 15:14 | 24.1648 | 19.7052 | 19.8657 | 19.641 | 19.6651 | 45.3146 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.4389 | 3.124 | 0.008286 | -0.7337 | -0.5732 | -0.7979 | -0.7738
2012-10-30 15:14 | 24.1633 | 19.6876 | 19.8842 | 19.5913 | 19.6716 | 45.3042 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014392 | 0.23461 | 0.003376 | 16.3013 | 20.4374 | 3.1233 | 0.008281 | -0.7498 | -0.5532 | -0.8461 | -0.7658
2012-10-30 15:14 | 24.1662 | 19.6865 | 19.9192 | 19.6143 | 19.6865 | 45.3054 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.4723 | 3.1234 | 0.008283 | -0.7858 | -0.5531 | -0.858 | -0.7858
2012-10-30 15:14 | 24.164 | 19.7124 | 19.8849 | 19.6683 | 19.6643 | 45.317 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4181 | 3.1242 | 0.008285 | -0.7057 | -0.5332 | -0.7498 | -0.7538
2012-10-30 15:15 | 24.1633 | 19.6957 | 19.8802 | 19.6234 | 19.6716 | 45.3031 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4614 | 3.1232 | 0.008287 | -0.7657 | -0.5812 | -0.838 | -0.7898
2012-10-30 15:15 | 24.1677 | 19.672 | 19.8846 | 19.6358 | 19.668 | 45.3037 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4778 | 3.1232 | 0.008287 | -0.8058 | -0.5932 | -0.842 | -0.8098
2012-10-30 15:15 | 24.1707 | 19.691 | 19.8956 | 19.6429 | 19.699 | 45.3129 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4929 | 3.1239 | 0.008288 | -0.8019 | -0.5973 | -0.85 | -0.7939
2012-10-30 15:15 | 24.1692 | 19.7136 | 19.8701 | 19.6494 | 19.6494 | 45.3135 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.4393 | 3.1239 | 0.00829 | -0.7257 | -0.5692 | -0.7899 | -0.7899
2012-10-30 15:15 | 24.1543 | 19.6746 | 19.8752 | 19.6345 | 19.6385 | 45.3187 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4404 | 3.1243 | 0.008287 | -0.7658 | -0.5652 | -0.8059 | -0.8019
2012-10-30 15:15 | 24.1662 | 19.7026 | 19.8991 | 19.6384 | 19.6263 | 45.3071 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014411 | 0.23461 | 0.003381 | 16.2799 | 20.4363 | 3.1235 | 0.008285 | -0.7337 | -0.5372 | -0.7979 | -0.81
2012-10-30 15:15 | 24.1573 | 19.6576 | 19.8862 | 19.6375 | 19.6696 | 45.3077 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.4634 | 3.1235 | 0.008283 | -0.8058 | -0.5772 | -0.8259 | -0.7938
2012-10-30 15:15 | 24.1573 | 19.7017 | 19.8943 | 19.6214 | 19.7057 | 45.3256 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.4634 | 3.1247 | 0.00829 | -0.7617 | -0.5691 | -0.842 | -0.7577
2012-10-30 15:15 | 24.158 | 19.6863 | 19.903 | 19.6502 | 19.6502 | 45.3106 | 0.014219 | 0.34546 | 0.004912 | 24.2954 | 0.014401 | 0.23462 | 0.003379 | 16.2922 | 20.4241 | 3.1237 | 0.008291 | -0.7378 | -0.5211 | -0.7739 | -0.7739
2012-10-30 15:15 | 24.1692 | 19.7176 | 19.8981 | 19.6614 | 19.6614 | 45.3094 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.4673 | 3.1236 | 0.008286 | -0.7497 | -0.5692 | -0.8059 | -0.8059
2012-10-30 15:15 | 24.1648 | 19.6771 | 19.9138 | 19.5888 | 19.657 | 45.3048 | 0.014212 | 0.34554 | 0.004911 | 24.3134 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.4509 | 3.1233 | 0.008289 | -0.7738 | -0.5371 | -0.8621 | -0.7939
2012-10-30 15:15 | 24.1677 | 19.688 | 19.8886 | 19.6599 | 19.6479 | 45.3025 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4177 | 3.1232 | 0.008288 | -0.7297 | -0.5291 | -0.7578 | -0.7698
2012-10-30 15:15 | 24.158 | 19.6984 | 19.8909 | 19.6181 | 19.6583 | 45.3199 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.4481 | 3.1244 | 0.008285 | -0.7497 | -0.5572 | -0.83 | -0.7898
2012-10-30 15:15 | 24.1648 | 19.7052 | 19.9018 | 19.637 | 19.645 | 45.3164 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4389 | 3.1241 | 0.008285 | -0.7337 | -0.5371 | -0.8019 | -0.7939
2012-10-30 15:15 | 24.1685 | 19.7169 | 19.9295 | 19.6607 | 19.6848 | 45.31 | 0.014219 | 0.3454 | 0.004911 | 24.2913 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.4826 | 3.1237 | 0.008288 | -0.7657 | -0.5531 | -0.8219 | -0.7978
2012-10-30 15:15 | 24.1595 | 19.6919 | 19.8924 | 19.6036 | 19.6919 | 45.3106 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.01443 | 0.23449 | 0.003384 | 16.2505 | 20.4736 | 3.1237 | 0.008288 | -0.7817 | -0.5812 | -0.87 | -0.7817
2012-10-30 15:15 | 24.167 | 19.7154 | 19.9601 | 19.6632 | 19.6793 | 45.3083 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4932 | 3.1236 | 0.008284 | -0.7778 | -0.5331 | -0.83 | -0.8139
2012-10-30 15:15 | 24.158 | 19.7024 | 19.927 | 19.6583 | 19.6904 | 45.3031 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4922 | 3.1232 | 0.008282 | -0.7898 | -0.5652 | -0.8339 | -0.8018
2012-10-30 15:15 | 24.1603 | 19.6927 | 19.9253 | 19.6445 | 19.6806 | 45.3089 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.5105 | 3.1236 | 0.008285 | -0.8178 | -0.5852 | -0.866 | -0.8299
2012-10-30 15:15 | 24.1633 | 19.6636 | 19.9243 | 19.6515 | 19.6997 | 45.317 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.5255 | 3.1242 | 0.008282 | -0.8619 | -0.6012 | -0.874 | -0.8258
2012-10-30 15:15 | 24.1528 | 19.6731 | 19.8938 | 19.6531 | 19.645 | 45.3187 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014408 | 0.23441 | 0.003377 | 16.2692 | 20.4709 | 3.1243 | 0.008286 | -0.7978 | -0.5771 | -0.8178 | -0.8259
2012-10-30 15:15 | 24.164 | 19.6843 | 19.929 | 19.6442 | 19.7004 | 45.3094 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4701 | 3.1236 | 0.008281 | -0.7858 | -0.5411 | -0.8259 | -0.7697
2012-10-30 15:15 | 24.1595 | 19.6919 | 19.9165 | 19.6317 | 19.6999 | 45.31 | 0.014219 | 0.34547 | 0.004912 | 24.2964 | 0.014414 | 0.23448 | 0.00338 | 16.2674 | 20.4616 | 3.1237 | 0.008292 | -0.7697 | -0.5451 | -0.8299 | -0.7617
2012-10-30 15:16 | 24.1595 | 19.6999 | 19.9245 | 19.6557 | 19.6919 | 45.306 | 0.01419 | 0.34543 | 0.004902 | 24.343 | 0.014414 | 0.23445 | 0.003379 | 16.2654 | 20.4937 | 3.1234 | 0.008281 | -0.7938 | -0.5692 | -0.838 | -0.8018
2012-10-30 15:16 | 24.1513 | 19.6957 | 19.9043 | 19.6516 | 19.6837 | 45.3123 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4534 | 3.1238 | 0.008286 | -0.7577 | -0.5491 | -0.8018 | -0.7697
2012-10-30 15:16 | 24.1648 | 19.6851 | 19.9098 | 19.641 | 19.6771 | 45.3106 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.499 | 3.1237 | 0.008287 | -0.8139 | -0.5892 | -0.858 | -0.8219
2012-10-30 15:16 | 24.1625 | 19.6587 | 19.9356 | 19.6467 | 19.6708 | 45.3014 | 0.014219 | 0.34548 | 0.004912 | 24.2974 | 0.014414 | 0.23441 | 0.003379 | 16.2625 | 20.4726 | 3.1231 | 0.008291 | -0.8139 | -0.537 | -0.8259 | -0.8018
2012-10-30 15:16 | 24.1595 | 19.74 | 19.9045 | 19.6638 | 19.6758 | 45.3019 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.4736 | 3.1231 | 0.008289 | -0.7336 | -0.5691 | -0.8098 | -0.7978
2012-10-30 15:16 | 24.17 | 19.7144 | 19.931 | 19.6542 | 19.7064 | 45.3204 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.5282 | 3.1244 | 0.008288 | -0.8138 | -0.5972 | -0.874 | -0.8218
2012-10-30 15:16 | 24.1505 | 19.6869 | 19.8955 | 19.6548 | 19.6668 | 45.3094 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4566 | 3.1236 | 0.008285 | -0.7697 | -0.5611 | -0.8018 | -0.7898
2012-10-30 15:16 | 24.1685 | 19.7129 | 19.9175 | 19.6647 | 19.6928 | 45.3146 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.4666 | 3.124 | 0.008284 | -0.7537 | -0.5491 | -0.8019 | -0.7738
2012-10-30 15:16 | 24.1662 | 19.6624 | 19.9312 | 19.6785 | 19.6825 | 45.3146 | 0.014219 | 0.34541 | 0.004911 | 24.2923 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.4803 | 3.124 | 0.008289 | -0.8179 | -0.5491 | -0.8018 | -0.7978
2012-10-30 15:16 | 24.161 | 19.6773 | 19.922 | 19.6332 | 19.6853 | 45.3066 | 0.0142 | 0.34531 | 0.004903 | 24.3177 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4551 | 3.1234 | 0.008283 | -0.7778 | -0.5331 | -0.8219 | -0.7698
2012-10-30 15:16 | 24.1633 | 19.6957 | 19.8922 | 19.6395 | 19.7077 | 45.3048 | 0.014216 | 0.34534 | 0.004909 | 24.2923 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4935 | 3.1233 | 0.008288 | -0.7978 | -0.6013 | -0.854 | -0.7858
2012-10-30 15:16 | 24.164 | 19.6763 | 19.9411 | 19.6522 | 19.6803 | 45.3152 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4741 | 3.124 | 0.008285 | -0.7978 | -0.533 | -0.8219 | -0.7938
2012-10-30 15:16 | 24.1662 | 19.7387 | 19.9352 | 19.6464 | 19.7106 | 45.3077 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.5164 | 3.1235 | 0.008281 | -0.7777 | -0.5812 | -0.87 | -0.8058
2012-10-30 15:16 | 24.1662 | 19.6825 | 19.9192 | 19.6785 | 19.7226 | 45.2996 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.01442 | 0.23448 | 0.003381 | 16.2607 | 20.5164 | 3.123 | 0.00829 | -0.8339 | -0.5972 | -0.8379 | -0.7938
2012-10-30 15:16 | 24.164 | 19.7204 | 19.901 | 19.6362 | 19.7365 | 45.3164 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4781 | 3.1241 | 0.008286 | -0.7577 | -0.5771 | -0.8419 | -0.7416
2012-10-30 15:16 | 24.1677 | 19.7081 | 19.9047 | 19.6559 | 19.7001 | 45.3089 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4818 | 3.1236 | 0.008287 | -0.7737 | -0.5771 | -0.8259 | -0.7817
2012-10-30 15:16 | 24.164 | 19.7204 | 19.8929 | 19.6281 | 19.6843 | 45.3037 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4821 | 3.1232 | 0.008289 | -0.7617 | -0.5892 | -0.854 | -0.7978
2012-10-30 15:16 | 24.1655 | 19.7019 | 19.9225 | 19.6818 | 19.6698 | 45.3141 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.5117 | 3.124 | 0.008281 | -0.8098 | -0.5892 | -0.8299 | -0.8419
2012-10-30 15:16 | 24.1648 | 19.6972 | 19.9058 | 19.6651 | 19.6891 | 45.3204 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4589 | 3.1244 | 0.00829 | -0.7617 | -0.5531 | -0.7938 | -0.7698
2012-10-30 15:16 | 24.1655 | 19.6898 | 19.9265 | 19.6336 | 19.6979 | 45.3019 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014386 | 0.23449 | 0.003373 | 16.3002 | 20.4716 | 3.1231 | 0.008284 | -0.7818 | -0.5451 | -0.838 | -0.7737
2012-10-30 15:16 | 24.1662 | 19.7146 | 19.9232 | 19.7066 | 19.7106 | 45.3002 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014392 | 0.23447 | 0.003374 | 16.2913 | 20.4523 | 3.123 | 0.008282 | -0.7377 | -0.5291 | -0.7457 | -0.7417
2012-10-30 15:16 | 24.1648 | 19.7052 | 19.9379 | 19.661 | 19.7132 | 45.3141 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014426 | 0.23452 | 0.003383 | 16.257 | 20.4629 | 3.124 | 0.008291 | -0.7577 | -0.525 | -0.8019 | -0.7497
2012-10-30 15:16 | 24.164 | 19.6883 | 19.905 | 19.6482 | 19.7124 | 45.3083 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014395 | 0.23448 | 0.003375 | 16.2889 | 20.4741 | 3.1236 | 0.008283 | -0.7858 | -0.5691 | -0.8259 | -0.7617
2012-10-30 15:16 | 24.1633 | 19.7398 | 19.9283 | 19.6515 | 19.6997 | 45.3042 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.4454 | 3.1233 | 0.008283 | -0.7056 | -0.5171 | -0.7939 | -0.7457
2012-10-30 15:17 | 24.1685 | 19.6728 | 19.9095 | 19.6527 | 19.7249 | 45.306 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4746 | 3.1234 | 0.008291 | -0.8018 | -0.5651 | -0.8219 | -0.7497
2012-10-30 15:17 | 24.1633 | 19.7037 | 19.9524 | 19.6997 | 19.7197 | 45.3089 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.5015 | 3.1236 | 0.008285 | -0.7978 | -0.5491 | -0.8018 | -0.7818
2012-10-30 15:17 | 24.1633 | 19.7278 | 19.9243 | 19.6796 | 19.6515 | 45.3066 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.01442 | 0.23447 | 0.003381 | 16.2597 | 20.5015 | 3.1234 | 0.008289 | -0.7737 | -0.5772 | -0.8219 | -0.85
2012-10-30 15:17 | 24.1677 | 19.7001 | 19.9167 | 19.692 | 19.7121 | 45.3141 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014395 | 0.23447 | 0.003375 | 16.2879 | 20.4899 | 3.124 | 0.008284 | -0.7898 | -0.5732 | -0.7979 | -0.7778
2012-10-30 15:17 | 24.1685 | 19.737 | 19.9175 | 19.6447 | 19.7209 | 45.3152 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014411 | 0.23444 | 0.003378 | 16.2679 | 20.5147 | 3.124 | 0.008285 | -0.7777 | -0.5972 | -0.87 | -0.7938
2012-10-30 15:17 | 24.1685 | 19.7009 | 19.8734 | 19.6808 | 19.7089 | 45.3083 | 0.014206 | 0.34537 | 0.004906 | 24.3115 | 0.014423 | 0.23445 | 0.003381 | 16.2553 | 20.4987 | 3.1236 | 0.008287 | -0.7978 | -0.6253 | -0.8179 | -0.7898
2012-10-30 15:17 | 24.167 | 19.6913 | 19.8839 | 19.6913 | 19.6913 | 45.306 | 0.014197 | 0.34537 | 0.004903 | 24.3269 | 0.014404 | 0.23444 | 0.003377 | 16.2758 | 20.5132 | 3.1234 | 0.00828 | -0.8219 | -0.6293 | -0.8219 | -0.8219
2012-10-30 15:17 | 24.1692 | 19.7377 | 19.9102 | 19.6935 | 19.6895 | 45.3054 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.5274 | 3.1234 | 0.008286 | -0.7897 | -0.6172 | -0.8339 | -0.8379
2012-10-30 15:17 | 24.1595 | 19.6517 | 19.8563 | 19.6437 | 19.7199 | 45.3204 | 0.014194 | 0.34536 | 0.004902 | 24.3311 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.5217 | 3.1244 | 0.008281 | -0.87 | -0.6654 | -0.878 | -0.8018
2012-10-30 15:17 | 24.1662 | 19.6665 | 19.8991 | 19.6464 | 19.7066 | 45.3048 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.5244 | 3.1233 | 0.008286 | -0.8579 | -0.6253 | -0.878 | -0.8178
2012-10-30 15:17 | 24.164 | 19.6522 | 19.8689 | 19.6562 | 19.6723 | 45.2967 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23442 | 0.003378 | 16.2702 | 20.5382 | 3.1228 | 0.008284 | -0.886 | -0.6693 | -0.882 | -0.8659
2012-10-30 15:17 | 24.1737 | 19.702 | 19.9387 | 19.674 | 19.7221 | 45.3308 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 20.5519 | 3.1251 | 0.008287 | -0.8499 | -0.6132 | -0.8779 | -0.8298
2012-10-30 15:17 | 24.1655 | 19.7059 | 19.9345 | 19.6818 | 19.6698 | 45.3025 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.5077 | 3.1232 | 0.008286 | -0.8018 | -0.5732 | -0.8259 | -0.8379
2012-10-30 15:17 | 24.1625 | 19.7029 | 19.9035 | 19.6226 | 19.7189 | 45.2979 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.5047 | 3.1228 | 0.008287 | -0.8018 | -0.6012 | -0.8821 | -0.7858
2012-10-30 15:17 | 24.1655 | 19.738 | 19.8623 | 19.6858 | 19.6979 | 45.3199 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.4796 | 3.1244 | 0.008286 | -0.7416 | -0.6173 | -0.7938 | -0.7817
2012-10-30 15:17 | 24.1685 | 19.6888 | 19.9095 | 19.6607 | 19.6808 | 45.3042 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.5147 | 3.1233 | 0.008282 | -0.8259 | -0.6052 | -0.854 | -0.8339
2012-10-30 15:17 | 24.17 | 19.6943 | 19.9029 | 19.6542 | 19.6903 | 45.3054 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.5202 | 3.1234 | 0.008282 | -0.8259 | -0.6173 | -0.866 | -0.8299
2012-10-30 15:17 | 24.1648 | 19.7453 | 19.8977 | 19.6891 | 19.7052 | 45.3314 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.511 | 3.1251 | 0.008288 | -0.7657 | -0.6133 | -0.8219 | -0.8058
2012-10-30 15:17 | 24.1648 | 19.7132 | 19.8737 | 19.6891 | 19.6811 | 45.3042 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.499 | 3.1233 | 0.008283 | -0.7858 | -0.6253 | -0.8099 | -0.8179
2012-10-30 15:17 | 24.1685 | 19.745 | 19.8854 | 19.6366 | 19.7009 | 45.3158 | 0.01419 | 0.34544 | 0.004902 | 24.344 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.5347 | 3.1241 | 0.008279 | -0.7897 | -0.6493 | -0.8981 | -0.8338
2012-10-30 15:17 | 24.1707 | 19.7311 | 19.9157 | 19.6629 | 19.6589 | 45.3129 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.5329 | 3.1239 | 0.008283 | -0.8018 | -0.6172 | -0.87 | -0.874
2012-10-30 15:17 | 24.1648 | 19.7614 | 19.8817 | 19.6731 | 19.7052 | 45.3164 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.491 | 3.1241 | 0.008285 | -0.7296 | -0.6093 | -0.8179 | -0.7858
2012-10-30 15:17 | 24.164 | 19.7084 | 19.8849 | 19.6442 | 19.6643 | 45.2962 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014395 | 0.23457 | 0.003377 | 16.295 | 20.5262 | 3.1227 | 0.008286 | -0.8178 | -0.6413 | -0.882 | -0.8619
2012-10-30 15:17 | 24.173 | 19.7375 | 19.9059 | 19.6813 | 19.6893 | 45.3118 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.5352 | 3.1238 | 0.008286 | -0.7977 | -0.6293 | -0.8539 | -0.8459
2012-10-30 15:18 | 24.167 | 19.7114 | 19.916 | 19.6713 | 19.7034 | 45.3193 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.5733 | 3.1243 | 0.008283 | -0.8619 | -0.6573 | -0.902 | -0.8699
2012-10-30 15:18 | 24.1692 | 19.6614 | 19.9302 | 19.6654 | 19.6935 | 45.3019 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.5514 | 3.1231 | 0.008284 | -0.89 | -0.6212 | -0.886 | -0.8579
2012-10-30 15:18 | 24.173 | 19.7375 | 19.926 | 19.7013 | 19.7294 | 45.3008 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.5793 | 3.123 | 0.008287 | -0.8418 | -0.6533 | -0.878 | -0.8499
2012-10-30 15:18 | 24.1722 | 19.7326 | 19.9051 | 19.6725 | 19.7888 | 45.3146 | 0.0142 | 0.34537 | 0.004904 | 24.3218 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.5744 | 3.124 | 0.00828 | -0.8418 | -0.6693 | -0.9019 | -0.7856
2012-10-30 15:18 | 24.17 | 19.7465 | 19.9551 | 19.6783 | 19.7144 | 45.3135 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.5402 | 3.1239 | 0.008285 | -0.7937 | -0.5851 | -0.8619 | -0.8258
2012-10-30 15:18 | 24.1655 | 19.7179 | 19.8784 | 19.6658 | 19.738 | 45.2898 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.5517 | 3.1223 | 0.008284 | -0.8338 | -0.6733 | -0.8859 | -0.8137
2012-10-30 15:18 | 24.1655 | 19.7059 | 19.9386 | 19.6778 | 19.6979 | 45.3141 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.5838 | 3.124 | 0.008284 | -0.8779 | -0.6452 | -0.906 | -0.8859
2012-10-30 15:18 | 24.1767 | 19.7532 | 19.9016 | 19.693 | 19.7251 | 45.3141 | 0.01419 | 0.3454 | 0.004901 | 24.3409 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.5669 | 3.124 | 0.008276 | -0.8137 | -0.6653 | -0.8739 | -0.8418
2012-10-30 15:18 | 24.1685 | 19.7049 | 19.9014 | 19.7009 | 19.6527 | 45.2962 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.5147 | 3.1227 | 0.008285 | -0.8098 | -0.6133 | -0.8138 | -0.862
2012-10-30 15:18 | 24.1812 | 19.7256 | 19.9543 | 19.6734 | 19.7055 | 45.3042 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.5554 | 3.1233 | 0.008284 | -0.8298 | -0.6011 | -0.882 | -0.8499
2012-10-30 15:18 | 24.1737 | 19.7221 | 19.9548 | 19.686 | 19.7382 | 45.3071 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.5159 | 3.1235 | 0.008284 | -0.7938 | -0.5611 | -0.8299 | -0.7777
2012-10-30 15:18 | 24.1782 | 19.7427 | 19.9513 | 19.6584 | 19.7266 | 45.2944 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.5004 | 3.1226 | 0.008285 | -0.7577 | -0.5491 | -0.842 | -0.7738
2012-10-30 15:18 | 24.1685 | 19.7169 | 19.9014 | 19.6768 | 19.7049 | 45.3112 | 0.014209 | 0.34536 | 0.004907 | 24.3054 | 0.014392 | 0.23457 | 0.003376 | 16.2984 | 20.4987 | 3.1238 | 0.008283 | -0.7818 | -0.5973 | -0.8219 | -0.7938
2012-10-30 15:18 | 24.1745 | 19.6988 | 19.9395 | 19.6828 | 19.7189 | 45.3037 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.4806 | 3.1232 | 0.008284 | -0.7818 | -0.5411 | -0.7978 | -0.7617
2012-10-30 15:18 | 24.1812 | 19.6935 | 19.9262 | 19.6453 | 19.7136 | 45.2915 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4953 | 3.1224 | 0.008282 | -0.8018 | -0.5691 | -0.85 | -0.7817
2012-10-30 15:18 | 24.1812 | 19.6895 | 19.8941 | 19.6574 | 19.7095 | 45.3083 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014414 | 0.23441 | 0.003379 | 16.2625 | 20.4793 | 3.1236 | 0.008287 | -0.7898 | -0.5852 | -0.8219 | -0.7698
2012-10-30 15:18 | 24.1797 | 19.7321 | 19.8966 | 19.6719 | 19.692 | 45.3054 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4498 | 3.1234 | 0.008282 | -0.7177 | -0.5532 | -0.7779 | -0.7578
2012-10-30 15:18 | 24.1782 | 19.7025 | 19.8991 | 19.6664 | 19.6825 | 45.2985 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.4323 | 3.1229 | 0.008289 | -0.7298 | -0.5332 | -0.7659 | -0.7498
2012-10-30 15:18 | 24.1805 | 19.6647 | 19.9255 | 19.6286 | 19.7329 | 45.3274 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.4506 | 3.1249 | 0.008281 | -0.7859 | -0.5251 | -0.822 | -0.7177
2012-10-30 15:18 | 24.182 | 19.7304 | 19.9029 | 19.6863 | 19.6823 | 45.2985 | 0.014197 | 0.34543 | 0.004904 | 24.331 | 0.014417 | 0.23458 | 0.003382 | 16.2711 | 20.412 | 3.1229 | 0.008286 | -0.6816 | -0.5091 | -0.7257 | -0.7297
2012-10-30 15:18 | 24.1812 | 19.7095 | 19.9141 | 19.6815 | 19.7095 | 45.3002 | 0.014194 | 0.34546 | 0.004903 | 24.3382 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4473 | 3.123 | 0.008284 | -0.7378 | -0.5332 | -0.7658 | -0.7378
2012-10-30 15:18 | 24.1812 | 19.7055 | 19.9141 | 19.6734 | 19.6614 | 45.3037 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.3952 | 3.1232 | 0.008286 | -0.6897 | -0.4811 | -0.7218 | -0.7338
2012-10-30 15:18 | 24.1812 | 19.7296 | 19.9462 | 19.6774 | 19.7055 | 45.2956 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.4032 | 3.1227 | 0.008285 | -0.6736 | -0.457 | -0.7258 | -0.6977
2012-10-30 15:18 | 24.182 | 19.6742 | 19.939 | 19.6421 | 19.7384 | 45.3227 | 0.0142 | 0.3454 | 0.004905 | 24.3238 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4641 | 3.1245 | 0.008282 | -0.7899 | -0.5251 | -0.822 | -0.7257
2012-10-30 15:18 | 24.1827 | 19.7151 | 19.8795 | 19.6549 | 19.687 | 45.2979 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4408 | 3.1228 | 0.008285 | -0.7257 | -0.5613 | -0.7859 | -0.7538
2012-10-30 15:19 | 24.1857 | 19.7261 | 19.9106 | 19.69 | 19.7421 | 45.3008 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4197 | 3.123 | 0.008284 | -0.6936 | -0.5091 | -0.7297 | -0.6776
2012-10-30 15:19 | 24.1812 | 19.7296 | 19.8821 | 19.6373 | 19.6855 | 45.306 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014401 | 0.23444 | 0.003376 | 16.2791 | 20.4152 | 3.1234 | 0.008287 | -0.6856 | -0.5331 | -0.7779 | -0.7297
2012-10-30 15:19 | 24.182 | 19.6903 | 19.931 | 19.6702 | 19.6823 | 45.3129 | 0.014197 | 0.3454 | 0.004904 | 24.3289 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4521 | 3.1239 | 0.008282 | -0.7618 | -0.5211 | -0.7819 | -0.7698
2012-10-30 15:19 | 24.1797 | 19.704 | 19.8966 | 19.704 | 19.708 | 45.3014 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014408 | 0.23442 | 0.003378 | 16.2702 | 20.4378 | 3.1231 | 0.00829 | -0.7338 | -0.5412 | -0.7338 | -0.7298
2012-10-30 15:19 | 24.1835 | 19.7239 | 19.9365 | 19.6958 | 19.6918 | 45.2962 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4736 | 3.1227 | 0.00829 | -0.7497 | -0.5371 | -0.7778 | -0.7818
2012-10-30 15:19 | 24.1782 | 19.7266 | 19.9152 | 19.6423 | 19.6825 | 45.3089 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4563 | 3.1236 | 0.008285 | -0.7297 | -0.5411 | -0.814 | -0.7738
2012-10-30 15:19 | 24.1767 | 19.701 | 19.9257 | 19.6609 | 19.7131 | 45.2886 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.4227 | 3.1222 | 0.008283 | -0.7217 | -0.497 | -0.7618 | -0.7096
2012-10-30 15:19 | 24.1827 | 19.7391 | 19.8876 | 19.6629 | 19.7191 | 45.2938 | 0.0142 | 0.34556 | 0.004907 | 24.335 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.4448 | 3.1226 | 0.008283 | -0.7057 | -0.5572 | -0.7819 | -0.7257
2012-10-30 15:19 | 24.1917 | 19.7361 | 19.9166 | 19.696 | 19.6839 | 45.3141 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.4337 | 3.124 | 0.008292 | -0.6976 | -0.5171 | -0.7377 | -0.7498
2012-10-30 15:19 | 24.1902 | 19.7145 | 19.8951 | 19.7065 | 19.6784 | 45.2921 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014411 | 0.23448 | 0.003379 | 16.2708 | 20.4242 | 3.1224 | 0.008286 | -0.7097 | -0.5291 | -0.7177 | -0.7458
2012-10-30 15:19 | 24.1864 | 19.7308 | 19.8792 | 19.6867 | 19.6786 | 45.291 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014389 | 0.23452 | 0.003375 | 16.2988 | 20.4084 | 3.1224 | 0.008285 | -0.6776 | -0.5292 | -0.7217 | -0.7298
2012-10-30 15:19 | 24.1917 | 19.7642 | 19.8685 | 19.7 | 19.696 | 45.3233 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4137 | 3.1246 | 0.008286 | -0.6495 | -0.5452 | -0.7137 | -0.7177
2012-10-30 15:19 | 24.1805 | 19.6928 | 19.8653 | 19.6888 | 19.6968 | 45.3008 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014398 | 0.23458 | 0.003377 | 16.2925 | 20.4386 | 3.123 | 0.008286 | -0.7458 | -0.5733 | -0.7498 | -0.7418
2012-10-30 15:19 | 24.1857 | 19.7381 | 19.8505 | 19.702 | 19.694 | 45.2979 | 0.014187 | 0.34546 | 0.004901 | 24.3502 | 0.014392 | 0.23452 | 0.003375 | 16.2954 | 20.4518 | 3.1228 | 0.008276 | -0.7137 | -0.6013 | -0.7498 | -0.7578
2012-10-30 15:19 | 24.1827 | 19.7191 | 19.8996 | 19.6227 | 19.687 | 45.3118 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4408 | 3.1238 | 0.008289 | -0.7217 | -0.5412 | -0.8181 | -0.7538
2012-10-30 15:19 | 24.1812 | 19.7256 | 19.8981 | 19.6855 | 19.6855 | 45.317 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014411 | 0.23459 | 0.003381 | 16.2789 | 20.4353 | 3.1242 | 0.008286 | -0.7097 | -0.5372 | -0.7498 | -0.7498
2012-10-30 15:19 | 24.1842 | 19.7487 | 19.8971 | 19.6523 | 19.6925 | 45.2915 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4222 | 3.1224 | 0.008283 | -0.6735 | -0.5251 | -0.7699 | -0.7297
2012-10-30 15:19 | 24.1872 | 19.7717 | 19.864 | 19.6754 | 19.6875 | 45.2915 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4453 | 3.1224 | 0.008288 | -0.6736 | -0.5813 | -0.7699 | -0.7578
2012-10-30 15:19 | 24.1887 | 19.7532 | 19.9136 | 19.689 | 19.6849 | 45.3083 | 0.014206 | 0.34553 | 0.004909 | 24.3227 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4187 | 3.1236 | 0.008287 | -0.6655 | -0.5051 | -0.7297 | -0.7338
2012-10-30 15:19 | 24.1894 | 19.7057 | 19.8622 | 19.6656 | 19.7258 | 45.3066 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4034 | 3.1234 | 0.008285 | -0.6977 | -0.5412 | -0.7378 | -0.6776
2012-10-30 15:19 | 24.1917 | 19.6839 | 19.8885 | 19.6679 | 19.6839 | 45.2921 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.3776 | 3.1224 | 0.008286 | -0.6937 | -0.4891 | -0.7097 | -0.6937
2012-10-30 15:19 | 24.1909 | 19.7594 | 19.8677 | 19.6791 | 19.6631 | 45.3019 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.3889 | 3.1231 | 0.008291 | -0.6295 | -0.5212 | -0.7098 | -0.7258
2012-10-30 15:19 | 24.1887 | 19.7491 | 19.8896 | 19.6167 | 19.697 | 45.2985 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23442 | 0.003377 | 16.2747 | 20.4267 | 3.1229 | 0.008284 | -0.6776 | -0.5371 | -0.81 | -0.7297
2012-10-30 15:19 | 24.1939 | 19.7744 | 19.8546 | 19.6741 | 19.6741 | 45.2898 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.3999 | 3.1223 | 0.008285 | -0.6255 | -0.5453 | -0.7258 | -0.7258
2012-10-30 15:20 | 24.1932 | 19.7256 | 19.882 | 19.6493 | 19.6894 | 45.3054 | 0.014194 | 0.34551 | 0.004904 | 24.3422 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.4312 | 3.1234 | 0.008284 | -0.7056 | -0.5492 | -0.7819 | -0.7418
2012-10-30 15:20 | 24.1939 | 19.7383 | 19.8627 | 19.654 | 19.6821 | 45.3031 | 0.014222 | 0.34548 | 0.004913 | 24.2923 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.3879 | 3.1232 | 0.008289 | -0.6496 | -0.5252 | -0.7339 | -0.7058
2012-10-30 15:20 | 24.1962 | 19.7847 | 19.8449 | 19.6924 | 19.6965 | 45.3014 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4022 | 3.1231 | 0.008285 | -0.6175 | -0.5573 | -0.7098 | -0.7057
2012-10-30 15:20 | 24.2007 | 19.729 | 19.9256 | 19.6729 | 19.6969 | 45.3031 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4347 | 3.1232 | 0.008282 | -0.7057 | -0.5091 | -0.7618 | -0.7378
2012-10-30 15:20 | 24.1992 | 19.7195 | 19.88 | 19.6714 | 19.6954 | 45.3025 | 0.014225 | 0.34544 | 0.004914 | 24.2841 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.3972 | 3.1232 | 0.008292 | -0.6777 | -0.5172 | -0.7258 | -0.7018
2012-10-30 15:20 | 24.2029 | 19.7513 | 19.8957 | 19.6791 | 19.663 | 45.2973 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.4249 | 3.1228 | 0.008286 | -0.6736 | -0.5292 | -0.7458 | -0.7619
2012-10-30 15:20 | 24.2051 | 19.7415 | 19.8618 | 19.6973 | 19.6652 | 45.3019 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.375 | 3.1231 | 0.008286 | -0.6335 | -0.5132 | -0.6777 | -0.7098
2012-10-30 15:20 | 24.2096 | 19.7259 | 19.8904 | 19.6657 | 19.7018 | 45.2979 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.4076 | 3.1228 | 0.008288 | -0.6817 | -0.5172 | -0.7419 | -0.7058
2012-10-30 15:20 | 24.2074 | 19.7598 | 19.8882 | 19.6715 | 19.7036 | 45.31 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4174 | 3.1237 | 0.008286 | -0.6576 | -0.5292 | -0.7459 | -0.7138
2012-10-30 15:20 | 24.2186 | 19.7309 | 19.8874 | 19.6707 | 19.6627 | 45.2979 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.4086 | 3.1228 | 0.008282 | -0.6777 | -0.5212 | -0.7379 | -0.7459
2012-10-30 15:20 | 24.2089 | 19.7011 | 19.8737 | 19.6409 | 19.6449 | 45.2956 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.01442 | 0.23452 | 0.003382 | 16.2637 | 20.4069 | 3.1227 | 0.008293 | -0.7058 | -0.5332 | -0.766 | -0.762
2012-10-30 15:20 | 24.2156 | 19.776 | 19.8683 | 19.6998 | 19.6918 | 45.3025 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014417 | 0.23458 | 0.003382 | 16.2711 | 20.4296 | 3.1232 | 0.008288 | -0.6536 | -0.5613 | -0.7298 | -0.7378
2012-10-30 15:20 | 24.2126 | 19.7129 | 19.8974 | 19.6687 | 19.6928 | 45.2985 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 20.3865 | 3.1229 | 0.008285 | -0.6736 | -0.4891 | -0.7178 | -0.6937
2012-10-30 15:20 | 24.2074 | 19.6996 | 19.8722 | 19.6555 | 19.6515 | 45.3014 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.3693 | 3.1231 | 0.008284 | -0.6697 | -0.4971 | -0.7138 | -0.7178
2012-10-30 15:20 | 24.2193 | 19.7557 | 19.9001 | 19.6794 | 19.7035 | 45.3031 | 0.014203 | 0.34557 | 0.004908 | 24.3308 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.4253 | 3.1232 | 0.008286 | -0.6696 | -0.5252 | -0.7459 | -0.7218
2012-10-30 15:20 | 24.2238 | 19.7521 | 19.9367 | 19.704 | 19.7281 | 45.3031 | 0.014212 | 0.34548 | 0.00491 | 24.3094 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4138 | 3.1232 | 0.008288 | -0.6617 | -0.4771 | -0.7098 | -0.6857
2012-10-30 15:20 | 24.2261 | 19.7986 | 19.8548 | 19.6983 | 19.6661 | 45.3118 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4 | 3.1238 | 0.008282 | -0.6014 | -0.5452 | -0.7017 | -0.7339
2012-10-30 15:20 | 24.2268 | 19.7431 | 19.8996 | 19.6668 | 19.6668 | 45.2915 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.3607 | 3.1224 | 0.008289 | -0.6176 | -0.4611 | -0.6939 | -0.6939
2012-10-30 15:20 | 24.2141 | 19.7505 | 19.8708 | 19.6822 | 19.7184 | 45.2962 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4561 | 3.1227 | 0.008285 | -0.7056 | -0.5853 | -0.7739 | -0.7377
2012-10-30 15:20 | 24.2276 | 19.772 | 19.9565 | 19.6636 | 19.7279 | 45.3123 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.3815 | 3.1238 | 0.008285 | -0.6095 | -0.425 | -0.7179 | -0.6536
2012-10-30 15:20 | 24.2291 | 19.7414 | 19.8537 | 19.6892 | 19.7253 | 45.2944 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.415 | 3.1226 | 0.008288 | -0.6736 | -0.5613 | -0.7258 | -0.6897
2012-10-30 15:20 | 24.2261 | 19.7304 | 19.8307 | 19.6621 | 19.6782 | 45.299 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4241 | 3.1229 | 0.008286 | -0.6937 | -0.5934 | -0.762 | -0.7459
2012-10-30 15:20 | 24.2291 | 19.7655 | 19.8698 | 19.7253 | 19.7013 | 45.2962 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.391 | 3.1227 | 0.008284 | -0.6255 | -0.5212 | -0.6657 | -0.6897
2012-10-30 15:20 | 24.2336 | 19.7619 | 19.8863 | 19.6817 | 19.6736 | 45.2962 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.3955 | 3.1227 | 0.008289 | -0.6336 | -0.5092 | -0.7138 | -0.7219
2012-10-30 15:20 | 24.2276 | 19.788 | 19.8803 | 19.6757 | 19.6998 | 45.2979 | 0.014209 | 0.34553 | 0.00491 | 24.3175 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.4216 | 3.1228 | 0.008286 | -0.6336 | -0.5413 | -0.7459 | -0.7218
2012-10-30 15:21 | 24.2373 | 19.8138 | 19.9261 | 19.7175 | 19.6773 | 45.3094 | 0.014219 | 0.34547 | 0.004912 | 24.2964 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.4152 | 3.1236 | 0.008288 | -0.6014 | -0.4891 | -0.6977 | -0.7379
2012-10-30 15:21 | 24.2321 | 19.7404 | 19.8648 | 19.6441 | 19.6882 | 45.3077 | 0.014219 | 0.34546 | 0.004912 | 24.2954 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.418 | 3.1235 | 0.008292 | -0.6776 | -0.5532 | -0.7739 | -0.7298
2012-10-30 15:21 | 24.2336 | 19.7579 | 19.8743 | 19.6576 | 19.7017 | 45.2915 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.3995 | 3.1224 | 0.008287 | -0.6416 | -0.5252 | -0.7419 | -0.6978
2012-10-30 15:21 | 24.2313 | 19.7797 | 19.888 | 19.6954 | 19.6754 | 45.3123 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.3972 | 3.1238 | 0.008291 | -0.6175 | -0.5092 | -0.7018 | -0.7218
2012-10-30 15:21 | 24.2313 | 19.7596 | 19.9201 | 19.6593 | 19.6954 | 45.2973 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014404 | 0.23457 | 0.003379 | 16.2848 | 20.4533 | 3.1228 | 0.00829 | -0.6937 | -0.5332 | -0.794 | -0.7579
2012-10-30 15:21 | 24.2321 | 19.7324 | 19.9289 | 19.6601 | 19.7002 | 45.2956 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014411 | 0.23447 | 0.003379 | 16.2699 | 20.5102 | 3.1227 | 0.008288 | -0.7778 | -0.5813 | -0.8501 | -0.81
2012-10-30 15:21 | 24.2433 | 19.7074 | 19.9361 | 19.6793 | 19.7155 | 45.2973 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.4773 | 3.1228 | 0.008286 | -0.7699 | -0.5412 | -0.798 | -0.7618
2012-10-30 15:21 | 24.2388 | 19.7631 | 19.8915 | 19.711 | 19.6989 | 45.2892 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.4648 | 3.1222 | 0.008283 | -0.7017 | -0.5733 | -0.7538 | -0.7659
2012-10-30 15:21 | 24.241 | 19.7854 | 19.9218 | 19.677 | 19.7573 | 45.2811 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.491 | 3.1217 | 0.008283 | -0.7056 | -0.5692 | -0.814 | -0.7337
2012-10-30 15:21 | 24.2418 | 19.726 | 19.9186 | 19.6738 | 19.718 | 45.2938 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4999 | 3.1226 | 0.008288 | -0.7739 | -0.5813 | -0.8261 | -0.7819
2012-10-30 15:21 | 24.2358 | 19.7521 | 19.8845 | 19.6558 | 19.708 | 45.2863 | 0.014209 | 0.34553 | 0.00491 | 24.3175 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.4979 | 3.122 | 0.008286 | -0.7458 | -0.6134 | -0.8421 | -0.7899
2012-10-30 15:21 | 24.2388 | 19.719 | 19.9196 | 19.6588 | 19.7511 | 45.299 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014386 | 0.23454 | 0.003374 | 16.3031 | 20.4888 | 3.1229 | 0.008285 | -0.7698 | -0.5692 | -0.83 | -0.7377
2012-10-30 15:21 | 24.2358 | 19.7361 | 19.8725 | 19.6758 | 19.7039 | 45.2985 | 0.014203 | 0.34538 | 0.004905 | 24.3176 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4458 | 3.1229 | 0.008284 | -0.7097 | -0.5733 | -0.77 | -0.7419
2012-10-30 15:21 | 24.241 | 19.7573 | 19.9058 | 19.6891 | 19.7132 | 45.2811 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.5071 | 3.1217 | 0.008285 | -0.7498 | -0.6013 | -0.818 | -0.7939
2012-10-30 15:21 | 24.2463 | 19.7425 | 19.9151 | 19.6663 | 19.7185 | 45.2967 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.4363 | 3.1228 | 0.008286 | -0.6938 | -0.5212 | -0.77 | -0.7178
2012-10-30 15:21 | 24.2455 | 19.7337 | 19.9183 | 19.6855 | 19.7016 | 45.3019 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4875 | 3.1231 | 0.008284 | -0.7538 | -0.5692 | -0.802 | -0.7859
2012-10-30 15:21 | 24.2373 | 19.6934 | 19.894 | 19.6733 | 19.7376 | 45.2823 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014401 | 0.23458 | 0.003378 | 16.2891 | 20.4433 | 3.1218 | 0.008283 | -0.7499 | -0.5493 | -0.77 | -0.7057
2012-10-30 15:21 | 24.2493 | 19.7174 | 19.886 | 19.6813 | 19.7375 | 45.3031 | 0.014225 | 0.34546 | 0.004914 | 24.2851 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4593 | 3.1232 | 0.008293 | -0.7419 | -0.5733 | -0.778 | -0.7218
2012-10-30 15:21 | 24.2373 | 19.7697 | 19.8619 | 19.6773 | 19.6854 | 45.3014 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4553 | 3.1231 | 0.008286 | -0.6856 | -0.5934 | -0.778 | -0.7699
2012-10-30 15:21 | 24.241 | 19.7653 | 19.8777 | 19.6931 | 19.673 | 45.2886 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.3869 | 3.1222 | 0.008283 | -0.6216 | -0.5092 | -0.6938 | -0.7139
2012-10-30 15:21 | 24.2425 | 19.7628 | 19.9233 | 19.6866 | 19.6866 | 45.2933 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.4084 | 3.1225 | 0.008287 | -0.6456 | -0.4851 | -0.7218 | -0.7218
2012-10-30 15:21 | 24.2373 | 19.7335 | 19.878 | 19.6493 | 19.7175 | 45.2886 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.4192 | 3.1222 | 0.008286 | -0.6857 | -0.5412 | -0.7699 | -0.7017
2012-10-30 15:21 | 24.2455 | 19.7096 | 19.9143 | 19.6294 | 19.7217 | 45.2788 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4635 | 3.1215 | 0.008285 | -0.7539 | -0.5492 | -0.8341 | -0.7418
2012-10-30 15:21 | 24.2478 | 19.72 | 19.9005 | 19.6878 | 19.6999 | 45.2921 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4458 | 3.1224 | 0.008284 | -0.7258 | -0.5453 | -0.758 | -0.7459
2012-10-30 15:21 | 24.2478 | 19.724 | 19.9045 | 19.6638 | 19.7079 | 45.2881 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4297 | 3.1222 | 0.008283 | -0.7057 | -0.5252 | -0.7659 | -0.7218
2012-10-30 15:21 | 24.2403 | 19.7406 | 19.897 | 19.6322 | 19.6924 | 45.2927 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4463 | 3.1225 | 0.008289 | -0.7057 | -0.5493 | -0.8141 | -0.7539
2012-10-30 15:21 | 24.2403 | 19.7125 | 19.9251 | 19.6563 | 19.6884 | 45.2892 | 0.01419 | 0.34546 | 0.004902 | 24.345 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.4142 | 3.1222 | 0.008281 | -0.7017 | -0.4891 | -0.7579 | -0.7258
2012-10-30 15:22 | 24.2485 | 19.7367 | 19.9012 | 19.7006 | 19.7126 | 45.3089 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4705 | 3.1236 | 0.008285 | -0.7338 | -0.5693 | -0.7699 | -0.7579
2012-10-30 15:22 | 24.253 | 19.7412 | 19.9017 | 19.665 | 19.7211 | 45.2863 | 0.014222 | 0.34553 | 0.004914 | 24.2953 | 0.014382 | 0.23449 | 0.003372 | 16.3047 | 20.479 | 3.122 | 0.008286 | -0.7378 | -0.5773 | -0.814 | -0.7579
2012-10-30 15:22 | 24.256 | 19.7282 | 19.8967 | 19.676 | 19.688 | 45.2933 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.45 | 3.1225 | 0.008288 | -0.7218 | -0.5533 | -0.774 | -0.762
2012-10-30 15:22 | 24.2575 | 19.7056 | 19.9102 | 19.6935 | 19.6855 | 45.2927 | 0.014197 | 0.34548 | 0.004905 | 24.3351 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4595 | 3.1225 | 0.008286 | -0.7539 | -0.5493 | -0.766 | -0.774
2012-10-30 15:22 | 24.2672 | 19.7474 | 19.9039 | 19.7153 | 19.6711 | 45.3031 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4892 | 3.1232 | 0.008285 | -0.7418 | -0.5853 | -0.7739 | -0.8181
2012-10-30 15:22 | 24.265 | 19.7693 | 19.9217 | 19.681 | 19.705 | 45.295 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.507 | 3.1226 | 0.00828 | -0.7377 | -0.5853 | -0.826 | -0.802
2012-10-30 15:22 | 24.2695 | 19.7537 | 19.8701 | 19.7055 | 19.6975 | 45.2852 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.4755 | 3.122 | 0.008283 | -0.7218 | -0.6054 | -0.77 | -0.778
2012-10-30 15:22 | 24.265 | 19.7331 | 19.9177 | 19.681 | 19.7131 | 45.2927 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.499 | 3.1225 | 0.008283 | -0.7659 | -0.5813 | -0.818 | -0.7859
2012-10-30 15:22 | 24.2695 | 19.7296 | 19.9182 | 19.6413 | 19.7015 | 45.2973 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.01442 | 0.23454 | 0.003382 | 16.2647 | 20.5075 | 3.1228 | 0.008288 | -0.7779 | -0.5893 | -0.8662 | -0.806
2012-10-30 15:22 | 24.268 | 19.7402 | 19.9127 | 19.684 | 19.7281 | 45.2869 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.502 | 3.1221 | 0.008285 | -0.7618 | -0.5893 | -0.818 | -0.7739
2012-10-30 15:22 | 24.2725 | 19.7286 | 19.9052 | 19.6604 | 19.6845 | 45.3066 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.5105 | 3.1234 | 0.008286 | -0.7819 | -0.6053 | -0.8501 | -0.826
2012-10-30 15:22 | 24.2695 | 19.7417 | 19.8901 | 19.6815 | 19.6654 | 45.2944 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.4715 | 3.1226 | 0.008288 | -0.7298 | -0.5814 | -0.79 | -0.8061
2012-10-30 15:22 | 24.2717 | 19.7398 | 19.9084 | 19.6596 | 19.7037 | 45.2852 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014398 | 0.23455 | 0.003377 | 16.2906 | 20.4977 | 3.122 | 0.008285 | -0.7579 | -0.5893 | -0.8381 | -0.794
2012-10-30 15:22 | 24.2687 | 19.7007 | 19.8974 | 19.7208 | 19.7409 | 45.2886 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.4947 | 3.1222 | 0.008281 | -0.794 | -0.5973 | -0.7739 | -0.7538
2012-10-30 15:22 | 24.2792 | 19.7514 | 19.9038 | 19.6711 | 19.7072 | 45.3118 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4611 | 3.1238 | 0.008288 | -0.7097 | -0.5573 | -0.79 | -0.7539
2012-10-30 15:22 | 24.2777 | 19.7298 | 19.9144 | 19.7177 | 19.7177 | 45.2875 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.5077 | 3.1221 | 0.008291 | -0.7779 | -0.5933 | -0.79 | -0.79
2012-10-30 15:22 | 24.2837 | 19.7438 | 19.9204 | 19.6716 | 19.6796 | 45.2846 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.4496 | 3.1219 | 0.008285 | -0.7058 | -0.5292 | -0.778 | -0.77
2012-10-30 15:22 | 24.2837 | 19.7237 | 19.9284 | 19.6957 | 19.6916 | 45.2915 | 0.0142 | 0.34551 | 0.004906 | 24.3319 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.4576 | 3.1224 | 0.008285 | -0.7339 | -0.5292 | -0.7619 | -0.766
2012-10-30 15:22 | 24.2874 | 19.7435 | 19.9241 | 19.6994 | 19.7194 | 45.2863 | 0.014222 | 0.34548 | 0.004913 | 24.2923 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.4533 | 3.122 | 0.008289 | -0.7098 | -0.5292 | -0.7539 | -0.7339
2012-10-30 15:22 | 24.2754 | 19.7114 | 19.9081 | 19.6753 | 19.7074 | 45.2933 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.01442 | 0.23454 | 0.003382 | 16.2647 | 20.4493 | 3.1225 | 0.008289 | -0.7379 | -0.5412 | -0.774 | -0.7419
2012-10-30 15:22 | 24.2837 | 19.7599 | 19.8762 | 19.6756 | 19.6796 | 45.3037 | 0.014203 | 0.34553 | 0.004908 | 24.3278 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4897 | 3.1232 | 0.008288 | -0.7298 | -0.6135 | -0.8141 | -0.8101
2012-10-30 15:22 | 24.2837 | 19.7157 | 19.9083 | 19.6555 | 19.7077 | 45.2852 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4296 | 3.122 | 0.008288 | -0.7139 | -0.5213 | -0.7741 | -0.7219
2012-10-30 15:23 | 24.2814 | 19.7375 | 19.9181 | 19.6653 | 19.7014 | 45.2788 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4553 | 3.1215 | 0.008286 | -0.7178 | -0.5372 | -0.79 | -0.7539
2012-10-30 15:23 | 24.2889 | 19.7651 | 19.9135 | 19.6888 | 19.737 | 45.2898 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.4829 | 3.1223 | 0.008285 | -0.7178 | -0.5694 | -0.7941 | -0.7459
2012-10-30 15:23 | 24.2837 | 19.7318 | 19.9124 | 19.6796 | 19.7197 | 45.2996 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4696 | 3.123 | 0.008289 | -0.7378 | -0.5572 | -0.79 | -0.7499
2012-10-30 15:23 | 24.2882 | 19.7885 | 19.8888 | 19.7082 | 19.7363 | 45.2719 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.5022 | 3.121 | 0.00829 | -0.7137 | -0.6134 | -0.794 | -0.7659
2012-10-30 15:23 | 24.2829 | 19.7229 | 19.9156 | 19.7029 | 19.7029 | 45.2967 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.4929 | 3.1228 | 0.008289 | -0.77 | -0.5773 | -0.79 | -0.79
2012-10-30 15:23 | 24.2784 | 19.7064 | 19.9472 | 19.6582 | 19.7104 | 45.2962 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.4804 | 3.1227 | 0.008289 | -0.774 | -0.5332 | -0.8222 | -0.77
2012-10-30 15:23 | 24.2807 | 19.6967 | 19.9414 | 19.7167 | 19.7248 | 45.2944 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014417 | 0.23452 | 0.003381 | 16.2671 | 20.5267 | 3.1226 | 0.008292 | -0.83 | -0.5853 | -0.81 | -0.8019
2012-10-30 15:23 | 24.2792 | 19.7233 | 19.9359 | 19.6992 | 19.7514 | 45.2927 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23459 | 0.003379 | 16.2868 | 20.4932 | 3.1225 | 0.008287 | -0.7699 | -0.5573 | -0.794 | -0.7418
2012-10-30 15:23 | 24.2814 | 19.7616 | 20.0023 | 19.6612 | 19.7255 | 45.3019 | 0.014216 | 0.34536 | 0.00491 | 24.2934 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.4593 | 3.1231 | 0.008286 | -0.6977 | -0.457 | -0.7981 | -0.7338
2012-10-30 15:23 | 24.2867 | 19.7267 | 19.9154 | 19.6585 | 19.6746 | 45.2765 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014395 | 0.23455 | 0.003376 | 16.294 | 20.5287 | 3.1214 | 0.008283 | -0.802 | -0.6133 | -0.8702 | -0.8541
2012-10-30 15:23 | 24.2852 | 19.7212 | 19.958 | 19.6891 | 19.6972 | 45.2904 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4912 | 3.1223 | 0.008285 | -0.77 | -0.5332 | -0.8021 | -0.794
2012-10-30 15:23 | 24.2882 | 19.7403 | 19.9169 | 19.6921 | 19.7724 | 45.291 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.5062 | 3.1224 | 0.008289 | -0.7659 | -0.5893 | -0.8141 | -0.7338
2012-10-30 15:23 | 24.2867 | 19.7067 | 19.9434 | 19.6826 | 19.6906 | 45.2904 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.5167 | 3.1223 | 0.008286 | -0.81 | -0.5733 | -0.8341 | -0.8261
2012-10-30 15:23 | 24.2837 | 19.7237 | 19.9485 | 19.6435 | 19.7077 | 45.3094 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4937 | 3.1236 | 0.008286 | -0.77 | -0.5452 | -0.8502 | -0.786
2012-10-30 15:23 | 24.2814 | 19.6893 | 19.9101 | 19.6934 | 19.7134 | 45.2811 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4834 | 3.1217 | 0.008282 | -0.7941 | -0.5733 | -0.79 | -0.77
2012-10-30 15:23 | 24.2874 | 19.7154 | 19.9522 | 19.6833 | 19.7194 | 45.2956 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.4974 | 3.1227 | 0.008287 | -0.782 | -0.5452 | -0.8141 | -0.778
2012-10-30 15:23 | 24.2867 | 19.7348 | 19.8993 | 19.6946 | 19.6987 | 45.3042 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.4807 | 3.1233 | 0.008287 | -0.7459 | -0.5814 | -0.7861 | -0.782
2012-10-30 15:23 | 24.2852 | 19.6931 | 19.9459 | 19.7212 | 19.7413 | 45.2869 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014411 | 0.23459 | 0.003381 | 16.2789 | 20.4591 | 3.1221 | 0.008288 | -0.766 | -0.5132 | -0.7379 | -0.7178
2012-10-30 15:23 | 24.2897 | 19.7458 | 19.9384 | 19.7057 | 19.7257 | 45.3014 | 0.014222 | 0.34541 | 0.004912 | 24.2872 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.5157 | 3.1231 | 0.008293 | -0.7699 | -0.5773 | -0.81 | -0.79
2012-10-30 15:23 | 24.2919 | 19.752 | 19.9406 | 19.6878 | 19.7079 | 45.28 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.4778 | 3.1216 | 0.008284 | -0.7258 | -0.5372 | -0.79 | -0.7699
2012-10-30 15:23 | 24.2904 | 19.7345 | 19.9271 | 19.7024 | 19.7345 | 45.2788 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.5164 | 3.1215 | 0.008283 | -0.7819 | -0.5893 | -0.814 | -0.7819
2012-10-30 15:23 | 24.2911 | 19.7392 | 19.9518 | 19.695 | 19.7592 | 45.3048 | 0.014222 | 0.34546 | 0.004913 | 24.2903 | 0.014417 | 0.23457 | 0.003382 | 16.2701 | 20.4931 | 3.1233 | 0.008295 | -0.7539 | -0.5413 | -0.7981 | -0.7339
2012-10-30 15:23 | 24.2911 | 19.7512 | 19.9398 | 19.7071 | 19.7352 | 45.2985 | 0.0142 | 0.34544 | 0.004905 | 24.3268 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4931 | 3.1229 | 0.008282 | -0.7419 | -0.5533 | -0.786 | -0.7579
2012-10-30 15:23 | 24.2867 | 19.7187 | 19.9194 | 19.6786 | 19.7067 | 45.2875 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014395 | 0.23452 | 0.003376 | 16.292 | 20.4967 | 3.1221 | 0.008287 | -0.778 | -0.5773 | -0.8181 | -0.79
2012-10-30 15:23 | 24.2926 | 19.6684 | 19.9533 | 19.7246 | 19.7286 | 45.2927 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4906 | 3.1225 | 0.008287 | -0.8222 | -0.5373 | -0.766 | -0.762
2012-10-30 15:23 | 24.2897 | 19.7217 | 19.9224 | 19.6736 | 19.7097 | 45.2794 | 0.014194 | 0.34544 | 0.004903 | 24.3371 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.5277 | 3.1216 | 0.008282 | -0.806 | -0.6053 | -0.8541 | -0.818
2012-10-30 15:24 | 24.2867 | 19.6906 | 19.9354 | 19.7027 | 19.7508 | 45.2944 | 0.014206 | 0.34553 | 0.004909 | 24.3227 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.5408 | 3.1226 | 0.008289 | -0.8502 | -0.6054 | -0.8381 | -0.79
2012-10-30 15:24 | 24.2852 | 19.7293 | 19.9018 | 19.6972 | 19.7212 | 45.2823 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.5112 | 3.1218 | 0.00829 | -0.7819 | -0.6094 | -0.814 | -0.79
2012-10-30 15:24 | 24.2867 | 19.7308 | 19.9595 | 19.6946 | 19.7267 | 45.2817 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4967 | 3.1217 | 0.008291 | -0.7659 | -0.5372 | -0.8021 | -0.77
2012-10-30 15:24 | 24.2874 | 19.7395 | 19.9321 | 19.7194 | 19.7355 | 45.3048 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014398 | 0.23451 | 0.003376 | 16.2875 | 20.4814 | 3.1233 | 0.008282 | -0.7419 | -0.5493 | -0.762 | -0.7459
2012-10-30 15:24 | 24.2874 | 19.7154 | 19.9522 | 19.6953 | 19.7194 | 45.2875 | 0.014209 | 0.34551 | 0.004909 | 24.3165 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.5054 | 3.1221 | 0.008288 | -0.79 | -0.5532 | -0.8101 | -0.786
2012-10-30 15:24 | 24.2829 | 19.7711 | 19.9316 | 19.727 | 19.727 | 45.2846 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.5009 | 3.1219 | 0.00829 | -0.7298 | -0.5693 | -0.7739 | -0.7739
2012-10-30 15:24 | 24.2889 | 19.6888 | 19.9456 | 19.7169 | 19.7289 | 45.2904 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014395 | 0.23449 | 0.003376 | 16.29 | 20.4708 | 3.1223 | 0.008281 | -0.782 | -0.5252 | -0.7539 | -0.7419
2012-10-30 15:24 | 24.2971 | 19.7612 | 19.9298 | 19.705 | 19.7532 | 45.28 | 0.014225 | 0.34546 | 0.004914 | 24.2851 | 0.01442 | 0.23452 | 0.003382 | 16.2637 | 20.5111 | 3.1216 | 0.008296 | -0.7499 | -0.5813 | -0.8061 | -0.7579
2012-10-30 15:24 | 24.2919 | 19.7319 | 19.9567 | 19.6958 | 19.7039 | 45.291 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.5139 | 3.1224 | 0.008291 | -0.782 | -0.5572 | -0.8181 | -0.81
2012-10-30 15:24 | 24.2934 | 19.7214 | 19.9822 | 19.7054 | 19.7254 | 45.2869 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.4673 | 3.1221 | 0.008282 | -0.7459 | -0.4851 | -0.7619 | -0.7419
2012-10-30 15:24 | 24.2904 | 19.7666 | 19.9672 | 19.6823 | 19.7545 | 45.2771 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014401 | 0.23454 | 0.003378 | 16.2862 | 20.4924 | 3.1214 | 0.008284 | -0.7258 | -0.5252 | -0.8101 | -0.7379
2012-10-30 15:24 | 24.2882 | 19.7082 | 19.9128 | 19.652 | 19.7363 | 45.2921 | 0.0142 | 0.34553 | 0.004906 | 24.333 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 20.4862 | 3.1224 | 0.008284 | -0.778 | -0.5734 | -0.8342 | -0.7499
2012-10-30 15:24 | 24.2859 | 19.7139 | 19.9386 | 19.7179 | 19.7059 | 45.2782 | 0.014219 | 0.34544 | 0.004912 | 24.2943 | 0.014404 | 0.23458 | 0.003379 | 16.2858 | 20.5039 | 3.1215 | 0.008291 | -0.79 | -0.5653 | -0.786 | -0.798
2012-10-30 15:24 | 24.2769 | 19.7169 | 19.9176 | 19.725 | 19.741 | 45.2944 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4308 | 3.1226 | 0.008289 | -0.7139 | -0.5132 | -0.7058 | -0.6898
2012-10-30 15:24 | 24.2867 | 19.7348 | 19.9314 | 19.7227 | 19.7348 | 45.2962 | 0.014197 | 0.34547 | 0.004905 | 24.334 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.5127 | 3.1227 | 0.008284 | -0.7779 | -0.5813 | -0.79 | -0.7779
2012-10-30 15:24 | 24.2859 | 19.734 | 19.9667 | 19.6457 | 19.6898 | 45.2875 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4839 | 3.1221 | 0.008283 | -0.7499 | -0.5172 | -0.8382 | -0.7941
2012-10-30 15:24 | 24.2919 | 19.7641 | 19.9045 | 19.7279 | 19.7641 | 45.2898 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.5059 | 3.1223 | 0.008286 | -0.7418 | -0.6014 | -0.778 | -0.7418
2012-10-30 15:24 | 24.2799 | 19.736 | 19.9447 | 19.6999 | 19.748 | 45.2811 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.4859 | 3.1217 | 0.008287 | -0.7499 | -0.5412 | -0.786 | -0.7379
2012-10-30 15:24 | 24.2897 | 19.7257 | 19.9625 | 19.6615 | 19.7017 | 45.2777 | 0.014209 | 0.34537 | 0.004907 | 24.3064 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.4716 | 3.1214 | 0.008288 | -0.7459 | -0.5091 | -0.8101 | -0.7699
2012-10-30 15:24 | 24.2882 | 19.7563 | 19.965 | 19.7403 | 19.7162 | 45.2881 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014392 | 0.23449 | 0.003375 | 16.2934 | 20.4701 | 3.1222 | 0.00828 | -0.7138 | -0.5051 | -0.7298 | -0.7539
2012-10-30 15:24 | 24.2829 | 19.7069 | 19.9517 | 19.6949 | 19.747 | 45.2834 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4729 | 3.1218 | 0.008289 | -0.766 | -0.5212 | -0.778 | -0.7259
2012-10-30 15:24 | 24.2904 | 19.7024 | 19.9752 | 19.6743 | 19.7545 | 45.2794 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.01442 | 0.23458 | 0.003383 | 16.2677 | 20.5084 | 3.1216 | 0.008289 | -0.806 | -0.5332 | -0.8341 | -0.7539
2012-10-30 15:24 | 24.2941 | 19.7622 | 20.007 | 19.678 | 19.702 | 45.2985 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014417 | 0.23458 | 0.003382 | 16.2711 | 20.436 | 3.1229 | 0.008287 | -0.6738 | -0.429 | -0.758 | -0.734
2012-10-30 15:24 | 24.3024 | 19.7184 | 19.9431 | 19.7184 | 19.7063 | 45.2667 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.01442 | 0.23451 | 0.003382 | 16.2627 | 20.5124 | 3.1207 | 0.008287 | -0.794 | -0.5693 | -0.794 | -0.8061
2012-10-30 15:24 | 24.2979 | 19.7339 | 19.9586 | 19.7139 | 19.742 | 45.2927 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.4638 | 3.1225 | 0.008288 | -0.7299 | -0.5052 | -0.7499 | -0.7218
2012-10-30 15:24 | 24.2994 | 19.7234 | 19.9561 | 19.7154 | 19.6752 | 45.2979 | 0.0142 | 0.34538 | 0.004904 | 24.3227 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.4934 | 3.1228 | 0.008282 | -0.77 | -0.5373 | -0.778 | -0.8182
2012-10-30 15:24 | 24.2986 | 19.7146 | 19.9714 | 19.7025 | 19.7266 | 45.2754 | 0.014212 | 0.34553 | 0.004911 | 24.3124 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4886 | 3.1213 | 0.00829 | -0.774 | -0.5172 | -0.7861 | -0.762
2012-10-30 15:25 | 24.3001 | 19.7281 | 19.9568 | 19.6759 | 19.7161 | 45.2817 | 0.014187 | 0.34547 | 0.004901 | 24.3512 | 0.014433 | 0.23448 | 0.003384 | 16.246 | 20.47 | 3.1217 | 0.008285 | -0.7419 | -0.5132 | -0.7941 | -0.7539
2012-10-30 15:25 | 24.3024 | 19.7826 | 19.9752 | 19.7023 | 19.7585 | 45.284 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4723 | 3.1219 | 0.008288 | -0.6897 | -0.4971 | -0.77 | -0.7138
2012-10-30 15:25 | 24.3016 | 19.7457 | 19.9142 | 19.7055 | 19.7497 | 45.2956 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014401 | 0.23445 | 0.003376 | 16.2801 | 20.4675 | 3.1227 | 0.008283 | -0.7218 | -0.5533 | -0.762 | -0.7178
2012-10-30 15:25 | 24.2986 | 19.7587 | 19.9433 | 19.6784 | 19.7427 | 45.295 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.4845 | 3.1226 | 0.008283 | -0.7258 | -0.5412 | -0.8061 | -0.7418
2012-10-30 15:25 | 24.2986 | 19.7667 | 19.9914 | 19.7186 | 19.7146 | 45.2881 | 0.014203 | 0.34551 | 0.004907 | 24.3268 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4845 | 3.1222 | 0.008286 | -0.7178 | -0.4931 | -0.7659 | -0.7699
2012-10-30 15:25 | 24.2956 | 19.7236 | 19.9443 | 19.7035 | 19.7437 | 45.2846 | 0.014222 | 0.34554 | 0.004914 | 24.2963 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4815 | 3.1219 | 0.008295 | -0.7579 | -0.5372 | -0.778 | -0.7378
2012-10-30 15:25 | 24.3046 | 19.7647 | 19.9653 | 19.7005 | 19.7567 | 45.299 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014401 | 0.23447 | 0.003377 | 16.2812 | 20.4825 | 3.1229 | 0.008285 | -0.7178 | -0.5172 | -0.782 | -0.7258
2012-10-30 15:25 | 24.3061 | 19.7542 | 19.9949 | 19.7301 | 19.71 | 45.2933 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.5041 | 3.1225 | 0.008285 | -0.7499 | -0.5092 | -0.774 | -0.7941
2012-10-30 15:25 | 24.2926 | 19.6925 | 19.9493 | 19.6684 | 19.7407 | 45.2921 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4986 | 3.1224 | 0.008287 | -0.8061 | -0.5493 | -0.8302 | -0.7579
2012-10-30 15:25 | 24.2964 | 19.7043 | 19.9491 | 19.7043 | 19.7204 | 45.3014 | 0.014212 | 0.34544 | 0.004909 | 24.3063 | 0.014395 | 0.23462 | 0.003377 | 16.299 | 20.4743 | 3.1231 | 0.008286 | -0.77 | -0.5252 | -0.77 | -0.7539
2012-10-30 15:25 | 24.2971 | 19.7331 | 19.9699 | 19.7291 | 19.7452 | 45.2823 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4991 | 3.1218 | 0.008287 | -0.766 | -0.5292 | -0.77 | -0.7539
2012-10-30 15:25 | 24.2964 | 19.7364 | 19.9772 | 19.7164 | 19.6923 | 45.2863 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.01442 | 0.23459 | 0.003383 | 16.2687 | 20.5184 | 3.122 | 0.008287 | -0.782 | -0.5412 | -0.802 | -0.8261
2012-10-30 15:25 | 24.2979 | 19.762 | 19.9145 | 19.6938 | 19.746 | 45.2944 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.5279 | 3.1226 | 0.00829 | -0.7659 | -0.6134 | -0.8341 | -0.7819
2012-10-30 15:25 | 24.3054 | 19.7615 | 19.9782 | 19.7374 | 19.6933 | 45.2678 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.4994 | 3.1208 | 0.008289 | -0.7379 | -0.5212 | -0.762 | -0.8061
2012-10-30 15:25 | 24.2971 | 19.7452 | 19.8977 | 19.7291 | 19.7291 | 45.2754 | 0.014212 | 0.34553 | 0.004911 | 24.3124 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.5071 | 3.1213 | 0.00829 | -0.7619 | -0.6094 | -0.778 | -0.778
2012-10-30 15:25 | 24.3046 | 19.7808 | 19.9774 | 19.7085 | 19.7647 | 45.2904 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.5346 | 3.1223 | 0.008286 | -0.7538 | -0.5572 | -0.8261 | -0.7699
2012-10-30 15:25 | 24.3113 | 19.7875 | 19.964 | 19.7353 | 19.7674 | 45.2794 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.5333 | 3.1216 | 0.008288 | -0.7458 | -0.5693 | -0.798 | -0.7659
2012-10-30 15:25 | 24.3024 | 19.7585 | 19.9431 | 19.6983 | 19.7424 | 45.2713 | 0.0142 | 0.34548 | 0.004906 | 24.3299 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.5164 | 3.121 | 0.008286 | -0.7579 | -0.5733 | -0.8181 | -0.774
2012-10-30 15:25 | 24.3024 | 19.7545 | 19.9511 | 19.7424 | 19.7344 | 45.2794 | 0.014216 | 0.34551 | 0.004912 | 24.3045 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.5084 | 3.1216 | 0.008289 | -0.7539 | -0.5573 | -0.766 | -0.774
2012-10-30 15:25 | 24.3031 | 19.7672 | 19.9638 | 19.7151 | 19.7271 | 45.2933 | 0.014212 | 0.34553 | 0.004911 | 24.3124 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.5251 | 3.1225 | 0.008291 | -0.7579 | -0.5613 | -0.81 | -0.798
2012-10-30 15:25 | 24.3016 | 19.7697 | 19.9503 | 19.7095 | 19.7216 | 45.2858 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.5036 | 3.122 | 0.008286 | -0.7339 | -0.5533 | -0.7941 | -0.782
2012-10-30 15:25 | 24.3069 | 19.763 | 19.9436 | 19.6867 | 19.7349 | 45.2881 | 0.014194 | 0.34547 | 0.004904 | 24.3392 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 20.4808 | 3.1222 | 0.008284 | -0.7178 | -0.5372 | -0.7941 | -0.7459
2012-10-30 15:25 | 24.3069 | 19.7148 | 19.9476 | 19.6908 | 19.7429 | 45.2846 | 0.014194 | 0.3455 | 0.004904 | 24.3412 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.5249 | 3.1219 | 0.008284 | -0.8101 | -0.5773 | -0.8341 | -0.782
2012-10-30 15:25 | 24.3076 | 19.7476 | 19.9764 | 19.7316 | 19.7436 | 45.2834 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4895 | 3.1218 | 0.008286 | -0.7419 | -0.5131 | -0.7579 | -0.7459
2012-10-30 15:25 | 24.3091 | 19.7451 | 19.9618 | 19.705 | 19.7451 | 45.2829 | 0.014197 | 0.34553 | 0.004905 | 24.3381 | 0.014414 | 0.23447 | 0.00338 | 16.2665 | 20.471 | 3.1218 | 0.008285 | -0.7259 | -0.5092 | -0.766 | -0.7259
2012-10-30 15:25 | 24.3151 | 19.7712 | 19.9478 | 19.711 | 19.7551 | 45.28 | 0.014219 | 0.34551 | 0.004913 | 24.2994 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.5131 | 3.1216 | 0.008293 | -0.7419 | -0.5653 | -0.8021 | -0.758
2012-10-30 15:26 | 24.3128 | 19.7609 | 19.9374 | 19.7488 | 19.7328 | 45.2852 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.01442 | 0.23457 | 0.003382 | 16.2667 | 20.4987 | 3.122 | 0.008291 | -0.7378 | -0.5613 | -0.7499 | -0.7659
2012-10-30 15:26 | 24.3143 | 19.7784 | 19.947 | 19.7303 | 19.7142 | 45.2754 | 0.014197 | 0.34544 | 0.004904 | 24.332 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.4842 | 3.1213 | 0.008281 | -0.7058 | -0.5372 | -0.7539 | -0.77
2012-10-30 15:26 | 24.3061 | 19.7341 | 19.9267 | 19.71 | 19.7221 | 45.2852 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.464 | 3.122 | 0.008285 | -0.7299 | -0.5373 | -0.754 | -0.7419
2012-10-30 15:26 | 24.3039 | 19.764 | 19.9165 | 19.6998 | 19.7159 | 45.2904 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014411 | 0.23455 | 0.00338 | 16.2759 | 20.4698 | 3.1223 | 0.008288 | -0.7058 | -0.5533 | -0.77 | -0.7539
2012-10-30 15:26 | 24.3128 | 19.7328 | 19.9535 | 19.7167 | 19.7488 | 45.2719 | 0.014203 | 0.34554 | 0.004908 | 24.3288 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.5268 | 3.121 | 0.008286 | -0.794 | -0.5733 | -0.8101 | -0.778
2012-10-30 15:26 | 24.3166 | 19.7486 | 19.9091 | 19.7005 | 19.7245 | 45.2915 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014401 | 0.23457 | 0.003378 | 16.2882 | 20.5106 | 3.1224 | 0.008288 | -0.762 | -0.6015 | -0.8101 | -0.7861
2012-10-30 15:26 | 24.3136 | 19.7938 | 19.9543 | 19.6975 | 19.7095 | 45.2834 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4715 | 3.1218 | 0.008285 | -0.6777 | -0.5172 | -0.774 | -0.762
2012-10-30 15:26 | 24.3143 | 19.7543 | 19.951 | 19.6701 | 19.7263 | 45.2782 | 0.014206 | 0.3455 | 0.004908 | 24.3206 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.5323 | 3.1215 | 0.008285 | -0.778 | -0.5813 | -0.8622 | -0.806
2012-10-30 15:26 | 24.3166 | 19.7566 | 19.9252 | 19.7286 | 19.7205 | 45.2782 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.4985 | 3.1215 | 0.008291 | -0.7419 | -0.5733 | -0.7699 | -0.778
2012-10-30 15:26 | 24.3233 | 19.7513 | 19.964 | 19.7152 | 19.7152 | 45.2979 | 0.014212 | 0.34551 | 0.00491 | 24.3114 | 0.014414 | 0.23459 | 0.003381 | 16.2755 | 20.4612 | 3.1228 | 0.008291 | -0.7099 | -0.4972 | -0.746 | -0.746
2012-10-30 15:26 | 24.3211 | 19.7692 | 19.9297 | 19.6929 | 19.705 | 45.284 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.471 | 3.1219 | 0.008287 | -0.7018 | -0.5413 | -0.7781 | -0.766
2012-10-30 15:26 | 24.3196 | 19.7797 | 19.9041 | 19.6593 | 19.7356 | 45.2696 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014392 | 0.23455 | 0.003376 | 16.2974 | 20.4334 | 3.1209 | 0.008285 | -0.6537 | -0.5293 | -0.7741 | -0.6978
2012-10-30 15:26 | 24.3233 | 19.7513 | 19.952 | 19.7031 | 19.7152 | 45.2915 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.4371 | 3.1224 | 0.008287 | -0.6858 | -0.4851 | -0.734 | -0.7219
2012-10-30 15:26 | 24.3211 | 19.7812 | 19.9538 | 19.7371 | 19.7692 | 45.2771 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014414 | 0.23458 | 0.003381 | 16.2745 | 20.451 | 3.1214 | 0.008289 | -0.6698 | -0.4972 | -0.7139 | -0.6818
2012-10-30 15:26 | 24.3211 | 19.7692 | 19.9498 | 19.6889 | 19.7852 | 45.2782 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.4349 | 3.1215 | 0.008288 | -0.6657 | -0.4851 | -0.746 | -0.6497
2012-10-30 15:26 | 24.3211 | 19.7611 | 19.9417 | 19.713 | 19.7331 | 45.2782 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.475 | 3.1215 | 0.008288 | -0.7139 | -0.5333 | -0.762 | -0.7419
2012-10-30 15:26 | 24.3256 | 19.7777 | 19.9422 | 19.6854 | 19.7656 | 45.2829 | 0.014212 | 0.34546 | 0.00491 | 24.3073 | 0.014401 | 0.23448 | 0.003377 | 16.2821 | 20.4394 | 3.1218 | 0.008287 | -0.6617 | -0.4972 | -0.754 | -0.6738
2012-10-30 15:26 | 24.3256 | 19.8218 | 19.9583 | 19.6934 | 19.7697 | 45.2823 | 0.014222 | 0.34538 | 0.004912 | 24.2851 | 0.014414 | 0.23451 | 0.00338 | 16.2695 | 20.4595 | 3.1218 | 0.008292 | -0.6377 | -0.5012 | -0.7661 | -0.6898
2012-10-30 15:26 | 24.3248 | 19.7608 | 19.9214 | 19.7127 | 19.7448 | 45.28 | 0.014203 | 0.34541 | 0.004906 | 24.3197 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.4186 | 3.1216 | 0.008286 | -0.6578 | -0.4972 | -0.7059 | -0.6738
2012-10-30 15:26 | 24.33 | 19.766 | 19.9627 | 19.7018 | 19.758 | 45.2881 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4679 | 3.1222 | 0.008284 | -0.7019 | -0.5052 | -0.7661 | -0.7099
2012-10-30 15:26 | 24.3248 | 19.7247 | 19.9535 | 19.7287 | 19.7488 | 45.2684 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014411 | 0.23457 | 0.00338 | 16.2769 | 20.4547 | 3.1208 | 0.00829 | -0.73 | -0.5012 | -0.726 | -0.7059
2012-10-30 15:26 | 24.3248 | 19.7849 | 19.9735 | 19.7408 | 19.7167 | 45.2771 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.4907 | 3.1214 | 0.008289 | -0.7058 | -0.5172 | -0.7499 | -0.774
2012-10-30 15:26 | 24.3263 | 19.8025 | 19.967 | 19.7222 | 19.7423 | 45.2938 | 0.014203 | 0.3454 | 0.004906 | 24.3187 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.4401 | 3.1226 | 0.008284 | -0.6376 | -0.4731 | -0.7179 | -0.6978
2012-10-30 15:26 | 24.3226 | 19.7707 | 19.9513 | 19.6824 | 19.7265 | 45.2719 | 0.014206 | 0.34551 | 0.004908 | 24.3216 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.4685 | 3.121 | 0.008289 | -0.6978 | -0.5172 | -0.7861 | -0.742
2012-10-30 15:26 | 24.3345 | 19.8067 | 19.9792 | 19.7264 | 19.7585 | 45.2754 | 0.014197 | 0.34546 | 0.004904 | 24.333 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.5124 | 3.1213 | 0.008283 | -0.7057 | -0.5332 | -0.786 | -0.7539
2012-10-30 15:26 | 24.3278 | 19.7518 | 19.9765 | 19.7157 | 19.7357 | 45.2869 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4617 | 3.1221 | 0.008289 | -0.7099 | -0.4852 | -0.746 | -0.726
2012-10-30 15:27 | 24.3323 | 19.7282 | 19.8887 | 19.7001 | 19.7563 | 45.2684 | 0.014197 | 0.34554 | 0.004906 | 24.3391 | 0.014398 | 0.23448 | 0.003376 | 16.2855 | 20.4662 | 3.1208 | 0.008282 | -0.738 | -0.5775 | -0.7661 | -0.7099
2012-10-30 15:27 | 24.3308 | 19.7628 | 19.9915 | 19.7147 | 19.7668 | 45.2944 | 0.014203 | 0.34547 | 0.004907 | 24.3237 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4927 | 3.1226 | 0.008287 | -0.7299 | -0.5012 | -0.778 | -0.7259
2012-10-30 15:27 | 24.3241 | 19.8003 | 19.9207 | 19.728 | 19.7601 | 45.2846 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4299 | 3.1219 | 0.00829 | -0.6296 | -0.5092 | -0.7019 | -0.6698
2012-10-30 15:27 | 24.3323 | 19.7563 | 19.9529 | 19.7322 | 19.7844 | 45.269 | 0.014222 | 0.34548 | 0.004913 | 24.2923 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.4862 | 3.1208 | 0.008292 | -0.7299 | -0.5333 | -0.754 | -0.7018
2012-10-30 15:27 | 24.3233 | 19.7633 | 19.9479 | 19.7031 | 19.7593 | 45.2794 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4491 | 3.1216 | 0.008287 | -0.6858 | -0.5012 | -0.746 | -0.6898
2012-10-30 15:27 | 24.3233 | 19.7674 | 19.9399 | 19.671 | 19.7232 | 45.291 | 0.014203 | 0.3455 | 0.004907 | 24.3258 | 0.014401 | 0.23452 | 0.003377 | 16.2852 | 20.4932 | 3.1224 | 0.008284 | -0.7258 | -0.5533 | -0.8222 | -0.77
2012-10-30 15:27 | 24.3308 | 19.7508 | 19.9996 | 19.6625 | 19.7708 | 45.273 | 0.0142 | 0.3455 | 0.004906 | 24.3309 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4566 | 3.1211 | 0.008284 | -0.7058 | -0.457 | -0.7941 | -0.6858
2012-10-30 15:27 | 24.3293 | 19.7693 | 19.9459 | 19.6931 | 19.7573 | 45.2713 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.5152 | 3.121 | 0.008287 | -0.7459 | -0.5693 | -0.8221 | -0.7579
2012-10-30 15:27 | 24.3323 | 19.7523 | 20.0011 | 19.7242 | 19.7443 | 45.2817 | 0.014194 | 0.34548 | 0.004904 | 24.3402 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.5223 | 3.1217 | 0.008283 | -0.77 | -0.5212 | -0.7981 | -0.778
2012-10-30 15:27 | 24.3248 | 19.7327 | 19.9735 | 19.6846 | 19.6926 | 45.2834 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.5348 | 3.1218 | 0.008287 | -0.8021 | -0.5613 | -0.8502 | -0.8422
2012-10-30 15:27 | 24.3345 | 19.7585 | 19.9431 | 19.7585 | 19.7224 | 45.2748 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014414 | 0.23449 | 0.00338 | 16.2685 | 20.5325 | 3.1212 | 0.008291 | -0.774 | -0.5894 | -0.774 | -0.8101
2012-10-30 15:27 | 24.33 | 19.7781 | 19.9747 | 19.6898 | 19.77 | 45.2996 | 0.014216 | 0.34556 | 0.004912 | 24.3076 | 0.014408 | 0.23461 | 0.00338 | 16.2832 | 20.54 | 3.123 | 0.008292 | -0.7619 | -0.5653 | -0.8502 | -0.77
2012-10-30 15:27 | 24.33 | 19.754 | 19.9787 | 19.7058 | 19.7219 | 45.2615 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014408 | 0.23445 | 0.003378 | 16.2722 | 20.528 | 3.1203 | 0.008289 | -0.774 | -0.5493 | -0.8222 | -0.8061
2012-10-30 15:27 | 24.3233 | 19.7633 | 19.956 | 19.7393 | 19.7353 | 45.273 | 0.014216 | 0.34546 | 0.004911 | 24.3005 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.5213 | 3.1211 | 0.008289 | -0.758 | -0.5653 | -0.782 | -0.786
2012-10-30 15:27 | 24.3233 | 19.7513 | 19.984 | 19.7152 | 19.7312 | 45.2858 | 0.014209 | 0.34548 | 0.004909 | 24.3145 | 0.014392 | 0.23459 | 0.003376 | 16.3004 | 20.5373 | 3.122 | 0.008285 | -0.786 | -0.5533 | -0.8221 | -0.8061
2012-10-30 15:27 | 24.3256 | 19.8058 | 19.9743 | 19.7054 | 19.8098 | 45.2846 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014417 | 0.23454 | 0.003381 | 16.2681 | 20.5196 | 3.1219 | 0.008286 | -0.7138 | -0.5453 | -0.8142 | -0.7098
2012-10-30 15:27 | 24.3345 | 19.7906 | 19.9752 | 19.7264 | 19.7946 | 45.2597 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.014398 | 0.23457 | 0.003377 | 16.2916 | 20.5485 | 3.1202 | 0.008288 | -0.7579 | -0.5733 | -0.8221 | -0.7539
2012-10-30 15:27 | 24.3293 | 19.7613 | 19.966 | 19.7413 | 19.7453 | 45.2777 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014404 | 0.23454 | 0.003378 | 16.2828 | 20.5353 | 3.1214 | 0.008287 | -0.774 | -0.5693 | -0.794 | -0.79
2012-10-30 15:27 | 24.3263 | 19.7623 | 19.967 | 19.678 | 19.7342 | 45.2794 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.5483 | 3.1216 | 0.008288 | -0.786 | -0.5813 | -0.8703 | -0.8141
2012-10-30 15:27 | 24.3241 | 19.7682 | 19.9808 | 19.708 | 19.712 | 45.2869 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.5421 | 3.1221 | 0.008285 | -0.7739 | -0.5613 | -0.8341 | -0.8301
2012-10-30 15:27 | 24.3293 | 19.7533 | 19.978 | 19.7051 | 19.7091 | 45.2742 | 0.014197 | 0.34551 | 0.004905 | 24.337 | 0.014398 | 0.23454 | 0.003377 | 16.2896 | 20.5233 | 3.1212 | 0.008282 | -0.77 | -0.5453 | -0.8182 | -0.8142
2012-10-30 15:27 | 24.3285 | 19.7806 | 20.0053 | 19.7003 | 19.7083 | 45.2869 | 0.014206 | 0.34553 | 0.004909 | 24.3227 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.5064 | 3.1221 | 0.008287 | -0.7258 | -0.5011 | -0.8061 | -0.7981
2012-10-30 15:27 | 24.3278 | 19.7598 | 20.0006 | 19.7438 | 19.7518 | 45.2777 | 0.014222 | 0.34537 | 0.004912 | 24.2841 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.5618 | 3.1214 | 0.008291 | -0.802 | -0.5612 | -0.818 | -0.81
2012-10-30 15:27 | 24.3278 | 19.7598 | 19.9885 | 19.7438 | 19.7317 | 45.2638 | 0.014206 | 0.34547 | 0.004908 | 24.3186 | 0.014404 | 0.23451 | 0.003378 | 16.2808 | 20.5538 | 3.1205 | 0.008286 | -0.794 | -0.5653 | -0.81 | -0.8221
2012-10-30 15:27 | 24.3256 | 19.7335 | 19.9984 | 19.7335 | 19.7536 | 45.2806 | 0.014197 | 0.34541 | 0.004904 | 24.3299 | 0.014401 | 0.23451 | 0.003377 | 16.2841 | 20.5396 | 3.1216 | 0.008281 | -0.8061 | -0.5412 | -0.8061 | -0.786
2012-10-30 15:27 | 24.3308 | 19.7949 | 19.9554 | 19.7347 | 19.7628 | 45.2858 | 0.014212 | 0.3454 | 0.004909 | 24.3033 | 0.014404 | 0.23455 | 0.003378 | 16.2838 | 20.5969 | 3.122 | 0.008287 | -0.802 | -0.6415 | -0.8622 | -0.8341
2012-10-30 15:28 | 24.3211 | 19.7852 | 19.9939 | 19.705 | 19.725 | 45.2748 | 0.014209 | 0.3455 | 0.004909 | 24.3155 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.5311 | 3.1212 | 0.00829 | -0.7459 | -0.5372 | -0.8261 | -0.8061
2012-10-30 15:28 | 24.3256 | 19.7416 | 19.9543 | 19.6894 | 19.8018 | 45.2707 | 0.014206 | 0.34538 | 0.004907 | 24.3125 | 0.014417 | 0.23444 | 0.00338 | 16.2611 | 20.5676 | 3.121 | 0.008287 | -0.826 | -0.6133 | -0.8782 | -0.7658
2012-10-30 15:28 | 24.3248 | 19.7568 | 19.9454 | 19.6886 | 19.7167 | 45.2748 | 0.014209 | 0.34538 | 0.004908 | 24.3073 | 0.014408 | 0.23455 | 0.003379 | 16.2793 | 20.5148 | 3.1212 | 0.008287 | -0.758 | -0.5694 | -0.8262 | -0.7981
2012-10-30 15:28 | 24.3196 | 19.7637 | 19.9362 | 19.7195 | 19.7677 | 45.2713 | 0.014216 | 0.34543 | 0.004911 | 24.2985 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.5136 | 3.121 | 0.008292 | -0.7499 | -0.5774 | -0.7941 | -0.7459
2012-10-30 15:28 | 24.3218 | 19.7217 | 19.9545 | 19.7137 | 19.7819 | 45.2725 | 0.014209 | 0.3454 | 0.004908 | 24.3084 | 0.014408 | 0.23447 | 0.003378 | 16.2733 | 20.4356 | 3.1211 | 0.008286 | -0.7139 | -0.4811 | -0.7219 | -0.6537
2012-10-30 15:28 | 24.3218 | 19.7699 | 20.0066 | 19.7498 | 19.7177 | 45.2782 | 0.014203 | 0.34537 | 0.004905 | 24.3166 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4316 | 3.1215 | 0.008285 | -0.6617 | -0.425 | -0.6818 | -0.7139
2012-10-30 15:28 | 24.3308 | 19.7949 | 19.9875 | 19.7267 | 19.7307 | 45.2592 | 0.014206 | 0.34541 | 0.004907 | 24.3145 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.4727 | 3.1202 | 0.008284 | -0.6778 | -0.4852 | -0.746 | -0.742
2012-10-30 15:28 | 24.3233 | 19.7754 | 19.9881 | 19.7353 | 19.7633 | 45.2782 | 0.014219 | 0.34553 | 0.004913 | 24.3004 | 0.014404 | 0.23448 | 0.003377 | 16.2787 | 20.4692 | 3.1215 | 0.00829 | -0.6938 | -0.4811 | -0.7339 | -0.7059
2012-10-30 15:28 | 24.3143 | 19.7503 | 19.9269 | 19.6861 | 19.7423 | 45.2655 | 0.014203 | 0.34543 | 0.004906 | 24.3207 | 0.014404 | 0.23449 | 0.003378 | 16.2798 | 20.4722 | 3.1206 | 0.008284 | -0.7219 | -0.5453 | -0.7861 | -0.7299
2012-10-30 15:28 | 24.3278 | 19.7438 | 19.9565 | 19.7117 | 19.6795 | 45.273 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014398 | 0.23449 | 0.003376 | 16.2866 | 20.4977 | 3.1211 | 0.008286 | -0.7539 | -0.5412 | -0.786 | -0.8182
2012-10-30 15:28 | 24.3076 | 19.7436 | 19.9403 | 19.7196 | 19.7396 | 45.2742 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.5336 | 3.1212 | 0.008291 | -0.79 | -0.5933 | -0.814 | -0.794
2012-10-30 15:28 | 24.3218 | 19.7538 | 19.9665 | 19.7338 | 19.7899 | 45.2881 | 0.014212 | 0.34537 | 0.004908 | 24.3012 | 0.014408 | 0.23458 | 0.00338 | 16.2812 | 20.5158 | 3.1222 | 0.008288 | -0.762 | -0.5493 | -0.782 | -0.7259
2012-10-30 15:28 | 24.3211 | 19.7772 | 19.9698 | 19.7371 | 19.7652 | 45.2626 | 0.014216 | 0.34548 | 0.004911 | 24.3025 | 0.014414 | 0.23455 | 0.003381 | 16.2725 | 20.499 | 3.1204 | 0.008292 | -0.7218 | -0.5292 | -0.7619 | -0.7338
2012-10-30 15:28 | 24.3158 | 19.7398 | 19.9725 | 19.7237 | 19.7558 | 45.2777 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014411 | 0.23454 | 0.00338 | 16.2749 | 20.4817 | 3.1214 | 0.008289 | -0.7419 | -0.5092 | -0.758 | -0.7259
2012-10-30 15:28 | 24.3054 | 19.7816 | 19.9341 | 19.7093 | 19.7495 | 45.295 | 0.014216 | 0.3455 | 0.004912 | 24.3035 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.4793 | 3.1226 | 0.008293 | -0.6977 | -0.5452 | -0.77 | -0.7298
2012-10-30 15:28 | 24.3203 | 19.8045 | 20.0091 | 19.7563 | 19.7764 | 45.2621 | 0.014222 | 0.34546 | 0.004913 | 24.2903 | 0.014408 | 0.23449 | 0.003379 | 16.2753 | 20.4542 | 3.1204 | 0.008292 | -0.6497 | -0.4451 | -0.6979 | -0.6778
2012-10-30 15:28 | 24.3121 | 19.7802 | 19.9648 | 19.72 | 19.7883 | 45.284 | 0.014216 | 0.34547 | 0.004911 | 24.3015 | 0.01443 | 0.23454 | 0.003384 | 16.2534 | 20.458 | 3.1219 | 0.008295 | -0.6778 | -0.4932 | -0.738 | -0.6697
2012-10-30 15:28 | 24.3121 | 19.8083 | 19.9889 | 19.7521 | 19.7682 | 45.2967 | 0.014203 | 0.34544 | 0.004906 | 24.3217 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.442 | 3.1228 | 0.008285 | -0.6337 | -0.4531 | -0.6899 | -0.6738
2012-10-30 15:28 | 24.3166 | 19.8128 | 20.0014 | 19.7406 | 19.7807 | 45.2701 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.4705 | 3.1209 | 0.008286 | -0.6577 | -0.4691 | -0.7299 | -0.6898
2012-10-30 15:28 | 24.3143 | 19.7744 | 19.967 | 19.7624 | 19.7584 | 45.2673 | 0.014222 | 0.34537 | 0.004912 | 24.2841 | 0.014408 | 0.23452 | 0.003379 | 16.2773 | 20.4041 | 3.1207 | 0.008291 | -0.6297 | -0.4371 | -0.6417 | -0.6457
2012-10-30 15:28 | 24.3106 | 19.7627 | 20.0235 | 19.7426 | 19.7707 | 45.28 | 0.014197 | 0.3455 | 0.004905 | 24.3361 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4084 | 3.1216 | 0.008283 | -0.6457 | -0.3849 | -0.6658 | -0.6377
2012-10-30 15:28 | 24.3084 | 19.7806 | 20.0012 | 19.7525 | 19.7846 | 45.2655 | 0.0142 | 0.34546 | 0.004905 | 24.3279 | 0.014395 | 0.23451 | 0.003376 | 16.2909 | 20.3701 | 3.1206 | 0.008281 | -0.5895 | -0.3689 | -0.6176 | -0.5855
2012-10-30 15:28 | 24.3098 | 19.7739 | 20.0147 | 19.7418 | 19.7539 | 45.2817 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014414 | 0.23452 | 0.00338 | 16.2705 | 20.4477 | 3.1217 | 0.008291 | -0.6738 | -0.433 | -0.7059 | -0.6938
2012-10-30 15:28 | 24.3136 | 19.7978 | 19.9423 | 19.7135 | 19.7456 | 45.2811 | 0.014219 | 0.34547 | 0.004912 | 24.2964 | 0.014417 | 0.23455 | 0.003382 | 16.2691 | 20.4234 | 3.1217 | 0.008294 | -0.6256 | -0.4811 | -0.7099 | -0.6778
2012-10-30 15:28 | 24.3076 | 19.7597 | 20.0044 | 19.7316 | 19.7798 | 45.2759 | 0.0142 | 0.34547 | 0.004906 | 24.3289 | 0.014411 | 0.23452 | 0.00338 | 16.2739 | 20.4575 | 3.1213 | 0.008286 | -0.6978 | -0.4531 | -0.7259 | -0.6777
2012-10-30 15:28 | 24.3121 | 19.7562 | 19.9889 | 19.7481 | 19.7401 | 45.2806 | 0.014222 | 0.34541 | 0.004912 | 24.2872 | 0.014414 | 0.23454 | 0.003381 | 16.2715 | 20.45 | 3.1216 | 0.008293 | -0.6938 | -0.4611 | -0.7019 | -0.7099
2012-10-30 15:28 | 24.2979 | 19.75 | 20.0108 | 19.6617 | 19.766 | 45.284 | 0.014212 | 0.34547 | 0.00491 | 24.3083 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4237 | 3.1219 | 0.008289 | -0.6737 | -0.4129 | -0.762 | -0.6577
2012-10-30 15:29 | 24.3009 | 19.761 | 19.9416 | 19.7249 | 19.757 | 45.2696 | 0.014209 | 0.34544 | 0.004908 | 24.3114 | 0.014411 | 0.23449 | 0.003379 | 16.2719 | 20.4107 | 3.1209 | 0.008287 | -0.6497 | -0.4691 | -0.6858 | -0.6537
2012-10-30 15:29 | 24.3031 | 19.8154 | 19.9799 | 19.7592 | 19.7431 | 45.2777 | 0.014216 | 0.34544 | 0.004911 | 24.2995 | 0.014408 | 0.23459 | 0.00338 | 16.2823 | 20.437 | 3.1214 | 0.008291 | -0.6216 | -0.4571 | -0.6778 | -0.6939
2012-10-30 15:29 | 24.3016 | 19.7818 | 19.9503 | 19.6855 | 19.7497 | 45.2829 | 0.014209 | 0.34541 | 0.004908 | 24.3094 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.4475 | 3.1218 | 0.008286 | -0.6657 | -0.4972 | -0.762 | -0.6978
2012-10-30 15:29 | 24.2986 | 19.7868 | 19.9433 | 19.7708 | 19.7186 | 45.2719 | 0.0142 | 0.34541 | 0.004905 | 24.3248 | 0.014389 | 0.23461 | 0.003376 | 16.3047 | 20.4244 | 3.121 | 0.008281 | -0.6376 | -0.4811 | -0.6536 | -0.7058
2012-10-30 15:29 | 24.3084 | 19.8006 | 19.9892 | 19.7685 | 19.7605 | 45.2586 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014408 | 0.23457 | 0.00338 | 16.2803 | 20.4262 | 3.1201 | 0.008288 | -0.6256 | -0.437 | -0.6577 | -0.6657
2012-10-30 15:29 | 24.3054 | 19.7896 | 19.9942 | 19.7053 | 19.7655 | 45.2904 | 0.014209 | 0.34547 | 0.004909 | 24.3135 | 0.014398 | 0.23452 | 0.003377 | 16.2886 | 20.4713 | 3.1223 | 0.008286 | -0.6817 | -0.4771 | -0.766 | -0.7058
2012-10-30 15:29 | 24.2979 | 19.7379 | 19.9546 | 19.7379 | 19.7299 | 45.2759 | 0.014206 | 0.34548 | 0.004908 | 24.3197 | 0.014411 | 0.23451 | 0.003379 | 16.2728 | 20.4798 | 3.1213 | 0.008287 | -0.7419 | -0.5252 | -0.7419 | -0.7499
2012-10-30 15:29 | 24.3016 | 19.7216 | 19.9944 | 19.7095 | 19.7738 | 45.2684 | 0.014219 | 0.34546 | 0.004912 | 24.2954 | 0.014401 | 0.23459 | 0.003378 | 16.2902 | 20.4956 | 3.1208 | 0.00829 | -0.774 | -0.5012 | -0.7861 | -0.7218
2012-10-30 15:29 | 24.3001 | 19.7763 | 19.9969 | 19.7482 | 19.7562 | 45.2765 | 0.014209 | 0.34543 | 0.004908 | 24.3104 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.478 | 3.1214 | 0.008287 | -0.7017 | -0.4811 | -0.7298 | -0.7218
2012-10-30 15:29 | 24.2964 | 19.8047 | 20.0093 | 19.7244 | 19.7846 | 45.2806 | 0.014206 | 0.34543 | 0.004907 | 24.3156 | 0.014401 | 0.23455 | 0.003378 | 16.2872 | 20.4623 | 3.1216 | 0.008285 | -0.6576 | -0.453 | -0.7379 | -0.6777
2012-10-30 15:29 | 24.3024 | 19.7826 | 19.9752 | 19.7144 | 19.7585 | 45.2794 | 0.014219 | 0.34543 | 0.004912 | 24.2933 | 0.014404 | 0.23447 | 0.003377 | 16.2778 | 20.4723 | 3.1216 | 0.008289 | -0.6897 | -0.4971 | -0.7579 | -0.7138
2012-10-30 15:29 | 24.3061 | 19.7863 | 20.027 | 19.706 | 19.7863 | 45.2788 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014417 | 0.23451 | 0.003381 | 16.2661 | 20.4279 | 3.1215 | 0.008288 | -0.6416 | -0.4009 | -0.7219 | -0.6416
2012-10-30 15:29 | 24.3046 | 19.7727 | 19.9854 | 19.7768 | 19.7687 | 45.2505 | 0.014206 | 0.34546 | 0.004908 | 24.3176 | 0.014404 | 0.23452 | 0.003378 | 16.2818 | 20.4385 | 3.1196 | 0.008286 | -0.6658 | -0.4531 | -0.6617 | -0.6698
2012-10-30 15:29 | 24.3039 | 19.756 | 19.9767 | 19.764 | 19.7761 | 45.2823 | 0.014209 | 0.34546 | 0.004909 | 24.3125 | 0.014398 | 0.23447 | 0.003376 | 16.2846 | 20.4257 | 3.1218 | 0.008285 | -0.6697 | -0.449 | -0.6617 | -0.6496
2012-10-30 15:29 | 24.3031 | 19.7993 | 20.012 | 19.7191 | 19.7271 | 45.273 | 0.014206 | 0.3454 | 0.004907 | 24.3135 | 0.014401 | 0.23449 | 0.003377 | 16.2832 | 20.4209 | 3.1211 | 0.008284 | -0.6216 | -0.4089 | -0.7018 | -0.6938
2012-10-30 15:29 | 24.3024 | 19.7786 | 19.9832 | 19.7465 | 19.7786 | 45.2701 | 0.014206 | 0.34544 | 0.004907 | 24.3166 | 0.014408 | 0.23448 | 0.003378 | 16.2742 | 20.4563 | 3.1209 | 0.008285 | -0.6777 | -0.4731 | -0.7098 | -0.6777
2012-10-30 15:29 | 24.3024 | 19.7465 | 20.0073 | 19.7545 | 19.7625 | 45.2678 | 0.0142 | 0.34543 | 0.004905 | 24.3258 | 0.014414 | 0.23457 | 0.003381 | 16.2735 | 20.4523 | 3.1208 | 0.008286 | -0.7058 | -0.445 | -0.6978 | -0.6898
2012-10-30 15:29 | 24.2971 | 19.7492 | 20.006 | 19.7492 | 19.7693 | 45.2707 | 0.014212 | 0.34543 | 0.004909 | 24.3053 | 0.014414 | 0.23458 | 0.003381 | 16.2745 | 20.467 | 3.121 | 0.00829 | -0.7178 | -0.461 | -0.7178 | -0.6977
2012-10-30 15:30 | 24.2971 | 19.7773 | 19.9939 | 19.701 | 19.7452 | 45.28 | 0.014203 | 0.34548 | 0.004907 | 24.3248 | 0.014411 | 0.23458 | 0.003381 | 16.2778 | 20.451 | 3.1216 | 0.008288 | -0.6737 | -0.4571 | -0.75 | -0.7058
2012-10-30 15:30 | 24.3039 | 19.7921 | 19.9807 | 19.6918 | 19.764 | 45.2707 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014408 | 0.23451 | 0.003379 | 16.2762 | 20.4939 | 3.121 | 0.008288 | -0.7018 | -0.5132 | -0.8021 | -0.7299
2012-10-30 15:30 | 24.2971 | 19.7612 | 19.9939 | 19.689 | 19.7733 | 45.2684 | 0.014212 | 0.34541 | 0.004909 | 24.3042 | 0.014417 | 0.23457 | 0.003382 | 16.2701 | 20.471 | 3.1208 | 0.008291 | -0.7098 | -0.4771 | -0.782 | -0.6977
2012-10-30 15:30 | 24.3024 | 19.7103 | 20.0033 | 19.6943 | 19.7665 | 45.2811 | 0.014219 | 0.34548 | 0.004912 | 24.2974 | 0.014408 | 0.23454 | 0.003379 | 16.2782 | 20.5204 | 3.1217 | 0.008291 | -0.8101 | -0.5171 | -0.8261 | -0.7539
2012-10-30 15:30 | 24.2986 | 19.7788 | 19.9995 | 19.7386 | 19.7667 | 45.2921 | 0.014203 | 0.34546 | 0.004907 | 24.3227 | 0.01442 | 0.23447 | 0.003381 | 16.2597 | 20.5006 | 3.1224 | 0.008288 | -0.7218 | -0.5011 | -0.762 | -0.7339
2012-10-30 15:30 | 0 | -15.6017 | -14.0295 | -14.2743 | -14.1804 | 43.6791 | 0.006707 | 0.31152 | 0.002089 | 46.4465 | 0.006902 | 0.21546 | 0.001487 | 31.2178 | -13.5157 | 3.0112 | 0.003576 | -2.086 | -0.5138 | -0.7586 | -0.6647
2012-10-30 15:30 | 26.3397 | 10.7707 | 12.3469 | 12.1185 | 12.1307 | 43.6947 | 0.006703 | 0.32579 | 0.002184 | 48.6039 | 0.006895 | 0.21551 | 0.001486 | 31.2558 | 12.8606 | 3.0123 | 0.00367 | -2.0899 | -0.5137 | -0.7421 | -0.7299
2012-10-30 15:30 | 26.3352 | 10.713 | 12.3465 | 12.1017 | 12.1099 | 43.6751 | 0.006707 | 0.32575 | 0.002185 | 48.5684 | 0.006902 | 0.21545 | 0.001487 | 31.2157 | 12.848 | 3.011 | 0.003672 | -2.135 | -0.5015 | -0.7463 | -0.7381
2012-10-30 15:30 | 26.3344 | 10.7613 | 12.3579 | 12.1091 | 12.1376 | 43.6768 | 0.006722 | 0.32579 | 0.00219 | 48.4665 | 0.006921 | 0.21549 | 0.001491 | 31.1363 | 12.8431 | 3.0111 | 0.003681 | -2.0818 | -0.4852 | -0.734 | -0.7055
2012-10-30 15:30 | 26.3329 | 10.7557 | 12.3564 | 12.0709 | 12.128 | 43.6866 | 0.00671 | 0.32585 | 0.002186 | 48.5617 | 0.006921 | 0.21551 | 0.001492 | 31.1384 | 12.8416 | 3.0118 | 0.003678 | -2.0859 | -0.4852 | -0.7707 | -0.7136
2012-10-30 15:30 | 26.3337 | 10.7606 | 12.3735 | 12.1002 | 12.1369 | 43.6895 | 0.00671 | 0.32582 | 0.002186 | 48.5575 | 0.006908 | 0.21557 | 0.001489 | 31.2054 | 12.8465 | 3.012 | 0.003675 | -2.0859 | -0.473 | -0.7463 | -0.7096
2012-10-30 15:30 | 26.3314 | 10.7256 | 12.3916 | 12.0979 | 12.102 | 43.6953 | 0.00671 | 0.32581 | 0.002186 | 48.5553 | 0.006911 | 0.2156 | 0.00149 | 31.1961 | 12.8442 | 3.0124 | 0.003676 | -2.1186 | -0.4526 | -0.7463 | -0.7422
2012-10-30 15:30 | 26.3322 | 10.7264 | 12.3639 | 12.0742 | 12.1069 | 43.6768 | 0.006707 | 0.32579 | 0.002185 | 48.5749 | 0.006911 | 0.21558 | 0.00149 | 31.1939 | 12.8694 | 3.0111 | 0.003675 | -2.143 | -0.5055 | -0.7952 | -0.7625
2012-10-30 15:30 | 26.3344 | 10.7449 | 12.3702 | 12.0846 | 12.1458 | 43.6849 | 0.006694 | 0.32589 | 0.002182 | 48.6843 | 0.006905 | 0.21557 | 0.001488 | 31.219 | 12.8961 | 3.0116 | 0.00367 | -2.1512 | -0.5259 | -0.8115 | -0.7503
2012-10-30 15:30 | 26.3337 | 10.7279 | 12.3368 | 12.1247 | 12.1084 | 43.6768 | 0.0067 | 0.32579 | 0.002183 | 48.6257 | 0.006898 | 0.21554 | 0.001487 | 31.2464 | 12.8954 | 3.0111 | 0.00367 | -2.1675 | -0.5586 | -0.7707 | -0.787
2012-10-30 15:30 | 26.3389 | 10.7699 | 12.3869 | 12.1258 | 12.1544 | 43.6889 | 0.006703 | 0.32581 | 0.002184 | 48.606 | 0.006917 | 0.21555 | 0.001491 | 31.1626 | 12.8802 | 3.0119 | 0.003675 | -2.1103 | -0.4933 | -0.7544 | -0.7258
2012-10-30 15:30 | 26.3352 | 10.7376 | 12.3383 | 12.1099 | 12.1262 | 43.6889 | 0.00671 | 0.32586 | 0.002187 | 48.5639 | 0.006898 | 0.21554 | 0.001487 | 31.2464 | 12.8684 | 3.0119 | 0.003674 | -2.1308 | -0.5301 | -0.7585 | -0.7422
2012-10-30 15:30 | 26.3344 | 10.7572 | 12.3498 | 12.0846 | 12.1336 | 43.6791 | 0.006707 | 0.32586 | 0.002186 | 48.5857 | 0.006905 | 0.21558 | 0.001489 | 31.221 | 12.9083 | 3.0112 | 0.003675 | -2.1511 | -0.5585 | -0.8237 | -0.7747
2012-10-30 15:30 | 26.3284 | 10.7635 | 12.3315 | 12.099 | 12.1235 | 43.6901 | 0.00671 | 0.32589 | 0.002187 | 48.5683 | 0.006917 | 0.21558 | 0.001491 | 31.1668 | 12.8901 | 3.012 | 0.003678 | -2.1266 | -0.5586 | -0.7911 | -0.7666
2012-10-30 15:30 | 26.3352 | 10.7539 | 12.3343 | 12.114 | 12.1221 | 43.6739 | 0.00671 | 0.32589 | 0.002187 | 48.5683 | 0.006908 | 0.21557 | 0.001489 | 31.2054 | 12.9173 | 3.0109 | 0.003676 | -2.1634 | -0.583 | -0.8033 | -0.7952
2012-10-30 15:30 | 26.3337 | 10.7647 | 12.3164 | 12.1002 | 12.1451 | 43.7121 | 0.018537 | 0.61218 | 0.011348 | 33.0249 | 0.040104 | 0.74736 | 0.029972 | 18.6355 | 12.8954 | 3.0135 | 0.04132 | -2.1307 | -0.579 | -0.7952 | -0.7503
2012-10-30 15:30 | 26.3322 | 10.7795 | 12.4699 | 12.0946 | 12.1232 | 44.0605 | 0.90931 | 22.3261 | 20.3013 | 24.5527 | 1.4066 | 22.6084 | 31.8012 | 16.073 | 12.8613 | 3.0375 | 52.1025 | -2.0818 | -0.3914 | -0.7667 | -0.7381
2012-10-30 15:31 | 26.3307 | 12.1584 | 13.8657 | 12.1584 | 12.138 | 45.7394 | 1.151 | 28.3265 | 32.6031 | 24.6109 | 1.7649 | 28.3212 | 49.9853 | 16.0465 | 12.8435 | 3.1533 | 82.5884 | -0.6851 | 1.0222 | -0.6851 | -0.7055
2012-10-30 15:31 | 26.3389 | 14.1628 | 15.0652 | 12.3951 | 12.3706 | 46.2786 | 1.2667 | 31.2392 | 39.5719 | 24.6611 | 1.9459 | 31.1985 | 60.7096 | 16.0329 | 12.8435 | 3.1904 | 100.2815 | 1.3193 | 2.2217 | -0.4484 | -0.4729
2012-10-30 15:31 | 26.3314 | 16.5064 | 16.182 | 12.7464 | 12.6241 | 46.6924 | 1.2684 | 31.2958 | 39.6955 | 24.6736 | 1.9495 | 31.2537 | 60.9279 | 16.0319 | 12.8768 | 3.219 | 100.6234 | 3.6296 | 3.3052 | -0.1304 | -0.2527
2012-10-30 15:31 | 26.3352 | 18.9816 | 17.3081 | 13.341 | 13.2432 | 46.9063 | 1.2678 | 31.2895 | 39.6692 | 24.68 | 1.9492 | 31.2481 | 60.9082 | 16.0314 | 12.8806 | 3.2337 | 100.5774 | 6.101 | 4.4275 | 0.4604 | 0.3626
2012-10-30 15:31 | 26.3359 | 21.1253 | 18.3523 | 14.1557 | 13.8791 | 47.0837 | 1.2656 | 31.2437 | 39.5434 | 24.686 | 1.9463 | 31.2014 | 60.7276 | 16.031 | 12.8976 | 3.246 | 100.271 | 8.2277 | 5.4547 | 1.2581 | 0.9815
2012-10-30 15:31 | 26.3322 | 23.5265 | 19.4063 | 15.156 | 14.6481 | 47.2865 | 1.2669 | 31.2829 | 39.6315 | 24.693 | 1.9487 | 31.2385 | 60.8749 | 16.0303 | 12.898 | 3.2599 | 100.5064 | 10.6285 | 6.5083 | 2.258 | 1.7501
2012-10-30 15:31 | 26.3277 |  | 20.5782 | 16.3405 | 15.4723 | 47.4957 | 1.2679 | 31.3155 | 39.7037 | 24.6995 | 1.9508 | 31.2709 | 61.0031 | 16.0298 | 12.8731 | 3.2744 | 100.7067 | -12.8731 | 7.7051 | 3.4674 | 2.5992
2012-10-30 15:31 | 26.3352 | 28.0209 | 21.6548 | 17.5833 | 16.275 | 47.6772 | 1.2675 | 31.3151 | 39.6934 | 24.7052 | 1.9508 | 31.2699 | 61.0006 | 16.0295 | 12.8806 | 3.2869 | 100.694 | 15.1403 | 8.7742 | 4.7027 | 3.3944
2012-10-30 15:31 | 26.3314 | 30.086 | 22.6896 | 19.0626 | 17.2072 | 47.8876 | 1.2673 | 31.3146 | 39.6839 | 24.7103 | 1.9508 | 31.269 | 60.9983 | 16.0291 | 12.8768 | 3.3014 | 100.6822 | 17.2092 | 9.8128 | 6.1858 | 4.3304
2012-10-30 15:31 | 26.3337 | 32.5122 | 24.0592 | 20.8939 | 18.3016 | 48.08 | 1.2667 | 31.3077 | 39.6563 | 24.7167 | 1.9504 | 31.2616 | 60.9716 | 16.0286 | 12.8465 | 3.3146 | 100.6279 | 19.6657 | 11.2127 | 8.0474 | 5.4551
2012-10-30 15:31 | 26.3322 |  |  |  |  | 827385.764 | 1.2663 | 31.3068 | 39.6452 | 24.7221 | 1.9503 | 31.2602 | 60.9673 | 16.0283 |  | 57039.9746 | 100.6125 | 0 | 0 | 0 | 0
2012-10-30 15:31 | 24.7878 | 46.4656 | 34.5123 | 32.8483 | 28.4871 | 50.1091 | 1.2662 | 31.3092 | 39.6435 | 24.727 | 1.9505 | 31.262 | 60.9757 | 16.0279 | 21.0419 | 3.4545 | 100.6192 | 25.4237 | 13.4704 | 11.8064 | 7.4452
2012-10-30 15:31 | 24.2897 | 47.487 | 34.9009 | 33.7383 | 28.7392 | 50.208 | 1.2661 | 31.3132 | 39.6456 | 24.732 | 1.9508 | 31.2658 | 60.9919 | 16.0276 | 20.5037 | 3.4613 | 100.6375 | 26.9833 | 14.3972 | 13.2346 | 8.2355
2012-10-30 15:31 | 24.2964 | 49.4554 | 36.0578 | 35.4986 | 29.6933 | 50.4027 | 1.2661 | 31.3202 | 39.6554 | 24.737 | 1.9512 | 31.2725 | 61.0198 | 16.0271 | 20.5144 | 3.4748 | 100.6752 | 28.941 | 15.5434 | 14.9842 | 9.1789
2012-10-30 15:31 | 24.3001 | 51.3978 | 37.2384 | 37.3447 | 30.7311 | 50.5755 | 1.266 | 31.3248 | 39.6585 | 24.7424 | 1.9515 | 31.2768 | 61.038 | 16.0267 | 20.5101 | 3.4867 | 100.6965 | 30.8877 | 16.7283 | 16.8346 | 10.221
2012-10-30 15:31 | 24.2949 | 53.2812 | 38.4486 | 39.2269 | 31.7479 | 50.7321 | 1.2657 | 31.3232 | 39.6472 | 24.7469 | 1.9515 | 31.2749 | 61.0319 | 16.0263 | 20.5369 | 3.4975 | 100.6791 | 32.7443 | 17.9117 | 18.69 | 11.211
2012-10-30 15:31 | 24.2964 | 55.137 | 39.4917 | 41.1331 | 32.6883 | 50.8916 | 1.2651 | 31.3123 | 39.6125 | 24.7514 | 1.9508 | 31.2635 | 60.9885 | 16.0261 | 20.5024 | 3.5085 | 100.601 | 34.6346 | 18.9893 | 20.6307 | 12.1859
2012-10-30 15:31 | 24.2979 | 56.9547 | 40.5654 | 42.9268 | 33.5928 | 51.0656 | 1.2648 | 31.3103 | 39.5998 | 24.756 | 1.9507 | 31.2613 | 60.9816 | 16.0257 | 20.5359 | 3.5205 | 100.5814 | 36.4188 | 20.0295 | 22.3909 | 13.0569
2012-10-30 15:31 | 24.2971 | 58.8056 | 41.8086 | 44.9425 | 34.5616 | 51.2297 | 1.2647 | 31.3133 | 39.6004 | 24.7605 | 1.9509 | 31.2642 | 60.9933 | 16.0255 | 20.4911 | 3.5318 | 100.5937 | 38.3145 | 21.3175 | 24.4514 | 14.0705
2012-10-30 15:31 | 24.2949 | 60.5786 | 43.0374 | 46.8987 | 35.5168 | 51.4031 | 1.2589 | 31.1763 | 39.2493 | 24.7638 | 1.9423 | 31.1257 | 60.4545 | 16.0255 | 20.4808 | 3.5437 | 99.7038 | 40.0978 | 22.5566 | 26.4179 | 15.036
2012-10-30 15:31 | 24.2956 | 62.2439 | 44.1191 | 48.6594 | 36.5216 | 51.558 | 1.4188 | 35.1919 | 49.929 | 24.8046 | 2.1958 | 35.1642 | 77.2145 | 16.0141 | 20.4775 | 3.5544 | 127.1435 | 41.7664 | 23.6416 | 28.1819 | 16.0441
2012-10-30 15:31 | 24.2934 | 64.1575 | 45.807 | 50.43 | 37.5229 | 52.5855 | 1.5382 | 38.2706 | 58.8688 | 24.8797 | 2.3871 | 38.2067 | 91.2036 | 16.0054 | 20.5314 | 3.6252 | 150.0724 | 43.6261 | 25.2756 | 29.8986 | 16.9915
2012-10-30 15:32 | 24.2874 | 67.3949 | 48.0271 | 52.5669 | 38.7557 | 52.9889 | 1.5386 | 38.2986 | 58.9276 | 24.8913 | 2.3889 | 38.2342 | 91.337 | 16.005 | 20.5735 | 3.6531 | 150.2646 | 46.8214 | 27.4536 | 31.9934 | 18.1822
2012-10-30 15:32 | 24.2904 | 70.7321 | 49.8779 | 54.8713 | 40.1221 | 53.2583 | 1.5383 | 38.303 | 58.9226 | 24.8991 | 2.3892 | 38.2382 | 91.3579 | 16.0048 | 20.5965 | 3.6716 | 150.2804 | 50.1356 | 29.2814 | 34.2748 | 19.5256
2012-10-30 15:32 | 24.2934 | 73.1338 | 51.3017 | 56.7528 | 41.3341 | 53.4611 | 1.5385 | 38.3152 | 58.9466 | 24.9049 | 2.39 | 38.2504 | 91.4169 | 16.0046 | 20.5915 | 3.6856 | 150.3635 | 52.5423 | 30.7102 | 36.1613 | 20.7426
2012-10-30 15:32 | 24.3001 | 75.318 | 52.7583 | 58.6734 | 42.4743 | 53.679 | 1.5381 | 38.3155 | 58.9333 | 24.9108 | 2.3899 | 38.2499 | 91.4153 | 16.0045 | 20.5942 | 3.7006 | 150.3486 | 54.7238 | 32.1641 | 38.0792 | 21.8801
2012-10-30 15:32 | 24.2934 | 77.7483 | 54.312 | 60.6928 | 43.7293 | 53.9119 | 1.5381 | 38.3261 | 58.951 | 24.9171 | 2.3906 | 38.2601 | 91.4644 | 16.0044 | 20.6115 | 3.7167 | 150.4154 | 57.1368 | 33.7005 | 40.0813 | 23.1178
2012-10-30 15:32 | 24.3001 | 80.3746 | 55.9659 | 63.0528 | 45.2154 | 54.1402 | 1.5379 | 38.3282 | 58.9432 | 24.9231 | 2.3907 | 38.2621 | 91.4744 | 16.0043 | 20.6263 | 3.7324 | 150.4176 | 59.7483 | 35.3396 | 42.4265 | 24.5891
2012-10-30 15:32 | 24.2897 | 82.7374 | 57.488 | 65.2641 | 46.6747 | 54.3534 | 1.5374 | 38.3246 | 58.9191 | 24.9287 | 2.3905 | 38.2581 | 91.4563 | 16.0042 | 20.5878 | 3.7471 | 150.3754 | 62.1496 | 36.9002 | 44.6763 | 26.0869
2012-10-30 15:32 | 24.2934 | 84.8843 | 58.8984 | 67.3779 | 47.9434 | 54.5626 | 1.5371 | 38.3257 | 58.9097 | 24.9341 | 2.3906 | 38.259 | 91.4609 | 16.0041 | 20.5755 | 3.7615 | 150.3706 | 64.3088 | 38.3229 | 46.8024 | 27.3679
2012-10-30 15:32 | 24.2926 | 87.1649 | 60.4606 | 69.8126 | 49.2801 | 54.7938 | 1.5372 | 38.3375 | 58.9324 | 24.9398 | 2.3913 | 38.2703 | 91.5157 | 16.004 | 20.5547 | 3.7775 | 150.4481 | 66.6102 | 39.9059 | 49.2579 | 28.7254
2012-10-30 15:32 | 24.2926 | 89.7589 | 62.2216 | 72.4879 | 50.8223 | 55.04 | 1.5368 | 38.337 | 58.9163 | 24.946 | 2.3913 | 38.2693 | 91.5118 | 16.0038 | 20.5547 | 3.7945 | 150.4281 | 69.2042 | 41.6669 | 51.9332 | 30.2676
2012-10-30 15:32 | 24.2882 | 92.4903 | 64.0562 | 75.2794 | 52.5016 | 55.303 | 1.5361 | 38.3292 | 58.8775 | 24.9523 | 2.3908 | 38.2614 | 91.4749 | 16.0037 | 20.5503 | 3.8126 | 150.3523 | 71.94 | 43.5059 | 54.7291 | 31.9513
2012-10-30 15:32 | 24.2919 | 94.7899 | 65.7155 | 77.8001 | 53.9924 | 55.5272 | 1.5359 | 38.3317 | 58.8723 | 24.9578 | 2.3909 | 38.2637 | 91.4861 | 16.0036 | 20.57 | 3.828 | 150.3583 | 74.2199 | 45.1455 | 57.2301 | 33.4224
2012-10-30 15:32 | 24.2919 | 96.9451 | 67.2307 | 79.8763 | 55.2759 | 55.748 | 1.5364 | 38.3533 | 58.926 | 24.9632 | 2.3923 | 38.285 | 91.5885 | 16.0035 | 20.546 | 3.8433 | 150.5145 | 76.3991 | 46.6847 | 59.3303 | 34.7299
2012-10-30 15:32 | 24.2904 | 99.3972 | 68.9187 | 82.4497 | 56.7885 | 55.9647 | 1.5369 | 38.3736 | 58.9746 | 24.9689 | 2.3935 | 38.3048 | 91.6838 | 16.0034 | 20.5164 | 3.8582 | 150.6584 | 78.8808 | 48.4023 | 61.9333 | 36.2721
2012-10-30 15:32 | 24.2874 | 101.621 | 70.4347 | 84.8479 | 58.197 | 56.1889 | 1.5618 | 39.0135 | 60.9325 | 24.9793 | 2.432 | 38.9193 | 94.6536 | 16.0027 | 20.5174 | 3.8737 | 155.5862 | 81.1036 | 49.9173 | 64.3305 | 37.6796
2012-10-30 15:32 | 24.2882 | 103.637 | 71.7999 | 86.7173 | 59.3141 | 56.5813 | 1.6179 | 40.4678 | 65.4719 | 25.013 | 2.5243 | 40.3927 | 101.9626 | 16.0016 | 20.5062 | 3.9007 | 167.4345 | 83.1308 | 51.2937 | 66.2111 | 38.8079
2012-10-30 15:32 | 24.2971 | 106.0624 | 73.6145 | 88.9759 | 60.6888 | 56.909 | 1.6176 | 40.4732 | 65.4687 | 25.0208 | 2.5247 | 40.3994 | 101.9961 | 16.0017 | 20.5111 | 3.9233 | 167.4648 | 85.5513 | 53.1034 | 68.4648 | 40.1777
2012-10-30 15:32 | 24.2956 | 108.4065 | 75.4126 | 91.4233 | 62.2208 | 57.1627 | 1.617 | 40.4696 | 65.4383 | 25.028 | 2.5244 | 40.3953 | 101.975 | 16.0018 | 20.5497 | 3.9408 | 167.4133 | 87.8568 | 54.8629 | 70.8736 | 41.6711
2012-10-30 15:32 | 24.2979 | 111.278 | 77.5055 | 94.3761 | 64.0659 | 57.4367 | 1.6165 | 40.4696 | 65.4201 | 25.0349 | 2.5244 | 40.3949 | 101.9723 | 16.0018 | 20.552 | 3.9597 | 167.3924 | 90.726 | 56.9535 | 73.8241 | 43.5139
2012-10-30 15:32 | 24.3061 | 114.2258 | 79.6396 | 97.1786 | 65.8794 | 57.7158 | 1.6162 | 40.4701 | 65.4073 | 25.0404 | 2.5244 | 40.3951 | 101.9724 | 16.002 | 20.5802 | 3.9789 | 167.3797 | 93.6456 | 59.0594 | 76.5984 | 45.2992
2012-10-30 15:32 | 24.3039 | 116.4491 | 81.0546 | 99.2633 | 67.189 | 57.9441 | 1.6159 | 40.4714 | 65.3985 | 25.0454 | 2.5245 | 40.3963 | 101.9785 | 16.002 | 20.562 | 3.9947 | 167.377 | 95.8871 | 60.4926 | 78.7013 | 46.627
2012-10-30 15:32 | 24.3136 | 118.2989 | 82.3894 | 101.3901 | 68.5169 | 58.1221 | 1.6155 | 40.4678 | 65.3743 | 25.0503 | 2.5242 | 40.3924 | 101.958 | 16.0022 | 20.6237 | 4.0069 | 167.3323 | 97.6752 | 61.7657 | 80.7664 | 47.8932
2012-10-30 15:32 | 24.3098 | 120.7709 | 84.2143 | 103.98 | 70.1741 | 58.392 | 1.6179 | 40.5401 | 65.5917 | 25.0565 | 2.529 | 40.4691 | 102.3459 | 16.0021 | 20.5999 | 4.0255 | 167.9376 | 100.171 | 63.6144 | 83.3801 | 49.5742
2012-10-30 15:32 | 24.3173 | 123.2302 | 86.0719 | 106.5744 | 71.8023 | 58.7098 | 1.6505 | 41.3863 | 68.3093 | 25.0746 | 2.5813 | 41.3071 | 106.6279 | 16.0022 | 20.6114 | 4.0475 | 174.9372 | 102.6188 | 65.4605 | 85.963 | 51.1909
2012-10-30 15:33 | 24.3308 | 125.1082 | 87.6577 | 108.665 | 73.1445 | 59.0335 | 1.6601 | 41.6459 | 69.1353 | 25.0868 | 2.5975 | 41.5671 | 107.9716 | 16.0026 | 20.6009 | 4.0698 | 177.1069 | 104.5073 | 67.0568 | 88.0641 | 52.5436
2012-10-30 15:33 | 24.3233 | 127.4126 | 89.3694 | 110.9134 | 74.5598 | 59.2514 | 1.6597 | 41.6471 | 69.1234 | 25.0925 | 2.5976 | 41.5682 | 107.977 | 16.0027 | 20.5894 | 4.0848 | 177.1004 | 106.8232 | 68.78 | 90.324 | 53.9704
2012-10-30 15:33 | 24.3181 | 129.6801 | 90.969 | 113.1768 | 76.0294 | 59.4848 | 1.6594 | 41.6475 | 69.1078 | 25.0987 | 2.5975 | 41.5682 | 107.9751 | 16.0029 | 20.5481 | 4.1009 | 177.083 | 109.132 | 70.4209 | 92.6287 | 55.4813
2012-10-30 15:33 | 24.3256 | 132.0699 | 92.732 | 115.5986 | 77.6169 | 59.7264 | 1.6589 | 41.6457 | 69.0875 | 25.1039 | 2.5974 | 41.5659 | 107.9616 | 16.0031 | 20.5837 | 4.1175 | 177.0491 | 111.4862 | 72.1483 | 95.0149 | 57.0332
2012-10-30 15:33 | 24.3263 | 134.3719 | 94.465 | 118.0118 | 79.1998 | 59.972 | 1.6585 | 41.6435 | 69.0657 | 25.1091 | 2.5972 | 41.5632 | 107.9464 | 16.0033 | 20.5964 | 4.1345 | 177.0121 | 113.7755 | 73.8686 | 97.4154 | 58.6034
2012-10-30 15:33 | 24.3293 | 136.3191 | 95.973 | 120.0763 | 80.5482 | 60.1847 | 1.658 | 41.6386 | 69.036 | 25.114 | 2.5968 | 41.5581 | 107.9184 | 16.0036 | 20.5713 | 4.1491 | 176.9544 | 115.7478 | 75.4017 | 99.505 | 59.9769
2012-10-30 15:33 | 24.3315 | 138.4096 | 97.5481 | 122.1151 | 82.0315 | 60.3916 | 1.6574 | 41.6328 | 69.003 | 25.119 | 2.5964 | 41.552 | 107.8858 | 16.0036 | 20.5815 | 4.1634 | 176.8888 | 117.8281 | 76.9666 | 101.5336 | 61.45
2012-10-30 15:33 | 24.3315 | 140.7959 | 99.3363 | 124.7054 | 83.6178 | 60.647 | 1.6568 | 41.6255 | 68.9641 | 25.1244 | 2.5959 | 41.5446 | 107.8468 | 16.0038 | 20.5936 | 4.181 | 176.8109 | 120.2023 | 78.7427 | 104.1118 | 63.0242
2012-10-30 15:33 | 24.3285 | 143.0134 | 101.0269 | 127.2501 | 85.1961 | 60.8799 | 1.6563 | 41.6219 | 68.9383 | 25.1295 | 2.5957 | 41.5408 | 107.8255 | 16.004 | 20.5745 | 4.1971 | 176.7638 | 122.4389 | 80.4524 | 106.6756 | 64.6216
2012-10-30 15:33 | 24.3338 | 144.8995 | 102.3895 | 129.074 | 86.4446 | 61.0689 | 1.6559 | 41.6178 | 68.9133 | 25.1336 | 2.5953 | 41.5364 | 107.8012 | 16.0042 | 20.5478 | 4.2101 | 176.7145 | 124.3517 | 81.8417 | 108.5262 | 65.8968
2012-10-30 15:33 | 24.3383 | 146.7291 | 103.6606 | 130.9102 | 87.6576 | 61.2724 | 1.6555 | 41.6147 | 68.8917 | 25.1377 | 2.5951 | 41.5334 | 107.7846 | 16.0044 | 20.5964 | 4.2241 | 176.6763 | 126.1327 | 83.0642 | 110.3138 | 67.0612
2012-10-30 15:33 | 24.336 | 148.4999 | 105.1709 | 132.9057 | 88.9126 | 61.4544 | 1.655 | 41.6116 | 68.8668 | 25.1431 | 2.5949 | 41.5299 | 107.765 | 16.0046 | 20.586 | 4.2367 | 176.6317 | 127.9139 | 84.5849 | 112.3197 | 68.3266
2012-10-30 15:33 | 24.3413 | 150.6005 | 106.9123 | 135.3434 | 90.5343 | 61.6925 | 1.6546 | 41.6104 | 68.8487 | 25.1483 | 2.5947 | 41.5286 | 107.7561 | 16.0048 | 20.5833 | 4.2531 | 176.6048 | 130.0172 | 86.329 | 114.7601 | 69.951
2012-10-30 15:33 | 24.3457 | 152.6293 | 108.5588 | 137.753 | 92.1506 | 61.9121 | 1.6543 | 41.609 | 68.8321 | 25.1526 | 2.5946 | 41.5269 | 107.7466 | 16.005 | 20.5998 | 4.2682 | 176.5787 | 132.0295 | 87.959 | 117.1532 | 71.5508
2012-10-30 15:33 | 24.3383 | 154.3817 | 109.9593 | 139.6858 | 93.4067 | 62.0913 | 1.654 | 41.6082 | 68.8188 | 25.1565 | 2.5945 | 41.526 | 107.7408 | 16.0051 | 20.6004 | 4.2806 | 176.5596 | 133.7813 | 89.3589 | 119.0854 | 72.8063
2012-10-30 15:33 | 24.3338 | 156.2471 | 111.3934 | 141.5854 | 94.6389 | 62.293 | 1.6536 | 41.6077 | 68.8044 | 25.1611 | 2.5944 | 41.5253 | 107.7351 | 16.0055 | 20.5358 | 4.2945 | 176.5395 | 135.7113 | 90.8576 | 121.0496 | 74.1031
2012-10-30 15:33 | 24.3368 | 158.4089 | 113.1198 | 143.9445 | 96.2338 | 62.5311 | 1.6533 | 41.6065 | 68.787 | 25.1661 | 2.5943 | 41.5239 | 107.7261 | 16.0057 | 20.5227 | 4.3109 | 176.5132 | 137.8862 | 92.5971 | 123.4218 | 75.7111
2012-10-30 15:33 | 24.3375 | 160.3536 | 114.7007 | 146.2622 | 97.728 | 62.749 | 1.6528 | 41.6027 | 68.7614 | 25.1709 | 2.594 | 41.5198 | 107.7035 | 16.0059 | 20.5234 | 4.3259 | 176.4649 | 139.8302 | 94.1773 | 125.7388 | 77.2046
2012-10-30 15:33 | 24.3495 | 162.1677 | 116.0244 | 148.1162 | 98.9913 | 62.9489 | 1.6525 | 41.6011 | 68.7443 | 25.1752 | 2.5939 | 41.5181 | 107.6931 | 16.0062 | 20.5354 | 4.3397 | 176.4374 | 141.6323 | 95.489 | 127.5808 | 78.4559
2012-10-30 15:33 | 24.3428 | 163.9338 | 117.4819 | 149.8724 | 100.3683 | 63.1188 | 1.6521 | 41.5988 | 68.7256 | 25.1793 | 2.5937 | 41.5158 | 107.6795 | 16.0064 | 20.5047 | 4.3514 | 176.4051 | 143.4291 | 96.9772 | 129.3677 | 79.8636
2012-10-30 15:33 | 24.3435 | 165.5648 | 118.9096 | 151.8039 | 101.5939 | 63.2911 | 1.6517 | 41.5973 | 68.7076 | 25.184 | 2.5936 | 41.5142 | 107.6704 | 16.0065 | 20.5254 | 4.3633 | 176.378 | 145.0394 | 98.3842 | 131.2785 | 81.0685
2012-10-30 15:33 | 24.345 | 167.4823 | 120.411 | 153.986 | 103.0018 | 63.4876 | 1.6514 | 41.5956 | 68.6903 | 25.1884 | 2.5934 | 41.5123 | 107.6593 | 16.0067 | 20.5109 | 4.3768 | 176.3496 | 146.9714 | 99.9001 | 133.4751 | 82.4909
2012-10-30 15:33 | 24.3472 | 169.3047 | 121.899 | 156.0513 | 104.4069 | 63.7066 | 1.6511 | 41.5941 | 68.6742 | 25.1924 | 2.5933 | 41.5106 | 107.6491 | 16.0069 | 20.5051 | 4.3919 | 176.3233 | 148.7996 | 101.3939 | 135.5462 | 83.9018
2012-10-30 15:34 | 24.3532 | 171.0311 | 123.2965 | 157.801 | 105.7441 | 63.8834 | 1.6508 | 41.5947 | 68.6655 | 25.1963 | 2.5933 | 41.511 | 107.6507 | 16.007 | 20.5071 | 4.4041 | 176.3162 | 150.524 | 102.7894 | 137.2939 | 85.237
2012-10-30 15:34 | 24.3472 | 172.8912 | 124.6754 | 159.6258 | 106.9788 | 64.0741 | 1.6506 | 41.5954 | 68.6565 | 25.2005 | 2.5933 | 41.5116 | 107.6522 | 16.0072 | 20.5091 | 4.4173 | 176.3086 | 152.3821 | 104.1663 | 139.1167 | 86.4697
2012-10-30 15:34 | 24.3405 | 174.8186 | 126.2107 | 161.8673 | 108.4968 | 64.2862 | 1.6503 | 41.5964 | 68.6476 | 25.205 | 2.5933 | 41.5126 | 107.6553 | 16.0075 | 20.4864 | 4.4319 | 176.3029 | 154.3322 | 105.7243 | 141.3809 | 88.0104
2012-10-30 15:34 | 24.3375 | 176.6417 | 127.8041 | 163.9207 | 109.9358 | 64.4804 | 1.6501 | 41.599 | 68.6437 | 25.2096 | 2.5934 | 41.515 | 107.6663 | 16.0077 | 20.4994 | 4.4453 | 176.31 | 156.1423 | 107.3047 | 143.4213 | 89.4364
2012-10-30 15:34 | 24.3323 | 178.139 | 129.0381 | 165.5575 | 111.0966 | 64.6347 | 1.65 | 41.6018 | 68.642 | 25.2135 | 2.5936 | 41.5176 | 107.6795 | 16.0078 | 20.4461 | 4.4559 | 176.3215 | 157.6929 | 108.592 | 145.1114 | 90.6505
2012-10-30 15:34 | 24.3368 | 179.78 | 130.3819 | 167.3767 | 112.3546 | 64.822 | 1.6498 | 41.6042 | 68.64 | 25.2172 | 2.5937 | 41.52 | 107.6902 | 16.0081 | 20.5027 | 4.4688 | 176.3302 | 159.2773 | 109.8792 | 146.874 | 91.8519
2012-10-30 15:34 | 24.3338 | 181.4105 | 131.6656 | 168.9872 | 113.651 | 64.9746 | 1.6498 | 41.6102 | 68.6489 | 25.2213 | 2.594 | 41.526 | 107.72 | 16.0083 | 20.4957 | 4.4793 | 176.3689 | 160.9148 | 111.1699 | 148.4915 | 93.1553
2012-10-30 15:34 | 24.3368 | 183.1421 | 133.0976 | 170.722 | 114.8478 | 65.1416 | 1.6497 | 41.6155 | 68.6547 | 25.2255 | 2.5943 | 41.5312 | 107.745 | 16.0086 | 20.4707 | 4.4909 | 176.3997 | 162.6714 | 112.6269 | 150.2513 | 94.3771
2012-10-30 15:34 | 24.3338 | 184.8764 | 134.5133 | 172.581 | 116.077 | 65.3069 | 1.6497 | 41.6205 | 68.6617 | 25.229 | 2.5946 | 41.5361 | 107.769 | 16.0087 | 20.4997 | 4.5023 | 176.4306 | 164.3767 | 114.0136 | 152.0813 | 95.5773
2012-10-30 15:34 | 24.3383 | 186.3484 | 135.7041 | 174.1246 | 117.2422 | 65.467 | 1.6496 | 41.6231 | 68.6605 | 25.2326 | 2.5947 | 41.5385 | 107.78 | 16.009 | 20.4762 | 4.5133 | 176.4405 | 165.8722 | 115.2279 | 153.6484 | 96.766
2012-10-30 15:34 | 24.3353 | 187.7074 | 136.8843 | 175.6774 | 118.3282 | 65.6074 | 1.6495 | 41.6263 | 68.6625 | 25.2358 | 2.5949 | 41.5416 | 107.7953 | 16.0091 | 20.4852 | 4.523 | 176.4578 | 167.2222 | 116.3991 | 155.1922 | 97.843
2012-10-30 15:34 | 24.33 | 189.3082 | 138.1742 | 177.3423 | 119.3934 | 65.7721 | 1.6494 | 41.6293 | 68.663 | 25.2392 | 2.595 | 41.5445 | 107.8096 | 16.0092 | 20.4599 | 4.5343 | 176.4726 | 168.8483 | 117.7143 | 156.8824 | 98.9335
2012-10-30 15:34 | 24.336 | 190.7087 | 139.369 | 179.0274 | 120.554 | 65.8969 | 1.6492 | 41.6303 | 68.6554 | 25.2432 | 2.595 | 41.5454 | 107.8122 | 16.0095 | 20.4699 | 4.5429 | 176.4677 | 170.2388 | 118.8991 | 158.5575 | 100.0841
2012-10-30 15:34 | 24.3368 | 192.1478 | 140.3904 | 180.4928 | 121.6871 | 66.0599 | 1.6491 | 41.634 | 68.6593 | 25.2463 | 2.5952 | 41.5489 | 107.8296 | 16.0096 | 20.4626 | 4.5542 | 176.4889 | 171.6852 | 119.9278 | 160.0302 | 101.2245
2012-10-30 15:34 | 24.3405 | 193.5233 | 141.5652 | 181.9894 | 122.8047 | 66.2021 | 1.649 | 41.6369 | 68.6601 | 25.2495 | 2.5954 | 41.5517 | 107.8426 | 16.0099 | 20.4463 | 4.564 | 176.5027 | 173.077 | 121.1189 | 161.5431 | 102.3584
2012-10-30 15:34 | 24.339 | 194.7721 | 142.6318 | 183.3678 | 123.8111 | 66.3177 | 1.649 | 41.6413 | 68.6658 | 25.2527 | 2.5956 | 41.5562 | 107.865 | 16.01 | 20.5049 | 4.5719 | 176.5308 | 174.2672 | 122.1269 | 162.8629 | 103.3062
2012-10-30 15:34 | 24.3435 | 196.3694 | 143.982 | 184.9371 | 124.8772 | 66.4673 | 1.649 | 41.6483 | 68.6784 | 25.2566 | 2.596 | 41.563 | 107.8992 | 16.0101 | 20.4533 | 4.5823 | 176.5777 | 175.9161 | 123.5287 | 164.4838 | 104.4239
2012-10-30 15:34 | 24.3472 | 197.8444 | 145.1785 | 186.683 | 126.0574 | 66.6453 | 1.6491 | 41.6552 | 68.6923 | 25.2599 | 2.5964 | 41.5696 | 107.9325 | 16.0103 | 20.5091 | 4.5945 | 176.6248 | 177.3353 | 124.6694 | 166.1739 | 105.5483
2012-10-30 15:34 | 24.3398 | 199.0803 | 146.1682 | 188.1203 | 127.1052 | 66.7748 | 1.6491 | 41.6606 | 68.7009 | 25.2632 | 2.5967 | 41.5751 | 107.9588 | 16.0107 | 20.4857 | 4.6035 | 176.6597 | 178.5946 | 125.6825 | 167.6346 | 106.6195
2012-10-30 15:34 | 24.3398 | 200.4262 | 147.2662 | 189.5458 | 128.0847 | 66.9112 | 1.649 | 41.6644 | 68.7052 | 25.2662 | 2.5969 | 41.5786 | 107.9767 | 16.0107 | 20.4977 | 4.6129 | 176.6819 | 179.9285 | 126.7685 | 169.0481 | 107.587
2012-10-30 15:34 | 24.3443 | 201.8279 | 148.404 | 190.941 | 129.1112 | 67.0268 | 1.6488 | 41.6632 | 68.6926 | 25.2694 | 2.5968 | 41.5773 | 107.9692 | 16.0108 | 20.5142 | 4.6208 | 176.6618 | 181.3137 | 127.8898 | 170.4268 | 108.597
2012-10-30 15:34 | 24.3495 | 203.3999 | 149.7093 | 192.6321 | 130.2801 | 67.1828 | 1.6485 | 41.6626 | 68.6814 | 25.2728 | 2.5968 | 41.5767 | 107.9642 | 16.011 | 20.5234 | 4.6316 | 176.6456 | 182.8765 | 129.1859 | 172.1087 | 109.7567
2012-10-30 15:34 | 24.3375 | 204.7698 | 150.8554 | 194.0943 | 131.3065 | 67.3238 | 1.6482 | 41.6604 | 68.6653 | 25.2761 | 2.5966 | 41.5743 | 107.9517 | 16.0111 | 20.4994 | 4.6413 | 176.6171 | 184.2704 | 130.356 | 173.5949 | 110.8071
2012-10-30 15:34 | 24.3532 | 205.9982 | 151.833 | 195.388 | 132.1968 | 67.444 | 1.6478 | 41.6556 | 68.6421 | 25.2788 | 2.5962 | 41.5695 | 107.9245 | 16.0114 | 20.5071 | 4.6496 | 176.5666 | 185.4911 | 131.3259 | 174.8809 | 111.6897
2012-10-30 15:34 | 24.3472 | 207.268 | 152.9092 | 196.7232 | 133.2341 | 67.566 | 1.6473 | 41.6479 | 68.6083 | 25.2819 | 2.5957 | 41.5618 | 107.8836 | 16.0115 | 20.5091 | 4.658 | 176.492 | 186.7589 | 132.4001 | 176.2141 | 112.725
2012-10-30 15:35 | 24.3547 | 208.5789 | 153.9763 | 198.2099 | 134.2435 | 67.7139 | 1.6465 | 41.631 | 68.5435 | 25.2852 | 2.5947 | 41.545 | 107.7958 | 16.0117 | 20.5366 | 4.6682 | 176.3393 | 188.0423 | 133.4397 | 177.6733 | 113.7069
2012-10-30 15:35 | 24.3607 | 210.0695 | 155.0813 | 199.7977 | 135.4164 | 67.8422 | 1.6456 | 41.6144 | 68.4807 | 25.2882 | 2.5936 | 41.5281 | 107.7069 | 16.0118 | 20.4865 | 4.677 | 176.1876 | 189.583 | 134.5948 | 179.3112 | 114.9299
2012-10-30 15:35 | 24.3562 | 211.4669 | 156.2579 | 201.3006 | 136.4647 | 67.9671 | 1.6453 | 41.612 | 68.4657 | 25.2909 | 2.5934 | 41.5256 | 107.6926 | 16.012 | 20.5101 | 4.6857 | 176.1583 | 190.9568 | 135.7478 | 180.7905 | 115.9546
2012-10-30 15:35 | 24.3682 | 212.8415 | 157.4439 | 202.8517 | 137.5417 | 68.1011 | 1.6453 | 41.616 | 68.4709 | 25.2939 | 2.5937 | 41.5298 | 107.7138 | 16.0121 | 20.5181 | 4.6949 | 176.1847 | 192.3234 | 136.9258 | 182.3336 | 117.0236
2012-10-30 15:35 | 24.3667 | 214.1473 | 158.4931 | 204.2517 | 138.5253 | 68.2271 | 1.6451 | 41.6167 | 68.464 | 25.2972 | 2.5936 | 41.5301 | 107.7143 | 16.0123 | 20.5006 | 4.7036 | 176.1783 | 193.6467 | 137.9925 | 183.7511 | 118.0247
2012-10-30 15:35 | 24.3667 | 215.5058 | 159.6647 | 205.7479 | 139.5608 | 68.3728 | 1.6449 | 41.6168 | 68.4559 | 25.3004 | 2.5936 | 41.5302 | 107.7133 | 16.0125 | 20.5206 | 4.7136 | 176.1692 | 194.9852 | 139.1441 | 185.2273 | 119.0402
2012-10-30 15:35 | 24.3659 | 217.0052 | 160.8751 | 207.2512 | 140.6115 | 68.5132 | 1.6447 | 41.6168 | 68.4475 | 25.3034 | 2.5936 | 41.5301 | 107.7121 | 16.0126 | 20.4837 | 4.7233 | 176.1597 | 196.5215 | 140.3914 | 186.7675 | 120.1278
2012-10-30 15:35 | 24.3757 | 218.1724 | 161.8559 | 208.5133 | 141.5735 | 68.6363 | 1.6445 | 41.6161 | 68.4378 | 25.3061 | 2.5935 | 41.5293 | 107.707 | 16.0127 | 20.4815 | 4.7318 | 176.1448 | 197.6909 | 141.3744 | 188.0318 | 121.092
2012-10-30 15:35 | 24.3712 | 219.2111 | 162.613 | 209.5956 | 142.4219 | 68.7184 | 1.6443 | 41.6154 | 68.4286 | 25.3087 | 2.5934 | 41.5286 | 107.7022 | 16.0129 | 20.4329 | 4.7374 | 176.1308 | 198.7782 | 142.1801 | 189.1627 | 121.989
2012-10-30 15:35 | 24.3749 | 220.3996 | 163.5847 | 210.765 | 143.2481 | 68.8507 | 1.6441 | 41.616 | 68.4227 | 25.3117 | 2.5935 | 41.5291 | 107.7039 | 16.013 | 20.4887 | 4.7466 | 176.1266 | 199.9109 | 143.096 | 190.2763 | 122.7594
2012-10-30 15:35 | 24.3839 | 221.7507 | 164.7762 | 212.2429 | 144.3224 | 68.9727 | 1.6439 | 41.6148 | 68.4101 | 25.3149 | 2.5933 | 41.5278 | 107.6956 | 16.0133 | 20.5017 | 4.755 | 176.1057 | 201.249 | 144.2745 | 191.7412 | 123.8207
2012-10-30 15:35 | 24.3787 | 223.048 | 165.7946 | 213.7065 | 145.3562 | 69.0854 | 1.6437 | 41.6154 | 68.4045 | 25.3176 | 2.5934 | 41.5283 | 107.698 | 16.0133 | 20.5366 | 4.7627 | 176.1025 | 202.5114 | 145.258 | 193.1699 | 124.8196
2012-10-30 15:35 | 24.3787 | 224.193 | 166.694 | 214.9114 | 146.2688 | 69.2125 | 1.6437 | 41.6179 | 68.4056 | 25.3203 | 2.5935 | 41.5308 | 107.7093 | 16.0135 | 20.5086 | 4.7715 | 176.1149 | 203.6844 | 146.1854 | 194.4028 | 125.7602
2012-10-30 15:35 | 24.3839 | 225.3075 | 167.7355 | 216.1609 | 147.133 | 69.3229 | 1.6436 | 41.6206 | 68.4077 | 25.3228 | 2.5936 | 41.5333 | 107.7214 | 16.0137 | 20.5258 | 4.7791 | 176.1291 | 204.7817 | 147.2097 | 195.6351 | 126.6072
2012-10-30 15:35 | 24.3824 | 226.5764 | 168.7824 | 217.5295 | 148.1144 | 69.4529 | 1.6435 | 41.6221 | 68.4044 | 25.3259 | 2.5937 | 41.5347 | 107.7271 | 16.0139 | 20.4882 | 4.7881 | 176.1315 | 206.0882 | 148.2942 | 197.0413 | 127.6262
2012-10-30 15:35 | 24.3929 | 227.8571 | 169.8691 | 218.9022 | 149.1662 | 69.5806 | 1.6432 | 41.621 | 68.3934 | 25.3285 | 2.5936 | 41.5333 | 107.719 | 16.014 | 20.5147 | 4.7969 | 176.1124 | 207.3424 | 149.3544 | 198.3875 | 128.6515
2012-10-30 15:35 | 24.3914 | 229.0389 | 170.8234 | 220.0383 | 149.921 | 69.6673 | 1.643 | 41.6197 | 68.382 | 25.3312 | 2.5935 | 41.5321 | 107.7115 | 16.0142 | 20.5092 | 4.8029 | 176.0935 | 208.5297 | 150.3142 | 199.5291 | 129.4118
2012-10-30 15:35 | 24.3966 | 230.1683 | 171.7263 | 221.1691 | 150.7601 | 69.7794 | 1.6428 | 41.6175 | 68.368 | 25.3337 | 2.5933 | 41.5298 | 107.6991 | 16.0143 | 20.4984 | 4.8106 | 176.0671 | 209.6699 | 151.2279 | 200.6707 | 130.2617
2012-10-30 15:35 | 24.3959 | 231.2915 | 172.6782 | 222.3963 | 151.6516 | 69.8881 | 1.6425 | 41.6163 | 68.3566 | 25.3365 | 2.5932 | 41.5284 | 107.6914 | 16.0144 | 20.4897 | 4.8181 | 176.048 | 210.8018 | 152.1885 | 201.9066 | 131.1619
2012-10-30 15:35 | 24.4011 | 232.5539 | 173.6756 | 223.7118 | 152.6306 | 70.0019 | 1.6422 | 41.6126 | 68.3368 | 25.3393 | 2.5929 | 41.5246 | 107.6697 | 16.0147 | 20.5189 | 4.8259 | 176.0065 | 212.035 | 153.1567 | 203.1929 | 132.1117
2012-10-30 15:35 | 24.4011 | 233.6066 | 174.5939 | 224.8802 | 153.5122 | 70.1071 | 1.642 | 41.6097 | 68.3215 | 25.3415 | 2.5927 | 41.5219 | 107.6546 | 16.0148 | 20.5149 | 4.8332 | 175.976 | 213.0917 | 154.079 | 204.3653 | 132.9973
2012-10-30 15:35 | 24.4033 | 234.4962 | 175.4443 | 225.9641 | 154.1952 | 70.2007 | 1.6417 | 41.6068 | 68.3059 | 25.3437 | 2.5925 | 41.5188 | 107.6378 | 16.0149 | 20.5051 | 4.8396 | 175.9437 | 213.9911 | 154.9392 | 205.459 | 133.6901
2012-10-30 15:35 | 24.4071 | 235.5365 | 176.2654 | 227.073 | 154.9032 | 70.314 | 1.6414 | 41.6027 | 68.2865 | 25.3459 | 2.5922 | 41.5146 | 107.6148 | 16.0151 | 20.4929 | 4.8474 | 175.9013 | 215.0436 | 155.7725 | 206.5801 | 134.4103
2012-10-30 15:36 | 24.4063 | 236.7172 | 177.1995 | 228.2558 | 155.7967 | 70.4007 | 1.6411 | 41.599 | 68.2679 | 25.3484 | 2.5919 | 41.5108 | 107.5938 | 16.0153 | 20.5241 | 4.8534 | 175.8617 | 216.1931 | 156.6754 | 207.7317 | 135.2726
2012-10-30 15:36 | 24.4048 | 237.8536 | 178.1958 | 229.41 | 156.7365 | 70.5059 | 1.6408 | 41.5963 | 68.2518 | 25.351 | 2.5918 | 41.508 | 107.5787 | 16.0154 | 20.5347 | 4.8607 | 175.8306 | 217.3189 | 157.6611 | 208.8753 | 136.2018
2012-10-30 15:36 | 24.4101 | 238.8435 | 179.0153 | 230.5512 | 157.4548 | 70.607 | 1.6406 | 41.5948 | 68.2415 | 25.353 | 2.5916 | 41.5065 | 107.57 | 16.0155 | 20.544 | 4.8676 | 175.8115 | 218.2995 | 158.4713 | 210.0072 | 136.9108
2012-10-30 15:36 | 24.4041 | 239.7278 | 179.7533 | 231.5551 | 158.1078 | 70.7041 | 1.6404 | 41.5933 | 68.2305 | 25.3553 | 2.5915 | 41.5049 | 107.5609 | 16.0157 | 20.5219 | 4.8743 | 175.7914 | 219.2059 | 159.2314 | 211.0332 | 137.5859
2012-10-30 15:36 | 24.4071 | 240.7306 | 180.5788 | 232.501 | 158.8903 | 70.8018 | 1.6403 | 41.5928 | 68.2235 | 25.3572 | 2.5915 | 41.5042 | 107.5568 | 16.0157 | 20.5089 | 4.8811 | 175.7803 | 220.2217 | 160.0699 | 211.9921 | 138.3814
2012-10-30 15:36 | 24.4146 | 241.8434 | 181.5188 | 233.7104 | 159.7436 | 70.9006 | 1.64 | 41.5904 | 68.2088 | 25.3598 | 2.5913 | 41.5019 | 107.5428 | 16.016 | 20.5565 | 4.8879 | 175.7517 | 221.2869 | 160.9623 | 213.1539 | 139.1871
2012-10-30 15:36 | 24.4026 | 242.8541 | 182.3847 | 234.896 | 160.5041 | 71.0012 | 1.6398 | 41.5894 | 68.1993 | 25.3622 | 2.5912 | 41.5009 | 107.5363 | 16.0162 | 20.5164 | 4.8948 | 175.7356 | 222.3377 | 161.8683 | 214.3796 | 139.9877
2012-10-30 15:36 | 24.4161 | 243.7921 | 183.1312 | 235.8831 | 161.3291 | 71.1081 | 1.6396 | 41.5865 | 68.1835 | 25.3644 | 2.591 | 41.498 | 107.5208 | 16.0163 | 20.566 | 4.9022 | 175.7043 | 223.2261 | 162.5652 | 215.3171 | 140.7631
2012-10-30 15:36 | 24.4123 | 244.7047 | 183.9548 | 236.8762 | 162.0518 | 71.1844 | 1.6396 | 41.5906 | 68.191 | 25.3667 | 2.5912 | 41.5016 | 107.5379 | 16.0165 | 20.5301 | 4.9075 | 175.7288 | 224.1746 | 163.4247 | 216.3461 | 141.5217
2012-10-30 15:36 | 24.4183 | 245.8735 | 184.8982 | 238.0946 | 162.8856 | 71.2873 | 1.6396 | 41.5957 | 68.2006 | 25.3693 | 2.5915 | 41.5068 | 107.5637 | 16.0167 | 20.5361 | 4.9145 | 175.7643 | 225.3374 | 164.3621 | 217.5585 | 142.3495
2012-10-30 15:36 | 24.4213 | 247.0584 | 185.8938 | 239.3175 | 163.7595 | 71.3878 | 1.6394 | 41.5954 | 68.1926 | 25.3719 | 2.5914 | 41.5064 | 107.5594 | 16.017 | 20.5351 | 4.9215 | 175.752 | 226.5233 | 165.3587 | 218.7824 | 143.2244
2012-10-30 15:36 | 24.4325 | 247.9303 | 186.6545 | 240.2698 | 164.4786 | 71.4953 | 1.6394 | 41.5984 | 68.1976 | 25.3738 | 2.5916 | 41.5093 | 107.5744 | 16.017 | 20.4862 | 4.9289 | 175.772 | 227.4441 | 166.1683 | 219.7836 | 143.9924
2012-10-30 15:36 | 24.4288 | 248.8105 | 187.373 | 241.1835 | 165.218 | 71.5652 | 1.6392 | 41.5954 | 68.1825 | 25.3757 | 2.5913 | 41.5064 | 107.5572 | 16.0173 | 20.5346 | 4.9337 | 175.7398 | 228.2759 | 166.8384 | 220.6489 | 144.6834
2012-10-30 15:36 | 24.4273 | 249.7005 | 188.1489 | 242.0599 | 165.8082 | 71.6444 | 1.6391 | 41.5953 | 68.1767 | 25.3776 | 2.5913 | 41.5061 | 107.5554 | 16.0174 | 20.5131 | 4.9392 | 175.7321 | 229.1874 | 167.6358 | 221.5468 | 145.2951
2012-10-30 15:36 | 24.4392 | 250.7091 | 189.0286 | 243.2316 | 166.7078 | 71.7369 | 1.6388 | 41.5926 | 68.1619 | 25.3799 | 2.5911 | 41.5034 | 107.5402 | 16.0175 | 20.521 | 4.9455 | 175.702 | 230.1881 | 168.5076 | 222.7106 | 146.1868
2012-10-30 15:36 | 24.4288 | 251.6873 | 189.7801 | 244.1299 | 167.4141 | 71.8357 | 1.6386 | 41.5899 | 68.1474 | 25.3821 | 2.5909 | 41.5007 | 107.5242 | 16.0179 | 20.5106 | 4.9524 | 175.6716 | 231.1767 | 169.2695 | 223.6193 | 146.9035
2012-10-30 15:36 | 24.4325 | 252.4332 | 190.4556 | 244.9559 | 168.0412 | 71.8999 | 1.6383 | 41.5878 | 68.1348 | 25.3841 | 2.5907 | 41.4985 | 107.512 | 16.018 | 20.5103 | 4.9568 | 175.6468 | 231.9229 | 169.9453 | 224.4456 | 147.5309
2012-10-30 15:36 | 24.443 | 253.3185 | 191.1654 | 245.9999 | 168.7884 | 71.9964 | 1.6382 | 41.5863 | 68.1254 | 25.3858 | 2.5906 | 41.4968 | 107.5026 | 16.0181 | 20.4927 | 4.9634 | 175.628 | 232.8258 | 170.6727 | 225.5072 | 148.2957
2012-10-30 15:36 | 24.4437 | 254.2132 | 191.897 | 246.8851 | 169.3583 | 72.0929 | 1.638 | 41.5859 | 68.1188 | 25.3878 | 2.5905 | 41.4962 | 107.4979 | 16.0183 | 20.4934 | 4.9701 | 175.6167 | 233.7198 | 171.4036 | 226.3917 | 148.8649
2012-10-30 15:36 | 24.4452 | 255.0773 | 192.6964 | 247.7239 | 170.0111 | 72.1807 | 1.638 | 41.5897 | 68.1255 | 25.3899 | 2.5908 | 41.5 | 107.5162 | 16.0185 | 20.527 | 4.9761 | 175.6417 | 234.5503 | 172.1694 | 227.1969 | 149.4841
2012-10-30 15:36 | 24.44 | 255.9421 | 193.3987 | 248.6614 | 170.7354 | 72.2426 | 1.6378 | 41.5875 | 68.1126 | 25.3921 | 2.5906 | 41.4977 | 107.503 | 16.0187 | 20.5097 | 4.9804 | 175.6155 | 235.4324 | 172.889 | 228.1517 | 150.2257
2012-10-30 15:36 | 24.4385 | 256.7831 | 194.0891 | 249.524 | 171.37 | 72.3212 | 1.6373 | 41.5775 | 68.0756 | 25.3937 | 2.59 | 41.4881 | 107.4524 | 16.0188 | 20.5163 | 4.9858 | 175.528 | 236.2668 | 173.5728 | 229.0077 | 150.8537
2012-10-30 15:36 | 24.4415 | 257.6205 | 194.7645 | 250.4222 | 171.982 | 72.4148 | 1.6369 | 41.5701 | 68.0465 | 25.3955 | 2.5895 | 41.4805 | 107.4121 | 16.0189 | 20.4952 | 4.9923 | 175.4586 | 237.1253 | 174.2693 | 229.927 | 151.4868
2012-10-30 15:36 | 24.449 | 258.4544 | 195.5113 | 251.2426 | 172.747 | 72.4812 | 1.6364 | 41.5616 | 68.0134 | 25.3974 | 2.5889 | 41.4715 | 107.3643 | 16.0192 | 20.5508 | 4.9969 | 175.3777 | 237.9036 | 174.9605 | 230.6918 | 152.1962
2012-10-30 15:36 | 24.443 | 259.3564 | 196.3352 | 252.1938 | 173.6628 | 72.5668 | 1.6361 | 41.5555 | 67.9879 | 25.3995 | 2.5885 | 41.4656 | 107.3336 | 16.0192 | 20.5248 | 5.0028 | 175.3215 | 238.8316 | 175.8104 | 231.669 | 153.138
2012-10-30 15:37 | 24.4377 | 260.247 | 197.1441 | 253.2468 | 174.4624 | 72.6697 | 1.6357 | 41.5495 | 67.9636 | 25.4012 | 2.5881 | 41.4597 | 107.3017 | 16.0194 | 20.5475 | 5.0098 | 175.2654 | 239.6995 | 176.5966 | 232.6993 | 153.9149
2012-10-30 15:37 | 24.4557 | 261.0244 | 197.827 | 254.1433 | 174.9025 | 72.7621 | 1.6354 | 41.5435 | 67.9396 | 25.4029 | 2.5877 | 41.4535 | 107.2677 | 16.0197 | 20.5896 | 5.0162 | 175.2073 | 240.4348 | 177.2374 | 233.5537 | 154.3129
2012-10-30 15:37 | 24.446 | 261.8673 | 198.5412 | 254.8595 | 175.3385 | 72.8326 | 1.6351 | 41.5387 | 67.9192 | 25.4046 | 2.5874 | 41.4489 | 107.2435 | 16.0198 | 20.5799 | 5.0211 | 175.1627 | 241.2874 | 177.9613 | 234.2796 | 154.7586
2012-10-30 15:37 | 24.4609 | 262.645 | 199.2289 | 255.7094 | 175.979 | 72.8991 | 1.6348 | 41.5342 | 67.9005 | 25.4062 | 2.587 | 41.4442 | 107.2181 | 16.0199 | 20.5908 | 5.0257 | 175.1186 | 242.0542 | 178.6381 | 235.1186 | 155.3882
2012-10-30 15:37 | 24.4557 | 263.4179 | 199.8258 | 256.5273 | 176.6699 | 72.9892 | 1.6344 | 41.5269 | 67.872 | 25.4079 | 2.5865 | 41.4368 | 107.1782 | 16.0201 | 20.5936 | 5.0319 | 175.0502 | 242.8243 | 179.2322 | 235.9337 | 156.0763
2012-10-30 15:37 | 24.4527 | 264.0061 | 200.3935 | 257.2691 | 177.2537 | 73.0418 | 1.634 | 41.5204 | 67.8463 | 25.4095 | 2.5861 | 41.4302 | 107.1436 | 16.0202 | 20.6066 | 5.0355 | 174.9899 | 243.3995 | 179.7869 | 236.6625 | 156.6471
2012-10-30 15:37 | 24.4617 | 264.6801 | 200.9259 | 257.9799 | 177.7675 | 73.1239 | 1.6338 | 41.5166 | 67.8305 | 25.4108 | 2.5859 | 41.4263 | 107.1226 | 16.0203 | 20.6036 | 5.0412 | 174.953 | 244.0765 | 180.3223 | 237.3763 | 157.1639
2012-10-30 15:37 | 24.4557 | 265.44 | 201.5851 | 258.6989 | 178.3054 | 73.1933 | 1.6337 | 41.5167 | 67.8267 | 25.4124 | 2.5858 | 41.4265 | 107.1223 | 16.0205 | 20.5856 | 5.0459 | 174.9491 | 244.8544 | 180.9995 | 238.1133 | 157.7198
2012-10-30 15:37 | 24.4512 | 266.2322 | 202.2537 | 259.5516 | 178.8372 | 73.2522 | 1.6336 | 41.5169 | 67.8226 | 25.4141 | 2.5858 | 41.4265 | 107.121 | 16.0207 | 20.6011 | 5.05 | 174.9436 | 245.6311 | 181.6526 | 238.9505 | 158.2361
2012-10-30 15:37 | 24.4527 | 266.902 | 202.8693 | 260.3399 | 179.4299 | 73.3291 | 1.6332 | 41.5097 | 67.7948 | 25.4158 | 2.5854 | 41.4192 | 107.0832 | 16.0207 | 20.5906 | 5.0553 | 174.8779 | 246.3114 | 182.2787 | 239.7493 | 158.8393
2012-10-30 15:37 | 24.4572 | 267.5979 | 203.4013 | 261.0609 | 179.92 | 73.4002 | 1.633 | 41.5071 | 67.7821 | 25.4173 | 2.5852 | 41.4166 | 107.0684 | 16.0209 | 20.6031 | 5.0602 | 174.8505 | 246.9948 | 182.7982 | 240.4578 | 159.3169
2012-10-30 15:37 | 24.449 | 268.2188 | 203.9442 | 261.7963 | 180.4014 | 73.4504 | 1.6327 | 41.502 | 67.7618 | 25.4187 | 2.5848 | 41.4119 | 107.0434 | 16.0211 | 20.6189 | 5.0637 | 174.8051 | 247.5999 | 183.3253 | 241.1774 | 159.7825
2012-10-30 15:37 | 24.4549 | 268.8653 | 204.521 | 262.4016 | 180.94 | 73.5192 | 1.6324 | 41.4945 | 67.7341 | 25.4199 | 2.5843 | 41.4039 | 107.0017 | 16.0211 | 20.5527 | 5.0684 | 174.7358 | 248.3126 | 183.9683 | 241.8489 | 160.3873
2012-10-30 15:37 | 24.4594 | 269.7586 | 205.2657 | 263.301 | 181.6028 | 73.5903 | 1.6319 | 41.4871 | 67.7048 | 25.4218 | 2.5838 | 41.3963 | 106.9616 | 16.0212 | 20.5612 | 5.0733 | 174.6663 | 249.1974 | 184.7045 | 242.7398 | 161.0416
2012-10-30 15:37 | 24.4512 | 270.5109 | 205.9111 | 264.1174 | 182.202 | 73.688 | 1.6315 | 41.4775 | 67.6694 | 25.4233 | 2.5832 | 41.3869 | 106.9121 | 16.0214 | 20.545 | 5.0801 | 174.5814 | 249.9659 | 185.3661 | 243.5724 | 161.657
2012-10-30 15:37 | 24.4647 | 271.1566 | 206.4877 | 264.862 | 182.7211 | 73.7411 | 1.631 | 41.4688 | 67.6369 | 25.4249 | 2.5827 | 41.3782 | 106.8665 | 16.0214 | 20.5384 | 5.0837 | 174.5034 | 250.6182 | 185.9493 | 244.3236 | 162.1827
2012-10-30 15:37 | 24.4662 | 271.7631 | 206.9578 | 265.4971 | 183.2048 | 73.8134 | 1.6308 | 41.4638 | 67.6173 | 25.4261 | 2.5823 | 41.3731 | 106.8394 | 16.0216 | 20.568 | 5.0887 | 174.4567 | 251.1951 | 186.3898 | 244.9291 | 162.6368
2012-10-30 15:37 | 24.4654 | 272.4214 | 207.5516 | 266.1609 | 183.7255 | 73.8694 | 1.6305 | 41.4604 | 67.6014 | 25.4279 | 2.5821 | 41.3696 | 106.8199 | 16.0218 | 20.5351 | 5.0926 | 174.4213 | 251.8863 | 187.0165 | 245.6258 | 163.1904
2012-10-30 15:37 | 24.4654 | 273.1849 | 208.1698 | 267.0041 | 184.3255 | 73.9503 | 1.6303 | 41.458 | 67.59 | 25.4293 | 2.5819 | 41.3672 | 106.8074 | 16.0218 | 20.5792 | 5.0981 | 174.3973 | 252.6057 | 187.5906 | 246.4249 | 163.7463
2012-10-30 15:37 | 24.4609 | 273.8895 | 208.7678 | 267.7259 | 184.8623 | 74.011 | 1.6301 | 41.4537 | 67.5718 | 25.4309 | 2.5816 | 41.363 | 106.784 | 16.0221 | 20.6108 | 5.1023 | 174.3558 | 253.2787 | 188.157 | 247.1151 | 164.2515
2012-10-30 15:37 | 24.4744 | 274.546 | 209.3444 | 268.3762 | 185.374 | 74.0827 | 1.6299 | 41.4517 | 67.5613 | 25.4324 | 2.5815 | 41.3606 | 106.7718 | 16.022 | 20.6123 | 5.1073 | 174.333 | 253.9337 | 188.7321 | 247.7639 | 164.7617
2012-10-30 15:37 | 24.4722 | 275.206 | 209.8699 | 269.034 | 185.9015 | 74.1301 | 1.6298 | 41.4506 | 67.5544 | 25.4337 | 2.5814 | 41.3598 | 106.7663 | 16.0222 | 20.6141 | 5.1105 | 174.3206 | 254.5919 | 189.2558 | 248.4199 | 165.2874
2012-10-30 15:37 | 24.4699 | 275.8116 | 210.3402 | 269.6915 | 186.3858 | 74.2075 | 1.6297 | 41.4509 | 67.5514 | 25.4351 | 2.5814 | 41.3599 | 106.7658 | 16.0224 | 20.6078 | 5.1159 | 174.3172 | 255.2038 | 189.7324 | 249.0837 | 165.778
2012-10-30 15:37 | 24.4751 | 276.4981 | 210.9597 | 270.4145 | 186.9483 | 74.278 | 1.6294 | 41.4467 | 67.5331 | 25.4368 | 2.5811 | 41.3556 | 106.7421 | 16.0226 | 20.609 | 5.1207 | 174.2752 | 255.8891 | 190.3507 | 249.8055 | 166.3393
2012-10-30 15:37 | 24.4766 | 277.1032 | 211.5361 | 271.1181 | 187.5033 | 74.3445 | 1.6292 | 41.4445 | 67.5228 | 25.438 | 2.5809 | 41.3533 | 106.7304 | 16.0226 | 20.6024 | 5.1253 | 174.2532 | 256.5008 | 190.9337 | 250.5157 | 166.9009
2012-10-30 15:38 | 24.4759 | 277.6829 | 212.0671 | 271.6487 | 187.9462 | 74.378 | 1.629 | 41.4397 | 67.5043 | 25.4391 | 2.5806 | 41.3486 | 106.7048 | 16.0227 | 20.6418 | 5.1276 | 174.2092 | 257.0411 | 191.4253 | 251.0069 | 167.3044
2012-10-30 15:38 | 24.4751 | 278.3435 | 212.5782 | 272.3807 | 188.448 | 74.4427 | 1.6289 | 41.4409 | 67.5039 | 25.4407 | 2.5807 | 41.3496 | 106.7091 | 16.0229 | 20.5889 | 5.1321 | 174.2131 | 257.7546 | 191.9893 | 251.7918 | 167.8591
2012-10-30 15:38 | 24.4729 | 279.0373 | 213.2336 | 273.0645 | 188.9602 | 74.5086 | 1.629 | 41.4451 | 67.5127 | 25.4425 | 2.5809 | 41.3537 | 106.7298 | 16.0229 | 20.6108 | 5.1366 | 174.2425 | 258.4265 | 192.6228 | 252.4537 | 168.3494
2012-10-30 15:38 | 24.4722 | 279.5856 | 213.7409 | 273.6993 | 189.4072 | 74.5664 | 1.6291 | 41.4496 | 67.5237 | 25.4439 | 2.5812 | 41.3583 | 106.7522 | 16.0231 | 20.59 | 5.1406 | 174.2759 | 258.9956 | 193.1509 | 253.1093 | 168.8172
2012-10-30 15:38 | 24.4841 | 280.1154 | 214.1742 | 274.2536 | 189.7372 | 74.6271 | 1.6291 | 41.4537 | 67.5339 | 25.4451 | 2.5814 | 41.3623 | 106.7726 | 16.0232 | 20.642 | 5.1448 | 174.3065 | 259.4734 | 193.5322 | 253.6116 | 169.0952
2012-10-30 15:38 | 24.4938 | 280.6738 | 214.6879 | 274.8172 | 190.1515 | 74.7028 | 1.6291 | 41.4532 | 67.5293 | 25.4462 | 2.5813 | 41.3617 | 106.7686 | 16.0234 | 20.6076 | 5.15 | 174.2979 | 260.0662 | 194.0803 | 254.2096 | 169.5439
2012-10-30 15:38 | 24.4938 | 281.2108 | 215.0895 | 275.4058 | 190.5994 | 74.7421 | 1.6289 | 41.4514 | 67.5198 | 25.4477 | 2.5812 | 41.36 | 106.759 | 16.0235 | 20.6196 | 5.1527 | 174.2788 | 260.5912 | 194.4699 | 254.7862 | 169.9798
2012-10-30 15:38 | 24.4849 | 281.7466 | 215.5531 | 276.0008 | 191.0384 | 74.7958 | 1.6287 | 41.4485 | 67.5072 | 25.4487 | 2.581 | 41.357 | 106.7429 | 16.0235 | 20.6468 | 5.1564 | 174.2502 | 261.0998 | 194.9063 | 255.354 | 170.3916
2012-10-30 15:38 | 24.4826 | 282.2309 | 215.984 | 276.552 | 191.4211 | 74.8479 | 1.6285 | 41.4457 | 67.4957 | 25.4497 | 2.5808 | 41.3541 | 106.7272 | 16.0237 | 20.6205 | 5.16 | 174.2229 | 261.6104 | 195.3635 | 255.9315 | 170.8006
2012-10-30 15:38 | 24.4916 | 282.792 | 216.4143 | 277.0292 | 191.8428 | 74.8981 | 1.6284 | 41.4447 | 67.4891 | 25.4509 | 2.5808 | 41.3532 | 106.7224 | 16.0237 | 20.6335 | 5.1635 | 174.2115 | 262.1585 | 195.7808 | 256.3957 | 171.2093
2012-10-30 15:38 | 24.4923 | 283.2521 | 216.7969 | 277.5368 | 192.205 | 74.9629 | 1.6283 | 41.4443 | 67.4853 | 25.452 | 2.5807 | 41.3525 | 106.7173 | 16.0239 | 20.5941 | 5.1679 | 174.2026 | 262.658 | 196.2028 | 256.9427 | 171.6109
2012-10-30 15:38 | 24.4946 | 283.7098 | 217.2441 | 278.0303 | 192.734 | 75.0027 | 1.6282 | 41.4417 | 67.4735 | 25.4532 | 2.5805 | 41.35 | 106.7041 | 16.0239 | 20.5683 | 5.1707 | 174.1775 | 263.1415 | 196.6758 | 257.462 | 172.1657
2012-10-30 15:38 | 24.4886 | 284.2595 | 217.679 | 278.5465 | 193.1879 | 75.0472 | 1.6281 | 41.4435 | 67.4754 | 25.4546 | 2.5806 | 41.3516 | 106.7119 | 16.024 | 20.6104 | 5.1738 | 174.1873 | 263.6491 | 197.0686 | 257.9361 | 172.5775
2012-10-30 15:38 | 24.4946 | 284.8673 | 218.2205 | 279.2252 | 193.7049 | 75.1154 | 1.628 | 41.4412 | 67.4651 | 25.4557 | 2.5804 | 41.3492 | 106.6987 | 16.0242 | 20.6204 | 5.1785 | 174.1639 | 264.2469 | 197.6001 | 258.6048 | 173.0845
2012-10-30 15:38 | 24.4938 | 285.4605 | 218.6645 | 279.8468 | 194.0383 | 75.1744 | 1.6276 | 41.4346 | 67.4406 | 25.4568 | 2.58 | 41.3428 | 106.6648 | 16.0243 | 20.6116 | 5.1825 | 174.1054 | 264.8489 | 198.0529 | 259.2352 | 173.4267
2012-10-30 15:38 | 24.4901 | 285.9235 | 219.145 | 280.3803 | 194.1879 | 75.2148 | 1.6273 | 41.4288 | 67.4186 | 25.458 | 2.5796 | 41.3369 | 106.6337 | 16.0243 | 20.5679 | 5.1853 | 174.0523 | 265.3556 | 198.5771 | 259.8124 | 173.62
2012-10-30 15:38 | 24.4879 | 286.4418 | 219.505 | 280.9345 | 194.5357 | 75.2767 | 1.6271 | 41.4246 | 67.4023 | 25.459 | 2.5794 | 41.3327 | 106.612 | 16.0244 | 20.6097 | 5.1896 | 174.0143 | 265.8321 | 198.8953 | 260.3248 | 173.926
2012-10-30 15:38 | 24.4946 | 286.9843 | 219.9525 | 281.4048 | 194.9985 | 75.3275 | 1.6268 | 41.4192 | 67.3812 | 25.4603 | 2.579 | 41.3271 | 106.583 | 16.0244 | 20.6044 | 5.1931 | 173.9642 | 266.3799 | 199.3481 | 260.8004 | 174.3941
2012-10-30 15:38 | 24.4983 | 287.5314 | 220.4246 | 281.9994 | 195.4741 | 75.3726 | 1.6265 | 41.412 | 67.3553 | 25.4614 | 2.5785 | 41.32 | 106.5452 | 16.0246 | 20.6001 | 5.1962 | 173.9005 | 266.9313 | 199.8245 | 261.3993 | 174.874
2012-10-30 15:38 | 24.5028 | 288.0985 | 220.8857 | 282.6024 | 195.9859 | 75.4397 | 1.6263 | 41.4087 | 67.3416 | 25.4625 | 2.5783 | 41.3168 | 106.5285 | 16.0246 | 20.6166 | 5.2008 | 173.87 | 267.4819 | 200.2691 | 261.9858 | 175.3693
2012-10-30 15:38 | 24.4998 | 288.5694 | 221.3156 | 283.1554 | 196.3448 | 75.4853 | 1.6261 | 41.4073 | 67.3338 | 25.4636 | 2.5782 | 41.3153 | 106.5203 | 16.0247 | 20.5816 | 5.204 | 173.8541 | 267.9878 | 200.734 | 262.5738 | 175.7632
2012-10-30 15:38 | 24.5148 | 289.0466 | 221.6967 | 283.6374 | 196.6626 | 75.5391 | 1.626 | 41.4043 | 67.3222 | 25.4643 | 2.578 | 41.3122 | 106.5041 | 16.0248 | 20.6366 | 5.2077 | 173.8263 | 268.41 | 201.0601 | 263.0008 | 176.026
2012-10-30 15:38 | 24.5118 | 289.5866 | 222.158 | 284.1013 | 197.1553 | 75.5778 | 1.6258 | 41.401 | 67.3077 | 25.4658 | 2.5778 | 41.3089 | 106.4863 | 16.0248 | 20.5976 | 5.2103 | 173.794 | 268.989 | 201.5604 | 263.5037 | 176.5577
2012-10-30 15:39 | 24.5215 | 291.1169 | 223.5214 | 285.7505 | 198.5498 | 75.7448 | 1.6251 | 41.3903 | 67.2642 | 25.4691 | 2.5771 | 41.2981 | 106.4284 | 16.0251 | 20.5912 | 5.2218 | 173.6926 | 270.5257 | 202.9302 | 265.1593 | 177.9586
2012-10-30 15:39 | 24.5335 | 292.5333 | 224.8081 | 287.293 | 199.8288 | 75.8858 | 1.6247 | 41.3842 | 67.2355 | 25.4725 | 2.5767 | 41.2918 | 106.3952 | 16.0253 | 20.6233 | 5.2316 | 173.6308 | 271.91 | 204.1848 | 266.6697 | 179.2055
2012-10-30 15:39 | 24.5268 | 292.9073 | 225.1003 | 287.6871 | 200.0307 | 75.9465 | 1.6246 | 41.384 | 67.2331 | 25.4731 | 2.5766 | 41.2916 | 106.3935 | 16.0254 | 20.6286 | 5.2358 | 173.6266 | 272.2787 | 204.4717 | 267.0585 | 179.4021
2012-10-30 15:39 | 24.5402 | 293.3015 | 225.4795 | 288.0665 | 200.418 | 75.9702 | 1.6245 | 41.3831 | 67.2275 | 25.4741 | 2.5766 | 41.2908 | 106.3882 | 16.0256 | 20.65 | 5.2374 | 173.6157 | 272.6515 | 204.8295 | 267.4165 | 179.768
2012-10-30 15:39 | 24.5342 | 293.6723 | 225.7646 | 288.4689 | 200.6757 | 76.0182 | 1.6244 | 41.3822 | 67.2223 | 25.475 | 2.5765 | 41.2899 | 106.3831 | 16.0256 | 20.5759 | 5.2407 | 173.6054 | 273.0964 | 205.1887 | 267.893 | 180.0998
2012-10-30 15:39 | 24.5484 | 294.0864 | 226.1839 | 288.8491 | 200.989 | 76.058 | 1.6244 | 41.3827 | 67.2217 | 25.4758 | 2.5765 | 41.2903 | 106.3843 | 16.0257 | 20.6181 | 5.2434 | 173.606 | 273.4683 | 205.5658 | 268.231 | 180.3709
2012-10-30 15:39 | 24.5447 | 294.4594 | 226.5145 | 289.2382 | 201.3671 | 76.0887 | 1.6244 | 41.384 | 67.2231 | 25.4769 | 2.5766 | 41.2916 | 106.3905 | 16.0258 | 20.5904 | 5.2456 | 173.6136 | 273.869 | 205.9241 | 268.6478 | 180.7767
2012-10-30 15:39 | 24.5447 | 294.8591 | 226.8213 | 289.7081 | 201.5875 | 76.1251 | 1.6244 | 41.3862 | 67.228 | 25.4777 | 2.5767 | 41.2937 | 106.4011 | 16.0259 | 20.6345 | 5.2481 | 173.6291 | 274.2246 | 206.1868 | 269.0736 | 180.953
2012-10-30 15:39 | 24.5462 | 295.1911 | 227.1571 | 290.0368 | 201.8291 | 76.1609 | 1.6244 | 41.3873 | 67.2292 | 25.4787 | 2.5767 | 41.2946 | 106.4057 | 16.0259 | 20.5919 | 5.2505 | 173.6349 | 274.5992 | 206.5652 | 269.4449 | 181.2372
2012-10-30 15:39 | 24.5559 | 295.5966 | 227.505 | 290.4007 | 202.1261 | 76.21 | 1.6241 | 41.3817 | 67.2092 | 25.4794 | 2.5764 | 41.2891 | 106.3763 | 16.026 | 20.5695 | 5.2539 | 173.5855 | 275.0271 | 206.9355 | 269.8312 | 181.5566
2012-10-30 15:39 | 24.5649 | 295.9706 | 227.8482 | 290.8408 | 202.4579 | 76.2493 | 1.6239 | 41.3776 | 67.1938 | 25.4801 | 2.5761 | 41.2851 | 106.355 | 16.0261 | 20.6066 | 5.2566 | 173.5488 | 275.364 | 207.2416 | 270.2342 | 181.8513
2012-10-30 15:39 | 24.5657 | 296.3825 | 228.1557 | 291.2688 | 202.7972 | 76.2828 | 1.6238 | 41.3769 | 67.1886 | 25.4812 | 2.5761 | 41.2843 | 106.3512 | 16.0261 | 20.6154 | 5.2589 | 173.5398 | 275.7671 | 207.5403 | 270.6534 | 182.1818
2012-10-30 15:39 | 24.5664 | 296.7596 | 228.4748 | 291.6736 | 203.042 | 76.3268 | 1.6238 | 41.3778 | 67.1899 | 25.4819 | 2.5761 | 41.2851 | 106.3544 | 16.0262 | 20.592 | 5.262 | 173.5443 | 276.1676 | 207.8828 | 271.0816 | 182.45
2012-10-30 15:39 | 24.5716 | 297.1412 | 228.767 | 292.0135 | 203.3346 | 76.3666 | 1.6238 | 41.3783 | 67.189 | 25.4828 | 2.5761 | 41.2857 | 106.3573 | 16.0262 | 20.6413 | 5.2647 | 173.5462 | 276.4999 | 208.1257 | 271.3722 | 182.6933
2012-10-30 15:39 | 24.5716 | 297.4715 | 229.113 | 292.3713 | 203.7598 | 76.3944 | 1.6237 | 41.3787 | 67.1883 | 25.4836 | 2.5761 | 41.286 | 106.3582 | 16.0263 | 20.6213 | 5.2666 | 173.5465 | 276.8502 | 208.4917 | 271.75 | 183.1385
2012-10-30 15:39 | 24.5784 | 297.8392 | 229.4107 | 292.7282 | 204.1445 | 76.4424 | 1.6236 | 41.3777 | 67.1828 | 25.4845 | 2.5761 | 41.2851 | 106.353 | 16.0264 | 20.6161 | 5.2699 | 173.5358 | 277.2231 | 208.7946 | 272.1121 | 183.5284
2012-10-30 15:39 | 24.5709 | 298.1964 | 229.7845 | 293.1053 | 204.4756 | 76.4724 | 1.6235 | 41.3753 | 67.1727 | 25.4853 | 2.5759 | 41.2826 | 106.3401 | 16.0265 | 20.5965 | 5.272 | 173.5128 | 277.5999 | 209.188 | 272.5088 | 183.8791
2012-10-30 15:39 | 24.5716 | 298.5733 | 230.1546 | 293.5136 | 204.7204 | 76.5094 | 1.6233 | 41.372 | 67.1592 | 25.4863 | 2.5757 | 41.2791 | 106.3219 | 16.0265 | 20.6093 | 5.2746 | 173.481 | 277.964 | 209.5453 | 272.9043 | 184.1111
2012-10-30 15:39 | 24.5829 | 298.907 | 230.4135 | 293.8979 | 204.9994 | 76.5562 | 1.6232 | 41.3706 | 67.1533 | 25.4869 | 2.5756 | 41.2777 | 106.3135 | 16.0267 | 20.5805 | 5.2778 | 173.4668 | 278.3265 | 209.833 | 273.3174 | 184.4189
2012-10-30 15:39 | 24.5814 | 299.2356 | 230.6832 | 294.2347 | 205.3326 | 76.5886 | 1.6231 | 41.3697 | 67.1492 | 25.4873 | 2.5755 | 41.2768 | 106.309 | 16.0266 | 20.6311 | 5.28 | 173.4582 | 278.6045 | 210.0521 | 273.6036 | 184.7015
2012-10-30 15:39 | 24.5776 | 299.535 | 231.0017 | 294.573 | 205.5178 | 76.6175 | 1.623 | 41.369 | 67.1438 | 25.4885 | 2.5755 | 41.2761 | 106.3053 | 16.0266 | 20.5672 | 5.282 | 173.4491 | 278.9678 | 210.4345 | 274.0058 | 184.9506
2012-10-30 15:39 | 24.5806 | 299.9178 | 231.3819 | 295.0065 | 205.883 | 76.6724 | 1.623 | 41.368 | 67.1393 | 25.489 | 2.5754 | 41.2751 | 106.299 | 16.0268 | 20.6303 | 5.2858 | 173.4383 | 279.2875 | 210.7516 | 274.3762 | 185.2527
2012-10-30 15:39 | 24.5821 | 300.2723 | 231.682 | 295.2924 | 206.231 | 76.7047 | 1.6229 | 41.3676 | 67.1355 | 25.4899 | 2.5753 | 41.2746 | 106.2964 | 16.0268 | 20.6118 | 5.288 | 173.4319 | 279.6605 | 211.0702 | 274.6806 | 185.6192
2012-10-30 15:39 | 24.5858 | 300.6097 | 231.9097 | 295.6727 | 206.5103 | 76.744 | 1.6228 | 41.3668 | 67.1316 | 25.4904 | 2.5753 | 41.2737 | 106.2914 | 16.0269 | 20.5994 | 5.2907 | 173.4229 | 280.0103 | 211.3103 | 275.0733 | 185.9109
2012-10-30 15:39 | 24.5911 | 300.918 | 232.2293 | 295.9738 | 206.7322 | 76.7776 | 1.6226 | 41.3624 | 67.1148 | 25.4913 | 2.575 | 41.2693 | 106.2687 | 16.0269 | 20.6087 | 5.293 | 173.3835 | 280.3093 | 211.6206 | 275.3651 | 186.1235
2012-10-30 15:39 | 24.5873 | 301.2939 | 232.5751 | 296.4195 | 207.1064 | 76.8041 | 1.6223 | 41.3551 | 67.0903 | 25.4917 | 2.5745 | 41.2621 | 106.2312 | 16.0269 | 20.6089 | 5.2949 | 173.3215 | 280.685 | 211.9662 | 275.8106 | 186.4975
2012-10-30 15:39 | 24.5926 | 301.7326 | 232.9418 | 296.8319 | 207.4504 | 76.8377 | 1.6221 | 41.351 | 67.0744 | 25.4927 | 2.5743 | 41.2579 | 106.2092 | 16.027 | 20.6182 | 5.2972 | 173.2836 | 281.1144 | 212.3236 | 276.2137 | 186.8322
2012-10-30 15:39 | 24.5903 | 302.0408 | 233.2498 | 297.2022 | 207.8222 | 76.8798 | 1.6219 | 41.347 | 67.0588 | 25.4936 | 2.574 | 41.2538 | 106.1876 | 16.0271 | 20.644 | 5.3001 | 173.2464 | 281.3968 | 212.6058 | 276.5582 | 187.1782
2012-10-30 15:39 | 24.5978 | 302.3321 | 233.5833 | 297.5053 | 207.9005 | 76.9041 | 1.6217 | 41.3454 | 67.0517 | 25.4945 | 2.5739 | 41.2522 | 106.1788 | 16.0271 | 20.7196 | 5.3018 | 173.2305 | 281.6125 | 212.8637 | 276.7857 | 187.1809
2012-10-30 15:40 | 24.5986 | 302.674 | 233.8551 | 297.8479 | 208.1337 | 76.9394 | 1.6216 | 41.3419 | 67.0387 | 25.495 | 2.5737 | 41.2486 | 106.1607 | 16.0271 | 20.6523 | 5.3042 | 173.1993 | 282.0217 | 213.2028 | 277.1956 | 187.4814
2012-10-30 15:40 | 24.6016 | 302.9875 | 234.1212 | 298.1926 | 208.4202 | 76.9885 | 1.6213 | 41.3368 | 67.0209 | 25.4955 | 2.5734 | 41.2435 | 106.1341 | 16.0271 | 20.6794 | 5.3076 | 173.155 | 282.3081 | 213.4418 | 277.5132 | 187.7408
2012-10-30 15:40 | 24.606 | 303.3062 | 234.3966 | 298.5348 | 208.6176 | 77.0122 | 1.6211 | 41.3318 | 67.0027 | 25.4962 | 2.573 | 41.2385 | 106.1082 | 16.0272 | 20.6276 | 5.3092 | 173.1109 | 282.6786 | 213.769 | 277.9072 | 187.99
2012-10-30 15:40 | 24.5993 | 303.6099 | 234.653 | 298.7891 | 208.8117 | 77.0446 | 1.621 | 41.3296 | 66.9936 | 25.497 | 2.5729 | 41.2363 | 106.0965 | 16.0272 | 20.637 | 5.3115 | 173.0901 | 282.9729 | 214.016 | 278.1521 | 188.1747
2012-10-30 15:40 | 24.618 | 303.8892 | 234.919 | 299.111 | 209.0745 | 77.0862 | 1.6209 | 41.3283 | 66.9873 | 25.4978 | 2.5728 | 41.2351 | 106.0893 | 16.0273 | 20.6236 | 5.3143 | 173.0766 | 283.2656 | 214.2954 | 278.4874 | 188.4509
2012-10-30 15:40 | 24.6075 | 304.2121 | 235.1441 | 299.4728 | 209.3239 | 77.1203 | 1.6208 | 41.3279 | 66.9845 | 25.4983 | 2.5728 | 41.2347 | 106.0877 | 16.0273 | 20.6011 | 5.3167 | 173.0722 | 283.611 | 214.543 | 278.8717 | 188.7228
2012-10-30 15:40 | 24.6195 | 304.5383 | 235.4859 | 299.8071 | 209.521 | 77.1411 | 1.6208 | 41.3283 | 66.9841 | 25.499 | 2.5728 | 41.2351 | 106.0894 | 16.0274 | 20.6492 | 5.3181 | 173.0735 | 283.8891 | 214.8367 | 279.1579 | 188.8718
2012-10-30 15:40 | 24.6135 | 304.7353 | 235.6997 | 300.0543 | 209.8379 | 77.1711 | 1.6206 | 41.325 | 66.9715 | 25.4998 | 2.5726 | 41.2318 | 106.0724 | 16.0274 | 20.6872 | 5.3202 | 173.0438 | 284.0481 | 215.0125 | 279.3671 | 189.1507
2012-10-30 15:40 | 24.6218 | 305.479 | 236.2811 | 300.7723 | 210.4724 | 77.2549 | 1.6209 | 41.3347 | 66.9985 | 25.5015 | 2.5731 | 41.2413 | 106.1193 | 16.0277 | 20.6555 | 5.326 | 173.1178 | 284.8235 | 215.6256 | 280.1168 | 189.8169
2012-10-30 15:40 | 24.624 | 306.1514 | 236.9231 | 301.461 | 210.9708 | 77.3191 | 1.6208 | 41.3351 | 66.9955 | 25.5031 | 2.5731 | 41.2417 | 106.1199 | 16.0279 | 20.6376 | 5.3304 | 173.1154 | 285.5138 | 216.2855 | 280.8234 | 190.3332
2012-10-30 15:40 | 24.6247 | 306.4546 | 237.2063 | 301.7992 | 211.1487 | 77.3537 | 1.6208 | 41.3367 | 66.9988 | 25.5038 | 2.5732 | 41.2432 | 106.1279 | 16.0278 | 20.6503 | 5.3328 | 173.1267 | 285.8043 | 216.556 | 281.1489 | 190.4984
2012-10-30 15:40 | 24.6345 | 306.7784 | 237.4594 | 302.0735 | 211.3869 | 77.3988 | 1.621 | 41.3434 | 67.0186 | 25.5045 | 2.5736 | 41.2498 | 106.1607 | 16.028 | 20.6682 | 5.3359 | 173.1793 | 286.1102 | 216.7912 | 281.4053 | 190.7187
2012-10-30 15:40 | 24.639 | 306.9896 | 237.6248 | 302.3618 | 211.6474 | 77.4352 | 1.6211 | 41.347 | 67.0283 | 25.5053 | 2.5738 | 41.2535 | 106.1791 | 16.0281 | 20.6927 | 5.3384 | 173.2074 | 286.2969 | 216.9321 | 281.6691 | 190.9547
2012-10-30 15:40 | 24.6412 | 307.3287 | 237.9017 | 302.6668 | 211.7834 | 77.4693 | 1.6211 | 41.3474 | 67.0275 | 25.5061 | 2.5738 | 41.2539 | 106.1805 | 16.0282 | 20.6709 | 5.3407 | 173.208 | 286.6578 | 217.2308 | 281.9959 | 191.1125
2012-10-30 15:40 | 24.633 | 307.554 | 238.1446 | 302.9729 | 211.9918 | 77.4832 | 1.621 | 41.3465 | 67.0237 | 25.5064 | 2.5738 | 41.2528 | 106.1746 | 16.0283 | 20.6667 | 5.3417 | 173.1983 | 286.8873 | 217.4779 | 282.3062 | 191.3251
2012-10-30 15:40 | 24.6427 | 307.9005 | 238.4093 | 303.2548 | 212.3048 | 77.5271 | 1.6208 | 41.3408 | 67.0031 | 25.5072 | 2.5734 | 41.2472 | 106.1461 | 16.0282 | 20.6563 | 5.3447 | 173.1492 | 287.2442 | 217.753 | 282.5985 | 191.6485
2012-10-30 15:40 | 24.6412 | 308.1899 | 238.702 | 303.5637 | 212.6262 | 77.5624 | 1.6206 | 41.3374 | 66.9908 | 25.5077 | 2.5732 | 41.2438 | 106.1281 | 16.0283 | 20.6628 | 5.3472 | 173.1189 | 287.5271 | 218.0392 | 282.9009 | 191.9634
2012-10-30 15:40 | 24.6449 | 308.45 | 238.8979 | 303.8855 | 212.7835 | 77.5797 | 1.6205 | 41.3372 | 66.9885 | 25.5084 | 2.5732 | 41.2435 | 106.1264 | 16.0283 | 20.6505 | 5.3483 | 173.1149 | 287.7995 | 218.2474 | 283.235 | 192.133
2012-10-30 15:40
2012-10-30 15:43 |  | 171.3219 | 137.3287 | 172.14 | 123.8669 | 64.9081 | 0.52482 | 12.7567 | 6.695 | 24.307 | 0.8091 | 13.0179 | 10.5328 | 16.0893 | -13.5565 | 4.4748 | 17.2277 | 184.8784 | 150.8852 | 185.6965 | 137.4234
2012-10-30 15:43 | 26.3187 | 196.7439 | 161.5274 | 197.3928 | 149.484 | 63.476 | 0.023666 | 0.75167 | 0.017789 | 31.7614 | 0.040563 | 0.75437 | 0.030599 | 18.5974 | 12.7622 | 4.376 | 0.048388 | 183.9817 | 148.7652 | 184.6306 | 136.7218
2012-10-30 15:43 | 26.318 | 195.7839 | 160.8943 | 196.8101 | 149.2059 | 63.3639 | 0.023692 | 0.7518 | 0.017812 | 31.7321 | 0.040557 | 0.75419 | 0.030588 | 18.5959 | 12.7778 | 4.3683 | 0.0484 | 183.0061 | 148.1165 | 184.0323 | 136.4281
2012-10-30 15:43 | 26.3135 | 194.9263 | 160.3042 | 196.1686 | 148.843 | 63.2708 | 0.023682 | 0.75165 | 0.017801 | 31.7393 | 0.040563 | 0.75434 | 0.030598 | 18.5967 | 12.7815 | 4.3619 | 0.048399 | 182.1448 | 147.5227 | 183.3871 | 136.0615
2012-10-30 15:43 | 26.3052 | 193.9511 | 159.7414 | 195.5273 | 148.357 | 63.1772 | 0.023666 | 0.7511 | 0.017776 | 31.7376 | 0.0405 | 0.75305 | 0.030499 | 18.5939 | 12.7854 | 4.3554 | 0.048275 | 181.1657 | 146.956 | 182.7419 | 135.5716
2012-10-30 15:43 | 26.3075 | 193.0653 | 159.1933 | 194.9164 | 147.9395 | 63.0668 | 0.023657 | 0.75084 | 0.017763 | 31.7387 | 0.040526 | 0.75353 | 0.030538 | 18.5937 | 12.7673 | 4.3478 | 0.048301 | 180.298 | 146.426 | 182.1491 | 135.1722
2012-10-30 15:43 | 26.3097 | 192.101 | 158.5368 | 194.16 | 147.1871 | 62.9934 | 0.023682 | 0.7512 | 0.01779 | 31.7204 | 0.0405 | 0.75311 | 0.030501 | 18.5953 | 12.7532 | 4.3428 | 0.048291 | 179.3478 | 145.7836 | 181.4068 | 134.4339
2012-10-30 15:43 | 26.2963 | 191.137 | 157.7874 | 193.2781 | 146.4811 | 62.9108 | 0.023651 | 0.75047 | 0.017749 | 31.7309 | 0.040459 | 0.75249 | 0.030445 | 18.5988 | 12.7765 | 4.3371 | 0.048194 | 178.3605 | 145.0109 | 180.5016 | 133.7046
2012-10-30 15:43 | 26.2963 | 190.3948 | 157.1679 | 192.5787 | 145.9348 | 62.827 | 0.023647 | 0.75083 | 0.017755 | 31.7515 | 0.040481 | 0.75282 | 0.030475 | 18.5969 | 12.7683 | 4.3313 | 0.04823 | 177.6265 | 144.3996 | 179.8104 | 133.1665
2012-10-30 15:43 | 26.303 | 189.5888 | 156.6134 | 191.9922 | 145.4608 | 62.7588 | 0.023638 | 0.75013 | 0.017732 | 31.7342 | 0.04051 | 0.75328 | 0.030516 | 18.595 | 12.7587 | 4.3266 | 0.048248 | 176.8301 | 143.8547 | 179.2335 | 132.7021
2012-10-30 15:44 | 26.3015 | 188.8689 | 156.0663 | 191.4093 | 145.021 | 62.6883 | 0.023676 | 0.751 | 0.017781 | 31.7199 | 0.040513 | 0.75343 | 0.030524 | 18.5972 | 12.6879 | 4.3217 | 0.048305 | 176.181 | 143.3784 | 178.7214 | 132.3331
2012-10-30 15:44 | 26.288 | 188.0391 | 155.4571 | 190.6496 | 144.577 | 62.5947 | 0.023616 | 0.74945 | 0.017699 | 31.7351 | 0.040428 | 0.7522 | 0.03041 | 18.6059 | 12.703 | 4.3153 | 0.048109 | 175.3361 | 142.7541 | 177.9466 | 131.874
2012-10-30 15:44 | 26.2955 | 187.3207 | 154.9232 | 189.8993 | 144.1656 | 62.5444 | 0.023638 | 0.75028 | 0.017735 | 31.7404 | 0.040447 | 0.75259 | 0.03044 | 18.6068 | 12.7308 | 4.3118 | 0.048175 | 174.5899 | 142.1924 | 177.1685 | 131.4348
2012-10-30 15:44 | 26.2873 | 186.6024 | 154.4047 | 189.2119 | 143.6618 | 62.4444 | 0.023607 | 0.74873 | 0.017675 | 31.7166 | 0.040396 | 0.75157 | 0.03036 | 18.6049 | 12.7267 | 4.3049 | 0.048035 | 173.8757 | 141.678 | 176.4852 | 130.9351
2012-10-30 15:44 | 26.2746 | 185.7504 | 153.7582 | 188.4417 | 143.0577 | 62.3467 | 0.023651 | 0.75021 | 0.017743 | 31.7199 | 0.040469 | 0.75271 | 0.030461 | 18.5996 | 12.7099 | 4.2982 | 0.048204 | 173.0405 | 141.0483 | 175.7318 | 130.3478
2012-10-30 15:44 | 26.282 | 185.0834 | 153.1823 | 187.7035 | 142.6043 | 62.2785 | 0.023663 | 0.75038 | 0.017756 | 31.7111 | 0.040475 | 0.75282 | 0.03047 | 18.5997 | 12.6929 | 4.2935 | 0.048226 | 172.3905 | 140.4894 | 175.0106 | 129.9114
2012-10-30 15:44 | 26.279 | 184.2611 | 152.6656 | 187.0256 | 142.0908 | 62.2022 | 0.023629 | 0.74964 | 0.017713 | 31.7255 | 0.040396 | 0.75145 | 0.030356 | 18.6021 | 12.7388 | 4.2882 | 0.048069 | 171.5223 | 139.9268 | 174.2868 | 129.352
2012-10-30 15:44 | 26.2768 | 183.4829 | 152.0959 | 186.1762 | 141.4325 | 62.1173 | 0.02366 | 0.75006 | 0.017746 | 31.7017 | 0.040441 | 0.75214 | 0.030417 | 18.5985 | 12.6958 | 4.2824 | 0.048163 | 170.7871 | 139.4001 | 173.4804 | 128.7367
2012-10-30 15:44 | 26.2828 | 182.7953 | 151.4496 | 185.4371 | 140.8593 | 62.0647 | 0.023581 | 0.7483 | 0.017646 | 31.7331 | 0.040444 | 0.75207 | 0.030417 | 18.5953 | 12.657 | 4.2787 | 0.048063 | 170.1383 | 138.7926 | 172.7801 | 128.2023
2012-10-30 15:44 | 26.2611 | 182.0725 | 150.9148 | 184.7292 | 140.2932 | 61.9774 | 0.023622 | 0.74885 | 0.017689 | 31.7013 | 0.040381 | 0.75115 | 0.030332 | 18.6015 | 12.7087 | 4.2727 | 0.048021 | 169.3638 | 138.2061 | 172.0205 | 127.5845
2012-10-30 15:44 | 26.2656 | 181.3642 | 150.2829 | 183.9772 | 139.7802 | 61.9052 | 0.023594 | 0.74823 | 0.017654 | 31.7126 | 0.04039 | 0.75132 | 0.030346 | 18.6016 | 12.7173 | 4.2677 | 0.048 | 168.6469 | 137.5656 | 171.2599 | 127.0629
2012-10-30 15:44 | 26.2656 | 180.589 | 149.7354 | 183.2679 | 139.2245 | 61.7989 | 0.023641 | 0.74899 | 0.017707 | 31.682 | 0.040412 | 0.7517 | 0.030377 | 18.6008 | 12.6928 | 4.2604 | 0.048084 | 167.8962 | 137.0426 | 170.5751 | 126.5317
2012-10-30 15:44 | 26.2558 | 179.8159 | 149.1205 | 182.5411 | 138.7818 | 61.7549 | 0.023569 | 0.74748 | 0.017617 | 31.7144 | 0.040327 | 0.75019 | 0.030253 | 18.6027 | 12.6667 | 4.2574 | 0.04787 | 167.1492 | 136.4538 | 169.8744 | 126.1151
2012-10-30 15:44 | 26.2529 | 179.2182 | 148.6206 | 181.794 | 138.4189 | 61.7075 | 0.64363 | 16.1121 | 10.3703 | 25.0332 | 1.0424 | 16.7325 | 17.4412 | 16.0526 | 12.7005 | 4.2541 | 27.8114 | 166.5177 | 135.9201 | 169.0935 | 125.7184
2012-10-30 15:44 | 26.2618 | 178.9924 | 148.9493 | 181.1215 | 138.0104 | 63.0847 | 1.2424 | 31.1284 | 38.6747 | 25.0545 | 1.9416 | 31.0615 | 60.3087 | 15.998 | 12.7053 | 4.3491 | 98.9834 | 166.2871 | 136.244 | 168.4162 | 125.3051
2012-10-30 15:44 | 26.2596 | 179.7571 | 149.7757 | 180.5947 | 137.5794 | 63.6679 | 1.2401 | 31.0911 | 38.556 | 25.0714 | 1.9392 | 31.0226 | 60.1577 | 15.998 | 12.646 | 4.3893 | 98.7137 | 167.1111 | 137.1297 | 167.9487 | 124.9334
2012-10-30 15:44 | 26.2536 | 180.5339 | 150.028 | 180.1425 | 137.3017 | 63.7603 | 1.2407 | 31.1119 | 38.6008 | 25.0759 | 1.9405 | 31.0439 | 60.2412 | 15.9977 | 12.6604 | 4.3956 | 98.842 | 167.8735 | 137.3676 | 167.4821 | 124.6413
2012-10-30 15:44 | 26.2514 | 181.5223 | 150.1685 | 179.788 | 137.1655 | 63.8355 | 1.2408 | 31.1196 | 38.6144 | 25.0795 | 1.941 | 31.0519 | 60.273 | 15.9976 | 12.7031 | 4.4008 | 98.8874 | 168.8192 | 137.4654 | 167.0849 | 124.4624
2012-10-30 15:44 | 26.2573 | 182.3742 | 150.2708 | 179.5239 | 137.0681 | 63.9037 | 1.2409 | 31.1239 | 38.6206 | 25.0825 | 1.9413 | 31.0563 | 60.2902 | 15.9976 | 12.6682 | 4.4055 | 98.9107 | 169.706 | 137.6026 | 166.8557 | 124.3999
2012-10-30 15:44 | 26.2588 | 183.0887 | 150.361 | 179.3767 | 136.9624 | 63.9378 | 1.241 | 31.1302 | 38.6331 | 25.0844 | 1.9417 | 31.0623 | 60.3132 | 15.9976 | 12.6982 | 4.4079 | 98.9463 | 170.3905 | 137.6628 | 166.6785 | 124.2642
2012-10-30 15:44 | 26.2573 | 183.7338 | 150.4405 | 179.3674 | 136.9418 | 63.9892 | 1.2403 | 31.1161 | 38.5946 | 25.0867 | 1.9408 | 31.0484 | 60.2599 | 15.9974 | 12.66 | 4.4114 | 98.8545 | 171.0738 | 137.7805 | 166.7074 | 124.2818
2012-10-30 15:44 | 26.2521 | 184.2067 | 150.5201 | 179.4013 | 136.9519 | 64.0071 | 1.2411 | 31.1382 | 38.6463 | 25.0888 | 1.9423 | 31.0708 | 60.3474 | 15.9973 | 12.6997 | 4.4126 | 98.9937 | 171.507 | 137.8204 | 166.7016 | 124.2522
2012-10-30 15:44 | 26.2514 | 184.7627 | 150.5734 | 179.4398 | 136.9474 | 64.0788 | 1.2411 | 31.1396 | 38.6474 | 25.0902 | 1.9423 | 31.0718 | 60.3509 | 15.9974 | 12.7153 | 4.4176 | 98.9984 | 172.0474 | 137.8581 | 166.7245 | 124.2321
2012-10-30 15:44 | 26.2416 | 185.2586 | 150.6832 | 179.6217 | 137.06 | 64.1244 | 1.241 | 31.1382 | 38.6414 | 25.092 | 1.9422 | 31.0703 | 60.3455 | 15.9973 | 12.6606 | 4.4207 | 98.9869 | 172.598 | 138.0226 | 166.9611 | 124.3994
2012-10-30 15:44 | 26.2536 | 185.8157 | 150.78 | 179.8255 | 137.0567 | 64.1475 | 1.2407 | 31.1343 | 38.6291 | 25.0936 | 1.942 | 31.0666 | 60.3314 | 15.9972 | 12.6767 | 4.4223 | 98.9605 | 173.139 | 138.1033 | 167.1488 | 124.38
2012-10-30 15:44 | 26.2506 | 186.3696 | 150.9622 | 180.069 | 137.0805 | 64.2024 | 1.2412 | 31.1482 | 38.6606 | 25.0955 | 1.9429 | 31.0805 | 60.3854 | 15.9972 | 12.633 | 4.4261 | 99.046 | 173.7366 | 138.3292 | 167.436 | 124.4475
2012-10-30 15:44 | 26.2506 | 186.8442 | 151.0432 | 180.3391 | 137.1418 | 64.229 | 1.2409 | 31.1434 | 38.6467 | 25.0969 | 1.9426 | 31.0755 | 60.3667 | 15.997 | 12.6737 | 4.4279 | 99.0134 | 174.1705 | 138.3695 | 167.6654 | 124.4681
2012-10-30 15:44 | 26.2461 | 187.3105 | 151.2046 | 180.5851 | 137.1908 | 64.2874 | 1.2408 | 31.1424 | 38.6426 | 25.0979 | 1.9425 | 31.0745 | 60.3622 | 15.9971 | 12.6407 | 4.432 | 99.0049 | 174.6698 | 138.5639 | 167.9444 | 124.5501
2012-10-30 15:44 | 26.2536 | 187.7653 | 151.2854 | 180.8119 | 137.2519 | 64.3047 | 1.3586 | 34.1239 | 46.3596 | 25.1175 | 2.1301 | 34.0687 | 72.5709 | 15.9937 | 12.6767 | 4.4332 | 118.9305 | 175.0886 | 138.6087 | 168.1352 | 124.5752
2012-10-30 15:44 | 26.2499 | 188.2207 | 151.6367 | 181.2114 | 137.3248 | 64.9058 | 1.5085 | 37.9884 | 57.3046 | 25.1833 | 2.3701 | 37.9043 | 89.8364 | 15.9928 | 12.6852 | 4.4746 | 147.141 | 175.5355 | 138.9515 | 168.5262 | 124.6396
2012-10-30 15:45 | 26.2543 | 189.1868 | 152.4286 | 181.5134 | 137.3483 | 65.2884 | 1.5082 | 37.9966 | 57.3075 | 25.1928 | 2.3705 | 37.9115 | 89.8694 | 15.993 | 12.657 | 4.501 | 147.1769 | 176.5298 | 139.7716 | 168.8564 | 124.6913
2012-10-30 15:45 | 26.2394 | 190.3065 | 152.9697 | 181.9724 | 137.5018 | 65.393 | 1.5087 | 38.0139 | 57.3504 | 25.1971 | 2.3716 | 37.9289 | 89.9513 | 15.9931 | 12.7033 | 4.5082 | 147.3017 | 177.6032 | 140.2664 | 169.2691 | 124.7985
2012-10-30 15:45 | 26.2491 | 191.6437 | 153.5318 | 182.6167 | 137.7527 | 65.5063 | 1.5083 | 38.0114 | 57.3343 | 25.2007 | 2.3714 | 37.9261 | 89.9377 | 15.9932 | 12.7171 | 4.516 | 147.272 | 178.9266 | 140.8147 | 169.8996 | 125.0356
2012-10-30 15:45 | 26.2499 | 192.8662 | 154.0349 | 183.2992 | 138.0138 | 65.5866 | 1.5081 | 38.0094 | 57.3205 | 25.2041 | 2.3713 | 37.9243 | 89.9298 | 15.993 | 12.7423 | 4.5215 | 147.2504 | 180.1239 | 141.2926 | 170.5569 | 125.2715
2012-10-30 15:45 | 26.2566 | 193.9536 | 154.4474 | 184.0348 | 138.1813 | 65.6952 | 1.5073 | 37.9934 | 57.267 | 25.2064 | 2.3703 | 37.9084 | 89.8537 | 15.9932 | 12.7694 | 4.529 | 147.1207 | 181.1842 | 141.678 | 171.2654 | 125.4119
2012-10-30 15:45 | 26.2461 | 195.0633 | 154.8738 | 184.6672 | 138.4466 | 65.7785 | 1.5074 | 38.0005 | 57.2824 | 25.2092 | 2.3707 | 37.9157 | 89.8883 | 15.9932 | 12.767 | 4.5348 | 147.1707 | 182.2963 | 142.1068 | 171.9002 | 125.6796
2012-10-30 15:45 | 26.2386 | 195.96 | 155.2878 | 185.2791 | 138.7991 | 65.8449 | 1.5074 | 38.0037 | 57.286 | 25.2118 | 2.3709 | 37.9187 | 89.9024 | 15.9933 | 12.7392 | 4.5393 | 147.1884 | 183.2208 | 142.5486 | 172.5399 | 126.0599
2012-10-30 15:45 | 26.2476 | 197.0386 | 155.7609 | 186.2019 | 139.1605 | 65.9426 | 1.507 | 37.9991 | 57.2657 | 25.2147 | 2.3706 | 37.9143 | 89.8806 | 15.9934 | 12.8174 | 4.5461 | 147.1463 | 184.2212 | 142.9435 | 173.3845 | 126.3431
2012-10-30 15:45 | 26.2514 | 198.1042 | 156.2831 | 187.1118 | 139.5628 | 66.0501 | 1.507 | 38.002 | 57.2684 | 25.2173 | 2.3708 | 37.9169 | 89.8928 | 15.9934 | 12.7397 | 4.5535 | 147.1611 | 185.3645 | 143.5434 | 174.3721 | 126.8231
2012-10-30 15:45 | 26.2581 | 199.2005 | 156.7659 | 187.9542 | 139.995 | 66.1275 | 1.5067 | 37.9989 | 57.2515 | 25.2205 | 2.3706 | 37.9137 | 89.8763 | 15.9936 | 12.7464 | 4.5588 | 147.1278 | 186.4541 | 144.0195 | 175.2078 | 127.2486
2012-10-30 15:45 | 26.2514 | 200.3348 | 157.3514 | 188.9012 | 140.4061 | 66.2385 | 1.5063 | 37.994 | 57.231 | 25.2231 | 2.3703 | 37.9089 | 89.8536 | 15.9936 | 12.8131 | 4.5665 | 147.0846 | 187.5217 | 144.5383 | 176.0881 | 127.593
2012-10-30 15:45 | 26.2536 | 201.3364 | 157.8028 | 190.0144 | 140.8378 | 66.3379 | 1.5061 | 37.9929 | 57.2219 | 25.2257 | 2.3702 | 37.9077 | 89.8479 | 15.9936 | 12.7542 | 4.5733 | 147.0698 | 188.5822 | 145.0486 | 177.2602 | 128.0836
2012-10-30 15:45 | 26.2588 | 202.4199 | 158.3348 | 191.0014 | 141.2035 | 66.398 | 1.5063 | 38.0014 | 57.2414 | 25.2283 | 2.3707 | 37.9159 | 89.8865 | 15.9937 | 12.7634 | 4.5775 | 147.1279 | 189.6565 | 145.5714 | 178.238 | 128.4401
2012-10-30 15:45 | 26.2506 | 203.376 | 158.9078 | 192.0144 | 141.5866 | 66.4743 | 1.5065 | 38.0108 | 57.2633 | 25.2312 | 2.3713 | 37.9252 | 89.9304 | 15.9937 | 12.7715 | 4.5827 | 147.1937 | 190.6045 | 146.1363 | 179.2429 | 128.8151
2012-10-30 15:45 | 26.2573 | 204.2684 | 159.3485 | 193.0544 | 142.0193 | 66.5667 | 1.5414 | 38.9037 | 59.9668 | 25.2389 | 2.4287 | 38.8444 | 94.3409 | 15.994 | 12.7579 | 4.5891 | 154.3077 | 191.5105 | 146.5906 | 180.2965 | 129.2614
2012-10-30 15:45 | 26.2648 | 205.3506 | 160.081 | 193.9657 | 142.472 | 66.9106 | 1.6121 | 40.7419 | 65.6809 | 25.2723 | 2.5411 | 40.6498 | 103.296 | 15.9968 | 12.7735 | 4.6128 | 168.9769 | 192.5771 | 147.3075 | 181.1922 | 129.6985
2012-10-30 15:45 | 26.2753 | 206.7784 | 160.9991 | 195.0532 | 142.9125 | 67.1597 | 1.6141 | 40.8045 | 65.8621 | 25.2801 | 2.545 | 40.713 | 103.6158 | 15.9971 | 12.7922 | 4.63 | 169.4779 | 193.9862 | 148.2069 | 182.261 | 130.1203
2012-10-30 15:45 | 26.2656 | 207.9696 | 161.6335 | 196.1482 | 143.256 | 67.2822 | 1.6133 | 40.7903 | 65.8081 | 25.2833 | 2.544 | 40.6973 | 103.5358 | 15.9971 | 12.7988 | 4.6384 | 169.3439 | 195.1708 | 148.8347 | 183.3494 | 130.4572
2012-10-30 15:45 | 26.2701 | 209.1633 | 162.2085 | 197.1908 | 143.6561 | 67.3695 | 1.6138 | 40.8077 | 65.8558 | 25.2866 | 2.5452 | 40.7153 | 103.6268 | 15.9971 | 12.8277 | 4.6445 | 169.4826 | 196.3356 | 149.3808 | 184.3631 | 130.8284
2012-10-30 15:45 | 26.2693 | 210.3872 | 162.8717 | 198.3109 | 144.247 | 67.4776 | 1.6141 | 40.8195 | 65.8853 | 25.2898 | 2.5459 | 40.7273 | 103.6884 | 15.9971 | 12.7902 | 4.6519 | 169.5737 | 197.597 | 150.0815 | 185.5207 | 131.4568
2012-10-30 15:45 | 26.2813 | 211.5924 | 163.5517 | 199.5542 | 144.7663 | 67.6064 | 1.6135 | 40.8094 | 65.8447 | 25.293 | 2.5453 | 40.7167 | 103.6347 | 15.9971 | 12.8552 | 4.6608 | 169.4793 | 198.7372 | 150.6965 | 186.699 | 131.9111
2012-10-30 15:45 | 26.2686 | 212.9107 | 164.2034 | 200.7769 | 145.3149 | 67.729 | 1.6133 | 40.8095 | 65.8375 | 25.2958 | 2.5453 | 40.7169 | 103.6358 | 15.997 | 12.7977 | 4.6692 | 169.4733 | 200.113 | 151.4057 | 187.9792 | 132.5172
2012-10-30 15:45 | 26.2723 | 214.0801 | 164.9106 | 201.9061 | 145.9646 | 67.8249 | 1.6128 | 40.8015 | 65.8044 | 25.2986 | 2.5448 | 40.7088 | 103.5952 | 15.9969 | 12.8462 | 4.6758 | 169.3996 | 201.2339 | 152.0644 | 189.0599 | 133.1184
2012-10-30 15:45 | 26.2858 | 215.2434 | 165.554 | 203.1436 | 146.5475 | 67.9295 | 1.6127 | 40.8019 | 65.7993 | 25.3011 | 2.5448 | 40.7093 | 103.5981 | 15.9969 | 12.8271 | 4.6831 | 169.3974 | 202.4163 | 152.7269 | 190.3165 | 133.7204
2012-10-30 15:45 | 26.285 | 216.42 | 166.2221 | 204.1898 | 147.0046 | 68.0312 | 1.6118 | 40.7847 | 65.7373 | 25.3036 | 2.5437 | 40.6917 | 103.5087 | 15.9969 | 12.8222 | 4.6901 | 169.246 | 203.5978 | 153.3999 | 191.3676 | 134.1824
2012-10-30 15:45 | 26.2858 | 217.4761 | 166.8766 | 205.3716 | 147.5136 | 68.1323 | 1.6124 | 40.8039 | 65.7918 | 25.3064 | 2.5449 | 40.7104 | 103.6042 | 15.9968 | 12.8271 | 4.697 | 169.396 | 204.649 | 154.0495 | 192.5445 | 134.6865
2012-10-30 15:45 | 26.2865 | 218.4808 | 167.5039 | 206.5573 | 148.0995 | 68.2283 | 1.6125 | 40.8116 | 65.8098 | 25.3091 | 2.5454 | 40.7189 | 103.6476 | 15.9968 | 12.7871 | 4.7037 | 169.4574 | 205.6937 | 154.7168 | 193.7702 | 135.3124
2012-10-30 15:45 | 26.2903 | 219.4729 | 168.1539 | 207.636 | 148.7119 | 68.3416 | 1.6119 | 40.7989 | 65.7619 | 25.3118 | 2.5447 | 40.7066 | 103.5851 | 15.9968 | 12.8357 | 4.7115 | 169.347 | 206.6372 | 155.3182 | 194.8003 | 135.8762
2012-10-30 15:45 | 26.2985 | 220.6503 | 168.906 | 208.794 | 149.3058 | 68.4323 | 1.6114 | 40.7921 | 65.7339 | 25.3141 | 2.5442 | 40.6993 | 103.549 | 15.9966 | 12.8602 | 4.7177 | 169.2829 | 207.7901 | 156.0458 | 195.9338 | 136.4456
2012-10-30 15:45 | 26.2985 | 221.6659 | 169.5137 | 210.1092 | 149.9264 | 68.5213 | 1.611 | 40.7833 | 65.6999 | 25.3163 | 2.5437 | 40.6909 | 103.5057 | 15.9967 | 12.8683 | 4.7239 | 169.2056 | 208.7976 | 156.6454 | 197.2409 | 137.0581
2012-10-30 15:46 | 26.2948 | 222.5203 | 170.1296 | 211.1412 | 150.4317 | 68.6149 | 1.6104 | 40.7734 | 65.6629 | 25.3182 | 2.5431 | 40.6809 | 103.4568 | 15.9964 | 12.8769 | 4.7303 | 169.1197 | 209.6434 | 157.2527 | 198.2643 | 137.5548
2012-10-30 15:46 | 26.3045 | 223.4352 | 170.7474 | 212.2535 | 150.9659 | 68.7126 | 1.6099 | 40.7632 | 65.6243 | 25.3205 | 2.5425 | 40.6705 | 103.4046 | 15.9963 | 12.911 | 4.737 | 169.0289 | 210.5242 | 157.8364 | 199.3425 | 138.0549
2012-10-30 15:46 | 26.3022 | 224.5702 | 171.4625 | 213.4484 | 151.608 | 68.7952 | 1.6096 | 40.7586 | 65.6033 | 25.3229 | 2.5422 | 40.6657 | 103.3805 | 15.9963 | 12.9047 | 4.7427 | 168.9838 | 211.6655 | 158.5578 | 200.5437 | 138.7033
2012-10-30 15:46 | 26.3187 | 225.6845 | 172.1928 | 214.6542 | 152.1996 | 68.9073 | 1.6095 | 40.7605 | 65.6033 | 25.3252 | 2.5423 | 40.6677 | 103.3906 | 15.9963 | 12.9252 | 4.7505 | 168.9938 | 212.7593 | 159.2676 | 201.729 | 139.2744
2012-10-30 15:46 | 26.3135 | 226.6235 | 172.8548 | 215.7516 | 152.7735 | 69.0131 | 1.6094 | 40.7617 | 65.6009 | 25.3277 | 2.5424 | 40.6691 | 103.3971 | 15.9963 | 12.916 | 4.7578 | 168.998 | 213.7075 | 159.9388 | 202.8356 | 139.8575
2012-10-30 15:46 | 26.3202 | 227.4798 | 173.4352 | 216.8805 | 153.2977 | 69.1038 | 1.6098 | 40.7777 | 65.646 | 25.3302 | 2.5434 | 40.6848 | 103.4776 | 15.9963 | 12.9349 | 4.764 | 169.1236 | 214.5449 | 160.5003 | 203.9456 | 140.3628
2012-10-30 15:46 | 26.3247 | 228.4715 | 174.0565 | 217.9875 | 153.8354 | 69.19 | 1.6095 | 40.7714 | 65.6196 | 25.3325 | 2.543 | 40.6784 | 103.4453 | 15.9962 | 12.9435 | 4.77 | 169.0649 | 215.528 | 161.113 | 205.044 | 140.8919
2012-10-30 15:46 | 26.3247 | 229.494 | 174.8218 | 219.1214 | 154.4962 | 69.2766 | 1.6091 | 40.7667 | 65.5989 | 25.3346 | 2.5427 | 40.6737 | 103.4215 | 15.9962 | 12.9516 | 4.7759 | 169.0204 | 216.5424 | 161.8702 | 206.1698 | 141.5446
2012-10-30 15:46 | 26.321 | 230.5282 | 175.4822 | 220.3146 | 155.0724 | 69.3795 | 1.6091 | 40.7685 | 65.5986 | 25.337 | 2.5429 | 40.6756 | 103.432 | 15.9961 | 12.9438 | 4.783 | 169.0306 | 217.5844 | 162.5384 | 207.3708 | 142.1286
2012-10-30 15:46 | 26.3389 | 231.5289 | 176.1838 | 221.4032 | 155.7052 | 69.476 | 1.6086 | 40.7614 | 65.5701 | 25.3392 | 2.5424 | 40.668 | 103.3936 | 15.996 | 12.9291 | 4.7897 | 168.9637 | 218.5998 | 163.2547 | 208.4741 | 142.7761
2012-10-30 15:46 | 26.3449 | 232.5174 | 176.866 | 222.5468 | 156.3844 | 69.5737 | 1.6082 | 40.7549 | 65.5429 | 25.3416 | 2.542 | 40.6618 | 103.3627 | 15.9959 | 12.9555 | 4.7964 | 168.9056 | 219.5619 | 163.9105 | 209.5913 | 143.4289
2012-10-30 15:46 | 26.3472 | 233.4903 | 177.5133 | 223.6511 | 156.9672 | 69.6627 | 1.608 | 40.7531 | 65.5318 | 25.3436 | 2.5419 | 40.66 | 103.3537 | 15.9959 | 12.917 | 4.8025 | 168.8855 | 220.5733 | 164.5963 | 210.7341 | 144.0502
2012-10-30 15:46 | 26.3591 | 234.3666 | 178.2057 | 224.7018 | 157.5133 | 69.7609 | 1.6076 | 40.7446 | 65.4992 | 25.3457 | 2.5414 | 40.6514 | 103.3106 | 15.9958 | 12.9493 | 4.8093 | 168.8098 | 221.4173 | 165.2564 | 211.7525 | 144.564
2012-10-30 15:46 | 26.3569 | 235.2717 | 178.8645 | 225.8053 | 158.0494 | 69.8326 | 1.6072 | 40.7382 | 65.4735 | 25.3477 | 2.541 | 40.6453 | 103.28 | 15.9957 | 12.9145 | 4.8143 | 168.7535 | 222.3572 | 165.95 | 212.8908 | 145.1349
2012-10-30 15:46 | 26.3614 | 236.0774 | 179.4872 | 226.8051 | 158.6078 | 69.9008 | 1.607 | 40.7376 | 65.4671 | 25.3494 | 2.541 | 40.6443 | 103.2759 | 15.9956 | 12.9353 | 4.819 | 168.743 | 223.1421 | 166.5519 | 213.8698 | 145.6725
2012-10-30 15:46 | 26.3599 | 236.9672 | 180.104 | 227.7319 | 159.1566 | 70.0002 | 1.6069 | 40.7383 | 65.4637 | 25.3516 | 2.541 | 40.6452 | 103.2805 | 15.9956 | 12.9094 | 4.8258 | 168.7442 | 224.0578 | 167.1946 | 214.8225 | 146.2472
2012-10-30 15:46 | 26.3733 | 237.8796 | 180.7711 | 228.7404 | 159.7281 | 70.0823 | 1.6068 | 40.7378 | 65.4559 | 25.354 | 2.541 | 40.6445 | 103.2768 | 15.9956 | 12.8983 | 4.8315 | 168.7327 | 224.9813 | 167.8728 | 215.8421 | 146.8298
2012-10-30 15:46 | 26.3726 | 238.7031 | 181.389 | 229.6834 | 160.3478 | 70.1719 | 1.6064 | 40.7304 | 65.4278 | 25.3557 | 2.5406 | 40.6373 | 103.2414 | 15.9955 | 12.9424 | 4.8377 | 168.6692 | 225.7607 | 168.4466 | 216.741 | 147.4054
2012-10-30 15:46 | 26.3786 | 239.537 | 182.045 | 230.6487 | 161.048 | 70.2609 | 1.6065 | 40.7373 | 65.4453 | 25.3574 | 2.541 | 40.6439 | 103.2749 | 15.9955 | 12.9566 | 4.8438 | 168.7202 | 226.5804 | 169.0884 | 217.6921 | 148.0914
2012-10-30 15:46 | 26.3921 | 240.4215 | 182.7009 | 231.5506 | 161.6203 | 70.3418 | 1.6065 | 40.7392 | 65.446 | 25.3595 | 2.5411 | 40.6459 | 103.2844 | 15.9955 | 12.9416 | 4.8494 | 168.7304 | 227.4799 | 169.7593 | 218.609 | 148.6787
2012-10-30 15:46 | 26.404 | 241.3316 | 183.3593 | 232.6827 | 162.1406 | 70.429 | 1.6064 | 40.7409 | 65.4457 | 25.3618 | 2.5412 | 40.6473 | 103.2917 | 15.9955 | 12.9535 | 4.8554 | 168.7374 | 228.3781 | 170.4058 | 219.7292 | 149.1871
2012-10-30 15:46 | 26.401 | 242.1677 | 184.046 | 233.7013 | 162.8946 | 70.5209 | 1.6064 | 40.7436 | 65.449 | 25.3639 | 2.5414 | 40.6501 | 103.3064 | 15.9954 | 12.9016 | 4.8617 | 168.7555 | 229.2661 | 171.1444 | 220.7997 | 149.993
2012-10-30 15:46 | 26.4085 | 243.0926 | 184.7865 | 234.6988 | 163.4653 | 70.6169 | 1.6062 | 40.7437 | 65.4438 | 25.3661 | 2.5414 | 40.6503 | 103.3081 | 15.9953 | 12.9417 | 4.8683 | 168.7518 | 230.1509 | 171.8448 | 221.7571 | 150.5236
2012-10-30 15:46 | 26.413 | 243.9555 | 185.4065 | 235.6656 | 164.0331 | 70.7145 | 1.6061 | 40.7444 | 65.4406 | 25.3681 | 2.5414 | 40.6508 | 103.3097 | 15.9954 | 12.9665 | 4.8751 | 168.7503 | 230.989 | 172.44 | 222.6991 | 151.0666
2012-10-30 15:46 | 26.4085 | 244.6994 | 186.0255 | 236.5603 | 164.5571 | 70.7891 | 1.606 | 40.7445 | 65.4351 | 25.3705 | 2.5414 | 40.6509 | 103.3102 | 15.9955 | 12.9661 | 4.8802 | 168.7453 | 231.7333 | 173.0594 | 223.5942 | 151.591
2012-10-30 15:46 | 26.4183 | 245.622 | 186.6863 | 237.5948 | 165.1577 | 70.8792 | 1.6058 | 40.7421 | 65.422 | 25.3724 | 2.5412 | 40.6485 | 103.2981 | 15.9955 | 12.927 | 4.8864 | 168.7201 | 232.695 | 173.7593 | 224.6678 | 152.2307
2012-10-30 15:47 | 26.416 | 246.4813 | 187.3745 | 238.6013 | 165.7192 | 70.9711 | 1.6055 | 40.7393 | 65.4085 | 25.3743 | 2.5411 | 40.6457 | 103.2836 | 15.9955 | 12.9329 | 4.8927 | 168.6921 | 233.5484 | 174.4416 | 225.6684 | 152.7863
2012-10-30 15:47 | 26.4272 | 247.2521 | 187.9625 | 239.4012 | 166.1776 | 71.0682 | 1.6055 | 40.7407 | 65.4083 | 25.376 | 2.5411 | 40.6469 | 103.2898 | 15.9955 | 12.9807 | 4.8994 | 168.698 | 234.2714 | 174.9818 | 226.4205 | 153.1969
2012-10-30 15:47 | 26.4317 | 247.9808 | 188.536 | 240.2139 | 166.6412 | 71.126 | 1.6053 | 40.7379 | 65.3955 | 25.3776 | 2.541 | 40.6442 | 103.2752 | 15.9956 | 12.9852 | 4.9034 | 168.6707 | 234.9956 | 175.5508 | 227.2287 | 153.656
2012-10-30 15:47 | 26.4377 | 248.7696 | 189.1503 | 241.1064 | 167.188 | 71.2023 | 1.6051 | 40.7373 | 65.3881 | 25.3796 | 2.5409 | 40.6435 | 103.2718 | 15.9956 | 12.9627 | 4.9087 | 168.6599 | 235.8069 | 176.1876 | 228.1437 | 154.2253
2012-10-30 15:47 | 26.4459 | 249.5603 | 189.7827 | 241.9539 | 167.819 | 71.2942 | 1.605 | 40.7358 | 65.3791 | 25.3813 | 2.5408 | 40.642 | 103.2644 | 15.9956 | 12.9872 | 4.915 | 168.6435 | 236.5731 | 176.7955 | 228.9667 | 154.8318
2012-10-30 15:47 | 26.4377 | 250.3188 | 190.2928 | 242.7298 | 168.1066 | 71.3705 | 1.6049 | 40.7376 | 65.3796 | 25.3833 | 2.5409 | 40.6439 | 103.2738 | 15.9956 | 13.0157 | 4.9203 | 168.6534 | 237.3031 | 177.2771 | 229.7141 | 155.0909
2012-10-30 15:47 | 26.4527 | 251.0729 | 190.9243 | 243.6855 | 168.5928 | 71.4398 | 1.6048 | 40.7372 | 65.3744 | 25.3848 | 2.5409 | 40.6435 | 103.272 | 15.9956 | 13.0388 | 4.9251 | 168.6464 | 238.0341 | 177.8855 | 230.6467 | 155.554
2012-10-30 15:47 | 0 |  |  |  |  | 709.5269 | 1.6048 | 40.7396 | 65.3777 | 25.3866 | 2.5411 | 40.6458 | 103.2831 | 15.9956 |  | 48.9148 | 168.6608 | 0 | 0 | 0 | 0
2012-10-30 15:47 | 24.9277 | 261.987 | 200.168 | 253.483 | 177.7991 | 73.2094 | 1.6046 | 40.7373 | 65.3651 | 25.3886 | 2.5409 | 40.6436 | 103.2712 | 15.9958 | 21.1497 | 5.0471 | 168.6363 | 240.8373 | 179.0183 | 232.3333 | 156.6494
2012-10-30 15:47 | 24.4243 | 262.0441 | 200.1722 | 253.6044 | 177.7496 | 73.2851 | 1.6045 | 40.7385 | 65.3635 | 25.3907 | 2.541 | 40.6447 | 103.2768 | 15.9958 | 20.6663 | 5.0523 | 168.6403 | 241.3778 | 179.5059 | 232.9381 | 157.0833
2012-10-30 15:47 | 24.4295 | 262.7498 | 200.6694 | 254.4371 | 178.2988 | 73.3574 | 1.6045 | 40.743 | 65.3737 | 25.3924 | 2.5412 | 40.6491 | 103.2984 | 15.9959 | 20.6154 | 5.0573 | 168.672 | 242.1344 | 180.054 | 233.8217 | 157.6834
2012-10-30 15:47 | 24.4407 | 263.4146 | 201.1962 | 255.143 | 178.7367 | 73.4169 | 1.6046 | 40.7476 | 65.3839 | 25.3941 | 2.5415 | 40.6536 | 103.3214 | 15.9959 | 20.6787 | 5.0614 | 168.7053 | 242.7359 | 180.5175 | 234.4643 | 158.058
2012-10-30 15:47 | 24.4475 | 264.1137 | 201.7934 | 255.8599 | 179.0958 | 73.5094 | 1.6047 | 40.7522 | 65.3942 | 25.3959 | 2.5417 | 40.6581 | 103.3427 | 15.9961 | 20.6735 | 5.0677 | 168.7368 | 243.4402 | 181.1199 | 235.1864 | 158.4223
2012-10-30 15:47 | 24.449 | 264.8502 | 202.3892 | 256.7078 | 179.5946 | 73.588 | 1.6044 | 40.7486 | 65.3778 | 25.3977 | 2.5415 | 40.6546 | 103.3252 | 15.9961 | 20.5909 | 5.0732 | 168.7031 | 244.2593 | 181.7983 | 236.1169 | 159.0037
2012-10-30 15:47 | 24.4602 | 265.5766 | 202.9752 | 257.5496 | 180.1541 | 73.6643 | 1.6041 | 40.7433 | 65.3565 | 25.3994 | 2.5412 | 40.6493 | 103.2974 | 15.9962 | 20.6301 | 5.0784 | 168.6538 | 244.9465 | 182.3451 | 236.9195 | 159.524
2012-10-30 15:47 | 24.4632 | 266.2209 | 203.4506 | 258.2971 | 180.6741 | 73.7186 | 1.604 | 40.7431 | 65.3525 | 25.4007 | 2.5412 | 40.6491 | 103.2964 | 15.9962 | 20.6411 | 5.0822 | 168.6489 | 245.5798 | 182.8095 | 237.656 | 160.033
2012-10-30 15:47 | 24.4624 | 266.8418 | 203.997 | 259.0134 | 181.1473 | 73.8047 | 1.6039 | 40.7415 | 65.3436 | 25.4022 | 2.5411 | 40.6475 | 103.2878 | 15.9963 | 20.6604 | 5.0881 | 168.6314 | 246.1814 | 183.3366 | 238.353 | 160.4869
2012-10-30 15:47 | 24.4766 | 267.4697 | 204.5467 | 259.6899 | 181.5847 | 73.8671 | 1.6037 | 40.7411 | 65.3384 | 25.4037 | 2.541 | 40.6469 | 103.2844 | 15.9963 | 20.7186 | 5.0924 | 168.6228 | 246.7511 | 183.8281 | 238.9713 | 160.8661
2012-10-30 15:47 | 24.4774 | 268.1385 | 205.0396 | 260.3412 | 182.0479 | 73.9422 | 1.6036 | 40.74 | 65.3312 | 25.4051 | 2.5409 | 40.6459 | 103.279 | 15.9964 | 20.6674 | 5.0976 | 168.6103 | 247.4711 | 184.3722 | 239.6738 | 161.3805
2012-10-30 15:47 | 24.4744 | 268.7878 | 205.5524 | 261.0898 | 182.4917 | 74.0104 | 1.6034 | 40.737 | 65.3173 | 25.4067 | 2.5407 | 40.6428 | 103.2629 | 15.9964 | 20.6563 | 5.1023 | 168.5801 | 248.1315 | 184.8961 | 240.4335 | 161.8354
2012-10-30 15:47 | 24.4796 | 269.3053 | 206.0065 | 261.714 | 182.932 | 74.0746 | 1.6032 | 40.7357 | 65.3092 | 25.4084 | 2.5407 | 40.6416 | 103.2574 | 15.9963 | 20.6736 | 5.1067 | 168.5665 | 248.6317 | 185.3329 | 241.0404 | 162.2584
2012-10-30 15:47 | 24.4819 | 269.9712 | 206.5049 | 262.4442 | 183.3813 | 74.1572 | 1.6033 | 40.7396 | 65.3176 | 25.4099 | 2.5409 | 40.6454 | 103.2751 | 15.9966 | 20.6518 | 5.1124 | 168.5927 | 249.3194 | 185.8531 | 241.7924 | 162.7295
2012-10-30 15:47 | 24.4886 | 270.6957 | 207.0471 | 263.1513 | 183.9409 | 74.2197 | 1.6034 | 40.7437 | 65.3267 | 25.4115 | 2.5411 | 40.6496 | 103.2954 | 15.9967 | 20.7106 | 5.1167 | 168.6221 | 249.9851 | 186.3365 | 242.4407 | 163.2303
2012-10-30 15:47 | 24.4946 | 271.3455 | 207.5532 | 263.8264 | 184.4253 | 74.2873 | 1.6033 | 40.7449 | 65.3264 | 25.4132 | 2.5412 | 40.6506 | 103.2995 | 15.9969 | 20.6725 | 5.1214 | 168.6258 | 250.673 | 186.8807 | 243.1539 | 163.7528
2012-10-30 15:47 | 24.5103 | 271.9119 | 208.1241 | 264.5382 | 184.8843 | 74.3584 |  |  |  |  | 1.9558 | 31.2565 | 61.1299 | 15.9818 | 20.6682 | 5.1263 | 61.1299 | 251.2437 | 187.4559 | 243.87 | 164.2161
2012-10-30 15:48 | 0 | 221.1381 | 167.0171 | 227.5554 | 149.1377 | 70.3782 | 1.2409 | 31.3238 | 38.8704 | 25.2423 | 1.9554 | 31.2532 | 61.1137 | 15.9827 | -13.5483 | 4.8519 | 99.9841 | 234.6864 | 180.5654 | 241.1037 | 162.686
2012-10-30 15:48 | 26.5365 | 247.7803 | 194.1038 | 253.3613 | 175.2992 | 70.4949 | 1.2401 | 31.3163 | 38.8349 | 25.2533 | 1.9549 | 31.2455 | 61.0826 | 15.983 | 12.8862 | 4.8599 | 99.9175 | 234.8941 | 181.2176 | 240.4751 | 162.413
2012-10-30 15:48 | 26.5313 | 247.9199 | 194.1536 | 252.4299 | 175.087 | 70.5458 | 1.2391 | 31.2933 | 38.7761 | 25.2544 | 1.9535 | 31.2226 | 60.9933 | 15.9829 | 12.9503 | 4.8634 | 99.7694 | 234.9696 | 181.2033 | 239.4796 | 162.1367
2012-10-30 15:48 | 26.5403 | 247.9759 | 194.0919 | 251.5164 | 174.9203 | 70.5458 | 1.2387 | 31.2839 | 38.7515 | 25.2553 | 1.9529 | 31.2133 | 60.9572 | 15.9829 | 12.9349 | 4.8634 | 99.7086 | 235.041 | 181.157 | 238.5815 | 161.9854
2012-10-30 15:48 | 26.5433 | 248.022 | 194.0045 | 250.8351 | 174.9428 | 70.5585 | 1.2384 | 31.2771 | 38.7342 | 25.2556 | 1.9525 | 31.2068 | 60.9311 | 15.983 | 12.942 | 4.8643 | 99.6653 | 235.08 | 181.0625 | 237.8931 | 162.0008
2012-10-30 15:48 | 26.5485 | 247.9919 | 193.9547 | 250.2262 | 174.7488 | 70.5429 | 1.2388 | 31.2879 | 38.7609 | 25.2557 | 1.9532 | 31.2172 | 60.9722 | 15.9829 | 12.9145 | 4.8632 | 99.7331 | 235.0774 | 181.0402 | 237.3117 | 161.8343
2012-10-30 15:48 | 26.5477 | 247.9324 | 193.9067 | 249.666 | 174.5294 | 70.5359 | 1.2384 | 31.2775 | 38.7352 | 25.2557 | 1.9525 | 31.2071 | 60.9334 | 15.9828 | 12.983 | 4.8627 | 99.6685 | 234.9494 | 180.9237 | 236.683 | 161.5464
2012-10-30 15:48 | 26.559 | 247.9437 | 193.7883 | 249.2352 | 174.4743 | 70.5232 | 1.2377 | 31.2586 | 38.6882 | 25.2559 | 1.9514 | 31.1882 | 60.8592 | 15.9828 | 12.925 | 4.8619 | 99.5474 | 235.0187 | 180.8633 | 236.3102 | 161.5493
2012-10-30 15:48 | 26.5537 | 247.8405 | 193.7241 | 248.7799 | 174.5354 | 70.5313 | 1.2376 | 31.2568 | 38.6837 | 25.2559 | 1.9512 | 31.1866 | 60.8529 | 15.9829 | 12.9524 | 4.8624 | 99.5366 | 234.8881 | 180.7717 | 235.8275 | 161.583
2012-10-30 15:48 | 26.5575 | 247.77 | 193.6768 | 248.4119 | 174.4884 | 70.5342 | 1.2377 | 31.2607 | 38.6927 | 25.2561 | 1.9515 | 31.1902 | 60.8667 | 15.9829 | 12.9765 | 4.8626 | 99.5595 | 234.7935 | 180.7003 | 235.4354 | 161.5119
2012-10-30 15:48 | 26.5627 | 247.7008 | 193.6191 | 248.0022 | 174.4272 | 70.5325 | 1.2376 | 31.2554 | 38.6803 | 25.2557 | 1.9511 | 31.1849 | 60.846 | 15.9829 | 12.9124 | 4.8625 | 99.5263 | 234.7884 | 180.7067 | 235.0898 | 161.5148
2012-10-30 15:48 | 26.5672 | 247.5996 | 193.5883 | 247.6661 | 174.3888 | 70.5296 | 1.2369 | 31.2378 | 38.6373 | 25.2555 | 1.9501 | 31.1673 | 60.778 | 15.9827 | 12.9373 | 4.8623 | 99.4153 | 234.6623 | 180.651 | 234.7288 | 161.4515
2012-10-30 15:48 | 26.5672 | 247.5487 | 193.5725 | 247.3334 | 174.3068 | 70.5758 | 1.3425 | 33.9299 | 45.5509 | 25.2737 | 2.1208 | 33.8938 | 71.8814 | 15.9817 | 12.9496 | 4.8655 | 117.4323 | 234.5991 | 180.6229 | 234.3838 | 161.3572
2012-10-30 15:48 | 26.5769 | 247.6367 | 193.8495 | 247.0103 | 174.1876 | 70.9971 | 1.4117 | 35.7236 | 50.4304 | 25.3057 | 2.2299 | 35.642 | 79.4793 | 15.9834 | 12.9266 | 4.8945 | 129.9098 | 234.7101 | 180.9229 | 234.0837 | 161.261
2012-10-30 15:48 | 26.5762 | 248.141 | 194.226 | 246.8569 | 174.1674 | 71.1104 | 1.4115 | 35.7237 | 50.4245 | 25.3088 | 2.2299 | 35.6422 | 79.4798 | 15.9835 | 12.983 | 4.9024 | 129.9043 | 235.158 | 181.243 | 233.8739 | 161.1844
2012-10-30 15:48 | 26.5844 | 248.5601 | 194.415 | 246.7437 | 174.1288 | 71.1497 | 1.4112 | 35.7187 | 50.4075 | 25.3102 | 2.2296 | 35.6372 | 79.457 | 15.9836 | 12.9831 | 4.9051 | 129.8646 | 235.577 | 181.4319 | 233.7606 | 161.1457
2012-10-30 15:48 | 26.5889 | 248.9755 | 194.5099 | 246.6973 | 173.9693 | 71.1942 | 1.4114 | 35.7239 | 50.4192 | 25.3117 | 2.2299 | 35.6425 | 79.4807 | 15.9836 | 13.0202 | 4.9081 | 129.8999 | 235.9553 | 181.4897 | 233.6771 | 160.9491
2012-10-30 15:48 | 26.6001 | 249.4641 | 194.6823 | 246.7437 | 174.0937 | 71.23 | 1.4113 | 35.7246 | 50.4184 | 25.3131 | 2.23 | 35.643 | 79.4827 | 15.9836 | 12.9988 | 4.9106 | 129.9011 | 236.4653 | 181.6835 | 233.7449 | 161.0949
2012-10-30 15:48 | 26.6114 | 249.7923 | 194.78 | 246.7668 | 174.1597 | 71.2659 | 1.4115 | 35.7321 | 50.437 | 25.3144 | 2.2304 | 35.6505 | 79.516 | 15.9836 | 13.0182 | 4.9131 | 129.9531 | 236.7741 | 181.7618 | 233.7486 | 161.1415
2012-10-30 15:49 | 26.6091 | 250.1577 | 194.9349 | 246.8584 | 174.2355 | 71.2901 | 1.4089 | 35.6664 | 50.25 | 25.3152 | 2.2266 | 35.5884 | 79.2395 | 15.9836 | 13.0241 | 4.9147 | 129.4894 | 237.1336 | 181.9108 | 233.8343 | 161.2114
2012-10-30 15:49 | 26.6099 | 250.4402 | 195.0733 | 246.9336 | 174.5564 | 71.3254 | 1.4043 | 35.5474 | 49.9178 | 25.314 | 2.2189 | 35.4654 | 78.6933 | 15.9835 | 13.0004 | 4.9172 | 128.6111 | 237.4398 | 182.0729 | 233.9332 | 161.556
2012-10-30 15:49 | 26.6256 | 250.8 | 195.203 | 247.1412 | 174.6736 | 71.341 | 1.4028 | 35.5103 | 49.8124 | 25.3146 | 2.2166 | 35.4288 | 78.5304 | 15.9836 | 13.012 | 4.9182 | 128.3428 | 237.788 | 182.191 | 234.1292 | 161.6616
2012-10-30 15:49 | 26.6323 | 251.0883 | 195.3866 | 247.3045 | 174.778 | 71.371 | 1.4021 | 35.4938 | 49.765 | 25.3152 | 2.2156 | 35.4127 | 78.4591 | 15.9836 | 13.0391 | 4.9203 | 128.2241 | 238.0492 | 182.3475 | 234.2654 | 161.7389
2012-10-30 15:49 | 26.6398 | 251.4086 | 195.4884 | 247.4999 | 174.6176 | 71.4075 | 1.4018 | 35.4878 | 49.7469 | 25.3159 | 2.2152 | 35.4068 | 78.4329 | 15.9836 | 13.0588 | 4.9228 | 128.1798 | 238.3498 | 182.4296 | 234.4411 | 161.5588
2012-10-30 15:49 | 26.6383 | 251.6965 | 195.6678 | 247.6942 | 175.1745 | 71.4439 | 1.4016 | 35.4837 | 49.733 | 25.317 | 2.215 | 35.4034 | 78.418 | 15.9836 | 13.0492 | 4.9253 | 128.151 | 238.6473 | 182.6186 | 234.645 | 162.1253
2012-10-30 15:49 | 26.6488 | 251.9455 | 195.8591 | 247.9278 | 175.0952 | 71.4589 | 1.4016 | 35.4847 | 49.7352 | 25.3174 | 2.2151 | 35.4048 | 78.4239 | 15.9836 | 13.0352 | 4.9264 | 128.1591 | 238.9103 | 182.8239 | 234.8926 | 162.06
2012-10-30 15:49 | 26.6533 | 252.2627 | 195.9894 | 248.1359 | 174.6311 | 71.4855 | 1.422 | 36.0049 | 51.198 | 25.3204 | 2.2448 | 35.8801 | 80.544 | 15.9836 | 13.0234 | 4.9282 | 131.7419 | 239.2393 | 182.966 | 235.1125 | 161.6077
2012-10-30 15:49 | 26.6548 | 252.6199 | 196.219 | 248.4153 | 174.8668 | 71.7259 | 1.5007 | 38.0399 | 57.0854 | 25.3487 | 2.3737 | 37.9473 | 90.0756 | 15.9866 | 13.0249 | 4.9448 | 147.161 | 239.595 | 183.1941 | 235.3904 | 161.8419
2012-10-30 15:49 | 26.6802 | 255.1457 | 197.774 | 249.9237 | 176.322 | 72.0877 | 1.507 | 38.2199 | 57.5986 | 25.361 | 2.3852 | 38.1325 | 90.9535 | 15.9872 | 13.0829 | 4.9697 | 148.5521 | 242.0628 | 184.6911 | 236.8408 | 163.2391
2012-10-30 15:49 | 26.6832 | 257.487 | 199.0906 | 251.714 | 177.3177 | 72.305 | 1.5082 | 38.2586 | 57.7005 | 25.3675 | 2.3876 | 38.1713 | 91.1388 | 15.9872 | 13.0859 | 4.9847 | 148.8393 | 244.4011 | 186.0047 | 238.6281 | 164.2318
2012-10-30 15:49 | 26.6855 | 257.9809 | 199.3958 | 252.2323 | 177.6131 | 72.3483 | 1.5082 | 38.2609 | 57.7041 | 25.369 | 2.3878 | 38.1736 | 91.15 | 15.9871 | 13.1168 | 4.9877 | 148.8541 | 244.8641 | 186.279 | 239.1155 | 164.4963
2012-10-30 15:49 | 26.6945 | 258.4892 | 199.7746 | 252.7143 | 178.1343 | 72.4165 | 1.5171 | 38.492 | 58.3981 | 25.3713 | 2.4024 | 38.4076 | 92.2704 | 15.9872 | 13.1095 | 4.9924 | 150.6685 | 245.3797 | 186.6651 | 239.6048 | 165.0248
2012-10-30 15:49 | 26.693 | 258.9908 | 200.1587 | 253.1935 | 178.4418 | 72.5448 | 1.5487 | 39.3134 | 60.886 | 25.3842 | 2.4534 | 39.2272 | 96.2398 | 15.989 | 13.059 | 5.0012 | 157.1258 | 245.9318 | 187.0997 | 240.1345 | 165.3828
2012-10-30 15:49 | 26.6922 | 259.6451 | 200.6497 | 253.728 | 178.8282 | 72.639 | 1.5492 | 39.3316 | 60.9345 | 25.3876 | 2.4542 | 39.2416 | 96.3084 | 15.9893 | 13.0909 | 5.0077 | 157.2429 | 246.5542 | 187.5588 | 240.6371 | 165.7373
2012-10-30 15:49 | 26.6974 | 260.3014 | 201.1152 | 254.37 | 179.0642 | 72.7263 | 1.5492 | 39.3321 | 60.9319 | 25.3892 | 2.4543 | 39.2421 | 96.3101 | 15.9894 | 13.0838 | 5.0138 | 157.2421 | 247.2176 | 188.0314 | 241.2862 | 165.9804
2012-10-30 15:49 | 26.6952 | 260.8838 | 201.4396 | 254.86 | 179.2302 | 72.8043 | 1.5491 | 39.3324 | 60.9295 | 25.3906 | 2.4543 | 39.2425 | 96.3119 | 15.9895 | 13.0898 | 5.0191 | 157.2414 | 247.794 | 188.3498 | 241.7702 | 166.1404
2012-10-30 15:49 | 26.6937 | 261.4863 | 201.8356 | 255.3623 | 179.4321 | 72.8488 | 1.549 | 39.3328 | 60.9278 | 25.3919 | 2.4543 | 39.2428 | 96.313 | 15.9895 | 13.072 | 5.0222 | 157.2407 | 248.4143 | 188.7636 | 242.2903 | 166.3601
2012-10-30 15:49 | 26.6997 | 262.0611 | 202.2115 | 255.9696 | 179.9272 | 72.9083 | 1.5489 | 39.3328 | 60.9242 | 25.3934 | 2.4543 | 39.2428 | 96.3124 | 15.9896 | 13.131 | 5.0263 | 157.2365 | 248.9301 | 189.0805 | 242.8386 | 166.7962
2012-10-30 15:50 | 26.7064 | 262.6598 | 202.6078 | 256.5657 | 180.4309 | 72.9586 | 1.5489 | 39.3329 | 60.9209 | 25.3949 | 2.4543 | 39.2429 | 96.3118 | 15.9897 | 13.0888 | 5.0298 | 157.2327 | 249.571 | 189.519 | 243.4769 | 167.3421
2012-10-30 15:50 | 26.7147 | 263.8072 | 203.4347 | 257.7565 | 181.0381 | 73.0904 | 1.5486 | 39.3321 | 60.9107 | 25.398 | 2.4542 | 39.2419 | 96.3064 | 15.9899 | 13.1215 | 5.0389 | 157.2171 | 250.6857 | 190.3132 | 244.635 | 167.9166
2012-10-30 15:50 | 26.7289 | 265.018 | 204.2519 | 259.0267 | 181.7335 | 73.2123 | 1.5484 | 39.331 | 60.9004 | 25.401 | 2.4541 | 39.2408 | 96.3008 | 15.9899 | 13.1153 | 5.0473 | 157.2011 | 251.9027 | 191.1366 | 245.9114 | 168.6182
2012-10-30 15:50 | 26.7221 | 265.5559 | 204.6899 | 259.6204 | 182.0831 | 73.2742 | 1.5484 | 39.3317 | 60.8999 | 25.402 | 2.4541 | 39.2415 | 96.3033 | 15.9901 | 13.0881 | 5.0515 | 157.2032 | 252.4678 | 191.6018 | 246.5323 | 168.995
2012-10-30 15:50 | 26.7274 | 266.0747 | 205.0416 | 260.1364 | 182.4643 | 73.3152 | 1.5483 | 39.3327 | 60.9002 | 25.4032 | 2.4542 | 39.2424 | 96.3073 | 15.9901 | 13.0771 | 5.0543 | 157.2074 | 252.9976 | 191.9645 | 247.0593 | 169.3872
2012-10-30 15:50 | 26.7326 | 266.5349 | 205.3854 | 260.6328 | 182.6654 | 73.3609 | 1.5483 | 39.3329 | 60.8979 | 25.4044 | 2.4542 | 39.2425 | 96.3082 | 15.9901 | 13.1068 | 5.0575 | 157.2061 | 253.4281 | 192.2786 | 247.526 | 169.5586
2012-10-30 15:50 | 26.7386 | 266.9648 | 205.7063 | 261.1454 | 182.9338 | 73.4152 | 1.5482 | 39.3333 | 60.8962 | 25.4057 | 2.4542 | 39.243 | 96.3098 | 15.9902 | 13.1046 | 5.0612 | 157.206 | 253.8602 | 192.6017 | 248.0408 | 169.8292
2012-10-30 15:50 | 26.7483 | 267.4372 | 206.0861 | 261.6889 | 183.2569 | 73.4799 | 1.5481 | 39.3333 | 60.8932 | 25.407 | 2.4542 | 39.2431 | 96.3095 | 15.9903 | 13.0777 | 5.0657 | 157.2026 | 254.3595 | 193.0084 | 248.6112 | 170.1792
2012-10-30 15:50 | 26.7476 | 267.9224 | 206.4239 | 262.1674 | 183.3346 | 73.5117 | 1.5481 | 39.3333 | 60.8909 | 25.4079 | 2.4542 | 39.243 | 96.3093 | 15.9903 | 13.0607 | 5.0679 | 157.2002 | 254.8617 | 193.3632 | 249.1067 | 170.2739
2012-10-30 15:50 | 26.7573 | 268.4179 | 206.7447 | 262.6874 | 183.6499 | 73.5764 | 1.548 | 39.333 | 60.8875 | 25.4089 | 2.4541 | 39.2426 | 96.3064 | 15.9905 | 13.0296 | 5.0724 | 157.1938 | 255.3883 | 193.7151 | 249.6578 | 170.6203
2012-10-30 15:50 | 26.7671 | 268.8318 | 207.0734 | 263.1995 | 184.2827 | 73.6117 | 1.5478 | 39.3302 | 60.8759 | 25.4101 | 2.454 | 39.2399 | 96.2934 | 15.9904 | 13.0435 | 5.0748 | 157.1692 | 255.7883 | 194.0299 | 250.156 | 171.2392
2012-10-30 15:50 | 26.7708 | 269.3056 | 207.4669 | 263.7756 | 184.7411 | 73.6591 | 1.5477 | 39.3288 | 60.8686 | 25.4113 | 2.4538 | 39.2384 | 96.285 | 15.9906 | 13.0186 | 5.0781 | 157.1537 | 256.287 | 194.4483 | 250.757 | 171.7225
2012-10-30 15:50 | 26.7723 | 269.8548 | 207.8582 | 264.357 | 184.9935 | 73.7197 | 1.5477 | 39.3296 | 60.8686 | 25.4125 | 2.4539 | 39.2393 | 96.2889 | 15.9907 | 13.0405 | 5.0822 | 157.1575 | 256.8143 | 194.8177 | 251.3165 | 171.953
2012-10-30 15:50 | 26.785 | 270.2947 | 208.2134 | 264.8601 | 185.1943 | 73.7764 | 1.5476 | 39.3305 | 60.8676 | 25.414 | 2.4539 | 39.24 | 96.2919 | 15.9907 | 13.045 | 5.0861 | 157.1595 | 257.2497 | 195.1684 | 251.8151 | 172.1493
2012-10-30 15:50 | 26.791 | 270.8248 | 208.6447 | 265.4148 | 185.4395 | 73.8295 | 1.5475 | 39.3297 | 60.8629 | 25.4149 | 2.4539 | 39.2393 | 96.2881 | 15.9908 | 13.0551 | 5.0898 | 157.151 | 257.7697 | 195.5896 | 252.3597 | 172.3844
2012-10-30 15:50 | 26.7925 | 271.331 | 209.0202 | 265.992 | 185.7625 | 73.8931 | 1.5474 | 39.3276 | 60.8537 | 25.4161 | 2.4537 | 39.2371 | 96.2765 | 15.991 | 13.0729 | 5.0942 | 157.1301 | 258.2581 | 195.9473 | 252.9191 | 172.6896
2012-10-30 15:50 | 26.8015 | 271.7631 | 209.3758 | 266.5027 | 185.7675 | 73.9359 | 1.5473 | 39.3269 | 60.8487 | 25.4172 | 2.4537 | 39.2365 | 96.2729 | 15.991 | 13.0575 | 5.0971 | 157.1216 | 258.7056 | 196.3183 | 253.4452 | 172.71
2012-10-30 15:50 | 26.7992 | 272.2109 | 209.6491 | 266.9399 | 186.2162 | 73.9844 | 1.5472 | 39.3265 | 60.845 | 25.4183 | 2.4536 | 39.2361 | 96.2707 | 15.991 | 13.0674 | 5.1005 | 157.1157 | 259.1435 | 196.5817 | 253.8725 | 173.1488
2012-10-30 15:50 | 26.8037 | 272.6228 | 209.9804 | 267.4653 | 186.6011 | 74.033 | 1.5608 | 39.6786 | 61.9294 | 25.4224 | 2.4747 | 39.5755 | 97.9387 | 15.9918 | 13.0352 | 5.1038 | 159.8681 | 259.5876 | 196.9452 | 254.4301 | 173.5659
2012-10-30 15:50 | 26.809 | 273.1131 | 210.3716 | 267.9488 | 186.9358 | 74.1393 | 1.5673 | 39.852 | 62.4593 | 25.4274 | 2.4861 | 39.7602 | 98.8495 | 15.9927 | 13.0976 | 5.1112 | 161.3088 | 260.0155 | 197.274 | 254.8512 | 173.8382
2012-10-30 15:50 | 26.8127 | 273.683 | 210.7809 | 268.5005 | 187.2847 | 74.2006 | 1.5672 | 39.8525 | 62.4578 | 25.4287 | 2.4861 | 39.7606 | 98.8501 | 15.993 | 13.0972 | 5.1154 | 161.3079 | 260.5858 | 197.6837 | 255.4033 | 174.1875
2012-10-30 15:50 | 26.8247 | 274.1448 | 211.1749 | 268.9943 | 187.5164 | 74.2601 | 1.5671 | 39.8522 | 62.4528 | 25.4303 | 2.4861 | 39.7605 | 98.8471 | 15.9933 | 13.1174 | 5.1195 | 161.2999 | 261.0274 | 198.0575 | 255.8769 | 174.399
2012-10-30 15:50 | 26.8254 | 274.6147 | 211.4828 | 269.399 | 187.7133 | 74.3144 | 1.567 | 39.8498 | 62.4429 | 25.4313 | 2.4859 | 39.758 | 98.8344 | 15.9934 | 13.1792 | 5.1232 | 161.2773 | 261.4355 | 198.3036 | 256.2198 | 174.5341
2012-10-30 15:51 | 26.8374 | 275.1035 | 211.8413 | 269.8849 | 188.149 | 74.367 | 1.5669 | 39.8502 | 62.4415 | 25.4324 | 2.4859 | 39.7584 | 98.8354 | 15.9936 | 13.2034 | 5.1269 | 161.2769 | 261.9001 | 198.6379 | 256.6815 | 174.9456
2012-10-30 15:51 | 26.8359 | 275.5167 | 212.1627 | 270.3378 | 188.5163 | 74.4156 | 1.5668 | 39.8503 | 62.4386 | 25.4337 | 2.4859 | 39.7582 | 98.8335 | 15.9937 | 13.1653 | 5.1302 | 161.2721 | 262.3514 | 198.9974 | 257.1725 | 175.351
2012-10-30 15:51 | 26.8427 | 275.9458 | 212.5043 | 270.8222 | 188.8253 | 74.4635 | 1.5666 | 39.8454 | 62.4207 | 25.4347 | 2.4855 | 39.7532 | 98.8066 | 15.994 | 13.1558 | 5.1335 | 161.2273 | 262.79 | 199.3485 | 257.6664 | 175.6695
2012-10-30 15:51 | 26.8457 | 276.4564 | 212.905 | 271.3764 | 189.2718 | 74.5242 | 1.5662 | 39.8371 | 62.3927 | 25.4356 | 2.485 | 39.7453 | 98.7661 | 15.9942 | 13.1547 | 5.1377 | 161.1588 | 263.3017 | 199.7503 | 258.2217 | 176.1171
2012-10-30 15:51 | 26.8509 | 276.9496 | 213.304 | 271.9095 | 189.7441 | 74.567 | 1.5663 | 39.8403 | 62.4003 | 25.4366 | 2.4852 | 39.7484 | 98.7809 | 15.9944 | 13.1191 | 5.1406 | 161.1812 | 263.8305 | 200.1849 | 258.7904 | 176.625
2012-10-30 15:51 | 26.8516 | 277.3763 | 213.7379 | 272.3177 | 190.051 | 74.5878 | 1.5661 | 39.8379 | 62.3897 | 25.4379 | 2.485 | 39.746 | 98.7676 | 15.9946 | 13.132 | 5.1421 | 161.1572 | 264.2443 | 200.6059 | 259.1857 | 176.919
2012-10-30 15:51 | 26.8591 | 277.8756 | 214.1116 | 272.8412 | 190.3137 | 74.6519 | 1.5661 | 39.8407 | 62.3949 | 25.4393 | 2.4851 | 39.7487 | 98.7802 | 15.9947 | 13.1518 | 5.1465 | 161.1751 | 264.7238 | 200.9598 | 259.6894 | 177.1619
2012-10-30 15:51 | 26.8651 | 278.3191 | 214.4838 | 273.398 | 190.6809 | 74.704 | 1.5661 | 39.842 | 62.3968 | 25.4402 | 2.4851 | 39.7499 | 98.7842 | 15.995 | 13.1782 | 5.1501 | 161.181 | 265.1409 | 201.3056 | 260.2198 | 177.5027
2012-10-30 15:51 | 26.8674 | 278.7084 | 214.7894 | 273.7726 | 190.8795 | 74.7178 | 1.5662 | 39.8455 | 62.4049 | 25.4413 | 2.4853 | 39.7535 | 98.8013 | 15.9951 | 13.1601 | 5.151 | 161.2062 | 265.5483 | 201.6293 | 260.6125 | 177.7194
2012-10-30 15:51 | 26.8734 | 279.1053 | 215.075 | 274.2052 | 191.235 | 74.7918 | 1.5619 | 39.7378 | 62.0677 | 25.4414 | 2.4789 | 39.6503 | 98.2887 | 15.9952 | 13.1701 | 5.1561 | 160.3564 | 265.9352 | 201.9049 | 261.0351 | 178.0649
2012-10-30 15:51 | 26.8741 | 279.5472 | 215.4458 | 274.675 | 191.5067 | 74.8086 | 1.5592 | 39.6672 | 61.8491 | 25.4408 | 2.4743 | 39.5757 | 97.9208 | 15.995 | 13.2035 | 5.1573 | 159.7699 | 266.3437 | 202.2423 | 261.4715 | 178.3032
2012-10-30 15:51 | 26.8883 | 279.9366 | 215.7593 | 275.1505 | 191.784 | 74.8502 | 1.5585 | 39.6514 | 61.7979 | 25.4416 | 2.4733 | 39.5602 | 97.843 | 15.9951 | 13.181 | 5.1602 | 159.6409 | 266.7556 | 202.5783 | 261.9695 | 178.603
2012-10-30 15:51 | 26.8891 | 280.3127 | 216.0279 | 275.5815 | 192.0558 | 74.8906 | 1.5581 | 39.6425 | 61.7677 | 25.4425 | 2.4727 | 39.5508 | 97.7967 | 15.9951 | 13.2388 | 5.163 | 159.5644 | 267.0739 | 202.7891 | 262.3427 | 178.817
2012-10-30 15:51 | 26.8876 | 280.6322 | 216.3217 | 275.9907 | 192.1643 | 74.9288 | 1.5579 | 39.6369 | 61.749 | 25.4431 | 2.4723 | 39.5453 | 97.769 | 15.9952 | 13.2373 | 5.1656 | 159.518 | 267.3949 | 203.0844 | 262.7534 | 178.927
2012-10-30 15:51 | 26.8943 | 281.0102 | 216.6513 | 276.3926 | 191.8764 | 74.9796 | 1.5579 | 39.6382 | 61.751 | 25.4439 | 2.4724 | 39.5467 | 97.7758 | 15.9952 | 13.2644 | 5.1691 | 159.5268 | 267.7458 | 203.3869 | 263.1282 | 178.612
2012-10-30 15:51 | 26.9018 | 281.4237 | 216.8203 | 276.8185 | 190.6783 | 75.016 | 1.5578 | 39.6383 | 61.7493 | 25.4447 | 2.4724 | 39.5466 | 97.7743 | 15.9953 | 13.2882 | 5.1716 | 159.5237 | 268.1355 | 203.5321 | 263.5303 | 177.3901
2012-10-30 15:51 | 26.9018 | 281.8141 | 216.9108 | 277.2871 | 191.503 | 75.06 | 1.5579 | 39.6425 | 61.7595 | 25.4459 | 2.4726 | 39.551 | 97.7949 | 15.9956 | 13.3453 | 5.1746 | 159.5545 | 268.4688 | 203.5655 | 263.9418 | 178.1577
2012-10-30 15:51 | 26.8951 | 282.1631 | 217.1246 | 277.6638 | 191.8537 | 75.0912 | 1.5582 | 39.6499 | 61.7806 | 25.4468 | 2.4731 | 39.5584 | 97.831 | 15.9956 | 13.3019 | 5.1768 | 159.6116 | 268.8612 | 203.8227 | 264.3619 | 178.5518
2012-10-30 15:51 | 26.8988 | 282.4914 | 217.341 | 278.0779 | 192.8355 | 75.1443 | 1.5591 | 39.6765 | 61.8614 | 25.4476 | 2.4749 | 39.5883 | 97.9777 | 15.9958 | 13.3627 | 5.1804 | 159.8391 | 269.1287 | 203.9783 | 264.7152 | 179.4728
2012-10-30 15:51 | 26.9153 | 282.9137 | 217.6882 | 278.4583 | 193.8421 | 75.2033 | 1.5638 | 39.7992 | 62.238 | 25.4503 | 2.4823 | 39.7071 | 98.5636 | 15.9963 | 13.3425 | 5.1845 | 160.8016 | 269.5712 | 204.3457 | 265.1158 | 180.4996
2012-10-30 15:51 | 26.9093 | 283.2632 | 218.072 | 278.8587 | 193.8989 | 75.2565 | 1.5645 | 39.8193 | 62.2979 | 25.4515 | 2.4835 | 39.7267 | 98.6601 | 15.9964 | 13.308 | 5.1882 | 160.9579 | 269.9552 | 204.764 | 265.5507 | 180.5909
2012-10-30 15:51 | 26.9078 | 283.5746 | 218.3698 | 279.2365 | 194.0428 | 75.3079 | 1.565 | 39.8327 | 62.3373 | 25.4526 | 2.4843 | 39.7401 | 98.7255 | 15.9966 | 13.3065 | 5.1917 | 161.0627 | 270.2681 | 205.0633 | 265.93 | 180.7363
2012-10-30 15:51 | 26.916 | 283.9421 | 218.6261 | 279.5581 | 193.9724 | 75.3379 | 1.5653 | 39.8413 | 62.3621 | 25.4534 | 2.4848 | 39.7489 | 98.7677 | 15.9968 | 13.3677 | 5.1938 | 161.1298 | 270.5744 | 205.2584 | 266.1904 | 180.6047
2012-10-30 15:51 | 26.9168 | 284.2403 | 218.9025 | 279.9187 | 194.6059 | 75.3732 | 1.5654 | 39.8462 | 62.3749 | 25.4544 | 2.4851 | 39.7539 | 98.7913 | 15.9971 | 13.3725 | 5.1962 | 161.1663 | 270.8678 | 205.53 | 266.5462 | 181.2334
2012-10-30 15:52 | 26.9175 | 284.6542 | 219.2339 | 280.3063 | 194.9407 | 75.431 | 1.5654 | 39.8488 | 62.3811 | 25.4553 | 2.4852 | 39.7565 | 98.803 | 15.9973 | 13.3732 | 5.2002 | 161.184 | 271.281 | 205.8607 | 266.9331 | 181.5675
2012-10-30 15:52 | 26.925 | 285.0285 | 219.58 | 280.6889 | 195.2037 | 75.4639 | 1.5655 | 39.8521 | 62.3882 | 25.4566 | 2.4854 | 39.7597 | 98.8182 | 15.9974 | 13.3726 | 5.2025 | 161.2064 | 271.6559 | 206.2074 | 267.3163 | 181.8311
2012-10-30 15:52 | 26.9265 | 285.3273 | 219.9004 | 281.0501 | 195.4764 | 75.5015 | 1.5654 | 39.8518 | 62.3849 | 25.4575 | 2.4853 | 39.7592 | 98.8146 | 15.9976 | 13.3904 | 5.2051 | 161.1995 | 271.9369 | 206.51 | 267.6597 | 182.086
2012-10-30 15:52 | 26.94 | 285.6959 | 220.1619 | 281.4619 | 195.6864 | 75.5483 | 1.5652 | 39.847 | 62.369 | 25.458 | 2.485 | 39.7546 | 98.7909 | 15.9977 | 13.4283 | 5.2083 | 161.1599 | 272.2676 | 206.7336 | 268.0336 | 182.2581
2012-10-30 15:52 | 26.94 | 286.0664 | 220.5123 | 281.8176 | 196.0167 | 75.5957 | 1.565 | 39.844 | 62.3564 | 25.4592 | 2.4848 | 39.7516 | 98.7753 | 15.9978 | 13.465 | 5.2116 | 161.1317 | 272.6014 | 207.0473 | 268.3526 | 182.5517
2012-10-30 15:52 | 26.9332 | 286.38 | 220.8047 | 282.1896 | 196.2772 | 75.6431 | 1.565 | 39.8434 | 62.3532 | 25.4598 | 2.4847 | 39.751 | 98.7713 | 15.998 | 13.4664 | 5.2148 | 161.1245 | 272.9136 | 207.3383 | 268.7232 | 182.8108
2012-10-30 15:52 | 26.9407 | 286.7155 | 221.0209 | 282.5024 | 196.3476 | 75.6691 | 1.5649 | 39.8427 | 62.3486 | 25.4607 | 2.4847 | 39.7503 | 98.7676 | 15.998 | 13.482 | 5.2166 | 161.1162 | 273.2335 | 207.5389 | 269.0204 | 182.8656
2012-10-30 15:52 | 26.946 | 286.9985 | 221.3332 | 282.8401 | 196.6084 | 75.7147 | 1.5648 | 39.8428 | 62.3472 | 25.4614 | 2.4847 | 39.7504 | 98.7665 | 15.9983 | 13.5281 | 5.2198 | 161.1137 | 273.4704 | 207.8051 | 269.312 | 183.0803
2012-10-30 15:52 | 26.9445 | 287.3057 | 221.5443 | 283.2134 | 196.8979 | 75.7552 | 1.5649 | 39.8458 | 62.3548 | 25.4622 | 2.4848 | 39.7534 | 98.7807 | 15.9984 | 13.547 | 5.2226 | 161.1355 | 273.7587 | 207.9973 | 269.6664 | 183.3509
2012-10-30 15:52 | 26.9579 | 287.6238 | 221.8805 | 283.5822 | 197.2219 | 75.7974 | 1.5649 | 39.8472 | 62.3569 | 25.463 | 2.4849 | 39.7547 | 98.7863 | 15.9985 | 13.5726 | 5.2255 | 161.1432 | 274.0512 | 208.3079 | 270.0096 | 183.6493
2012-10-30 15:52 | 26.9602 | 288.0234 | 222.1741 | 283.909 | 197.5388 | 75.839 | 1.5648 | 39.8463 | 62.3522 | 25.4639 | 2.4848 | 39.7537 | 98.7812 | 15.9986 | 13.5382 | 5.2283 | 161.1334 | 274.4852 | 208.6359 | 270.3708 | 184.0006
2012-10-30 15:52 | 26.9609 | 288.2554 | 222.4385 | 284.2573 | 197.799 | 75.8771 | 1.5648 | 39.8465 | 62.351 | 25.4646 | 2.4848 | 39.7541 | 98.7817 | 15.9988 | 13.5552 | 5.231 | 161.1327 | 274.7002 | 208.8833 | 270.7021 | 184.2438
2012-10-30 15:52 | 26.9684 | 288.5984 | 222.7452 | 284.5467 | 198.0543 | 75.9228 | 1.5648 | 39.8478 | 62.3532 | 25.4653 | 2.4849 | 39.7553 | 98.7867 | 15.9989 | 13.5342 | 5.2341 | 161.1398 | 275.0642 | 209.211 | 271.0125 | 184.5201
2012-10-30 15:52 | 26.9722 | 288.849 | 222.9418 | 284.8517 | 198.3295 | 75.9586 | 1.5647 | 39.8476 | 62.3504 | 25.4663 | 2.4848 | 39.7551 | 98.785 | 15.9991 | 13.5339 | 5.2366 | 161.1354 | 275.3151 | 209.4079 | 271.3178 | 184.7956
2012-10-30 15:52 | 26.9737 | 289.1589 | 223.1834 | 285.1583 | 198.5669 | 75.995 | 1.5647 | 39.847 | 62.3469 | 25.467 | 2.4848 | 39.7545 | 98.7813 | 15.9992 | 13.5191 | 5.2391 | 161.1282 | 275.6398 | 209.6643 | 271.6392 | 185.0478
2012-10-30 15:52 | 26.9841 | 289.4275 | 223.4851 | 285.4775 | 198.8212 | 76.0361 | 1.5646 | 39.8476 | 62.3474 | 25.4675 | 2.4848 | 39.755 | 98.7831 | 15.9993 | 13.5376 | 5.2419 | 161.1305 | 275.8899 | 209.9475 | 271.9399 | 185.2836
2012-10-30 15:52 | 26.9864 | 289.6919 | 223.7117 | 285.8272 | 198.9651 | 76.0667 | 1.5646 | 39.8477 | 62.3455 | 25.4684 | 2.4848 | 39.755 | 98.7818 | 15.9995 | 13.5359 | 5.244 | 161.1273 | 276.156 | 210.1758 | 272.2913 | 185.4292
2012-10-30 15:52 | 26.9946 | 289.9969 | 223.9994 | 286.04 | 199.1542 | 76.1754 | 1.6284 | 41.5005 | 67.5811 | 25.4848 | 2.5836 | 41.349 | 106.8311 | 16.0042 | 13.5155 | 5.2515 | 174.4122 | 276.4814 | 210.4839 | 272.5245 | 185.6387
2012-10-30 15:52 | 27.0029 | 290.4291 | 224.4956 | 286.3493 | 199.3749 | 76.388 | 1.6573 | 42.2657 | 70.0483 | 25.5022 | 2.6337 | 42.1671 | 111.0549 | 16.0107 | 13.5116 | 5.2662 | 181.1032 | 276.9175 | 210.984 | 272.8377 | 185.8633
2012-10-30 15:52 | 27.0044 | 291.0817 | 225.0009 | 286.7251 | 199.6046 | 76.503 | 1.657 | 42.2614 | 70.0264 | 25.505 | 2.633 | 42.163 | 111.0171 | 16.013 | 13.5009 | 5.2741 | 181.0435 | 277.5808 | 211.5 | 273.2242 | 186.1037
2012-10-30 15:52 | 27.0103 | 291.7424 | 225.4121 | 287.1824 | 200.0196 | 76.5776 | 1.657 | 42.264 | 70.0309 | 25.5065 | 2.6329 | 42.1653 | 111.0183 | 16.0146 | 13.5231 | 5.2793 | 181.0491 | 278.2193 | 211.889 | 273.6593 | 186.4965
2012-10-30 15:52 | 27.0103 | 292.3161 | 225.8055 | 287.5065 | 200.3855 | 76.6365 | 1.6569 | 42.264 | 70.0267 | 25.5081 | 2.6327 | 42.1655 | 111.011 | 16.0158 | 13.5679 | 5.2833 | 181.0377 | 278.7482 | 212.2376 | 273.9386 | 186.8176
2012-10-30 15:52 | 27.0118 | 292.8797 | 226.1258 | 287.859 | 200.7096 | 76.6938 | 1.657 | 42.2682 | 70.0371 | 25.5094 | 2.6328 | 42.1695 | 111.0243 | 16.0169 | 13.5287 | 5.2873 | 181.0614 | 279.351 | 212.5971 | 274.3303 | 187.1809
2012-10-30 15:52 | 27.0186 | 293.437 | 226.526 | 288.3401 | 201.1492 | 76.7631 | 1.6603 | 42.357 | 70.3268 | 25.5111 | 2.6368 | 42.2371 | 111.3728 | 16.018 | 13.5803 | 5.292 | 181.6996 | 279.8567 | 212.9457 | 274.7598 | 187.5689
2012-10-30 15:52 | 27.0261 | 293.9678 | 227.0095 | 288.845 | 201.5266 | 76.9469 | 1.7359 | 44.3271 | 76.9459 | 25.5361 | 2.7578 | 44.2038 | 121.9049 | 16.0287 | 13.6 | 5.3047 | 198.8507 | 280.3678 | 213.4095 | 275.245 | 187.9266
2012-10-30 15:52 | 27.0365 | 294.7439 | 227.7241 | 289.3335 | 201.6826 | 77.1596 | 1.7393 | 44.4311 | 77.2805 | 25.5449 | 2.7643 | 44.3273 | 122.5356 | 16.0355 | 13.6186 | 5.3194 | 199.8161 | 281.1253 | 214.1055 | 275.7149 | 188.064
2012-10-30 15:53 | 27.0463 | 295.6268 | 228.2885 | 289.8559 | 201.9639 | 77.2532 | 1.739 | 44.4272 | 77.2611 | 25.5469 | 2.7635 | 44.3235 | 122.4898 | 16.0387 | 13.6488 | 5.3258 | 199.7509 | 281.978 | 214.6397 | 276.2071 | 188.3151
2012-10-30 15:53 | 27.0395 | 296.4045 | 228.7812 | 290.3848 | 202.3112 | 77.3329 | 1.7383 | 44.4127 | 77.2047 | 25.5488 | 2.7622 | 44.3091 | 122.3912 | 16.0411 | 13.6338 | 5.3313 | 199.5959 | 282.7707 | 215.1474 | 276.751 | 188.6774
2012-10-30 15:53 | 27.0418 | 297.2448 | 229.2673 | 291.042 | 202.77 | 77.4063 | 1.7382 | 44.4134 | 77.2015 | 25.5506 | 2.7619 | 44.3095 | 122.3768 | 16.0433 | 13.6116 | 5.3364 | 199.5783 | 283.6332 | 215.6557 | 277.4304 | 189.1584
2012-10-30 15:53 | 27.0493 | 298.0824 | 229.8568 | 291.6928 | 203.1986 | 77.4965 | 1.7383 | 44.4192 | 77.2147 | 25.553 | 2.7619 | 44.3151 | 122.3934 | 16.0452 | 13.6273 | 5.3426 | 199.6081 | 284.4551 | 216.2295 | 278.0655 | 189.5713
2012-10-30 15:53 | 27.0598 | 298.8267 | 230.2998 | 292.3117 | 203.6224 | 77.5705 | 1.737 | 44.3872 | 77.0993 | 25.5544 | 2.7597 | 44.2832 | 122.2082 | 16.0464 | 13.6093 | 5.3477 | 199.3075 | 285.2174 | 216.6905 | 278.7024 | 190.0131
2012-10-30 15:53 | 27.065 | 299.5655 | 230.824 | 293.0138 | 204.0015 | 77.6421 | 1.7381 | 44.419 | 77.2033 | 25.5565 | 2.7614 | 44.3145 | 122.3698 | 16.0479 | 13.5656 | 5.3526 | 199.5731 | 285.9999 | 217.2584 | 279.4482 | 190.4359
2012-10-30 15:53 | 0 | 0 | 0 | 0 | 0 | 0 | 1.737 | 44.3952 | 77.1158 | 25.5581 | 2.7598 | 44.2921 | 122.2365 | 16.0491 | 0 | 0 | 199.3523 | 0 | 0 | 0 | 0
2012-10-30 15:53 | [5 empty] | 71.5745 | 1.4454 | 34.2332 | 49.4807 | 23.6842 | 2.2769 | 36.4317 | 82.9509 | 16.0006 | -13.312 | 4.9343 | 132.4317 | 13.312 | 13.312 | 13.312 | 13.312
2012-10-30 15:54 | 27.0141 | 243.1572 | 198.3006 | 248.0426 | 180.2181 | 71.8299 | 1.444 | 36.5123 | 52.7222 | 25.2862 | 2.2766 | 36.4284 | 82.9322 | 16.0014 | 13.6818 | 4.952 | 135.6544 | 229.4754 | 184.6188 | 234.3608 | 166.5363
2012-10-30 15:54 | 27.0156 | 244.2563 | 198.5852 | 247.1514 | 179.8636 | 71.9467 | 1.4433 | 36.5042 | 52.6874 | 25.2917 | 2.2761 | 36.4204 | 82.896 | 16.0013 | 13.7036 | 4.96 | 135.5834 | 230.5527 | 184.8816 | 233.4478 | 166.16
2012-10-30 15:54 | 27.0103 | 245.2032 | 198.6861 | 246.2218 | 179.5375 | 72.0108 | 1.443 | 36.5016 | 52.6727 | 25.2953 | 2.2759 | 36.4177 | 82.8838 | 16.0013 | 13.7146 | 4.9644 | 135.5565 | 231.4886 | 184.9715 | 232.5072 | 165.8229
2012-10-30 15:54 | 27.0201 | 246.0945 | 198.7628 | 245.5578 | 179.4651 | 72.0848 | 1.4428 | 36.5011 | 52.6651 | 25.2982 | 2.2759 | 36.4171 | 82.8811 | 16.0013 | 13.7611 | 4.9695 | 135.5461 | 232.3334 | 185.0017 | 231.7967 | 165.704
2012-10-30 15:54 | 27.0156 | 246.8498 | 198.7662 | 245.1223 | 179.3472 | 72.1362 | 1.4427 | 36.5025 | 52.6638 | 25.3007 | 2.276 | 36.4183 | 82.8871 | 16.0012 | 13.7647 | 4.9731 | 135.5509 | 233.0851 | 185.0015 | 231.3576 | 165.5825
2012-10-30 15:54 | 27.0201 | 247.5631 | 198.9044 | 244.8251 | 179.3204 | 72.1842 | 1.4427 | 36.5033 | 52.6623 | 25.3025 | 2.276 | 36.4193 | 82.8914 | 16.0012 | 13.7163 | 4.9764 | 135.5538 | 233.8468 | 185.1881 | 231.1088 | 165.6041
2012-10-30 15:54 | 27.0178 | 248.1833 | 198.9572 | 244.6151 | 179.4706 | 72.2455 | 1.4428 | 36.5088 | 52.6735 | 25.3049 | 2.2764 | 36.4247 | 82.917 | 16.001 | 13.714 | 4.9806 | 135.5904 | 234.4693 | 185.2432 | 230.9011 | 165.7566
2012-10-30 15:54 | 27.0126 | 248.8592 | 199.0778 | 244.4923 | 179.4576 | 72.2998 | 1.4426 | 36.507 | 52.6633 | 25.3073 | 2.2763 | 36.4231 | 82.9093 | 16.0011 | 13.7454 | 4.9843 | 135.5726 | 235.1138 | 185.3324 | 230.7469 | 165.7122
2012-10-30 15:55 | 27.0163 | 249.5125 | 199.2035 | 244.5431 | 179.3322 | 72.3373 | 1.4423 | 36.5044 | 52.6519 | 25.3091 | 2.2761 | 36.4204 | 82.8982 | 16.0009 | 13.7206 | 4.9869 | 135.5501 | 235.7919 | 185.4829 | 230.8225 | 165.6116
2012-10-30 15:55 | 27.0186 | 250.0861 | 199.3434 | 244.5963 | 179.2876 | 72.3633 | 1.4426 | 36.5131 | 52.6743 | 25.3105 | 2.2767 | 36.429 | 82.9381 | 16.0008 | 13.7311 | 4.9887 | 135.6124 | 236.355 | 185.6123 | 230.8652 | 165.5565
2012-10-30 15:55 | 27.0178 | 250.5899 | 199.4449 | 244.7679 | 179.3259 | 72.4096 | 1.4378 | 36.3935 | 52.3276 | 25.3114 | 2.2694 | 36.3124 | 82.4091 | 16.0006 | 13.7588 | 4.9919 | 134.7367 | 236.8311 | 185.6861 | 231.0091 | 165.5671
2012-10-30 15:55 | 27.0156 | 251.104 | 199.545 | 244.9969 | 179.3237 | 72.4385 | 1.4358 | 36.3432 | 52.1829 | 25.3116 | 2.2662 | 36.2594 | 82.1698 | 16.0003 | 13.7444 | 4.9939 | 134.3527 | 237.3596 | 185.8006 | 231.2525 | 165.5793
2012-10-30 15:55 | 27.0186 | 251.6076 | 199.6031 | 245.3173 | 179.3815 | 72.4737 | 1.4353 | 36.3327 | 52.1496 | 25.3131 | 2.2655 | 36.2489 | 82.1234 | 16.0001 | 13.7759 | 4.9963 | 134.2731 | 237.8317 | 185.8272 | 231.5414 | 165.6056
2012-10-30 15:55 | 27.0171 | 252.1847 | 199.7747 | 245.6371 | 179.1257 | 72.5142 | 1.4349 | 36.3249 | 52.1242 | 25.3145 | 2.2651 | 36.2411 | 82.0883 | 16 | 13.7377 | 4.9991 | 134.2125 | 238.447 | 186.037 | 231.8994 | 165.388
2012-10-30 15:55 | 27.0163 | 252.5905 | 199.9037 | 245.9654 | 178.988 | 72.5396 | 1.4345 | 36.3134 | 52.0902 | 25.3151 | 2.2644 | 36.2297 | 82.0383 | 15.9998 | 13.7817 | 5.0009 | 134.1285 | 238.8088 | 186.122 | 232.1837 | 165.2063
2012-10-30 15:55 | 27.0111 | 252.9723 | 199.9535 | 246.3362 | 178.3649 | 72.572 | 1.4343 | 36.3106 | 52.0795 | 25.3162 | 2.2642 | 36.2266 | 82.0248 | 15.9997 | 13.7562 | 5.0031 | 134.1043 | 239.2161 | 186.1973 | 232.58 | 164.6087
2012-10-30 15:55 | 27.0163 | 253.4738 | 200.0374 | 246.7722 | 178.8981 | 72.6228 | 1.4341 | 36.3078 | 52.0694 | 25.3173 | 2.2641 | 36.2239 | 82.0134 | 15.9995 | 13.7573 | 5.0066 | 134.0828 | 239.7165 | 186.2801 | 233.0149 | 165.1408
2012-10-30 15:55 | 27.0193 | 254.0083 | 200.1309 | 247.2021 | 178.6547 | 72.6777 | 1.4336 | 36.2966 | 52.0352 | 25.3183 | 2.2634 | 36.213 | 81.9653 | 15.9992 | 13.7888 | 5.0104 | 134.0006 | 240.2195 | 186.3421 | 233.4133 | 164.8659
2012-10-30 15:55 | 27.0118 | 254.3524 | 200.2021 | 247.5979 | 178.4516 | 72.6945 | 1.4334 | 36.2919 | 52.0195 | 25.3194 | 2.2631 | 36.2078 | 81.9416 | 15.9993 | 13.8017 | 5.0116 | 133.9611 | 240.5507 | 186.4004 | 233.7962 | 164.6499
2012-10-30 15:55 | 27.0141 | 254.7727 | 200.2162 | 248.0426 | 178.7512 | 72.7367 | 1.4333 | 36.2922 | 52.0178 | 25.3207 | 2.2631 | 36.2084 | 81.9447 | 15.9992 | 13.8081 | 5.0145 | 133.9625 | 240.9646 | 186.4081 | 234.2345 | 164.9431
2012-10-30 15:55 | 27.0014 | 255.1193 | 200.2625 | 248.484 | 178.4491 | 72.776 | 1.4369 | 36.3863 | 52.2846 | 25.3223 | 2.269 | 36.3017 | 82.3683 | 15.999 | 13.7628 | 5.0172 | 134.6528 | 241.3565 | 186.4997 | 234.7212 | 164.6863
2012-10-30 15:55 | 27.0096 | 255.5805 | 200.31 | 248.9345 | 178.9657 | 72.8216 | 1.4389 | 36.439 | 52.4325 | 25.324 | 2.2723 | 36.3544 | 82.6084 | 15.9989 | 13.8076 | 5.0203 | 135.0409 | 241.7729 | 186.5024 | 235.1269 | 165.1581
2012-10-30 15:55 | 27.0036 | 255.9494 | 200.4693 | 249.3433 | 179.1161 | 72.8627 | 1.4393 | 36.4521 | 52.4672 | 25.3255 | 2.2731 | 36.3675 | 82.6685 | 15.9988 | 13.8179 | 5.0232 | 135.1357 | 242.1315 | 186.6514 | 235.5254 | 165.2982
2012-10-30 15:55 | 27.0081 | 256.3092 | 200.6312 | 249.7665 | 179.4023 | 72.8945 | 1.4398 | 36.465 | 52.5031 | 25.3261 | 2.274 | 36.3804 | 82.728 | 15.9986 | 13.8469 | 5.0253 | 135.231 | 242.4623 | 186.7843 | 235.9196 | 165.5554
2012-10-30 15:55 | 27.0006 | 256.6569 | 200.7732 | 250.1659 | 179.9816 | 72.9482 | 1.44 | 36.4713 | 52.5189 | 25.3271 | 2.2744 | 36.3868 | 82.7578 | 15.9984 | 13.8271 | 5.029 | 135.2767 | 242.8298 | 186.9461 | 236.3388 | 166.1545
2012-10-30 15:55 | 26.9984 | 256.9943 | 200.9559 | 250.5079 | 180.1359 | 72.9557 | 1.4402 | 36.477 | 52.5333 | 25.3282 | 2.2748 | 36.3926 | 82.7863 | 15.9981 | 13.7924 | 5.0296 | 135.3196 | 243.2019 | 187.1635 | 236.7155 | 166.3435
2012-10-30 15:55 | 26.9924 | 257.3123 | 201.1545 | 251.0104 | 180.3021 | 72.9632 | 1.4404 | 36.4842 | 52.552 | 25.3291 | 2.2752 | 36.3996 | 82.8182 | 15.9981 | 13.8149 | 5.0301 | 135.3702 | 243.4974 | 187.3396 | 237.1955 | 166.4872
2012-10-30 15:55 | 27.0006 | 257.7536 | 201.3673 | 251.4253 | 180.3572 | 73.0129 | 1.4408 | 36.4957 | 52.5837 | 25.3299 | 2.2763 | 36.4164 | 82.8964 | 15.9978 | 13.7864 | 5.0335 | 135.48 | 243.9672 | 187.5809 | 237.6389 | 166.5708
2012-10-30 15:55 | 26.9954 | 258.1113 | 201.6297 | 251.8424 | 180.6534 | 73.0921 | 1.4721 | 37.3031 | 54.9152 | 25.3394 | 2.3268 | 37.2264 | 86.6201 | 15.9986 | 13.8423 | 5.039 | 141.5353 | 244.269 | 187.7874 | 238.0001 | 166.8111
2012-10-30 15:55 | 27.0044 | 258.5572 | 201.9889 | 252.258 | 180.7798 | 73.1811 | 1.4754 | 37.3912 | 55.166 | 25.3435 | 2.3325 | 37.3164 | 87.0394 | 15.9987 | 13.8798 | 5.0451 | 142.2055 | 244.6774 | 188.1091 | 238.3782 | 166.9
2012-10-30 15:55 | 27.0051 | 258.9714 | 202.2178 | 252.7239 | 180.9644 | 73.2724 | 1.5065 | 38.1931 | 57.5366 | 25.3528 | 2.3827 | 38.1225 | 90.835 | 15.9996 | 13.8683 | 5.0514 | 148.3716 | 245.1031 | 188.3495 | 238.8556 | 167.0961
2012-10-30 15:55 | 27.0103 | 259.546 | 202.648 | 253.2451 | 181.0244 | 73.4782 | 1.5503 | 39.3323 | 60.977 | 25.3708 | 2.453 | 39.2527 | 96.2863 | 16.002 | 13.9183 | 5.0656 | 157.2634 | 245.6277 | 188.7297 | 239.3268 | 167.1061
2012-10-30 15:55 | 27.0141 | 260.3842 | 203.132 | 253.8038 | 181.13 | 73.6672 | 1.5732 | 39.933 | 62.8243 | 25.3825 | 2.4895 | 39.8405 | 99.1847 | 16.0031 | 13.9221 | 5.0786 | 162.009 | 246.4621 | 189.2099 | 239.8817 | 167.2079
2012-10-30 15:55 | 27.0111 | 261.0984 | 203.495 | 254.2423 | 181.4872 | 73.7781 | 1.6009 | 40.653 | 65.0824 | 25.3935 | 2.5331 | 40.5413 | 102.6949 | 16.0046 | 13.9191 | 5.0863 | 167.7773 | 247.1793 | 189.5759 | 240.3232 | 167.5681
2012-10-30 15:55 | 27.0126 | 261.8402 | 204.0475 | 254.7399 | 181.7157 | 73.9434 | 1.6144 | 41.0109 | 66.2073 | 25.4034 | 2.5563 | 40.9153 | 104.5928 | 16.0055 | 13.9573 | 5.0977 | 170.8 | 247.8829 | 190.0902 | 240.7826 | 167.7584
2012-10-30 15:56 | 27.0178 | 262.8076 | 204.525 | 255.3818 | 182.0537 | 74.0434 | 1.6144 | 41.0141 | 66.2123 | 25.4055 | 2.5565 | 40.9179 | 104.6082 | 16.0052 | 13.9299 | 5.1046 | 170.8205 | 248.8777 | 190.5951 | 241.4519 | 168.1238
2012-10-30 15:56 | 27.0193 | 263.6464 | 204.9202 | 255.9885 | 182.4742 | 74.1162 | 1.6141 | 41.0113 | 66.1975 | 25.4078 | 2.5564 | 40.9153 | 104.5963 | 16.005 | 13.9558 | 5.1096 | 170.7938 | 249.6906 | 190.9644 | 242.0327 | 168.5184
2012-10-30 15:56 | 27.0276 | 264.4255 | 205.4402 | 256.6566 | 182.9055 | 74.185 | 1.6138 | 41.0063 | 66.1755 | 25.41 | 2.5562 | 40.91 | 104.5721 | 16.0045 | 13.9153 | 5.1143 | 170.7476 | 250.5102 | 191.5249 | 242.7413 | 168.9902
2012-10-30 15:56 | 27.0268 | 265.1292 | 205.9433 | 257.2647 | 183.3944 | 74.2595 | 1.6135 | 41.0026 | 66.1596 | 25.4115 | 2.556 | 40.9064 | 104.5565 | 16.0041 | 13.9593 | 5.1194 | 170.7161 | 251.1699 | 191.984 | 243.3054 | 169.4351
2012-10-30 15:56 | 27.0298 | 265.9064 | 206.3597 | 257.8843 | 183.75 | 74.3387 | 1.6132 | 40.9977 | 66.1385 | 25.4135 | 2.5557 | 40.9015 | 104.5327 | 16.0039 | 13.946 | 5.1249 | 170.6712 | 251.9604 | 192.4137 | 243.9383 | 169.804
2012-10-30 15:56 | 27.0276 | 266.6977 | 206.8063 | 258.6741 | 183.8888 | 74.3942 | 1.6129 | 40.9918 | 66.1142 | 25.4155 | 2.5554 | 40.8954 | 104.5052 | 16.0034 | 13.8908 | 5.1287 | 170.6194 | 252.8069 | 192.9155 | 244.7833 | 169.998
2012-10-30 15:56 | 0 | 251.3998 | 189.9263 | 242.1346 | 167.0918 | 76.0898 | 1.6126 | 40.9878 | 66.0963 | 25.4175 | 2.5553 | 40.8918 | 104.4893 | 16.003 | -3.5377 | 5.2456 | 170.5856 | 254.9375 | 193.464 | 245.6723 | 170.6295
2012-10-30 15:56 | 24.992 | 276.8679 | 215.2609 | 267.6161 | 192.5082 | 76.147 | 1.6126 | 40.9916 | 66.1039 | 25.4192 | 2.5555 | 40.8952 | 104.5092 | 16.0026 | 21.5224 | 5.2496 | 170.6131 | 255.3455 | 193.7385 | 246.0937 | 170.9858
2012-10-30 15:56 | 25.0003 | 277.5805 | 215.7654 | 268.3702 | 192.7523 | 76.2187 | 1.6127 | 40.9949 | 66.1109 | 25.4207 | 2.5558 | 40.8987 | 104.5279 | 16.0024 | 21.5627 | 5.2545 | 170.6388 | 256.0178 | 194.2027 | 246.8075 | 171.1896
2012-10-30 15:56 | 24.9913 | 278.2601 | 216.1698 | 269.1144 | 193.1049 | 76.28 | 1.6123 | 40.988 | 66.0841 | 25.4224 | 2.5554 | 40.8916 | 104.4936 | 16.0022 | 21.5217 | 5.2587 | 170.5778 | 256.7384 | 194.6481 | 247.5927 | 171.5832
2012-10-30 15:56 | 24.998 | 279.0211 | 216.6805 | 269.8664 | 193.4378 | 76.347 | 1.6123 | 40.9917 | 66.0907 | 25.4245 | 2.5556 | 40.8952 | 104.5137 | 16.0019 | 21.5284 | 5.2634 | 170.6044 | 257.4927 | 195.1521 | 248.338 | 171.9094
2012-10-30 15:56 | 25.0055 | 279.7864 | 217.1802 | 270.7432 | 193.803 | 76.3909 | 1.6122 | 40.9921 | 66.0874 | 25.4262 | 2.5557 | 40.8955 | 104.5164 | 16.0017 | 21.5038 | 5.2664 | 170.6038 | 258.2826 | 195.6764 | 249.2394 | 172.2992
2012-10-30 15:56 | 25.013 | 280.4549 | 217.6641 | 271.48 | 194.2312 | 76.4684 | 1.6121 | 40.9916 | 66.0825 | 25.4275 | 2.5557 | 40.8951 | 104.5157 | 16.0015 | 21.5394 | 5.2717 | 170.5982 | 258.9155 | 196.1247 | 249.9406 | 172.6918
2012-10-30 15:56 | 25.0175 | 280.9889 | 218.0268 | 272.074 | 194.4794 | 76.5215 | 1.612 | 40.9908 | 66.0763 | 25.4288 | 2.5557 | 40.8943 | 104.5124 | 16.0014 | 21.5279 | 5.2754 | 170.5887 | 259.461 | 196.4989 | 250.5461 | 172.9515
2012-10-30 15:56 | 25.0257 | 281.5959 | 218.413 | 272.7413 | 194.8926 | 76.592 | 1.6119 | 40.991 | 66.0732 | 25.4303 | 2.5557 | 40.8944 | 104.5146 | 16.0012 | 21.5 | 5.2803 | 170.5878 | 260.0959 | 196.913 | 251.2413 | 173.3926
2012-10-30 15:56 | 25.0265 | 282.2959 | 218.8941 | 273.498 | 195.3534 | 76.6654 | 1.6118 | 40.9913 | 66.0698 | 25.4319 | 2.5558 | 40.8948 | 104.5178 | 16.0009 | 21.5008 | 5.2853 | 170.5876 | 260.7951 | 197.3933 | 251.9972 | 173.8526
2012-10-30 15:56 | 25.0182 | 282.9093 | 219.3267 | 274.2066 | 195.7148 | 76.7099 | 1.6117 | 40.9907 | 66.0643 | 25.4334 | 2.5557 | 40.894 | 104.5144 | 16.0009 | 21.5326 | 5.2884 | 170.5787 | 261.3767 | 197.7941 | 252.674 | 174.1822
2012-10-30 15:56 | 25.0287 | 283.4642 | 219.723 | 274.8253 | 196.1107 | 76.777 | 1.6116 | 40.9893 | 66.0567 | 25.4346 | 2.5557 | 40.8926 | 104.5084 | 16.0007 | 21.5311 | 5.293 | 170.5651 | 261.9331 | 198.1919 | 253.2942 | 174.5796
2012-10-30 15:56 | 25.0354 | 284.0615 | 220.1469 | 275.5059 | 196.2904 | 76.8417 | 1.6115 | 40.9888 | 66.0514 | 25.436 | 2.5557 | 40.8921 | 104.5066 | 16.0006 | 21.5297 | 5.2975 | 170.558 | 262.5318 | 198.6172 | 253.9762 | 174.7607
2012-10-30 15:56 | 25.0377 | 284.6272 | 220.5468 | 276.1355 | 196.6545 | 76.8966 | 1.6113 | 40.9876 | 66.0443 | 25.4373 | 2.5556 | 40.8911 | 104.5016 | 16.0005 | 21.5481 | 5.3013 | 170.5459 | 263.0791 | 198.9987 | 254.5874 | 175.1064
2012-10-30 15:56 | 25.0429 | 285.1495 | 220.8984 | 276.8104 | 197.0452 | 76.9613 | 1.6112 | 40.9868 | 66.0384 | 25.4386 | 2.5556 | 40.8902 | 104.4989 | 16.0003 | 21.5372 | 5.3057 | 170.5372 | 263.6123 | 199.3612 | 255.2732 | 175.508
2012-10-30 15:56 | 25.0526 | 285.6992 | 221.345 | 277.4393 | 197.3971 | 77.0099 | 1.6111 | 40.9864 | 66.0342 | 25.4396 | 2.5556 | 40.8897 | 104.4965 | 16.0002 | 21.595 | 5.3091 | 170.5307 | 264.1042 | 199.75 | 255.8443 | 175.8021
2012-10-30 15:56 | 25.0556 | 286.3232 | 221.773 | 278.0266 | 197.7738 | 77.0769 | 1.611 | 40.9859 | 66.0291 | 25.441 | 2.5555 | 40.8891 | 104.494 | 16.0002 | 21.574 | 5.3137 | 170.5231 | 264.7492 | 200.199 | 256.4526 | 176.1998
2012-10-30 15:56 | 25.0579 | 286.846 | 222.1571 | 278.6594 | 198.0436 | 77.126 | 1.6109 | 40.985 | 66.0234 | 25.442 | 2.5555 | 40.8882 | 104.4896 | 16.0001 | 21.5803 | 5.3171 | 170.513 | 265.2657 | 200.5768 | 257.0791 | 176.4633
2012-10-30 15:56 | 25.0639 | 287.3917 | 222.5566 | 279.2803 | 198.3683 | 77.1653 | 1.6108 | 40.9849 | 66.0192 | 25.4435 | 2.5555 | 40.8881 | 104.4893 | 16.0001 | 21.5663 | 5.3198 | 170.5085 | 265.8254 | 200.9903 | 257.714 | 176.802
2012-10-30 15:56 | 25.0691 | 287.8671 | 222.8806 | 279.8809 | 198.6764 | 77.2301 | 1.6108 | 40.9861 | 66.0198 | 25.4448 | 2.5556 | 40.8893 | 104.4952 | 16.0001 | 21.5474 | 5.3242 | 170.5151 | 266.3197 | 201.3332 | 258.3335 | 177.129
2012-10-30 15:57 | 25.0713 | 288.4011 | 223.2763 | 280.5093 | 199.0209 | 77.296 | 1.6108 | 40.9896 | 66.0274 | 25.4462 | 2.5558 | 40.8927 | 104.5137 | 16 | 21.5897 | 5.3288 | 170.5411 | 266.8114 | 201.6866 | 258.9196 | 177.4312
2012-10-30 15:57 | 25.0796 | 288.972 | 223.7096 | 281.1437 | 199.3873 | 77.3497 | 1.6108 | 40.9907 | 66.0279 | 25.4474 | 2.5559 | 40.8939 | 104.5188 | 16.0001 | 21.63 | 5.3325 | 170.5468 | 267.342 | 202.0796 | 259.5137 | 177.7573
2012-10-30 15:57 | 25.087 | 289.5109 | 224.075 | 281.7268 | 199.6622 | 77.4023 | 1.6106 | 40.9864 | 66.011 | 25.4486 | 2.5556 | 40.8896 | 104.4966 | 16.0001 | 21.6214 | 5.3361 | 170.5076 | 267.8895 | 202.4536 | 260.1054 | 178.0408
2012-10-30 15:57 | 25.0863 | 289.9724 | 224.4796 | 282.2978 | 199.9763 | 77.4572 | 1.6105 | 40.9871 | 66.0104 | 25.4497 | 2.5556 | 40.8902 | 104.501 | 15.9999 | 21.5806 | 5.3399 | 170.5114 | 268.3918 | 202.899 | 260.7172 | 178.3957
2012-10-30 15:57 | 25.0938 | 290.5613 | 224.8136 | 282.8614 | 200.4797 | 77.5034 | 1.6105 | 40.989 | 66.0126 | 25.4511 | 2.5557 | 40.892 | 104.5095 | 16.0001 | 21.6242 | 5.3431 | 170.5221 | 268.9371 | 203.1894 | 261.2372 | 178.8555
2012-10-30 15:57 | 25.0893 | 290.988 | 225.2065 | 283.4167 | 200.8844 | 77.5791 | 1.6104 | 40.989 | 66.0104 | 25.452 | 2.5557 | 40.892 | 104.5093 | 16.0001 | 21.5756 | 5.3483 | 170.5196 | 269.4124 | 203.6309 | 261.8411 | 179.3088
2012-10-30 15:57 | 25.1058 | 291.4857 | 225.6361 | 284.0084 | 201.2788 | 77.6236 | 1.6104 | 40.9889 | 66.007 | 25.4532 | 2.5557 | 40.8919 | 104.509 | 16.0001 | 21.6482 | 5.3514 | 170.516 | 269.8375 | 203.9879 | 262.3602 | 179.6306
2012-10-30 15:57 | 25.1125 | 291.9465 | 225.9653 | 284.4589 | 201.5728 | 77.6595 | 1.6102 | 40.9872 | 65.999 | 25.4542 | 2.5556 | 40.8903 | 104.4996 | 16.0002 | 21.6269 | 5.3538 | 170.4986 | 270.3196 | 204.3384 | 262.832 | 179.9459
2012-10-30 15:57 | 25.1087 | 292.3429 | 226.3313 | 285.0069 | 201.8956 | 77.7173 | 1.6101 | 40.9856 | 65.9908 | 25.4553 | 2.5555 | 40.8885 | 104.4902 | 16.0003 | 21.6191 | 5.3578 | 170.481 | 270.7238 | 204.7122 | 263.3878 | 180.2765
2012-10-30 15:57 | 25.1162 | 292.8428 | 226.7242 | 285.5468 | 202.2377 | 77.7768 | 1.6099 | 40.9827 | 65.9789 | 25.4563 | 2.5553 | 40.8858 | 104.4766 | 16.0002 | 21.6386 | 5.3619 | 170.4555 | 271.2042 | 205.0856 | 263.9082 | 180.5991
2012-10-30 15:57 | 25.1312 | 293.4156 | 227.1482 | 286.1095 | 202.7723 | 77.823 | 1.6097 | 40.9791 | 65.964 | 25.4577 | 2.5551 | 40.8823 | 104.4585 | 16.0002 | 21.6616 | 5.3651 | 170.4225 | 271.754 | 205.4866 | 264.4479 | 181.1107
2012-10-30 15:57 | 25.1215 | 293.8097 | 227.5318 | 286.6628 | 202.9437 | 77.8641 | 1.6095 | 40.9767 | 65.9538 | 25.4586 | 2.5549 | 40.8796 | 104.4441 | 16.0003 | 21.6439 | 5.368 | 170.3979 | 272.1658 | 205.8879 | 265.0189 | 181.2998
2012-10-30 15:57 | 25.1327 | 294.3131 | 227.8851 | 287.1675 | 203.2028 | 77.9276 | 1.6095 | 40.9769 | 65.9519 | 25.4595 | 2.5549 | 40.8799 | 104.4453 | 16.0004 | 21.6551 | 5.3723 | 170.3972 | 272.658 | 206.23 | 265.5124 | 181.5477
2012-10-30 15:57 | 25.1417 | 294.7681 | 228.252 | 287.6545 | 203.4362 | 77.9663 | 1.6094 | 40.9768 | 65.9488 | 25.4606 | 2.5549 | 40.8796 | 104.445 | 16.0002 | 21.6481 | 5.375 | 170.3938 | 273.12 | 206.6039 | 266.0064 | 181.7881
2012-10-30 15:57 | 25.1439 | 295.197 | 228.5766 | 288.2425 | 203.7258 | 78.0132 | 1.6092 | 40.9739 | 65.9364 | 25.4618 | 2.5547 | 40.8769 | 104.4299 | 16.0004 | 21.7023 | 5.3782 | 170.3663 | 273.4947 | 206.8743 | 266.5402 | 182.0235
2012-10-30 15:57 | 25.1514 | 295.6003 | 228.8829 | 288.6623 | 203.9695 | 78.0455 | 1.6022 | 40.7924 | 65.3566 | 25.4606 | 2.5434 | 40.6951 | 103.5059 | 16 | 21.6898 | 5.3805 | 168.8625 | 273.9105 | 207.1931 | 266.9725 | 182.2797
2012-10-30 15:57 | 25.1544 | 295.9146 | 229.1808 | 289.1045 | 204.2835 | 78.0744 | 1.6024 | 40.7983 | 65.3746 | 25.4609 | 2.5439 | 40.7015 | 103.5388 | 15.9999 | 21.6768 | 5.3824 | 168.9133 | 274.2378 | 207.504 | 267.4277 | 182.6067
2012-10-30 15:57 | 25.1604 | 296.3049 | 229.4502 | 289.5227 | 204.7698 | 78.1062 | 1.6018 | 40.785 | 65.3293 | 25.462 | 2.543 | 40.6883 | 103.472 | 15.9999 | 21.6628 | 5.3846 | 168.8013 | 274.6421 | 207.7874 | 267.8599 | 183.107
2012-10-30 15:57 | 25.1649 | 296.7243 | 229.836 | 289.9971 | 205.1562 | 78.1594 | 1.6014 | 40.7771 | 65.3023 | 25.4628 | 2.5426 | 40.6804 | 103.4324 | 15.9998 | 21.7073 | 5.3883 | 168.7347 | 275.017 | 208.1287 | 268.2898 | 183.4489
2012-10-30 15:57 | 25.1761 | 297.1273 | 230.2049 | 290.4627 | 205.3052 | 78.2062 | 1.6012 | 40.7719 | 65.2832 | 25.4637 | 2.5422 | 40.6752 | 103.4053 | 15.9998 | 21.6905 | 5.3915 | 168.6885 | 275.4368 | 208.5144 | 268.7722 | 183.6147
2012-10-30 15:57 | 25.1791 | 297.4876 | 230.4949 | 290.9392 | 205.2255 | 78.2599 | 1.6011 | 40.7716 | 65.2801 | 25.4644 | 2.5422 | 40.6747 | 103.4032 | 15.9998 | 21.7255 | 5.3952 | 168.6833 | 275.7621 | 208.7694 | 269.2137 | 183.5
2012-10-30 15:57 | 25.1783 | 297.8516 | 230.7456 | 291.3465 | 205.1775 | 78.2883 | 1.6011 | 40.7729 | 65.283 | 25.4649 | 2.5423 | 40.6762 | 103.4102 | 15.9999 | 21.6887 | 5.3972 | 168.6932 | 276.1629 | 209.0569 | 269.6578 | 183.4888
2012-10-30 15:57 | 25.1836 | 298.1756 | 230.9553 | 291.6751 | 205.9624 | 78.331 | 1.6012 | 40.7752 | 65.2877 | 25.466 | 2.5424 | 40.6784 | 103.4206 | 16 | 21.706 | 5.4001 | 168.7083 | 276.4696 | 209.2493 | 269.9691 | 184.2564
2012-10-30 15:57 | 25.1821 | 298.5235 | 231.1857 | 292.0738 | 206.1223 | 78.3651 | 1.6011 | 40.7751 | 65.2853 | 25.4669 | 2.5424 | 40.6784 | 103.4205 | 16 | 21.6965 | 5.4025 | 168.7059 | 276.827 | 209.4892 | 270.3773 | 184.4258
2012-10-30 15:57 | 25.1851 | 298.9335 | 231.507 | 292.5232 | 206.834 | 78.4298 | 1.601 | 40.7746 | 65.281 | 25.4679 | 2.5423 | 40.6779 | 103.4171 | 16.0001 | 21.7395 | 5.407 | 168.6982 | 277.194 | 209.7675 | 270.7837 | 185.0945
2012-10-30 15:57 | 25.1918 | 299.2473 | 231.8438 | 292.8684 | 207.1912 | 78.4397 | 1.601 | 40.7763 | 65.2843 | 25.4687 | 2.5424 | 40.6796 | 103.4254 | 16.0002 | 21.7502 | 5.4076 | 168.7097 | 277.4971 | 210.0936 | 271.1182 | 185.441
2012-10-30 15:57 | 25.2015 | 299.6062 | 232.1796 | 293.2243 | 207.2324 | 78.4847 | 1.6044 | 40.8635 | 65.5597 | 25.4703 | 2.5478 | 40.7668 | 103.8671 | 16.0005 | 21.7759 | 5.4107 | 169.4268 | 277.8303 | 210.4037 | 271.4484 | 185.4565
2012-10-30 15:58 | 25.2045 | 300.0045 | 232.4734 | 293.6427 | 206.7944 | 78.561 | 1.6062 | 40.9123 | 65.712 | 25.472 | 2.5508 | 40.8148 | 104.1094 | 16.0009 | 21.7949 | 5.416 | 169.8213 | 278.2096 | 210.6785 | 271.8478 | 184.9995
2012-10-30 15:58 | 25.1963 | 300.2496 | 232.5791 | 294.0076 | 206.4278 | 78.5899 | 1.6064 | 40.9205 | 65.7359 | 25.4729 | 2.5513 | 40.8231 | 104.1512 | 16.001 | 21.7707 | 5.418 | 169.8871 | 278.4789 | 210.8084 | 272.2369 | 184.6571
2012-10-30 15:58 | 25.1948 | 300.6394 | 232.7583 | 294.3752 | 207.332 | 78.62 | 1.6066 | 40.9274 | 65.7547 | 25.4742 | 2.5517 | 40.8301 | 104.1855 | 16.0012 | 21.7692 | 5.4201 | 169.9402 | 278.8702 | 210.9891 | 272.606 | 185.5628
2012-10-30 15:58 | 25.194 | 301.053 | 233.0247 | 294.7589 | 207.8628 | 78.6882 | 1.6067 | 40.9315 | 65.7659 | 25.475 | 2.5519 | 40.8342 | 104.2056 | 16.0014 | 21.7564 | 5.4248 | 169.9715 | 279.2966 | 211.2683 | 273.0025 | 186.1064
2012-10-30 15:58 | 25.2 | 301.3467 | 233.3135 | 295.1647 | 208.109 | 78.7125 | 1.6067 | 40.9326 | 65.7679 | 25.4756 | 2.5519 | 40.8352 | 104.2089 | 16.0016 | 21.8024 | 5.4264 | 169.9769 | 279.5443 | 211.5111 | 273.3623 | 186.3066
2012-10-30 15:58 | 25.203 | 301.668 | 233.6111 | 295.4982 | 208.4979 | 78.7627 | 1.6067 | 40.9322 | 65.7639 | 25.4767 | 2.5519 | 40.8348 | 104.2069 | 16.0016 | 21.7854 | 5.4299 | 169.9708 | 279.8826 | 211.8257 | 273.7128 | 186.7125
2012-10-30 15:58 | 25.2023 | 302.0086 | 233.9207 | 295.8396 | 208.8752 | 78.8026 | 1.6065 | 40.9283 | 65.75 | 25.4773 | 2.5517 | 40.8309 | 104.1866 | 16.0017 | 21.8287 | 5.4327 | 169.9366 | 280.1799 | 212.092 | 274.0109 | 187.0465
2012-10-30 15:58 | 25.1963 | 302.3439 | 234.174 | 296.164 | 209.0425 | 78.8529 | 1.6055 | 40.9041 | 65.6709 | 25.4777 | 2.5502 | 40.8075 | 104.0678 | 16.0016 | 21.7787 | 5.4361 | 169.7387 | 280.5652 | 212.3953 | 274.3853 | 187.2638
2012-10-30 15:58 | 25.203 | 302.6497 | 234.5106 | 296.5319 | 209.1358 | 78.8864 | 1.6046 | 40.8827 | 65.6012 | 25.4782 | 2.5488 | 40.7856 | 103.9551 | 16.0018 | 21.7814 | 5.4384 | 169.5563 | 280.8683 | 212.7292 | 274.7505 | 187.3544
2012-10-30 15:58 | 25.2067 | 302.9371 | 234.7342 | 296.8468 | 209.0883 | 78.9055 | 1.6045 | 40.8797 | 65.5901 | 25.4788 | 2.5486 | 40.7823 | 103.9388 | 16.0017 | 21.7571 | 5.4397 | 169.5289 | 281.18 | 212.9771 | 275.0897 | 187.3312
2012-10-30 15:58 | 25.209 | 303.2576 | 234.9407 | 297.1871 | 209.2993 | 78.954 | 1.6043 | 40.8774 | 65.5807 | 25.4795 | 2.5485 | 40.7802 | 103.9275 | 16.0018 | 21.8154 | 5.4431 | 169.5083 | 281.4422 | 213.1253 | 275.3717 | 187.4839
2012-10-30 15:58 | 25.2127 | 303.5181 | 235.18 | 297.5173 | 209.6889 | 78.9887 | 1.6041 | 40.8734 | 65.5662 | 25.4801 | 2.5482 | 40.7761 | 103.9059 | 16.0019 | 21.7631 | 5.4455 | 169.472 | 281.755 | 213.4169 | 275.7542 | 187.9258
2012-10-30 15:58 | 25.2225 | 303.8844 | 235.4568 | 297.8267 | 209.7224 | 79.0344 | 1.604 | 40.8711 | 65.5569 | 25.4809 | 2.548 | 40.7737 | 103.8929 | 16.002 | 21.8049 | 5.4486 | 169.4498 | 282.0795 | 213.6519 | 276.0218 | 187.9175
2012-10-30 15:58 | 25.2142 | 304.1482 | 235.688 | 298.1832 | 209.9976 | 79.0656 | 1.6038 | 40.8679 | 65.545 | 25.4815 | 2.5478 | 40.7707 | 103.8767 | 16.0021 | 21.7526 | 5.4508 | 169.4218 | 282.3956 | 213.9354 | 276.4306 | 188.245
2012-10-30 15:58 | 25.218 | 304.4049 | 235.9156 | 298.498 | 210.3479 | 79.0945 | 1.6038 | 40.867 | 65.5405 | 25.4821 | 2.5478 | 40.7696 | 103.8711 | 16.0021 | 21.7644 | 5.4528 | 169.4116 | 282.6405 | 214.1512 | 276.7336 | 188.5835
2012-10-30 15:58 | 25.2255 | 304.719 | 236.0801 | 298.7858 | 210.6114 | 79.1361 | 1.6036 | 40.8658 | 65.5342 | 25.483 | 2.5477 | 40.7684 | 103.8639 | 16.0023 | 21.7719 | 5.4556 | 169.3981 | 282.9471 | 214.3082 | 277.0139 | 188.8395
2012-10-30 15:58 | 25.227 | 304.9465 | 236.3407 | 299.0713 | 210.8452 | 79.1875 | 1.6036 | 40.8663 | 65.5343 | 25.4837 | 2.5477 | 40.769 | 103.8655 | 16.0025 | 21.7774 | 5.4592 | 169.3997 | 283.1691 | 214.5633 | 277.2939 | 189.0678
2012-10-30 15:58 | 25.2262 | 305.279 | 236.6069 | 299.3853 | 211.0335 | 79.2135 | 1.6036 | 40.8661 | 65.5318 | 25.4844 | 2.5476 | 40.7688 | 103.8637 | 16.0026 | 21.7526 | 5.461 | 169.3955 | 283.5264 | 214.8543 | 277.6327 | 189.2809
2012-10-30 15:58 | 25.2329 | 305.5654 | 236.8608 | 299.7528 | 211.3119 | 79.2384 | 1.6036 | 40.8682 | 65.5368 | 25.485 | 2.5477 | 40.7708 | 103.8729 | 16.0028 | 21.7993 | 5.4627 | 169.4097 | 283.7661 | 215.0615 | 277.9535 | 189.5126
2012-10-30 15:58 | 25.2374 | 305.8074 | 237.0891 | 299.9683 | 211.6157 | 79.284 | 1.6036 | 40.8683 | 65.5351 | 25.4859 | 2.5477 | 40.7711 | 103.8733 | 16.003 | 21.7598 | 5.4658 | 169.4084 | 284.0476 | 215.3293 | 278.2085 | 189.8559
2012-10-30 15:58 | 25.2382 | 306.0686 | 237.3371 | 300.2761 | 211.9079 | 79.3279 | 1.6036 | 40.8695 | 65.5375 | 25.4864 | 2.5478 | 40.7721 | 103.8777 | 16.0031 | 21.7486 | 5.4689 | 169.4152 | 284.32 | 215.5885 | 278.5275 | 190.1593
2012-10-30 15:58 | 25.2516 | 306.331 | 237.637 | 300.6118 | 212.0552 | 79.3464 | 1.6035 | 40.8697 | 65.5355 | 25.4875 | 2.5478 | 40.7723 | 103.8783 | 16.0031 | 21.774 | 5.4701 | 169.4138 | 284.557 | 215.863 | 278.8378 | 190.2812
2012-10-30 15:58 | 25.2449 | 306.577 | 237.803 | 300.8507 | 212.399 | 79.3753 | 1.6036 | 40.8726 | 65.5433 | 25.488 | 2.5479 | 40.7753 | 103.8919 | 16.0034 | 21.7553 | 5.4721 | 169.4352 | 284.8217 | 216.0477 | 279.0954 | 190.6437
2012-10-30 15:58 | 25.2494 | 306.8151 | 238.0979 | 301.1122 | 212.7658 | 79.4181 | 1.6037 | 40.8754 | 65.5511 | 25.4885 | 2.5481 | 40.778 | 103.9047 | 16.0035 | 21.7598 | 5.4751 | 169.4558 | 285.0553 | 216.3381 | 279.3524 | 191.006
2012-10-30 15:58 | 25.2554 | 307.0892 | 238.3432 | 301.4519 | 212.9135 | 79.4377 | 1.6036 | 40.875 | 65.5473 | 25.4894 | 2.548 | 40.7775 | 103.9013 | 16.0037 | 21.7778 | 5.4764 | 169.4486 | 285.3114 | 216.5654 | 279.6741 | 191.1357
2012-10-30 15:58 | 25.2554 | 307.3533 | 238.5472 | 301.7204 | 213.008 | 79.4695 | 1.6035 | 40.8717 | 65.5358 | 25.4898 | 2.5478 | 40.7742 | 103.8848 | 16.0036 | 21.7858 | 5.4786 | 169.4206 | 285.5675 | 216.7614 | 279.9346 | 191.2222
2012-10-30 15:58 | 25.2666 | 307.5943 | 238.7742 | 302.0001 | 213.1846 | 79.5065 | 1.6035 | 40.8727 | 65.5373 | 25.4904 | 2.5479 | 40.7752 | 103.8893 | 16.0038 | 21.789 | 5.4812 | 169.4266 | 285.8053 | 216.9852 | 280.2111 | 191.3956
2012-10-30 15:58 | 25.2614 | 307.8264 | 238.9848 | 302.2365 | 213.4787 | 79.5285 | 1.6035 | 40.875 | 65.543 | 25.4911 | 2.548 | 40.7773 | 103.8997 | 16.0038 | 21.7517 | 5.4827 | 169.4427 | 286.0747 | 217.2331 | 280.4848 | 191.727
2012-10-30 15:58 | 25.2636 | 308.0812 | 239.2538 | 302.5071 | 213.4061 | 79.5608 | 1.6035 | 40.875 | 65.5417 | 25.4916 | 2.5479 | 40.777 | 103.8969 | 16.004 | 21.778 | 5.4849 | 169.4387 | 286.3032 | 217.4758 | 280.7291 | 191.6281
2012-10-30 15:59 | 25.2771 | 308.3856 | 239.4752 | 302.7775 | 213.6637 | 79.6342 | 1.6424 | 41.885 | 68.7913 | 25.5026 | 2.6113 | 41.8013 | 109.1563 | 16.0078 | 21.7915 | 5.49 | 177.9476 | 286.5941 | 217.6837 | 280.986 | 191.8722
2012-10-30 15:59 | 25.2846 | 309.3308 | 240.2985 | 303.475 | 214.1989 | 79.8406 | 1.6518 | 42.1402 | 69.6054 | 25.5123 | 2.6254 | 42.0396 | 110.3716 | 16.0125 | 21.803 | 5.5042 | 179.977 | 287.5278 | 218.4955 | 281.672 | 192.3959
2012-10-30 15:59 | 25.2958 | 310.746 | 241.2664 | 304.5785 | 214.9189 | 79.9885 | 1.6508 | 42.121 | 69.5331 | 25.5155 | 2.6238 | 42.0202 | 110.2529 | 16.015 | 21.7861 | 5.5144 | 179.7861 | 288.9599 | 219.4803 | 282.7924 | 193.1328
2012-10-30 15:59 | 25.3107 | 312.0344 | 242.1908 | 305.7811 | 215.6662 | 80.1249 | 1.6509 | 42.1294 | 69.552 | 25.5189 | 2.624 | 42.0287 | 110.2846 | 16.0169 | 21.745 | 5.5238 | 179.8366 | 290.2894 | 220.4458 | 284.0361 | 193.9212
2012-10-30 15:59 | 25.3265 | 313.2315 | 243.1002 | 306.9803 | 216.7216 | 80.2781 | 1.6513 | 42.1435 | 69.5895 | 25.5221 | 2.6246 | 42.0427 | 110.3467 | 16.0185 | 21.8128 | 5.5344 | 179.9362 | 291.4187 | 221.2874 | 285.1675 | 194.9088
2012-10-30 15:59 | 25.3452 | 314.37 | 244.118 | 308.1552 | 217.6144 | 80.4075 | 1.6521 | 42.1694 | 69.6668 | 25.5252 | 2.626 | 42.0683 | 110.4713 | 16.02 | 21.8155 | 5.5433 | 180.1382 | 292.5545 | 222.3025 | 286.3397 | 195.7989
2012-10-30 15:59 | 25.3616 | 315.4638 | 245.0314 | 309.3197 | 218.3513 | 80.5306 | 1.6461 | 42.0183 | 69.1662 | 25.526 | 2.6166 | 41.9185 | 109.6828 | 16.0204 | 21.8239 | 5.5518 | 178.849 | 293.6399 | 223.2075 | 287.4958 | 196.5274
2012-10-30 15:59 | 25.3743 | 316.42 | 245.7333 | 310.4496 | 218.4073 | 80.6352 | 1.6453 | 42.0016 | 69.1064 | 25.5278 | 2.6154 | 41.9011 | 109.5878 | 16.021 | 21.8606 | 5.559 | 178.6942 | 294.5594 | 223.8727 | 288.589 | 196.5467
2012-10-30 15:59 | 25.3811 | 317.4197 | 246.4213 | 311.4891 | 219.3865 | 80.76 | 1.6458 | 42.0167 | 69.1489 | 25.5304 | 2.6161 | 41.9157 | 109.6574 | 16.022 | 21.8754 | 5.5676 | 178.8064 | 295.5443 | 224.5459 | 289.6137 | 197.5111
2012-10-30 16:00 | 25.3908 | 318.3878 | 247.163 | 312.5236 | 219.4985 | 80.8831 | 1.6485 | 42.0917 | 69.3882 | 25.5334 | 2.6207 | 41.9917 | 110.0465 | 16.0232 | 21.9092 | 5.5761 | 179.4347 | 296.4786 | 225.2538 | 290.6144 | 197.5893
2012-10-30 16:00 | 25.3953 | 319.2931 | 247.7037 | 313.4837 | 219.9282 | 81.0091 | 1.6529 | 42.2103 | 69.7683 | 25.5376 | 2.6277 | 42.109 | 110.6496 | 16.0251 | 21.9137 | 5.5848 | 180.4179 | 297.3794 | 225.79 | 291.57 | 198.0145
2012-10-30 16:00 | 25.4013 | 319.8258 | 248.1949 | 314.0668 | 221.3591 | 81.0981 | 1.6532 | 42.2203 | 69.7977 | 25.5389 | 2.6282 | 42.119 | 110.6973 | 16.0258 | 21.9077 | 5.5909 | 180.4951 | 297.9181 | 226.2872 | 292.1591 | 199.4514
2012-10-30 16:00 | 25.4095 | 320.1165 | 248.4535 | 314.3808 | 221.5248 | 81.116 | 1.6595 | 42.3841 | 70.3355 | 25.5406 | 2.6374 | 42.269 | 111.482 | 16.0265 | 21.9038 | 5.5921 | 181.8175 | 298.2127 | 226.5497 | 292.477 | 199.621
2012-10-30 16:00 | 25.3975 | 320.4136 | 248.8014 | 314.6515 | 221.8631 | 81.268 | 1.7144 | 43.8201 | 75.1262 | 25.5596 | 2.7255 | 43.7052 | 119.1204 | 16.0354 | 21.8999 | 5.6026 | 194.2467 | 298.5137 | 226.9015 | 292.7516 | 199.9632
2012-10-30 16:00 | 25.4095 | 320.7842 | 249.2085 | 314.9577 | 222.0561 | 81.4299 | 1.7166 | 43.8852 | 75.334 | 25.565 | 2.7294 | 43.7795 | 119.493 | 16.0398 | 21.8958 | 5.6138 | 194.8269 | 298.8884 | 227.3127 | 293.0619 | 200.1603
2012-10-30 16:00 | 25.4162 | 321.3365 | 249.6024 | 315.2625 | 222.2557 | 81.4686 | 1.7163 | 43.8797 | 75.3104 | 25.5665 | 2.7287 | 43.774 | 119.4481 | 16.0418 | 21.9186 | 5.6164 | 194.7585 | 299.4179 | 227.6838 | 293.3439 | 200.3371
2012-10-30 16:00 | 25.411 | 321.7624 | 249.9491 | 315.5285 | 222.3882 | 81.5414 | 1.7164 | 43.8831 | 75.3191 | 25.5676 | 2.7287 | 43.7772 | 119.4542 | 16.0433 | 21.9614 | 5.6215 | 194.7733 | 299.801 | 227.9877 | 293.5671 | 200.4268
2012-10-30 16:00 | 25.4147 | 322.1438 | 250.2539 | 315.8455 | 222.6005 | 81.5616 | 1.716 | 43.8763 | 75.2932 | 25.5685 | 2.7281 | 43.7711 | 119.4121 | 16.0445 | 21.9451 | 5.6229 | 194.7053 | 300.1987 | 228.3088 | 293.9004 | 200.6554
2012-10-30 16:00 | 25.4267 | 322.6402 | 250.5631 | 316.2547 | 222.7975 | 81.6212 | 1.7163 | 43.8861 | 75.3229 | 25.5697 | 2.7285 | 43.7805 | 119.454 | 16.0458 | 21.9491 | 5.627 | 194.7769 | 300.6911 | 228.614 | 294.3056 | 200.8484
2012-10-30 16:00 | 25.417 | 323.0997 | 250.827 | 316.5888 | 223.0317 | 81.6865 | 1.7162 | 43.8856 | 75.3172 | 25.5712 | 2.7282 | 43.7796 | 119.4417 | 16.0468 | 21.9474 | 5.6315 | 194.7589 | 301.1523 | 228.8796 | 294.6414 | 201.0843
2012-10-30 16:00 | 25.4379 | 323.4676 | 251.1724 | 317.0336 | 223.3596 | 81.7136 | 1.7158 | 43.8759 | 75.2821 | 25.5717 | 2.7275 | 43.7702 | 119.3833 | 16.0477 | 21.9803 | 5.6333 | 194.6654 | 301.4873 | 229.1921 | 295.0533 | 201.3793
2012-10-30 16:00 | 25.432 | 323.9041 | 251.4596 | 317.4057 | 223.8259 | 81.761 | 1.7158 | 43.8783 | 75.2872 | 25.5728 | 2.7275 | 43.7722 | 119.3891 | 16.0484 | 21.9384 | 5.6366 | 194.6763 | 301.9657 | 229.5212 | 295.4673 | 201.8875
2012-10-30 16:00 | 25.4402 | 324.3317 | 251.7766 | 317.7729 | 224.0072 | 81.798 | 1.7161 | 43.8877 | 75.316 | 25.5739 | 2.728 | 43.782 | 119.4366 | 16.0492 | 21.9746 | 5.6392 | 194.7526 | 302.3571 | 229.802 | 295.7983 | 202.0326
2012-10-30 16:00 | 25.4387 | 324.7039 | 252.0721 | 318.1341 | 224.1867 | 81.8413 | 1.7161 | 43.8899 | 75.321 | 25.5749 | 2.728 | 43.7838 | 119.4406 | 16.05 | 22.0051 | 5.6421 | 194.7616 | 302.6988 | 230.067 | 296.129 | 202.1816
2012-10-30 16:00 | 25.4522 | 325.1482 | 252.3708 | 318.5484 | 224.3773 | 81.9015 | 1.716 | 43.8893 | 75.3158 | 25.5759 | 2.7278 | 43.7835 | 119.4344 | 16.0506 | 21.9946 | 5.6463 | 194.7502 | 303.1536 | 230.3762 | 296.5538 | 202.3827
2012-10-30 16:00 | 25.4469 | 325.5432 | 252.6937 | 318.9401 | 224.8323 | 81.9523 | 1.7161 | 43.8937 | 75.3282 | 25.5769 | 2.728 | 43.7879 | 119.4529 | 16.0513 | 21.9693 | 5.6498 | 194.7811 | 303.5739 | 230.7244 | 296.9708 | 202.863
2012-10-30 16:00 | 25.4514 | 325.8717 | 252.9756 | 319.3416 | 225.0021 | 81.9835 | 1.7161 | 43.8952 | 75.3301 | 25.578 | 2.728 | 43.7893 | 119.4563 | 16.0519 | 21.9658 | 5.6519 | 194.7864 | 303.9059 | 231.0098 | 297.3758 | 203.0363
2012-10-30 16:00 | 25.4634 | 326.3449 | 253.3235 | 319.6971 | 225.199 | 82.0413 | 1.7161 | 43.896 | 75.3302 | 25.5788 | 2.7279 | 43.7902 | 119.4557 | 16.0527 | 22.0298 | 5.6559 | 194.7859 | 304.3151 | 231.2937 | 297.6673 | 203.1692
2012-10-30 16:00 | 25.4671 | 326.684 | 253.6397 | 320.1931 | 225.4663 | 82.0806 | 1.716 | 43.8937 | 75.3196 | 25.5798 | 2.7277 | 43.7878 | 119.4383 | 16.0532 | 22.0575 | 5.6586 | 194.7579 | 304.6265 | 231.5822 | 298.1356 | 203.4088
2012-10-30 16:00 | 25.4649 | 327.1011 | 253.946 | 320.5573 | 225.9283 | 82.1251 | 1.7157 | 43.8891 | 75.3017 | 25.5805 | 2.7273 | 43.7831 | 119.409 | 16.0537 | 21.9953 | 5.6617 | 194.7107 | 305.1058 | 231.9507 | 298.562 | 203.933
2012-10-30 16:00 | 25.4716 | 327.3974 | 254.1753 | 320.9303 | 226.1317 | 82.1488 | 1.7155 | 43.8849 | 75.2842 | 25.5815 | 2.727 | 43.779 | 119.3834 | 16.0542 | 21.9619 | 5.6633 | 194.6676 | 305.4355 | 232.2134 | 298.9684 | 204.1698
2012-10-30 16:00 | 25.4806 | 327.7227 | 254.485 | 321.3246 | 226.4829 | 82.1893 | 1.715 | 43.8725 | 75.2403 | 25.582 | 2.7261 | 43.7669 | 119.314 | 16.0547 | 21.999 | 5.6661 | 194.5544 | 305.7237 | 232.486 | 299.3256 | 204.4839
2012-10-30 16:01 | 25.4821 | 328.0367 | 254.7207 | 321.6352 | 226.6142 | 82.228 | 1.7149 | 43.8712 | 75.2347 | 25.5824 | 2.7259 | 43.765 | 119.3 | 16.0551 | 21.9804 | 5.6688 | 194.5346 | 306.0563 | 232.7403 | 299.6548 | 204.6338
2012-10-30 16:01 | 25.4881 | 328.3933 | 254.9961 | 321.9731 | 226.8129 | 82.2667 | 1.715 | 43.8762 | 75.2478 | 25.5838 | 2.7262 | 43.7701 | 119.3248 | 16.0555 | 21.9905 | 5.6715 | 194.5726 | 306.4028 | 233.0056 | 299.9826 | 204.8224
2012-10-30 16:01 | 25.4926 | 328.775 | 255.3207 | 322.3705 | 227.1006 | 82.2985 | 1.7153 | 43.8852 | 75.2766 | 25.5845 | 2.7266 | 43.7792 | 119.3702 | 16.0561 | 21.9469 | 5.6737 | 194.6468 | 306.8281 | 233.3738 | 300.4236 | 205.1537
2012-10-30 16:01 | 25.4978 | 329.0659 | 255.5913 | 322.7342 | 227.4676 | 82.3395 | 1.7152 | 43.8826 | 75.2657 | 25.5851 | 2.7264 | 43.7764 | 119.3514 | 16.0566 | 22.0202 | 5.6765 | 194.6171 | 307.0457 | 233.5711 | 300.714 | 205.4474
2012-10-30 16:01 | 25.4918 | 329.3952 | 255.8663 | 323.102 | 227.5718 | 82.3927 | 1.715 | 43.8809 | 75.257 | 25.5861 | 2.7262 | 43.775 | 119.3403 | 16.057 | 22.0182 | 5.6802 | 194.5973 | 307.377 | 233.8481 | 301.0838 | 205.5536
2012-10-30 16:01 | 25.5045 | 329.7089 | 256.1444 | 323.477 | 227.8912 | 82.4251 | 1.7153 | 43.8886 | 75.281 | 25.5869 | 2.7266 | 43.7826 | 119.3778 | 16.0575 | 22.0429 | 5.6824 | 194.6588 | 307.666 | 234.1015 | 301.4341 | 205.8483
2012-10-30 16:01 | 25.5083 | 330.1393 | 256.4408 | 323.8736 | 228.072 | 82.4678 | 1.7152 | 43.8895 | 75.281 | 25.588 | 2.7266 | 43.7836 | 119.3796 | 16.058 | 22.0747 | 5.6853 | 194.6607 | 308.0646 | 234.3661 | 301.7989 | 205.9973
2012-10-30 16:01 | 25.5075 | 330.5271 | 256.7717 | 324.2732 | 228.2481 | 82.5077 | 1.7148 | 43.8793 | 75.2451 | 25.5883 | 2.7259 | 43.7735 | 119.3217 | 16.0584 | 22.0819 | 5.6881 | 194.5668 | 308.4452 | 234.6898 | 302.1913 | 206.1662
2012-10-30 16:01 | 25.5098 | 330.8151 | 256.9612 | 324.6034 | 228.5218 | 82.5499 | 1.7153 | 43.8927 | 75.2875 | 25.5894 | 2.7266 | 43.7867 | 119.3896 | 16.059 | 22.0402 | 5.691 | 194.6771 | 308.7749 | 234.921 | 302.5632 | 206.4816
2012-10-30 16:01 | 25.5232 | 331.1446 | 257.2321 | 325.0362 | 228.9992 | 82.6007 | 1.7151 | 43.8909 | 75.2789 | 25.5903 | 2.7264 | 43.7847 | 119.3747 | 16.0595 | 22.0135 | 5.6945 | 194.6536 | 309.1311 | 235.2186 | 303.0227 | 206.9857
2012-10-30 16:01 | 25.5217 | 331.5049 | 257.5543 | 325.3816 | 229.4223 | 82.6215 | 1.7149 | 43.8873 | 75.2636 | 25.5913 | 2.7261 | 43.781 | 119.3509 | 16.06 | 21.992 | 5.6959 | 194.6146 | 309.5129 | 235.5623 | 303.3896 | 207.4303
2012-10-30 16:01 | 25.53 | 331.8865 | 257.9487 | 325.8131 | 229.749 | 82.6799 | 1.7148 | 43.8872 | 75.2597 | 25.5925 | 2.726 | 43.7811 | 119.3475 | 16.0605 | 21.9723 | 5.7 | 194.6072 | 309.9142 | 235.9764 | 303.8408 | 207.7767
2012-10-30 16:01 | 25.5345 | 332.2794 | 258.3159 | 326.2483 | 229.8715 | 82.7244 | 1.7149 | 43.8892 | 75.2646 | 25.5932 | 2.7261 | 43.7832 | 119.3557 | 16.061 | 21.9688 | 5.703 | 194.6203 | 310.3106 | 236.3471 | 304.2795 | 207.9027
2012-10-30 16:01 | 25.5404 | 332.6395 | 258.5596 | 326.6049 | 230.1014 | 82.7724 | 1.715 | 43.8926 | 75.2743 | 25.5939 | 2.7262 | 43.7866 | 119.3704 | 16.0615 | 21.9627 | 5.7063 | 194.6447 | 310.6768 | 236.5969 | 304.6422 | 208.1387
2012-10-30 16:01 | 25.5449 | 332.922 | 258.8643 | 326.9677 | 230.499 | 82.8065 | 1.7148 | 43.8906 | 75.2645 | 25.5949 | 2.726 | 43.7844 | 119.3557 | 16.0619 | 21.9792 | 5.7087 | 194.6202 | 310.9428 | 236.8851 | 304.9885 | 208.5198
2012-10-30 16:01 | 25.5487 | 333.2952 | 259.1176 | 327.2764 | 230.7937 | 82.8417 | 1.7145 | 43.8827 | 75.2364 | 25.5952 | 2.7253 | 43.7742 | 119.2977 | 16.0622 | 21.983 | 5.7111 | 194.5341 | 311.3122 | 237.1346 | 305.2934 | 208.8107
2012-10-30 16:01 |  |  | 203.2754 | 268.1305 | 181.8694 | 76.5377 | 1.2389 | 31.052 | 38.4717 | 25.0633 | 1.9546 | 31.3031 | 61.1852 | 16.0151 | -13.2223 | 5.2765 | 99.657 | 13.2223 | 216.4977 | 281.3528 | 195.0917
2012-10-30 16:02 | 27.5973 | 287.7077 | 230.6722 | 292.7029 | 207.7579 | 77.185 | 1.5049 | 38.2485 | 57.5583 | 25.4168 | 2.3815 | 38.1569 | 90.8693 | 16.0224 | 14.2853 | 5.3211 | 148.4276 | 273.4224 | 216.3869 | 278.4176 | 193.4726
2012-10-30 16:02 | 27.5958 | 288.2501 | 230.8909 | 291.6808 | 207.4611 | 77.5786 | 1.619 | 41.2186 | 66.731 | 25.46 | 2.566 | 41.1317 | 105.5447 | 16.0294 | 14.3572 | 5.3483 | 172.2757 | 273.8929 | 216.5337 | 277.3236 | 193.1039
2012-10-30 16:02 | 27.5966 | 289.1302 | 231.2691 | 290.8532 | 207.1666 | 78.0132 | 1.7035 | 43.4316 | 73.9839 | 25.4962 | 2.6984 | 43.2721 | 116.7674 | 16.0359 | 14.3335 | 5.3782 | 190.7514 | 274.7967 | 216.9356 | 276.5197 | 192.8331
2012-10-30 16:02 | 27.6011 | 290.2141 | 231.7297 | 290.1679 | 206.868 | 78.3923 | 1.817 | 46.409 | 84.3258 | 25.5414 | 2.8847 | 46.3075 | 133.5834 | 16.0528 | 14.3421 | 5.4044 | 217.9091 | 275.872 | 217.3876 | 275.8258 | 192.5259
2012-10-30 16:02 | 27.6011 | 291.9364 | 232.6339 | 289.794 | 206.3759 | 78.9939 | 1.9176 | 49.0788 | 94.1125 | 25.5942 | 3.0461 | 48.9971 | 149.2495 | 16.0853 | 14.3054 | 5.4458 | 243.3619 | 277.631 | 218.3285 | 275.4886 | 192.0705
2012-10-30 16:02 | 27.6026 | 294.0825 | 233.4136 | 289.5835 | 206.7002 | 79.4146 | 1.97 | 50.4751 | 99.4349 | 25.6222 | 3.1251 | 50.3435 | 157.3303 | 16.1092 | 14.3232 | 5.4748 | 256.7652 | 279.7593 | 219.0904 | 275.2603 | 192.377
2012-10-30 16:02 | 27.5981 | 295.9556 | 234.2461 | 289.6522 | 207.5067 | 79.8174 | 2.0066 | 51.4648 | 103.2683 | 25.648 | 3.1822 | 51.3417 | 163.3794 | 16.134 | 14.282 | 5.5026 | 266.6477 | 281.6736 | 219.9641 | 275.3702 | 193.2247
2012-10-30 16:02 | 27.6093 | 297.7662 | 235.1375 | 289.9988 | 207.9589 | 79.9798 | 1.8321 | 46.9499 | 86.019 | 25.6256 | 2.8986 | 46.7389 | 135.4775 | 16.1247 | 14.334 | 5.5138 | 221.4965 | 283.4322 | 220.8035 | 275.6648 | 193.6249
2012-10-30 16:02 | 27.5981 | 299.2382 | 235.0163 | 290.4848 | 208.3217 | 78.7754 | 1.5579 | 39.7029 | 61.8549 | 25.4842 | 2.4668 | 39.6384 | 97.7806 | 16.0687 | 14.2902 | 5.4308 | 159.6355 | 284.948 | 220.7261 | 276.1946 | 194.0315
2012-10-30 16:02 | 27.61 | 299.3653 | 234.3445 | 291.0362 | 208.0698 | 78.5807 | 1.568 | 39.9594 | 62.6583 | 25.4835 | 2.4816 | 39.8649 | 98.9304 | 16.0639 | 14.3062 | 5.4174 | 161.5887 | 285.0591 | 220.0383 | 276.73 | 193.7636
2012-10-30 16:02 | 27.6078 | 299.4861 | 234.2205 | 291.5117 | 207.7527 | 78.5587 | 1.5692 | 39.9885 | 62.7504 | 25.4832 | 2.4841 | 39.8938 | 99.1006 | 16.0596 | 14.304 | 5.4158 | 161.851 | 285.1821 | 219.9165 | 277.2077 | 193.4487
2012-10-30 16:02 | 27.616 | 299.6325 | 234.2601 | 292.0515 | 208.8515 | 78.594 | 1.5702 | 40.014 | 62.8286 | 25.4839 | 2.4861 | 39.9188 | 99.2436 | 16.0565 | 14.3081 | 5.4183 | 162.0722 | 285.3244 | 219.952 | 277.7434 | 194.5434
2012-10-30 16:02 | 27.616 | 299.8131 | 234.3898 | 292.5676 | 209.0405 | 78.6038 | 1.5704 | 40.0199 | 62.8462 | 25.4843 | 2.4869 | 39.9249 | 99.289 | 16.0541 | 14.3244 | 5.4189 | 162.1352 | 285.4887 | 220.0654 | 278.2432 | 194.7161
2012-10-30 16:02 | 27.6213 | 300.0834 | 234.5523 | 292.9618 | 209.3411 | 78.6396 | 1.5346 | 39.1011 | 60.0051 | 25.4794 | 2.4329 | 39.049 | 95.0007 | 16.0507 | 14.3215 | 5.4214 | 155.0058 | 285.7619 | 220.2308 | 278.6403 | 195.0196
2012-10-30 16:02 | 27.6228 | 300.1118 | 234.4477 | 293.4715 | 209.3505 | 78.3143 | 1.3914 | 35.3887 | 49.2396 | 25.434 | 2.2024 | 35.3254 | 77.7997 | 16.0397 | 14.3475 | 5.399 | 127.0393 | 285.7643 | 220.1002 | 279.124 | 195.003
2012-10-30 16:02 | 27.6175 | 299.8107 | 234.0612 | 293.7895 | 209.4672 | 78.0559 | 1.396 | 35.4924 | 49.5464 | 25.4248 | 2.2083 | 35.4162 | 78.2094 | 16.0378 | 14.3829 | 5.3812 | 127.7557 | 285.4278 | 219.6783 | 279.4066 | 195.0843
2012-10-30 16:03 | 27.6235 | 299.5056 | 233.9218 | 294.1227 | 209.4575 | 78.0837 | 1.4027 | 35.6662 | 50.0291 | 25.4268 | 2.2187 | 35.5821 | 78.9466 | 16.0373 | 14.3686 | 5.3831 | 128.9757 | 285.137 | 219.5532 | 279.7541 | 195.0889
2012-10-30 16:03 | 27.6265 | 298.8747 | 233.8894 | 294.7183 | 209.2479 | 78.2484 | 1.4819 | 37.7146 | 55.8894 | 25.4501 | 2.3449 | 37.6108 | 88.1942 | 16.0392 | 14.3593 | 5.3944 | 144.0837 | 284.5154 | 219.5301 | 280.359 | 194.8886
2012-10-30 16:03 | 27.6332 | 299.3194 | 234.348 | 295.1559 | 209.1364 | 78.5391 | 1.5476 | 39.4245 | 61.0141 | 25.4743 | 2.452 | 39.3308 | 96.4381 | 16.0405 | 14.3823 | 5.4145 | 157.4521 | 284.9371 | 219.9657 | 280.7736 | 194.7541
2012-10-30 16:03 | 27.6332 | 299.5499 | 234.3834 | 295.3098 | 209.1797 | 78.5472 | 1.5476 | 39.4253 | 61.0142 | 25.4753 | 2.4522 | 39.3317 | 96.4487 | 16.0394 | 14.3823 | 5.415 | 157.4629 | 285.1676 | 220.0011 | 280.9275 | 194.7974
2012-10-30 16:03 | 27.6475 | 299.883 | 234.4881 | 295.5357 | 209.2492 | 78.5818 | 1.5476 | 39.4275 | 61.0197 | 25.4758 | 2.4525 | 39.3337 | 96.465 | 16.0384 | 14.3844 | 5.4174 | 157.4847 | 285.4986 | 220.1037 | 281.1513 | 194.8648
2012-10-30 16:03 | 27.6422 | 300.0928 | 234.5535 | 295.7035 | 209.4092 | 78.6038 | 1.5476 | 39.4264 | 61.0146 | 25.4765 | 2.4526 | 39.3326 | 96.4656 | 16.0374 | 14.375 | 5.4189 | 157.4802 | 285.7178 | 220.1785 | 281.3285 | 195.0342
2012-10-30 16:03 | 27.6377 | 300.2841 | 234.7023 | 295.8798 | 209.4953 | 78.6339 | 1.5475 | 39.4247 | 61.0082 | 25.4769 | 2.4526 | 39.331 | 96.4629 | 16.0365 | 14.3257 | 5.421 | 157.4712 | 285.9584 | 220.3766 | 281.5541 | 195.1696
2012-10-30 16:03 | 27.6497 | 300.4497 | 234.8086 | 296.1225 | 209.5546 | 78.635 | 1.5597 | 39.7407 | 61.9831 | 25.4799 | 2.4714 | 39.6315 | 97.9446 | 16.0361 | 14.3785 | 5.4211 | 159.9277 | 286.0712 | 220.4301 | 281.744 | 195.1761
2012-10-30 16:03 | 27.6475 | 300.7586 | 234.9989 | 296.2973 | 209.7532 | 78.7587 | 1.5775 | 40.2088 | 63.43 | 25.4887 | 2.5012 | 40.113 | 100.332 | 16.0373 | 14.3314 | 5.4296 | 163.7619 | 286.4272 | 220.6675 | 281.9659 | 195.4218
2012-10-30 16:03 | 27.6542 | 301.0763 | 235.1942 | 296.5501 | 209.8583 | 78.8182 | 1.5777 | 40.2152 | 63.4477 | 25.4896 | 2.5017 | 40.1193 | 100.3669 | 16.0367 | 14.3585 | 5.4337 | 163.8146 | 286.7178 | 220.8357 | 282.1916 | 195.4998
2012-10-30 16:03 | 27.6572 | 301.4209 | 235.37 | 296.7453 | 209.9046 | 78.8431 | 1.5775 | 40.2103 | 63.4303 | 25.4904 | 2.5015 | 40.1148 | 100.3481 | 16.0361 | 14.333 | 5.4354 | 163.7784 | 287.0879 | 221.037 | 282.4123 | 195.5716
2012-10-30 16:03 | 27.6669 | 301.6456 | 235.5172 | 297.0011 | 209.9734 | 78.8529 | 1.5772 | 40.2055 | 63.4136 | 25.4912 | 2.5013 | 40.1096 | 100.3262 | 16.0355 | 14.3794 | 5.4361 | 163.7398 | 287.2662 | 221.1378 | 282.6217 | 195.594
2012-10-30 16:03 | 27.6639 | 301.919 | 235.6281 | 297.1865 | 209.931 | 78.8962 | 1.5825 | 40.341 | 63.839 | 25.4922 | 2.509 | 40.2312 | 100.9398 | 16.0348 | 14.3764 | 5.4391 | 164.7788 | 287.5426 | 221.2517 | 282.8101 | 195.5546
2012-10-30 16:03 | 27.6654 | 302.262 | 235.8732 | 297.4801 | 210.1609 | 79.0078 | 1.648 | 42.048 | 69.2949 | 25.5147 | 2.6143 | 41.9366 | 109.6359 | 16.0411 | 14.3534 | 5.4468 | 178.9308 | 287.9086 | 221.5198 | 283.1267 | 195.8075
2012-10-30 16:03 | 27.6639 | 302.744 | 236.2802 | 297.7516 | 210.3484 | 79.2077 | 1.6487 | 42.0746 | 69.3694 | 25.5195 | 2.6164 | 41.9739 | 109.8206 | 16.0426 | 14.356 | 5.4606 | 179.19 | 288.388 | 221.9242 | 283.3956 | 195.9924
2012-10-30 16:03 | 27.6684 | 303.2512 | 236.5439 | 298.0751 | 210.4238 | 79.2626 | 1.6484 | 42.0674 | 69.3425 | 25.5206 | 2.6159 | 41.9666 | 109.7814 | 16.0427 | 14.3238 | 5.4644 | 179.1239 | 288.9274 | 222.2201 | 283.7513 | 196.1
2012-10-30 16:03 | 27.6782 | 303.6983 | 236.7894 | 298.3347 | 210.5754 | 79.3089 | 1.6477 | 42.0533 | 69.2931 | 25.5217 | 2.6151 | 41.9528 | 109.7112 | 16.0425 | 14.3907 | 5.4676 | 179.0042 | 289.3076 | 222.3987 | 283.944 | 196.1847
2012-10-30 16:03 | 27.6812 | 304.1999 | 237.0202 | 298.7258 | 210.8186 | 79.3551 | 1.6478 | 42.0582 | 69.3054 | 25.5231 | 2.6154 | 41.9572 | 109.7362 | 16.0421 | 14.3285 | 5.4707 | 179.0416 | 289.8714 | 222.6917 | 284.3973 | 196.4901
2012-10-30 16:03 | 27.6812 | 304.7061 | 237.3069 | 299.1023 | 211.1257 | 79.3961 | 1.6481 | 42.067 | 69.3307 | 25.5245 | 2.616 | 41.9663 | 109.7854 | 16.0419 | 14.3448 | 5.4736 | 179.1161 | 290.3613 | 222.9621 | 284.7575 | 196.7809
2012-10-30 16:03 | 27.6864 | 305.1369 | 237.5909 | 299.4686 | 211.4223 | 79.4487 | 1.6482 | 42.0721 | 69.3446 | 25.5256 | 2.6164 | 41.9712 | 109.8143 | 16.0415 | 14.3378 | 5.4772 | 179.1589 | 290.7991 | 223.2531 | 285.1308 | 197.0845
2012-10-30 16:04 | 27.6849 | 305.5457 | 237.821 | 299.8781 | 211.7673 | 79.4955 | 1.648 | 42.0677 | 69.3279 | 25.5264 | 2.6162 | 41.967 | 109.7956 | 16.041 | 14.3729 | 5.4804 | 179.1235 | 291.1728 | 223.4481 | 285.5052 | 197.3944
2012-10-30 16:04 | 27.6954 | 305.9855 | 238.1339 | 300.2266 | 211.8133 | 79.5239 | 1.6478 | 42.0649 | 69.3161 | 25.5274 | 2.6161 | 41.964 | 109.7815 | 16.0407 | 14.3508 | 5.4824 | 179.0975 | 291.6347 | 223.7831 | 285.8758 | 197.4625
2012-10-30 16:04 | 27.7021 | 306.433 | 238.439 | 300.6903 | 211.9381 | 79.5765 | 1.6478 | 42.0653 | 69.3143 | 25.5285 | 2.6162 | 41.9645 | 109.7881 | 16.0402 | 14.4023 | 5.486 | 179.1024 | 292.0307 | 224.0367 | 286.288 | 197.5358
2012-10-30 16:04 | 27.7164 | 306.934 | 238.7241 | 301.196 | 212.1729 | 79.6198 | 1.6479 | 42.0691 | 69.3241 | 25.5295 | 2.6165 | 41.9682 | 109.8098 | 16.0398 | 14.4166 | 5.489 | 179.1339 | 292.5174 | 224.3075 | 286.7794 | 197.7563
2012-10-30 16:04 | 27.7029 | 307.2577 | 238.9501 | 301.5664 | 212.4115 | 79.6707 | 1.6478 | 42.0681 | 69.3186 | 25.5303 | 2.6165 | 41.9671 | 109.8055 | 16.0396 | 14.3746 | 5.4925 | 179.1242 | 292.8831 | 224.5755 | 287.1918 | 198.0369
2012-10-30 16:04 | 27.7141 | 307.6099 | 239.185 | 301.9001 | 212.5251 | 79.7013 | 1.648 | 42.0744 | 69.3368 | 25.5313 | 2.6169 | 41.9734 | 109.8414 | 16.0391 | 14.4143 | 5.4946 | 179.1782 | 293.1956 | 224.7707 | 287.4858 | 198.1108
2012-10-30 16:04 | 27.7194 | 308.0481 | 239.4376 | 302.3467 | 212.4713 | 79.7533 | 1.6613 | 42.4223 | 70.4744 | 25.5362 | 2.6389 | 42.3292 | 111.7024 | 16.0405 | 14.4115 | 5.4982 | 182.1768 | 293.6366 | 225.0261 | 287.9352 | 198.0598
2012-10-30 16:04 | 27.7306 | 308.4921 | 239.7353 | 302.7954 | 212.4904 | 79.8527 | 1.6667 | 42.5684 | 70.9493 | 25.5403 | 2.6473 | 42.467 | 112.4226 | 16.0417 | 14.3901 | 5.505 | 183.3719 | 294.102 | 225.3452 | 288.4053 | 198.1003
2012-10-30 16:04 | 27.7268 | 308.9402 | 240.018 | 303.2405 | 212.5929 | 79.8822 | 1.6592 | 42.3736 | 70.3046 | 25.5392 | 2.6353 | 42.2721 | 111.3979 | 16.041 | 14.4148 | 5.5071 | 181.7024 | 294.5254 | 225.6032 | 288.8257 | 198.1781
2012-10-30 16:04 | 27.7403 | 309.3788 | 240.1767 | 303.6837 | 212.8702 | 79.918 | 1.6589 | 42.3678 | 70.2844 | 25.5395 | 2.6349 | 42.2647 | 111.3617 | 16.0405 | 14.4772 | 5.5095 | 181.6461 | 294.9016 | 225.6995 | 289.2065 | 198.393
2012-10-30 16:04 | 27.7426 | 309.7142 | 240.3674 | 304.0312 | 213.1561 | 79.9596 | 1.6603 | 42.4046 | 70.4025 | 25.541 | 2.6372 | 42.3025 | 111.5605 | 16.0407 | 14.4347 | 5.5124 | 181.963 | 295.2795 | 225.9327 | 289.5965 | 198.7214
2012-10-30 16:04 | 27.7351 | 310.1317 | 240.6071 | 304.4954 | 212.3334 | 79.9856 | 1.6597 | 42.3913 | 70.3564 | 25.5418 | 2.6364 | 42.2892 | 111.4908 | 16.0405 | 14.4231 | 5.5142 | 181.8471 | 295.7086 | 226.184 | 290.0723 | 197.9103
2012-10-30 16:04 | 27.7336 | 310.4977 | 240.7116 | 304.9119 | 213.2928 | 80.0313 | 1.6592 | 42.3807 | 70.3191 | 25.5425 | 2.6358 | 42.2788 | 111.4375 | 16.0404 | 14.4827 | 5.5174 | 181.7567 | 296.015 | 226.2289 | 290.4292 | 198.8101
2012-10-30 16:04 | 27.7456 | 310.8886 | 240.9786 | 305.2268 | 213.4111 | 80.0475 | 1.6592 | 42.3817 | 70.3205 | 25.5432 | 2.6358 | 42.2797 | 111.4426 | 16.0403 | 14.4662 | 5.5185 | 181.7631 | 296.4224 | 226.5124 | 290.7606 | 198.9449
2012-10-30 16:04 | 27.7448 | 311.3126 | 241.1818 | 305.7551 | 212.7251 | 80.0937 | 1.6594 | 42.3879 | 70.3379 | 25.5443 | 2.6362 | 42.2858 | 111.4754 | 16.0402 | 14.4695 | 5.5217 | 181.8133 | 296.8431 | 226.7123 | 291.2856 | 198.2556
2012-10-30 16:04 | 27.7463 | 311.6777 | 241.3402 | 306.2587 | 211.9193 | 80.1417 | 1.6597 | 42.3985 | 70.3698 | 25.5455 | 2.6369 | 42.2964 | 111.5318 | 16.0402 | 14.5199 | 5.525 | 181.9016 | 297.1578 | 226.8203 | 291.7388 | 197.3994
2012-10-30 16:04 | 27.7456 | 312.0519 | 241.3317 | 306.6911 | 211.8517 | 80.1891 | 1.6603 | 42.4139 | 70.4183 | 25.5464 | 2.6379 | 42.3116 | 111.6118 | 16.0402 | 14.5273 | 5.5282 | 182.0301 | 297.5246 | 226.8044 | 292.1638 | 197.3244
2012-10-30 16:04 | 27.7471 | 312.4016 | 241.4195 | 307.0835 | 212.9952 | 80.2203 | 1.6587 | 42.374 | 70.2845 | 25.547 | 2.6354 | 42.2718 | 111.4019 | 16.0401 | 14.5166 | 5.5304 | 181.6865 | 297.885 | 226.9029 | 292.5669 | 198.4786
2012-10-30 16:04 | 27.7403 | 312.7506 | 241.6206 | 307.4752 | 212.9214 | 80.2544 | 1.6623 | 42.4695 | 70.5976 | 25.5485 | 2.6413 | 42.3673 | 111.9038 | 16.0405 | 14.5302 | 5.5327 | 182.5014 | 298.2204 | 227.0904 | 292.945 | 198.3912
2012-10-30 16:04 | 27.7448 | 313.0994 | 241.786 | 307.8245 | 213.7765 | 80.3185 | 1.6781 | 42.8806 | 71.956 | 25.5538 | 2.6677 | 42.7968 | 114.1695 | 16.0425 | 14.6162 | 5.5372 | 186.1255 | 298.4832 | 227.1698 | 293.2083 | 199.1603
2012-10-30 16:04 | 27.7411 | 313.5355 | 242.1706 | 308.2689 | 214.4856 | 80.4485 | 1.6916 | 43.24 | 73.1454 | 25.5614 | 2.6882 | 43.1356 | 115.959 | 16.046 | 14.5636 | 5.5461 | 189.1044 | 298.9719 | 227.607 | 293.7053 | 199.922
2012-10-30 16:04 | 27.7426 | 313.9577 | 242.5604 | 308.6381 | 214.6998 | 80.5 | 1.6919 | 43.2509 | 73.1775 | 25.5631 | 2.6887 | 43.1463 | 116.0089 | 16.0471 | 14.5773 | 5.5497 | 189.1864 | 299.3804 | 227.9831 | 294.0608 | 200.1225
2012-10-30 16:04 | 27.7381 | 314.37 | 242.8971 | 309.0434 | 215.0891 | 80.5526 | 1.6921 | 43.2567 | 73.1938 | 25.5642 | 2.689 | 43.1521 | 116.0351 | 16.0478 | 14.5117 | 5.5533 | 189.2289 | 299.8583 | 228.3854 | 294.5317 | 200.5774
2012-10-30 16:04 | 27.7426 | 314.7607 | 243.2192 | 309.45 | 215.511 | 80.6115 | 1.6924 | 43.2651 | 73.2198 | 25.565 | 2.6895 | 43.162 | 116.0848 | 16.0483 | 14.508 | 5.5574 | 189.3047 | 300.2527 | 228.7112 | 294.942 | 201.003
2012-10-30 16:04 | 27.7456 | 315.1766 | 243.5712 | 309.8053 | 215.833 | 80.6693 | 1.7066 | 43.6374 | 74.47 | 25.5704 | 2.7124 | 43.5358 | 118.0863 | 16.0507 | 14.5029 | 5.5613 | 192.5563 | 300.6737 | 229.0683 | 295.3024 | 201.3301
2012-10-30 16:05 | 27.7501 | 315.6589 | 244.0108 | 310.2156 | 216.2667 | 80.7404 | 1.7068 | 43.649 | 74.5016 | 25.5731 | 2.7126 | 43.5436 | 118.1177 | 16.0521 | 14.5237 | 5.5662 | 192.6193 | 301.1352 | 229.4871 | 295.6919 | 201.743
2012-10-30 16:05 | 27.7531 | 316.1779 | 244.3666 | 310.6473 | 216.691 | 80.793 | 1.7067 | 43.6475 | 74.4938 | 25.574 | 2.7124 | 43.542 | 118.1036 | 16.0529 | 14.5511 | 5.5699 | 192.5973 | 301.6268 | 229.8155 | 296.0962 | 202.1399
2012-10-30 16:05 | 27.759 | 316.5316 | 244.7292 | 311.1087 | 216.9765 | 80.8473 | 1.7067 | 43.6498 | 74.4982 | 25.5751 | 2.7124 | 43.544 | 118.1099 | 16.0535 | 14.5407 | 5.5736 | 192.6081 | 301.9909 | 230.1885 | 296.568 | 202.4358
2012-10-30 16:05 | 27.7643 | 316.9878 | 245.0834 | 311.5235 | 217.2811 | 80.9132 | 1.7071 | 43.662 | 74.5361 | 25.5764 | 2.7131 | 43.5568 | 118.1758 | 16.054 | 14.5379 | 5.5782 | 192.7119 | 302.4499 | 230.5455 | 296.9856 | 202.7432
2012-10-30 16:05 | 27.7643 | 317.4042 | 245.4596 | 311.975 | 217.7103 | 80.9687 | 1.7171 | 43.9254 | 75.4263 | 25.5805 | 2.7287 | 43.8128 | 119.5533 | 16.0561 | 14.5542 | 5.582 | 194.9796 | 302.85 | 230.9054 | 297.4208 | 203.1561
2012-10-30 16:05 | 27.7673 | 317.8504 | 245.8348 | 312.3682 | 217.8945 | 81.0455 | 1.7189 | 43.974 | 75.5857 | 25.5831 | 2.7319 | 43.8676 | 119.8411 | 16.0576 | 14.5531 | 5.5873 | 195.4268 | 303.2973 | 231.2817 | 297.8151 | 203.3414
2012-10-30 16:05 | 27.7755 | 318.3476 | 246.2701 | 312.7781 | 218.1153 | 81.0964 | 1.7187 | 43.9726 | 75.577 | 25.5843 | 2.7317 | 43.8662 | 119.8277 | 16.0584 | 14.5165 | 5.5908 | 195.4047 | 303.8311 | 231.7536 | 298.2616 | 203.5988
2012-10-30 16:05 | 27.7845 | 318.7882 | 246.6043 | 313.2347 | 218.4354 | 81.1426 | 1.7186 | 43.9713 | 75.5703 | 25.5852 | 2.7315 | 43.8649 | 119.8158 | 16.0591 | 14.5336 | 5.594 | 195.3861 | 304.2546 | 232.0707 | 298.7011 | 203.9018
2012-10-30 16:05 | 27.7875 | 319.1388 | 246.9285 | 313.6507 | 218.7061 | 81.1975 | 1.7185 | 43.9715 | 75.5671 | 25.5864 | 2.7314 | 43.8651 | 119.8136 | 16.0595 | 14.5611 | 5.5978 | 195.3807 | 304.5777 | 232.3674 | 299.0896 | 204.145
2012-10-30 16:05 | 27.7778 | 319.4766 | 247.1734 | 314.0349 | 219.1886 | 81.2311 | 1.7185 | 43.9708 | 75.5618 | 25.5874 | 2.7313 | 43.8644 | 119.8057 | 16.06 | 14.5514 | 5.6001 | 195.3675 | 304.9252 | 232.622 | 299.4835 | 204.6372
2012-10-30 16:05 | 27.7972 | 319.958 | 247.5727 | 314.4864 | 219.3931 | 81.2819 | 1.7186 | 43.9766 | 75.5786 | 25.5885 | 2.7316 | 43.8701 | 119.8341 | 16.0604 | 14.53 | 5.6036 | 195.4127 | 305.428 | 233.0427 | 299.9564 | 204.8631
2012-10-30 16:05 | 27.807 | 320.4756 | 247.9506 | 314.9818 | 219.6116 | 81.3362 | 1.7185 | 43.9761 | 75.5738 | 25.5895 | 2.7315 | 43.8697 | 119.8285 | 16.0609 | 14.5561 | 5.6073 | 195.4022 | 305.9195 | 233.3945 | 300.4257 | 205.0555
2012-10-30 16:05 | 27.8047 | 320.8399 | 248.242 | 315.4 | 219.9675 | 81.3894 | 1.7184 | 43.9756 | 75.5698 | 25.5903 | 2.7314 | 43.8692 | 119.8228 | 16.0612 | 14.5049 | 5.611 | 195.3926 | 306.335 | 233.7371 | 300.8951 | 205.4626
2012-10-30 16:05 | 27.8115 | 321.2017 | 248.4915 | 315.7585 | 220.0807 | 81.4218 | 1.7184 | 43.9757 | 75.5677 | 25.5911 | 2.7313 | 43.8693 | 119.8209 | 16.0616 | 14.5688 | 5.6132 | 195.3886 | 306.6329 | 233.9227 | 301.1897 | 205.5119
2012-10-30 16:05 | 27.8219 | 321.5594 | 248.7877 | 316.1358 | 220.3824 | 81.4518 | 1.7184 | 43.9766 | 75.5683 | 25.592 | 2.7313 | 43.8702 | 119.8231 | 16.062 | 14.5751 | 5.6153 | 195.3914 | 306.9843 | 234.2126 | 301.5607 | 205.8073
2012-10-30 16:05 | 27.8122 | 321.9543 | 249.1304 | 316.5886 | 220.7704 | 81.5067 | 1.7183 | 43.9764 | 75.5642 | 25.5932 | 2.7312 | 43.8701 | 119.8186 | 16.0625 | 14.545 | 5.6191 | 195.3829 | 307.4093 | 234.5854 | 302.0436 | 206.2254
2012-10-30 16:05 | 27.8242 | 322.3861 | 249.4594 | 317.0515 | 221.2312 | 81.5599 | 1.7183 | 43.9778 | 75.5672 | 25.5938 | 2.7312 | 43.8713 | 119.8227 | 16.0629 | 14.5733 | 5.6227 | 195.3899 | 307.8128 | 234.8861 | 302.4782 | 206.6579
2012-10-30 16:05 | 27.8234 | 322.7288 | 249.803 | 317.429 | 221.4902 | 81.6015 | 1.7183 | 43.9785 | 75.5671 | 25.5946 | 2.7312 | 43.872 | 119.8242 | 16.0632 | 14.5399 | 5.6256 | 195.3913 | 308.1889 | 235.2631 | 302.8891 | 206.9503
2012-10-30 16:05 | 27.8339 | 323.1323 | 250.1501 | 317.8826 | 222.103 | 81.6269 | 1.7182 | 43.9788 | 75.5652 | 25.5955 | 2.7312 | 43.8724 | 119.8223 | 16.0637 | 14.5382 | 5.6274 | 195.3876 | 308.5941 | 235.6119 | 303.3444 | 207.5648
2012-10-30 16:05 | 27.8377 | 323.5558 | 250.5335 | 318.299 | 222.1108 | 81.6789 | 1.7182 | 43.9793 | 75.5638 | 25.5966 | 2.7311 | 43.8728 | 119.8217 | 16.064 | 14.5583 | 5.6309 | 195.3854 | 308.9975 | 235.9752 | 303.7407 | 207.5525
2012-10-30 16:05 | 27.8452 | 323.9677 | 250.8461 | 318.7229 | 222.2561 | 81.7211 | 1.7183 | 43.9838 | 75.5764 | 25.5976 | 2.7313 | 43.8772 | 119.8425 | 16.0645 | 14.5291 | 5.6339 | 195.4189 | 309.4386 | 236.317 | 304.1938 | 207.727
2012-10-30 16:05 | 27.8489 | 324.2499 | 251.1041 | 319.0894 | 222.4763 | 81.7599 | 1.7184 | 43.9866 | 75.5845 | 25.5982 | 2.7314 | 43.88 | 119.8557 | 16.0648 | 14.5491 | 5.6365 | 195.4402 | 309.7008 | 236.555 | 304.5403 | 207.9272
2012-10-30 16:05 | 27.8482 | 324.5849 | 251.3577 | 319.4401 | 222.8692 | 81.7859 | 1.7183 | 43.9865 | 75.5811 | 25.5991 | 2.7314 | 43.8799 | 119.8525 | 16.0651 | 14.577 | 5.6383 | 195.4336 | 310.0079 | 236.7807 | 304.8631 | 208.2922
2012-10-30 16:05 | 27.8519 | 324.9319 | 251.6743 | 319.818 | 223.2114 | 81.8096 | 1.7183 | 43.9886 | 75.5862 | 25.5999 | 2.7314 | 43.8821 | 119.8608 | 16.0656 | 14.5562 | 5.64 | 195.447 | 310.3757 | 237.1181 | 305.2618 | 208.6552
2012-10-30 16:05 | 27.8579 | 325.2736 | 252.0441 | 320.1715 | 223.4654 | 81.861 | 1.7183 | 43.9889 | 75.5841 | 25.6009 | 2.7314 | 43.8824 | 119.8596 | 16.066 | 14.607 | 5.6435 | 195.4437 | 310.6666 | 237.4371 | 305.5645 | 208.8584
2012-10-30 16:06 | 27.8669 | 325.6373 | 252.3151 | 320.6272 | 223.7106 | 81.9072 | 1.7183 | 43.9909 | 75.5885 | 25.6017 | 2.7314 | 43.8841 | 119.866 | 16.0664 | 14.616 | 5.6467 | 195.4544 | 311.0213 | 237.6991 | 306.0112 | 209.0946
2012-10-30 16:06 | 27.8639 | 325.9851 | 252.6249 | 320.9792 | 224.0225 | 81.9523 | 1.7183 | 43.9935 | 75.5954 | 25.6024 | 2.7315 | 43.8868 | 119.8771 | 16.0669 | 14.6049 | 5.6498 | 195.4726 | 311.3802 | 238.02 | 306.3743 | 209.4176
2012-10-30 16:06 | 27.8631 | 326.3085 | 252.8665 | 321.3487 | 224.183 | 81.9835 | 1.7183 | 43.9951 | 75.5981 | 25.6034 | 2.7315 | 43.8885 | 119.8835 | 16.0673 | 14.6082 | 5.6519 | 195.4816 | 311.7003 | 238.2583 | 306.7405 | 209.5748
2012-10-30 16:06 | 27.8691 | 326.6119 | 253.1501 | 321.6982 | 224.6849 | 82.0274 | 1.7184 | 43.9975 | 75.6038 | 25.6042 | 2.7316 | 43.8908 | 119.8934 | 16.0676 | 14.6101 | 5.655 | 195.4972 | 312.0018 | 238.54 | 307.0881 | 210.0748
2012-10-30 16:06 | 27.8699 | 326.9749 | 253.5145 | 322.0731 | 224.855 | 82.0656 | 1.7182 | 43.9956 | 75.5952 | 25.605 | 2.7315 | 43.889 | 119.8815 | 16.0679 | 14.5864 | 5.6576 | 195.4766 | 312.3885 | 238.9281 | 307.4867 | 210.2686
2012-10-30 16:06 | 27.8848 | 327.3177 | 253.7365 | 322.4124 | 225.1572 | 82.1089 | 1.7182 | 43.9945 | 75.5896 | 25.6056 | 2.7313 | 43.8878 | 119.8729 | 16.0682 | 14.6176 | 5.6606 | 195.4625 | 312.7001 | 239.1189 | 307.7948 | 210.5396
2012-10-30 16:06 | 27.8833 | 327.6022 | 254.0712 | 322.762 | 225.5807 | 82.15 | [8 empty] | 14.5917 | 5.6634 | 0 | 313.0105 | 239.4795 | 308.1703 | 210.989
2012-10-30 16:06 |  | 253.996 | 199.9559 | 262.6053 | 179.4441 | 76.8758 | 1.3503 | 33.1537 | 44.766 | 24.5536 | 2.1326 | 34.1776 | 72.8864 | 16.0264 | -13.3894 | 5.2998 | 117.6524 | 267.3854 | 213.3453 | 275.9947 | 192.8335
2012-10-30 16:07 | 27.8856 | 282.2335 | 228.0933 | 289.1416 | 206.6998 | 76.9983 | 1.3489 | 34.2412 | 46.1894 | 25.3837 | 2.1315 | 34.1603 | 72.8121 | 16.0265 | 14.5614 | 5.3083 | 119.0015 | 267.6721 | 213.5319 | 274.5802 | 192.1384
2012-10-30 16:07 | 27.8856 | 282.5389 | 227.9792 | 288.0925 | 206.3022 | 77.0081 | 1.3485 | 34.2306 | 46.1586 | 25.3849 | 2.1309 | 34.1498 | 72.7684 | 16.0264 | 14.5573 | 5.3089 | 118.927 | 267.9816 | 213.4219 | 273.5352 | 191.7449
2012-10-30 16:07 | 27.8781 | 282.7325 | 227.7474 | 286.9621 | 205.7475 | 76.9977 | 1.3515 | 34.3114 | 46.3732 | 25.3869 | 2.1361 | 34.2338 | 73.1278 | 16.0261 | 14.5458 | 5.3082 | 119.501 | 268.1867 | 213.2016 | 272.4163 | 191.2017
2012-10-30 16:07 | 27.8856 | 282.8675 | 227.5189 | 285.9659 | 204.9796 | 77.0683 | 1.4461 | 36.7397 | 53.1296 | 25.406 | 2.2874 | 36.662 | 83.8604 | 16.0279 | 14.5614 | 5.3131 | 136.99 | 268.3061 | 212.9575 | 271.4045 | 190.4182
2012-10-30 16:07 | 27.8961 | 283.253 | 227.6041 | 285.1655 | 204.8562 | 77.4682 | 1.5371 | 39.1109 | 60.1189 | 25.4439 | 2.4334 | 39.0144 | 94.9379 | 16.0328 | 14.6167 | 5.3407 | 155.0568 | 268.6363 | 212.9874 | 270.5488 | 190.2395
2012-10-30 16:07 | 27.8818 | 283.745 | 227.7944 | 284.5216 | 204.0705 | 77.6606 | 1.5523 | 39.513 | 61.3363 | 25.4544 | 2.4582 | 39.4116 | 96.8814 | 16.0328 | 14.6065 | 5.3539 | 158.2176 | 269.1385 | 213.1879 | 269.9151 | 189.464
2012-10-30 16:07 | 27.8833 | 284.3956 | 227.8864 | 283.9783 | 203.8359 | 77.7918 | 1.5697 | 39.969 | 62.7375 | 25.4636 | 2.487 | 39.873 | 99.1624 | 16.0329 | 14.5835 | 5.363 | 161.8999 | 269.8121 | 213.3029 | 269.3948 | 189.2524
2012-10-30 16:07 | 27.8744 | 285.055 | 227.8854 | 283.5559 | 204.6023 | 77.8496 | 1.5697 | 39.9736 | 62.7448 | 25.4665 | 2.4874 | 39.8778 | 99.1939 | 16.0316 | 14.6195 | 5.367 | 161.9388 | 270.4355 | 213.2659 | 268.9364 | 189.9828
2012-10-30 16:07 | 27.8774 | 285.6797 | 228.0143 | 283.2729 | 205.1012 | 77.9005 | 1.5696 | 39.976 | 62.7461 | 25.469 | 2.4878 | 39.8801 | 99.2127 | 16.0304 | 14.6143 | 5.3705 | 161.9588 | 271.0654 | 213.4 | 268.6586 | 190.4869
2012-10-30 16:07 | 27.8639 | 286.2492 | 228.1227 | 283.0739 | 204.95 | 77.9467 | 1.5696 | 39.9786 | 62.7512 | 25.4703 | 2.4881 | 39.8825 | 99.2313 | 16.0293 | 14.609 | 5.3736 | 161.9826 | 271.6402 | 213.5137 | 268.4649 | 190.341
2012-10-30 16:07 | 27.8714 | 286.6582 | 228.2482 | 282.9809 | 204.8709 | 77.9848 | 1.5696 | 39.982 | 62.7577 | 25.472 | 2.4885 | 39.886 | 99.2546 | 16.0284 | 14.7265 | 5.3763 | 162.0124 | 271.9317 | 213.5217 | 268.2544 | 190.1444
2012-10-30 16:07 | 27.8766 | 287.1188 | 228.3675 | 282.8895 | 204.9902 | 78.0108 | 1.5741 | 40.0992 | 63.1199 | 25.4745 | 2.4959 | 40.0043 | 99.8483 | 16.0278 | 14.6746 | 5.3781 | 162.9682 | 272.4442 | 213.6929 | 268.2149 | 190.3156
2012-10-30 16:07 | 27.8676 | 287.6462 | 228.4451 | 282.9307 | 204.9183 | 78.049 | 1.5752 | 40.1301 | 63.2117 | 25.4767 | 2.4979 | 40.0338 | 100 | 16.0271 | 14.6982 | 5.3807 | 163.2118 | 272.948 | 213.7469 | 268.2325 | 190.2201
2012-10-30 16:07 | 27.8691 | 288.2727 | 228.5567 | 282.9631 | 203.4675 | 78.0912 | 1.5757 | 40.145 | 63.2545 | 25.4783 | 2.4989 | 40.0483 | 100.0776 | 16.0263 | 14.6794 | 5.3836 | 163.3322 | 273.5933 | 213.8773 | 268.2837 | 188.7881
2012-10-30 16:07 | 27.8609 | 288.8006 | 228.4895 | 283.1289 | 204.27 | 78.1328 | 1.576 | 40.1576 | 63.29 | 25.4801 | 2.4998 | 40.0607 | 100.1444 | 16.0254 | 14.663 | 5.3865 | 163.4344 | 274.1376 | 213.8265 | 268.4659 | 189.607
2012-10-30 16:07 | 27.8579 | 289.2411 | 228.6556 | 283.3153 | 204.389 | 78.1698 | 1.5762 | 40.1638 | 63.3056 | 25.4817 | 2.5003 | 40.067 | 100.18 | 16.0248 | 14.6885 | 5.389 | 163.4855 | 274.5526 | 213.9671 | 268.6268 | 189.7005
2012-10-30 16:07 | 27.8564 | 289.6946 | 228.7407 | 283.5186 | 204.4426 | 78.2016 | 1.5763 | 40.1684 | 63.3176 | 25.4826 | 2.5007 | 40.0715 | 100.2068 | 16.0241 | 14.6707 | 5.3912 | 163.5244 | 275.0239 | 214.07 | 268.8479 | 189.7719
2012-10-30 16:07 | 27.8639 | 290.0761 | 228.8977 | 283.785 | 204.6154 | 78.2397 | 1.5764 | 40.1736 | 63.3308 | 25.4839 | 2.5011 | 40.0765 | 100.2353 | 16.0236 | 14.7149 | 5.3938 | 163.5661 | 275.3612 | 214.1828 | 269.0701 | 189.9005
2012-10-30 16:07 | 27.8684 | 290.5007 | 229.0752 | 284.1218 | 204.7144 | 78.2888 | 1.5766 | 40.1791 | 63.345 | 25.4852 | 2.5015 | 40.0822 | 100.2675 | 16.0229 | 14.7153 | 5.3972 | 163.6125 | 275.7854 | 214.3599 | 269.4065 | 189.9991
2012-10-30 16:07 | 27.8586 | 290.8493 | 229.2031 | 284.4056 | 204.7833 | 78.3177 | 1.5767 | 40.1833 | 63.3556 | 25.4862 | 2.5019 | 40.0864 | 100.2921 | 16.0223 | 14.624 | 5.3992 | 163.6477 | 276.2253 | 214.5791 | 269.7816 | 190.1593
2012-10-30 16:07 | 27.8549 | 291.2233 | 229.3567 | 284.7226 | 204.8898 | 78.3536 | 1.5767 | 40.1855 | 63.36 | 25.4873 | 2.5021 | 40.0885 | 100.3048 | 16.022 | 14.657 | 5.4017 | 163.6648 | 276.5663 | 214.6997 | 270.0656 | 190.2328
2012-10-30 16:07 | 27.8564 | 291.6062 | 229.5037 | 285.0022 | 204.9818 | 78.3778 | 1.582 | 40.3234 | 63.7912 | 25.489 | 2.5109 | 40.2284 | 101.008 | 16.0218 | 14.5933 | 5.4034 | 164.7992 | 277.0129 | 214.9104 | 270.4089 | 190.3885
2012-10-30 16:07 | 27.8519 | 291.9677 | 229.5897 | 285.3414 | 205.001 | 78.4691 | 1.601 | 40.8189 | 65.3494 | 25.4966 | 2.5418 | 40.727 | 103.5181 | 16.0231 | 14.6296 | 5.4097 | 168.8675 | 277.3381 | 214.9601 | 270.7118 | 190.3714
2012-10-30 16:07 | 27.8579 | 292.4821 | 229.8395 | 285.6911 | 205.1172 | 78.5408 | 1.6031 | 40.8766 | 65.5273 | 25.4992 | 2.5448 | 40.7756 | 103.7653 | 16.0232 | 14.6152 | 5.4146 | 169.2927 | 277.8669 | 215.2243 | 271.0759 | 190.502
2012-10-30 16:08 | 27.8579 | 292.9442 | 230.1109 | 286.1506 | 205.314 | 78.5899 | 1.609 | 41.0346 | 66.0262 | 25.5026 | 2.5547 | 40.9354 | 104.5769 | 16.0236 | 14.6396 | 5.418 | 170.6031 | 278.3046 | 215.4713 | 271.511 | 190.6744
2012-10-30 16:08 | 27.8601 | 293.447 | 230.4159 | 286.6237 | 205.6901 | 78.6599 | 1.6089 | 41.0346 | 66.0214 | 25.5044 | 2.5547 | 40.9353 | 104.5782 | 16.0234 | 14.6581 | 5.4228 | 170.5996 | 278.7889 | 215.7578 | 271.9656 | 191.032
2012-10-30 16:08 | 27.8489 | 293.9401 | 230.5738 | 287.0718 | 205.6789 | 78.6992 | 1.6089 | 41.0351 | 66.0201 | 25.5055 | 2.5548 | 40.9358 | 104.5821 | 16.0232 | 14.6021 | 5.4255 | 170.6022 | 279.338 | 215.9717 | 272.4697 | 191.0768
2012-10-30 16:08 | 27.8541 | 294.3302 | 230.7599 | 287.4629 | 205.7077 | 78.7269 | 1.6088 | 41.0354 | 66.0185 | 25.5065 | 2.5548 | 40.936 | 104.5844 | 16.023 | 14.6888 | 5.4274 | 170.603 | 279.6414 | 216.0711 | 272.7741 | 191.0189
2012-10-30 16:08 | 27.8541 | 294.7958 | 230.9919 | 287.9337 | 205.8022 | 78.7604 | 1.6086 | 41.0321 | 66.0055 | 25.5075 | 2.5547 | 40.9329 | 104.5695 | 16.0228 | 14.6603 | 5.4297 | 170.575 | 280.1355 | 216.3316 | 273.2734 | 191.1419
2012-10-30 16:08 | 27.8661 | 295.2272 | 231.2438 | 288.3777 | 206.0819 | 78.8194 | 1.6085 | 41.0299 | 65.9958 | 25.5085 | 2.5545 | 40.9305 | 104.559 | 16.0226 | 14.6234 | 5.4338 | 170.5548 | 280.6038 | 216.6204 | 273.7543 | 191.4585
2012-10-30 16:08 | 27.8729 | 295.6342 | 231.3842 | 288.7664 | 206.2383 | 78.8379 | 1.6084 | 41.0299 | 65.9939 | 25.5093 | 2.5546 | 40.9306 | 104.5618 | 16.0223 | 14.6669 | 5.4351 | 170.5557 | 280.9673 | 216.7173 | 274.0995 | 191.5714
2012-10-30 16:08 | 27.8721 | 296.0488 | 231.6075 | 289.2438 | 206.4579 | 78.8737 | 1.6084 | 41.0307 | 65.9948 | 25.5099 | 2.5547 | 40.9313 | 104.5663 | 16.0221 | 14.6661 | 5.4376 | 170.5611 | 281.3827 | 216.9414 | 274.5777 | 191.7918
2012-10-30 16:08 | 27.8654 | 296.3998 | 231.7974 | 289.7075 | 206.5733 | 78.9136 | 1.6084 | 41.0309 | 65.9925 | 25.511 | 2.5547 | 40.9316 | 104.5686 | 16.022 | 14.6268 | 5.4403 | 170.5611 | 281.773 | 217.1706 | 275.0807 | 191.9465
2012-10-30 16:08 | 27.8736 | 296.7733 | 232.0337 | 290.1012 | 206.6051 | 78.9361 | 1.6152 | 41.2089 | 66.5606 | 25.5132 | 2.5651 | 41.0987 | 105.4231 | 16.0222 | 14.6676 | 5.4419 | 171.9838 | 282.1057 | 217.3661 | 275.4336 | 191.9375
2012-10-30 16:08 | 27.8714 | 297.1787 | 232.2398 | 290.6039 | 206.6934 | 79.0245 | 1.6277 | 41.5382 | 67.6117 | 25.5196 | 2.586 | 41.4375 | 107.1571 | 16.0238 | 14.6409 | 5.4479 | 174.7688 | 282.5378 | 217.5989 | 275.963 | 192.0525
2012-10-30 16:08 | 27.8676 | 297.6401 | 232.5426 | 291.0549 | 206.9416 | 79.0864 | 1.6275 | 41.5359 | 67.6008 | 25.5209 | 2.5858 | 41.4352 | 107.1445 | 16.024 | 14.633 | 5.4522 | 174.7453 | 283.0071 | 217.9096 | 276.4219 | 192.3086
2012-10-30 16:08 | 27.8744 | 298.0543 | 232.7499 | 291.4817 | 207.1374 | 79.1251 | 1.6273 | 41.5322 | 67.5859 | 25.522 | 2.5856 | 41.4316 | 107.1235 | 16.0243 | 14.6969 | 5.4549 | 174.7095 | 283.3574 | 218.053 | 276.7848 | 192.4405
2012-10-30 16:08 | 27.8766 | 298.5523 | 233.0076 | 291.8961 | 207.3246 | 79.1621 | 1.6272 | 41.5314 | 67.5805 | 25.523 | 2.5855 | 41.4307 | 107.1183 | 16.0243 | 14.6543 | 5.4574 | 174.6988 | 283.898 | 218.3533 | 277.2418 | 192.6703
2012-10-30 16:08 | 27.8878 | 298.8901 | 233.2468 | 292.3464 | 207.5602 | 79.1898 | 1.6271 | 41.5304 | 67.5745 | 25.5241 | 2.5854 | 41.4297 | 107.1129 | 16.0244 | 14.6695 | 5.4593 | 174.6874 | 284.2206 | 218.5773 | 277.6769 | 192.8907
2012-10-30 16:08 | 27.8931 | 299.345 | 233.5469 | 292.9101 | 207.7663 | 79.2395 | 1.627 | 41.5301 | 67.5708 | 25.525 | 2.5854 | 41.4293 | 107.1113 | 16.0243 | 14.6137 | 5.4628 | 174.6821 | 284.7313 | 218.9332 | 278.2964 | 193.1526
2012-10-30 16:08 | 27.8901 | 299.7722 | 233.7875 | 293.4077 | 208.0074 | 79.3008 | 1.627 | 41.53 | 67.5676 | 25.5261 | 2.5854 | 41.4292 | 107.111 | 16.0243 | 14.6963 | 5.467 | 174.6787 | 285.0759 | 219.0912 | 278.7114 | 193.3111
2012-10-30 16:08 | 27.9021 | 300.2029 | 234.0511 | 293.7854 | 208.2792 | 79.3233 | 1.6269 | 41.5292 | 67.5634 | 25.5268 | 2.5854 | 41.4284 | 107.107 | 16.0243 | 14.6431 | 5.4685 | 174.6705 | 285.5598 | 219.408 | 279.1423 | 193.6361
2012-10-30 16:08 | 27.8983 | 300.5524 | 234.2948 | 294.2782 | 208.3148 | 79.3482 | 1.6268 | 41.5292 | 67.5609 | 25.5277 | 2.5853 | 41.4284 | 107.1063 | 16.0243 | 14.676 | 5.4703 | 174.6672 | 285.8764 | 219.6188 | 279.6022 | 193.6388
2012-10-30 16:08 | 27.9021 | 300.9479 | 234.5226 | 294.6976 | 208.5115 | 79.3973 | 1.6267 | 41.5282 | 67.555 | 25.5287 | 2.5853 | 41.4274 | 107.1027 | 16.0242 | 14.6309 | 5.4736 | 174.6577 | 286.317 | 219.8917 | 280.0667 | 193.8806
2012-10-30 16:08 | 27.9058 | 301.3624 | 234.7817 | 295.1746 | 208.6412 | 79.4325 | 1.6266 | 41.526 | 67.5458 | 25.5294 | 2.5852 | 41.4251 | 107.0915 | 16.024 | 14.6631 | 5.4761 | 174.6373 | 286.6993 | 220.1186 | 280.5115 | 193.9781
2012-10-30 16:08 | 27.9051 | 301.7801 | 235.0364 | 295.6664 | 208.928 | 79.4632 | 1.6264 | 41.5233 | 67.5345 | 25.5304 | 2.585 | 41.4225 | 107.0775 | 16.0242 | 14.6257 | 5.4782 | 174.612 | 287.1544 | 220.4107 | 281.0407 | 194.3023
2012-10-30 16:08 | 27.914 | 302.1307 | 235.2418 | 296.1676 | 209.1889 | 79.5244 | 1.6265 | 41.526 | 67.542 | 25.5309 | 2.5852 | 41.4252 | 107.0917 | 16.0241 | 14.5979 | 5.4824 | 174.6337 | 287.5328 | 220.6439 | 281.5697 | 194.591
2012-10-30 16:08 | 27.9163 | 302.5168 | 235.468 | 296.5891 | 209.5416 | 79.558 | 1.6266 | 41.5301 | 67.5522 | 25.5321 | 2.5855 | 41.4293 | 107.1137 | 16.024 | 14.5962 | 5.4847 | 174.6658 | 287.9206 | 220.8718 | 281.9929 | 194.9454
2012-10-30 16:08 | 27.9178 | 302.8636 | 235.7642 | 296.9944 | 209.6297 | 79.5845 | 1.6266 | 41.5312 | 67.5535 | 25.5329 | 2.5855 | 41.4302 | 107.1179 | 16.0241 | 14.6384 | 5.4866 | 174.6714 | 288.2252 | 221.1258 | 282.356 | 194.9913
2012-10-30 16:08 | 27.9193 | 303.2335 | 235.9975 | 297.4342 | 209.7139 | 79.6285 | 1.6266 | 41.5321 | 67.5545 | 25.5336 | 2.5856 | 41.4311 | 107.1229 | 16.024 | 14.6236 | 5.4896 | 174.6774 | 288.6099 | 221.3739 | 282.8106 | 195.0903
2012-10-30 16:08 | 27.9268 | 303.5825 | 236.2132 | 297.8107 | 209.7883 | 79.6678 | 1.6265 | 41.5307 | 67.5484 | 25.5343 | 2.5855 | 41.4298 | 107.1159 | 16.0241 | 14.6596 | 5.4923 | 174.6642 | 288.9229 | 221.5536 | 283.1511 | 195.1287
2012-10-30 16:08 | 27.929 | 303.9261 | 236.4039 | 298.1973 | 209.885 | 79.7076 | 1.6263 | 41.5286 | 67.5397 | 25.535 | 2.5853 | 41.4277 | 107.1044 | 16.0241 | 14.6537 | 5.495 | 174.6441 | 289.2724 | 221.7502 | 283.5436 | 195.2313
2012-10-30 16:09 | 27.9343 | 304.2382 | 236.6174 | 298.6177 | 210.103 | 79.729 | 1.6263 | 41.5296 | 67.5408 | 25.5358 | 2.5854 | 41.4286 | 107.1088 | 16.0241 | 14.659 | 5.4965 | 174.6496 | 289.5792 | 221.9584 | 283.9587 | 195.444
2012-10-30 16:09 | 27.9313 | 304.5996 | 236.8697 | 299.0412 | 210.3835 | 79.788 | 1.6263 | 41.5303 | 67.5408 | 25.5367 | 2.5854 | 41.4293 | 107.1134 | 16.0241 | 14.6519 | 5.5006 | 174.6543 | 289.9477 | 222.2178 | 284.3893 | 195.7316
2012-10-30 16:09 | 27.938 | 304.9284 | 237.112 | 299.4167 | 210.5989 | 79.8174 | 1.6262 | 41.5293 | 67.536 | 25.5372 | 2.5854 | 41.4282 | 107.1073 | 16.0241 | 14.6382 | 5.5026 | 174.6434 | 290.2902 | 222.4738 | 284.7785 | 195.9607
2012-10-30 16:09 | 27.938 | 305.2774 | 237.277 | 299.7971 | 210.8391 | 79.8319 | 1.6262 | 41.5289 | 67.5327 | 25.538 | 2.5854 | 41.4279 | 107.1061 | 16.024 | 14.6464 | 5.5036 | 174.6387 | 290.631 | 222.6306 | 285.1507 | 196.1927
2012-10-30 16:09 | 27.9448 | 305.5947 | 237.543 | 300.138 | 211.0428 | 79.8932 | 1.6261 | 41.5291 | 67.5318 | 25.5386 | 2.5854 | 41.4281 | 107.1066 | 16.0241 | 14.6654 | 5.5078 | 174.6385 | 290.9293 | 222.8776 | 285.4726 | 196.3774
2012-10-30 16:09 | 27.953 | 305.9518 | 237.7829 | 300.5418 | 211.3109 | 79.9088 | 1.6255 | 41.5132 | 67.4807 | 25.5383 | 2.5837 | 41.3998 | 106.9633 | 16.0237 | 14.6818 | 5.5089 | 174.444 | 291.27 | 223.1011 | 285.86 | 196.6291
2012-10-30 16:09 | 27.962 | 306.2483 | 237.9843 | 300.9579 | 211.3435 | 79.9683 | 1.6441 | 42.0006 | 69.0543 | 25.5458 | 2.6143 | 41.8982 | 109.5348 | 16.0265 | 14.6948 | 5.513 | 178.5891 | 291.5535 | 223.2895 | 286.2631 | 196.6487
2012-10-30 16:09 | 27.9642 | 306.6108 | 238.2535 | 301.3056 | 211.4284 | 80.0434 | 1.6428 | 41.9676 | 68.9429 | 25.5469 | 2.6122 | 41.865 | 109.3578 | 16.027 | 14.6848 | 5.5182 | 178.3007 | 291.926 | 223.5687 | 286.6208 | 196.7436
2012-10-30 16:09 | 27.9612 | 306.8914 | 238.4743 | 301.6789 | 211.5829 | 80.0729 | 1.6423 | 41.9565 | 68.9035 | 25.548 | 2.6114 | 41.8543 | 109.2987 | 16.0275 | 14.7022 | 5.5202 | 178.2022 | 292.1892 | 223.7721 | 286.9767 | 196.8807
2012-10-30 16:09 | 27.9635 | 307.1964 | 238.6415 | 301.9575 | 211.3686 | 80.1041 | 1.6421 | 41.9521 | 68.8879 | 25.5485 | 2.6111 | 41.8501 | 109.2749 | 16.0278 | 14.7371 | 5.5224 | 178.1627 | 292.4593 | 223.9044 | 287.2204 | 196.6315
2012-10-30 16:09 | 27.9642 | 307.5764 | 238.8071 | 302.2922 | 211.2984 | 80.1498 | 1.6421 | 41.9552 | 68.8963 | 25.5492 | 2.6113 | 41.8532 | 109.2893 | 16.028 | 14.7011 | 5.5255 | 178.1856 | 292.8753 | 224.106 | 287.5911 | 196.5973
2012-10-30 16:09 | 27.9747 | 309.1805 | 239.8342 | 303.8874 | 212.8959 | 80.3567 | 1.6632 | 42.5125 | 70.7073 | 25.5605 | 2.6453 | 42.4123 | 112.1917 | 16.0333 | 14.789 | 5.5398 | 182.8989 | 294.3915 | 225.0452 | 289.0984 | 198.1069
2012-10-30 16:09 | 27.9717 | 310.9042 | 241.2204 | 305.4376 | 214.4799 | 80.5751 | 1.6798 | 42.9534 | 72.1516 | 25.5711 | 2.6715 | 42.8485 | 114.4708 | 16.039 | 14.8553 | 5.5548 | 186.6223 | 296.0489 | 226.3651 | 290.5823 | 199.6246
2012-10-30 16:09 | 27.9889 | 311.3233 | 241.4534 | 305.8688 | 214.2451 | 80.6502 | 1.6788 | 42.931 | 72.074 | 25.5719 | 2.6701 | 42.8264 | 114.3488 | 16.0395 | 14.8888 | 5.56 | 186.4228 | 296.4345 | 226.5646 | 290.98 | 199.3563
2012-10-30 16:09 | 27.9822 | 311.7414 | 241.6781 | 306.2225 | 214.1596 | 80.6757 | 1.6917 | 43.2654 | 73.192 | 25.5751 | 2.6919 | 43.1817 | 116.2414 | 16.0413 | 14.8658 | 5.5618 | 189.4334 | 296.8756 | 226.8123 | 291.3567 | 199.2938
2012-10-30 16:09 | 27.9919 | 312.1414 | 242.0448 | 306.604 | 214.7128 | 80.8086 | 1.7132 | 43.8333 | 75.0941 | 25.586 | 2.7249 | 43.7263 | 119.1503 | 16.0469 | 14.8103 | 5.5709 | 194.2444 | 297.3311 | 227.2345 | 291.7937 | 199.9025
2012-10-30 16:09 | 27.9919 | 312.6808 | 242.4331 | 307.0179 | 214.8664 | 80.8982 | 1.7133 | 43.8387 | 75.1082 | 25.5875 | 2.7249 | 43.7315 | 119.1642 | 16.0488 | 14.8144 | 5.5771 | 194.2724 | 297.8664 | 227.6187 | 292.2035 | 200.052
2012-10-30 16:09 | 27.9897 | 313.13 | 242.6545 | 307.3989 | 215.0335 | 80.956 | 1.7133 | 43.8415 | 75.1139 | 25.5889 | 2.7249 | 43.7345 | 119.1701 | 16.0502 | 14.8285 | 5.5811 | 194.2839 | 298.3015 | 227.826 | 292.5704 | 200.205
2012-10-30 16:09 | 27.9957 | 313.5721 | 242.986 | 307.8455 | 215.0749 | 81.0016 | 1.7129 | 43.8334 | 75.0833 | 25.5898 | 2.7242 | 43.7264 | 119.1192 | 16.0511 | 14.8793 | 5.5843 | 194.2025 | 298.6928 | 228.1067 | 292.9662 | 200.1956
2012-10-30 16:09 | 27.9987 | 313.9729 | 243.24 | 308.1856 | 215.2079 | 81.0479 | 1.7129 | 43.8361 | 75.0887 | 25.5911 | 2.7242 | 43.7291 | 119.1264 | 16.0521 | 14.8864 | 5.5874 | 194.2151 | 299.0865 | 228.3536 | 293.2992 | 200.3215
2012-10-30 16:09 | 28.0002 | 314.4371 | 243.5042 | 308.6659 | 215.5205 | 81.0929 | 1.7127 | 43.8325 | 75.0734 | 25.5921 | 2.7239 | 43.7256 | 119.1022 | 16.0529 | 14.8064 | 5.5905 | 194.1756 | 299.6307 | 228.6978 | 293.8595 | 200.7141
2012-10-30 16:10 | 28.0099 | 314.9477 | 243.8668 | 309.1812 | 215.924 | 81.1421 | 1.7125 | 43.8279 | 75.0549 | 25.5931 | 2.7235 | 43.7214 | 119.0743 | 16.0535 | 14.8202 | 5.5939 | 194.1291 | 300.1275 | 229.0466 | 294.361 | 201.1038
2012-10-30 16:10 | 28.0129 | 315.8338 | 244.5206 | 310.1147 | 216.5925 | 81.2553 | 1.7111 | 43.7963 | 74.9414 | 25.5949 | 2.7213 | 43.6894 | 118.8921 | 16.0546 | 14.7865 | 5.6017 | 193.8335 | 301.0473 | 229.7341 | 295.3282 | 201.806
2012-10-30 16:10 | 28.0106 | 316.6685 | 245.0788 | 310.9507 | 217.2084 | 81.3368 | 1.7106 | 43.7863 | 74.9006 | 25.5971 | 2.7205 | 43.6792 | 118.8272 | 16.0559 | 14.7964 | 5.6074 | 193.7278 | 301.8721 | 230.2824 | 296.1543 | 202.412
2012-10-30 16:10 | 28.0129 | 317.0912 | 245.4887 | 311.4429 | 217.451 | 81.3958 | 1.7104 | 43.7828 | 74.8862 | 25.5979 | 2.7201 | 43.6758 | 118.8044 | 16.0564 | 14.7539 | 5.6114 | 193.6906 | 302.3373 | 230.7348 | 296.689 | 202.6971
2012-10-30 16:10 | 28.0301 | 317.4981 | 245.7763 | 311.8543 | 217.8068 | 81.4391 | 1.7102 | 43.7801 | 74.8747 | 25.5987 | 2.7199 | 43.6734 | 118.7875 | 16.057 | 14.7629 | 5.6144 | 193.6623 | 302.7352 | 231.0134 | 297.0914 | 203.0439
2012-10-30 16:10 | 28.0196 | 317.8735 | 246.1067 | 312.2915 | 217.9262 | 81.4807 | 1.7104 | 43.7851 | 74.8886 | 25.5998 | 2.7201 | 43.6782 | 118.8102 | 16.0574 | 14.7484 | 5.6173 | 193.6988 | 303.1251 | 231.3583 | 297.5431 | 203.1778
2012-10-30 16:10 | 28.0264 | 318.2471 | 246.3604 | 312.6771 | 218.2047 | 81.5235 | 1.7103 | 43.7861 | 74.8892 | 25.6008 | 2.7201 | 43.6792 | 118.8118 | 16.0579 | 14.7959 | 5.6202 | 193.701 | 303.4512 | 231.5645 | 297.8812 | 203.4088
2012-10-30 16:10 | 28.0376 | 318.7052 | 246.732 | 313.1894 | 218.527 | 81.5668 | 1.7105 | 43.791 | 74.9031 | 25.6018 | 2.7203 | 43.6841 | 118.8333 | 16.0586 | 14.8316 | 5.6232 | 193.7364 | 303.8736 | 231.9004 | 298.3578 | 203.6954
2012-10-30 16:10 | 28.0369 | 319.1399 | 247.0447 | 313.6439 | 218.7862 | 81.6154 | 1.7106 | 43.7969 | 74.92 | 25.6029 | 2.7206 | 43.6899 | 118.8612 | 16.0591 | 14.8227 | 5.6266 | 193.7812 | 304.3172 | 232.222 | 298.8212 | 203.9635
2012-10-30 16:10 | 28.0346 | 319.5233 | 247.3127 | 314.0891 | 219.158 | 81.6691 | 1.7106 | 43.799 | 74.9243 | 25.6039 | 2.7206 | 43.692 | 118.8693 | 16.0596 | 14.7797 | 5.6303 | 193.7936 | 304.7436 | 232.533 | 299.3094 | 204.3783
2012-10-30 16:10 | 28.0399 | 319.8151 | 247.6313 | 314.4462 | 219.4822 | 81.7136 | 1.7107 | 43.8009 | 74.9292 | 25.6044 | 2.7207 | 43.6939 | 118.8758 | 16.0601 | 14.842 | 5.6333 | 193.8051 | 304.9731 | 232.7893 | 299.6042 | 204.6402
2012-10-30 16:10 | 28.0436 | 320.2273 | 247.8896 | 314.8628 | 219.7064 | 81.7558 | 1.7108 | 43.8055 | 74.9424 | 25.6053 | 2.7208 | 43.6985 | 118.897 | 16.0606 | 14.8213 | 5.6362 | 193.8394 | 305.406 | 233.0683 | 300.0415 | 204.8851
2012-10-30 16:10 | 28.0601 | 320.5302 | 248.2664 | 315.2273 | 220.0851 | 81.7749 | 1.7108 | 43.8085 | 74.9494 | 25.6064 | 2.7209 | 43.7015 | 118.9093 | 16.0611 | 14.8378 | 5.6376 | 193.8587 | 305.6924 | 233.4286 | 300.3895 | 205.2473
2012-10-30 16:10 | 28.0601 | 320.9502 | 248.5405 | 315.6554 | 220.3883 | 81.8298 | 1.7108 | 43.8095 | 74.9495 | 25.6075 | 2.7209 | 43.7025 | 118.9119 | 16.0615 | 14.8052 | 5.6413 | 193.8613 | 306.145 | 233.7353 | 300.8502 | 205.5831
2012-10-30 16:10 | 28.0691 | 321.2799 | 248.8744 | 316.0773 | 220.5548 | 81.8789 | 1.7109 | 43.8132 | 74.96 | 25.6083 | 2.7211 | 43.7063 | 118.9291 | 16.062 | 14.7693 | 5.6447 | 193.8891 | 306.5106 | 234.1051 | 301.308 | 205.7855
2012-10-30 16:10 | 28.0691 | 321.6349 | 249.1563 | 316.4633 | 220.7477 | 81.9078 | 1.7112 | 43.8214 | 74.9857 | 25.6091 | 2.7215 | 43.7145 | 118.9687 | 16.0627 | 14.8182 | 5.6467 | 193.9544 | 306.8167 | 234.3381 | 301.6451 | 205.9295
2012-10-30 16:10 | 28.0765 | 322.0202 | 249.4612 | 316.807 | 221.011 | 81.946 | 1.7115 | 43.8329 | 75.0214 | 25.6103 | 2.7221 | 43.7256 | 119.0263 | 16.0631 | 14.8053 | 5.6494 | 194.0477 | 307.2149 | 234.6559 | 302.0017 | 206.2057
2012-10-30 16:10 | 28.0803 | 322.3675 | 249.7977 | 317.2465 | 221.381 | 81.9939 | 1.7114 | 43.8309 | 75.0126 | 25.6111 | 2.7219 | 43.7239 | 119.0122 | 16.0637 | 14.8254 | 5.6527 | 194.0248 | 307.5421 | 234.9723 | 302.4211 | 206.5556
2012-10-30 16:10 | 28.087 | 322.7596 | 250.094 | 317.6696 | 221.7735 | 82.0396 | 1.7112 | 43.8268 | 74.9958 | 25.6119 | 2.7216 | 43.7197 | 118.9856 | 16.0642 | 14.8035 | 5.6558 | 193.9814 | 307.9561 | 235.2905 | 302.8661 | 206.97
2012-10-30 16:10 | 28.087 | 323.1145 | 250.3915 | 318.1205 | 221.9467 | 82.0847 | 1.711 | 43.8246 | 74.9851 | 25.613 | 2.7213 | 43.7175 | 118.9706 | 16.0647 | 14.8239 | 5.6589 | 193.9557 | 308.2906 | 235.5676 | 303.2966 | 207.1228
2012-10-30 16:10 | 28.0923 | 323.4937 | 250.6942 | 318.4619 | 222.2473 | 82.1136 | 1.7112 | 43.8298 | 75.001 | 25.6137 | 2.7216 | 43.7228 | 118.9961 | 16.0651 | 14.8414 | 5.6609 | 193.9971 | 308.6523 | 235.8528 | 303.6205 | 207.4059
2012-10-30 16:10 | 28.0968 | 323.8187 | 251.0195 | 318.8408 | 222.6257 | 82.1528 | 1.7112 | 43.8324 | 75.0076 | 25.6145 | 2.7217 | 43.7253 | 119.0073 | 16.0654 | 14.8785 | 5.6636 | 194.0148 | 308.9402 | 236.141 | 303.9623 | 207.7472
2012-10-30 16:10 | 28.0915 | 324.1683 | 251.2959 | 319.1945 | 222.648 | 82.1869 | 1.7112 | 43.8333 | 75.0089 | 25.6151 | 2.7217 | 43.7262 | 119.008 | 16.066 | 14.8692 | 5.666 | 194.0169 | 309.2991 | 236.4267 | 304.3253 | 207.7788
2012-10-30 16:10 | 28.108 | 324.5472 | 251.5667 | 319.5891 | 223.0267 | 82.2378 | 1.7113 | 43.8354 | 75.0138 | 25.6159 | 2.7217 | 43.7283 | 119.0171 | 16.0663 | 14.8612 | 5.6695 | 194.0309 | 309.686 | 236.7055 | 304.7279 | 208.1655
2012-10-30 16:11 | 28.096 | 324.8518 | 251.8558 | 319.9246 | 223.0973 | 82.2696 | 1.7113 | 43.8374 | 75.0197 | 25.6161 | 2.7218 | 43.7301 | 119.0241 | 16.0666 | 14.8288 | 5.6717 | 194.0438 | 310.023 | 237.027 | 305.0958 | 208.2685
2012-10-30 16:11 | 28.1035 | 325.1225 | 252.0902 | 320.2376 | 223.4079 | 82.2933 | 1.7112 | 43.8368 | 75.0138 | 25.6175 | 2.7217 | 43.7296 | 119.0178 | 16.0671 | 14.873 | 5.6733 | 194.0316 | 310.2495 | 237.2172 | 305.3646 | 208.5349
2012-10-30 16:11 | 28.108 | 325.4283 | 252.3411 | 320.5017 | 223.6958 | 82.3216 | 1.7111 | 43.8356 | 75.0082 | 25.6179 | 2.7215 | 43.7284 | 119.0088 | 16.0675 | 14.8694 | 5.6753 | 194.017 | 310.5589 | 237.4717 | 305.6323 | 208.8264
2012-10-30 16:11 | 28.111 | 325.7479 | 252.5748 | 320.9171 | 223.8484 | 82.3684 | 1.7109 | 43.8303 | 74.9872 | 25.619 | 2.7211 | 43.7233 | 118.9766 | 16.0681 | 14.8479 | 5.6785 | 193.9638 | 310.9 | 237.7269 | 306.0692 | 209.0005
2012-10-30 16:11 | 28.1132 | 326.0171 | 252.8429 | 321.2209 | 224.1537 | 82.4037 |  |  |  |  | 2.07 | 33.1749 | 68.6737 | 16.0262 | 14.8786 | 5.6809 | 68.6737 | 311.1385 | 237.9643 | 306.3423 | 209.2751
2012-10-30 16:11 | 0 | 282.7028 | 216.6527 | 288.9636 | 193.1904 | 79.8088 | 1.3065 | 33.2445 | 43.4329 | 25.4462 | 2.0693 | 33.165 | 68.6291 | 16.0269 | -13.3935 | 5.502 | 112.062 | 296.0963 | 230.0462 | 302.3571 | 206.5839
2012-10-30 16:11 | 28.0945 | 310.4758 | 244.8805 | 315.8504 | 220.8282 | 79.8735 | 1.3054 | 33.2186 | 43.3632 | 25.4472 | 2.0678 | 33.1391 | 68.5254 | 16.0262 | 14.7499 | 5.5065 | 111.8886 | 295.7259 | 230.1306 | 301.1005 | 206.0783
2012-10-30 16:11 | 28.093 | 310.1067 | 244.636 | 314.7899 | 220.3582 | 79.8267 | 1.3049 | 33.2054 | 43.3303 | 25.4463 | 2.0671 | 33.1259 | 68.4733 | 16.0255 | 14.7525 | 5.5033 | 111.8036 | 295.3542 | 229.8835 | 300.0374 | 205.6057
2012-10-30 16:11 | 28.0878 | 309.6803 | 244.2623 | 313.8019 | 219.4159 | 79.7886 | 1.3047 | 33.1993 | 43.3153 | 25.4459 | 2.0668 | 33.1198 | 68.452 | 16.0246 | 14.6821 | 5.5006 | 111.7673 | 294.9982 | 229.5802 | 299.1198 | 204.7338
2012-10-30 16:11 | 28.0818 | 309.2378 | 243.7388 | 312.8664 | 218.5515 | 79.7562 | 1.3047 | 33.1971 | 43.3116 | 25.4447 | 2.0668 | 33.1177 | 68.4462 | 16.024 | 14.7454 | 5.4984 | 111.7578 | 294.4924 | 228.9934 | 298.121 | 203.8061
2012-10-30 16:11 | 28.0915 | 308.7266 | 243.3603 | 311.9922 | 217.6162 | 79.7048 | 1.3047 | 33.1976 | 43.3142 | 25.4438 | 2.0669 | 33.118 | 68.4507 | 16.0232 | 14.8284 | 5.4948 | 111.765 | 293.8982 | 228.5319 | 297.1638 | 202.7878
2012-10-30 16:11 | 28.0736 | 308.2912 | 242.8051 | 311.0444 | 216.2555 | 79.6568 | 1.3048 | 33.1988 | 43.3184 | 25.4432 | 2.067 | 33.1191 | 68.4583 | 16.0225 | 14.7616 | 5.4915 | 111.7767 | 293.5296 | 228.0435 | 296.2828 | 201.4939
2012-10-30 16:11 | 28.0825 | 307.8863 | 242.2923 | 310.2494 | 216.099 | 79.6048 | 1.3049 | 33.2001 | 43.3239 | 25.4421 | 2.0672 | 33.1207 | 68.4676 | 16.0219 | 14.8276 | 5.488 | 111.7915 | 293.0587 | 227.4647 | 295.4218 | 201.2714
2012-10-30 16:11 | 28.084 | 307.3935 | 241.8388 | 309.5004 | 216.6242 | 79.5759 | 1.305 | 33.2011 | 43.3277 | 25.4413 | 2.0673 | 33.1216 | 68.4738 | 16.0213 | 14.7965 | 5.486 | 111.8015 | 292.597 | 227.0423 | 294.7039 | 201.8277
2012-10-30 16:11 | 28.0803 | 306.8265 | 241.3682 | 308.6771 | 216.5457 | 79.5256 | 1.3052 | 33.204 | 43.3377 | 25.4398 | 2.0676 | 33.1243 | 68.4884 | 16.0206 | 14.8498 | 5.4825 | 111.8261 | 291.9767 | 226.5184 | 293.8273 | 201.6959
2012-10-30 16:11 | 28.0885 | 306.3288 | 240.9958 | 307.8463 | 216.7429 | 79.4551 | 1.3059 | 33.2197 | 43.3806 | 25.4387 | 2.0686 | 33.1393 | 68.5524 | 16.02 | 14.7887 | 5.4776 | 111.933 | 291.5401 | 226.2071 | 293.0576 | 201.9542
2012-10-30 16:11 | 28.084 | 305.8183 | 240.7637 | 307.1291 | 217.2425 | 79.4487 | 1.3103 | 33.3324 | 43.6749 | 25.4391 | 2.0757 | 33.2522 | 69.0231 | 16.0194 | 14.7598 | 5.4772 | 112.698 | 291.0585 | 226.0039 | 292.3693 | 202.4827
2012-10-30 16:11 | 28.0855 | 305.3443 | 240.4827 | 306.4638 | 216.5233 | 79.4008 | 1.3109 | 33.3472 | 43.7155 | 25.4381 | 2.0767 | 33.267 | 69.087 | 16.0188 | 14.7532 | 5.4739 | 112.8025 | 290.5911 | 225.7295 | 291.7106 | 201.7701
2012-10-30 16:12 | 28.081 | 304.8912 | 240.1368 | 305.7616 | 215.9636 | 79.3655 | 1.3115 | 33.3611 | 43.754 | 25.4369 | 2.0777 | 33.2811 | 69.1484 | 16.0182 | 14.7527 | 5.4715 | 112.9025 | 290.1385 | 225.3841 | 291.0089 | 201.2109
2012-10-30 16:12 | 28.0863 | 304.4247 | 239.7927 | 305.1342 | 215.6735 | 79.317 | 1.312 | 33.3721 | 43.7839 | 25.4363 | 2.0785 | 33.2922 | 69.197 | 16.0176 | 14.7988 | 5.4681 | 112.9809 | 289.6259 | 224.9939 | 290.3354 | 200.8747
2012-10-30 16:12 | 28.0691 | 303.9779 | 239.4655 | 304.5571 | 215.5185 | 79.2638 | 1.3123 | 33.3795 | 43.8046 | 25.4355 | 2.079 | 33.2993 | 69.2292 | 16.017 | 14.749 | 5.4644 | 113.0338 | 289.2289 | 224.7165 | 289.8081 | 200.7695
2012-10-30 16:12 | 28.0833 | 303.5663 | 239.1774 | 303.9499 | 215.068 | 79.2274 | 1.3125 | 33.3836 | 43.8173 | 25.4343 | 2.0793 | 33.3034 | 69.2491 | 16.0163 | 14.7347 | 5.4619 | 113.0665 | 288.8316 | 224.4427 | 289.2152 | 200.3333
2012-10-30 16:12 | 28.0893 | 303.108 | 238.8379 | 303.3536 | 214.7629 | 79.1944 | 1.3127 | 33.3859 | 43.8256 | 25.4331 | 2.0796 | 33.3058 | 69.2623 | 16.0156 | 14.7692 | 5.4597 | 113.0879 | 288.3388 | 224.0687 | 288.5844 | 199.9937
2012-10-30 16:12 | 28.0893 | 302.6399 | 238.5317 | 302.8202 | 214.6133 | 79.154 | 1.3128 | 33.3874 | 43.8313 | 25.432 | 2.0797 | 33.3072 | 69.2697 | 16.0152 | 14.7244 | 5.4569 | 113.1011 | 287.9155 | 223.8073 | 288.0958 | 199.8889
2012-10-30 16:12 | 28.0848 | 302.1825 | 238.1935 | 302.2362 | 214.2347 | 79.1101 | 1.3129 | 33.3883 | 43.8353 | 25.431 | 2.0799 | 33.3082 | 69.2765 | 16.0146 | 14.7199 | 5.4539 | 113.1118 | 287.4626 | 223.4736 | 287.5163 | 199.5148
2012-10-30 16:12 | 28.0855 | 301.7801 | 237.9075 | 301.7379 | 213.9164 | 79.0777 | 1.313 | 33.3892 | 43.8385 | 25.4306 | 2.08 | 33.309 | 69.2816 | 16.0142 | 14.6839 | 5.4516 | 113.1201 | 287.0962 | 223.2236 | 287.054 | 199.2325
2012-10-30 16:12 | 28.0945 | 301.3092 | 237.6338 | 301.2939 | 213.6694 | 79.028 | 1.313 | 33.39 | 43.8425 | 25.4294 | 2.0801 | 33.3098 | 69.2869 | 16.0137 | 14.7092 | 5.4482 | 113.1294 | 286.6 | 222.9246 | 286.5847 | 198.9602
2012-10-30 16:12 | 28.1005 | 300.912 | 237.2942 | 300.7546 | 213.3249 | 78.9951 | 1.3131 | 33.3894 | 43.8424 | 25.4287 | 2.0801 | 33.3093 | 69.2866 | 16.0134 | 14.7437 | 5.4459 | 113.129 | 286.1683 | 222.5505 | 286.0109 | 198.5812
2012-10-30 16:12 | 28.0968 | 300.4859 | 236.9566 | 300.2593 | 213.1952 | 78.9696 | 1.3131 | 33.3894 | 43.8443 | 25.4276 | 2.0801 | 33.3092 | 69.2882 | 16.0129 | 14.7319 | 5.4442 | 113.1325 | 285.754 | 222.2247 | 285.5274 | 198.4633
2012-10-30 16:12 | 28.102 | 300.0917 | 236.7222 | 299.8036 | 213.2044 | 78.9205 | 1.3132 | 33.3898 | 43.8468 | 25.4267 | 2.0802 | 33.3096 | 69.2921 | 16.0124 | 14.7371 | 5.4408 | 113.1389 | 285.3546 | 221.9851 | 285.0665 | 198.4673
2012-10-30 16:12 | 28.0968 | 299.6793 | 236.446 | 299.3412 | 212.7502 | 78.8789 | 1.3132 | 33.3899 | 43.8481 | 25.4261 | 2.0803 | 33.3098 | 69.2949 | 16.0119 | 14.6993 | 5.4379 | 113.143 | 284.98 | 221.7467 | 284.6419 | 198.0509
2012-10-30 16:12 | 28.102 | 299.2811 | 236.188 | 298.893 | 212.1687 | 78.8517 | 1.3133 | 33.3893 | 43.8486 | 25.4249 | 2.0803 | 33.3091 | 69.2945 | 16.0113 | 14.733 | 5.436 | 113.1431 | 284.5481 | 221.455 | 284.16 | 197.4357
2012-10-30 16:12 | 28.1095 | 298.8889 | 235.8223 | 298.447 | 211.7981 | 78.8003 | 1.3133 | 33.3892 | 43.8502 | 25.4238 | 2.0804 | 33.309 | 69.2961 | 16.0109 | 14.716 | 5.4325 | 113.1463 | 284.1729 | 221.1063 | 283.731 | 197.0821
2012-10-30 16:12 | 28.1065 | 298.4171 | 235.4343 | 297.9558 | 211.4644 | 78.7639 | 1.3134 | 33.3896 | 43.8531 | 25.4227 | 2.0805 | 33.3095 | 69.2997 | 16.0105 | 14.709 | 5.43 | 113.1529 | 283.7081 | 220.7253 | 283.2468 | 196.7554
2012-10-30 16:12 | 28.1088 | 298.0196 | 235.1223 | 297.5237 | 211.1044 | 78.7142 | 1.3134 | 33.3893 | 43.8542 | 25.4216 | 2.0805 | 33.3092 | 69.3008 | 16.01 | 14.7276 | 5.4266 | 113.155 | 283.292 | 220.3947 | 282.7961 | 196.3768
2012-10-30 16:12 | 28.114 | 297.6212 | 234.8013 | 297.0599 | 211.0505 | 78.694 | 1.3134 | 33.3874 | 43.8501 | 25.4211 | 2.0805 | 33.3072 | 69.294 | 16.0096 | 14.7205 | 5.4252 | 113.1441 | 282.9007 | 220.0808 | 282.3394 | 196.33
2012-10-30 16:12 | 28.1125 | 297.266 | 234.5012 | 296.6969 | 210.9585 | 78.6529 | 1.3134 | 33.3867 | 43.8503 | 25.4199 | 2.0804 | 33.3065 | 69.2923 | 16.0093 | 14.7394 | 5.4223 | 113.1427 | 282.5266 | 219.7618 | 281.9575 | 196.2191
2012-10-30 16:12 | 28.1147 | 296.949 | 234.2244 | 296.2837 | 210.7323 | 78.6125 | 1.3135 | 33.3871 | 43.8529 | 25.4191 | 2.0805 | 33.3068 | 69.2956 | 16.0089 | 14.7579 | 5.4195 | 113.1485 | 282.1911 | 219.4665 | 281.5258 | 195.9744
2012-10-30 16:12 | 28.1147 | 296.6106 | 234.0239 | 295.8567 | 210.4173 | 78.5882 | 1.3135 | 33.3879 | 43.8563 | 25.4183 | 2.0806 | 33.3077 | 69.3013 | 16.0084 | 14.7824 | 5.4179 | 113.1577 | 281.8282 | 219.2415 | 281.0743 | 195.6349
2012-10-30 16:12 | 28.1185 | 296.2721 | 233.7526 | 295.4989 | 210.0037 | 78.5529 | 1.3135 | 33.3865 | 43.8544 | 25.4172 | 2.0806 | 33.3064 | 69.2971 | 16.0081 | 14.7413 | 5.4154 | 113.1515 | 281.5308 | 219.0113 | 280.7576 | 195.2624
2012-10-30 16:12 | 28.129 | 295.9133 | 233.488 | 295.1169 | 209.8488 | 78.5264 | 1.3135 | 33.3856 | 43.8537 | 25.4163 | 2.0806 | 33.3055 | 69.2953 | 16.0076 | 14.7681 | 5.4136 | 113.149 | 281.1452 | 218.7199 | 280.3488 | 195.0807
2012-10-30 16:12 | 28.1357 | 295.5507 | 233.1567 | 294.7426 | 209.5484 | 78.4847 | 1.3136 | 33.3867 | 43.858 | 25.4155 | 2.0807 | 33.3065 | 69.3017 | 16.0072 | 14.7382 | 5.4107 | 113.1597 | 280.8125 | 218.4185 | 280.0044 | 194.8102
2012-10-30 16:12 | 28.129 | 295.1746 | 232.9338 | 294.378 | 209.396 | 78.4483 | 1.3137 | 33.3873 | 43.8608 | 25.4148 | 2.0808 | 33.3071 | 69.3059 | 16.0068 | 14.7355 | 5.4082 | 113.1667 | 280.4391 | 218.1983 | 279.6425 | 194.6605
2012-10-30 16:13 | 28.1387 | 294.888 | 232.6604 | 294.022 | 209.0868 | 78.4096 | 1.3137 | 33.3867 | 43.8609 | 25.4138 | 2.0808 | 33.3065 | 69.3041 | 16.0066 | 14.733 | 5.4056 | 113.165 | 280.155 | 217.9274 | 279.289 | 194.3538
2012-10-30 16:13 | 28.1357 | 294.5771 | 232.3901 | 293.6687 | 208.7176 | 78.3859 | 1.3138 | 33.3865 | 43.8619 | 25.4128 | 2.0808 | 33.3064 | 69.3052 | 16.0062 | 14.73 | 5.4039 | 113.1672 | 279.8471 | 217.6601 | 278.9387 | 193.9876
2012-10-30 16:13 | 28.1357 | 294.2038 | 232.0795 | 293.3068 | 208.4971 | 78.3541 | 1.3138 | 33.3868 | 43.8641 | 25.412 | 2.0809 | 33.3066 | 69.3085 | 16.0057 | 14.7218 | 5.4017 | 113.1726 | 279.482 | 217.3577 | 278.585 | 193.7753
2012-10-30 16:13 | 28.138 | 293.875 | 231.783 | 292.9163 | 208.3616 | 78.3189 | 1.3139 | 33.3868 | 43.8658 | 25.411 | 2.081 | 33.3067 | 69.3101 | 16.0054 | 14.7119 | 5.3993 | 113.1759 | 279.1631 | 217.0711 | 278.2044 | 193.6497
2012-10-30 16:13 | 28.1387 | 293.4946 | 231.5321 | 292.5318 | 208.2402 | 78.2842 | 1.3139 | 33.3865 | 43.8665 | 25.4103 | 2.081 | 33.3064 | 69.3098 | 16.0052 | 14.7289 | 5.3969 | 113.1762 | 278.7657 | 216.8032 | 277.8029 | 193.5113
2012-10-30 16:13 | 28.1469 | 293.2101 | 231.269 | 292.1779 | 207.9177 | 78.2553 | 1.3139 | 33.3853 | 43.8646 | 25.4096 | 2.0809 | 33.3052 | 69.3063 | 16.0048 | 14.7127 | 5.3949 | 113.1709 | 278.4974 | 216.5563 | 277.4652 | 193.205
2012-10-30 16:13 | 28.1484 | 292.8959 | 230.9991 | 291.8635 | 207.6633 | 78.2189 | 1.3139 | 33.3842 | 43.8626 | 25.409 | 2.0809 | 33.3039 | 69.3024 | 16.0045 | 14.6408 | 5.3924 | 113.165 | 278.2551 | 216.3583 | 277.2227 | 193.0225
2012-10-30 16:13 | 28.1544 | 292.5591 | 230.781 | 291.4919 | 207.4213 | 78.1802 | 1.3139 | 33.3834 | 43.8625 | 25.4079 | 2.0809 | 33.3033 | 69.3013 | 16.0041 | 14.7324 | 5.3897 | 113.1638 | 277.8267 | 216.0486 | 276.7595 | 192.6889
2012-10-30 16:13 | 28.1537 | 292.2964 | 230.5208 | 291.1868 | 207.2041 | 78.164 | 1.3139 | 33.383 | 43.8625 | 25.4072 | 2.0809 | 33.3029 | 69.3004 | 16.004 | 14.6991 | 5.3886 | 113.1629 | 277.5973 | 215.8217 | 276.4877 | 192.505
2012-10-30 16:13 | 28.1612 | 291.9958 | 230.3198 | 290.8783 | 207.0777 | 78.1201 | 1.3139 | 33.3828 | 43.8625 | 25.407 | 2.0809 | 33.3027 | 69.3001 | 16.0038 | 14.6903 | 5.3856 | 113.1626 | 277.3055 | 215.6295 | 276.188 | 192.3874
2012-10-30 16:13 | 28.1567 | 291.7331 | 230.0636 | 290.5655 | 206.9079 | 78.0871 | 1.3139 | 33.3821 | 43.8621 | 25.4061 | 2.0809 | 33.3019 | 69.2986 | 16.0034 | 14.6939 | 5.3833 | 113.1607 | 277.0392 | 215.3697 | 275.8716 | 192.214
2012-10-30 16:13 | 28.1717 | 291.4322 | 229.8584 | 290.2913 | 206.7221 | 78.0606 | 1.3139 | 33.3811 | 43.8602 | 25.4057 | 2.0809 | 33.301 | 69.296 | 16.0032 | 14.713 | 5.3815 | 113.1562 | 276.7192 | 215.1454 | 275.5783 | 192.0091
2012-10-30 16:13 | 28.1679 | 291.1509 | 229.6737 | 290.0215 | 206.675 | 78.0409 | 1.3139 | 33.3809 | 43.8609 | 25.405 | 2.0809 | 33.3007 | 69.296 | 16.0029 | 14.7744 | 5.3801 | 113.1568 | 276.3765 | 214.8993 | 275.2471 | 191.9006
2012-10-30 16:13 | 28.1627 | 290.8413 | 229.4757 | 289.6539 | 206.4966 | 78.0155 | 1.314 | 33.3807 | 43.8622 | 25.4039 | 2.0809 | 33.3005 | 69.2966 | 16.0026 | 14.7203 | 5.3784 | 113.1588 | 276.121 | 214.7554 | 274.9336 | 191.7763
2012-10-30 16:13 | 28.1709 | 290.5989 | 229.2715 | 289.3228 | 206.182 | 77.9796 | 1.314 | 33.3812 | 43.8642 | 25.4035 | 2.081 | 33.301 | 69.2997 | 16.0024 | 14.7652 | 5.3759 | 113.164 | 275.8337 | 214.5063 | 274.5576 | 191.4168
2012-10-30 16:13 | 28.1732 | 290.3353 | 229.0732 | 289.0435 | 205.9363 | 77.9403 | 1.3141 | 33.3813 | 43.8658 | 25.4027 | 2.0811 | 33.3011 | 69.3018 | 16.002 | 14.7797 | 5.3732 | 113.1676 | 275.5556 | 214.2935 | 274.2638 | 191.1566
2012-10-30 16:13 | 28.1702 | 290.0393 | 228.8106 | 288.732 | 205.7877 | 77.9218 | 1.3141 | 33.3811 | 43.8667 | 25.4019 | 2.0811 | 33.301 | 69.3015 | 16.0019 | 14.74 | 5.3719 | 113.1682 | 275.2993 | 214.0706 | 273.992 | 191.0477
2012-10-30 16:13 | 28.1717 | 289.8056 | 228.6036 | 288.4056 | 205.5884 | 77.8941 | 1.3141 | 33.3803 | 43.8657 | 25.4014 | 2.0811 | 33.3003 | 69.3002 | 16.0015 | 14.7293 | 5.37 | 113.1659 | 275.0763 | 213.8743 | 273.6763 | 190.8591
2012-10-30 16:13 | 28.1724 | 289.5672 | 228.3879 | 288.2134 | 205.4002 | 77.8548 | 1.3124 | 33.3372 | 43.7529 | 25.4011 | 2.0786 | 33.261 | 69.1373 | 16.0015 | 14.7504 | 5.3673 | 112.8901 | 274.8168 | 213.6375 | 273.463 | 190.6498
2012-10-30 16:13 | 28.1776 | 289.314 | 228.2004 | 287.9138 | 205.2873 | 77.8236 | 1.3078 | 33.2164 | 43.4406 | 25.3986 | 2.0709 | 33.136 | 68.6206 | 16.0009 | 14.7678 | 5.3652 | 112.0612 | 274.5462 | 213.4326 | 273.146 | 190.5195
2012-10-30 16:13 | 28.1844 | 289.0586 | 228.0144 | 287.635 | 204.9792 | 77.7941 | 1.3075 | 33.2081 | 43.4204 | 25.3977 | 2.0704 | 33.128 | 68.588 | 16.0008 | 14.7379 | 5.3631 | 112.0085 | 274.3207 | 213.2765 | 272.8971 | 190.2413
2012-10-30 16:13 | 28.1851 | 288.8279 | 227.7791 | 287.3617 | 204.8461 | 77.76 | 1.3069 | 33.1917 | 43.3788 | 25.3969 | 2.0694 | 33.1119 | 68.5226 | 16.0006 | 14.7264 | 5.3608 | 111.9014 | 274.1015 | 213.0527 | 272.6353 | 190.1197
2012-10-30 16:13 | 28.1732 | 288.5305 | 227.5547 | 287.0681 | 204.2045 | 77.7306 | 1.3067 | 33.1843 | 43.3611 | 25.396 | 2.069 | 33.1046 | 68.4933 | 16.0004 | 14.7553 | 5.3587 | 111.8544 | 273.7752 | 212.7994 | 272.3128 | 189.4492
2012-10-30 16:13 | 28.1739 | 288.3113 | 227.3036 | 286.8526 | 203.1111 | 77.7161 | 1.3067 | 33.1844 | 43.3623 | 25.3954 | 2.069 | 33.1048 | 68.4945 | 16.0002 | 14.7682 | 5.3577 | 111.8568 | 273.5431 | 212.5354 | 272.0844 | 188.3429
2012-10-30 16:13 | 28.1799 | 288.0318 | 226.9633 | 286.6 | 203.1722 | 77.682 | 1.3068 | 33.1855 | 43.366 | 25.395 | 2.0691 | 33.1059 | 68.4997 | 16 | 14.7457 | 5.3554 | 111.8657 | 273.2861 | 212.2176 | 271.8543 | 188.4265
2012-10-30 16:14 | 28.1881 | 287.7931 | 226.7355 | 286.365 | 203.086 | 77.6491 | 1.3069 | 33.1869 | 43.3704 | 25.3945 | 2.0692 | 33.1072 | 68.5059 | 15.9999 | 14.7987 | 5.3531 | 111.8763 | 272.9944 | 211.9368 | 271.5663 | 188.2873
2012-10-30 16:14 | 28.1911 | 287.58 | 226.4748 | 286.1054 | 203.392 | 77.6219 | 1.3069 | 33.1861 | 43.3696 | 25.3937 | 2.0692 | 33.1064 | 68.5033 | 15.9998 | 14.8017 | 5.3513 | 111.8729 | 272.7783 | 211.6731 | 271.3037 | 188.5903
2012-10-30 16:14 | 28.1844 | 287.3417 | 226.2674 | 285.8207 | 203.3302 | 77.6109 | 1.3068 | 33.1848 | 43.367 | 25.3933 | 2.0691 | 33.1052 | 68.4989 | 15.9996 | 14.7542 | 5.3505 | 111.8659 | 272.5875 | 211.5132 | 271.0665 | 188.576
2012-10-30 16:14 | 28.1934 | 287.0767 | 226.1072 | 285.5555 | 203.3549 | 77.5768 | 1.3069 | 33.1857 | 43.3708 | 25.3925 | 2.0692 | 33.106 | 68.5031 | 15.9994 | 14.7632 | 5.3481 | 111.8739 | 272.3135 | 211.344 | 270.7923 | 188.5917
2012-10-30 16:14 | 28.1859 | 286.8183 | 225.8675 | 285.2815 | 203.312 | 77.5479 | 1.3081 | 33.2167 | 43.4518 | 25.3925 | 2.0713 | 33.1393 | 68.6418 | 15.9992 | 14.7557 | 5.3462 | 112.0936 | 272.0626 | 211.1118 | 270.5258 | 188.5563
2012-10-30 16:14 | 28.1836 | 286.5187 | 225.7551 | 285.0474 | 203.3609 | 77.5358 | 1.3121 | 33.317 | 43.714 | 25.3929 | 2.0774 | 33.2369 | 69.0468 | 15.9992 | 14.7494 | 5.3453 | 112.7608 | 271.7693 | 211.0057 | 270.298 | 188.6115
2012-10-30 16:14 | 28.1836 | 286.372 | 225.5858 | 284.8041 | 203.1248 | 77.5289 | 1.3126 | 33.3308 | 43.751 | 25.3925 | 2.0783 | 33.2507 | 69.1052 | 15.9989 | 14.8064 | 5.3448 | 112.8562 | 271.5656 | 210.7794 | 269.9977 | 188.3184
2012-10-30 16:14 | 28.1844 | 286.1605 | 225.4646 | 284.5615 | 202.7557 | 77.4994 | 1.3131 | 33.343 | 43.7832 | 25.3923 | 2.0791 | 33.2627 | 69.1561 | 15.9987 | 14.8195 | 5.3428 | 112.9393 | 271.341 | 210.6451 | 269.742 | 187.9362
2012-10-30 16:14 | 28.1919 | 285.9324 | 225.2832 | 284.3333 | 202.0392 | 77.4948 | 1.3136 | 33.3542 | 43.8143 | 25.3913 | 2.0798 | 33.274 | 69.2036 | 15.9985 | 14.8473 | 5.3425 | 113.0179 | 271.0851 | 210.4359 | 269.486 | 187.1919
2012-10-30 16:14 | 28.1904 | 285.6838 | 224.9905 | 284.0883 | 201.4987 | 77.4849 | 1.3139 | 33.3626 | 43.8368 | 25.3911 | 2.0804 | 33.2825 | 69.2404 | 15.9982 | 14.854 | 5.3418 | 113.0772 | 270.8298 | 210.1365 | 269.2343 | 186.6447
2012-10-30 16:14 | 28.1881 | 285.4614 | 224.7127 | 283.7691 | 201.6971 | 77.4578 | 1.3142 | 33.3683 | 43.8523 | 25.3907 | 2.0808 | 33.2881 | 69.2644 | 15.9981 | 14.9209 | 5.3399 | 113.1168 | 270.5405 | 209.7918 | 268.8482 | 186.7762
2012-10-30 16:14 | 28.1821 | 285.2275 | 224.4942 | 283.6008 | 201.6556 | 77.4214 | 1.3144 | 33.3717 | 43.8623 | 25.3901 | 2.081 | 33.2915 | 69.2794 | 15.9979 | 14.8905 | 5.3374 | 113.1417 | 270.337 | 209.6037 | 268.7103 | 186.7651
2012-10-30 16:14 | 28.1994 | 285.0439 | 224.3304 | 283.3669 | 201.4212 | 77.4231 | 1.3145 | 33.3742 | 43.8697 | 25.3897 | 2.0812 | 33.2939 | 69.2901 | 15.9977 | 14.9648 | 5.3375 | 113.1598 | 270.0791 | 209.3656 | 268.4021 | 186.4564
2012-10-30 16:14 | 28.2039 | 284.8012 | 224.1184 | 283.1549 | 201.524 | 77.3861 | 1.3145 | 33.375 | 43.8727 | 25.3892 | 2.0813 | 33.2947 | 69.2953 | 15.9973 | 15.0182 | 5.335 | 113.168 | 269.783 | 209.1002 | 268.1367 | 186.5058
2012-10-30 16:14 | 28.2061 | 284.6025 | 223.9789 | 282.8826 | 201.4161 | 77.3792 | 1.3146 | 33.3751 | 43.8738 | 25.3887 | 2.0813 | 33.2949 | 69.2967 | 15.9972 | 15.0001 | 5.3345 | 113.1705 | 269.6024 | 208.9788 | 267.8825 | 186.416
2012-10-30 16:14 | 28.2009 | 284.3732 | 223.773 | 282.5836 | 201.2928 | 77.3422 | 1.3146 | 33.3768 | 43.8785 | 25.3885 | 2.0814 | 33.2965 | 69.3035 | 15.9971 | 15.0071 | 5.332 | 113.182 | 269.3661 | 208.7659 | 267.5765 | 186.2857
2012-10-30 16:14 | 28.2016 | 284.1459 | 223.632 | 282.4257 | 201.1126 | 77.3364 | 1.3147 | 33.3779 | 43.8828 | 25.3878 | 2.0815 | 33.2976 | 69.3094 | 15.9968 | 15.0241 | 5.3316 | 113.1921 | 269.1218 | 208.6079 | 267.4016 | 186.0885
2012-10-30 16:14 | 28.2121 | 283.9593 | 223.4693 | 282.1733 | 200.946 | 77.3133 | 1.3148 | 33.3787 | 43.8854 | 25.3874 | 2.0816 | 33.2985 | 69.3135 | 15.9967 | 14.9979 | 5.33 | 113.1989 | 268.9614 | 208.4714 | 267.1754 | 185.9481
2012-10-30 16:14 | 28.2121 | 283.7931 | 223.3236 | 282.0263 | 200.7966 | 77.3006 | 1.3148 | 33.3796 | 43.8891 | 25.3867 | 2.0817 | 33.2994 | 69.3184 | 15.9965 | 15.0142 | 5.3291 | 113.2075 | 268.7789 | 208.3094 | 267.0121 | 185.7824
2012-10-30 16:14 | 28.2218 | 283.598 | 223.1916 | 281.7924 | 200.7551 | 77.2861 | 1.3149 | 33.38 | 43.8904 | 25.3865 | 2.0817 | 33.2997 | 69.321 | 15.9961 | 14.9791 | 5.3281 | 113.2114 | 268.6189 | 208.2125 | 266.8133 | 185.776
2012-10-30 16:14 | 28.2263 | 283.3706 | 223.0505 | 281.5957 | 200.7754 | 77.2653 | 1.3149 | 33.3803 | 43.8928 | 25.3856 | 2.0818 | 33.2999 | 69.323 | 15.9959 | 14.9917 | 5.3267 | 113.2158 | 268.3789 | 208.0588 | 266.604 | 185.7837
2012-10-30 16:14 | 28.2241 | 283.1867 | 222.9105 | 281.3885 | 200.667 | 77.2549 | 1.315 | 33.3806 | 43.8941 | 25.3853 | 2.0818 | 33.3004 | 69.3257 | 15.9957 | 15.0018 | 5.326 | 113.2199 | 268.1849 | 207.9087 | 266.3867 | 185.6652
2012-10-30 16:14 | 28.2331 | 283.0333 | 222.7542 | 281.2002 | 200.4242 | 77.2272 | 1.315 | 33.3809 | 43.8954 | 25.385 | 2.0819 | 33.3006 | 69.3269 | 15.9957 | 15.023 | 5.324 | 113.2223 | 268.0103 | 207.7312 | 266.1772 | 185.4012
2012-10-30 16:14 | 28.2346 | 282.8106 | 222.6376 | 281.0121 | 200.17 | 77.2023 | 1.315 | 33.3812 | 43.8969 | 25.3845 | 2.0819 | 33.3009 | 69.329 | 15.9954 | 14.9593 | 5.3223 | 113.2259 | 267.8513 | 207.6783 | 266.0528 | 185.2107
2012-10-30 16:14 | 28.2383 | 282.6751 | 222.4838 | 280.7682 | 200.0361 | 77.1966 | 1.315 | 33.3809 | 43.897 | 25.3841 | 2.0819 | 33.3006 | 69.3287 | 15.9953 | 14.9834 | 5.3219 | 113.2257 | 267.6917 | 207.5004 | 265.7848 | 185.0527
2012-10-30 16:14 | 28.2473 | 282.4714 | 222.4023 | 280.634 | 199.8523 | 77.1694 | 1.315 | 33.3808 | 43.897 | 25.3839 | 2.0819 | 33.3006 | 69.3287 | 15.9952 | 14.9312 | 5.3201 | 113.2258 | 267.5402 | 207.4711 | 265.7028 | 184.9211
2012-10-30 16:14 | 28.2503 | 282.312 | 222.2084 | 280.4551 | 199.7374 | 77.1584 | 1.3151 | 33.3807 | 43.8981 | 25.3832 | 2.0819 | 33.3004 | 69.3286 | 15.9951 | 14.9342 | 5.3193 | 113.2267 | 267.3778 | 207.2742 | 265.5209 | 184.8032
2012-10-30 16:15 | 28.2473 | 282.0925 | 222.0637 | 280.2083 | 199.7304 | 77.1434 | 1.3151 | 33.3801 | 43.8971 | 25.3827 | 2.0819 | 33.2996 | 69.3264 | 15.9948 | 14.9394 | 5.3183 | 113.2235 | 267.1531 | 207.1243 | 265.2689 | 184.791
2012-10-30 16:15 | 28.2443 | 281.9077 | 221.8875 | 280.0466 | 199.5662 | 77.1416 | 1.3151 | 33.3802 | 43.8981 | 25.3824 | 2.0819 | 33.2998 | 69.328 | 15.9947 | 14.9894 | 5.3181 | 113.2261 | 266.9183 | 206.8981 | 265.0572 | 184.5768
2012-10-30 16:15 | 28.2585 | 281.7556 | 221.7993 | 279.8712 | 199.3955 | 77.0954 | 1.3152 | 33.3807 | 43.9007 | 25.3817 | 2.082 | 33.3003 | 69.3308 | 15.9945 | 15.0199 | 5.315 | 113.2316 | 266.7357 | 206.7794 | 264.8513 | 184.3756
2012-10-30 16:15 | 28.2698 | 281.5657 | 221.6571 | 279.6967 | 199.3164 | 77.0937 | 1.315 | 33.3777 | 43.8931 | 25.3815 | 2.0818 | 33.2971 | 69.3183 | 15.9943 | 15.0026 | 5.3148 | 113.2114 | 266.5631 | 206.6545 | 264.6941 | 184.3138
2012-10-30 16:15 | 28.2735 | 281.3992 | 221.523 | 279.5223 | 199.1667 | 77.0827 | 1.315 | 33.376 | 43.8895 | 25.381 | 2.0817 | 33.2956 | 69.3124 | 15.9942 | 15.0267 | 5.3141 | 113.2019 | 266.3725 | 206.4963 | 264.4956 | 184.14
2012-10-30 16:15 | 28.2758 | 281.239 | 221.4347 | 279.2923 | 198.8151 | 77.0573 | 1.3151 | 33.3771 | 43.8933 | 25.3805 | 2.0818 | 33.2968 | 69.318 | 15.994 | 14.9679 | 5.3123 | 113.2113 | 266.2711 | 206.4668 | 264.3244 | 183.8472
2012-10-30 16:15 | 28.2915 | 281.1038 | 221.2063 | 279.157 | 198.5044 | 77.0342 | 1.3147 | 33.3673 | 43.8683 | 25.38 | 2.0815 | 33.2908 | 69.2941 | 15.9938 | 14.9958 | 5.3107 | 113.1624 | 266.108 | 206.2105 | 264.1612 | 183.5086
2012-10-30 16:15 | 28.2952 | 280.9373 | 221.0604 | 279.0019 | 198.225 | 77.0342 | 1.3147 | 33.3656 | 43.8644 | 25.3797 | 2.0817 | 33.2931 | 69.3048 | 15.9936 | 14.9669 | 5.3107 | 113.1692 | 265.9704 | 206.0935 | 264.035 | 183.2581
2012-10-30 16:15 | 28.2945 | 280.778 | 220.8471 | 278.796 | 198.2676 | 77.0087 | 1.3147 | 33.366 | 43.8657 | 25.3796 | 2.0817 | 33.2932 | 69.3059 | 15.9934 | 14.9784 | 5.309 | 113.1716 | 265.7996 | 205.8687 | 263.8176 | 183.2892
2012-10-30 16:15 | 28.3042 | 280.5864 | 220.7466 | 278.6004 | 198.407 | 76.9954 | 1.3147 | 33.3662 | 43.8679 | 25.3786 | 2.0817 | 33.2933 | 69.3067 | 15.9933 | 14.9963 | 5.3081 | 113.1746 | 265.5901 | 205.7503 | 263.6041 | 183.4107
2012-10-30 16:15 | 28.3125 | 280.4361 | 220.6053 | 278.4345 | 198.4782 | 76.992 | 1.3148 | 33.3672 | 43.8705 | 25.3785 | 2.0818 | 33.2941 | 69.3108 | 15.9932 | 14.9598 | 5.3078 | 113.1813 | 265.4763 | 205.6455 | 263.4747 | 183.5184
2012-10-30 16:15 | 28.3102 | 280.248 | 220.4376 | 278.2462 | 198.2675 | 76.955 | 1.3148 | 33.3677 | 43.8734 | 25.3776 | 2.0818 | 33.2945 | 69.3135 | 15.9929 | 14.9371 | 5.3053 | 113.1869 | 265.3109 | 205.5005 | 263.3091 | 183.3304
2012-10-30 16:15 | 28.3327 | 280.1234 | 220.3735 | 278.1564 | 198.2507 | 76.944 | 1.3148 | 33.3673 | 43.8729 | 25.3773 | 2.0818 | 33.2941 | 69.312 | 15.9928 | 14.9515 | 5.3045 | 113.1849 | 265.1719 | 205.422 | 263.2049 | 183.2992
2012-10-30 16:15 | 28.3252 | 280.0153 | 220.2557 | 277.963 | 198.082 | 76.9324 | 1.3149 | 33.3671 | 43.8728 | 25.377 | 2.0818 | 33.2938 | 69.312 | 15.9926 | 14.9929 | 5.3037 | 113.1848 | 265.0224 | 205.2628 | 262.9701 | 183.0891
2012-10-30 16:15 | 28.3259 | 279.8612 | 220.1423 | 277.8126 | 197.9805 | 76.918 | 1.3149 | 33.3675 | 43.8745 | 25.3767 | 2.0819 | 33.2941 | 69.3139 | 15.9924 | 14.9487 | 5.3027 | 113.1884 | 264.9125 | 205.1936 | 262.8639 | 183.0318
2012-10-30 16:15 | 28.3327 | 279.6938 | 220.0428 | 277.6722 | 197.8969 | 76.9053 | 1.3149 | 33.3672 | 43.8745 | 25.3762 | 2.0819 | 33.2938 | 69.3135 | 15.9922 | 14.9718 | 5.3019 | 113.188 | 264.722 | 205.071 | 262.7004 | 182.9251
2012-10-30 16:15 | 28.3439 | 279.5385 | 219.9201 | 277.5129 | 197.7036 | 76.8891 | 1.3149 | 33.367 | 43.8748 | 25.3757 | 2.0819 | 33.2934 | 69.3127 | 15.992 | 14.9423 | 5.3007 | 113.1875 | 264.5962 | 204.9778 | 262.5706 | 182.7613
2012-10-30 16:15 | 28.3521 | 279.388 | 219.7668 | 277.3235 | 197.2912 | 76.8775 | 1.3149 | 33.3668 | 43.875 | 25.3754 | 2.0819 | 33.2932 | 69.3118 | 15.992 | 14.9464 | 5.2999 | 113.1868 | 264.4416 | 204.8204 | 262.3771 | 182.3448
2012-10-30 16:15 | 28.3551 | 279.209 | 219.5336 | 277.1521 | 197.2391 | 76.8377 | 1.315 | 33.3668 | 43.8759 | 25.3748 | 2.0819 | 33.2932 | 69.3131 | 15.9917 | 14.9331 | 5.2972 | 113.1891 | 264.2759 | 204.6005 | 262.219 | 182.306
2012-10-30 16:15 | 28.3656 | 279.0762 | 219.4693 | 276.9998 | 197.1671 | 76.851 | 1.315 | 33.3668 | 43.8758 | 25.3749 | 2.0819 | 33.2931 | 69.3142 | 15.9914 | 14.8621 | 5.2981 | 113.19 | 264.2141 | 204.6072 | 262.1377 | 182.305
2012-10-30 16:15 | 28.3739 | 278.9567 | 219.324 | 276.8841 | 197.3837 | 76.8134 | 1.315 | 33.3664 | 43.876 | 25.3742 | 2.0819 | 33.2928 | 69.313 | 15.9914 | 14.96 | 5.2955 | 113.189 | 263.9967 | 204.364 | 261.9241 | 182.4237
2012-10-30 16:15 | 28.3806 | 278.8008 | 219.2756 | 276.697 | 196.9895 | 76.8059 | 1.315 | 33.3666 | 43.877 | 25.3739 | 2.0819 | 33.293 | 69.3139 | 15.9913 | 14.9749 | 5.295 | 113.1909 | 263.8259 | 204.3007 | 261.7221 | 182.0146
2012-10-30 16:15 | 28.3761 | 278.6336 | 219.1057 | 276.5298 | 196.8906 | 76.7851 | 1.315 | 33.3663 | 43.8767 | 25.3736 | 2.0819 | 33.2927 | 69.3135 | 15.9912 | 14.9622 | 5.2936 | 113.1901 | 263.6714 | 204.1435 | 261.5676 | 181.9284
2012-10-30 16:15 | 28.3746 | 278.4695 | 218.994 | 276.3848 | 196.9009 | 76.7689 | 1.315 | 33.3661 | 43.8772 | 25.373 | 2.0819 | 33.2922 | 69.3122 | 15.991 | 14.973 | 5.2924 | 113.1893 | 263.4965 | 204.021 | 261.4118 | 181.9279
2012-10-30 16:15 | 28.3971 | 278.399 | 218.9062 | 276.2174 | 196.9549 | 76.7521 | 1.315 | 33.3654 | 43.8762 | 25.3725 | 2.0819 | 33.2917 | 69.3103 | 15.991 | 15.0199 | 5.2913 | 113.1865 | 263.3791 | 203.8863 | 261.1975 | 181.935
2012-10-30 16:15 | 28.3926 | 278.1969 | 218.8269 | 276.0927 | 196.86 | 76.7446 | 1.315 | 33.3652 | 43.8757 | 25.3725 | 2.0819 | 33.2914 | 69.3097 | 15.9908 | 15.0317 | 5.2908 | 113.1854 | 263.1652 | 203.7952 | 261.061 | 181.8283
2012-10-30 16:15 | 28.4053 | 278.0896 | 218.7175 | 275.9464 | 196.7076 | 76.7209 | 1.315 | 33.3636 | 43.8717 | 25.3725 | 2.0818 | 33.2898 | 69.3037 | 15.9907 | 15.0811 | 5.2891 | 113.1753 | 263.0085 | 203.6364 | 260.8653 | 181.6265
2012-10-30 16:16 | 28.4053 | 277.9424 | 218.5797 | 275.7642 | 196.5071 | 76.7221 | 1.3154 | 33.3744 | 43.9014 | 25.3716 | 2.0825 | 33.3005 | 69.3488 | 15.9905 | 15.073 | 5.2892 | 113.2503 | 262.8694 | 203.5067 | 260.6912 | 181.4341
2012-10-30 16:16 | 28.4091 | 277.8571 | 218.4772 | 275.6013 | 196.3733 | 76.7157 | 1.3155 | 33.3757 | 43.9054 | 25.3712 | 2.0826 | 33.3017 | 69.3541 | 15.9904 | 15.0727 | 5.2888 | 113.2595 | 262.7844 | 203.4045 | 260.5286 | 181.3006
2012-10-30 16:16 | 28.4225 | 277.6496 | 218.3567 | 275.4673 | 196.3081 | 76.6943 | 1.3155 | 33.3756 | 43.9055 | 25.3712 | 2.0826 | 33.3018 | 69.3555 | 15.9903 | 15.0616 | 5.2873 | 113.2609 | 262.588 | 203.2951 | 260.4057 | 181.2465
2012-10-30 16:16 | 28.4255 | 277.5131 | 218.281 | 275.3307 | 196.1225 | 76.6718 | 1.3154 | 33.3738 | 43.9016 | 25.3707 | 2.0825 | 33.3 | 69.3481 | 15.9902 | 14.9913 | 5.2858 | 113.2497 | 262.5218 | 203.2897 | 260.3394 | 181.1312
2012-10-30 16:16 | 28.4315 | 277.4106 | 218.1452 | 275.2165 | 195.932 | 76.6614 | 1.3155 | 33.374 | 43.902 | 25.3706 | 2.0826 | 33.3001 | 69.3497 | 15.99 | 14.9891 | 5.285 | 113.2517 | 262.4215 | 203.1561 | 260.2274 | 180.9429
2012-10-30 16:16 | 28.43 | 277.2696 | 218.0059 | 275.1064 | 196.068 | 76.6417 | 1.3099 | 33.2316 | 43.5312 | 25.369 | 2.0736 | 33.157 | 68.7558 | 15.9897 | 15.0651 | 5.2837 | 112.287 | 262.2045 | 202.9408 | 260.0413 | 181.0029
2012-10-30 16:16 | 28.4435 | 277.1398 | 217.9446 | 274.9299 | 196.2073 | 76.614 | 1.3091 | 33.21 | 43.4757 | 25.3684 | 2.0724 | 33.1364 | 68.6707 | 15.9897 | 15.1112 | 5.2818 | 112.1464 | 262.0286 | 202.8334 | 259.8187 | 181.0961
2012-10-30 16:16 | 28.4518 | 276.9814 | 217.8465 | 274.7908 | 195.0131 | 76.599 | 1.3085 | 33.1948 | 43.4369 | 25.3677 | 2.0714 | 33.1215 | 68.6088 | 15.9897 | 15.148 | 5.2807 | 112.0457 | 261.8334 | 202.6985 | 259.6428 | 179.8651
2012-10-30 16:16 | 28.4533 | 276.8434 | 217.5803 | 274.6449 | 194.8889 | 76.5897 | 1.3083 | 33.1888 | 43.4219 | 25.3673 | 2.0711 | 33.1154 | 68.5837 | 15.9897 | 15.1454 | 5.2801 | 112.0056 | 261.698 | 202.4349 | 259.4995 | 179.7435
2012-10-30 16:16 | 28.457 | 276.6765 | 217.3713 | 274.4934 | 195.0616 | 76.562 | 1.3082 | 33.1839 | 43.4102 | 25.3667 | 2.0708 | 33.1103 | 68.5633 | 15.9895 | 15.1491 | 5.2782 | 111.9735 | 261.5274 | 202.2222 | 259.3443 | 179.9125
2012-10-30 16:16 | 28.4675 | 276.5591 | 217.2519 | 274.3177 | 195.0328 | 76.5666 | 1.3081 | 33.183 | 43.4083 | 25.3663 | 2.0707 | 33.1094 | 68.5596 | 15.9895 | 15.1148 | 5.2785 | 111.9679 | 261.4443 | 202.1371 | 259.2029 | 179.918
2012-10-30 16:16 | 28.4622 | 276.4143 | 217.1166 | 274.1649 | 194.4224 | 76.551 | 1.3081 | 33.1812 | 43.404 | 25.3662 | 2.0706 | 33.1076 | 68.553 | 15.9892 | 15.1258 | 5.2774 | 111.9571 | 261.2885 | 201.9908 | 259.0391 | 179.2966
2012-10-30 16:16 | 28.454 | 276.2665 | 216.9273 | 274.0054 | 194.1078 | 76.5192 | 1.3082 | 33.1825 | 43.4079 | 25.3658 | 2.0707 | 33.1088 | 68.5579 | 15.9893 | 15.1135 | 5.2752 | 111.9658 | 261.153 | 201.8138 | 258.8919 | 178.9943
2012-10-30 16:16 | 28.463 | 276.1592 | 216.7197 | 273.867 | 194.2268 | 76.5117 | 1.3082 | 33.1822 | 43.4079 | 25.3654 | 2.0707 | 33.1085 | 68.5574 | 15.9892 | 15.1469 | 5.2747 | 111.9653 | 261.0123 | 201.5728 | 258.7201 | 179.0799
2012-10-30 16:16 | 28.4727 | 275.9634 | 216.5719 | 273.7641 | 194.6097 | 76.5025 | 1.3083 | 33.1844 | 43.4144 | 25.365 | 2.0708 | 33.1108 | 68.567 | 15.9891 | 15.1037 | 5.2741 | 111.9814 | 260.8597 | 201.4682 | 258.6604 | 179.506
2012-10-30 16:16 | 28.466 | 275.8443 | 216.4825 | 273.6255 | 193.837 | 76.4892 | 1.3084 | 33.1875 | 43.4229 | 25.3647 | 2.071 | 33.1138 | 68.5797 | 15.9891 | 15.1092 | 5.2732 | 112.0026 | 260.7351 | 201.3733 | 258.5163 | 178.7278
2012-10-30 16:16 | 28.4787 | 275.7135 | 216.251 | 273.4713 | 193.127 | 76.4863 | 1.3119 | 33.2768 | 43.6559 | 25.3652 | 2.0766 | 33.2029 | 68.9496 | 15.9889 | 15.0526 | 5.273 | 112.6056 | 260.6609 | 201.1984 | 258.4187 | 178.0744
2012-10-30 16:16 | 28.4705 | 275.6162 | 216.0302 | 273.3661 | 193.0599 | 76.4701 | 1.3138 | 33.3247 | 43.7811 | 25.3657 | 2.0796 | 33.2504 | 69.1479 | 15.9888 | 15.0811 | 5.2718 | 112.929 | 260.5351 | 200.9491 | 258.285 | 177.9788
2012-10-30 16:16 | 28.478 | 275.4918 | 215.8211 | 273.1756 | 192.176 | 76.4643 | 1.3143 | 33.3389 | 43.8185 | 25.3655 | 2.0805 | 33.2648 | 69.2076 | 15.9888 | 15.0438 | 5.2714 | 113.0261 | 260.448 | 200.7773 | 258.1318 | 177.1322
2012-10-30 16:16 | 28.4787 | 275.349 | 215.5816 | 273.0483 | 192.6283 | 76.4632 | 1.3149 | 33.3519 | 43.853 | 25.3654 | 2.0813 | 33.2777 | 69.2624 | 15.9886 | 15.1382 | 5.2714 | 113.1154 | 260.2108 | 200.4434 | 257.9101 | 177.4901
2012-10-30 16:16 | 28.4765 | 275.2344 | 215.4179 | 272.918 | 193.1641 | 76.436 | 1.3152 | 33.3614 | 43.8783 | 25.3652 | 2.0819 | 33.2871 | 69.3017 | 15.9885 | 15.1075 | 5.2695 | 113.1801 | 260.1269 | 200.3104 | 257.8105 | 178.0566
2012-10-30 16:16 | 28.475 | 275.163 | 215.3258 | 272.7884 | 193.2097 | 76.4418 | 1.3155 | 33.3675 | 43.8946 | 25.365 | 2.0823 | 33.2931 | 69.3276 | 15.9883 | 15.1752 | 5.2699 | 113.2223 | 259.9878 | 200.1506 | 257.6132 | 178.0345
2012-10-30 16:16 | 28.4825 | 274.9844 | 215.3018 | 272.6794 | 193.4607 | 76.4383 | 1.3157 | 33.3719 | 43.9072 | 25.3645 | 2.0826 | 33.2975 | 69.346 | 15.9883 | 15.1339 | 5.2697 | 113.2532 | 259.8505 | 200.1679 | 257.5455 | 178.3268
2012-10-30 16:16 | 28.4989 | 274.9077 | 215.2316 | 272.5444 | 193.536 | 76.4291 | 1.3158 | 33.3753 | 43.9161 | 25.3645 | 2.0829 | 33.301 | 69.3614 | 15.9881 | 15.1747 | 5.269 | 113.2774 | 259.733 | 200.0569 | 257.3697 | 178.3613
2012-10-30 16:16 | 28.493 | 274.7738 | 215.2257 | 272.3909 | 193.5615 | 76.4221 | 1.316 | 33.378 | 43.9238 | 25.3641 | 2.0831 | 33.3037 | 69.3733 | 15.9879 | 15.1769 | 5.2685 | 113.2971 | 259.5969 | 200.0488 | 257.214 | 178.3846
2012-10-30 16:16 | 28.4922 | 274.707 | 215.1737 | 272.3047 | 193.6236 | 76.4031 | 1.316 | 33.3795 | 43.9281 | 25.364 | 2.0832 | 33.3051 | 69.3796 | 15.9879 | 15.1028 | 5.2672 | 113.3077 | 259.6042 | 200.0709 | 257.2019 | 178.5208
2012-10-30 16:16 | 28.4982 | 274.5579 | 215.1009 | 272.1476 | 193.4253 | 76.3979 | 1.3161 | 33.3804 | 43.9309 | 25.3638 | 2.0832 | 33.3061 | 69.3843 | 15.9878 | 15.1862 | 5.2669 | 113.3152 | 259.3717 | 199.9147 | 256.9614 | 178.2391
2012-10-30 16:17 | 28.5034 | 274.4971 | 215.0589 | 272.0364 | 193.3599 | 76.3927 | 1.3161 | 33.3811 | 43.9332 | 25.3635 | 2.0833 | 33.3068 | 69.3879 | 15.9876 | 15.1548 | 5.2665 | 113.3211 | 259.3423 | 199.9041 | 256.8816 | 178.2051
2012-10-30 16:17 | 28.5042 | 274.3505 | 214.9927 | 271.9245 | 193.286 | 76.3742 | 1.3161 | 33.3818 | 43.9351 | 25.3634 | 2.0833 | 33.3074 | 69.3903 | 15.9876 | 15.1352 | 5.2652 | 113.3253 | 259.2153 | 199.8575 | 256.7893 | 178.1508
2012-10-30 16:17 | 28.5027 | 274.2365 | 214.9243 | 271.8493 | 193.2492 | 76.358 | 1.3162 | 33.3826 | 43.938 | 25.363 | 2.0834 | 33.3081 | 69.3942 | 15.9874 | 15.0888 | 5.2641 | 113.3322 | 259.1477 | 199.8355 | 256.7605 | 178.1604
2012-10-30 16:17 | 28.5064 | 274.1277 | 214.928 | 271.818 | 193.2175 | 76.3562 | 1.3162 | 33.383 | 43.9393 | 25.3628 | 2.0834 | 33.3087 | 69.3965 | 15.9874 | 15.1415 | 5.264 | 113.3358 | 258.9862 | 199.7865 | 256.6765 | 178.076
2012-10-30 16:17 | 28.5094 | 274.0492 | 214.8365 | 271.6695 | 193.1459 | 76.3406 | 1.3163 | 33.3836 | 43.9414 | 25.3625 | 2.0835 | 33.3092 | 69.3995 | 15.9872 | 15.1241 | 5.2629 | 113.3409 | 258.9251 | 199.7124 | 256.5454 | 178.0218
2012-10-30 16:17 | 28.5147 | 273.9342 | 214.7866 | 271.5117 | 193.1551 | 76.3447 | 1.3162 | 33.3827 | 43.9391 | 25.3625 | 2.0834 | 33.3084 | 69.3962 | 15.9872 | 15.1457 | 5.2632 | 113.3352 | 258.7885 | 199.6409 | 256.366 | 178.0094
2012-10-30 16:17 | 28.5154 | 273.834 | 214.6889 | 271.4037 | 193.0969 | 76.3343 | 1.3163 | 33.384 | 43.9433 | 25.362 | 2.0835 | 33.3095 | 69.4017 | 15.9869 | 15.1382 | 5.2625 | 113.3449 | 258.6958 | 199.5507 | 256.2655 | 177.9587
2012-10-30 16:17 | 28.5117 | 273.7372 | 214.6537 | 271.2912 | 192.9401 | 76.3083 | 1.3163 | 33.3845 | 43.9446 | 25.362 | 2.0836 | 33.31 | 69.4044 | 15.9869 | 15.0978 | 5.2607 | 113.349 | 258.6394 | 199.5559 | 256.1934 | 177.8423
2012-10-30 16:17 | 28.5237 | 273.6405 | 214.5869 | 271.2022 | 192.9953 | 76.2996 | 1.3163 | 33.3846 | 43.9457 | 25.3615 | 2.0836 | 33.3101 | 69.405 | 15.9867 | 15.118 | 5.2601 | 113.3507 | 258.5225 | 199.4689 | 256.0842 | 177.8773
2012-10-30 16:17 | 28.5282 | 273.5596 | 214.5914 | 271.0979 | 192.9801 | 76.3025 | 1.3164 | 33.3851 | 43.9474 | 25.3613 | 2.0837 | 33.3105 | 69.4077 | 15.9865 | 15.1429 | 5.2603 | 113.3551 | 258.4167 | 199.4485 | 255.955 | 177.8372
2012-10-30 16:17 | 28.5341 | 273.4685 | 214.4831 | 270.9562 | 192.8839 | 76.2823 | 1.3164 | 33.384 | 43.9452 | 25.361 | 2.0836 | 33.3096 | 69.4041 | 15.9865 | 15.0836 | 5.2589 | 113.3493 | 258.3849 | 199.3995 | 255.8726 | 177.8003
2012-10-30 16:17 | 28.5409 | 273.3899 | 214.4269 | 270.897 | 192.8358 | 76.2834 | 1.3164 | 33.3843 | 43.9462 | 25.3607 | 2.0836 | 33.3097 | 69.405 | 15.9864 | 15.0822 | 5.259 | 113.3513 | 258.3077 | 199.3447 | 255.8148 | 177.7536
2012-10-30 16:17 | 28.5416 | 273.2121 | 214.3134 | 270.7733 | 192.7029 | 76.2678 | 1.3164 | 33.3839 | 43.9454 | 25.3607 | 2.0836 | 33.3094 | 69.4047 | 15.9862 | 15.0829 | 5.2579 | 113.3501 | 258.1292 | 199.2305 | 255.6904 | 177.62
2012-10-30 16:17 | 28.5469 | 273.1514 | 214.2478 | 270.7087 | 192.6847 | 76.2672 | 1.3163 | 33.3825 | 43.9422 | 25.3605 | 2.0836 | 33.3081 | 69.3993 | 15.9862 | 15.0352 | 5.2579 | 113.3415 | 258.1162 | 199.2126 | 255.6735 | 177.6495
2012-10-30 16:17 | 28.5454 | 273.0257 | 214.1912 | 270.5673 | 192.4201 | 76.2638 | 1.3163 | 33.3825 | 43.9427 | 25.3601 | 2.0836 | 33.3081 | 69.4 | 15.986 | 15.0826 | 5.2576 | 113.3426 | 257.9431 | 199.1086 | 255.4847 | 177.3375
2012-10-30 16:17 | 28.5536 | 272.9757 | 214.0852 | 270.4667 | 192.3458 | 76.2615 | 1.3164 | 33.3829 | 43.9443 | 25.3598 | 2.0836 | 33.3082 | 69.4013 | 15.9858 | 15.099 | 5.2575 | 113.3457 | 257.8767 | 198.9862 | 255.3677 | 177.2468
2012-10-30 16:17 | 28.5544 | 272.8639 | 214.0545 | 270.3704 | 192.4566 | 76.2372 | 1.3164 | 33.3834 | 43.9466 | 25.3592 | 2.0836 | 33.3087 | 69.4037 | 15.9858 | 15.0794 | 5.2558 | 113.3503 | 257.7845 | 198.9751 | 255.291 | 177.3772
2012-10-30 16:17 | 28.5596 | 272.772 | 213.9652 | 270.2901 | 192.4657 | 76.2482 | 1.3164 | 33.3831 | 43.9458 | 25.3593 | 2.0837 | 33.3085 | 69.4039 | 15.9855 | 15.0887 | 5.2566 | 113.3497 | 257.6833 | 198.8765 | 255.2014 | 177.377
2012-10-30 16:17 | 28.5656 | 272.7004 | 213.9318 | 270.1756 | 192.3657 | 76.2199 | 1.3164 | 33.3832 | 43.9469 | 25.3587 | 2.0837 | 33.3086 | 69.4044 | 15.9854 | 15.058 | 5.2546 | 113.3513 | 257.6424 | 198.8738 | 255.1176 | 177.3077
2012-10-30 16:17 | 28.5686 | 272.6374 | 213.8915 | 270.0893 | 192.3608 | 76.2141 | 1.3164 | 33.3829 | 43.9462 | 25.3587 | 2.0837 | 33.3084 | 69.4036 | 15.9854 | 15.0936 | 5.2542 | 113.3497 | 257.5438 | 198.7979 | 254.9957 | 177.2672
2012-10-30 16:17 | 28.5701 | 272.4719 | 213.8142 | 269.9936 | 192.3113 | 76.2135 | 1.3164 | 33.3827 | 43.9459 | 25.3585 | 2.0837 | 33.308 | 69.403 | 15.9853 | 15.0625 | 5.2542 | 113.3489 | 257.4094 | 198.7517 | 254.9311 | 177.2488
2012-10-30 16:17 | 28.5746 | 272.3949 | 213.7242 | 269.9243 | 192.041 | 76.1962 | 1.3164 | 33.3829 | 43.9467 | 25.3584 | 2.0837 | 33.3083 | 69.4047 | 15.9851 | 15.014 | 5.253 | 113.3514 | 257.3809 | 198.7102 | 254.9103 | 177.027
2012-10-30 16:17 | 28.5806 | 272.3232 | 213.6633 | 269.8137 | 191.9763 | 76.199 | 1.3164 | 33.3824 | 43.9462 | 25.3579 | 2.0837 | 33.3078 | 69.4028 | 15.9851 | 15.0119 | 5.2532 | 113.3491 | 257.3113 | 198.6514 | 254.8018 | 176.9644
2012-10-30 16:17 | 28.5903 | 272.2514 | 213.5706 | 269.7651 | 191.9742 | 76.2025 | 1.3165 | 33.3823 | 43.9463 | 25.3577 | 2.0837 | 33.3075 | 69.4026 | 15.9848 | 15.046 | 5.2534 | 113.3489 | 257.2054 | 198.5246 | 254.7191 | 176.9282
2012-10-30 16:17 | 28.5986 | 272.182 | 213.4844 | 269.6762 | 191.9119 | 76.173 | 1.3164 | 33.3814 | 43.9431 | 25.3582 | 2.0837 | 33.3068 | 69.4001 | 15.9848 | 15.0747 | 5.2514 | 113.3432 | 257.1073 | 198.4097 | 254.6015 | 176.8372
2012-10-30 16:17 | 28.6008 | 272.091 | 213.4629 | 269.5307 | 191.9219 | 76.1782 | 1.3164 | 33.3809 | 43.9436 | 25.3572 | 2.0836 | 33.3063 | 69.3984 | 15.9847 | 15.0769 | 5.2517 | 113.342 | 257.0141 | 198.386 | 254.4538 | 176.845
2012-10-30 16:17 | 28.6008 | 271.9901 | 213.4275 | 269.4918 | 191.9651 | 76.1638 | 1.3164 | 33.3804 | 43.9423 | 25.3571 | 2.0836 | 33.3058 | 69.3964 | 15.9846 | 15.0484 | 5.2507 | 113.3387 | 256.9417 | 198.3791 | 254.4434 | 176.9167
2012-10-30 16:18 | 28.6031 | 271.9225 | 213.355 | 269.3697 | 191.9124 | 76.158 | 1.3164 | 33.38 | 43.9419 | 25.3569 | 2.0836 | 33.3053 | 69.3951 | 15.9845 | 15.0833 | 5.2503 | 113.337 | 256.8392 | 198.2717 | 254.2864 | 176.8291
2012-10-30 16:18 | 28.6083 | 271.8306 | 213.3326 | 269.3166 | 191.737 | 76.1569 | 1.3164 | 33.3786 | 43.9389 | 25.3563 | 2.0835 | 33.3039 | 69.3901 | 15.9843 | 15.1211 | 5.2503 | 113.329 | 256.7095 | 198.2115 | 254.1955 | 176.6159
2012-10-30 16:18 | 28.6143 | 271.7822 | 213.2559 | 269.1865 | 191.7273 | 76.1488 | 1.3164 | 33.3784 | 43.9384 | 25.3563 | 2.0835 | 33.3037 | 69.3892 | 15.9843 | 15.1352 | 5.2497 | 113.3276 | 256.647 | 198.1207 | 254.0513 | 176.5921
2012-10-30 16:18 | 28.6218 | 271.681 | 213.1689 | 269.159 | 191.7309 | 76.132 | 1.3163 | 33.3775 | 43.9359 | 25.3564 | 2.0835 | 33.3029 | 69.3866 | 15.9841 | 15.1264 | 5.2485 | 113.3225 | 256.5546 | 198.0425 | 254.0326 | 176.6045
2012-10-30 16:18 | 28.6233 | 271.5504 | 213.1232 | 269.0167 | 191.701 | 76.1332 | 1.3164 | 33.3785 | 43.9394 | 25.3559 | 2.0836 | 33.3039 | 69.3911 | 15.984 | 15.132 | 5.2486 | 113.3305 | 256.4184 | 197.9912 | 253.8847 | 176.569
2012-10-30 16:18 | 28.6308 | 271.5035 | 213.0873 | 268.962 | 191.5829 | 76.1297 | 1.3164 | 33.3785 | 43.9393 | 25.356 | 2.0836 | 33.3038 | 69.3921 | 15.9838 | 15.1151 | 5.2484 | 113.3314 | 256.3884 | 197.9722 | 253.8469 | 176.4678
2012-10-30 16:18 | 28.636 | 271.4349 | 213.0295 | 268.8816 | 191.5959 | 76.1233 | 1.3164 | 33.3782 | 43.9397 | 25.3553 | 2.0836 | 33.3035 | 69.3906 | 15.9838 | 15.108 | 5.2479 | 113.3303 | 256.3269 | 197.9215 | 253.7736 | 176.4879
2012-10-30 16:18 | 28.645 | 271.3352 | 212.9716 | 268.7818 | 191.5029 | 76.0835 | 1.3164 | 33.3788 | 43.941 | 25.3554 | 2.0836 | 33.3041 | 69.393 | 15.9838 | 15.1578 | 5.2452 | 113.334 | 256.1774 | 197.8138 | 253.624 | 176.3451
2012-10-30 16:18 | 28.6525 | 271.2261 | 212.924 | 268.7037 | 191.3534 | 76.0887 | 1.3165 | 33.3792 | 43.9426 | 25.3551 | 2.0837 | 33.3044 | 69.3952 | 15.9836 | 15.1327 | 5.2456 | 113.3378 | 256.0934 | 197.7913 | 253.571 | 176.2207
2012-10-30 16:18 | 28.6532 | 271.188 | 212.8459 | 268.615 | 191.2638 | 76.0898 | 1.3165 | 33.3797 | 43.9445 | 25.3549 | 2.0837 | 33.305 | 69.3985 | 15.9834 | 15.1293 | 5.2456 | 113.343 | 256.0587 | 197.7166 | 253.4857 | 176.1345
2012-10-30 16:18 | 28.6607 | 271.11 | 212.7825 | 268.5175 | 191.2006 | 76.0817 | 1.3165 | 33.3802 | 43.9461 | 25.3547 | 2.0838 | 33.3054 | 69.4008 | 15.9832 | 15.0757 | 5.2451 | 113.3469 | 256.0343 | 197.7068 | 253.4418 | 176.1249
2012-10-30 16:18 | 28.6622 | 271.0027 | 212.6895 | 268.4295 | 191.2021 | 76.0719 | 1.3165 | 33.3805 | 43.9466 | 25.3547 | 2.0838 | 33.3058 | 69.4023 | 15.9833 | 15.069 | 5.2444 | 113.3489 | 255.9337 | 197.6205 | 253.3605 | 176.1331
2012-10-30 16:18 | 28.6607 | 270.9235 | 212.6329 | 268.3619 | 190.9337 | 76.0765 | 1.3166 | 33.3803 | 43.9471 | 25.3543 | 2.0838 | 33.3055 | 69.4017 | 15.9831 | 15.1124 | 5.2447 | 113.3488 | 255.8111 | 197.5205 | 253.2495 | 175.8213
2012-10-30 16:18 | 28.6697 | 270.8781 | 212.5552 | 268.2193 | 190.7857 | 76.0719 | 1.3166 | 33.3802 | 43.9472 | 25.354 | 2.0838 | 33.3054 | 69.402 | 15.983 | 15.0887 | 5.2444 | 113.3492 | 255.7894 | 197.4665 | 253.1306 | 175.697
2012-10-30 16:18 | 28.6705 | 270.7896 | 212.4379 | 268.1812 | 190.7355 | 76.0459 | 1.3166 | 33.3804 | 43.9479 | 25.3539 | 2.0838 | 33.3056 | 69.4032 | 15.9829 | 15.1588 | 5.2426 | 113.3511 | 255.6308 | 197.2791 | 253.0224 | 175.5767
2012-10-30 16:18 | 28.678 | 270.7155 | 212.3706 | 268.0603 | 190.7548 | 76.0459 | 1.3166 | 33.3812 | 43.9505 | 25.3536 | 2.0839 | 33.3064 | 69.4068 | 15.9828 | 15.0767 | 5.2426 | 113.3572 | 255.6388 | 197.2939 | 252.9836 | 175.6781
2012-10-30 16:18 | 28.6854 | 270.6413 | 212.315 | 267.9899 | 190.7583 | 76.0471 | 1.3167 | 33.3817 | 43.9521 | 25.3535 | 2.0839 | 33.307 | 69.4099 | 15.9826 | 15.0922 | 5.2427 | 113.362 | 255.5491 | 197.2228 | 252.8977 | 175.6661
2012-10-30 16:18 | 28.6877 | 270.597 | 212.2306 | 267.9494 | 190.7841 | 76.032 | 1.3103 | 33.2195 | 43.529 | 25.3518 | 2.0741 | 33.1493 | 68.7549 | 15.9825 | 15.0864 | 5.2416 | 112.2839 | 255.5106 | 197.1442 | 252.863 | 175.6977
2012-10-30 16:18 | 28.6967 | 270.4699 | 212.1609 | 267.8573 | 190.6989 | 75.9985 | 1.3102 | 33.2151 | 43.5176 | 25.3516 | 2.0736 | 33.1409 | 68.7197 | 15.9826 | 15.1484 | 5.2393 | 112.2374 | 255.3215 | 197.0125 | 252.7089 | 175.5505
2012-10-30 16:18 | 28.7019 | 270.3935 | 212.0952 | 267.7808 | 190.763 | 76.002 | 1.3096 | 33.2001 | 43.48 | 25.3507 | 2.0726 | 33.1257 | 68.657 | 15.9825 | 15.1617 | 5.2396 | 112.1371 | 255.2318 | 196.9335 | 252.6191 | 175.6013
2012-10-30 16:18 | 28.6967 | 270.3067 | 212.0546 | 267.7289 | 189.9377 | 75.9898 | 1.3093 | 33.1919 | 43.4588 | 25.3505 | 2.0721 | 33.1173 | 68.6218 | 15.9826 | 15.1443 | 5.2387 | 112.0807 | 255.1624 | 196.9103 | 252.5846 | 174.7934
2012-10-30 16:18 | 28.7094 | 270.2145 | 211.9649 | 267.6444 | 190.0053 | 75.9806 | 1.3092 | 33.188 | 43.4483 | 25.3506 | 2.0719 | 33.1136 | 68.6067 | 15.9825 | 15.1733 | 5.2381 | 112.0549 | 255.0412 | 196.7916 | 252.4711 | 174.832
2012-10-30 16:18 | 28.7049 | 270.109 | 211.8462 | 267.5543 | 190.1734 | 75.9656 | 1.3091 | 33.1856 | 43.4419 | 25.3508 | 2.0717 | 33.1112 | 68.597 | 15.9825 | 15.2055 | 5.2371 | 112.0389 | 254.9035 | 196.6407 | 252.3488 | 174.9679
2012-10-30 16:18 | 28.7004 | 270.0656 | 211.7551 | 267.4837 | 190.3691 | 75.9488 | 1.3091 | 33.187 | 43.4466 | 25.3502 | 2.0718 | 33.1127 | 68.6035 | 15.9824 | 15.1643 | 5.2359 | 112.05 | 254.9013 | 196.5908 | 252.3194 | 175.2048
2012-10-30 16:18 | 28.7132 | 269.9657 | 211.7442 | 267.3953 | 190.4721 | 75.9378 | 1.3093 | 33.1909 | 43.4565 | 25.3502 | 2.072 | 33.1165 | 68.6187 | 15.9825 | 15.173 | 5.2352 | 112.0752 | 254.7927 | 196.5712 | 252.2223 | 175.2991
2012-10-30 16:18 | 28.7064 | 269.8695 | 211.7099 | 267.3185 | 190.4418 | 75.9199 | 1.3093 | 33.192 | 43.4597 | 25.3501 | 2.0721 | 33.1176 | 68.6237 | 15.9825 | 15.1744 | 5.2339 | 112.0833 | 254.6951 | 196.5355 | 252.1441 | 175.2674
2012-10-30 16:19 | 28.7004 | 269.8168 | 211.6172 | 267.2502 | 190.6948 | 75.9199 | 1.3094 | 33.1919 | 43.4604 | 25.3497 | 2.0721 | 33.1175 | 68.6239 | 15.9823 | 15.1806 | 5.2339 | 112.0843 | 254.6362 | 196.4366 | 252.0696 | 175.5142
2012-10-30 16:19 | 28.7012 | 269.635 | 211.6299 | 267.111 | 190.7584 | 75.9176 | 1.3094 | 33.1926 | 43.4621 | 25.3496 | 2.0722 | 33.1181 | 68.6263 | 15.9824 | 15.2425 | 5.2338 | 112.0884 | 254.3925 | 196.3874 | 251.8685 | 175.5159
2012-10-30 16:19 | 28.7094 | 269.6237 | 211.5672 | 267.0413 | 190.1387 | 75.906 | 1.3128 | 33.2801 | 43.6905 | 25.3503 | 2.0777 | 33.2061 | 68.9916 | 15.9823 | 15.214 | 5.233 | 112.6821 | 254.4097 | 196.3532 | 251.8273 | 174.9247
2012-10-30 16:19 | 28.7132 | 269.5537 | 211.508 | 266.9712 | 190.1543 | 75.8916 | 1.3147 | 33.3279 | 43.8148 | 25.351 | 2.0806 | 33.2533 | 69.188 | 15.9823 | 15.2219 | 5.232 | 113.0028 | 254.3318 | 196.2861 | 251.7493 | 174.9324
2012-10-30 16:19 | 28.7132 | 269.4876 | 211.4489 | 266.8778 | 190.2171 | 75.9066 | 1.3152 | 33.3427 | 43.8537 | 25.351 | 2.0815 | 33.2678 | 69.2484 | 15.9823 | 15.1934 | 5.233 | 113.1021 | 254.2942 | 196.2555 | 251.6844 | 175.0237
2012-10-30 16:19 | 28.7154 | 269.3926 | 211.3802 | 266.7983 | 190.0584 | 75.9107 | 1.3158 | 33.3565 | 43.8897 | 25.3511 | 2.0824 | 33.2814 | 69.3053 | 15.9822 | 15.1752 | 5.2333 | 113.195 | 254.2174 | 196.205 | 251.6231 | 174.8832
2012-10-30 16:19 | 28.7102 | 269.3135 | 211.3356 | 266.6919 | 189.9159 | 75.8974 | 1.3162 | 33.3664 | 43.9164 | 25.3508 | 2.0831 | 33.2913 | 69.3475 | 15.982 | 15.1619 | 5.2324 | 113.2639 | 254.1516 | 196.1737 | 251.53 | 174.754
2012-10-30 16:19 | 28.7177 | 269.2472 | 211.2408 | 266.6177 | 189.5506 | 75.8997 | 1.3164 | 33.3713 | 43.9295 | 25.3507 | 2.0834 | 33.2963 | 69.3684 | 15.982 | 15.1367 | 5.2325 | 113.2979 | 254.1105 | 196.1041 | 251.481 | 174.4139
2012-10-30 16:19 | 28.7072 | 269.1783 | 211.0964 | 266.5449 | 189.6029 | 75.8875 | 1.3165 | 33.3743 | 43.9378 | 25.3505 | 2.0836 | 33.2992 | 69.3818 | 15.9817 | 15.1385 | 5.2317 | 113.3196 | 254.0398 | 195.9579 | 251.4064 | 174.4644
2012-10-30 16:19 | 28.7244 | 269.04 | 211.1057 | 266.4687 | 189.522 | 75.8789 | 1.3166 | 33.3771 | 43.9458 | 25.3501 | 2.0838 | 33.3021 | 69.3936 | 15.9817 | 15.2005 | 5.2311 | 113.3394 | 253.8395 | 195.9052 | 251.2682 | 174.3215
2012-10-30 16:19 |  | 268.7095 | 210.7828 | 266.0406 | 189.6907 | 75.8598 | 1.3166 | 33.1409 | 43.6318 | 25.1724 | 2.0837 | 33.3 | 69.3869 | 15.9813 | 15.2005 | 5.2298 | 113.0186 | 253.509 | 195.5823 | 250.8401 | 174.4902
2012-10-30 16:19 | 28.7259 | 268.571 | 210.7174 | 265.8747 | 189.5118 | 75.8298 | 1.3167 | 33.3786 | 43.9513 | 25.3493 | 2.084 | 33.3036 | 69.4035 | 15.9808 | 15.1653 | 5.2277 | 113.3548 | 253.4057 | 195.5521 | 250.7094 | 174.3465
2012-10-30 16:19 | 28.7319 | 268.4876 | 210.7234 | 265.8184 | 189.4825 | 75.8402 | 1.3167 | 33.378 | 43.9502 | 25.349 | 2.0839 | 33.3029 | 69.4007 | 15.9808 | 15.1387 | 5.2284 | 113.351 | 253.3489 | 195.5847 | 250.6797 | 174.3438
2012-10-30 16:19 | 28.7416 | 268.3923 | 210.6858 | 265.7502 | 189.3156 | 75.8361 | 1.3167 | 33.3772 | 43.948 | 25.3489 | 2.0839 | 33.302 | 69.3977 | 15.9807 | 15.1362 | 5.2281 | 113.3457 | 253.2561 | 195.5496 | 250.614 | 174.1794
2012-10-30 16:19 | 28.7326 | 268.3716 | 210.5784 | 265.679 | 189.1222 | 75.8257 | 1.3167 | 33.377 | 43.9484 | 25.3485 | 2.0839 | 33.3018 | 69.3966 | 15.9808 | 15.1557 | 5.2274 | 113.345 | 253.2159 | 195.4227 | 250.5233 | 173.9665
2012-10-30 16:19 | 28.7506 | 268.3079 | 210.561 | 265.6074 | 189.2658 | 75.8141 | 1.3167 | 33.3771 | 43.9485 | 25.3485 | 2.0839 | 33.3017 | 69.3967 | 15.9806 | 15.1982 | 5.2266 | 113.3453 | 253.1097 | 195.3628 | 250.4092 | 174.0676
2012-10-30 16:19 | 28.7491 | 268.2637 | 210.4846 | 265.5553 | 189.28 | 75.8124 | 1.3167 | 33.3768 | 43.9482 | 25.3483 | 2.0839 | 33.3014 | 69.3964 | 15.9804 | 15.1641 | 5.2265 | 113.3446 | 253.0996 | 195.3205 | 250.3912 | 174.1159
2012-10-30 16:19 | 28.7469 | 268.2342 | 210.4588 | 265.5025 | 189.266 | 75.8274 | 1.3167 | 33.3764 | 43.9471 | 25.3484 | 2.0839 | 33.301 | 69.3947 | 15.9804 | 15.2271 | 5.2275 | 113.3418 | 253.0071 | 195.2317 | 250.2754 | 174.0389
2012-10-30 16:19 | 28.7566 | 268.1778 | 210.437 | 265.4499 | 189.3071 | 75.8124 | 1.3167 | 33.3765 | 43.9483 | 25.3477 | 2.0839 | 33.3009 | 69.3953 | 15.9802 | 15.1919 | 5.2265 | 113.3436 | 252.9859 | 195.2451 | 250.258 | 174.1152
2012-10-30 16:19 | 28.7559 | 268.0799 | 210.4284 | 265.4375 | 189.2397 | 75.8124 | 1.3167 | 33.3758 | 43.9462 | 25.348 | 2.0839 | 33.3003 | 69.3936 | 15.98 | 15.2239 | 5.2265 | 113.3398 | 252.856 | 195.2045 | 250.2136 | 174.0158
2012-10-30 16:20 | 0 | 239.1995 | 181.4993 | 236.4558 | 160.2171 | 75.8118 | 1.3165 | 32.4447 | 42.7149 | 24.6438 | 2.0836 | 33.2949 | 69.374 | 15.9793 | -13.581 | 5.2265 | 112.0889 | 252.7805 | 195.0803 | 250.0368 | 173.7981
2012-10-30 16:20 | 28.7753 | 267.9087 | 210.2234 | 265.2233 | 188.9218 | 75.7829 | 1.3166 | 33.3715 | 43.9357 | 25.3474 | 2.0836 | 33.2947 | 69.3738 | 15.9791 | 15.1943 | 5.2245 | 113.3096 | 252.7144 | 195.0291 | 250.029 | 173.7275
2012-10-30 16:20 | 28.7641 | 267.8625 | 210.1728 | 265.1225 | 188.785 | 75.7858 | 1.3166 | 33.3722 | 43.9378 | 25.3473 | 2.0837 | 33.2953 | 69.3771 | 15.979 | 15.1994 | 5.2247 | 113.3149 | 252.6631 | 194.9734 | 249.9231 | 173.5856
2012-10-30 16:20 | 28.7678 | 267.7456 | 210.0505 | 265.0172 | 188.6593 | 75.7731 | 1.3166 | 33.3725 | 43.939 | 25.3471 | 2.0837 | 33.2956 | 69.3783 | 15.979 | 15.2235 | 5.2238 | 113.3173 | 252.5221 | 194.827 | 249.7937 | 173.4358
2012-10-30 16:20 | 28.7768 | 267.6379 | 210.0005 | 264.8977 | 188.6252 | 75.7668 | 1.3167 | 33.3736 | 43.9427 | 25.3466 | 2.0838 | 33.2966 | 69.3839 | 15.9787 | 15.1755 | 5.2234 | 113.3266 | 252.4624 | 194.825 | 249.7222 | 173.4497
2012-10-30 16:20 | 28.7746 | 267.6045 | 209.888 | 264.8604 | 188.7132 | 75.7494 | 1.3167 | 33.3742 | 43.9444 | 25.3464 | 2.0839 | 33.2971 | 69.3864 | 15.9786 | 15.2222 | 5.2222 | 113.3308 | 252.3823 | 194.6658 | 249.6382 | 173.491
2012-10-30 16:20 | 28.7821 | 267.5537 | 209.927 | 264.8563 | 188.8227 | 75.7494 | 1.3167 | 33.3732 | 43.9419 | 25.3464 | 2.0838 | 33.2961 | 69.3822 | 15.9786 | 15.1808 | 5.2222 | 113.3241 | 252.3729 | 194.7462 | 249.6755 | 173.6419
2012-10-30 16:20 | 28.7873 | 267.5239 | 209.9125 | 264.7602 | 188.8985 | 75.7425 | 1.3167 | 33.3742 | 43.9446 | 25.3464 | 2.0839 | 33.297 | 69.3863 | 15.9785 | 15.1941 | 5.2217 | 113.3309 | 252.3298 | 194.7184 | 249.5661 | 173.7044
2012-10-30 16:20 | 28.7963 | 267.4979 | 209.9019 | 264.738 | 188.7231 | 75.739 | 1.3168 | 33.3752 | 43.9473 | 25.3463 | 2.0839 | 33.298 | 69.3909 | 15.9784 | 15.146 | 5.2214 | 113.3382 | 252.3519 | 194.7559 | 249.592 | 173.5771
2012-10-30 16:20 | 28.7836 | 267.4385 | 209.8695 | 264.6591 | 188.581 | 75.7413 | 1.3168 | 33.3751 | 43.9478 | 25.346 | 2.0839 | 33.2978 | 69.3905 | 15.9783 | 15.166 | 5.2216 | 113.3383 | 252.2725 | 194.7035 | 249.4931 | 173.415
2012-10-30 16:20 | 28.7926 | 267.3736 | 209.8312 | 264.6097 | 188.5272 | 75.7425 | 1.3168 | 33.375 | 43.947 | 25.3462 | 2.0839 | 33.2977 | 69.3899 | 15.9783 | 15.1994 | 5.2217 | 113.3369 | 252.1742 | 194.6318 | 249.4103 | 173.3278
2012-10-30 16:20 | 28.7963 | 267.35 | 209.7444 | 264.5745 | 188.5152 | 75.7153 | 1.3168 | 33.3746 | 43.9465 | 25.346 | 2.0839 | 33.2974 | 69.3891 | 15.9783 | 15.2439 | 5.2198 | 113.3356 | 252.1061 | 194.5005 | 249.3306 | 173.2713
2012-10-30 16:20 | 28.8075 | 267.3106 | 209.7674 | 264.4883 | 188.499 | 75.7171 | 1.3168 | 33.3747 | 43.9471 | 25.3457 | 2.0839 | 33.2974 | 69.3898 | 15.9781 | 15.2795 | 5.2199 | 113.3368 | 252.0311 | 194.4879 | 249.2088 | 173.2195
2012-10-30 16:20 | 28.806 | 267.2858 | 209.7029 | 264.4985 | 188.4857 | 75.7199 | 1.3168 | 33.3744 | 43.9466 | 25.3455 | 2.0839 | 33.2971 | 69.3889 | 15.978 | 15.2821 | 5.2201 | 113.3355 | 252.0037 | 194.4208 | 249.2164 | 173.2036
2012-10-30 16:20 | 28.8128 | 267.172 | 209.6782 | 264.4585 | 188.4337 | 75.7107 | 1.3167 | 33.373 | 43.9435 | 25.3452 | 2.0838 | 33.2958 | 69.383 | 15.9781 | 15.2318 | 5.2195 | 113.3266 | 251.9402 | 194.4464 | 249.2267 | 173.2019
2012-10-30 16:20 | 28.809 | 267.1526 | 209.6114 | 264.4002 | 188.277 | 75.6997 | 1.3167 | 33.3733 | 43.9437 | 25.3456 | 2.0839 | 33.296 | 69.3849 | 15.9779 | 15.2117 | 5.2187 | 113.3286 | 251.9409 | 194.3997 | 249.1885 | 173.0653
2012-10-30 16:20 | 28.8135 | 267.1143 | 209.608 | 264.3852 | 188.2697 | 75.7275 | 1.3167 | 33.3733 | 43.9436 | 25.3455 | 2.0839 | 33.296 | 69.3848 | 15.9779 | 15.2244 | 5.2207 | 113.3284 | 251.8899 | 194.3836 | 249.1608 | 173.0453
2012-10-30 16:20 | 28.8075 | 267.0694 | 209.5469 | 264.3559 | 188.1225 | 75.7182 | 1.3167 | 33.3728 | 43.9429 | 25.3452 | 2.0839 | 33.2953 | 69.3827 | 15.9778 | 15.1776 | 5.22 | 113.3256 | 251.8918 | 194.3693 | 249.1783 | 172.9449
2012-10-30 16:20 | 28.8165 | 267.0434 | 209.4889 | 264.3103 | 188.1354 | 75.6968 | 1.3167 | 33.3724 | 43.9423 | 25.345 | 2.0838 | 33.295 | 69.3817 | 15.9776 | 15.1785 | 5.2185 | 113.324 | 251.8649 | 194.3104 | 249.1318 | 172.9569
2012-10-30 16:20 | 28.8225 | 267.0105 | 209.4713 | 264.2384 | 187.8591 | 75.6945 | 1.3167 | 33.3718 | 43.9406 | 25.345 | 2.0838 | 33.2943 | 69.3789 | 15.9776 | 15.1682 | 5.2184 | 113.3195 | 251.8423 | 194.3031 | 249.0702 | 172.6909
2012-10-30 16:20 | 28.8255 | 266.9823 | 209.4034 | 264.183 | 187.6778 | 75.7003 | 1.3167 | 33.3717 | 43.941 | 25.3446 | 2.0838 | 33.2942 | 69.3794 | 15.9775 | 15.1549 | 5.2188 | 113.3204 | 251.8274 | 194.2485 | 249.0281 | 172.5229
2012-10-30 16:20 | 28.8143 | 266.9128 | 209.2899 | 264.1367 | 187.7215 | 75.676 | 1.3167 | 33.3715 | 43.9405 | 25.3447 | 2.0838 | 33.294 | 69.3781 | 15.9776 | 15.1233 | 5.2171 | 113.3186 | 251.7895 | 194.1666 | 249.0134 | 172.5982
2012-10-30 16:21 | 28.8158 | 266.8754 | 209.2205 | 264.1421 | 187.8014 | 75.6772 | 1.3167 | 33.3716 | 43.9414 | 25.3443 | 2.0838 | 33.2942 | 69.3799 | 15.9773 | 15.137 | 5.2172 | 113.3212 | 251.7384 | 194.0835 | 249.0051 | 172.6644
2012-10-30 16:21 | 28.8225 | 266.8782 | 209.1642 | 264.0787 | 187.9964 | 75.6726 | 1.3167 | 33.3709 | 43.9397 | 25.3442 | 2.0838 | 33.2934 | 69.3766 | 15.9773 | 15.1478 | 5.2169 | 113.3163 | 251.7304 | 194.0164 | 248.9309 | 172.8486
2012-10-30 16:21 | 28.8323 | 266.814 | 209.1937 | 264.0223 | 187.8924 | 75.6772 | 1.3167 | 33.3707 | 43.9387 | 25.3445 | 2.0838 | 33.2934 | 69.3768 | 15.9772 | 15.1535 | 5.2172 | 113.3156 | 251.6605 | 194.0402 | 248.8688 | 172.7389
2012-10-30 16:21 | 28.8218 | 266.7996 | 209.1478 | 263.9455 | 187.9917 | 75.6593 | 1.3167 | 33.3708 | 43.9395 | 25.3441 | 2.0838 | 33.2933 | 69.3763 | 15.9773 | 15.1104 | 5.216 | 113.3158 | 251.6892 | 194.0374 | 248.8351 | 172.8813
2012-10-30 16:21 | 28.8323 | 266.7051 | 209.1465 | 263.9405 | 188.0493 | 75.6789 | 1.3167 | 33.3704 | 43.9387 | 25.344 | 2.0838 | 33.2928 | 69.3747 | 15.9771 | 15.1494 | 5.2173 | 113.3134 | 251.5557 | 193.9971 | 248.7911 | 172.8999
2012-10-30 16:21 | 28.8248 | 266.6664 | 209.1114 | 263.9057 | 187.8457 | 75.6795 | 1.3126 | 33.2654 | 43.6633 | 25.3436 | 2.0773 | 33.1894 | 68.9448 | 15.9771 | 15.1908 | 5.2173 | 112.6081 | 251.4756 | 193.9206 | 248.7149 | 172.6549
2012-10-30 16:21 | 28.827 | 266.6453 | 209.09 | 263.8689 | 187.7303 | 75.6367 | 1.3106 | 33.2127 | 43.5274 | 25.3423 | 2.0739 | 33.1354 | 68.7205 | 15.9771 | 15.1727 | 5.2144 | 112.2479 | 251.4726 | 193.9173 | 248.6962 | 172.5576
2012-10-30 16:21 | 28.8338 | 266.5898 | 208.9826 | 263.8679 | 187.6509 | 75.6437 | 1.31 | 33.1984 | 43.4906 | 25.3419 | 2.0731 | 33.1213 | 68.6621 | 15.977 | 15.1713 | 5.2149 | 112.1527 | 251.4185 | 193.8113 | 248.6966 | 172.4796
2012-10-30 16:21 | 28.8278 | 266.5293 | 208.953 | 263.7957 | 187.0333 | 75.6373 | 1.3097 | 33.19 | 43.4688 | 25.3417 | 2.0725 | 33.1128 | 68.6271 | 15.977 | 15.1938 | 5.2144 | 112.0959 | 251.3355 | 193.7592 | 248.6019 | 171.8395
2012-10-30 16:21 | 28.8255 | 266.4725 | 208.8129 | 263.7583 | 186.7409 | 75.6165 | 1.3094 | 33.1832 | 43.4511 | 25.3417 | 2.0721 | 33.1061 | 68.5993 | 15.977 | 15.1834 | 5.213 | 112.0504 | 251.2891 | 193.6295 | 248.5749 | 171.5575
2012-10-30 16:21 |  | 266.1221 | 208.3167 | 263.3842 | 186.976 | 75.5951 | 1.313 | 33.2738 | 43.6885 | 25.3418 | 2.0778 | 33.1962 | 68.9742 | 15.9768 | 15.1709 | 5.2115 | 112.6627 | 250.9512 | 193.1458 | 248.2133 | 171.8051
2012-10-30 16:21 | 28.8382 | 265.7923 | 208.0342 | 263.0151 | 186.9692 | 75.5951 | 1.2712 | 32.2203 | 40.9599 | 25.3454 | 2.0372 | 32.5496 | 66.3112 | 15.9773 | 15.1676 | 5.2115 | 107.2711 | 250.6247 | 192.8666 | 247.8475 | 171.8016
2012-10-30 16:21 | 28.8345 | 265.5589 | 207.6487 | 262.9763 | 186.9381 | 74.6369 | 0.028378 | 0.87771 | 0.024908 | 30.9293 | 0.065275 | 1.1452 | 0.074756 | 17.545 | 15.1802 | 5.1455 | 0.099664 | 250.3787 | 192.4685 | 247.7961 | 171.7579
2012-10-30 16:21 | 28.8368 | 264.1553 | 205.8678 | 262.8656 | 186.7797 | 73.3793 | 0.02349 | 0.7533 | 0.017695 | 32.0689 | 0.040771 | 0.75512 | 0.030787 | 18.521 | 15.2232 | 5.0588 | 0.048482 | 248.9321 | 190.6446 | 247.6424 | 171.5565
2012-10-30 16:22 | 28.848 | 262.452 | 204.8084 | 262.5104 | 186.7321 | 73.1881 | 0.37622 | 9.5641 | 3.5982 | 25.4217 | 0.58524 | 9.4201 | 5.513 | 16.0961 | 15.1774 | 5.0456 | 9.1112 | 247.2746 | 189.631 | 247.333 | 171.5547
2012-10-30 16:22 | 28.8263 | 260.8004 | 204.4404 | 262.056 | 186.2949 | 73.8885 | 1.1851 | 29.9353 | 35.4766 | 25.2595 | 1.8527 | 29.5976 | 54.8356 | 15.9753 | 15.1638 | 5.0939 | 90.3122 | 245.6366 | 189.2766 | 246.8922 | 171.1311
2012-10-30 16:22 | 28.8427 | 260.3877 | 205.2833 | 261.5344 | 186.1624 | 75.1917 | 1.5671 | 39.7787 | 62.3371 | 25.3837 | 2.4734 | 39.5316 | 97.7785 | 15.9825 | 15.2087 | 5.1837 | 160.1156 | 245.179 | 190.0746 | 246.3257 | 170.9537
2012-10-30 16:22 | 28.842 | 261.1827 | 206.609 | 261.171 | 185.8247 | 76.2586 | 1.7986 | 45.8214 | 82.4135 | 25.4765 | 2.855 | 45.706 | 130.4918 | 16.009 | 15.2121 | 5.2573 | 212.9053 | 245.9706 | 191.3969 | 245.9589 | 170.6126
2012-10-30 16:22 | 28.8555 | 262.9155 | 208.2208 | 260.9622 | 185.3328 | 77.3925 | 2.0021 | 51.2127 | 102.5319 | 25.5797 | 3.1798 | 51.1181 | 162.5451 | 16.0759 | 15.2338 | 5.3354 | 265.077 | 247.6817 | 192.987 | 245.7284 | 170.099
2012-10-30 16:22 | 28.848 | 265.4284 | 209.8787 | 261.0093 | 185.1568 | 78.001 | 2.0032 | 51.2865 | 102.7357 | 25.6027 | 3.1773 | 51.157 | 162.5389 | 16.101 | 15.2304 | 5.3774 | 265.2746 | 250.198 | 194.6483 | 245.7789 | 169.9264
2012-10-30 16:22 | 28.8525 | 267.9781 | 211.0212 | 261.2868 | 185.7294 | 78.179 | 2.0146 | 51.5995 | 103.9537 | 25.6125 | 3.1952 | 51.4773 | 164.4812 | 16.1107 | 15.2389 | 5.3897 | 268.4349 | 252.7392 | 195.7823 | 246.0479 | 170.4905
2012-10-30 16:22 | 28.8585 | 270.5579 | 212.1179 | 261.8582 | 186.382 | 78.4408 | 2.0139 | 51.5945 | 103.9047 | 25.6196 | 3.1935 | 51.4721 | 164.3787 | 16.1175 | 15.2572 | 5.4077 | 268.2835 | 255.3007 | 196.8607 | 246.601 | 171.1248
2012-10-30 16:22 | 28.8577 | 272.744 | 213.0898 | 262.4539 | 187.1416 | 78.6367 | 2.0134 | 51.5921 | 103.8756 | 25.6243 | 3.1926 | 51.4695 | 164.3203 | 16.1216 | 15.2441 | 5.4212 | 268.196 | 257.4999 | 197.8457 | 247.2098 | 171.8975
2012-10-30 16:22 | 28.8495 | 274.7424 | 214.0031 | 263.2368 | 187.8351 | 78.7858 | 2.013 | 51.5895 | 103.8471 | 25.6288 | 3.1917 | 51.4669 | 164.2662 | 16.1253 | 15.2196 | 5.4315 | 268.1133 | 259.5228 | 198.7835 | 248.0172 | 172.6155
2012-10-30 16:22 | 28.8517 | 276.7766 | 215.0094 | 264.2559 | 188.5981 | 78.9558 | 2.0125 | 51.5881 | 103.8216 | 25.6337 | 3.1909 | 51.4655 | 164.2227 | 16.1287 | 15.2504 | 5.4432 | 268.0443 | 261.5262 | 199.759 | 249.0055 | 173.3477
2012-10-30 16:22 | 28.8517 | 278.8071 | 216.0215 | 265.2724 | 189.3865 | 79.1164 | 2.0121 | 51.5871 | 103.7989 | 25.6383 | 3.1902 | 51.4645 | 164.1811 | 16.1322 | 15.1811 | 5.4543 | 267.98 | 263.626 | 200.8404 | 250.0913 | 174.2054
2012-10-30 16:22 | 28.857 | 280.6903 | 217.0783 | 266.4417 | 190.0587 | 79.2863 | 2.0091 | 51.5195 | 103.5103 | 25.6425 | 3.1857 | 51.4015 | 163.7514 | 16.1349 | 15.2149 | 5.466 | 267.2617 | 265.4754 | 201.8634 | 251.2268 | 174.8438
2012-10-30 16:22 | 28.8622 | 282.7422 | 218.0915 | 267.8983 | 190.939 | 79.4638 | 2.0114 | 51.5859 | 103.7611 | 25.6465 | 3.189 | 51.4632 | 164.1158 | 16.1377 | 15.2282 | 5.4782 | 267.8769 | 267.514 | 202.8633 | 252.6701 | 175.7108
2012-10-30 16:22 | 28.857 | 284.5347 | 219.1141 | 269.2154 | 191.8169 | 79.6314 | 2.011 | 51.5852 | 103.7388 | 25.6513 | 3.1884 | 51.4623 | 164.0802 | 16.1407 | 15.2027 | 5.4898 | 267.819 | 269.332 | 203.9114 | 254.0127 | 176.6142
2012-10-30 16:22 | 28.8615 | 286.2854 | 220.2644 | 270.7591 | 192.7872 | 79.7839 | 2.0107 | 51.5849 | 103.7219 | 25.6551 | 3.1878 | 51.4619 | 164.049 | 16.1435 | 15.2194 | 5.5003 | 267.7709 | 271.066 | 205.045 | 255.5397 | 177.5678
2012-10-30 16:22 | 28.8645 | 288.1916 | 221.4329 | 272.4324 | 193.7642 | 79.9492 | 2.0104 | 51.5849 | 103.7053 | 25.6593 | 3.1872 | 51.4619 | 164.0192 | 16.1464 | 15.2346 | 5.5117 | 267.7245 | 272.957 | 206.1983 | 257.1978 | 178.5296
2012-10-30 16:22 | 28.8705 | 290.0301 | 222.5491 | 274.1193 | 194.5243 | 80.1255 | 2.0101 | 51.5855 | 103.6925 | 25.6631 | 3.1867 | 51.4626 | 163.9975 | 16.149 | 15.261 | 5.5239 | 267.6899 | 274.7691 | 207.2881 | 258.8583 | 179.2633
2012-10-30 16:22 | 28.872 | 291.724 | 223.5702 | 275.6105 | 195.4058 | 80.2804 | 2.0098 | 51.5855 | 103.6784 | 25.6665 | 3.1863 | 51.4624 | 163.9724 | 16.1514 | 15.234 | 5.5345 | 267.6508 | 276.49 | 208.3362 | 260.3765 | 180.1718
2012-10-30 16:22 |  | 245.0935 | 185.3229 | 249.5409 | 165.5242 | 75.1004 | 0.79019 | 19.3805 | 15.3143 | 24.5262 | 1.2378 | 19.9246 | 24.6632 | 16.0964 | -13.6747 | 5.1774 | 39.9775 | 258.7682 | 198.9976 | 263.2156 | 179.1989
2012-10-30 16:23 | 28.854 | 270.2386 | 212.3458 | 275.3326 | 192.9172 | 74.8265 | 0.79217 | 20.0123 | 15.8532 | 25.2626 | 1.2406 | 19.9689 | 24.7738 | 16.0959 | 15.1181 | 5.1585 | 40.627 | 255.1205 | 197.2277 | 260.2145 | 177.7991
2012-10-30 16:23 | 28.854 | 266.6878 | 210.3927 | 272.0025 | 191.2958 | 74.4791 | 0.79273 | 20.0186 | 15.8694 | 25.2526 | 1.241 | 19.9757 | 24.7907 | 16.0959 | 15.1711 | 5.1346 | 40.6601 | 251.5167 | 195.2216 | 256.8314 | 176.1247
2012-10-30 16:23 | 28.8457 | 265.6714 | 209.8134 | 271.0036 | 190.7538 | 74.3982 | 0.79209 | 20.0004 | 15.8422 | 25.25 | 1.2399 | 19.9572 | 24.7446 | 16.0961 | 15.1098 | 5.129 | 40.5868 | 250.5616 | 194.7036 | 255.8938 | 175.644
2012-10-30 16:23 | 28.857 | 264.7793 | 209.2814 | 270.0706 | 190.5257 | 74.3034 | 0.79165 | 19.9872 | 15.8229 | 25.2474 | 1.2391 | 19.9442 | 24.7124 | 16.096 | 15.1415 | 5.1225 | 40.5353 | 249.6378 | 194.1399 | 254.9291 | 175.3842
2012-10-30 16:23 | 28.8465 | 263.8806 | 208.7354 | 269.096 | 190.0992 | 74.2277 | 0.79135 | 19.9778 | 15.8094 | 25.2451 | 1.2385 | 19.9345 | 24.6882 | 16.0961 | 15.0902 | 5.1173 | 40.4976 | 248.7904 | 193.6452 | 254.0058 | 175.009
2012-10-30 16:23 | 28.8502 | 262.8283 | 208.1092 | 267.9991 | 189.7028 | 74.1353 | 0.79188 | 19.9883 | 15.8283 | 25.2415 | 1.2392 | 19.9455 | 24.7156 | 16.0961 | 15.098 | 5.1109 | 40.5439 | 247.7303 | 193.0112 | 252.9011 | 174.6048
2012-10-30 16:23 | 28.8517 | 261.8787 | 207.5517 | 266.9463 | 189.0648 | 74.0399 | 0.79317 | 20.0193 | 15.8787 | 25.2397 | 1.2411 | 19.9766 | 24.7929 | 16.096 | 15.1199 | 5.1043 | 40.6717 | 246.7588 | 192.4318 | 251.8264 | 173.9449
2012-10-30 16:23 | 28.8517 | 260.8687 | 206.973 | 265.872 | 188.5197 | 73.944 | 0.79361 | 20.0277 | 15.8941 | 25.2363 | 1.2416 | 19.9845 | 24.8123 | 16.096 | 15.1444 | 5.0977 | 40.7065 | 245.7243 | 191.8286 | 250.7276 | 173.3753
2012-10-30 16:23 | 28.8577 | 259.9112 | 206.3295 | 264.8033 | 187.9806 | 73.8492 | 0.79325 | 20.0167 | 15.8782 | 25.2339 | 1.2409 | 19.9731 | 24.7843 | 16.096 | 15.1504 | 5.0912 | 40.6625 | 244.7608 | 191.1791 | 249.6529 | 172.8302
2012-10-30 16:23 | 28.8525 | 259.0006 | 205.8087 | 263.8243 | 187.497 | 73.7458 | 0.79245 | 19.995 | 15.845 | 25.2319 | 1.2395 | 19.9512 | 24.7296 | 16.0961 | 15.137 | 5.084 | 40.5746 | 243.8636 | 190.6717 | 248.6873 | 172.36
2012-10-30 16:23 | 28.866 | 258.1202 | 205.2554 | 262.8285 | 186.9931 | 73.6984 | 0.7924 | 19.9917 | 15.8415 | 25.2293 | 1.2393 | 19.9485 | 24.7224 | 16.0965 | 15.1383 | 5.0808 | 40.5639 | 242.9819 | 190.1171 | 247.6902 | 171.8548
2012-10-30 16:23 | 28.8495 | 257.2642 | 204.6171 | 261.7712 | 186.467 | 73.6047 | 0.79251 | 19.9927 | 15.8444 | 25.2271 | 1.2394 | 19.9496 | 24.725 | 16.0965 | 15.1299 | 5.0743 | 40.5694 | 242.1343 | 189.4872 | 246.6413 | 171.3371
2012-10-30 16:23 | 28.8585 | 256.2421 | 203.9886 | 260.7117 | 186.0097 | 73.5134 | 0.79249 | 19.9896 | 15.8415 | 25.2238 | 1.2392 | 19.9465 | 24.7171 | 16.0966 | 15.196 | 5.068 | 40.5587 | 241.0461 | 188.7926 | 245.5157 | 170.8137
2012-10-30 16:23 | 28.8592 | 255.3871 | 203.4384 | 259.7449 | 185.4618 | 73.4238 | 0.79285 | 19.997 | 15.8546 | 25.2218 | 1.2397 | 19.9539 | 24.736 | 16.0963 | 15.0989 | 5.0618 | 40.5905 | 240.2882 | 188.3395 | 244.646 | 170.3629
2012-10-30 16:23 | 28.863 | 254.4958 | 202.8559 | 258.8004 | 185.1091 | 73.3389 | 0.79311 | 20.0015 | 15.8633 | 25.2192 | 1.2399 | 19.9584 | 24.7471 | 16.0963 | 15.1394 | 5.056 | 40.6105 | 239.3564 | 187.7165 | 243.661 | 169.9697
2012-10-30 16:23 | 28.8735 | 253.6306 | 202.2093 | 257.7998 | 184.5321 | 73.243 | 0.79338 | 20.0065 | 15.8726 | 25.2169 | 1.2402 | 19.9628 | 24.7578 | 16.0965 | 15.1621 | 5.0494 | 40.6304 | 238.4685 | 187.0472 | 242.6377 | 169.37
2012-10-30 16:23 | 28.8727 | 252.789 | 201.7247 | 256.8969 | 183.9282 | 73.1759 | 0.79035 | 19.9283 | 15.7504 | 25.2144 | 1.2354 | 19.8858 | 24.5671 | 16.0965 | 15.2021 | 5.0447 | 40.3175 | 237.5869 | 186.5226 | 241.6948 | 168.7261
2012-10-30 16:23 | 28.8735 | 251.9057 | 201.1354 | 255.8273 | 183.4983 | 73.0696 | 0.78798 | 19.8663 | 15.6542 | 25.2118 | 1.2314 | 19.8225 | 24.4099 | 16.0972 | 15.158 | 5.0374 | 40.0641 | 236.7477 | 185.9774 | 240.6693 | 168.3403
2012-10-30 16:23 | 28.8794 | 251.0508 | 200.5238 | 254.7468 | 183.1597 | 72.9875 | 0.78682 | 19.8355 | 15.6069 | 25.2098 | 1.2295 | 19.7925 | 24.3356 | 16.0975 | 15.1965 | 5.0318 | 39.9425 | 235.8543 | 185.3273 | 239.5503 | 167.9632
2012-10-30 16:23 | 28.875 | 250.1736 | 199.8783 | 253.8159 | 181.7973 | 72.895 | 0.78662 | 19.8287 | 15.5976 | 25.2076 | 1.2291 | 19.7859 | 24.3194 | 16.0976 | 15.2737 | 5.0254 | 39.917 | 234.8999 | 184.6046 | 238.5422 | 166.5236
2012-10-30 16:23 | 28.8742 | 249.3546 | 199.2798 | 252.9626 | 181.5383 | 72.8216 | 0.78601 | 19.8115 | 15.572 | 25.2052 | 1.228 | 19.7684 | 24.2763 | 16.0976 | 15.2484 | 5.0203 | 39.8482 | 234.1062 | 184.0314 | 237.7142 | 166.2899
2012-10-30 16:24 | 28.875 | 248.4508 | 198.6672 | 251.9425 | 181.0188 | 72.7066 | 0.78598 | 19.809 | 15.5694 | 25.2031 | 1.2279 | 19.7662 | 24.2702 | 16.0981 | 15.3348 | 5.0124 | 39.8396 | 233.116 | 183.3324 | 236.6077 | 165.684
2012-10-30 16:24 | 28.8817 | 247.5369 | 198.0448 | 250.9866 | 180.13 | 72.6298 | 0.78613 | 19.8111 | 15.5741 | 25.2009 | 1.228 | 19.7682 | 24.2753 | 16.0979 | 15.3415 | 5.0071 | 39.8494 | 232.1954 | 182.7033 | 235.6451 | 164.7885
2012-10-30 16:24 | 28.878 | 246.8202 | 197.3689 | 250.2001 | 179.4655 | 72.5622 | 0.78616 | 19.8104 | 15.5742 | 25.1989 | 1.2279 | 19.7676 | 24.2735 | 16.0982 | 15.3215 | 5.0024 | 39.8477 | 231.4987 | 182.0474 | 234.8786 | 164.144
2012-10-30 16:24 | 28.8765 | 246.0741 | 196.7857 | 249.3647 | 179.4914 | 72.4945 | 0.78622 | 19.8099 | 15.5749 | 25.1963 | 1.2279 | 19.7672 | 24.272 | 16.0984 | 15.3608 | 4.9978 | 39.8469 | 230.7133 | 181.4249 | 234.0039 | 164.1306
2012-10-30 16:24 | 28.8794 | 245.2695 | 196.1913 | 248.4199 | 177.6578 | 72.3917 | 0.78646 | 19.8146 | 15.5834 | 25.1946 | 1.2282 | 19.7715 | 24.2825 | 16.0986 | 15.3392 | 4.9907 | 39.8659 | 229.9303 | 180.8521 | 233.0807 | 162.3186
2012-10-30 16:24 | 28.878 | 244.4918 | 195.4433 | 247.4079 | 177.0979 | 72.309 | 0.78832 | 19.8596 | 15.6558 | 25.1922 | 1.231 | 19.8169 | 24.3942 | 16.0984 | 15.3582 | 4.985 | 40.05 | 229.1336 | 180.0851 | 232.0497 | 161.7397
2012-10-30 16:24 | 28.8697 | 243.7188 | 194.8221 | 246.5807 | 176.9216 | 72.2431 | 0.79104 | 19.9263 | 15.7625 | 25.1902 | 1.2351 | 19.8831 | 24.5582 | 16.0981 | 15.3662 | 4.9804 | 40.3206 | 228.3526 | 179.4559 | 231.2145 | 161.5554
2012-10-30 16:24 | 28.869 | 242.9847 | 194.2282 | 245.8784 | 176.1401 | 72.1663 | 0.79155 | 19.9376 | 15.7816 | 25.1882 | 1.2358 | 19.8945 | 24.5861 | 16.0982 | 15.3614 | 4.9751 | 40.3676 | 227.6233 | 178.8668 | 230.517 | 160.7787
2012-10-30 16:24 | 28.8652 | 242.314 | 193.6156 | 245.075 | 175.7226 | 72.105 | 0.79216 | 19.9517 | 15.805 | 25.1863 | 1.2367 | 19.9084 | 24.6209 | 16.0979 | 15.3291 | 4.9709 | 40.4259 | 226.9849 | 178.2865 | 229.7459 | 160.3935
2012-10-30 16:24 | 28.8712 | 241.6059 | 192.9894 | 244.3007 | 175.2291 | 72.0299 | 0.79263 | 19.9618 | 15.8223 | 25.1843 | 1.2373 | 19.9184 | 24.6459 | 16.0978 | 15.3514 | 4.9657 | 40.4681 | 226.2545 | 177.638 | 228.9493 | 159.8777
2012-10-30 16:24 | 28.8787 | 240.8599 | 192.4432 | 243.4337 | 175.5605 | 71.9524 | 0.79293 | 19.9674 | 15.8327 | 25.1817 | 1.2377 | 19.9241 | 24.6598 | 16.0978 | 15.3833 | 4.9604 | 40.4925 | 225.4766 | 177.0599 | 228.0504 | 160.1772
2012-10-30 16:24 | 28.8727 | 240.0689 | 191.919 | 242.5216 | 175.3281 | 71.8733 | 0.79294 | 19.9661 | 15.832 | 25.1798 | 1.2375 | 19.9214 | 24.6531 | 16.0979 | 15.361 | 4.9549 | 40.4851 | 224.7079 | 176.558 | 227.1606 | 159.9671
2012-10-30 16:24 | 28.8727 | 239.374 | 191.5461 | 241.784 | 175.0355 | 72.0172 | 1.1903 | 30.0005 | 35.7106 | 25.2034 | 1.8405 | 29.5764 | 54.435 | 16.0699 | 15.3529 | 4.9649 | 90.1456 | 224.0211 | 176.1932 | 226.4311 | 159.6826
2012-10-30 16:24 | 28.8622 | 239.3439 | 192.1519 | 241.0279 | 174.7051 | 73.4007 | 1.6783 | 42.562 | 71.431 | 25.3605 | 2.6252 | 42.197 | 110.7738 | 16.0741 | 15.3587 | 5.0602 | 182.2048 | 223.9852 | 176.7932 | 225.6692 | 159.3464
2012-10-30 16:24 | 28.8757 | 240.8844 | 193.8578 | 240.4722 | 174.4455 | 75.1264 | 2.0059 | 51.1924 | 102.6875 | 25.5208 | 3.1654 | 50.9849 | 161.3893 | 16.1068 | 15.3763 | 5.1792 | 264.0768 | 225.5081 | 178.4815 | 225.0959 | 159.0692
2012-10-30 16:24 | 28.8682 | 244.0781 | 196.2076 | 240.2214 | 174.4497 | 76.0702 | 2.0192 | 51.6122 | 104.2167 | 25.5604 | 3.1957 | 51.4893 | 164.5434 | 16.1121 | 15.3484 | 5.2443 | 268.7601 | 228.7297 | 180.8592 | 224.873 | 159.1013
2012-10-30 16:24 | 28.869 | 247.3754 | 197.7805 | 240.3949 | 174.8562 | 76.3487 | 2.0182 | 51.604 | 104.1467 | 25.5694 | 3.1951 | 51.4808 | 164.488 | 16.1123 | 15.3533 | 5.2635 | 268.6347 | 232.0221 | 182.4272 | 225.0416 | 159.5029
2012-10-30 16:24 | 28.8765 | 250.179 | 198.8692 | 240.8027 | 175.3085 | 76.5556 | 2.0175 | 51.5989 | 104.0995 | 25.5759 | 3.1946 | 51.4755 | 164.4426 | 16.1134 | 15.373 | 5.2777 | 268.5422 | 234.806 | 183.4962 | 225.4297 | 159.9355
2012-10-30 16:24 | 28.8742 | 252.7475 | 199.9247 | 241.4244 | 175.8018 | 76.7712 | 2.0169 | 51.596 | 104.0628 | 25.5821 | 3.1941 | 51.4724 | 164.4088 | 16.1148 | 15.3585 | 5.2926 | 268.4716 | 237.389 | 184.5662 | 226.0659 | 160.4433
2012-10-30 16:24 | 28.8787 | 255.473 | 201.0973 | 242.3432 | 176.7276 | 76.9636 | 2.0163 | 51.5939 | 104.0303 | 25.588 | 3.1937 | 51.4703 | 164.3784 | 16.1165 | 15.3059 | 5.3059 | 268.4087 | 240.1671 | 185.7914 | 227.0373 | 161.4217
2012-10-30 16:24 | 28.8787 | 257.8518 | 202.1791 | 243.4062 | 177.4891 | 77.1555 | 2.0159 | 51.5934 | 104.0051 | 25.5937 | 3.1932 | 51.4697 | 164.3545 | 16.1184 | 15.2977 | 5.3191 | 268.3597 | 242.5541 | 186.8814 | 228.1085 | 162.1914
2012-10-30 16:24 | 28.8772 | 260.1101 | 203.2793 | 244.5773 | 178.1673 | 77.3497 | 2.0154 | 51.5929 | 103.9802 | 25.5994 | 3.1928 | 51.4691 | 164.3304 | 16.1204 | 15.2555 | 5.3325 | 268.3106 | 244.8546 | 188.0238 | 229.3218 | 162.9118
2012-10-30 16:24 | 28.8824 | 262.6929 | 204.5359 | 246.1975 | 179.1729 | 77.5595 | 2.0149 | 51.5925 | 103.955 | 25.6052 | 3.1923 | 51.4686 | 164.3047 | 16.1226 | 15.2729 | 5.347 | 268.2598 | 247.42 | 189.263 | 230.9246 | 163.9
2012-10-30 16:24 | 28.8847 | 265.0251 | 205.7622 | 247.7867 | 180.2112 | 77.7485 | 2.0145 | 51.5925 | 103.9345 | 25.6102 | 3.1919 | 51.4685 | 164.2831 | 16.1246 | 15.2915 | 5.36 | 268.2176 | 249.7336 | 190.4707 | 232.4952 | 164.9197
2012-10-30 16:24 | 28.8854 | 267.0889 | 206.8532 | 249.2052 | 181.0605 | 77.9403 | 2.0141 | 51.5921 | 103.9139 | 25.615 | 3.1915 | 51.4681 | 164.2585 | 16.1268 | 15.2514 | 5.3732 | 268.1724 | 251.8375 | 191.6018 | 233.9538 | 165.8091
2012-10-30 16:24 | 28.8877 | 269.4794 | 208.0601 | 250.9534 | 182.0526 | 78.1357 | 2.0137 | 51.5923 | 103.893 | 25.6203 | 3.191 | 51.4683 | 164.2364 | 16.1291 | 15.2782 | 5.3867 | 268.1294 | 254.2012 | 192.7819 | 235.6752 | 166.7744
2012-10-30 16:25 | 28.8847 | 271.6292 | 209.376 | 252.9614 | 183.0906 | 78.3403 | 2.0134 | 51.5931 | 103.8774 | 25.6249 | 3.1906 | 51.4687 | 164.2158 | 16.1314 | 15.263 | 5.4008 | 268.0932 | 256.3662 | 194.113 | 237.6984 | 167.8276
2012-10-30 16:25 | 28.8877 | 273.5582 | 210.5169 | 254.8841 | 183.7161 | 78.4946 | 2.0131 | 51.5932 | 103.8607 | 25.6291 | 3.1902 | 51.4689 | 164.1962 | 16.1334 | 15.2741 | 5.4114 | 268.0569 | 258.2841 | 195.2428 | 239.61 | 168.442
2012-10-30 16:25 | 28.8922 | 275.5996 | 211.7618 | 256.8539 | 184.825 | 78.6818 | 2.0127 | 51.5932 | 103.8435 | 25.6334 | 3.1898 | 51.4688 | 164.1726 | 16.1357 | 15.2868 | 5.4243 | 268.0161 | 260.3128 | 196.475 | 241.5671 | 169.5382
2012-10-30 16:25 | 28.8982 | 277.6565 | 212.9491 | 258.8473 | 185.971 | 78.861 | 2.0124 | 51.593 | 103.8248 | 25.6378 | 3.1893 | 51.4684 | 164.1466 | 16.138 | 15.252 | 5.4367 | 267.9714 | 262.4045 | 197.6971 | 243.5953 | 170.719
2012-10-30 16:25 | 28.8967 | 279.6189 | 214.2157 | 260.902 | 186.9493 | 79.0552 | 2.0121 | 51.5929 | 103.8082 | 25.6418 | 3.1888 | 51.4683 | 164.1247 | 16.1401 | 15.2791 | 5.4501 | 267.9329 | 264.3398 | 198.9366 | 245.6229 | 171.6702
2012-10-30 16:25 | 28.9019 | 281.5904 | 215.2723 | 262.6969 | 187.8405 | 79.2089 | 2.0118 | 51.5938 | 103.7962 | 25.6456 | 3.1885 | 51.469 | 164.1083 | 16.1422 | 15.325 | 5.4607 | 267.9045 | 266.2654 | 199.9473 | 247.3719 | 172.5155
2012-10-30 16:25 | 0 | 242.8145 | 180.3338 | 241.0232 | 161.0958 | 73.944 | 0.023474 | 0.74411 | 0.017467 | 31.6993 | 0.040472 | 0.75528 | 0.030568 | 18.6617 | -13.5647 | 5.0977 | 0.048035 | 256.3792 | 193.8985 | 254.5879 | 174.6605
2012-10-30 16:25 | 28.8899 | 267.7665 | 207.7238 | 269.5593 | 189.0403 | 73.5966 | 0.023474 | 0.75314 | 0.017679 | 32.084 | 0.040481 | 0.75505 | 0.030565 | 18.6519 | 15.3212 | 5.0737 | 0.048244 | 252.4453 | 192.4026 | 254.2381 | 173.7191
2012-10-30 16:25 | 28.9012 | 266.1317 | 207.1013 | 269.1352 | 188.6829 | 73.462 | 0.023484 | 0.75304 | 0.017684 | 32.066 | 0.040475 | 0.7549 | 0.030555 | 18.6511 | 15.3325 | 5.0645 | 0.048239 | 250.7992 | 191.7688 | 253.8027 | 173.3504
2012-10-30 16:25 | 28.9087 | 264.4804 | 206.3687 | 268.5905 | 188.1178 | 73.3158 | 0.023465 | 0.75281 | 0.017665 | 32.0821 | 0.040437 | 0.75444 | 0.030507 | 18.6572 | 15.3277 | 5.0544 | 0.048172 | 249.1527 | 191.041 | 253.2628 | 172.7901
2012-10-30 16:25 | 28.9019 | 262.6891 | 205.5275 | 267.8252 | 187.5229 | 73.1771 | 0.023496 | 0.75276 | 0.017687 | 32.038 | 0.040444 | 0.75458 | 0.030518 | 18.6575 | 15.325 | 5.0448 | 0.048205 | 247.3641 | 190.2025 | 252.5002 | 172.1979
2012-10-30 16:25 | 28.9132 | 260.8912 | 204.7792 | 266.8443 | 187.0873 | 73.0193 | 0.023496 | 0.75268 | 0.017685 | 32.0343 | 0.040444 | 0.7545 | 0.030515 | 18.6554 | 15.3363 | 5.034 | 0.0482 | 245.5549 | 189.4429 | 251.508 | 171.751
2012-10-30 16:25 | 28.9109 | 259.3674 | 203.9623 | 265.7716 | 186.5441 | 72.8488 | 0.02349 | 0.75269 | 0.017681 | 32.0431 | 0.040431 | 0.75451 | 0.030506 | 18.6617 | 15.3462 | 5.0222 | 0.048187 | 244.0212 | 188.6161 | 250.4254 | 171.1979
2012-10-30 16:25 | 28.9124 | 257.8231 | 203.1453 | 264.5971 | 185.7776 | 72.6968 | 0.023503 | 0.75255 | 0.017687 | 32.0192 | 0.040447 | 0.75448 | 0.030517 | 18.6536 | 15.3559 | 5.0117 | 0.048204 | 242.4672 | 187.7894 | 249.2412 | 170.4217
2012-10-30 16:25 | 28.9154 | 256.1857 | 202.2591 | 263.3807 | 184.9774 | 72.5789 | 0.023506 | 0.75255 | 0.017689 | 32.0151 | 0.040444 | 0.75435 | 0.030509 | 18.6518 | 15.3507 | 5.0036 | 0.048198 | 240.835 | 186.9084 | 248.03 | 169.6267
2012-10-30 16:25 | 28.9117 | 254.5602 | 201.3742 | 262.0635 | 184.4019 | 72.4183 | 0.0235 | 0.75246 | 0.017683 | 32.0196 | 0.040447 | 0.75431 | 0.03051 | 18.6493 | 15.3715 | 4.9925 | 0.048193 | 239.1887 | 186.0027 | 246.692 | 169.0304
2012-10-30 16:25 | 28.9147 | 252.9835 | 200.5119 | 260.7016 | 183.7275 | 72.246 | 0.023506 | 0.75245 | 0.017687 | 32.0108 | 0.040437 | 0.75428 | 0.030501 | 18.6532 | 15.3378 | 4.9806 | 0.048188 | 237.6457 | 185.1741 | 245.3638 | 168.3897
2012-10-30 16:25 | 28.9236 | 251.4237 | 199.6162 | 259.2903 | 183.0552 | 72.0981 | 0.023522 | 0.7523 | 0.017696 | 31.9829 | 0.040434 | 0.75414 | 0.030493 | 18.651 | 15.363 | 4.9704 | 0.048189 | 236.0607 | 184.2532 | 243.9273 | 167.6922
2012-10-30 16:25 | 28.9229 | 249.9984 | 198.8763 | 258.0366 | 182.5848 | 71.9634 | 0.023518 | 0.75227 | 0.017692 | 31.9871 | 0.040428 | 0.75415 | 0.030489 | 18.6542 | 15.452 | 4.9612 | 0.048181 | 234.5464 | 183.4243 | 242.5846 | 167.1328
2012-10-30 16:25 | 28.9184 | 248.6587 | 198.0619 | 256.6887 | 181.8603 | 71.8253 | 0.023512 | 0.75221 | 0.017686 | 31.9928 | 0.040437 | 0.75409 | 0.030493 | 18.6486 | 15.3904 | 4.9516 | 0.048179 | 233.2683 | 182.6715 | 241.2983 | 166.4699
2012-10-30 16:25 | 28.9206 | 247.1527 | 197.1757 | 255.214 | 181.217 | 71.6982 | 0.023522 | 0.75196 | 0.017687 | 31.9682 | 0.040406 | 0.75372 | 0.030455 | 18.6536 | 15.3356 | 4.9429 | 0.048142 | 231.8171 | 181.8401 | 239.8784 | 165.8814
2012-10-30 16:26 | 28.9266 | 245.8733 | 196.4271 | 253.8479 | 180.4056 | 71.5623 | 0.023506 | 0.7518 | 0.017672 | 31.9832 | 0.040434 | 0.75382 | 0.03048 | 18.6432 | 15.3049 | 4.9335 | 0.048152 | 230.5684 | 181.1222 | 238.543 | 165.1007
2012-10-30 16:26 | 28.9289 | 244.5114 | 195.5924 | 252.5675 | 179.88 | 71.43 | 0.023534 | 0.75223 | 0.017703 | 31.9635 | 0.040434 | 0.75402 | 0.030488 | 18.6482 | 15.3194 | 4.9244 | 0.048191 | 229.192 | 180.273 | 237.2481 | 164.5606
2012-10-30 16:26 | 28.9326 | 243.0718 | 194.775 | 251.0688 | 179.2309 | 71.2832 | 0.023531 | 0.75184 | 0.017692 | 31.951 | 0.040415 | 0.7537 | 0.030461 | 18.6491 | 15.3516 | 4.9143 | 0.048153 | 227.7202 | 179.4234 | 235.7172 | 163.8793
2012-10-30 16:26 | 28.9379 | 241.5823 | 193.9279 | 249.5475 | 178.396 | 71.1416 | 0.023534 | 0.75178 | 0.017692 | 31.9445 | 0.040422 | 0.75389 | 0.030474 | 18.6505 | 15.3936 | 4.9045 | 0.048166 | 226.1887 | 178.5343 | 234.1539 | 163.0024
2012-10-30 16:26 | 28.9394 | 240.2101 | 193.0458 | 248.0804 | 177.4014 | 71.022 | 0.02355 | 0.7521 | 0.017712 | 31.9363 | 0.040419 | 0.75395 | 0.030474 | 18.6533 | 15.3503 | 4.8963 | 0.048186 | 224.8598 | 177.6955 | 232.7301 | 162.0511
2012-10-30 16:26 | 28.9394 | 238.8439 | 192.0995 | 246.6073 | 176.5775 | 70.8775 | 0.023563 | 0.75203 | 0.01772 | 31.9156 | 0.040415 | 0.75389 | 0.030469 | 18.6537 | 15.3014 | 4.8863 | 0.048189 | 223.5425 | 176.7981 | 231.3059 | 161.2761
2012-10-30 16:26 | 28.9559 | 237.5212 | 191.2368 | 245.1774 | 175.8835 | 70.74 | 0.023506 | 0.75096 | 0.017652 | 31.9475 | 0.040409 | 0.75349 | 0.030448 | 18.6465 | 15.3464 | 4.8768 | 0.0481 | 222.1748 | 175.8904 | 229.831 | 160.5371
2012-10-30 16:26 | 28.9506 | 236.2705 | 190.4231 | 243.6625 | 175.2656 | 70.6134 | 0.023559 | 0.75182 | 0.017712 | 31.9124 | 0.0404 | 0.75372 | 0.03045 | 18.6564 | 15.3859 | 4.8681 | 0.048162 | 220.8846 | 175.0372 | 228.2766 | 159.8797
2012-10-30 16:26 | 28.9521 | 234.9792 | 189.5693 | 242.2361 | 174.5141 | 70.4782 | 0.023547 | 0.75152 | 0.017696 | 31.9158 | 0.040403 | 0.75354 | 0.030445 | 18.6507 | 15.3915 | 4.8588 | 0.048141 | 219.5877 | 174.1778 | 226.8446 | 159.1226
2012-10-30 16:26 | 28.9618 | 233.6879 | 188.7239 | 240.786 | 174.009 | 70.3383 | 0.023547 | 0.75177 | 0.017702 | 31.9262 | 0.040387 | 0.75323 | 0.030421 | 18.6502 | 15.3564 | 4.8491 | 0.048123 | 218.3315 | 173.3675 | 225.4296 | 158.6526
2012-10-30 16:26 | 28.9633 | 232.3056 | 187.8745 | 239.335 | 173.3126 | 70.2204 | 0.023559 | 0.75138 | 0.017702 | 31.8934 | 0.04039 | 0.75353 | 0.030435 | 18.6564 | 15.3171 | 4.841 | 0.048137 | 216.9885 | 172.5574 | 224.0179 | 157.9955
2012-10-30 16:26 | 28.9618 | 230.9041 | 186.9791 | 237.8256 | 172.5043 | 70.0956 | 0.023572 | 0.75185 | 0.017723 | 31.8961 | 0.040422 | 0.75359 | 0.030461 | 18.643 | 15.2912 | 4.8324 | 0.048184 | 215.6129 | 171.6879 | 222.5344 | 157.2131
2012-10-30 16:26 | 28.9611 | 229.6015 | 186.0888 | 236.4107 | 171.6387 | 69.9632 | 0.023531 | 0.75061 | 0.017663 | 31.8988 | 0.040381 | 0.75289 | 0.030403 | 18.6448 | 15.3027 | 4.8233 | 0.048066 | 214.2988 | 170.7861 | 221.108 | 156.336
2012-10-30 16:26 | 28.9686 | 228.4406 | 185.2461 | 235.1372 | 170.7972 | 69.8187 | 0.023572 | 0.75172 | 0.01772 | 31.8905 | 0.040415 | 0.75354 | 0.030454 | 18.6452 | 15.2613 | 4.8133 | 0.048174 | 213.1793 | 169.9848 | 219.8759 | 155.5359
2012-10-30 16:26 | 28.9686 | 227.0831 | 184.3766 | 233.5257 | 170.0108 | 69.6847 | 0.023572 | 0.75143 | 0.017713 | 31.8783 | 0.0404 | 0.75337 | 0.030436 | 18.6478 | 15.2857 | 4.8041 | 0.048149 | 211.7974 | 169.0909 | 218.24 | 154.7251
2012-10-30 16:26 | 28.9806 | 225.7728 | 183.5075 | 232.0791 | 169.2912 | 69.5587 | 0.023581 | 0.75174 | 0.017727 | 31.879 | 0.040415 | 0.75353 | 0.030454 | 18.6448 | 15.2773 | 4.7954 | 0.048181 | 210.4955 | 168.2302 | 216.8018 | 154.0139
2012-10-30 16:26 | 28.9746 | 224.6097 | 182.7422 | 230.8501 | 168.7173 | 69.4414 |  |  |  |  | 0.040315 | 0.75227 | 0.030328 | 18.6599 | 15.2387 | 4.7873 | 0.030328 | 209.371 | 167.5035 | 215.6114 | 153.4786
2012-10-30 16:26 | 0 | 187.8314 | 147.9778 | 192.4818 | 134.8169 | 68.6606 | 0.023572 | 0.75061 | 0.017693 | 31.8433 | 0.040368 | 0.75299 | 0.030397 | 18.6533 | -13.8582 | 4.7335 | 0.04809 | 201.6896 | 161.836 | 206.34 | 148.6751
2012-10-30 16:26 | 28.9701 | 215.8288 | 176.2256 | 220.3258 | 163.0926 | 68.6178 | 0.46909 | 11.8234 | 5.5462 | 25.2052 | 0.72843 | 11.7716 | 8.5747 | 16.1602 | 15.0956 | 4.7305 | 14.121 | 200.7332 | 161.13 | 205.2302 | 147.997
2012-10-30 16:27 | 28.9596 | 214.9638 | 175.977 | 219.1498 | 162.2756 | 69.6153 | 1.2923 | 32.5168 | 42.0216 | 25.1619 | 2.0011 | 32.1426 | 64.3212 | 16.0623 | 15.0688 | 4.7993 | 106.3428 | 199.895 | 160.9082 | 204.081 | 147.2068
2012-10-30 16:27 | 28.9656 | 215.6432 | 177.3182 | 218.1438 | 161.5646 | 71.3468 | 0.88127 | 22.1944 | 19.5594 | 25.1845 | 1.3814 | 22.2175 | 30.6921 | 16.083 | 15.0503 | 4.9186 | 50.2515 | 200.5929 | 162.2679 | 203.0935 | 146.5143
2012-10-30 16:27 | 0 | 184.8267 | 145.328 | 185.2402 | 130.7284 | 69.5656 | 1.2789 | 32.224 | 41.2103 | 25.1973 | 2.0021 | 32.1497 | 64.367 | 16.058 | -13.7522 | 4.7959 | 105.5774 | 198.5789 | 159.0802 | 198.9924 | 144.4806
2012-10-30 16:27 | 28.9603 | 215.2204 | 174.9865 | 212.838 | 158.667 | 70.5383 | 1.4216 | 35.8718 | 50.9936 | 25.2343 | 2.2322 | 35.8402 | 80.0011 | 16.0563 | 15.257 | 4.8629 | 130.9946 | 199.9634 | 159.7295 | 197.581 | 143.41
2012-10-30 16:27 | 28.9603 | 216.2601 | 175.5757 | 212.6135 | 158.4272 | 71.0856 | 1.511 | 38.1847 | 57.6963 | 25.2715 | 2.3663 | 37.9971 | 89.914 | 16.0574 | 15.2203 | 4.9006 | 147.6103 | 201.0398 | 160.3554 | 197.3932 | 143.2069
2012-10-30 16:27 | 28.9641 | 217.587 | 176.2899 | 212.5779 | 158.5354 | 71.6531 | 1.6776 | 42.5098 | 71.3148 | 25.3395 | 2.6391 | 42.3946 | 111.8848 | 16.0639 | 15.216 | 4.9398 | 183.1996 | 202.371 | 161.0739 | 197.3619 | 143.3194
2012-10-30 16:27 | 28.9641 | 219.485 | 177.3519 | 212.7551 | 158.8217 | 72.1362 | 1.7388 | 44.1133 | 76.7056 | 25.3695 | 2.744 | 44.0885 | 120.9788 | 16.0673 | 15.2364 | 4.9731 | 197.6844 | 204.2486 | 162.1155 | 197.5187 | 143.5853
2012-10-30 16:27 | 28.9611 | 221.62 | 178.4505 | 213.1341 | 159.1786 | 72.8003 | 1.818 | 46.2131 | 84.0175 | 25.4191 | 2.8686 | 46.1034 | 132.2544 | 16.0715 | 15.1518 | 5.0189 | 216.272 | 206.4682 | 163.2987 | 197.9823 | 144.0268
2012-10-30 16:27 | 28.9656 | 224.026 | 179.5961 | 213.6939 | 159.4772 | 73.0366 | 1.8169 | 46.1983 | 83.9369 | 25.4272 | 2.8684 | 46.0877 | 132.197 | 16.0675 | 15.1278 | 5.0351 | 216.1339 | 208.8982 | 164.4683 | 198.5661 | 144.3494
2012-10-30 16:27 | 28.9663 | 226.2702 | 180.5469 | 214.431 | 159.8805 | 73.2123 | 1.8158 | 46.1799 | 83.8524 | 25.4326 | 2.8677 | 46.0685 | 132.1127 | 16.0644 | 15.1326 | 5.0473 | 215.9651 | 211.1376 | 165.4143 | 199.2984 | 144.7479
2012-10-30 16:27 | 28.9559 | 228.5066 | 181.4166 | 215.4247 | 160.3618 | 73.41 | 1.8158 | 46.1924 | 83.8781 | 25.4385 | 2.869 | 46.0816 | 132.2079 | 16.0619 | 15.0977 | 5.0609 | 216.086 | 213.4089 | 166.3189 | 200.327 | 145.2641
2012-10-30 16:27 | 28.9618 | 230.5383 | 182.3145 | 216.616 | 161.092 | 73.607 | 1.8147 | 46.1723 | 83.7902 | 25.4431 | 2.8681 | 46.0618 | 132.1101 | 16.06 | 15.1566 | 5.0745 | 215.9003 | 215.3817 | 167.1579 | 201.4594 | 145.9354
2012-10-30 16:27 | 28.9618 | 232.713 | 183.203 | 217.8682 | 161.8437 | 73.7469 | 1.8142 | 46.1674 | 83.757 | 25.4478 | 2.8681 | 46.0563 | 132.0939 | 16.0581 | 15.124 | 5.0841 | 215.8509 | 217.589 | 168.079 | 202.7442 | 146.7197
2012-10-30 16:27 | 28.9618 | 234.8121 | 184.1387 | 219.1874 | 162.4639 | 73.9272 | 1.8135 | 46.1583 | 83.7086 | 25.4524 | 2.8678 | 46.0469 | 132.0522 | 16.0567 | 15.0873 | 5.0965 | 215.7608 | 219.7248 | 169.0514 | 204.1001 | 147.3766
2012-10-30 16:27 | 28.9521 | 236.661 | 185.1708 | 220.5755 | 163.1949 | 74.0763 | 1.8144 | 46.1891 | 83.8049 | 25.4571 | 2.8699 | 46.0777 | 132.2367 | 16.0557 | 15.049 | 5.1068 | 216.0416 | 221.612 | 170.1218 | 205.5265 | 148.1459
2012-10-30 16:27 | 0 | 202.0131 | 151.303 | 197.5085 | 135.8103 | 69.7581 | 0.024677 | 0.74534 | 0.018393 | 30.2038 | 0.041677 | 0.75532 | 0.03148 | 18.1232 | -13.9887 | 4.8091 | 0.049873 | 216.0018 | 165.2917 | 211.4972 | 149.799
2012-10-30 16:28 | 28.9236 | 227.1207 | 178.7373 | 226.5265 | 163.8181 | 69.4182 | 0.024673 | 0.7522 | 0.018559 | 30.4868 | 0.041624 | 0.75454 | 0.031407 | 18.1275 | 14.9186 | 4.7857 | 0.049966 | 212.2021 | 163.8187 | 211.6079 | 148.8995
2012-10-30 16:28 | 28.9184 | 225.6673 | 178.0561 | 226.1357 | 163.359 | 69.2714 | 0.024683 | 0.7524 | 0.018572 | 30.4826 | 0.041658 | 0.75484 | 0.031445 | 18.12 | 14.9215 | 4.7756 | 0.050017 | 210.7458 | 163.1346 | 211.2142 | 148.4375
2012-10-30 16:28 | 28.9102 | 224.2579 | 177.4308 | 225.6788 | 162.8039 | 69.1611 | 0.024689 | 0.75226 | 0.018572 | 30.4694 | 0.041646 | 0.75476 | 0.031433 | 18.1232 | 14.8848 | 4.768 | 0.050005 | 209.3731 | 162.546 | 210.794 | 147.9191
2012-10-30 16:28 | 28.9064 | 223.022 | 176.7943 | 225.1516 | 162.4046 | 69.027 | 0.024673 | 0.75223 | 0.01856 | 30.4879 | 0.041643 | 0.75461 | 0.031424 | 18.121 | 14.8932 | 4.7587 | 0.049984 | 208.1288 | 161.9011 | 210.2584 | 147.5114
2012-10-30 16:28 | 28.9109 | 221.6919 | 176.1117 | 224.4476 | 161.789 | 68.8923 | 0.02468 | 0.75198 | 0.018559 | 30.4694 | 0.04164 | 0.75448 | 0.031417 | 18.1192 | 14.8569 | 4.7494 | 0.049976 | 206.835 | 161.2548 | 209.5907 | 146.9321
2012-10-30 16:28 | 28.9079 | 220.3896 | 175.4102 | 223.7282 | 161.2822 | 68.7652 | 0.024677 | 0.75196 | 0.018556 | 30.4719 | 0.04163 | 0.75438 | 0.031405 | 18.1211 | 14.9029 | 4.7407 | 0.049961 | 205.4867 | 160.5073 | 208.8253 | 146.3793
2012-10-30 16:28 | 28.9094 | 219.1114 | 174.7523 | 222.9345 | 160.8576 | 68.6398 | 0.024677 | 0.75148 | 0.018544 | 30.4526 | 0.041624 | 0.75402 | 0.031385 | 18.1151 | 14.8717 | 4.732 | 0.049929 | 204.2397 | 159.8806 | 208.0628 | 145.9859
2012-10-30 16:28 | 28.9079 | 217.7868 | 174.0136 | 222.0472 | 160.2712 | 68.5092 | 0.02468 | 0.75177 | 0.018554 | 30.4606 | 0.041605 | 0.75416 | 0.031377 | 18.1268 | 14.8621 | 4.723 | 0.049931 | 202.9247 | 159.1515 | 207.1851 | 145.4091
2012-10-30 16:28 | 28.9012 | 216.5869 | 173.2583 | 221.0089 | 159.6915 | 68.4098 | 0.024661 | 0.751 | 0.01852 | 30.453 | 0.041614 | 0.75391 | 0.031373 | 18.1166 | 14.8595 | 4.7162 | 0.049893 | 201.7274 | 158.3988 | 206.1494 | 144.832
2012-10-30 16:28 | 28.9034 | 215.2541 | 172.5355 | 219.8496 | 159.1286 | 68.2624 | 0.024708 | 0.75152 | 0.018569 | 30.4161 | 0.041602 | 0.75404 | 0.031369 | 18.125 | 14.878 | 4.706 | 0.049938 | 200.3761 | 157.6575 | 204.9716 | 144.2506
2012-10-30 16:28 | 28.8982 | 213.8549 | 171.7589 | 218.7103 | 158.516 | 68.1046 | 0.024692 | 0.75152 | 0.018557 | 30.4358 | 0.041611 | 0.75404 | 0.031376 | 18.1211 | 14.8769 | 4.6951 | 0.049933 | 198.978 | 156.882 | 203.8334 | 143.6391
2012-10-30 16:28 | 28.8989 | 212.6466 | 170.9962 | 217.5769 | 157.747 | 67.9884 | 0.024708 | 0.75152 | 0.018569 | 30.4161 | 0.041589 | 0.75392 | 0.031355 | 18.1279 | 14.8898 | 4.6871 | 0.049924 | 197.7568 | 156.1064 | 202.6871 | 142.8572
2012-10-30 16:28 | 28.9004 | 211.4825 | 170.2852 | 216.4325 | 157.2149 | 67.8607 | 0.024689 | 0.75116 | 0.018545 | 30.4249 | 0.041605 | 0.75366 | 0.031356 | 18.1147 | 14.8791 | 4.6783 | 0.049901 | 196.6034 | 155.4061 | 201.5534 | 142.3358
2012-10-30 16:28 | 28.8847 | 210.1871 | 169.4521 | 215.2118 | 156.6543 | 67.7347 | 0.024695 | 0.75123 | 0.018552 | 30.4204 | 0.041596 | 0.7538 | 0.031355 | 18.122 | 14.8715 | 4.6696 | 0.049907 | 195.3156 | 154.5806 | 200.3403 | 141.7828
2012-10-30 16:28 | 28.8757 | 208.8867 | 168.6923 | 213.8639 | 155.7991 | 67.6082 | 0.024705 | 0.75103 | 0.018554 | 30.3999 | 0.041574 | 0.75341 | 0.031322 | 18.1222 | 14.838 | 4.6609 | 0.049876 | 194.0487 | 153.8543 | 199.0259 | 140.9611
2012-10-30 16:28 | 28.8735 | 207.8058 | 168.0134 | 212.5425 | 155.237 | 67.4966 | 0.02467 | 0.75068 | 0.018519 | 30.429 | 0.041574 | 0.75336 | 0.03132 | 18.1208 | 14.8562 | 4.6532 | 0.049839 | 192.9496 | 153.1572 | 197.6863 | 140.3808
2012-10-30 16:28 | 28.8712 | 206.7248 | 167.3035 | 211.4888 | 154.6362 | 67.3857 | 0.024673 | 0.75025 | 0.018511 | 30.4077 | 0.04157 | 0.75314 | 0.031308 | 18.1174 | 14.8825 | 4.6456 | 0.049819 | 191.8423 | 152.421 | 196.6063 | 139.7537
2012-10-30 16:28 | 28.8645 | 205.6081 | 166.6244 | 210.3874 | 154.1317 | 67.2591 | 0.024695 | 0.75099 | 0.018546 | 30.4105 | 0.041577 | 0.75344 | 0.031326 | 18.1216 | 14.8676 | 4.6368 | 0.049872 | 190.7405 | 151.7568 | 195.5198 | 139.2641
2012-10-30 16:28 | 28.8735 | 204.4994 | 165.973 | 209.2231 | 153.4925 | 67.147 | 0.024714 | 0.75078 | 0.018555 | 30.3789 | 0.041558 | 0.75288 | 0.031288 | 18.1163 | 14.9786 | 4.6291 | 0.049843 | 189.5208 | 150.9944 | 194.2445 | 138.5139
2012-10-30 16:28 | 28.8592 | 203.34 | 165.2208 | 207.8072 | 152.9575 | 67.0383 | 0.024718 | 0.75081 | 0.018559 | 30.3752 | 0.041577 | 0.7533 | 0.03132 | 18.1181 | 14.9847 | 4.6216 | 0.049879 | 188.3553 | 150.2361 | 192.8225 | 137.9728
2012-10-30 16:28 | 28.8615 | 202.2367 | 164.4545 | 206.5144 | 152.1309 | 66.9031 | 0.024711 | 0.75083 | 0.018554 | 30.3844 | 0.041583 | 0.75324 | 0.031322 | 18.1142 | 14.9584 | 4.6123 | 0.049876 | 187.2783 | 149.4961 | 191.556 | 137.1725
2012-10-30 16:28 | 28.8525 | 201.1419 | 163.6771 | 205.297 | 151.4244 | 66.7864 | 0.024692 | 0.75023 | 0.018525 | 30.3837 | 0.041539 | 0.75248 | 0.031257 | 18.1149 | 15.0025 | 4.6043 | 0.049782 | 186.1394 | 148.6746 | 190.2945 | 136.4219
2012-10-30 16:29 | 28.857 | 200.0688 | 162.933 |  | 150.8626 | 66.669 | 0.024708 | 0.75055 | 0.018545 | 30.3769 | 0.041551 | 0.75266 | 0.031274 | 18.1142 | 15.0314 | 4.5962 | 0.049819 | 185.0374 | 147.9016 | -15.0314 | 135.8312
2012-10-30 16:29 | 28.8615 | 198.8739 | 162.1814 | 19.1783 | 150.224 | 66.5673 | 0.024714 | 0.75068 | 0.018552 | 30.3748 | 0.04157 | 0.75317 | 0.031309 | 18.1181 | 14.9666 | 4.5891 | 0.049861 | 183.9073 | 147.2148 | 4.2117 | 135.2574
2012-10-30 16:29 | 28.854 | 197.8638 | 161.4182 | 18.9641 | 149.6352 | 66.4494 | 0.024724 | 0.7507 | 0.01856 | 30.3631 | 0.041574 | 0.75311 | 0.03131 | 18.115 | 14.9999 | 4.581 | 0.04987 | 182.8639 | 146.4183 | 3.9642 | 134.6353
2012-10-30 16:29 | 28.8495 | 196.8845 | 160.7705 | 18.9961 | 148.8688 | 66.3356 | 0.02468 | 0.74958 | 0.0185 | 30.3722 | 0.041523 | 0.7523 | 0.031238 | 18.1177 | 14.8975 | 4.5732 | 0.049738 | 181.987 | 145.873 | 4.0986 | 133.9713
2012-10-30 16:29 | 28.8502 | 195.887 | 160.078 | 19.1387 | 148.281 | 66.2379 | 0.024714 | 0.75036 | 0.018545 | 30.3619 | 0.041548 | 0.75266 | 0.031272 | 18.1155 | 14.9349 | 4.5664 | 0.049817 | 180.9521 | 145.1431 | 4.2038 | 133.3461
2012-10-30 16:29 | 28.8457 | 194.8256 | 159.3612 | 19.0085 | 147.692 | 66.1269 | 0.024721 | 0.75057 | 0.018555 | 30.3616 | 0.041564 | 0.75301 | 0.031298 | 18.1169 | 14.91 | 4.5588 | 0.049853 | 179.9156 | 144.4512 | 4.0985 | 132.782
2012-10-30 16:29 | 28.8397 | 193.8179 | 158.6625 | 18.889 | 147.048 | 66.0212 | 0.024718 | 0.75009 | 0.018541 | 30.3459 | 0.041533 | 0.75262 | 0.031259 | 18.121 | 14.8836 | 4.5515 | 0.0498 | 178.9343 | 143.7789 | 4.0054 | 132.1644
2012-10-30 16:29 | 28.8338 | 192.8891 | 157.9951 | 19.0979 | 146.5886 | 65.9414 | 0.024724 | 0.75042 | 0.018553 | 30.352 | 0.041555 | 0.75291 | 0.031287 | 18.1183 | 14.9022 | 4.546 | 0.04984 | 177.9869 | 143.0929 | 4.1957 | 131.6864
2012-10-30 16:29 | 28.827 | 191.8929 | 157.3464 | 18.9331 | 146.0517 | 65.8455 | 0.024736 | 0.75051 | 0.018565 | 30.3408 | 0.041555 | 0.75288 | 0.031286 | 18.1177 | 14.9035 | 4.5394 | 0.049851 | 176.9894 | 142.4429 | 4.0296 | 131.1482
2012-10-30 16:29 | 28.8293 | 191.0631 | 156.7496 | 19.0367 | 145.6277 | 65.7374 | 0.02474 | 0.75052 | 0.018568 | 30.3365 | 0.041577 | 0.75295 | 0.031305 | 18.1098 | 14.8569 | 4.5319 | 0.049873 | 176.2062 | 141.8927 | 4.1798 | 130.7708
2012-10-30 16:29 | 28.8105 | 187.212 | 154.0275 |  | 143.1685 | 65.3248 | 0.02473 | 0.75013 | 0.018551 | 30.333 | 0.041542 | 0.7526 | 0.031265 | 18.1167 | 14.7932 | 4.5035 | 0.049816 | 172.4188 | 139.2343 | -14.7932 | 128.3753
2012-10-30 16:29 | 28.8135 | 183.4736 | 151.3546 | 186.7132 | 140.9795 | 64.9324 | 0.024746 | 0.75008 | 0.018561 | 30.311 | 0.041507 | 0.75216 | 0.03122 | 18.1212 | 14.7514 | 4.4764 | 0.049781 | 168.7222 | 136.6032 | 171.9618 | 126.2281
2012-10-30 16:29 | 28.8045 | 182.6777 | 150.8987 | 186.0066 | 140.3425 | 64.8463 | 0.024746 | 0.7497 | 0.018552 | 30.2958 | 0.041542 | 0.75252 | 0.031261 | 18.1146 | 14.7628 | 4.4705 | 0.049813 | 167.9149 | 136.1359 | 171.2438 | 125.5797
2012-10-30 16:29 | 28.7978 | 181.8883 | 150.2604 | 185.1066 | 139.7234 | 64.7272 | 0.024708 | 0.74896 | 0.018505 | 30.3126 | 0.041526 | 0.75216 | 0.031234 | 18.1129 | 14.7153 | 4.4623 | 0.049739 | 167.173 | 135.5451 | 170.3913 | 125.0081
2012-10-30 16:29 | 28.8045 | 181.0461 | 149.5588 | 184.0989 | 19.0565 | 64.6313 | 0.024733 | 0.74958 | 0.018539 | 30.3071 | 0.041545 | 0.75255 | 0.031265 | 18.114 | 14.7709 | 4.4557 | 0.049804 | 166.2752 | 134.7879 | 169.328 | 4.2856
2012-10-30 16:29 | 28.8053 | 180.2412 | 149.0401 | 183.2853 | 19.1626 | 64.585 | 0.024733 | 0.74977 | 0.018544 | 30.3147 | 0.041517 | 0.75206 | 0.031223 | 18.1144 | 14.7554 | 4.4525 | 0.049767 | 165.4858 | 134.2847 | 168.5299 | 4.4072
2012-10-30 16:29 | 28.806 | 179.5343 | 148.4906 | 182.5109 | 18.8431 | 64.4804 | 0.024755 | 0.75011 | 0.018569 | 30.3012 | 0.041545 | 0.75258 | 0.031266 | 18.1147 | 14.7276 | 4.4453 | 0.049835 | 164.8067 | 133.763 | 167.7833 | 4.1155
2012-10-30 16:29 | 28.8008 | 178.7747 | 147.9009 | 181.7466 | 18.992 | 64.3978 | 0.024762 | 0.75006 | 0.018573 | 30.2908 | 0.041542 | 0.75253 | 0.031262 | 18.115 | 14.6938 | 4.4396 | 0.049835 | 164.0809 | 133.2071 | 167.0528 | 4.2982
2012-10-30 16:29 | 28.8045 | 178.0321 | 147.3894 | 181.1243 | 19.0686 | 64.3042 | 0.024746 | 0.7492 | 0.01854 | 30.2754 | 0.041485 | 0.75171 | 0.031185 | 18.12 | 14.7261 | 4.4331 | 0.049725 | 163.306 | 132.6633 | 166.3982 | 4.3425
2012-10-30 16:30 | 28.8023 | 177.1393 | 146.8299 | 180.3595 | 18.884 | 64.2169 | 0.024752 | 0.74966 | 0.018556 | 30.2868 | 0.041514 | 0.75204 | 0.03122 | 18.1154 | 14.7402 | 4.4271 | 0.049776 | 162.3991 | 132.0897 | 165.6193 | 4.1438
2012-10-30 16:30 | 28.7963 | 176.3329 | 146.1977 | 179.5285 | 18.9388 | 64.1129 | 0.024718 | 0.74873 | 0.018507 | 30.291 | 0.041482 | 0.75159 | 0.031178 | 18.1186 | 14.6771 | 4.4199 | 0.049685 | 161.6558 | 131.5206 | 164.8514 | 4.2617
2012-10-30 16:30 | 28.7963 | 175.6615 | 145.6792 | 178.7976 | 19.1253 | 64.0464 | 0.024758 | 0.7496 | 0.018559 | 30.2771 | 0.041533 | 0.7522 | 0.031241 | 18.1109 | 14.6771 | 4.4154 | 0.0498 | 160.9844 | 131.0021 | 164.1205 | 4.4482
2012-10-30 16:30 | 28.7971 | 174.9833 | 145.1847 | 178.0677 | 18.8545 | 63.9626 | 0.024755 | 0.7495 | 0.018554 | 30.2767 | 0.041463 | 0.75129 | 0.031151 | 18.1195 | 14.6942 | 4.4096 | 0.049705 | 160.2891 | 130.4905 | 163.3735 | 4.1603
2012-10-30 16:30 | 28.7881 | 173.3401 | 143.94 | 176.3871 | 18.9833 | 63.7811 | 0.024768 | 0.74967 | 0.018568 | 30.2678 | 0.041517 | 0.75203 | 0.031222 | 18.1137 | 14.6648 | 4.3971 | 0.04979 | 158.6753 | 129.2752 | 161.7223 | 4.3185
2012-10-30 16:30 | 28.7768 | 171.8245 | 142.8242 | 174.7718 | 18.7977 | 63.5945 | 0.024774 | 0.7493 | 0.018563 | 30.2453 | 0.041529 | 0.75206 | 0.031232 | 18.1092 | 14.7065 | 4.3842 | 0.049795 | 157.118 | 128.1177 | 160.0653 | 4.0912
2012-10-30 16:30 | 28.7851 | 171.1706 | 142.2997 | 174.0624 | 19.0168 | 63.5297 | 0.024774 | 0.74986 | 0.018577 | 30.268 | 0.041542 | 0.7522 | 0.031248 | 18.107 | 14.6903 | 4.3797 | 0.049825 | 156.4803 | 127.6094 | 159.3721 | 4.3265
2012-10-30 16:30 | 28.7836 | 170.3941 | 141.7962 | 173.32 | 19.1328 | 63.4459 | 0.024736 | 0.74873 | 0.018521 | 30.269 | 0.041473 | 0.75136 | 0.031161 | 18.1169 | 14.7215 | 4.374 | 0.049682 | 155.6726 | 127.0747 | 158.5985 | 4.4113
2012-10-30 16:30 | 28.7836 | 169.7168 | 141.2483 | 172.5989 |  | 63.3569 | 0.024746 | 0.7487 | 0.018527 | 30.2556 | 0.04146 | 0.75106 | 0.031139 | 18.1153 | 14.7296 | 4.3678 | 0.049666 | 154.9872 | 126.5187 | 157.8693 | -14.7296
2012-10-30 16:30 | 28.7746 | 169.0074 | 140.7721 | 171.8651 | 131.3723 | 63.2835 | 0.02478 | 0.74931 | 0.018568 | 30.2385 | 0.04152 | 0.75178 | 0.031214 | 18.1065 | 14.6594 | 4.3628 | 0.049782 | 154.348 | 126.1127 | 157.2057 | 116.7129
2012-10-30 16:30 | 28.7746 | 168.2606 | 140.2016 | 171.1445 | 130.82 | 63.2246 | 0.024771 | 0.74905 | 0.018555 | 30.239 | 0.041476 | 0.75138 | 0.031164 | 18.1159 | 14.6309 | 4.3587 | 0.049719 | 153.6297 | 125.5707 | 156.5136 | 116.1891
2012-10-30 16:30 | 28.7716 | 167.616 | 139.6628 | 170.4367 | 130.2914 | 63.1437 | 0.024784 | 0.74963 | 0.018579 | 30.2465 | 0.041504 | 0.75198 | 0.03121 | 18.1183 | 14.6442 | 4.3531 | 0.049789 | 152.9718 | 125.0186 | 155.7925 | 115.6472
2012-10-30 16:30 | 28.7776 | 166.9534 | 139.1445 | 169.7147 | 129.8214 | 63.0535 | 0.024727 | 0.74833 | 0.018504 | 30.2636 | 0.041407 | 0.75016 | 0.031062 | 18.1168 | 14.691 | 4.3469 | 0.049566 | 152.2624 | 124.4535 | 155.0237 | 115.1304
2012-10-30 16:30 | 28.7753 | 166.3371 | 138.6679 | 169.0976 | 129.4042 | 62.9749 | 0.024774 | 0.74906 | 0.018557 | 30.2359 | 0.041504 | 0.75168 | 0.031198 | 18.111 | 14.6805 | 4.3415 | 0.049755 | 151.6566 | 123.9874 | 154.4171 | 114.7237
2012-10-30 16:30 | 28.7701 | 165.6443 | 138.1349 | 168.3338 | 128.8167 | 62.8992 | 0.024774 | 0.7492 | 0.018561 | 30.2412 | 0.041501 | 0.75158 | 0.031191 | 18.1099 | 14.6509 | 4.3363 | 0.049752 | 150.9934 | 123.484 | 153.6829 | 114.1658
2012-10-30 16:30 | 28.7618 | 164.8982 | 137.5073 | 167.4662 | 128.135 | 62.8148 | 0.02479 | 0.74947 | 0.018579 | 30.2328 | 0.041473 | 0.75129 | 0.031158 | 18.1152 | 14.5936 | 4.3305 | 0.049737 | 150.3046 | 122.9137 | 152.8726 | 113.5414
2012-10-30 16:30 | 28.7663 | 164.2506 | 136.9768 | 166.8604 | 127.4891 | 62.7495 | 0.024765 | 0.74857 | 0.018538 | 30.2271 | 0.041476 | 0.75145 | 0.031167 | 18.1177 | 14.6185 | 4.326 | 0.049705 | 149.6321 | 122.3583 | 152.2419 | 112.8706
2012-10-30 16:30 | 28.7618 | 163.5903 | 136.445 | 166.1837 | 127.2184 | 62.6721 | 0.024752 | 0.74816 | 0.018518 | 30.226 | 0.041463 | 0.75107 | 0.031142 | 18.1143 | 14.5895 | 4.3206 | 0.04966 | 149.0008 | 121.8555 | 151.5942 | 112.6289
2012-10-30 16:30 | 28.7686 | 162.9881 | 136.0049 | 165.503 | 126.8526 | 62.5987 | 0.024774 | 0.74869 | 0.018548 | 30.2208 | 0.041517 | 0.75178 | 0.031212 | 18.1078 | 14.5718 | 4.3156 | 0.04976 | 148.4163 | 121.4331 | 150.9312 | 112.2808
2012-10-30 16:30 | 28.7693 | 162.3567 | 135.5282 | 164.9407 | 126.5909 | 62.5363 | 0.024802 | 0.74956 | 0.01859 | 30.2216 | 0.041489 | 0.75154 | 0.03118 | 18.1141 | 14.6215 | 4.3113 | 0.04977 | 147.7352 | 120.9067 | 150.3192 | 111.9694
2012-10-30 16:30 | 28.7746 | 161.7108 | 135.064 | 164.2278 | 126.1363 | 62.4588 | 0.024796 | 0.74935 | 0.018581 | 30.2208 | 0.04152 | 0.75191 | 0.031219 | 18.1096 | 14.5942 | 4.3059 | 0.0498 | 147.1166 | 120.4698 | 149.6336 | 111.5421
2012-10-30 16:30 | 28.7701 | 161.1018 | 134.5556 | 163.5171 | 125.6834 | 62.3722 | 0.02479 | 0.74927 | 0.018574 | 30.2246 | 0.041511 | 0.75184 | 0.03121 | 18.1118 | 14.6427 | 4.2999 | 0.049784 | 146.4591 | 119.9129 | 148.8744 | 111.0407
2012-10-30 16:30 | 28.7731 | 160.4889 | 134.1351 | 162.9267 | 125.2266 | 62.3034 | 0.024758 | 0.74852 | 0.018532 | 30.2333 | 0.041467 | 0.75122 | 0.031151 | 18.116 | 14.6375 | 4.2952 | 0.049683 | 145.8514 | 119.4976 | 148.2892 | 110.5891
2012-10-30 16:30 | 28.7738 | 159.839 | 133.6398 | 162.3961 | 124.7867 | 62.2537 | 0.024784 | 0.74839 | 0.018548 | 30.1963 | 0.041444 | 0.75063 | 0.031109 | 18.1118 | 14.6056 | 4.2918 | 0.049657 | 145.2334 | 119.0342 | 147.7905 | 110.1811
2012-10-30 16:30 | 28.7768 | 159.3194 | 133.1279 | 161.7789 | 124.3339 | 62.1601 | 0.02479 | 0.7492 | 0.018573 | 30.2217 | 0.041504 | 0.75157 | 0.031193 | 18.1083 | 14.58 | 4.2853 | 0.049766 | 144.7394 | 118.5479 | 147.1989 | 109.7539
2012-10-30 16:30 | 28.7716 | 158.695 | 132.6728 | 161.2001 | 123.8692 | 62.1046 | 0.024793 | 0.74855 | 0.018559 | 30.1918 | 0.041448 | 0.75086 | 0.031122 | 18.1156 | 14.6442 | 4.2815 | 0.049681 | 144.0508 | 118.0286 | 146.5559 | 109.225
2012-10-30 16:31 | 28.7738 | 158.1092 | 132.2328 | 160.6019 | 123.5639 | 62.037 | 0.024762 | 0.74798 | 0.018522 | 30.2068 | 0.041457 | 0.75084 | 0.031128 | 18.1113 | 14.6382 | 4.2768 | 0.04965 | 143.471 | 117.5946 | 145.9637 | 108.9257
2012-10-30 16:31 | 28.7761 | 157.4579 | 131.7168 | 159.8761 | 123.0916 | 61.941 | 0.024806 | 0.74904 | 0.018581 | 30.1958 | 0.041504 | 0.75146 | 0.031189 | 18.1058 | 14.6732 | 4.2702 | 0.04977 | 142.7847 | 117.0436 | 145.2029 | 108.4184
2012-10-30 16:31 | 28.7716 | 156.7112 | 131.1712 | 159.1052 | 122.4078 | 61.8642 | 0.024802 | 0.74885 | 0.018573 | 30.193 | 0.041479 | 0.75109 | 0.031154 | 18.1077 | 14.6605 | 4.2649 | 0.049727 | 142.0507 | 116.5107 | 144.4447 | 107.7473
2012-10-30 16:31 | 28.7746 | 156.0959 | 130.6943 | 158.4272 | 121.9744 | 61.8046 | 0.024755 | 0.74775 | 0.018511 | 30.206 | 0.041501 | 0.75141 | 0.031184 | 18.1057 | 14.7125 | 4.2608 | 0.049695 | 141.3834 | 115.9818 | 143.7147 | 107.2619
2012-10-30 16:31 | 28.7701 | 155.5081 | 130.1604 | 157.7651 | 121.4501 | 61.733 | 0.024787 | 0.74862 | 0.018556 | 30.202 | 0.041482 | 0.75132 | 0.031166 | 18.1119 | 14.6917 | 4.2559 | 0.049722 | 140.8164 | 115.4687 | 143.0734 | 106.7584
2012-10-30 16:31 | 28.7671 | 154.9607 | 129.6777 | 157.2053 | 121.1626 | 61.6434 | 0.024787 | 0.74869 | 0.018558 | 30.2049 | 0.041448 | 0.75093 | 0.031125 | 18.1174 | 14.7213 | 4.2497 | 0.049683 | 140.2394 | 114.9564 | 142.484 | 106.4413
2012-10-30 16:31 |  | 121.1585 | 97.0463 | 0 | 0 | 0 | 0.023616 | 0.67775 | 0.016006 | 28.6987 | 0.040173 | 0.74743 | 0.030027 | 18.6054 | 0 | 0 | 0.046033 | 121.1585 | 97.0463 | 0 | 0
2012-10-30 16:31 | 28.7334 | 149.5916 | 125.4833 | 151.7446 | 117.4217 | 61.0614 | 0.023622 | 0.74554 | 0.017611 | 31.5613 | 0.04023 | 0.74847 | 0.030111 | 18.6048 | 14.659 | 4.2096 | 0.047722 | 134.9326 | 110.8243 | 137.0856 | 102.7627
2012-10-30 16:31 | 28.7177 | 149.0948 | 125.1142 | 151.1817 | 116.8908 | 60.9776 | 0.023666 | 0.74645 | 0.017665 | 31.541 | 0.040261 | 0.74888 | 0.030151 | 18.6006 | 14.7127 | 4.2038 | 0.047816 | 134.3821 | 110.4015 | 136.469 | 102.1781
2012-10-30 16:31 | 28.7206 | 148.5476 | 124.6537 | 150.576 | 116.549 | 60.9244 | 0.023666 | 0.74696 | 0.017677 | 31.5624 | 0.040305 | 0.74933 | 0.030202 | 18.5914 | 14.6993 | 4.2001 | 0.047879 | 133.8483 | 109.9544 | 135.8767 | 101.8497
2012-10-30 16:31 | 28.7162 | 147.9663 | 124.2353 | 150.0595 | 116.181 | 60.8499 | 0.023676 | 0.74674 | 0.01768 | 31.5399 | 0.040245 | 0.74863 | 0.030129 | 18.6019 | 14.7234 | 4.195 | 0.047809 | 133.2429 | 109.5119 | 135.3361 | 101.4576
2012-10-30 16:31 | 28.7169 | 147.4556 | 123.7917 | 149.4673 | 115.8295 | 60.7985 | 0.023676 | 0.7469 | 0.017684 | 31.5466 | 0.040283 | 0.74905 | 0.030174 | 18.5947 | 14.7323 | 4.1914 | 0.047858 | 132.7233 | 109.0594 | 134.735 | 101.0972
2012-10-30 16:31 | 28.7229 | 146.9158 | 123.308 | 148.9538 | 115.1691 | 60.732 | 0.023638 | 0.7458 | 0.017629 | 31.5509 | 0.040289 | 0.74917 | 0.030183 | 18.5948 | 14.7627 | 4.1869 | 0.047812 | 132.1531 | 108.5453 | 134.1911 | 100.4064
2012-10-30 16:31 | 28.7154 | 146.2705 | 122.8032 | 148.377 | 114.6278 | 60.6488 | 0.023688 | 0.747 | 0.017695 | 31.5349 | 0.040289 | 0.7493 | 0.030188 | 18.598 | 14.6981 | 4.1811 | 0.047883 | 131.5724 | 108.1051 | 133.6789 | 99.9297
2012-10-30 16:31 | 28.7072 | 145.7399 | 122.3168 | 147.6958 | 114.2979 | 60.5927 | 0.023666 | 0.74655 | 0.017668 | 31.5453 | 0.040299 | 0.74941 | 0.030201 | 18.5963 | 14.6899 | 4.1773 | 0.047869 | 131.05 | 107.6269 | 133.0059 | 99.608
2012-10-30 16:31 | 28.7192 | 145.1797 | 121.8697 | 147.1427 | 113.78 | 60.5222 | 0.023644 | 0.74576 | 0.017633 | 31.5411 | 0.040261 | 0.7487 | 0.030144 | 18.5963 | 14.6938 | 4.1724 | 0.047777 | 130.4859 | 107.1759 | 132.4489 | 99.0862
2012-10-30 16:31 | 28.7192 | 144.6999 | 121.4258 | 146.6047 | 113.4015 | 60.4852 | 0.023679 | 0.74652 | 0.017677 | 31.5268 | 0.040233 | 0.74831 | 0.030107 | 18.5995 | 14.6693 | 4.1698 | 0.047784 | 130.0306 | 106.7565 | 131.9354 | 98.7322
2012-10-30 16:31 | 28.7094 | 144.2334 | 121.0443 | 146.1454 | 113.2517 | 60.3951 | 0.023698 | 0.74732 | 0.01771 | 31.535 | 0.040318 | 0.74986 | 0.030233 | 18.5986 | 14.635 | 4.1636 | 0.047943 | 129.5984 | 106.4093 | 131.5104 | 98.6167
2012-10-30 16:31 | 28.7094 | 143.6158 | 120.6612 | 145.5501 | 113.0171 | 60.3274 | 0.023701 | 0.747 | 0.017705 | 31.5176 | 0.040286 | 0.74925 | 0.030184 | 18.5983 | 14.6799 | 4.159 | 0.047889 | 128.9359 | 105.9813 | 130.8702 | 98.3372
2012-10-30 16:32 | 28.7072 | 143.0767 | 120.1926 | 144.9182 | 112.4624 | 60.2448 | 0.023673 | 0.74635 | 0.017668 | 31.5274 | 0.040296 | 0.7492 | 0.03019 | 18.5923 | 14.6369 | 4.1533 | 0.047858 | 128.4398 | 105.5557 | 130.2813 | 97.8255
2012-10-30 16:32 | 28.7019 | 142.4963 | 119.664 | 144.3449 | 111.9085 | 60.202 | 0.023695 | 0.74688 | 0.017697 | 31.5207 | 0.040271 | 0.74894 | 0.03016 | 18.5974 | 14.6235 | 4.1503 | 0.047857 | 127.8728 | 105.0405 | 129.7214 | 97.285
2012-10-30 16:32 | 28.6929 | 142.0427 | 119.2835 | 143.887 | 111.5855 | 60.1367 | 0.023622 | 0.74522 | 0.017604 | 31.5478 | 0.040227 | 0.74827 | 0.030101 | 18.6012 | 14.5492 | 4.1458 | 0.047705 | 127.4935 | 104.7343 | 129.3378 | 97.0363
2012-10-30 16:32 | 28.7027 | 141.5888 | 118.9256 | 143.4672 | 111.2019 | 60.0789 | 0.023613 | 0.74512 | 0.017595 | 31.5555 | 0.040261 | 0.74881 | 0.030148 | 18.5988 | 14.559 | 4.1418 | 0.047743 | 127.0298 | 104.3666 | 128.9082 | 96.6429
2012-10-30 16:32 | 28.6832 | 141.0024 | 118.4362 | 142.9031 | 110.8419 | 60.0165 | 0.023651 | 0.74606 | 0.017645 | 31.5445 | 0.040258 | 0.74876 | 0.030144 | 18.5991 | 14.5721 | 4.1375 | 0.047789 | 126.4303 | 103.8641 | 128.331 | 96.2698
2012-10-30 16:32 | 28.6847 | 140.5979 | 118.1042 | 142.4331 | 110.3668 | 59.9541 | 0.067097 | 1.813 | 0.12165 | 27.0204 | 0.11739 | 1.983 | 0.23278 | 16.893 | 14.5328 | 4.1332 | 0.35443 | 126.0651 | 103.5714 | 127.9003 | 95.834
2012-10-30 16:32 | 28.6862 | 140.1668 | 117.7723 | 141.8827 | 110.1338 | 60.0703 | 0.76177 | 18.951 | 14.4362 | 24.8778 | 1.2097 | 19.4422 | 23.5184 | 16.0726 | 14.571 | 4.1412 | 37.9546 | 125.5958 | 103.2013 | 127.3117 | 95.5628
2012-10-30 16:32 | 28.6847 | 140.1691 | 118.2103 | 141.4827 | 109.7428 | 61.4706 | 1.6334 | 40.9383 | 66.8688 | 25.0632 | 2.5787 | 41.3062 | 106.5177 | 16.018 | 14.541 | 4.2378 | 173.3866 | 125.6281 | 103.6693 | 126.9417 | 95.2018
2012-10-30 16:32 | 28.6899 | 142.6721 | 121.442 | 141.0742 | 109.5513 | 64.6076 | 2.0349 | 51.5152 | 104.8259 | 25.3164 | 3.2057 | 51.4237 | 164.8505 | 16.0412 | 14.5666 | 4.454 | 269.6764 | 128.1055 | 106.8754 | 126.5076 | 94.9847
2012-10-30 16:32 | 28.675 | 147.2292 | 124.5435 | 141.0401 | 109.7671 | 65.7264 | 2.0332 | 51.5373 | 104.7862 | 25.3478 | 3.2047 | 51.414 | 164.7642 | 16.0435 | 14.5843 | 4.5312 | 269.5504 | 132.6449 | 109.9592 | 126.4558 | 95.1828
2012-10-30 16:32 | 28.6705 | 151.7356 | 126.4427 | 141.5643 | 110.4737 | 66.1125 | 2.032 | 51.5303 | 104.7107 | 25.3592 | 3.2041 | 51.4068 | 164.7123 | 16.0441 | 14.5921 | 4.5578 | 269.423 | 137.1435 | 111.8506 | 126.9722 | 95.8816
2012-10-30 16:32 | 28.6675 | 156.2554 | 128.2918 | 142.3392 | 111.3672 | 66.4546 | 2.0311 | 51.5261 | 104.6527 | 25.3691 | 3.2037 | 51.4026 | 164.6795 | 16.0447 | 14.536 | 4.5814 | 269.3322 | 141.7194 | 113.7558 | 127.8032 | 96.8312
2012-10-30 16:32 | 28.6645 | 160.1827 | 129.9558 | 143.4251 | 112.3251 | 66.7742 | 2.0303 | 51.5243 | 104.6083 | 25.3781 | 3.2035 | 51.4006 | 164.6603 | 16.0453 | 14.5657 | 4.6034 | 269.2686 | 145.617 | 115.3901 | 128.8594 | 97.7594
2012-10-30 16:32 | 28.672 | 163.26 | 131.2926 | 144.8216 | 113.377 | 67.069 | 2.0296 | 51.5231 | 104.5715 | 25.3858 | 3.2032 | 51.3993 | 164.6443 | 16.0461 | 14.5446 | 4.6237 | 269.2158 | 148.7154 | 116.748 | 130.277 | 98.8324
2012-10-30 16:32 | 28.6735 | 166.2897 | 132.6662 | 146.1594 | 114.4799 | 67.3342 | 2.029 | 51.5238 | 104.5404 | 25.394 | 3.2031 | 51.3998 | 164.6401 | 16.0467 | 14.5339 | 4.642 | 269.1806 | 151.7558 | 118.1323 | 131.6255 | 99.946
2012-10-30 16:32 | 28.6742 | 170.0667 | 134.3681 | 148.1012 | 115.7376 | 67.6376 | 2.0283 | 51.5251 | 104.5104 | 25.4026 | 3.203 | 51.401 | 164.6386 | 16.0477 | 14.5305 | 4.6629 | 269.149 | 155.5362 | 119.8376 | 133.5707 | 101.2071
2012-10-30 16:32 | 28.6742 | 173.2067 | 135.8647 | 150.2639 | 117.0102 | 67.9122 | 2.0279 | 51.5262 | 104.4884 | 25.409 | 3.2029 | 51.4021 | 164.6367 | 16.0485 | 14.559 | 4.6819 | 269.1251 | 158.6477 | 121.3057 | 135.7049 | 102.4512
2012-10-30 16:32 | 28.669 | 175.6903 | 137.3342 | 151.9461 | 118.0696 | 68.137 | 2.0274 | 51.5271 | 104.4685 | 25.4148 | 3.2028 | 51.4029 | 164.6337 | 16.0493 | 14.5294 | 4.6974 | 269.1022 | 161.1609 | 122.8048 | 137.4167 | 103.5402
2012-10-30 16:32 | 28.6555 | 178.434 | 138.7432 | 153.7066 | 119.0945 | 68.3884 | 2.027 | 51.5282 | 104.4459 | 25.4213 | 3.2027 | 51.4039 | 164.63 | 16.0503 | 14.5199 | 4.7147 | 269.0759 | 163.9141 | 124.2233 | 139.1867 | 104.5746
2012-10-30 16:32 | 28.6705 | 181.1312 | 140.1473 | 155.791 | 120.3151 | 68.6271 | 2.0265 | 51.5298 | 104.4275 | 25.4274 | 3.2026 | 51.4054 | 164.6298 | 16.0513 | 14.5513 | 4.7312 | 269.0573 | 166.5799 | 125.596 | 141.2397 | 105.7638
2012-10-30 16:32 | 28.6682 | 184.2759 | 141.8838 | 158.6651 | 121.7883 | 68.9033 | 2.026 | 51.5318 | 104.4058 | 25.4346 | 3.2024 | 51.4072 | 164.6277 | 16.0526 | 14.5408 | 4.7502 | 269.0334 | 169.7351 | 127.343 | 144.1243 | 107.2475
2012-10-30 16:32 | 28.6742 | 187.5932 | 143.7187 | 161.5871 | 123.1264 | 69.2217 | 2.0255 | 51.5327 | 104.3797 | 25.4419 | 3.2022 | 51.408 | 164.6187 | 16.054 | 14.5019 | 4.7721 | 268.9985 | 173.0913 | 129.2168 | 147.0852 | 108.6245
2012-10-30 16:32 | 28.6622 | 190.3858 | 145.3378 | 163.933 | 124.5193 | 69.4674 | 2.0251 | 51.5344 | 104.3624 | 25.4478 | 3.2021 | 51.4098 | 164.6187 | 16.0551 | 14.5022 | 4.7891 | 268.9811 | 175.8836 | 130.8356 | 149.4308 | 110.0171
2012-10-30 16:32 | 28.675 | 192.8795 | 146.7334 | 166.5165 | 125.8277 | 69.6887 | 2.0247 | 51.536 | 104.3444 | 25.4538 | 3.2019 | 51.4113 | 164.6156 | 16.0563 | 14.5721 | 4.8043 | 268.9601 | 178.3074 | 132.1613 | 151.9444 | 111.2556
2012-10-30 16:32 | 28.6765 | 195.3203 | 148.1765 | 169.1778 | 127.1673 | 69.9586 | 2.0243 | 51.5375 | 104.3276 | 25.4593 | 3.2018 | 51.4126 | 164.6112 | 16.0576 | 14.6021 | 4.8229 | 268.9388 | 180.7182 | 133.5744 | 154.5757 | 112.5652
2012-10-30 16:32 | 28.672 | 198.0081 | 149.7843 | 171.8093 | 128.5702 | 70.1938 | 2.0239 | 51.5384 | 104.3089 | 25.4648 | 3.2016 | 51.4135 | 164.6034 | 16.0589 | 14.5528 | 4.8392 | 268.9123 | 183.4553 | 135.2315 | 157.2565 | 114.0174
2012-10-30 16:32 | 28.6675 | 200.5793 | 151.3319 | 174.21 | 129.8674 | 70.4256 | 2.0235 | 51.5396 | 104.2917 | 25.4702 | 3.2013 | 51.4146 | 164.5957 | 16.0603 | 14.5891 | 4.8551 | 268.8874 | 185.9902 | 136.7428 | 159.6209 | 115.2783
2012-10-30 16:32 | 28.681 | 203.1697 | 152.8179 | 176.69 | 131.153 | 70.6608 | 2.0232 | 51.5408 | 104.2752 | 25.4754 | 3.2012 | 51.4158 | 164.5902 | 16.0616 | 14.5577 | 4.8714 | 268.8654 | 188.612 | 138.2602 | 162.1323 | 116.5953
2012-10-30 16:32 | 28.678 | 205.7129 | 154.3658 | 179.4297 | 132.3505 | 70.9029 | 2.0228 | 51.5418 | 104.2587 | 25.4804 | 3.2009 | 51.4168 | 164.5806 | 16.0632 | 14.5628 | 4.888 | 268.8393 | 191.1501 | 139.803 | 164.8669 | 117.7877
2012-10-30 16:33 | 28.6757 | 207.9741 | 155.9005 | 182.1458 | 133.6295 | 71.1023 | 2.0224 | 51.5428 | 104.2422 | 25.4855 | 3.2007 | 51.4177 | 164.5723 | 16.0645 | 14.5401 | 4.9018 | 268.8145 | 193.434 | 141.3604 | 167.6057 | 119.0894
2012-10-30 16:33 | 28.678 | 210.4608 | 157.4836 | 185.1788 | 135.0628 | 71.356 | 2.0221 | 51.5439 | 104.2247 | 25.4908 | 3.2004 | 51.4188 | 164.5633 | 16.0661 | 14.5343 | 4.9193 | 268.7879 | 195.9265 | 142.9493 | 170.6445 | 120.5285
2012-10-30 16:33 | 28.678 | 212.6896 | 158.8529 | 187.5931 | 136.1778 | 71.5751 | 2.0217 | 51.5448 | 104.2085 | 25.4957 | 3.2002 | 51.4197 | 164.5525 | 16.0677 | 14.5547 | 4.9344 | 268.761 | 198.1349 | 144.2982 | 173.0384 | 121.6231
2012-10-30 16:33 | 28.6824 | 214.852 | 160.3052 | 189.7939 | 137.3515 | 71.7837 | 2.0214 | 51.546 | 104.1933 | 25.5006 | 3.2 | 51.4208 | 164.5452 | 16.0692 | 14.4979 | 4.9488 | 268.7385 | 200.3541 | 145.8073 | 175.296 | 122.8536
2012-10-30 16:33 | 28.6802 | 217.244 | 161.8141 | 192.4607 | 138.722 | 71.9889 | 2.021 | 51.5466 | 104.1762 | 25.5053 | 3.1997 | 51.4212 | 164.531 | 16.0708 | 14.4875 | 4.9629 | 268.7072 | 202.7565 | 147.3266 | 177.9732 | 124.2345
2012-10-30 16:33 | 28.6705 | 219.5379 | 163.2197 | 195.0707 | 140.0937 | 72.2119 | 2.0207 | 51.5474 | 104.1604 | 25.51 | 3.1994 | 51.422 | 164.5188 | 16.0725 | 14.5145 | 4.9783 | 268.6792 | 205.0234 | 148.7052 | 180.5562 | 125.5792
2012-10-30 16:33 | 28.6824 | 221.9004 | 164.7334 | 197.7355 | 141.4383 | 72.446 | 2.0203 | 51.5483 | 104.1421 | 25.5154 | 3.1991 | 51.4229 | 164.5048 | 16.0744 | 14.5468 | 4.9944 | 268.6469 | 207.3536 | 150.1866 | 183.1887 | 126.8915
2012-10-30 16:33 | 28.675 | 224.1802 | 166.4388 | 200.579 | 142.8757 | 72.6922 | 2.0199 | 51.5497 | 104.1262 | 25.5207 | 3.1988 | 51.4242 | 164.4946 | 16.0762 | 14.5272 | 5.0114 | 268.6209 | 209.653 | 151.9116 | 186.0518 | 128.3485
2012-10-30 16:33 | 28.6765 | 226.5982 | 168.108 | 203.354 | 144.2734 | 72.9193 | 2.0196 | 51.5508 | 104.1111 | 25.5255 | 3.1985 | 51.4252 | 164.4817 | 16.0781 | 14.5532 | 5.0271 | 268.5928 | 212.045 | 153.5548 | 188.8008 | 129.7202
2012-10-30 16:33 | 28.6757 | 229.0093 | 169.6946 | 206.1633 | 145.6585 | 73.1545 | 2.0193 | 51.5516 | 104.0957 | 25.53 | 3.1982 | 51.426 | 164.4686 | 16.0799 | 14.532 | 5.0433 | 268.5644 | 214.4773 | 155.1626 | 191.6313 | 131.1265
2012-10-30 16:33 | 28.6772 | 231.2291 | 171.1873 | 208.7158 | 147.0584 | 73.373 | 2.0189 | 51.5522 | 104.0793 | 25.5346 | 3.1978 | 51.4266 | 164.4538 | 16.0817 | 14.5457 | 5.0583 | 268.5331 | 216.6834 | 156.6416 | 194.1701 | 132.5127
2012-10-30 16:33 | 28.681 | 233.1752 | 172.5781 | 210.9875 | 148.2926 | 73.5556 | 2.0186 | 51.5525 | 104.0625 | 25.5391 | 3.1975 | 51.427 | 164.4363 | 16.0837 | 14.5332 | 5.0709 | 268.4988 | 218.642 | 158.0449 | 196.4543 | 133.7594
2012-10-30 16:33 | 28.6787 | 235.4839 | 174.1744 | 213.655 | 149.7756 | 73.7914 | 2.0182 | 51.553 | 104.0437 | 25.5442 | 3.1971 | 51.4273 | 164.4178 | 16.0857 | 14.5227 | 5.0872 | 268.4614 | 220.9612 | 159.6517 | 199.1323 | 135.2529
2012-10-30 16:33 | 28.6757 | 238.0265 | 175.9312 | 216.5662 | 151.3208 | 74.0191 | 2.0179 | 51.5535 | 104.0275 | 25.5487 | 3.1967 | 51.4278 | 164.4003 | 16.0876 | 14.4708 | 5.1029 | 268.4279 | 223.5557 | 161.4604 | 202.0954 | 136.85
2012-10-30 16:33 | 28.6705 | 240.0276 | 177.402 | 218.9591 | 152.5837 | 74.2156 | 2.0176 | 51.5538 | 104.0127 | 25.5526 | 3.1964 | 51.4282 | 164.3854 | 16.0894 | 14.4044 | 5.1164 | 268.3981 | 225.6232 | 162.9976 | 204.5547 | 138.1793
2012-10-30 16:33 | 28.681 | 242.1141 | 178.8503 | 221.151 | 153.7244 | 74.4202 | 2.0173 | 51.5543 | 103.9983 | 25.5566 | 3.1961 | 51.4285 | 164.3688 | 16.0912 | 14.4353 | 5.1305 | 268.367 | 227.6788 | 164.415 | 206.7157 | 139.2891
2012-10-30 16:33 | 28.6914 | 244.1249 | 180.3072 | 223.4841 | 154.9738 | 74.6074 | 2.0169 | 51.5549 | 103.9824 | 25.5611 | 3.1957 | 51.4291 | 164.3519 | 16.0932 | 14.4416 | 5.1434 | 268.3343 | 229.6833 | 165.8656 | 209.0425 | 140.5322
2012-10-30 16:33 | 28.6802 | 246.0424 | 181.7394 | 225.9801 | 156.3416 | 74.7993 | 2.0165 | 51.5545 | 103.9619 | 25.5657 | 3.1953 | 51.4288 | 164.3291 | 16.0953 | 14.4263 | 5.1567 | 268.291 | 231.6161 | 167.3131 | 211.5538 | 141.9153
2012-10-30 16:33 | 28.6757 | 248.0282 | 183.1322 | 228.2618 | 157.5973 | 75.0293 | 2.0163 | 51.5551 | 103.9481 | 25.5698 | 3.1949 | 51.4292 | 164.3119 | 16.0972 | 14.4014 | 5.1725 | 268.26 | 233.6268 | 168.7308 | 213.8604 | 143.1959
2012-10-30 16:33 | 28.6839 | 249.9864 | 184.5814 | 230.3548 | 158.5183 | 75.2079 | 2.016 | 51.5548 | 103.9327 | 25.5733 | 3.1945 | 51.429 | 164.292 | 16.099 | 14.4055 | 5.1848 | 268.2246 | 235.5809 | 170.1759 | 215.9493 | 144.1128
2012-10-30 16:33 | 28.675 | 251.8716 | 185.9712 | 232.5087 | 159.8641 | 75.405 | 2.0156 | 51.5544 | 103.915 | 25.5773 | 3.1941 | 51.4286 | 164.2698 | 16.101 | 14.3966 | 5.1984 | 268.1848 | 237.475 | 171.5746 | 218.1121 | 145.4675
2012-10-30 16:33 | 28.675 | 253.7644 | 187.3156 | 234.7454 | 161.3515 | 75.5945 | 2.0153 | 51.5548 | 103.8971 | 25.582 | 3.1937 | 51.4288 | 164.247 | 16.1033 | 14.364 | 5.2115 | 268.1441 | 239.4004 | 172.9516 | 220.3814 | 146.9875
2012-10-30 16:33 | 28.6765 | 255.9429 | 188.8818 | 237.1121 | 162.7137 | 75.7881 | 2.015 | 51.5549 | 103.8809 | 25.5861 | 3.1933 | 51.429 | 164.2276 | 16.1053 | 14.4185 | 5.2248 | 268.1085 | 241.5244 | 174.4633 | 222.6936 | 148.2952
2012-10-30 16:33 | 28.678 | 257.9244 | 190.4173 | 239.6465 | 163.9294 | 75.9939 | 2.0147 | 51.555 | 103.8667 | 25.5897 | 3.1929 | 51.429 | 164.2092 | 16.1071 | 14.4323 | 5.239 | 268.0759 | 243.4921 | 175.985 | 225.2142 | 149.4971
2012-10-30 16:33 | 28.681 | 259.633 | 191.743 | 241.7374 | 165.2096 | 76.1811 | 2.0144 | 51.5555 | 103.854 | 25.5933 | 3.1926 | 51.4294 | 164.193 | 16.109 | 14.4353 | 5.2519 | 268.047 | 245.1977 | 177.3077 | 227.3021 | 150.7743
2012-10-30 16:33 | 28.6765 | 261.2473 | 192.9753 | 243.3884 | 166.2577 | 76.3337 | [8 empty] | 14.4226 | 5.2624 | 0 | 246.8247 | 178.5527 | 228.9658 | 151.8351
2012-10-30 16:34 | 0 | 204.848 | 152.8714 | 213.1781 | 136.7072 | 70.3932 |  |  |  |  | 0.04017 | 0.74922 | 0.030096 | 18.6513 | -14.2661 | 4.8529 | 0.030096 | 219.1141 | 167.1375 | 227.4442 | 150.9733
2012-10-30 16:34 | 0 | 195.1706 | 147.8138 | 203.1732 | 133.1377 | 68.8305 | 0.12089 | 3.166 | 0.38273 | 26.19 | 0.20151 | 3.3302 | 0.67109 | 16.5261 | -14.2702 | 4.7452 | 1.0538 | 209.4408 | 162.084 | 217.4434 | 147.4079
2012-10-30 16:34 | 28.5963 | 222.5702 | 175.8323 | 230.4914 | 161.3581 | 68.8264 | 0.45361 | 11.432 | 5.1857 | 25.2023 | 0.71427 | 11.5309 | 8.2361 | 16.1437 | 14.3139 | 4.7449 | 13.4218 | 208.2563 | 161.5184 | 216.1775 | 147.0442
2012-10-30 16:34 | 28.5963 | 221.6214 | 175.442 | 229.3626 | 160.8892 | 68.9611 | 0.5368 | 13.5071 | 7.2507 | 25.1621 | 0.83471 | 13.456 | 11.2318 | 16.1207 | 14.2894 | 4.7542 | 18.4825 | 207.332 | 161.1526 | 215.0732 | 146.5998
2012-10-30 16:34 | 28.5851 | 220.6928 | 175.1264 | 228.2853 | 160.51 | 68.9888 | 0.64493 | 16.2057 | 10.4515 | 25.1279 | 1.0032 | 16.1485 | 16.1997 | 16.0975 | 14.3394 | 4.7561 | 26.6512 | 206.3534 | 160.787 | 213.9459 | 146.1706
2012-10-30 16:34 | 28.5918 | 219.9317 | 174.7624 | 227.1037 | 160.0636 | 69.0709 | 0.75263 | 18.9009 | 14.2254 | 25.1132 | 1.1719 | 18.8444 | 22.0834 | 16.0804 | 14.391 | 4.7617 | 36.3088 | 205.5407 | 160.3714 | 212.7127 | 145.6726
2012-10-30 16:34 | 28.5873 | 219.2736 | 174.4887 | 225.8479 | 159.5944 | 69.3003 | 0.91307 | 22.9333 | 20.9397 | 25.1169 | 1.4229 | 22.8551 | 32.5208 | 16.0621 | 14.3171 | 4.7776 | 53.4605 | 204.9565 | 160.1716 | 211.5308 | 145.2773
2012-10-30 16:35 | 28.5866 | 218.887 | 174.4373 | 224.7766 | 159.1911 | 69.4812 | 0.93539 | 23.5019 | 21.9835 | 25.1252 | 1.4616 | 23.472 | 34.3057 | 16.0595 | 14.3123 | 4.79 | 56.2892 | 204.5747 | 160.125 | 210.4643 | 144.8788
2012-10-30 16:35 | 28.5753 | 218.5883 | 174.2075 | 223.7813 | 158.7928 | 69.5714 | 1.0378 | 26.084 | 27.0698 | 25.1341 | 1.6268 | 26.1126 | 42.4801 | 16.0515 | 14.3377 | 4.7963 | 69.55 | 204.2506 | 159.8698 | 209.4436 | 144.4551
2012-10-30 16:35 | 28.5776 | 218.4803 | 174.2527 | 222.9018 | 158.4275 | 69.8719 | 1.1328 | 28.4985 | 32.2838 | 25.1571 | 1.768 | 28.3703 | 50.1595 | 16.0463 | 14.3156 | 4.817 | 82.4433 | 204.1647 | 159.9371 | 208.5862 | 144.1119
2012-10-30 16:35 | 28.5798 | 218.6164 | 174.3408 | 222.1324 | 158.213 | 70.2528 | 1.3743 | 34.646 | 47.6141 | 25.2099 | 2.1454 | 34.4106 | 73.824 | 16.0394 | 14.3178 | 4.8432 | 121.4381 | 204.2986 | 160.023 | 207.8146 | 143.8952
2012-10-30 16:35 | 28.5731 | 219.3263 | 175.1456 | 221.3698 | 158.0168 | 71.1052 | 1.5997 | 40.4707 | 64.7393 | 25.2996 | 2.509 | 40.2542 | 100.9988 | 16.0438 | 14.3478 | 4.902 | 165.7382 | 204.9785 | 160.7978 | 207.022 | 143.669
2012-10-30 16:35 | 28.5678 | 220.8527 | 176.3855 | 220.8094 | 157.9187 | 72.0212 | 1.837 | 46.6627 | 85.7183 | 25.4019 | 2.8929 | 46.4564 | 134.3954 | 16.0585 | 14.318 | 4.9651 | 220.1137 | 206.5347 | 162.0675 | 206.4914 | 143.6007
2012-10-30 16:35 | 28.5649 | 222.9246 | 177.8198 | 220.6214 | 157.9854 | 72.9771 | 1.9867 | 50.6199 | 100.5688 | 25.4789 | 3.1358 | 50.4054 | 158.0615 | 16.0742 | 14.258 | 5.031 | 258.6302 | 208.6666 | 163.5618 | 206.3634 | 143.7274
2012-10-30 16:35 | 28.5626 | 225.8035 | 179.5372 | 220.631 | 158.2152 | 73.677 | 2.0205 | 51.5569 | 104.1695 | 25.5172 | 3.1976 | 51.4313 | 164.4545 | 16.0846 | 14.321 | 5.0793 | 268.624 | 211.4825 | 165.2162 | 206.31 | 143.8942
2012-10-30 16:35 | 28.5671 | 228.7787 | 180.9652 | 220.9544 | 158.773 | 73.9498 | 2.0196 | 51.5521 | 104.1172 | 25.5253 | 3.1971 | 51.4264 | 164.4131 | 16.0856 | 14.3173 | 5.0981 | 268.5303 | 214.4614 | 166.6479 | 206.6371 | 144.4557
2012-10-30 16:35 | 28.5649 | 231.6594 | 182.1329 | 221.5191 | 159.3823 | 74.1804 | 2.019 | 51.5487 | 104.0765 | 25.5319 | 3.1966 | 51.4229 | 164.3779 | 16.0868 | 14.3192 | 5.114 | 268.4544 | 217.3402 | 167.8137 | 207.1999 | 145.0631
2012-10-30 16:35 | 28.5664 | 234.3342 | 183.3322 | 222.3119 | 160.2008 | 74.3722 | 2.0184 | 51.5462 | 104.0405 | 25.5382 | 3.1961 | 51.4203 | 164.3468 | 16.0882 | 14.3778 | 5.1272 | 268.3873 | 219.9564 | 168.9544 | 207.9341 | 145.823
2012-10-30 16:35 | 28.5641 | 236.6619 | 184.3676 | 223.1442 | 161.1011 | 74.5612 | 2.0179 | 51.5454 | 104.0159 | 25.5434 | 3.1958 | 51.4193 | 164.3242 | 16.0898 | 14.3673 | 5.1402 | 268.3401 | 222.2946 | 170.0003 | 208.7769 | 146.7338
2012-10-30 16:35 | 28.5581 | 238.9573 | 185.4624 | 224.1853 | 161.6842 | 74.7473 | 2.0175 | 51.545 | 103.9933 | 25.5487 | 3.1954 | 51.4188 | 164.3045 | 16.0914 | 14.3736 | 5.1531 | 268.2978 | 224.5837 | 171.0888 | 209.8117 | 147.3106
2012-10-30 16:35 | 28.5544 | 241.0811 | 186.5678 | 225.4096 | 162.452 | 74.9461 | 2.0171 | 51.545 | 103.9708 | 25.5541 | 3.1951 | 51.4188 | 164.2872 | 16.0931 | 14.4352 | 5.1668 | 268.258 | 226.6459 | 172.1326 | 210.9744 | 148.0168
2012-10-30 16:35 | 28.5596 | 243.2754 | 187.6354 | 226.8 | 163.3532 | 75.1513 | 2.0167 | 51.5451 | 103.9496 | 25.5594 | 3.1947 | 51.4189 | 164.2663 | 16.0952 | 14.5179 | 5.1809 | 268.216 | 228.7575 | 173.1175 | 212.2821 | 148.8353
2012-10-30 16:35 | 28.5566 | 245.5543 | 188.8364 | 228.3276 | 164.4407 | 75.3512 | 2.0163 | 51.5457 | 103.9301 | 25.5649 | 3.1943 | 51.4194 | 164.2484 | 16.0973 | 14.5026 | 5.1947 | 268.1785 | 231.0517 | 174.3338 | 213.825 | 149.9381
2012-10-30 16:35 | 28.5499 | 248.0082 | 190.2343 | 229.9298 | 165.5445 | 75.5391 | 2.0159 | 51.5464 | 103.9136 | 25.5696 | 3.1939 | 51.4201 | 164.232 | 16.0993 | 14.4633 | 5.2077 | 268.1456 | 233.5449 | 175.771 | 215.4665 | 151.0812
2012-10-30 16:35 | 28.5604 | 250.1565 | 191.434 | 231.6864 | 166.4758 | 75.7327 | 2.0156 | 51.5467 | 103.8952 | 25.5745 | 3.1936 | 51.4205 | 164.2139 | 16.1013 | 14.482 | 5.221 | 268.1091 | 235.6745 | 176.952 | 217.2044 | 151.9938
2012-10-30 16:35 | 28.5484 | 252.0892 | 192.5762 | 233.5261 | 167.2567 | 75.9026 | 2.0152 | 51.5475 | 103.8797 | 25.579 | 3.1932 | 51.4212 | 164.1978 | 16.1033 | 14.4822 | 5.2327 | 268.0775 | 237.607 | 178.094 | 219.0439 | 152.7745
2012-10-30 16:35 | 28.5431 | 254.215 | 193.812 | 235.4269 | 168.2624 | 76.0944 | 2.0148 | 51.5482 | 103.8613 | 25.5843 | 3.1927 | 51.4219 | 164.1771 | 16.1059 | 14.5136 | 5.2459 | 268.0384 | 239.7014 | 179.2984 | 220.9133 | 153.7488
2012-10-30 16:35 | 28.5566 | 256.6472 | 195.2437 | 237.6717 | 169.591 | 76.3169 | 2.0145 | 51.5491 | 103.8432 | 25.5896 | 3.1923 | 51.4226 | 164.1562 | 16.1084 | 14.4985 | 5.2613 | 267.9995 | 242.1487 | 180.7452 | 223.1732 | 155.0925
2012-10-30 16:35 | 0 | 219.7755 | 161.3548 | 216.9271 | 143.3087 | 71.3641 | 0.46243 | 11.55 | 5.341 | 24.977 | 0.72846 | 11.7371 | 8.5501 | 16.1122 | -14.1519 | 4.9198 | 13.8911 | 233.9274 | 175.5067 | 231.079 | 157.4606
2012-10-30 16:36 | 28.5297 | 244.9551 | 189.1586 | 245.4725 | 170.7321 | 71.9305 | 1.0312 | 25.9661 | 26.7772 | 25.1795 | 1.6148 | 25.9039 | 41.8308 | 16.0411 | 14.3493 | 4.9589 | 68.608 | 230.6058 | 174.8093 | 231.1232 | 156.3828
2012-10-30 16:36 | 28.5297 | 244.3592 | 189.4058 | 245.2883 | 170.7204 | 72.1819 | 1.0354 | 26.0783 | 27.0002 | 25.1879 | 1.624 | 26.0499 | 42.3041 | 16.0409 | 14.3941 | 4.9762 | 69.3043 | 229.9651 | 175.0117 | 230.8942 | 156.3263
2012-10-30 16:36 | 28.5222 | 244.0929 | 189.3473 | 245.0809 | 170.6701 | 72.2327 | 1.1152 | 28.1036 | 31.3416 | 25.2002 | 1.7495 | 28.0571 | 49.0847 | 16.0376 | 14.4111 | 4.9797 | 80.4263 | 229.6818 | 174.9362 | 230.6698 | 156.259
2012-10-30 16:36 | 28.5154 | 243.7998 | 189.4307 | 244.8114 | 170.5659 | 72.3373 | 1.1183 | 28.189 | 31.5246 | 25.2063 | 1.7536 | 28.123 | 49.3156 | 16.0376 | 14.4288 | 4.9869 | 80.8402 | 229.371 | 175.0019 | 230.3826 | 156.1371
2012-10-30 16:36 | 28.5252 | 243.594 | 189.3895 | 244.5742 | 170.3071 | 72.3229 | 1.1179 | 28.1788 | 31.5023 | 25.2059 | 1.7529 | 28.1129 | 49.28 | 16.0377 | 14.459 | 4.9859 | 80.7823 | 229.135 | 174.9305 | 230.1152 | 155.8481
2012-10-30 16:36 | 28.5259 | 243.3162 | 189.2255 | 244.277 | 170.483 | 72.3165 | 1.3504 | 34.1039 | 46.0538 | 25.2547 | 2.1075 | 33.8055 | 71.2463 | 16.0403 | 14.4475 | 4.9855 | 117.3002 | 228.8687 | 174.778 | 229.8295 | 156.0355
2012-10-30 16:36 | 28.5259 | 243.677 | 190.4889 | 243.928 | 170.6426 | 74.2214 | 2.0177 | 51.4829 | 103.8791 | 25.5152 | 3.1924 | 51.357 | 163.9527 | 16.0872 | 14.5168 | 5.1168 | 267.8318 | 229.1602 | 175.9721 | 229.4112 | 156.1258
2012-10-30 16:36 | 28.5237 | 245.7957 | 192.8539 | 243.8317 | 170.7573 | 75.509 | 2.0158 | 51.495 | 103.805 | 25.5453 | 3.1918 | 51.3648 | 163.944 | 16.093 | 14.5472 | 5.2056 | 267.749 | 231.2485 | 178.3067 | 229.2845 | 156.2101
2012-10-30 16:36 | 28.5259 | 248.4228 | 194.3447 | 244.0299 | 170.8529 | 75.8009 | 2.013 | 51.4414 | 103.5517 | 25.5546 | 3.1772 | 51.1404 | 162.4825 | 16.0962 | 14.5331 | 5.2257 | 266.0342 | 233.8897 | 179.8116 | 229.4968 | 156.3198
2012-10-30 16:36 | 28.5214 | 250.9002 | 194.8077 | 244.3626 | 171.3236 | 74.7207 | 0.6545 | 16.6474 | 10.8956 | 25.4355 | 0.97522 | 15.6898 | 15.3011 | 16.0885 | 14.5205 | 5.1512 | 26.1967 | 236.3797 | 180.2872 | 229.8421 | 156.8031
2012-10-30 16:36 | 28.5252 | 250.1604 | 191.9602 | 244.7781 | 171.4014 | 71.8918 | 0.73706 | 18.5577 | 13.6781 | 25.178 | 1.1653 | 18.7127 | 21.8055 | 16.0586 | 14.5202 | 4.9562 | 35.4836 | 235.6402 | 177.44 | 230.2579 | 156.8812
2012-10-30 16:36 | 28.5222 | 248.5053 | 191.07 | 244.9241 | 171.7062 | 72.1628 | 0.89103 | 22.4351 | 19.9902 | 25.1789 | 1.3951 | 22.3835 | 31.2265 | 16.0447 | 14.4519 | 4.9749 | 51.2167 | 234.0534 | 176.6181 | 230.4722 | 157.2543
2012-10-30 16:36 | 28.5334 | 247.5921 | 190.7947 | 244.9509 | 172.0174 | 72.1281 | 0.89163 | 22.4481 | 20.0153 | 25.1766 | 1.3959 | 22.3967 | 31.2631 | 16.0448 | 14.5243 | 4.9725 | 51.2784 | 233.0678 | 176.2704 | 230.4266 | 157.4931
2012-10-30 16:36 | 28.5259 | 246.6757 | 190.5124 | 244.8768 | 172.0957 | 72.0599 | 0.89206 | 22.4568 | 20.0328 | 25.1742 | 1.3964 | 22.4054 | 31.2875 | 16.0448 | 14.5168 | 4.9678 | 51.3203 | 232.1589 | 175.9956 | 230.36 | 157.5789
2012-10-30 16:36 | 28.5282 | 245.7023 | 190.2714 | 244.7144 | 171.9577 | 71.9854 | 0.89238 | 22.4627 | 20.0452 | 25.1716 | 1.3968 | 22.4112 | 31.3038 | 16.0448 | 14.5069 | 4.9627 | 51.349 | 231.1954 | 175.7645 | 230.2075 | 157.4508
2012-10-30 16:36 | 28.5326 | 244.9266 | 189.9972 | 244.5542 | 171.6738 | 71.9235 | 0.89266 | 22.4673 | 20.0556 | 25.169 | 1.3971 | 22.416 | 31.317 | 16.0448 | 14.4949 | 4.9584 | 51.3726 | 230.4317 | 175.5023 | 230.0593 | 157.1789
2012-10-30 16:36 | 28.5267 | 244.0503 | 189.6853 | 244.2778 | 171.3718 | 71.8513 | 0.89287 | 22.4702 | 20.063 | 25.1662 | 1.3973 | 22.4188 | 31.3251 | 16.0447 | 14.5176 | 4.9534 | 51.3881 | 229.5327 | 175.1677 | 229.7602 | 156.8542
2012-10-30 16:36 | 28.5304 | 243.2736 | 189.383 | 243.9521 | 171.4962 | 71.7727 | 0.89297 | 22.4709 | 20.066 | 25.1641 | 1.3973 | 22.4197 | 31.3278 | 16.0446 | 14.4438 | 4.948 | 51.3938 | 228.8298 | 174.9392 | 229.5083 | 157.0524
2012-10-30 16:36 | 28.5297 | 242.6768 | 189.1312 | 243.622 | 171.367 | 71.7219 | 0.89303 | 22.4709 | 20.0672 | 25.1625 | 1.3973 | 22.4196 | 31.3275 | 16.0446 | 14.4716 | 4.9445 | 51.3947 | 228.2052 | 174.6596 | 229.1504 | 156.8954
2012-10-30 16:36 | 28.5244 | 241.9143 | 188.816 | 243.2794 | 171.0773 | 71.6693 | 0.89314 | 22.4712 | 20.0698 | 25.1599 | 1.3973 | 22.4198 | 31.3279 | 16.0447 | 14.4827 | 4.9409 | 51.3977 | 227.4316 | 174.3333 | 228.7967 | 156.5946
2012-10-30 16:36 | 28.5214 | 241.1148 | 188.4796 | 242.8489 | 170.7433 | 71.5999 | 0.89321 | 22.471 | 20.0713 | 25.1577 | 1.3973 | 22.4198 | 31.3277 | 16.0448 | 14.4674 | 4.9361 | 51.399 | 226.6474 | 174.0122 | 228.3815 | 156.2759
2012-10-30 16:36 | 28.5282 | 240.4152 | 188.1256 | 242.3536 | 170.6761 | 71.5346 | 0.89328 | 22.471 | 20.0729 | 25.1556 | 1.3973 | 22.4198 | 31.3276 | 16.0448 | 14.4865 | 4.9316 | 51.4005 | 225.9287 | 173.6391 | 227.8671 | 156.1896
2012-10-30 16:36 | 28.5311 | 239.8411 | 187.8069 | 241.8504 | 170.3558 | 71.4768 | 0.89336 | 22.4712 | 20.075 | 25.1535 | 1.3973 | 22.42 | 31.3283 | 16.0448 | 14.4486 | 4.9276 | 51.4033 | 225.3925 | 173.3583 | 227.4018 | 155.9072
2012-10-30 16:36 | 28.5461 | 239.228 | 187.5317 | 241.3514 | 169.8219 | 71.4242 | 0.89341 | 22.4709 | 20.0758 | 25.1518 | 1.3973 | 22.4197 | 31.3273 | 16.0449 | 14.4309 | 4.924 | 51.4031 | 224.7971 | 173.1008 | 226.9205 | 155.391
2012-10-30 16:36 | 28.5341 | 238.5603 | 187.2178 | 240.845 | 169.6191 | 71.3612 | 0.89346 | 22.4705 | 20.0766 | 25.1499 | 1.3973 | 22.4193 | 31.3261 | 16.0449 | 14.4271 | 4.9196 | 51.4027 | 224.1332 | 172.7907 | 226.4179 | 155.192
2012-10-30 16:36 | 28.5326 | 237.903 | 186.8674 | 240.3058 | 169.4269 | 71.2982 | 0.89347 | 22.4697 | 20.076 | 25.1488 | 1.3972 | 22.4185 | 31.3238 | 16.0449 | 14.3889 | 4.9153 | 51.3998 | 223.5141 | 172.4785 | 225.9169 | 155.038
2012-10-30 16:36 | 28.5282 | 237.2938 | 186.4946 | 239.7126 | 169.2435 | 71.2485 | 0.89348 | 22.4685 | 20.0751 | 25.1471 | 1.3971 | 22.4173 | 31.3202 | 16.0451 | 14.4334 | 4.9119 | 51.3953 | 222.8604 | 172.0612 | 225.2792 | 154.8101
2012-10-30 16:36 | 28.5401 | 236.7322 | 186.1028 | 239.061 | 169.0064 | 71.1919 | 0.89354 | 22.468 | 20.0761 | 25.1448 | 1.3971 | 22.4169 | 31.3187 | 16.0453 | 14.4657 | 4.908 | 51.3948 | 222.2665 | 171.6371 | 224.5953 | 154.5407
2012-10-30 16:36 | 28.5349 | 236.1024 | 185.8546 | 238.51 | 168.7599 | 71.1387 | 0.89357 | 22.4676 | 20.0764 | 25.1437 | 1.3971 | 22.4164 | 31.3175 | 16.0452 | 14.485 | 4.9043 | 51.3939 | 221.6174 | 171.3696 | 224.025 | 154.2749
2012-10-30 16:37 | 28.5401 | 235.5654 | 185.5385 | 237.9733 | 168.4267 | 71.0965 | 0.89362 | 22.4675 | 20.0774 | 25.1421 | 1.3971 | 22.4164 | 31.3177 | 16.0451 | 14.482 | 4.9014 | 51.3951 | 221.0834 | 171.0565 | 223.4913 | 153.9447
2012-10-30 16:37 | 28.5386 | 234.9902 | 185.2235 | 237.3592 | 168.2074 | 71.0428 | 0.8937 | 22.4679 | 20.0795 | 25.1403 | 1.3971 | 22.4167 | 31.3181 | 16.0453 | 14.4928 | 4.8977 | 51.3976 | 220.4974 | 170.7307 | 222.8664 | 153.7146
2012-10-30 16:37 | 28.5409 | 234.4501 | 184.8811 | 236.7958 | 168.0269 | 70.9757 | 0.89375 | 22.4679 | 20.0808 | 25.1389 | 1.3971 | 22.4167 | 31.3179 | 16.0454 | 14.5236 | 4.8931 | 51.3987 | 219.9265 | 170.3575 | 222.2722 | 153.5033
2012-10-30 16:37 | 28.5401 | 233.8755 | 184.5551 | 236.2922 | 167.8238 | 70.9388 | 0.8938 | 22.468 | 20.0818 | 25.1376 | 1.3971 | 22.4168 | 31.3183 | 16.0454 | 14.5718 | 4.8905 | 51.4001 | 219.3037 | 169.9833 | 221.7204 | 153.252
2012-10-30 16:37 | 28.5476 | 233.376 | 184.2728 | 235.7497 | 167.5319 | 70.8943 | 0.89385 | 22.468 | 20.083 | 25.1361 | 1.3971 | 22.4168 | 31.3185 | 16.0453 | 14.567 | 4.8875 | 51.4016 | 218.809 | 169.7058 | 221.1827 | 152.9649
2012-10-30 16:37 | 28.5484 | 232.791 | 183.925 | 235.1375 | 167.2683 | 70.8347 | 0.89393 | 22.4686 | 20.0854 | 25.1346 | 1.3971 | 22.4175 | 31.3203 | 16.0453 | 14.5964 | 4.8833 | 51.4057 | 218.1946 | 169.3286 | 220.5411 | 152.6719
2012-10-30 16:37 | 28.5604 | 232.2841 | 183.6629 | 234.5679 | 167.0043 | 70.7729 | 0.89399 | 22.4685 | 20.0865 | 25.1329 | 1.3971 | 22.4174 | 31.3198 | 16.0454 | 14.5431 | 4.8791 | 51.4064 | 217.741 | 169.1198 | 220.0248 | 152.4612
2012-10-30 16:37 | 28.5566 | 231.6905 | 183.2597 | 233.8999 | 166.7712 | 70.722 | 0.89401 | 22.4679 | 20.0865 | 25.1317 | 1.3971 | 22.4168 | 31.3178 | 16.0456 | 14.5801 | 4.8756 | 51.4043 | 217.1104 | 168.6796 | 219.3198 | 152.1911
2012-10-30 16:37 | 28.5604 | 231.1752 | 182.9268 | 233.2748 | 166.4447 | 70.6648 | 0.97402 | 24.4763 | 23.8404 | 25.129 | 1.5292 | 24.5275 | 37.5064 | 16.0398 | 14.5391 | 4.8716 | 61.3469 | 216.6361 | 168.3877 | 218.7357 | 151.9056
2012-10-30 16:37 | 28.5604 | 230.7584 | 182.8485 | 232.748 | 166.2038 | 71.1555 | 1.3752 | 34.6733 | 47.6826 | 25.2133 | 2.1524 | 34.5008 | 74.2594 | 16.0291 | 14.6002 | 4.9055 | 121.9421 | 216.1582 | 168.2483 | 218.1478 | 151.6036
2012-10-30 16:37 | 28.5634 | 231.1428 | 183.7834 | 232.2674 | 165.9853 | 72.4125 | 1.7638 | 44.7342 | 78.9009 | 25.3628 | 2.787 | 44.7203 | 124.6339 | 16.0462 | 14.6726 | 4.9921 | 203.5348 | 216.4702 | 169.1108 | 217.5948 | 151.3127
2012-10-30 16:37 | 28.5619 | 232.7652 | 185.5838 | 232.0103 | 165.7973 | 74.1 | 1.7109 | 43.572 | 74.5459 | 25.4678 | 2.6665 | 42.8477 | 114.2556 | 16.0686 | 14.7608 | 5.1085 | 188.8015 | 218.0044 | 170.823 | 217.2495 | 151.0365
2012-10-30 16:37 | 28.5701 | 234.9509 | 186.2386 | 231.9202 | 165.9532 | 72.1131 | 1.0043 | 25.2969 | 25.4065 | 25.1878 | 1.5736 | 25.2291 | 39.7014 | 16.0324 | 14.7282 | 4.9715 | 65.1079 | 220.2227 | 171.5104 | 217.192 | 151.225
2012-10-30 16:37 | 28.5686 | 234.7176 | 184.427 | 231.8715 | 166.3285 | 71.2399 | 1.0059 | 25.3077 | 25.4558 | 25.1604 | 1.5746 | 25.2444 | 39.75 | 16.0322 | 14.8286 | 4.9113 | 65.2058 | 219.889 | 169.5984 | 217.0429 | 151.4999
2012-10-30 16:37 | 28.5708 | 233.9691 | 183.8456 | 231.7283 | 166.4318 | 71.1301 | 1.0042 | 25.2631 | 25.3694 | 25.1571 | 1.5721 | 25.2052 | 39.6259 | 16.0324 | 14.7982 | 4.9037 | 64.9953 | 219.1709 | 169.0474 | 216.9301 | 151.6336
2012-10-30 16:37 | 28.5693 | 233.4055 | 183.5308 | 231.5734 | 166.3409 | 71.0769 | 1.0054 | 25.2925 | 25.4295 | 25.1562 | 1.5741 | 25.2372 | 39.7271 | 16.0323 | 14.7804 | 4.9 | 65.1566 | 218.6251 | 168.7504 | 216.793 | 151.5605
2012-10-30 16:37 | 28.5664 | 232.8916 | 183.3478 | 231.37 | 166.2098 | 71.0307 | 1.0083 | 25.3635 | 25.5742 | 25.1547 | 1.5785 | 25.3066 | 39.9465 | 16.032 | 14.8101 | 4.8969 | 65.5207 | 218.0815 | 168.5377 | 216.5599 | 151.3997
2012-10-30 16:37 | 28.5753 | 232.5152 | 183.1492 | 231.1744 | 166.0438 | 71.0041 | 1.0101 | 25.4088 | 25.6662 | 25.154 | 1.5812 | 25.3508 | 40.0858 | 16.0321 | 14.8231 | 4.895 | 65.7521 | 217.6921 | 168.3261 | 216.3513 | 151.2207
2012-10-30 16:37 | 28.5723 | 232.0679 | 182.9035 | 230.9787 | 165.8582 | 70.9752 | 1.0086 | 25.37 | 25.5891 | 25.1528 | 1.5788 | 25.3118 | 39.9625 | 16.0322 | 14.8405 | 4.893 | 65.5515 | 217.2274 | 168.063 | 216.1382 | 151.0177
2012-10-30 16:37 | 28.5783 | 231.6886 | 182.7177 | 230.7645 | 165.8254 | 70.9341 | 1.009 | 25.3783 | 25.6064 | 25.1521 | 1.5792 | 25.3183 | 39.9826 | 16.0325 | 14.9036 | 4.8902 | 65.589 | 216.785 | 167.8141 | 215.8609 | 150.9218
2012-10-30 16:37 | 28.5866 | 231.2919 | 182.6125 | 230.5564 | 165.6123 | 70.8833 | 1.0079 | 25.3491 | 25.5491 | 25.1507 | 1.5775 | 25.2912 | 39.897 | 16.0324 | 14.9282 | 4.8867 | 65.4461 | 216.3637 | 167.6843 | 215.6282 | 150.6841
2012-10-30 16:37 | 28.5911 | 230.8992 | 182.3743 | 230.3132 | 165.4964 | 70.8561 | 1.0064 | 25.3093 | 25.47 | 25.1496 | 1.575 | 25.2519 | 39.7727 | 16.0325 | 14.9735 | 4.8848 | 65.2427 | 215.9257 | 167.4008 | 215.3397 | 150.5229
2012-10-30 16:37 | 28.5858 | 230.6304 | 182.189 | 230.1309 | 165.3551 | 70.8209 | 1.0047 | 25.2658 | 25.3836 | 25.1486 | 1.5723 | 25.2089 | 39.6367 | 16.0328 | 14.9682 | 4.8824 | 65.0204 | 215.6622 | 167.2208 | 215.1627 | 150.3869
2012-10-30 16:37 | 28.5866 | 230.2851 | 181.9511 | 229.8564 | 165.1928 | 70.7671 | 1.0042 | 25.2539 | 25.3609 | 25.1473 | 1.5714 | 25.1945 | 39.5919 | 16.0327 | 14.9445 | 4.8787 | 64.9528 | 215.3406 | 167.0066 | 214.9119 | 150.2483
2012-10-30 16:37 | 28.5978 | 229.8558 | 181.7275 | 229.5647 | 164.9711 | 70.7434 | 1.0068 | 25.3182 | 25.4908 | 25.1468 | 1.5756 | 25.2602 | 39.7989 | 16.0326 | 14.9353 | 4.877 | 65.2897 | 214.9205 | 166.7922 | 214.6294 | 150.0358
2012-10-30 16:37 | 28.5993 | 229.5426 | 181.5568 | 229.3026 | 164.6503 | 70.7128 | 1.0079 | 25.3438 | 25.5434 | 25.1456 | 1.5772 | 25.2859 | 39.8798 | 16.0326 | 14.9613 | 4.8749 | 65.4232 | 214.5813 | 166.5955 | 214.3413 | 149.689
2012-10-30 16:37 | 28.5986 | 229.2665 | 181.3018 | 229.0226 | 164.4749 | 70.6683 | 1.0086 | 25.3618 | 25.5806 | 25.145 | 1.5781 | 25.3014 | 39.9285 | 16.0327 | 15.0013 | 4.8719 | 65.5091 | 214.2652 | 166.3005 | 214.0213 | 149.4736
2012-10-30 16:37 | 28.6046 | 228.8831 | 181.0887 | 228.7375 | 164.2403 | 70.6579 | 1.0093 | 25.3795 | 25.6167 | 25.1444 | 1.5794 | 25.3209 | 39.9908 | 16.0324 | 15.0114 | 4.8712 | 65.6076 | 213.8717 | 166.0773 | 213.7261 | 149.2289
2012-10-30 16:37 | 28.6031 | 228.5944 | 180.8721 | 228.3584 | 164.099 | 70.6192 | 1.0104 | 25.4063 | 25.6718 | 25.1436 | 1.581 | 25.3469 | 40.0729 | 16.0324 | 15.0221 | 4.8685 | 65.7447 | 213.5723 | 165.85 | 213.3363 | 149.0769
2012-10-30 16:38 | 28.612 | 228.1981 | 180.6111 | 228.0407 | 163.9682 | 70.5793 | 1.0073 | 25.3266 | 25.5124 | 25.1421 | 1.5761 | 25.2686 | 39.8254 | 16.0325 | 15.0392 | 4.8657 | 65.3378 | 213.1589 | 165.5719 | 213.0015 | 148.929
2012-10-30 16:38 | 28.6143 | 227.9289 | 180.41 | 227.8227 | 163.9045 | 70.5533 | 1.0092 | 25.3722 | 25.6048 | 25.1418 | 1.5789 | 25.314 | 39.9685 | 16.0325 | 15.0048 | 4.8639 | 65.5733 | 212.9241 | 165.4052 | 212.8179 | 148.8997
2012-10-30 16:38 | 28.621 | 227.6523 | 180.2094 | 227.491 | 163.8103 | 70.5146 | 1.0109 | 25.4148 | 25.6913 | 25.1413 | 1.5816 | 25.3568 | 40.1045 | 16.0323 | 14.987 | 4.8613 | 65.7958 | 212.6653 | 165.2224 | 212.504 | 148.8233
2012-10-30 16:38 | 28.6143 | 227.3977 | 180.0541 | 227.2522 | 163.5359 | 70.514 | 1.0093 | 25.3741 | 25.6098 | 25.1406 | 1.5788 | 25.3124 | 39.9633 | 16.0326 | 14.9233 | 4.8612 | 65.5731 | 212.4744 | 165.1308 | 212.3289 | 148.6126
2012-10-30 16:38 | 28.621 | 227.0779 | 179.8966 | 226.9362 | 163.4262 | 70.4631 | 1.0078 | 25.3347 | 25.5321 | 25.1388 | 1.5766 | 25.2764 | 39.8496 | 16.0327 | 14.9626 | 4.8577 | 65.3817 | 212.1153 | 164.934 | 211.9736 | 148.4636
2012-10-30 16:38 | 28.6233 | 226.7693 | 179.6565 | 226.5922 | 163.3083 | 70.4244 | 1.0083 | 25.3467 | 25.5566 | 25.1386 | 1.5773 | 25.2881 | 39.887 | 16.0325 | 14.9323 | 4.8551 | 65.4436 | 211.837 | 164.7242 | 211.6599 | 148.376
2012-10-30 16:38 | 28.618 | 226.5397 | 179.4753 | 226.3429 | 163.2564 | 70.3996 | 1.0097 | 25.3813 | 25.6275 | 25.1374 | 1.5796 | 25.3243 | 40.0015 | 16.0324 | 14.9311 | 4.8533 | 65.629 | 211.6086 | 164.5442 | 211.4118 | 148.3253
2012-10-30 16:38 | 28.6248 | 226.2632 | 179.3335 | 226.1057 | 163.0188 | 70.3828 | 1.0102 | 25.394 | 25.6537 | 25.1369 | 1.5802 | 25.335 | 40.0349 | 16.0326 | 14.9501 | 4.8522 | 65.6886 | 211.3131 | 164.3834 | 211.1556 | 148.0687
2012-10-30 16:38 | 28.6255 | 226.0041 | 179.1427 | 225.8074 | 162.7713 | 70.3371 | 1.0102 | 25.3925 | 25.651 | 25.1366 | 1.5802 | 25.334 | 40.032 | 16.0325 | 15.0079 | 4.849 | 65.683 | 210.9962 | 164.1348 | 210.7995 | 147.7634
2012-10-30 16:38 | 28.6308 | 225.7536 | 178.9291 | 225.5293 | 162.5478 | 70.3233 | 1.0094 | 25.3734 | 25.6131 | 25.136 | 1.5789 | 25.3139 | 39.9679 | 16.0327 | 14.9887 | 4.8481 | 65.581 | 210.7649 | 163.9404 | 210.5406 | 147.5591
2012-10-30 16:38 | 28.6203 | 225.444 | 178.7036 | 225.1882 | 162.4404 | 70.2776 | 1.009 | 25.3601 | 25.5874 | 25.1349 | 1.5783 | 25.3035 | 39.9355 | 16.0325 | 15.0353 | 4.8449 | 65.5229 | 210.4087 | 163.6683 | 210.1529 | 147.4051
2012-10-30 16:38 | 28.6338 | 225.182 | 178.5335 | 224.8986 | 162.2871 | 70.2557 | 1.0061 | 25.2857 | 25.4387 | 25.1336 | 1.5736 | 25.2283 | 39.698 | 16.0327 | 15.0243 | 4.8434 | 65.1367 | 210.1577 | 163.5092 | 209.8743 | 147.2628
2012-10-30 16:38 | 28.6278 | 224.8926 | 178.3282 | 224.5974 | 162.064 | 70.2169 | 1.0074 | 25.3188 | 25.5062 | 25.1328 | 1.5755 | 25.2598 | 39.7976 | 16.0326 | 14.9816 | 4.8408 | 65.3038 | 209.911 | 163.3466 | 209.6158 | 147.0824
2012-10-30 16:38 | 28.6345 | 224.6278 | 178.1161 | 224.3877 | 161.9312 | 70.1823 | 1.0076 | 25.3246 | 25.5182 | 25.1325 | 1.576 | 25.2678 | 39.8215 | 16.033 | 14.9679 | 4.8384 | 65.3396 | 209.6599 | 163.1482 | 209.4198 | 146.9633
2012-10-30 16:38 | 28.6398 | 224.4166 | 177.9729 | 224.2001 | 161.8047 | 70.1667 | 1.0049 | 25.2554 | 25.3798 | 25.1315 | 1.5714 | 25.1943 | 39.5908 | 16.0328 | 14.9325 | 4.8373 | 64.9706 | 209.4841 | 163.0404 | 209.2676 | 146.8722
2012-10-30 16:38 | 28.6442 | 224.1297 | 177.7741 | 223.8502 | 161.7703 | 70.1458 | 1.006 | 25.2818 | 25.4342 | 25.1304 | 1.5733 | 25.2245 | 39.6856 | 16.0329 | 14.9206 | 4.8359 | 65.1198 | 209.2091 | 162.8535 | 208.9296 | 146.8497
2012-10-30 16:38 | 28.639 | 223.7978 | 177.6244 | 223.5104 | 161.7031 | 70.099 | 1.008 | 25.3303 | 25.5324 | 25.1297 | 1.5762 | 25.2715 | 39.8343 | 16.0327 | 14.9276 | 4.8326 | 65.3667 | 208.8702 | 162.6968 | 208.5828 | 146.7755
2012-10-30 16:38 | 28.645 | 223.5519 | 177.4038 | 223.2763 | 161.5153 | 70.0782 | 1.0085 | 25.343 | 25.5576 | 25.1302 | 1.5771 | 25.2846 | 39.8755 | 16.0327 | 14.9336 | 4.8312 | 65.4331 | 208.6183 | 162.4702 | 208.3427 | 146.5817
2012-10-30 16:38 | 28.6442 | 223.3857 | 177.278 | 223.0511 | 161.3246 | 70.0655 | 1.0049 | 25.2519 | 25.3752 | 25.1292 | 1.5714 | 25.1947 | 39.5918 | 16.0329 | 14.9409 | 4.8303 | 64.9669 | 208.4448 | 162.3371 | 208.1102 | 146.3837
2012-10-30 16:38 | 28.6413 | 223.1506 | 177.1502 | 222.7687 | 161.0233 | 70.017 | 1.0024 | 25.1883 | 25.2493 | 25.1274 | 1.5674 | 25.1308 | 39.3912 | 16.033 | 14.8769 | 4.827 | 64.6405 | 208.2737 | 162.2733 | 207.8918 | 146.1464
2012-10-30 16:38 | 28.642 | 222.86 | 176.9204 | 222.486 | 160.993 | 70.006 | 1.0055 | 25.2647 | 25.4034 | 25.1267 | 1.5722 | 25.2064 | 39.6289 | 16.0328 | 14.8327 | 4.8262 | 65.0324 | 208.0273 | 162.0877 | 207.6533 | 146.1603
2012-10-30 16:38 | 28.6435 | 222.6134 | 176.6954 | 222.2237 | 160.8744 | 69.9713 | 1.0093 | 25.3602 | 25.5957 | 25.1269 | 1.5781 | 25.3012 | 39.928 | 16.0326 | 14.8098 | 4.8238 | 65.5238 | 207.8036 | 161.8856 | 207.4139 | 146.0646
2012-10-30 16:38 | 28.6555 | 222.3892 | 176.5981 | 221.9641 | 160.6656 | 69.9563 | 1.0077 | 25.3198 | 25.5147 | 25.1263 | 1.5756 | 25.2616 | 39.8033 | 16.0326 | 14.8014 | 4.8228 | 65.3181 | 207.5878 | 161.7967 | 207.1627 | 145.8642
2012-10-30 16:38 | 28.6502 | 222.1438 | 176.3937 | 221.7422 | 160.4628 | 69.9251 | 1.0082 | 25.3309 | 25.538 | 25.1255 | 1.5763 | 25.2717 | 39.8347 | 16.0327 | 14.7349 | 4.8206 | 65.3726 | 207.4089 | 161.6588 | 207.0073 | 145.7279
2012-10-30 16:38 | 28.6487 | 221.9297 | 176.2087 | 221.4691 | 160.2947 | 69.8869 | 1.0082 | 25.3318 | 25.5406 | 25.1248 | 1.5764 | 25.2744 | 39.843 | 16.0328 | 14.6967 | 4.818 | 65.3835 | 207.233 | 161.512 | 206.7724 | 145.598
2012-10-30 16:38 | 28.6607 | 221.7291 | 176.0762 | 221.2251 | 160.1789 | 69.8829 | 1.0087 | 25.343 | 25.5638 | 25.1241 | 1.5771 | 25.2851 | 39.8772 | 16.0326 | 14.672 | 4.8177 | 65.441 | 207.0571 | 161.4042 | 206.5531 | 145.5069
2012-10-30 16:38 | 28.6637 | 221.4289 | 175.9348 | 220.9407 | 160.1122 | 69.8361 | 1.008 | 25.3237 | 25.5255 | 25.1234 | 1.5758 | 25.2645 | 39.8118 | 16.0327 | 14.6954 | 4.8145 | 65.3373 | 206.7335 | 161.2394 | 206.2453 | 145.4168
2012-10-30 16:38 | 28.6577 | 221.1867 | 175.6907 | 220.6355 | 159.8352 | 69.8286 | 1.0079 | 25.3214 | 25.5221 | 25.1222 | 1.5757 | 25.2634 | 39.8086 | 16.0327 | 14.6935 | 4.814 | 65.3307 | 206.4932 | 160.9972 | 205.942 | 145.1417
2012-10-30 16:38 | 28.6585 | 220.967 | 175.4964 | 220.404 | 159.6153 | 69.7956 | 1.0083 | 25.3305 | 25.5407 | 25.122 | 1.5763 | 25.2715 | 39.8343 | 16.0326 | 14.7065 | 4.8117 | 65.375 | 206.2605 | 160.7899 | 205.6975 | 144.9088
2012-10-30 16:39 | 28.6652 | 220.7611 | 175.347 | 220.1823 | 159.5058 | 69.7777 | 1.0099 | 25.3713 | 25.6234 | 25.1217 | 1.5788 | 25.3127 | 39.9645 | 16.0325 | 14.6439 | 4.8105 | 65.5879 | 206.1172 | 160.7031 | 205.5384 | 144.8619
2012-10-30 16:39 | 28.672 | 220.5238 | 175.1977 | 219.9096 | 159.3152 | 69.743 | 1.0089 | 25.3447 | 25.5704 | 25.121 | 1.5772 | 25.2863 | 39.8811 | 16.0326 | 14.7078 | 4.8081 | 65.4515 | 205.816 | 160.4899 | 205.2018 | 144.6074
2012-10-30 16:39 | 28.672 | 220.3348 | 175.0143 | 219.6694 | 159.114 | 69.7032 | 1.0077 | 25.3149 | 25.5106 | 25.1206 | 1.5753 | 25.2564 | 39.7863 | 16.0327 | 14.6751 | 4.8053 | 65.2969 | 205.6597 | 160.3392 | 204.9943 | 144.4389
2012-10-30 16:39 | 28.6832 | 220.0625 | 174.8031 | 219.4719 | 158.8813 | 69.6928 | 1.009 | 25.3456 | 25.5733 | 25.12 | 1.5773 | 25.288 | 39.8864 | 16.0327 | 14.6945 | 4.8046 | 65.4596 | 205.368 | 160.1086 | 204.7774 | 144.1868
2012-10-30 16:39 | 28.6757 | 219.8936 | 174.62 | 219.2439 | 158.5372 | 69.676 | 1.0091 | 25.3496 | 25.5815 | 25.1198 | 1.5775 | 25.2913 | 39.897 | 16.0326 | 14.638 | 4.8035 | 65.4785 | 205.2556 | 159.982 | 204.6059 | 143.8992
2012-10-30 16:39 | 28.6802 | 219.6461 | 174.4334 | 219.0042 | 158.4179 | 69.6558 | 1.0096 | 25.36 | 25.6033 | 25.119 | 1.5781 | 25.3014 | 39.929 | 16.0325 | 14.6629 | 4.8021 | 65.5323 | 204.9832 | 159.7705 | 204.3413 | 143.755
2012-10-30 16:39 | 28.6802 | 219.4492 | 174.3124 | 218.7877 | 158.4256 | 69.7904 | 1.3228 | 33.2966 | 44.0451 | 25.1712 | 2.0711 | 33.1851 | 68.7291 | 16.0231 | 14.6507 | 4.8114 | 112.7742 | 204.7985 | 159.6617 | 204.137 | 143.7749
2012-10-30 16:39
2012-10-30 16:41
2012-10-30 16:44
2012-10-30 16:46
2012-10-30 16:49 |  | -10.7742 | -9.7683 | -10.2589 | -10.0034 | -1.5939 | 1.2709 | 31.5138 | 40.0497 | 24.7972 | 2.0005 | 32.0495 | 64.1153 | 16.0207 | -13.2386 | -0.1099 | 104.165 | 2.4644 | 3.4703 | 2.9797 | 3.2352
2012-10-30 16:49 | 27.2671 | 16.1803 | 17.4502 |  | 17.4218 | -1.5962 | 1.2708 | 32.1188 | 40.8166 | 25.2745 | 2.0005 | 32.0484 | 64.1114 | 16.0205 | 13.951 | -0.11 | 104.928 | 2.2293 | 3.4992 | -13.951 | 3.4708
2012-10-30 16:50 | 27.2769 | 15.6213 | 17.4397 | 241.8397 | 17.8854 | -1.598 | 1.2708 | 32.1184 | 40.8149 | 25.2748 | 2.0004 | 32.0476 | 64.109 | 16.0204 | 13.912 | -0.1102 | 104.9238 | 1.7093 | 3.5277 | 227.9277 | 3.9734
2012-10-30 16:50 | 27.2828 | 17.3118 | 18.8181 | 241.8378 | 17.4537 | -1.5974 | 1.2707 | 32.1183 | 40.8141 | 25.2751 | 2.0004 | 32.0474 | 64.1087 | 16.0202 | 13.8812 | -0.1101 | 104.9228 | 3.4306 | 4.9369 | 227.9566 | 3.5725
2012-10-30 16:50 | 27.2731 | 15.7069 |  | 242.0987 | 17.1156 | -1.6009 | 1.2707 | 32.1179 | 40.8137 | 25.2747 | 2.0004 | 32.0471 | 64.1075 | 16.0202 | 13.8796 | -0.1104 | 104.9213 | 1.8273 | -13.8796 | 228.2191 | 3.236
2012-10-30 16:50 | 27.2769 | 16.0439 | 191.2262 | 242.0947 | 18.8567 | -1.5945 | 1.2707 | 32.1174 | 40.8119 | 25.2751 | 2.0004 | 32.0468 | 64.1065 | 16.0201 | 13.8875 | -0.1099 | 104.9184 | 2.1564 | 177.3387 | 228.2072 | 4.9692
2012-10-30 16:50 | 27.2836 | 16.7326 | 191.3114 | 242.1367 |  | -1.5957 | 1.2707 | 32.117 | 40.8109 | 25.2752 | 2.0004 | 32.0463 | 64.1043 | 16.0202 | 13.9553 | -0.11 | 104.9152 | 2.7773 | 177.3561 | 228.1814 | -13.9553
2012-10-30 16:50 | 27.2814 | 16.4504 | 191.3014 | 242.3384 |  | -1.5985 | 1.2707 | 32.1165 | 40.8094 | 25.2753 | 2.0003 | 32.0458 | 64.1027 | 16.0201 | 13.945 | -0.1102 | 104.9121 | 2.5054 | 177.3564 | 228.3934 | -13.945
2012-10-30 16:50 | 27.2896 | 15.3128 | 191.3606 | 242.3584 | 172.0951 | -1.5991 | 1.2707 | 32.1164 | 40.8091 | 25.2753 | 2.0003 | 32.0456 | 64.1022 | 16.0201 | 13.9165 | -0.1102 | 104.9114 | 1.3963 | 177.4441 | 228.4419 | 158.1786
2012-10-30 16:50 | 27.2941 | 16.4306 | 191.4044 | 242.3746 | 172.1035 | -1.5985 | 1.2707 | 32.1166 | 40.8096 | 25.2753 | 2.0004 | 32.0458 | 64.1029 | 16.0201 | 13.9414 | -0.1102 | 104.9125 | 2.4892 | 177.463 | 228.4332 | 158.1621
2012-10-30 16:50 | 27.2896 | 245.1847 | 191.4745 | 242.6093 | 172.1224 | -1.5997 | 1.2706 | 32.1161 | 40.8082 | 25.2754 | 2.0003 | 32.0453 | 64.1016 | 16.0199 | 13.9287 | -0.1103 | 104.9098 | 231.256 | 177.5458 | 228.6806 | 158.1937
2012-10-30 16:50 | 27.2903 | 245.2285 | 191.5302 | 242.5826 | 172.0919 | -1.5985 | 1.2706 | 32.1153 | 40.8055 | 25.2759 | 2.0003 | 32.0446 | 64.099 | 16.0198 | 13.9172 | -0.1102 | 104.9045 | 231.3113 | 177.613 | 228.6654 | 158.1747
2012-10-30 16:50 | 27.2993 | 245.312 | 191.5078 | 242.8151 | 172.21 | -1.5939 | 1.2706 | 32.1148 | 40.8039 | 25.2761 | 2.0003 | 32.044 | 64.0972 | 16.0197 | 13.9181 | -0.1099 | 104.901 | 231.3939 | 177.5897 | 228.897 | 158.2919
2012-10-30 16:50 | 27.3038 | 245.34 | 191.5555 | 242.9215 | 172.2964 | -1.5916 | 1.2705 | 32.1146 | 40.803 | 25.2762 | 2.0003 | 32.0437 | 64.0955 | 16.0198 | 13.9592 | -0.1097 | 104.8985 | 231.3808 | 177.5963 | 228.9623 | 158.3372
2012-10-30 16:50 | 27.2993 | 245.3277 | 191.6138 | 242.8465 | 172.3621 | -1.602 | 1.2706 | 32.1145 | 40.8031 | 25.276 | 2.0003 | 32.0437 | 64.0961 | 16.0197 | 13.9588 | -0.1104 | 104.8992 | 231.3689 | 177.655 | 228.8877 | 158.4033
2012-10-30 16:50 | 27.3098 | 245.4518 | 191.6361 | 242.7629 | 172.357 | -1.5997 | 1.2706 | 32.1144 | 40.8032 | 25.2758 | 2.0003 | 32.0437 | 64.0955 | 16.0198 | 13.9775 | -0.1103 | 104.8986 | 231.4743 | 177.6586 | 228.7854 | 158.3795
2012-10-30 16:50 | 27.3068 | 245.4449 | 191.6409 | 243.0931 | 172.3228 | -1.5962 | 1.2705 | 32.1143 | 40.8027 | 25.276 | 2.0003 | 32.0437 | 64.0958 | 16.0197 | 13.9826 | -0.11 | 104.8986 | 231.4623 | 177.6583 | 229.1105 | 158.3402
2012-10-30 16:50 | 27.312 | 245.4657 | 191.6618 | 243.0474 | 172.5542 | -1.6014 | 1.2705 | 32.1138 | 40.8007 | 25.2764 | 2.0002 | 32.0428 | 64.0926 | 16.0197 | 13.9471 | -0.1104 | 104.8932 | 231.5186 | 177.7147 | 229.1003 | 158.6071
2012-10-30 16:50 | 27.3173 | 245.5141 | 191.7418 | 243.0056 | 172.5127 | -1.5991 | 1.2705 | 32.1134 | 40.8002 | 25.2761 | 2.0002 | 32.0424 | 64.0914 | 16.0196 | 13.9972 | -0.1102 | 104.8916 | 231.5169 | 177.7446 | 229.0084 | 158.5155
2012-10-30 16:50 | 27.3203 | 245.5367 | 191.7919 | 243.2988 | 172.3558 | -1.5933 | 1.2705 | 32.1132 | 40.7991 | 25.2764 | 2.0002 | 32.0422 | 64.0904 | 16.0196 | 13.9676 | -0.1098 | 104.8895 | 231.5691 | 177.8243 | 229.3312 | 158.3882
2012-10-30 16:50 | 27.3293 | 245.5849 | 191.793 | 243.0764 | 172.6378 | -1.5933 | 1.2639 | 31.9443 | 40.373 | 25.2752 | 1.9903 | 31.8836 | 63.4576 | 16.0196 | 13.9644 | -0.1098 | 103.8305 | 231.6205 | 177.8286 | 229.112 | 158.6734
2012-10-30 16:50 | 27.342 | 245.5976 | 191.7979 | 243.246 | 172.5491 | -1.6014 | 1.2643 | 31.9561 | 40.4035 | 25.2748 | 1.9904 | 31.8857 | 63.4652 | 16.0198 | 13.9404 | -0.1104 | 103.8688 | 231.6572 | 177.8575 | 229.3056 | 158.6087
2012-10-30 16:50 | 27.345 | 245.5849 | 191.8127 | 243.4019 | 172.6418 | -1.6009 | 1.2636 | 31.9385 | 40.3591 | 25.2748 | 1.9893 | 31.8684 | 63.396 | 16.0199 | 13.9311 | -0.1104 | 103.7551 | 231.6538 | 177.8816 | 229.4708 | 158.7107
2012-10-30 16:50 | 27.339 | 245.575 | 191.8891 | 243.3959 | 172.5539 | -1.5991 | 1.2632 | 31.9269 | 40.3304 | 25.2744 | 1.9885 | 31.8565 | 63.3481 | 16.02 | 13.9455 | -0.1102 | 103.6784 | 231.6295 | 177.9436 | 229.4504 | 158.6084
2012-10-30 16:50 | 27.3405 | 245.6275 | 191.8867 | 243.2641 | 172.2668 | -1.5968 | 1.263 | 31.9224 | 40.3186 | 25.2747 | 1.9883 | 31.8519 | 63.33 | 16.02 | 13.9348 | -0.1101 | 103.6486 | 231.6927 | 177.9519 | 229.3293 | 158.332
2012-10-30 16:50 | 27.3442 | 245.6233 | 191.8668 | 243.4912 | 171.9313 | -1.6014 | 1.2629 | 31.92 | 40.3121 | 25.2749 | 1.9881 | 31.8495 | 63.3202 | 16.02 | 13.9711 | -0.1104 | 103.6323 | 231.6522 | 177.8957 | 229.5201 | 157.9602
2012-10-30 16:50 | 27.351 | 245.6419 | 191.8187 | 243.2942 | 172.3124 | -1.598 | 1.2629 | 31.9199 | 40.3121 | 25.2748 | 1.9881 | 31.8496 | 63.3199 | 16.0202 | 13.9901 | -0.1102 | 103.632 | 231.6518 | 177.8286 | 229.3041 | 158.3223
2012-10-30 16:50 | 27.3472 | 245.6538 | 191.8306 | 243.3845 | 172.6128 | -1.5997 | 1.263 | 31.9215 | 40.3155 | 25.2751 | 1.9882 | 31.8511 | 63.3258 | 16.0202 | 14.0271 | -0.1103 | 103.6413 | 231.6267 | 177.8035 | 229.3574 | 158.5857
2012-10-30 16:51 | 27.348 | 245.6898 | 191.8981 | 243.1892 | 172.7852 | -1.5957 | 1.263 | 31.9228 | 40.3186 | 25.2753 | 1.9883 | 31.8523 | 63.3306 | 16.0202 | 13.9749 | -0.11 | 103.6491 | 231.7149 | 177.9232 | 229.2143 | 158.8103
2012-10-30 16:51 | 27.345 | 245.7377 | 191.9148 | 243.4646 | 172.2323 | -1.6009 | 1.2631 | 31.9242 | 40.3222 | 25.2753 | 1.9883 | 31.8536 | 63.3354 | 16.0203 | 14.0004 | -0.1104 | 103.6576 | 231.7373 | 177.9144 | 229.4642 | 158.2319
2012-10-30 16:51 | 27.354 | 245.7154 | 191.9473 | 243.4697 | 172.5494 | -1.5985 | 1.2631 | 31.9265 | 40.3278 | 25.2754 | 1.9885 | 31.856 | 63.3446 | 16.0203 | 13.9931 | -0.1102 | 103.6724 | 231.7223 | 177.9542 | 229.4766 | 158.5563
2012-10-30 16:51 | 27.3547 | 245.7122 | 191.9402 | 243.3998 | 172.6788 | -1.6032 | 1.2676 | 32.0406 | 40.6149 | 25.2764 | 1.9957 | 31.9707 | 63.8024 | 16.0202 | 14.0305 | -0.1105 | 104.4173 | 231.6817 | 177.9097 | 229.3693 | 158.6483
2012-10-30 16:51 | 27.348 | 245.7564 | 192.0042 | 243.4284 | 172.5824 | -1.6037 | 1.2685 | 32.065 | 40.6757 | 25.2771 | 1.9971 | 31.9942 | 63.8961 | 16.0202 | 14.036 | -0.1106 | 104.5718 | 231.7204 | 177.9682 | 229.3924 | 158.5464
2012-10-30 16:51 | 27.351 | 245.8143 | 192.0032 | 243.5725 | 172.6439 | -1.6014 | 1.2691 | 32.0792 | 40.7113 | 25.2774 | 1.998 | 32.0082 | 63.9518 | 16.0203 | 14.0838 | -0.1104 | 104.663 | 231.7305 | 177.9194 | 229.4887 | 158.5601
2012-10-30 16:51 | 27.354 | 245.8369 | 192.0298 | 243.5442 | 172.7678 | -1.5997 | 1.2695 | 32.0901 | 40.7393 | 25.2772 | 1.9987 | 32.0191 | 63.9966 | 16.02 | 14.095 | -0.1103 | 104.7359 | 231.7419 | 177.9348 | 229.4492 | 158.6728
2012-10-30 16:51 | 27.3607 | 245.8906 | 192.115 | 243.4372 | 172.8135 | -1.5968 | 1.2699 | 32.0989 | 40.7613 | 25.2773 | 1.9992 | 32.0276 | 64.0303 | 16.0201 | 14.0813 | -0.1101 | 104.7917 | 231.8093 | 178.0337 | 229.3559 | 158.7322
2012-10-30 16:51 | 27.3592 | 245.9243 | 192.1489 | 243.5572 | 172.695 | -1.5991 | 1.2701 | 32.1053 | 40.777 | 25.2777 | 1.9996 | 32.034 | 64.0561 | 16.02 | 14.1206 | -0.1102 | 104.8331 | 231.8037 | 178.0283 | 229.4366 | 158.5744
2012-10-30 16:51 | 27.3637 | 245.9132 | 192.1023 | 243.6205 | 172.9413 | -1.5951 | 1.2702 | 32.1091 | 40.7865 | 25.2779 | 1.9999 | 32.0379 | 64.0722 | 16.0199 | 14.0639 | -0.11 | 104.8587 | 231.8493 | 178.0384 | 229.5566 | 158.8774
2012-10-30 16:51 | 27.3675 | 245.9522 | 192.1768 | 243.4361 | 173.1441 | -1.5968 | 1.2704 | 32.1122 | 40.7942 | 25.2779 | 2.0001 | 32.0409 | 64.084 | 16.0199 | 14.1126 | -0.1101 | 104.8782 | 231.8396 | 178.0642 | 229.3235 | 159.0315
2012-10-30 16:51 | 27.3719 | 245.988 | 192.1616 | 243.6326 | 173.0705 | -1.5968 | 1.2704 | 32.1142 | 40.7992 | 25.2781 | 2.0002 | 32.043 | 64.0927 | 16.0198 | 14.0518 | -0.1101 | 104.8919 | 231.9362 | 178.1098 | 229.5808 | 159.0187
2012-10-30 16:51 | 27.3757 | 246.0623 | 192.2439 | 243.6129 | 173.1055 | -1.6003 | 1.2705 | 32.1158 | 40.803 | 25.2781 | 2.0003 | 32.0445 | 64.0991 | 16.0197 | 14.0841 | -0.1103 | 104.9021 | 231.9782 | 178.1598 | 229.5288 | 159.0214
2012-10-30 16:51 | 27.3847 | 246.087 | 192.3393 | 243.7356 | 173.0169 | -1.5962 | 1.2705 | 32.117 | 40.8058 | 25.2783 | 2.0004 | 32.0457 | 64.1039 | 16.0197 | 14.0727 | -0.11 | 104.9097 | 232.0143 | 178.2666 | 229.6629 | 158.9442
2012-10-30 16:51 | 27.3922 | 246.114 | 192.339 | 243.7039 | 173.1064 | -1.5968 | 1.2706 | 32.1182 | 40.8088 | 25.2783 | 2.0005 | 32.047 | 64.1089 | 16.0198 | 14.1087 | -0.1101 | 104.9177 | 232.0053 | 178.2303 | 229.5952 | 158.9977
2012-10-30 16:51 | 27.3974 | 246.1702 | 192.3403 | 243.7836 | 172.936 | -1.5997 | 1.2706 | 32.1186 | 40.8097 | 25.2784 | 2.0005 | 32.0473 | 64.1111 | 16.0196 | 14.0976 | -0.1103 | 104.9208 | 232.0726 | 178.2427 | 229.686 | 158.8384
2012-10-30 16:51 | 27.3952 | 246.1445 | 192.3184 | 243.9343 | 173.1094 | -1.5933 | 1.2706 | 32.1192 | 40.8113 | 25.2784 | 2.0005 | 32.048 | 64.1134 | 16.0196 | 14.1158 | -0.1098 | 104.9248 | 232.0287 | 178.2026 | 229.8185 | 158.9936
2012-10-30 16:51 | 27.4049 | 246.2012 | 192.4067 | 243.8499 | 173.396 | -1.5951 | 1.2706 | 32.1194 | 40.8114 | 25.2786 | 2.0006 | 32.0481 | 64.114 | 16.0196 | 14.1092 | -0.11 | 104.9254 | 232.092 | 178.2975 | 229.7407 | 159.2868
2012-10-30 16:51 | 27.4094 | 246.2448 | 192.4504 | 243.9014 | 173.3537 | -1.6003 | 1.2706 | 32.1193 | 40.8112 | 25.2786 | 2.0006 | 32.0479 | 64.114 | 16.0194 | 14.1382 | -0.1103 | 104.9252 | 232.1066 | 178.3122 | 229.7632 | 159.2155
2012-10-30 16:51 | 27.4184 | 246.2538 | 192.4791 | 243.8242 | 173.4485 | -1.6032 | 1.2706 | 32.1192 | 40.8109 | 25.2786 | 2.0006 | 32.0478 | 64.1136 | 16.0194 | 14.1594 | -0.1105 | 104.9245 | 232.0944 | 178.3197 | 229.6648 | 159.2891
2012-10-30 16:51 | 27.4288 | 246.3269 | 192.5602 | 244.0933 | 173.6657 | -1.5957 | 1.2706 | 32.119 | 40.8103 | 25.2786 | 2.0005 | 32.0477 | 64.113 | 16.0194 | 14.1331 | -0.11 | 104.9233 | 232.1938 | 178.4271 | 229.9602 | 159.5326
2012-10-30 16:51 | 27.4303 | 246.3363 | 192.601 | 244.2751 | 173.5307 | -1.5991 | 1.2706 | 32.1189 | 40.8096 | 25.279 | 2.0005 | 32.0476 | 64.1126 | 16.0194 | 14.1224 | -0.1102 | 104.9222 | 232.2139 | 178.4786 | 230.1527 | 159.4083
2012-10-30 16:51 | 27.4386 | 246.3093 | 192.676 | 244.1345 | 173.4726 | -1.5985 | 1.2706 | 32.1184 | 40.809 | 25.2786 | 2.0005 | 32.0471 | 64.1111 | 16.0193 | 14.1551 | -0.1102 | 104.9202 | 232.1542 | 178.5209 | 229.9794 | 159.3175
2012-10-30 16:51 | 27.4416 | 246.3711 | 192.6594 | 244.11 | 173.3664 | -1.5974 | 1.2706 | 32.1187 | 40.8093 | 25.2789 | 2.0005 | 32.0474 | 64.112 | 16.0194 | 14.1867 | -0.1101 | 104.9213 | 232.1844 | 178.4727 | 229.9233 | 159.1797
2012-10-30 16:51 | 27.4521 | 246.4207 | 192.6974 | 244.0852 | 173.2365 | -1.5997 | 1.2706 | 32.1187 | 40.8093 | 25.2788 | 2.0005 | 32.0473 | 64.1123 | 16.0193 | 14.1849 | -0.1103 | 104.9216 | 232.2358 | 178.5125 | 229.9003 | 159.0516
2012-10-30 16:51 | 27.4558 | 246.4127 | 192.6932 | 244.0811 | 173.443 | -1.5968 | 1.2706 | 32.1186 | 40.8094 | 25.2787 | 2.0006 | 32.0473 | 64.1124 | 16.0192 | 14.1683 | -0.1101 | 104.9218 | 232.2444 | 178.5249 | 229.9128 | 159.2747
2012-10-30 16:52 | 27.4625 | 246.4586 | 192.751 | 244.1427 | 173.6136 | -1.5974 | 1.2706 | 32.1188 | 40.8093 | 25.279 | 2.0006 | 32.0474 | 64.1129 | 16.0192 | 14.1587 | -0.1101 | 104.9222 | 232.2999 | 178.5923 | 229.984 | 159.4549
2012-10-30 16:52 | 27.4753 | 246.5223 | 192.7913 | 244.2809 | 173.4859 | -1.6014 | 1.2706 | 32.1187 | 40.8086 | 25.2792 | 2.0006 | 32.0474 | 64.113 | 16.0192 | 14.1633 | -0.1104 | 104.9216 | 232.359 | 178.628 | 230.1176 | 159.3226
2012-10-30 16:52 | 27.4723 | 246.5193 | 192.7372 | 244.3681 | 173.2567 | -1.6026 | 1.2706 | 32.1186 | 40.809 | 25.2788 | 2.0006 | 32.0472 | 64.1126 | 16.0191 | 14.1196 | -0.1105 | 104.9216 | 232.3997 | 178.6176 | 230.2485 | 159.1371
2012-10-30 16:52 | 27.4813 | 246.6066 | 192.7423 | 244.428 | 173.3437 | -1.5957 | 1.2705 | 32.118 | 40.8072 | 25.2791 | 2.0005 | 32.0467 | 64.1105 | 16.019 | 14.1856 | -0.11 | 104.9178 | 232.421 | 178.5567 | 230.2424 | 159.1581
2012-10-30 16:52 | 27.4925 | 246.6335 | 192.7889 | 244.2393 | 173.3315 | -1.6043 | 1.2705 | 32.1178 | 40.8071 | 25.2788 | 2.0005 | 32.0465 | 64.1095 | 16.0191 | 14.1398 | -0.1106 | 104.9166 | 232.4937 | 178.6491 | 230.0995 | 159.1917
2012-10-30 16:52 | 27.491 | 246.632 | 192.7559 | 244.328 | 173.3651 | -1.5974 | 1.2705 | 32.118 | 40.807 | 25.2792 | 2.0005 | 32.0467 | 64.1107 | 16.019 | 14.1301 | -0.1101 | 104.9177 | 232.5019 | 178.6258 | 230.1979 | 159.235
2012-10-30 16:52 | 27.497 | 246.6772 | 192.8051 | 244.3653 | 173.4803 | -1.5962 | 1.2705 | 32.1177 | 40.806 | 25.2793 | 2.0005 | 32.0463 | 64.1094 | 16.019 | 14.0913 | -0.11 | 104.9153 | 232.5859 | 178.7138 | 230.274 | 159.389
2012-10-30 16:52 | 27.5007 | 246.6456 | 192.8638 | 244.4592 | 173.605 | -1.6014 | 1.2705 | 32.1171 | 40.8045 | 25.2793 | 2.0005 | 32.0457 | 64.1066 | 16.0191 | 14.1439 | -0.1104 | 104.911 | 232.5017 | 178.7199 | 230.3153 | 159.4611
2012-10-30 16:52 | 27.5172 | 246.7561 | 192.9196 | 244.4208 | 173.6683 | -1.6032 | 1.2705 | 32.117 | 40.8042 | 25.2794 | 2.0005 | 32.0456 | 64.1067 | 16.0189 | 14.1319 | -0.1105 | 104.911 | 232.6242 | 178.7877 | 230.2889 | 159.5364
2012-10-30 16:52 | 27.5142 | 246.7884 | 192.9205 | 244.4492 | 173.599 | -1.6003 | 1.2705 | 32.1169 | 40.8034 | 25.2796 | 2.0005 | 32.0455 | 64.1065 | 16.0189 | 14.0922 | -0.1103 | 104.9099 | 232.6962 | 178.8283 | 230.357 | 159.5068
2012-10-30 16:52 | 27.5254 | 246.78 | 192.9042 | 244.5192 | 173.5243 | -1.6003 | 1.2705 | 32.1168 | 40.8037 | 25.2794 | 2.0005 | 32.0455 | 64.1069 | 16.0188 | 14.0952 | -0.1103 | 104.9105 | 232.6848 | 178.809 | 230.424 | 159.4291
2012-10-30 16:52 | 27.5239 | 246.8177 | 192.9106 | 244.5765 | 173.515 | -1.5991 | 1.2705 | 32.1165 | 40.8027 | 25.2794 | 2.0005 | 32.0451 | 64.1058 | 16.0187 | 14.1223 | -0.1102 | 104.9085 | 232.6954 | 178.7883 | 230.4542 | 159.3927
2012-10-30 16:52 | 27.5307 | 246.8636 | 192.9724 | 244.6264 | 173.5803 | -1.6032 | 1.2705 | 32.1166 | 40.8026 | 25.2796 | 2.0005 | 32.0452 | 64.1055 | 16.0188 | 14.1128 | -0.1105 | 104.9081 | 232.7508 | 178.8596 | 230.5136 | 159.4675
2012-10-30 16:52 | 27.5352 | 246.876 | 192.9415 | 244.576 | 173.4717 | -1.5957 | 1.2705 | 32.1167 | 40.8031 | 25.2795 | 2.0005 | 32.0453 | 64.1063 | 16.0187 | 14.1213 | -0.11 | 104.9093 | 232.7547 | 178.8202 | 230.4547 | 159.3504
2012-10-30 16:52 | 27.5456 | 246.9334 | 192.9951 | 244.7784 | 173.4977 | -1.6003 | 1.2704 | 32.1161 | 40.8009 | 25.2799 | 2.0004 | 32.0446 | 64.1035 | 16.0188 | 14.0788 | -0.1103 | 104.9044 | 232.8546 | 178.9163 | 230.6996 | 159.4189
2012-10-30 16:52 | 27.5546 | 246.9306 | 192.9924 | 244.6816 | 173.7369 | -1.5991 | 1.2704 | 32.1156 | 40.7998 | 25.2798 | 2.0004 | 32.0442 | 64.1025 | 16.0186 | 14.1122 | -0.1102 | 104.9023 | 232.8184 | 178.8802 | 230.5694 | 159.6247
2012-10-30 16:52 | 27.5509 | 246.9543 | 193.0122 | 244.6348 | 173.8034 | -1.5991 | 1.2704 | 32.1157 | 40.8002 | 25.2797 | 2.0005 | 32.0443 | 64.1031 | 16.0185 | 14.0637 | -0.1102 | 104.9033 | 232.8906 | 178.9485 | 230.5711 | 159.7397
2012-10-30 16:52 | 27.5509 | 246.97 | 193.0397 | 244.7837 | 173.8971 | -1.5974 | 1.2704 | 32.1155 | 40.7999 | 25.2797 | 2.0004 | 32.0441 | 64.1025 | 16.0185 | 14.0148 | -0.1101 | 104.9023 | 232.9552 | 179.0249 | 230.7689 | 159.8823
2012-10-30 16:52 | 27.5689 | 247.0037 | 193.1324 | 244.739 | 173.8917 | -1.5968 | 1.2704 | 32.1152 | 40.7985 | 25.28 | 2.0004 | 32.0438 | 64.1015 | 16.0185 | 14.0287 | -0.1101 | 104.9001 | 232.975 | 179.1037 | 230.7103 | 159.863
2012-10-30 16:52 | 27.5704 | 247.0443 | 193.126 | 244.4897 | 173.9829 | -1.6014 | 1.2704 | 32.1152 | 40.7988 | 25.2799 | 2.0004 | 32.0438 | 64.1018 | 16.0184 | 13.9976 | -0.1104 | 104.9005 | 233.0467 | 179.1284 | 230.4921 | 159.9853
2012-10-30 16:52 | 27.5801 | 247.0658 | 193.1671 | 244.715 | 173.9692 | -1.6003 | 1.2704 | 32.1149 | 40.7982 | 25.2798 | 2.0004 | 32.0435 | 64.1006 | 16.0184 | 14.0318 | -0.1103 | 104.8987 | 233.034 | 179.1353 | 230.6832 | 159.9374
2012-10-30 16:52 | 27.5733 | 247.0825 | 193.2075 | 244.7865 | 173.9702 | -1.5922 | 1.2704 | 32.1146 | 40.7971 | 25.28 | 2.0004 | 32.0431 | 64.0995 | 16.0182 | 13.976 | -0.1098 | 104.8966 | 233.1065 | 179.2315 | 230.8105 | 159.9942
2012-10-30 16:52 | 27.5763 | 247.0776 | 193.2105 | 244.8287 | 174.0786 | -1.5957 | 1.2704 | 32.1146 | 40.797 | 25.28 | 2.0004 | 32.0431 | 64.0993 | 16.0183 | 14.0035 | -0.11 | 104.8963 | 233.0741 | 179.207 | 230.8252 | 160.0751
2012-10-30 16:52 | 27.5741 | 247.1499 | 193.224 | 244.7952 | 173.9515 | -1.6014 | 1.2704 | 32.1146 | 40.7968 | 25.2802 | 2.0004 | 32.0432 | 64.0999 | 16.0182 | 14.0094 | -0.1104 | 104.8967 | 233.1405 | 179.2146 | 230.7858 | 159.9421
2012-10-30 16:52 | 27.5891 | 247.157 | 193.2508 | 244.963 | 173.7987 | -1.5945 | 1.2704 | 32.1144 | 40.7967 | 25.2799 | 2.0004 | 32.043 | 64.099 | 16.0182 | 14.0122 | -0.1099 | 104.8958 | 233.1448 | 179.2386 | 230.9508 | 159.7865
2012-10-30 16:52 | 27.5831 | 247.2019 | 193.2998 | 244.8316 | 173.9995 | -1.5968 | 1.2703 | 32.1141 | 40.7955 | 25.2801 | 2.0004 | 32.0427 | 64.098 | 16.0182 | 13.994 | -0.1101 | 104.8935 | 233.2079 | 179.3058 | 230.8376 | 160.0055
2012-10-30 16:52 | 27.5913 | 247.2336 | 193.3158 | 244.6635 | 174.0662 | -1.5962 | 1.2703 | 32.1136 | 40.7948 | 25.2798 | 2.0004 | 32.0421 | 64.0955 | 16.0182 | 13.9859 | -0.11 | 104.8903 | 233.2477 | 179.3299 | 230.6776 | 160.0803
2012-10-30 16:53 | 27.5921 | 247.207 | 193.3363 | 244.9072 | 173.8758 | -1.5939 | 1.2703 | 32.1134 | 40.7942 | 25.2799 | 2.0004 | 32.0422 | 64.0963 | 16.0181 | 14.0071 | -0.1099 | 104.8904 | 233.1999 | 179.3292 | 230.9001 | 159.8687
2012-10-30 16:53 | 27.6018 | 247.2441 | 193.3695 | 244.9992 | 174.0182 | -1.6032 | 1.2703 | 32.1131 | 40.7933 | 25.28 | 2.0004 | 32.0417 | 64.0947 | 16.018 | 13.9964 | -0.1105 | 104.888 | 233.2477 | 179.3731 | 231.0028 | 160.0218
2012-10-30 16:53 | 27.6085 | 247.3096 | 193.3723 | 244.9511 | 174.0522 | -1.6003 | 1.2703 | 32.1125 | 40.7911 | 25.2804 | 2.0003 | 32.041 | 64.0924 | 16.0179 | 14.0031 | -0.1103 | 104.8835 | 233.3065 | 179.3692 | 230.948 | 160.0491
2012-10-30 16:53 | 27.6093 | 247.3182 | 193.4006 | 244.9323 | 174.0335 | -1.5957 | 1.2702 | 32.112 | 40.7902 | 25.2801 | 2.0003 | 32.0405 | 64.0905 | 16.0179 | 13.9957 | -0.11 | 104.8807 | 233.3225 | 179.4049 | 230.9366 | 160.0378
2012-10-30 16:53 | 27.6093 | 247.33 | 193.4477 | 245.0773 | 174.0179 | -1.6009 | 1.2702 | 32.1118 | 40.7893 | 25.2803 | 2.0003 | 32.0403 | 64.089 | 16.018 | 14.012 | -0.1104 | 104.8783 | 233.318 | 179.4357 | 231.0653 | 160.0059
2012-10-30 16:53 | 27.6153 | 247.3595 | 193.4773 | 244.954 | 174.098 | -1.598 | 1.2694 | 32.0903 | 40.735 | 25.2801 | 1.9989 | 32.0175 | 63.9983 | 16.0179 | 14.0099 | -0.1102 | 104.7333 | 233.3496 | 179.4674 | 230.9441 | 160.0881
2012-10-30 16:53 | 27.6265 | 247.3942 | 193.4649 | 245.0631 | 174.0585 | -1.598 | 1.2638 | 31.9477 | 40.3762 | 25.2787 | 1.9901 | 31.8776 | 63.44 | 16.0179 | 14.017 | -0.1102 | 103.8162 | 233.3772 | 179.4479 | 231.0461 | 160.0415
2012-10-30 16:53 | 27.625 | 247.3966 | 193.4595 | 245.1949 | 174.0765 | -1.602 | 1.2638 | 31.948 | 40.3768 | 25.2787 | 1.9901 | 31.8772 | 63.4384 | 16.018 | 14.0033 | -0.1104 | 103.8152 | 233.3933 | 179.4562 | 231.1916 | 160.0732
2012-10-30 16:53 | 27.6235 | 247.3677 | 193.4737 | 245.0954 | 174.2389 | -1.602 | 1.2631 | 31.9295 | 40.3299 | 25.2788 | 1.9889 | 31.8586 | 63.3636 | 16.0181 | 13.9855 | -0.1104 | 103.6935 | 233.3822 | 179.4882 | 231.1099 | 160.2534
2012-10-30 16:53 | 27.6235 | 247.3912 | 193.4541 | 245.1816 | 173.759 | -1.6032 | 1.2628 | 31.9223 | 40.3119 | 25.2787 | 1.9884 | 31.8512 | 63.3345 | 16.0181 | 14.014 | -0.1105 | 103.6464 | 233.3772 | 179.4401 | 231.1676 | 159.745
2012-10-30 16:53 | 27.6303 | 247.3392 | 193.4295 | 245.2002 | 173.2898 | -1.6026 | 1.2628 | 31.9209 | 40.3083 | 25.2787 | 1.9883 | 31.8498 | 63.3282 | 16.0183 | 14.0086 | -0.1105 | 103.6365 | 233.3306 | 179.4209 | 231.1916 | 159.2812
2012-10-30 16:53 | 27.6318 | 247.3995 | 193.3799 | 245.3232 | 174.0404 | -1.5968 | 1.2628 | 31.9207 | 40.3084 | 25.2784 | 1.9883 | 31.8497 | 63.3279 | 16.0183 | 13.9938 | -0.1101 | 103.6363 | 233.4057 | 179.3861 | 231.3294 | 160.0466
2012-10-30 16:53 | 27.6407 | 247.424 | 193.4241 | 245.0852 | 173.9049 | -1.6003 | 1.2627 | 31.9205 | 40.3075 | 25.2787 | 1.9883 | 31.8496 | 63.3273 | 16.0183 | 14.0108 | -0.1103 | 103.6348 | 233.4132 | 179.4133 | 231.0744 | 159.8941
2012-10-30 16:53 | 27.6422 | 247.4373 | 193.4139 | 245.1533 | 173.7699 | -1.5974 | 1.2628 | 31.9208 | 40.3082 | 25.2787 | 1.9883 | 31.8499 | 63.3277 | 16.0185 | 14.0449 | -0.1101 | 103.6359 | 233.3924 | 179.369 | 231.1084 | 159.725
2012-10-30 16:53 | 27.6303 | 247.4097 | 193.4216 | 245.1179 | 174.0974 | -1.5997 | 1.2628 | 31.9227 | 40.3124 | 25.2791 | 1.9884 | 31.8517 | 63.335 | 16.0184 | 14.0656 | -0.1103 | 103.6474 | 233.3441 | 179.356 | 231.0523 | 160.0318
2012-10-30 16:53 | 27.6377 | 247.4249 | 193.4722 | 244.9724 | 174.5575 | -1.5974 | 1.2629 | 31.9244 | 40.317 | 25.2788 | 1.9885 | 31.8533 | 63.3416 | 16.0185 | 14.0282 | -0.1101 | 103.6586 | 233.3967 | 179.444 | 230.9442 | 160.5293
2012-10-30 16:53 | 27.643 | 247.4616 | 193.5168 | 245.2168 | 174.3326 | -1.6049 | 1.2656 | 31.9934 | 40.4908 | 25.2793 | 1.9928 | 31.9216 | 63.614 | 16.0184 | 14.0947 | -0.1106 | 104.1048 | 233.3669 | 179.4221 | 231.1221 | 160.2379
2012-10-30 16:53 | 27.6407 | 247.4123 | 193.5184 | 245.2223 | 174.1742 | -1.6032 | 1.268 | 32.0555 | 40.6461 | 25.2805 | 1.9967 | 31.9844 | 63.8643 | 16.0184 | 14.0394 | -0.1105 | 104.5104 | 233.3729 | 179.479 | 231.1829 | 160.1348
2012-10-30 16:53 | 27.6392 | 247.4304 | 193.5876 | 245.1895 | 174.4771 | -1.5939 | 1.2686 | 32.071 | 40.6854 | 25.2805 | 1.9977 | 31.9999 | 63.9266 | 16.0183 | 14.0338 | -0.1099 | 104.6121 | 233.3966 | 179.5538 | 231.1557 | 160.4433
2012-10-30 16:53 | 27.6377 | 247.5072 | 193.6215 | 245.09 | 174.2726 | -1.5951 | 1.2691 | 32.0839 | 40.7175 | 25.2809 | 1.9985 | 32.0123 | 63.9765 | 16.0182 | 14.0934 | -0.11 | 104.6939 | 233.4138 | 179.5281 | 230.9966 | 160.1792
2012-10-30 16:53 | 27.64 | 247.4742 | 193.6749 | 245.1825 | 174.193 | -1.598 | 1.2695 | 32.0941 | 40.7438 | 25.2806 | 1.9991 | 32.0224 | 64.0172 | 16.0181 | 14.002 | -0.1102 | 104.761 | 233.4722 | 179.6729 | 231.1805 | 160.191
2012-10-30 16:53 | 27.6527 | 247.5183 | 193.6836 | 245.1756 | 174.4671 | -1.6026 | 1.2698 | 32.1015 | 40.7616 | 25.2812 | 1.9996 | 32.0301 | 64.0483 | 16.018 | 14.0473 | -0.1105 | 104.8099 | 233.471 | 179.6363 | 231.1283 | 160.4198
2012-10-30 16:53 | 27.6505 | 247.5513 | 193.6657 | 245.1616 | 174.4298 | -1.5951 | 1.27 | 32.107 | 40.7762 | 25.281 | 2 | 32.0355 | 64.0701 | 16.018 | 14.0492 | -0.11 | 104.8462 | 233.5021 | 179.6165 | 231.1124 | 160.3806
2012-10-30 16:53 | 27.643 | 247.536 | 193.7014 | 245.2873 | 174.235 | -1.5997 | 1.2701 | 32.1103 | 40.7841 | 25.2812 | 2.0002 | 32.0387 | 64.0825 | 16.018 | 14.0661 | -0.1103 | 104.8667 | 233.4699 | 179.6353 | 231.2212 | 160.1689
2012-10-30 16:53 | 27.655 | 247.595 | 193.7802 | 245.1583 | 174.3524 | -1.602 | 1.2702 | 32.1124 | 40.7895 | 25.2812 | 2.0003 | 32.0409 | 64.0919 | 16.018 | 14.0374 | -0.1104 | 104.8814 | 233.5576 | 179.7428 | 231.1209 | 160.315
2012-10-30 16:53 | 27.6572 | 247.652 | 193.7196 | 245.329 | 174.1673 | -1.5962 | 1.2702 | 32.1137 | 40.7923 | 25.2815 | 2.0004 | 32.0421 | 64.0967 | 16.0179 | 14.0518 | -0.11 | 104.889 | 233.6002 | 179.6678 | 231.2772 | 160.1155
2012-10-30 16:54 | 27.6639 | 247.6665 | 193.7145 | 245.1358 | 174.0764 | -1.6026 | 1.2703 | 32.1146 | 40.7944 | 25.2816 | 2.0005 | 32.0431 | 64.1006 | 16.0179 | 14.0422 | -0.1105 | 104.895 | 233.6243 | 179.6723 | 231.0936 | 160.0342
2012-10-30 16:54 | 27.6707 | 247.6459 | 193.7291 | 245.3307 | 174.3017 | -1.5957 | 1.2703 | 32.1157 | 40.7972 | 25.2817 | 2.0005 | 32.0441 | 64.1051 | 16.0178 | 14.0734 | -0.11 | 104.9023 | 233.5725 | 179.6557 | 231.2573 | 160.2283
2012-10-30 16:54 | 27.6684 | 247.6632 | 193.7426 | 245.2148 | 174.6077 | -1.5928 | 1.2704 | 32.1167 | 40.7996 | 25.2816 | 2.0006 | 32.045 | 64.109 | 16.0178 | 14.0589 | -0.1098 | 104.9086 | 233.6043 | 179.6837 | 231.1559 | 160.5488
2012-10-30 16:54 | 27.6737 | 247.6646 | 193.7596 | 245.4434 | 174.4101 | -1.6003 | 1.2704 | 32.1172 | 40.8011 | 25.2816 | 2.0006 | 32.0455 | 64.111 | 16.0178 | 14.0397 | -0.1103 | 104.9122 | 233.6249 | 179.7199 | 231.4037 | 160.3704
2012-10-30 16:54 | 27.6707 | 247.642 | 193.7802 | 245.2837 | 174.3329 | -1.5997 | 1.2704 | 32.1179 | 40.8029 | 25.2816 | 2.0007 | 32.0463 | 64.1145 | 16.0177 | 14.0857 | -0.1103 | 104.9174 | 233.5563 | 179.6945 | 231.198 | 160.2472
2012-10-30 16:54 | 27.6797 | 247.6784 | 193.7892 | 245.3005 | 174.4629 | -1.6003 | 1.2704 | 32.118 | 40.8028 | 25.2818 | 2.0007 | 32.0463 | 64.1149 | 16.0176 | 14.0743 | -0.1103 | 104.9177 | 233.6041 | 179.7149 | 231.2262 | 160.3886
2012-10-30 16:54 | 27.6842 | 247.726 | 193.8409 | 245.3599 | 174.4401 | -1.5962 | 1.2704 | 32.1179 | 40.8027 | 25.2817 | 2.0007 | 32.0463 | 64.1148 | 16.0176 | 14.0421 | -0.11 | 104.9175 | 233.6839 | 179.7988 | 231.3178 | 160.398
2012-10-30 16:54 | 27.6834 | 247.7409 | 193.8243 | 245.3669 | 174.4939 | -1.6009 | 1.2704 | 32.1181 | 40.8027 | 25.2819 | 2.0007 | 32.0464 | 64.115 | 16.0176 | 14.0739 | -0.1104 | 104.9177 | 233.667 | 179.7504 | 231.293 | 160.42
2012-10-30 16:54 | 27.7006 | 247.7228 | 193.8965 | 245.3293 | 174.519 | -1.5997 | 1.2704 | 32.1184 | 40.8034 | 25.282 | 2.0007 | 32.0467 | 64.1164 | 16.0175 | 14.0911 | -0.1103 | 104.9198 | 233.6317 | 179.8054 | 231.2382 | 160.4279
2012-10-30 16:54 | 27.6954 | 247.7411 | 193.8796 | 245.375 | 174.5411 | -1.6003 | 1.2704 | 32.118 | 40.8023 | 25.2821 | 2.0007 | 32.0464 | 64.1156 | 16.0175 | 14.0818 | -0.1103 | 104.9179 | 233.6593 | 179.7978 | 231.2932 | 160.4593
2012-10-30 16:54 | 27.6924 | 247.8008 | 193.908 | 245.4935 | 174.5732 | -1.6037 | 1.2704 | 32.1178 | 40.8022 | 25.2817 | 2.0007 | 32.0461 | 64.1149 | 16.0174 | 14.0748 | -0.1106 | 104.9171 | 233.726 | 179.8332 | 231.4187 | 160.4984
2012-10-30 16:54 | 27.7074 | 247.8197 | 193.9583 | 245.5986 | 174.6311 | -1.5962 | 1.2704 | 32.1181 | 40.8024 | 25.2821 | 2.0007 | 32.0464 | 64.1165 | 16.0173 | 14.0816 | -0.11 | 104.9189 | 233.7381 | 179.8767 | 231.517 | 160.5495
2012-10-30 16:54 | 27.7059 | 247.8652 | 193.9686 | 245.4835 | 174.7194 | -1.5962 | 1.2704 | 32.1183 | 40.8029 | 25.2821 | 2.0007 | 32.0466 | 64.1172 | 16.0173 | 14.1127 | -0.11 | 104.9202 | 233.7525 | 179.8559 | 231.3708 | 160.6067
2012-10-30 16:54 | 27.7029 | 247.8622 | 193.9971 | 245.5001 | 174.5056 | -1.6026 | 1.2704 | 32.1184 | 40.804 | 25.2816 | 2.0008 | 32.0466 | 64.1178 | 16.0172 | 14.0567 | -0.1105 | 104.9218 | 233.8055 | 179.9404 | 231.4434 | 160.4489
2012-10-30 16:54 | 27.7186 | 247.8622 | 193.9931 | 245.4021 | 174.5018 | -1.5997 | 1.2704 | 32.1184 | 40.8031 | 25.2822 | 2.0008 | 32.0467 | 64.118 | 16.0172 | 14.0928 | -0.1103 | 104.9211 | 233.7694 | 179.9003 | 231.3093 | 160.409
2012-10-30 16:54 | 27.7134 | 247.9706 | 194.0076 | 245.5184 | 174.7152 | -1.6032 | 1.2704 | 32.1183 | 40.8032 | 25.282 | 2.0008 | 32.0467 | 64.1181 | 16.0171 | 14.055 | -0.1105 | 104.9213 | 233.9156 | 179.9526 | 231.4634 | 160.6602
2012-10-30 16:54 | 27.7276 | 247.9104 | 194.0768 | 245.5404 | 174.827 | -1.5951 | 1.2704 | 32.118 | 40.8023 | 25.2821 | 2.0007 | 32.0463 | 64.1165 | 16.0172 | 14.061 | -0.11 | 104.9188 | 233.8494 | 180.0158 | 231.4794 | 160.766
2012-10-30 16:54 | 27.7224 | 247.8738 | 194.1226 | 245.5901 | 174.7983 | -1.6066 | 1.2704 | 32.1181 | 40.8025 | 25.2822 | 2.0008 | 32.0464 | 64.1174 | 16.017 | 14.0395 | -0.1108 | 104.9199 | 233.8343 | 180.0831 | 231.5506 | 160.7588
2012-10-30 16:54 | 27.7343 | 247.8897 | 194.1542 | 245.7352 | 174.9117 | -1.5957 | 1.2704 | 32.118 | 40.8024 | 25.282 | 2.0008 | 32.0462 | 64.1171 | 16.017 | 14.0759 | -0.11 | 104.9195 | 233.8138 | 180.0783 | 231.6593 | 160.8358
2012-10-30 16:54 | 27.7388 | 247.9764 | 194.1862 | 245.5556 | 174.9084 | -1.5962 | 1.2704 | 32.1181 | 40.8025 | 25.282 | 2.0008 | 32.0463 | 64.1172 | 16.0171 | 14.0804 | -0.11 | 104.9196 | 233.896 | 180.1058 | 231.4752 | 160.828
2012-10-30 16:54 | 27.7366 | 248.0251 | 194.2469 | 245.573 | 174.8164 | -1.6032 | 1.2704 | 32.1179 | 40.8025 | 25.2818 | 2.0008 | 32.0463 | 64.1173 | 16.017 | 14.07 | -0.1105 | 104.9199 | 233.9551 | 180.1769 | 231.503 | 160.7464
2012-10-30 16:54 | 27.7411 | 248.0296 | 194.216 | 245.7773 | 174.8756 | -1.598 | 1.2704 | 32.1178 | 40.8016 | 25.2822 | 2.0008 | 32.046 | 64.1168 | 16.0168 | 14.0542 | -0.1102 | 104.9183 | 233.9754 | 180.1618 | 231.7231 | 160.8214
2012-10-30 16:54 | 27.7381 | 247.9992 | 194.2562 | 245.7821 | 174.8843 | -1.5974 | 1.2704 | 32.1175 | 40.801 | 25.282 | 2.0008 | 32.0457 | 64.1158 | 16.0168 | 14.0389 | -0.1101 | 104.9169 | 233.9603 | 180.2173 | 231.7432 | 160.8454
2012-10-30 16:54 | 27.7471 | 248.0356 | 194.3124 | 245.7049 | 174.7606 | -1.5962 | 1.2703 | 32.1173 | 40.8001 | 25.2823 | 2.0008 | 32.0456 | 64.1155 | 16.0167 | 14.1009 | -0.11 | 104.9156 | 233.9347 | 180.2115 | 231.604 | 160.6597
2012-10-30 16:54 | 27.7523 | 248.0682 | 194.2822 | 245.7023 | 174.7463 | -1.5939 | 1.2703 | 32.1172 | 40.8 | 25.2823 | 2.0007 | 32.0455 | 64.1145 | 16.0168 | 14.1183 | -0.1099 | 104.9145 | 233.9499 | 180.1639 | 231.584 | 160.628
2012-10-30 16:55 | 27.7546 | 248.0627 | 194.2609 | 245.8613 | 174.7642 | -1.6014 | 1.2703 | 32.117 | 40.7991 | 25.2824 | 2.0007 | 32.0452 | 64.114 | 16.0167 | 14.0717 | -0.1104 | 104.9131 | 233.991 | 180.1892 | 231.7896 | 160.6925
2012-10-30 16:55 | 27.7568 | 248.0688 | 194.2867 | 245.7185 | 174.9264 | -1.5985 | 1.2703 | 32.1165 | 40.7984 | 25.2821 | 2.0007 | 32.0448 | 64.1127 | 16.0166 | 14.0658 | -0.1102 | 104.9112 | 234.003 | 180.2209 | 231.6527 | 160.8606
2012-10-30 16:55 | 27.7583 | 248.129 | 194.3354 | 245.7005 | 174.9006 | -1.6055 | 1.2703 | 32.1164 | 40.7978 | 25.2824 | 2.0007 | 32.0446 | 64.1118 | 16.0167 | 14.0673 | -0.1107 | 104.9096 | 234.0617 | 180.2681 | 231.6332 | 160.8333
2012-10-30 16:55 | 27.759 | 248.1415 | 194.3085 | 245.8109 | 174.7569 | -1.602 | 1.2703 | 32.1163 | 40.7975 | 25.2823 | 2.0007 | 32.0444 | 64.1116 | 16.0165 | 14.0476 | -0.1104 | 104.9091 | 234.0939 | 180.2609 | 231.7633 | 160.7093
2012-10-30 16:55 | 27.7628 | 248.157 | 194.2966 | 245.9283 | 174.6085 | -1.6009 | 1.2703 | 32.1159 | 40.7966 | 25.2823 | 2.0007 | 32.0441 | 64.1104 | 16.0165 | 14.0555 | -0.1104 | 104.907 | 234.1015 | 180.2411 | 231.8728 | 160.553
2012-10-30 16:55 | 27.7703 | 248.2037 | 194.2963 | 245.8378 | 174.9009 | -1.6032 | 1.2703 | 32.1154 | 40.7948 | 25.2826 | 2.0007 | 32.0437 | 64.1088 | 16.0165 | 14.0507 | -0.1105 | 104.9036 | 234.153 | 180.2456 | 231.7871 | 160.8502
2012-10-30 16:55 | 27.7785 | 248.208 | 194.3713 | 245.8421 | 174.9403 | -1.6014 | 1.2703 | 32.1154 | 40.795 | 25.2825 | 2.0007 | 32.0436 | 64.1084 | 16.0165 | 14.1038 | -0.1104 | 104.9034 | 234.1042 | 180.2675 | 231.7383 | 160.8365
2012-10-30 16:55 | 27.7733 | 248.2537 | 194.3425 | 246.072 | 174.8649 | -1.602 | 1.2702 | 32.1152 | 40.7942 | 25.2827 | 2.0007 | 32.0434 | 64.108 | 16.0164 | 14.0904 | -0.1104 | 104.9022 | 234.1633 | 180.2521 | 231.9816 | 160.7745
2012-10-30 16:55 | 27.7763 | 248.2527 | 194.3966 | 246.0084 | 174.9732 | -1.602 | 1.2702 | 32.115 | 40.7938 | 25.2826 | 2.0006 | 32.0432 | 64.107 | 16.0164 | 14.069 | -0.1104 | 104.9009 | 234.1837 | 180.3276 | 231.9394 | 160.9042
2012-10-30 16:55 | 27.7823 | 248.2392 | 194.3633 | 245.9086 | 174.9285 | -1.6003 | 1.2702 | 32.1149 | 40.7939 | 25.2824 | 2.0006 | 32.043 | 64.1068 | 16.0163 | 14.0831 | -0.1103 | 104.9007 | 234.1561 | 180.2802 | 231.8255 | 160.8454
2012-10-30 16:55 | 27.7838 | 248.2877 | 194.3766 | 245.9336 | 174.8597 | -1.6014 | 1.2702 | 32.1146 | 40.7933 | 25.2822 | 2.0006 | 32.0428 | 64.1062 | 16.0163 | 14.1009 | -0.1104 | 104.8995 | 234.1868 | 180.2757 | 231.8327 | 160.7588
2012-10-30 16:55 | 27.7882 | 248.2764 | 194.3928 | 246.0556 | 174.8602 | -1.6026 | 1.2702 | 32.1144 | 40.7924 | 25.2825 | 2.0006 | 32.0426 | 64.1055 | 16.0162 | 14.1257 | -0.1105 | 104.8979 | 234.1507 | 180.2671 | 231.9299 | 160.7345
2012-10-30 16:55 | 27.7912 | 248.3264 | 194.4115 | 245.9058 | 174.9101 | -1.5985 | 1.2702 | 32.1141 | 40.7914 | 25.2827 | 2.0006 | 32.0423 | 64.1039 | 16.0163 | 14.0431 | -0.1102 | 104.8953 | 234.2833 | 180.3684 | 231.8627 | 160.867
2012-10-30 16:55 | 27.7935 | 248.3091 | 194.3863 | 245.8806 | 174.7641 | -1.6003 | 1.2702 | 32.1138 | 40.7908 | 25.2826 | 2.0006 | 32.0421 | 64.1035 | 16.0162 | 14.1025 | -0.1103 | 104.8944 | 234.2066 | 180.2838 | 231.7781 | 160.6616
2012-10-30 16:55 | 27.7972 | 248.3206 | 194.386 | 246.0489 | 174.7053 | -1.5939 | 1.2702 | 32.1135 | 40.7902 | 25.2825 | 2.0006 | 32.0417 | 64.1023 | 16.0161 | 14.0613 | -0.1099 | 104.8925 | 234.2593 | 180.3247 | 231.9876 | 160.644
2012-10-30 16:55 | 27.7995 | 248.3308 | 194.3569 | 246.0512 | 174.7818 | -1.5962 | 1.2702 | 32.1132 | 40.7891 | 25.2826 | 2.0006 | 32.0414 | 64.1013 | 16.016 | 14.0759 | -0.11 | 104.8905 | 234.2549 | 180.281 | 231.9753 | 160.7059
2012-10-30 16:55 | 27.8047 | 248.3751 | 194.3739 | 245.8409 | 174.8963 | -1.5997 | 1.2702 | 32.1131 | 40.7887 | 25.2827 | 2.0006 | 32.0413 | 64.1006 | 16.0162 | 14.0444 | -0.1103 | 104.8892 | 234.3307 | 180.3295 | 231.7965 | 160.8519
2012-10-30 16:55 | 27.8122 | 248.3395 | 194.4207 | 246.0443 | 174.9974 | -1.602 | 1.2701 | 32.1127 | 40.7878 | 25.2827 | 2.0006 | 32.0408 | 64.0993 | 16.016 | 14.0763 | -0.1104 | 104.8871 | 234.2632 | 180.3444 | 231.968 | 160.9211
2012-10-30 16:55 | 27.8182 | 248.3651 | 194.4188 | 246.1443 | 174.9917 | -1.6014 | 1.2702 | 32.1127 | 40.788 | 25.2825 | 2.0006 | 32.0408 | 64.0995 | 16.0159 | 14.1027 | -0.1104 | 104.8875 | 234.2624 | 180.3161 | 232.0416 | 160.889
2012-10-30 16:55 | 27.8085 | 248.3946 | 194.4523 | 246.0249 | 174.9196 | -1.5974 | 1.2701 | 32.1125 | 40.7873 | 25.2827 | 2.0006 | 32.0407 | 64.099 | 16.0159 | 14.0685 | -0.1101 | 104.8864 | 234.3261 | 180.3838 | 231.9564 | 160.8511
2012-10-30 16:55 | 27.8085 | 248.3907 | 194.417 | 245.9583 | 174.7322 | -1.5991 | 1.2677 | 32.0504 | 40.6303 | 25.2824 | 1.9963 | 31.9733 | 63.8294 | 16.016 | 14.0889 | -0.1102 | 104.4598 | 234.3018 | 180.3281 | 231.8694 | 160.6433
2012-10-30 16:55 | 27.8264 | 248.4164 | 194.4152 | 246.18 | 174.9062 | -1.5997 | 1.2641 | 31.9582 | 40.3992 | 25.2809 | 1.9909 | 31.8871 | 63.4854 | 16.0161 | 14.0946 | -0.1103 | 103.8845 | 234.3218 | 180.3206 | 232.0854 | 160.8116
2012-10-30 16:55 | 27.8197 | 248.4019 | 194.4085 | 246.0361 | 175.0205 | -1.6037 | 1.2636 | 31.9453 | 40.3663 | 25.2811 | 1.9901 | 31.874 | 63.4332 | 16.0161 | 14.092 | -0.1106 | 103.7995 | 234.3099 | 180.3165 | 231.9441 | 160.9285
2012-10-30 16:55 | 27.8212 | 248.3681 | 194.4375 | 246.0102 | 175.2446 | -1.6009 | 1.2631 | 31.9329 | 40.3354 | 25.2808 | 1.9893 | 31.8616 | 63.3836 | 16.0161 | 14.0812 | -0.1104 | 103.719 | 234.2869 | 180.3563 | 231.929 | 161.1634
2012-10-30 16:55 | 27.8242 | 248.3907 | 194.4759 | 246.1778 | 175.2007 | -1.5991 | 1.2629 | 31.9259 | 40.3181 | 25.2806 | 1.9889 | 31.8547 | 63.3564 | 16.0161 | 14.0639 | -0.1102 | 103.6745 | 234.3268 | 180.412 | 232.1139 | 161.1368
2012-10-30 16:55 | 27.8309 | 248.3348 | 194.4905 | 246.1139 | 175.3714 | -1.6009 | 1.2628 | 31.9237 | 40.312 | 25.2808 | 1.9888 | 31.8523 | 63.3468 | 16.0162 | 14.1113 | -0.1104 | 103.6588 | 234.2235 | 180.3792 | 232.0026 | 161.2601
2012-10-30 16:56 | 27.8317 | 248.3434 | 194.5423 | 246.0325 | 175.2785 | -1.6009 | 1.2628 | 31.9239 | 40.3124 | 25.281 | 1.9888 | 31.8525 | 63.3475 | 16.0162 | 14.0877 | -0.1104 | 103.66 | 234.2557 | 180.4546 | 231.9448 | 161.1908
2012-10-30 16:56 | 27.8317 | 248.3904 | 194.5306 | 246.0364 | 175.216 | -1.5957 | 1.2628 | 31.9258 | 40.3172 | 25.2809 | 1.9889 | 31.8544 | 63.3549 | 16.0161 | 14.104 | -0.11 | 103.6721 | 234.2864 | 180.4266 | 231.9324 | 161.112
2012-10-30 16:56 | 27.8429 | 248.3781 | 194.5418 | 246.122 | 175.0945 | -1.6066 | 1.2629 | 31.9266 | 40.319 | 25.281 | 1.9889 | 31.8552 | 63.3579 | 16.0162 | 14.0907 | -0.1108 | 103.6769 | 234.2874 | 180.4511 | 232.0313 | 161.0038
2012-10-30 16:56 | 27.8377 | 248.3885 | 194.5641 | 246.1756 | 175.2064 | -1.5991 | 1.2629 | 31.9274 | 40.3214 | 25.2809 | 1.989 | 31.8559 | 63.3606 | 16.0163 | 14.0977 | -0.1102 | 103.6819 | 234.2908 | 180.4664 | 232.0779 | 161.1087
2012-10-30 16:56 | 27.8496 | 248.3691 | 194.5995 | 245.9602 | 175.2964 | -1.602 | 1.263 | 31.9301 | 40.3281 | 25.281 | 1.9891 | 31.8587 | 63.3713 | 16.0163 | 14.0566 | -0.1104 | 103.6994 | 234.3125 | 180.5429 | 231.9036 | 161.2398
2012-10-30 16:56 | 27.8496 | 248.3926 | 194.5249 | 246.2384 | 175.3042 | -1.6003 | 1.2631 | 31.9341 | 40.3376 | 25.2814 | 1.9893 | 31.8612 | 63.3816 | 16.0163 | 14.0526 | -0.1103 | 103.7192 | 234.34 | 180.4723 | 232.1858 | 161.2516
2012-10-30 16:56 | 27.8489 | 248.3919 | 194.5988 | 246.1202 | 175.3659 | -1.6055 | 1.2679 | 32.0549 | 40.6419 | 25.2822 | 1.9969 | 31.9832 | 63.8686 | 16.0161 | 14.0926 | -0.1107 | 104.5106 | 234.2993 | 180.5062 | 232.0276 | 161.2733
2012-10-30 16:56 | 27.8556 | 248.3947 | 194.6841 | 246.0289 | 175.3687 | -1.6026 | 1.2684 | 32.0691 | 40.6766 | 25.283 | 1.9978 | 31.9977 | 63.9266 | 16.0161 | 14.083 | -0.1105 | 104.6032 | 234.3117 | 180.6011 | 231.9459 | 161.2857
2012-10-30 16:56 | 27.8564 | 248.3877 | 194.6692 | 246.21 | 175.1119 | -1.5957 | 1.269 | 32.0839 | 40.7147 | 25.2827 | 1.9987 | 32.0122 | 63.9843 | 16.0161 | 14.0716 | -0.11 | 104.6989 | 234.3161 | 180.5976 | 232.1384 | 161.0403
2012-10-30 16:56 | 27.8631 | 248.4296 | 194.6405 | 246.1775 | 175.1576 | -1.598 | 1.2695 | 32.0965 | 40.7464 | 25.2828 | 1.9995 | 32.0246 | 64.034 | 16.0161 | 14.0742 | -0.1102 | 104.7804 | 234.3554 | 180.5663 | 232.1033 | 161.0834
2012-10-30 16:56 | 27.8601 | 248.4501 | 194.6454 | 245.9433 | 175.1976 | -1.6014 | 1.2698 | 32.1051 | 40.7682 | 25.2829 | 2.0001 | 32.0333 | 64.07 | 16.0158 | 14.0468 | -0.1104 | 104.8381 | 234.4033 | 180.5986 | 231.8965 | 161.1508
2012-10-30 16:56 | 27.8661 | 248.4718 | 194.62 | 246.1374 | 175.2192 | -1.6049 | 1.27 | 32.1103 | 40.7812 | 25.283 | 2.0004 | 32.0384 | 64.0896 | 16.016 | 14.065 | -0.1106 | 104.8708 | 234.4068 | 180.555 | 232.0724 | 161.1542
2012-10-30 16:56 | 27.8676 | 248.4458 | 194.6804 | 246.1193 | 175.3222 | -1.6089 | 1.2702 | 32.1139 | 40.7898 | 25.2833 | 2.0006 | 32.0419 | 64.1041 | 16.0159 | 14.0502 | -0.1109 | 104.8939 | 234.3956 | 180.6302 | 232.0691 | 161.272
2012-10-30 16:56 | 27.8721 | 248.466 | 194.7046 | 245.9749 | 175.3072 | -1.5997 | 1.2703 | 32.117 | 40.7981 | 25.2831 | 2.0008 | 32.045 | 64.1169 | 16.0158 | 14.0506 | -0.1103 | 104.915 | 234.4154 | 180.654 | 231.9243 | 161.2566
2012-10-30 16:56 | 27.8736 | 248.5106 | 194.7021 | 246.1331 | 175.2852 | -1.6026 | 1.2704 | 32.1187 | 40.802 | 25.2833 | 2.001 | 32.0468 | 64.1247 | 16.0156 | 14.0847 | -0.1105 | 104.9267 | 234.4259 | 180.6174 | 232.0484 | 161.2005
2012-10-30 16:56 | 27.8774 | 248.4909 | 194.7256 | 246.2545 | 175.2656 | -1.6014 | 1.2704 | 32.1202 | 40.8057 | 25.2834 | 2.0011 | 32.0482 | 64.1304 | 16.0156 | 14.0355 | -0.1104 | 104.9362 | 234.4554 | 180.6901 | 232.219 | 161.2301
2012-10-30 16:56 | 27.8856 | 248.5461 | 194.7023 | 246.0864 | 175.2426 | -1.6026 | 1.2705 | 32.1213 | 40.8088 | 25.2833 | 2.0011 | 32.0495 | 64.1353 | 16.0157 | 14.0437 | -0.1105 | 104.9441 | 234.5024 | 180.6586 | 232.0427 | 161.1989
2012-10-30 16:56 | 27.8826 | 248.5392 | 194.7111 | 246.1069 | 175.0795 | -1.6032 | 1.2705 | 32.1222 | 40.811 | 25.2833 | 2.0012 | 32.0501 | 64.138 | 16.0156 | 14.0326 | -0.1105 | 104.9489 | 234.5066 | 180.6785 | 232.0743 | 161.0469
2012-10-30 16:56 | 27.8931 | 248.5379 | 194.702 | 246.184 | 175.1759 | -1.6009 | 1.2705 | 32.1225 | 40.8112 | 25.2836 | 2.0012 | 32.0506 | 64.1402 | 16.0155 | 14.039 | -0.1104 | 104.9514 | 234.4989 | 180.663 | 232.145 | 161.1369
2012-10-30 16:56 | 27.8901 | 248.5662 | 194.7343 | 246.0987 | 175.3017 | -1.6003 | 1.2705 | 32.1231 | 40.8131 | 25.2834 | 2.0013 | 32.0512 | 64.143 | 16.0155 | 14.0401 | -0.1103 | 104.9561 | 234.5261 | 180.6942 | 232.0586 | 161.2616
2012-10-30 16:56 | 27.8908 | 248.5826 | 194.6997 | 246.1817 | 175.1736 | -1.6003 | 1.2705 | 32.1234 | 40.8135 | 25.2836 | 2.0013 | 32.0515 | 64.1442 | 16.0155 | 14.0571 | -0.1103 | 104.9577 | 234.5255 | 180.6426 | 232.1246 | 161.1165
2012-10-30 16:56 | 27.8871 | 248.5867 | 194.6724 | 246.3034 | 175.2675 | -1.598 | 1.2705 | 32.1235 | 40.8134 | 25.2838 | 2.0013 | 32.0514 | 64.1438 | 16.0155 | 14.0493 | -0.1102 | 104.9573 | 234.5374 | 180.6231 | 232.2541 | 161.2182
2012-10-30 16:56 | 27.8983 | 248.6058 | 194.7858 | 246.1657 | 175.3685 | -1.6003 | 1.2705 | 32.1237 | 40.8142 | 25.2836 | 2.0013 | 32.0516 | 64.1455 | 16.0153 | 14.0686 | -0.1103 | 104.9597 | 234.5372 | 180.7172 | 232.0971 | 161.2999
2012-10-30 16:56 | 27.8976 | 248.5855 | 194.7497 | 246.1924 | 175.2858 | -1.598 | 1.2706 | 32.1241 | 40.8153 | 25.2836 | 2.0014 | 32.0522 | 64.1478 | 16.0153 | 14.072 | -0.1102 | 104.9631 | 234.5135 | 180.6777 | 232.1204 | 161.2138
2012-10-30 16:56 | 27.9081 | 248.6547 | 194.7563 | 246.3479 | 175.2222 | -1.6009 | 1.2706 | 32.1247 | 40.8166 | 25.2837 | 2.0014 | 32.0526 | 64.1498 | 16.0152 | 14.0458 | -0.1104 | 104.9664 | 234.6089 | 180.7105 | 232.3021 | 161.1764
2012-10-30 16:56 | 27.9133 | 248.6286 | 194.7733 | 246.259 | 175.4382 | -1.5997 | 1.2706 | 32.1248 | 40.8169 | 25.2837 | 2.0014 | 32.0527 | 64.1503 | 16.0151 | 14.0918 | -0.1103 | 104.9671 | 234.5368 | 180.6815 | 232.1672 | 161.3464
2012-10-30 16:57 | 27.9148 | 248.6849 | 194.7433 | 246.2214 | 175.4357 | -1.6014 | 1.2706 | 32.1248 | 40.8167 | 25.2839 | 2.0014 | 32.0528 | 64.1507 | 16.0152 | 14.0688 | -0.1104 | 104.9674 | 234.6161 | 180.6745 | 232.1526 | 161.3669
2012-10-30 16:57 | 27.8991 | 248.6614 | 194.7984 | 246.3349 | 175.3185 | -1.5968 | 1.2706 | 32.1246 | 40.8165 | 25.2837 | 2.0014 | 32.0527 | 64.1501 | 16.0152 | 14.0858 | -0.1101 | 104.9666 | 234.5756 | 180.7126 | 232.2491 | 161.2327
2012-10-30 16:57 | 27.914 | 248.6763 | 194.8565 | 246.2793 | 175.3452 | -1.5997 | 1.2706 | 32.1248 | 40.8167 | 25.2838 | 2.0014 | 32.0528 | 64.1503 | 16.0152 | 14.0721 | -0.1103 | 104.967 | 234.6042 | 180.7844 | 232.2072 | 161.2731
2012-10-30 16:57 | 27.9208 | 248.6987 | 194.8318 | 246.2352 | 175.5237 | -1.6009 | 1.2706 | 32.1246 | 40.8164 | 25.2837 | 2.0014 | 32.0526 | 64.1498 | 16.0151 | 14.1075 | -0.1104 | 104.9662 | 234.5912 | 180.7243 | 232.1277 | 161.4162
2012-10-30 16:57 | 27.9133 | 248.6482 | 194.8243 | 246.357 | 175.446 | -1.602 | 1.2706 | 32.1245 | 40.8161 | 25.2837 | 2.0014 | 32.0524 | 64.1496 | 16.015 | 14.0388 | -0.1104 | 104.9658 | 234.6094 | 180.7855 | 232.3182 | 161.4072
2012-10-30 16:57 | 27.9223 | 248.6532 | 194.8608 | 246.3934 | 175.6267 | -1.5997 | 1.2705 | 32.1242 | 40.8153 | 25.2838 | 2.0014 | 32.0522 | 64.1491 | 16.015 | 14.06 | -0.1103 | 104.9643 | 234.5932 | 180.8008 | 232.3334 | 161.5667
2012-10-30 16:57 | 27.9358 | 248.7137 | 194.9451 | 246.2894 | 175.6988 | -1.6014 | 1.2705 | 32.124 | 40.8149 | 25.2836 | 2.0014 | 32.0519 | 64.1484 | 16.0148 | 14.0817 | -0.1104 | 104.9633 | 234.632 | 180.8634 | 232.2077 | 161.6171
2012-10-30 16:57 | 27.935 | 248.7442 | 194.9718 | 246.3199 | 175.6785 | -1.6009 | 1.2705 | 32.1238 | 40.8141 | 25.2839 | 2.0014 | 32.0517 | 64.1473 | 16.0149 | 14.0686 | -0.1104 | 104.9614 | 234.6756 | 180.9032 | 232.2513 | 161.6099
2012-10-30 16:57 | 27.9388 | 248.7637 | 194.9756 | 246.4843 | 175.5964 | -1.5974 | 1.2705 | 32.1235 | 40.8133 | 25.2839 | 2.0014 | 32.0514 | 64.1465 | 16.0148 | 14.0806 | -0.1101 | 104.9598 | 234.6831 | 180.895 | 232.4037 | 161.5158
2012-10-30 16:57 | 27.9492 | 248.735 | 195.0213 | 246.4556 | 175.5365 | -1.5991 | 1.2705 | 32.1235 | 40.8136 | 25.2838 | 2.0014 | 32.0515 | 64.1466 | 16.0149 | 14.0625 | -0.1102 | 104.9602 | 234.6725 | 180.9588 | 232.3931 | 161.474
2012-10-30 16:57 | 27.947 | 248.7602 | 195.0113 | 246.2536 | 175.6358 | -1.6055 | 1.2705 | 32.1232 | 40.8126 | 25.2839 | 2.0013 | 32.0512 | 64.1453 | 16.0149 | 14.0276 | -0.1107 | 104.9579 | 234.7326 | 180.9837 | 232.226 | 161.6082
2012-10-30 16:57 | 27.953 | 248.7818 | 195.0134 | 246.4163 | 175.5442 | -1.5991 | 1.2705 | 32.123 | 40.8116 | 25.2841 | 2.0013 | 32.0509 | 64.1446 | 16.0148 | 14.0581 | -0.1102 | 104.9562 | 234.7237 | 180.9553 | 232.3582 | 161.4861
2012-10-30 16:57 | 27.9657 | 248.7906 | 195.0575 | 246.5504 | 175.6077 | -1.6009 | 1.2705 | 32.1227 | 40.8112 | 25.2839 | 2.0013 | 32.0507 | 64.1441 | 16.0147 | 14.1034 | -0.1104 | 104.9553 | 234.6872 | 180.9541 | 232.447 | 161.5043
2012-10-30 16:57 | 27.965 | 248.8369 | 195.0725 | 246.3538 | 175.6968 | -1.6037 | 1.2705 | 32.1227 | 40.8111 | 25.284 | 2.0013 | 32.0507 | 64.1439 | 16.0147 | 14.0497 | -0.1106 | 104.955 | 234.7872 | 181.0228 | 232.3041 | 161.6471
2012-10-30 16:57 | 27.9605 | 248.8089 | 195.0209 | 246.4512 | 175.5517 | -1.6043 | 1.2705 | 32.1229 | 40.8118 | 25.2839 | 2.0014 | 32.0508 | 64.1451 | 16.0145 | 14.0167 | -0.1106 | 104.9569 | 234.7922 | 181.0042 | 232.4345 | 161.535
2012-10-30 16:57 | 27.9612 | 248.8213 | 195.0333 | 246.5342 | 175.611 | -1.5997 | 1.2705 | 32.1228 | 40.8115 | 25.2839 | 2.0014 | 32.0506 | 64.1446 | 16.0145 | 14.0377 | -0.1103 | 104.9561 | 234.7836 | 180.9956 | 232.4965 | 161.5733
2012-10-30 16:57 | 27.968 | 248.8555 | 195.0755 | 246.4587 | 175.7271 | -1.6014 | 1.2705 | 32.1225 | 40.8104 | 25.2841 | 2.0013 | 32.0505 | 64.1436 | 16.0146 | 14.0445 | -0.1104 | 104.954 | 234.811 | 181.031 | 232.4142 | 161.6826
2012-10-30 16:57 | 27.9799 | 248.8244 | 195.0992 | 246.4157 | 175.6687 | -1.5997 | 1.2705 | 32.122 | 40.8096 | 25.2839 | 2.0013 | 32.0501 | 64.142 | 16.0146 | 14.0687 | -0.1103 | 104.9516 | 234.7557 | 181.0305 | 232.347 | 161.6
2012-10-30 16:57 | 27.9717 | 248.8632 | 195.0792 | 246.5212 | 175.4966 | -1.5997 | 1.2704 | 32.1215 | 40.8075 | 25.2844 | 2.0013 | 32.0494 | 64.1397 | 16.0144 | 14.0075 | -0.1103 | 104.9472 | 234.8557 | 181.0717 | 232.5137 | 161.4891
2012-10-30 16:57 | 27.9874 | 248.8358 | 195.1185 | 246.576 | 175.6216 | -1.6014 | 1.2704 | 32.1209 | 40.8064 | 25.2842 | 2.0013 | 32.049 | 64.1385 | 16.0144 | 14.0476 | -0.1104 | 104.9448 | 234.7882 | 181.0709 | 232.5284 | 161.574
2012-10-30 16:57 | 27.9822 | 248.8972 | 195.1015 | 246.4376 | 175.4836 | -1.6055 | 1.2704 | 32.1208 | 40.8064 | 25.284 | 2.0013 | 32.0487 | 64.1377 | 16.0143 | 14.0302 | -0.1107 | 104.9441 | 234.867 | 181.0713 | 232.4074 | 161.4534
2012-10-30 16:57 | 27.9927 | 248.892 | 195.0806 | 246.503 | 175.459 | -1.5968 | 1.2704 | 32.1205 | 40.8055 | 25.2839 | 2.0012 | 32.0484 | 64.1365 | 16.0143 | 13.9999 | -0.1101 | 104.942 | 234.8921 | 181.0807 | 232.5031 | 161.4591
2012-10-30 16:57 | 28.0077 | 248.9109 | 195.1034 | 246.6394 | 175.4857 | -1.6037 | 1.2703 | 32.12 | 40.8035 | 25.2844 | 2.0012 | 32.0479 | 64.1345 | 16.0143 | 14.0557 | -0.1106 | 104.9381 | 234.8552 | 181.0477 | 232.5837 | 161.43
2012-10-30 16:57 | 28.0054 | 248.9399 | 195.0775 | 246.5431 | 175.663 | -1.5985 | 1.2703 | 32.1196 | 40.8029 | 25.2842 | 2.0012 | 32.0476 | 64.1337 | 16.0142 | 14.0738 | -0.1102 | 104.9367 | 234.8661 | 181.0037 | 232.4693 | 161.5892
2012-10-30 16:57 | 28.0002 | 248.923 | 195.0841 | 246.4204 | 175.65 | -1.5991 | 1.2703 | 32.1192 | 40.8021 | 25.2842 | 2.0012 | 32.0472 | 64.1321 | 16.0142 | 14.0971 | -0.1102 | 104.9342 | 234.8259 | 180.987 | 232.3233 | 161.5529
2012-10-30 16:57 | 28.0039 | 248.8993 | 195.0957 | 246.7375 | 175.5873 | -1.5962 | 1.2703 | 32.1181 | 40.799 | 25.2843 | 2.0011 | 32.046 | 64.128 | 16.014 | 14.0927 | -0.11 | 104.927 | 234.8066 | 181.003 | 232.6448 | 161.4946
2012-10-30 16:58 | 28.0099 | 248.9366 | 195.0938 | 246.5868 | 175.6245 | -1.6003 | 1.2703 | 32.1183 | 40.8 | 25.284 | 2.0011 | 32.0463 | 64.1292 | 16.014 | 14.0579 | -0.1103 | 104.9292 | 234.8787 | 181.0359 | 232.5289 | 161.5666
2012-10-30 16:58 | 28.0092 | 248.9555 | 195.0853 | 246.637 | 175.5926 | -1.6014 | 1.2703 | 32.1183 | 40.7999 | 25.284 | 2.0011 | 32.0461 | 64.1282 | 16.0141 | 14.0817 | -0.1104 | 104.9281 | 234.8738 | 181.0036 | 232.5553 | 161.5109
2012-10-30 16:58 | 28.0129 | 248.9357 | 195.1322 | 246.7191 | 175.5885 | -1.6032 | 1.2703 | 32.1175 | 40.7983 | 25.2838 | 2.0011 | 32.0456 | 64.1264 | 16.014 | 14.0201 | -0.1105 | 104.9246 | 234.9156 | 181.1121 | 232.699 | 161.5684
2012-10-30 16:58 | 28.0084 | 248.9899 | 195.0884 | 246.7773 | 175.8261 | -1.6026 | 1.2702 | 32.1169 | 40.7964 | 25.2839 | 2.001 | 32.0449 | 64.1231 | 16.0141 | 14.0482 | -0.1105 | 104.9196 | 234.9417 | 181.0402 | 232.7291 | 161.7779
2012-10-30 16:58 | 28.0092 | 248.9594 | 195.1403 | 246.5704 | 175.823 | -1.602 | 1.2703 | 32.1172 | 40.7971 | 25.2841 | 2.0011 | 32.0451 | 64.1244 | 16.014 | 14.049 | -0.1104 | 104.9214 | 234.9104 | 181.0913 | 232.5214 | 161.774
2012-10-30 16:58 | 28.0099 | 248.9992 | 195.1488 | 246.6299 | 175.7534 | -1.598 | 1.2702 | 32.117 | 40.7964 | 25.2841 | 2.0011 | 32.0448 | 64.1241 | 16.0138 | 14.0416 | -0.1102 | 104.9204 | 234.9576 | 181.1072 | 232.5883 | 161.7118
2012-10-30 16:58 | 28.0151 | 248.9927 | 195.2091 | 246.784 | 175.7664 | -1.5991 | 1.2702 | 32.1164 | 40.7949 | 25.2841 | 2.001 | 32.0443 | 64.1214 | 16.014 | 14.0345 | -0.1102 | 104.9163 | 234.9582 | 181.1746 | 232.7495 | 161.7319
2012-10-30 16:58 | 28.0271 | 249.0321 | 195.1936 | 246.5883 | 175.794 | -1.5951 | 1.2702 | 32.1159 | 40.7937 | 25.284 | 2.001 | 32.0438 | 64.1202 | 16.0138 | 14.0384 | -0.11 | 104.914 | 234.9937 | 181.1552 | 232.5499 | 161.7556
2012-10-30 16:58 | 28.0196 | 249.0363 | 195.1939 | 246.6709 | 175.6499 | -1.5962 | 1.2702 | 32.1158 | 40.7939 | 25.2838 | 2.001 | 32.0436 | 64.1193 | 16.0138 | 14.0635 | -0.11 | 104.9133 | 234.9728 | 181.1304 | 232.6074 | 161.5864
2012-10-30 16:58 | 28.0226 | 249.0002 | 195.1851 | 246.7797 | 175.4772 | -1.602 | 1.2702 | 32.1153 | 40.7923 | 25.2841 | 2.001 | 32.0433 | 64.1187 | 16.0136 | 14.0788 | -0.1104 | 104.911 | 234.9214 | 181.1063 | 232.7009 | 161.3984
2012-10-30 16:58 | 28.0226 | 249.0119 | 195.1891 | 246.6896 | 175.6919 | -1.6049 | 1.2657 | 32.001 | 40.5037 | 25.2832 | 1.9937 | 31.9266 | 63.6524 | 16.0137 | 14.0584 | -0.1106 | 104.1562 | 234.9535 | 181.1307 | 232.6312 | 161.6335
2012-10-30 16:58 | 28.0271 | 249.0791 | 195.17 | 246.6001 | 175.548 | -1.5985 | 1.264 | 31.9571 | 40.3938 | 25.2825 | 1.9911 | 31.8853 | 63.4869 | 16.0139 | 14.0506 | -0.1102 | 103.8807 | 235.0285 | 181.1194 | 232.5495 | 161.4974
2012-10-30 16:58 | 28.0294 | 248.9992 | 195.1369 | 246.8139 | 175.4059 | -1.6032 | 1.2631 | 31.9338 | 40.3348 | 25.2825 | 1.9896 | 31.8619 | 63.3933 | 16.014 | 14.1059 | -0.1105 | 103.7281 | 234.8933 | 181.031 | 232.708 | 161.3
2012-10-30 16:58 | 28.0294 | 249.0266 | 195.1566 | 246.7317 | 175.7456 | -1.6014 | 1.2629 | 31.9289 | 40.3228 | 25.2824 | 1.9893 | 31.8574 | 63.3754 | 16.014 | 14.0896 | -0.1104 | 103.6982 | 234.937 | 181.067 | 232.6421 | 161.656
2012-10-30 16:58 | 28.0294 | 249.0266 | 195.1055 | 246.6024 | 175.5464 | -1.598 | 1.2628 | 31.9277 | 40.3196 | 25.2824 | 1.9893 | 31.856 | 63.3703 | 16.0139 | 14.1182 | -0.1102 | 103.6899 | 234.9084 | 180.9873 | 232.4842 | 161.4282
2012-10-30 16:58 | 28.0331 | 249.0263 | 195.0699 | 246.7706 | 175.0856 | -1.5997 | 1.2628 | 31.9273 | 40.3184 | 25.2825 | 1.9892 | 31.8556 | 63.3685 | 16.014 | 14.1627 | -0.1103 | 103.6869 | 234.8636 | 180.9072 | 232.6079 | 160.9229
2012-10-30 16:58 | 28.0309 | 248.9928 | 195.0087 | 246.788 | 175.4777 | -1.6037 | 1.2627 | 31.9246 | 40.3121 | 25.2823 | 1.9891 | 31.853 | 63.3574 | 16.0141 | 14.1238 | -0.1106 | 103.6694 | 234.869 | 180.8849 | 232.6642 | 161.3539
2012-10-30 16:58 | 28.0309 | 248.9772 | 195.052 | 246.7802 | 175.5518 | -1.5968 | 1.2626 | 31.9225 | 40.3064 | 25.2824 | 1.9889 | 31.8508 | 63.3484 | 16.0142 | 14.0748 | -0.1101 | 103.6548 | 234.9024 | 180.9772 | 232.7054 | 161.477
2012-10-30 16:58 | 28.0294 | 248.9404 | 195.019 | 246.6651 | 175.7534 | -1.6043 | 1.2627 | 31.9243 | 40.3114 | 25.2823 | 1.989 | 31.8527 | 63.3559 | 16.0142 | 14.0407 | -0.1106 | 103.6673 | 234.8997 | 180.9783 | 232.6244 | 161.7127
2012-10-30 16:58 | 28.0339 | 248.9528 | 195.0078 | 246.7832 | 175.7579 | -1.6026 | 1.2628 | 31.928 | 40.32 | 25.2828 | 1.9893 | 31.8562 | 63.3702 | 16.0141 | 14.033 | -0.1105 | 103.6901 | 234.9198 | 180.9748 | 232.7502 | 161.7249
2012-10-30 16:58 | 28.0391 | 248.8797 | 195.0405 | 246.7414 | 175.7592 | -1.6009 | 1.2629 | 31.9293 | 40.3236 | 25.2824 | 1.9893 | 31.8573 | 63.3741 | 16.0143 | 14.0341 | -0.1104 | 103.6977 | 234.8456 | 181.0064 | 232.7073 | 161.7251
2012-10-30 16:58 | 28.0376 | 248.9369 | 195.0783 | 246.6537 | 175.6366 | -1.5997 | [8 empty] | 14.0448 | -0.1103 | 0 | 234.8921 | 181.0335 | 232.6089 | 161.5918
2012-10-30 16:59
2012-10-30 17:01
2012-10-30 17:04
2012-10-30 17:06
2012-10-30 17:08
2012-10-30 17:11
2012-10-30 17:14 | [6 empty] | 1.2974 | 31.0475 | 40.282 | 23.93 | 2.001 | 32.0497 | 64.1322 | 16.0166 | -12.5097 |  | 104.4142 | 12.5097 | 12.5097 | 12.5097 | 12.5097
2012-10-30 17:15 | 25.8585 | 53.4345 | 47.6496 | 50.5264 | 42.6206 | -1.6003 | 1.2959 | 32.1168 | 41.6205 | 24.7832 | 2.0011 | 32.0481 | 64.1316 | 16.0152 | 13.3121 | -0.1103 | 105.7521 | 40.1224 | 34.3375 | 37.2143 | 29.3085
2012-10-30 17:15 | 25.8637 | 55.277 | 48.4128 | 50.746 | 43.0771 | -1.6055 | 1.2955 | 32.1159 | 41.6064 | 24.7902 | 2.0011 | 32.0472 | 64.1302 | 16.0146 | 13.3417 | -0.1107 | 105.7366 | 41.9353 | 35.0711 | 37.4043 | 29.7354
2012-10-30 17:15 | 25.8532 | 57.2285 | 49.2144 | 51.1641 | 43.498 | -1.5997 | 1.2953 | 32.1163 | 41.5987 | 24.7955 | 2.0012 | 32.0474 | 64.1329 | 16.0142 | 13.3394 | -0.1103 | 105.7315 | 43.8891 | 35.875 | 37.8247 | 30.1586
2012-10-30 17:15 | 25.8577 | 58.9903 | 49.9017 | 51.8697 | 43.8594 | -1.5962 | 1.295 | 32.117 | 41.5914 | 24.8008 | 2.0013 | 32.0482 | 64.1368 | 16.014 | 13.3357 | -0.11 | 105.7282 | 45.6546 | 36.566 | 38.534 | 30.5237
2012-10-30 17:15 | 25.8517 | 60.7464 | 50.6677 | 52.4709 | 44.4649 | -1.598 | 1.2948 | 32.118 | 41.5865 | 24.8053 | 2.0014 | 32.049 | 64.1417 | 16.0136 | 13.3338 | -0.1102 | 105.7282 | 47.4126 | 37.3339 | 39.1371 | 31.1311
2012-10-30 17:15 | 25.8562 | 62.2325 | 51.3892 | 53.2145 | 44.9631 | -1.5985 | 1.2946 | 32.1191 | 41.5811 | 24.8102 | 2.0015 | 32.0501 | 64.1471 | 16.0133 | 13.3546 | -0.1102 | 105.7282 | 48.8779 | 38.0346 | 39.8599 | 31.6085
2012-10-30 17:15 | 25.8644 | 63.8858 | 52.145 | 54.2915 | 45.5391 | -1.5962 | 1.2944 | 32.1202 | 41.5765 | 24.8147 | 2.0016 | 32.0511 | 64.1524 | 16.013 | 13.375 | -0.11 | 105.7289 | 50.5108 | 38.77 | 40.9165 | 32.1641
2012-10-30 17:15 | 25.8652 | 65.5986 | 52.9746 | 55.4299 | 46.1115 | -1.6014 | 1.2942 | 32.1217 | 41.5721 | 24.8196 | 2.0017 | 32.0525 | 64.1588 | 16.0128 | 13.3717 | -0.1104 | 105.7308 | 52.2269 | 39.6029 | 42.0582 | 32.7398
2012-10-30 17:15 | 25.8607 | 67.4838 | 53.8449 | 56.6625 | 46.8386 | -1.5985 | 1.294 | 32.1232 | 41.5679 | 24.8244 | 2.0018 | 32.054 | 64.1664 | 16.0124 | 13.3306 | -0.1102 | 105.7342 | 54.1532 | 40.5143 | 43.3319 | 33.508
2012-10-30 17:16 | 25.8615 | 69.1261 | 54.7197 | 57.9223 | 47.6018 | -1.5945 | 1.2938 | 32.1244 | 41.5638 | 24.8288 | 2.0019 | 32.0552 | 64.1714 | 16.0123 | 13.3599 | -0.1099 | 105.7352 | 55.7662 | 41.3598 | 44.5624 | 34.2419
2012-10-30 17:16 | 25.86 | 70.4883 | 55.5023 | 59.097 | 48.2216 | -1.5985 | 1.2937 | 32.1256 | 41.5603 | 24.8327 | 2.002 | 32.0562 | 64.1765 | 16.0121 | 13.3665 | -0.1102 | 105.7368 | 57.1218 | 42.1358 | 45.7305 | 34.8551
2012-10-30 17:16 | 25.8615 | 71.8214 | 56.1516 | 60.3854 | 48.7152 | -1.5968 | 1.2936 | 32.1267 | 41.558 | 24.8359 | 2.0021 | 32.0575 | 64.1827 | 16.0118 | 13.3721 | -0.1101 | 105.7407 | 58.4493 | 42.7795 | 47.0133 | 35.3431
2012-10-30 17:16 | 25.8674 | 73.276 | 57.018 | 61.6999 | 49.4003 | -1.5991 | 1.2934 | 32.128 | 41.555 | 24.8396 | 2.0022 | 32.0587 | 64.1886 | 16.0115 | 13.3699 | -0.1102 | 105.7436 | 59.9061 | 43.6481 | 48.33 | 36.0304
2012-10-30 17:16 | 25.863 | 74.8217 | 57.9083 | 63.0872 | 50.1761 | -1.5991 | 1.2932 | 32.129 | 41.55 | 24.8441 | 2.0023 | 32.0596 | 64.1927 | 16.0114 | 13.3777 | -0.1102 | 105.7427 | 61.444 | 44.5306 | 49.7095 | 36.7984
2012-10-30 17:16 | 25.8674 | 76.3593 | 58.7563 | 64.5508 | 50.8471 | -1.5933 | 1.2931 | 32.13 | 41.5463 | 24.8478 | 2.0024 | 32.0606 | 64.1983 | 16.0111 | 13.3861 | -0.1098 | 105.7445 | 62.9732 | 45.3702 | 51.1647 | 37.461
2012-10-30 17:16 | 25.8667 | 77.7113 | 59.5676 | 66.0574 | 51.3763 | -1.5951 | 1.2929 | 32.1312 | 41.5439 | 24.8512 | 2.0025 | 32.0617 | 64.2033 | 16.0109 | 13.3366 | -0.11 | 105.7471 | 64.3747 | 46.231 | 52.7208 | 38.0397
2012-10-30 17:16 | 25.8727 | 78.9623 | 60.2692 | 67.1312 | 52.1377 | -1.602 | 1.2928 | 32.1322 | 41.5411 | 24.8544 | 2.0026 | 32.0626 | 64.2084 | 16.0106 | 13.363 | -0.1104 | 105.7494 | 65.5993 | 46.9062 | 53.7682 | 38.7747
2012-10-30 17:16 | 25.8749 | 80.3416 | 61.1209 | 68.5146 | 52.7937 | -1.5962 | 1.2927 | 32.1333 | 41.538 | 24.8579 | 2.0027 | 32.0637 | 64.2131 | 16.0105 | 13.3977 | -0.11 | 105.7511 | 66.9439 | 47.7232 | 55.1169 | 39.396
2012-10-30 17:16 | 25.8697 | 81.7239 | 61.9606 | 70.0998 | 53.3213 | -1.5968 | 1.2925 | 32.1338 | 41.5331 | 24.8617 | 2.0027 | 32.0642 | 64.2161 | 16.0102 | 13.4088 | -0.1101 | 105.7493 | 68.3151 | 48.5518 | 56.691 | 39.9125
2012-10-30 17:16 | 25.8682 | 83.0442 | 62.7302 | 71.5374 | 54.1049 | -1.5985 | 1.2924 | 32.1345 | 41.5299 | 24.8647 | 2.0028 | 32.065 | 64.2199 | 16.0101 | 13.4154 | -0.1102 | 105.7497 | 69.6288 | 49.3148 | 58.122 | 40.6895
2012-10-30 17:16 | 25.8824 | 84.2656 | 63.5149 | 72.8439 | 54.9153 | -1.598 | 1.2923 | 32.1351 | 41.5269 | 24.8675 | 2.0029 | 32.0656 | 64.2231 | 16.0098 | 13.4337 | -0.1102 | 105.75 | 70.8319 | 50.0812 | 59.4102 | 41.4816
2012-10-30 17:16 | 25.8847 | 85.4213 | 64.2064 | 74.1907 | 55.4106 | -1.6009 | 1.2921 | 32.1354 | 41.522 | 24.8708 | 2.0029 | 32.0658 | 64.2241 | 16.0098 | 13.4604 | -0.1104 | 105.746 | 71.9609 | 50.746 | 60.7303 | 41.9502
2012-10-30 17:16 | 25.8869 | 86.8304 | 65.0704 | 75.723 | 56.2197 | -1.5991 | 1.2919 | 32.1362 | 41.5177 | 24.8746 | 2.003 | 32.0665 | 64.2288 | 16.0094 | 13.4341 | -0.1102 | 105.7465 | 73.3963 | 51.6363 | 62.2889 | 42.7856
2012-10-30 17:16 | 25.8876 | 88.2598 | 65.9438 | 77.1951 | 57.108 | -1.5974 | 1.2918 | 32.1371 | 41.514 | 24.8782 | 2.0031 | 32.0673 | 64.2329 | 16.0091 | 13.4674 | -0.1101 | 105.7469 | 74.7924 | 52.4764 | 63.7277 | 43.6406
2012-10-30 17:16 | 25.8921 | 89.5409 | 66.7321 | 78.5402 | 57.7786 | -1.6009 | 1.2916 | 32.1378 | 41.5105 | 24.8813 | 2.0031 | 32.068 | 64.2353 | 16.0092 | 13.476 | -0.1104 | 105.7459 | 76.0649 | 53.2561 | 65.0642 | 44.3026
2012-10-30 17:16 | 25.8996 | 90.6087 | 67.5266 | 80.0126 | 58.4016 | -1.598 | 1.2915 | 32.1386 | 41.5084 | 24.8839 | 2.0032 | 32.0688 | 64.2395 | 16.0089 | 13.4916 | -0.1102 | 105.7479 | 77.1171 | 54.035 | 66.521 | 44.91
2012-10-30 17:16 | 25.8959 | 91.7066 | 68.2634 | 81.2218 | 59.2102 | -1.5997 | 1.2914 | 32.1391 | 41.5054 | 24.8864 | 2.0032 | 32.0693 | 64.2419 | 16.0088 | 13.4798 | -0.1103 | 105.7473 | 78.2268 | 54.7836 | 67.742 | 45.7304
2012-10-30 17:16 | 25.9086 | 92.9757 | 69.0889 | 82.4656 | 59.8144 | -1.6003 | 1.2913 | 32.1396 | 41.5007 | 24.89 | 2.0033 | 32.0697 | 64.2448 | 16.0085 | 13.4965 | -0.1103 | 105.7456 | 79.4792 | 55.5924 | 68.9691 | 46.3179
2012-10-30 17:16 | 25.9168 | 94.1833 | 69.9401 | 83.9281 | 60.518 | -1.5945 | 1.2911 | 32.1399 | 41.4961 | 24.8932 | 2.0033 | 32.0699 | 64.2461 | 16.0084 | 13.5332 | -0.1099 | 105.7422 | 80.6501 | 56.4069 | 70.3949 | 46.9848
2012-10-30 17:16 | 25.9198 | 95.4156 | 70.751 | 85.263 | 61.3047 | -1.5974 |  |  |  |  | 2.0035 | 32.074 | 64.2602 | 16.009 | 13.5118 | -0.1101 | 64.2602 | 81.9038 | 57.2392 | 71.7512 | 47.7929
2012-10-30 17:17 | 0 | 69.318 | 46.3461 | 65.8448 | 38.0484 | -1.5939 | 1.291 | 32.1414 | 41.496 | 24.8956 | 2.0034 | 32.0714 | 64.2531 | 16.0081 | -12.4039 | -0.1099 | 105.7491 | 81.7219 | 58.75 | 78.2487 | 50.4523
2012-10-30 17:17 | 25.9438 | 98.0828 | 74.1009 | 93.0034 | 64.8542 | -1.5968 | 1.2908 | 32.1404 | 41.4864 | 24.8999 | 2.0034 | 32.0704 | 64.2502 | 16.0079 | 13.5358 | -0.1101 | 105.7366 | 84.547 | 60.5651 | 79.4676 | 51.3184
2012-10-30 17:17 | 25.9475 | 99.3493 | 74.8261 | 93.8546 | 65.3887 | -1.602 | 1.2906 | 32.14 | 41.4796 | 24.9034 | 2.0034 | 32.07 | 64.2491 | 16.0078 | 13.5395 | -0.1104 | 105.7287 | 85.8098 | 61.2866 | 80.3151 | 51.8492
2012-10-30 17:17 | 25.964 | 100.4584 | 75.5332 | 94.6201 | 66.0318 | -1.5985 | 1.2904 | 32.1398 | 41.4738 | 24.9065 | 2.0034 | 32.0697 | 64.2485 | 16.0076 | 13.5438 | -0.1102 | 105.7223 | 86.9146 | 61.9894 | 81.0763 | 52.488
2012-10-30 17:17 | 25.9677 | 101.5811 | 76.2576 | 95.2819 | 66.6732 | -1.5968 | 1.2902 | 32.1396 | 41.4678 | 24.9098 | 2.0034 | 32.0694 | 64.2482 | 16.0074 | 13.5841 | -0.1101 | 105.716 | 87.997 | 62.6735 | 81.6978 | 53.0891
2012-10-30 17:17 | 25.961 | 102.8598 | 76.9256 | 96.2812 | 67.1888 | -1.5962 | 1.2901 | 32.1399 | 41.463 | 24.9131 | 2.0034 | 32.0698 | 64.2503 | 16.0073 | 13.5204 | -0.11 | 105.7134 | 89.3394 | 63.4052 | 82.7608 | 53.6684
2012-10-30 17:17 | 25.9737 | 104.0861 | 77.7726 | 97.2582 | 67.8655 | -1.5974 | 1.2899 | 32.1403 | 41.4591 | 24.9161 | 2.0035 | 32.07 | 64.252 | 16.0071 | 13.5616 | -0.1101 | 105.7111 | 90.5245 | 64.211 | 83.6966 | 54.3039
2012-10-30 17:17 | 25.9745 | 105.104 | 78.535 | 98.0757 | 68.5874 | -1.6049 | 1.2898 | 32.1403 | 41.4548 | 24.9188 | 2.0035 | 32.0702 | 64.2529 | 16.007 | 13.5462 | -0.1106 | 105.7077 | 91.5578 | 64.9888 | 84.5295 | 55.0412
2012-10-30 17:17 | 25.9782 | 106.174 | 79.1515 | 98.949 | 69.0895 | -1.5962 | 1.2897 | 32.1406 | 41.4502 | 24.9218 | 2.0035 | 32.0703 | 64.2535 | 16.007 | 13.5295 | -0.11 | 105.7037 | 92.6445 | 65.622 | 85.4195 | 55.56
2012-10-30 17:17 | 25.9902 | 107.2638 | 79.9054 | 99.9629 | 69.8219 | -1.5985 | 1.2895 | 32.141 | 41.4458 | 24.9252 | 2.0036 | 32.0707 | 64.2554 | 16.0069 | 13.5537 | -0.1102 | 105.7012 | 93.7101 | 66.3517 | 86.4092 | 56.2682
2012-10-30 17:17 | 25.9917 | 108.4719 | 80.7968 | 101.0947 | 70.5894 | -1.6009 | 1.2893 | 32.141 | 41.4401 | 24.9286 | 2.0036 | 32.0707 | 64.2554 | 16.0069 | 13.5471 | -0.1104 | 105.6955 | 94.9248 | 67.2497 | 87.5476 | 57.0423
2012-10-30 17:17 | 26.0021 | 109.7006 | 81.6853 | 102.1032 | 71.3805 | -1.598 | 1.2892 | 32.1412 | 41.4363 | 24.9311 | 2.0036 | 32.0709 | 64.2576 | 16.0066 | 13.5941 | -0.1102 | 105.694 | 96.1065 | 68.0912 | 88.5091 | 57.7864
2012-10-30 17:17 | 26.0066 | 110.5602 | 82.2597 | 103.0982 | 71.584 | -1.6049 | 1.2891 | 32.1415 | 41.4333 | 24.9335 | 2.0036 | 32.0712 | 64.2582 | 16.0067 | 13.5701 | -0.1106 | 105.6916 | 96.9901 | 68.6896 | 89.5281 | 58.0139
2012-10-30 17:17 | 26.0134 | 111.4262 | 82.8894 | 103.9065 | 72.3748 | -1.5951 | 1.289 | 32.1416 | 41.4298 | 24.9358 | 2.0036 | 32.0713 | 64.2594 | 16.0065 | 13.5973 | -0.11 | 105.6892 | 97.8289 | 69.2921 | 90.3092 | 58.7775
2012-10-30 17:17 | 26.0186 | 112.5406 | 83.692 | 104.8456 | 73.0565 | -1.5962 | 1.2888 | 32.142 | 41.4255 | 24.9389 | 2.0037 | 32.0717 | 64.2612 | 16.0065 | 13.574 | -0.11 | 105.6867 | 98.9666 | 70.118 | 91.2716 | 59.4825
2012-10-30 17:17 | 26.0246 | 113.5311 | 84.3433 | 105.986 | 73.6165 | -1.6049 | 1.2887 | 32.1423 | 41.4227 | 24.9412 | 2.0037 | 32.0719 | 64.2622 | 16.0064 | 13.6003 | -0.1106 | 105.6849 | 99.9308 | 70.743 | 92.3857 | 60.0162
2012-10-30 17:17 | 26.0186 | 114.3924 | 85.0431 | 106.9027 | 74.3399 | -1.6014 | 1.2886 | 32.1422 | 41.4179 | 24.9438 | 2.0037 | 32.0717 | 64.2617 | 16.0064 | 13.574 | -0.1104 | 105.6796 | 100.8184 | 71.4691 | 93.3287 | 60.7659
2012-10-30 17:17 | 26.0268 | 115.3592 | 85.6772 | 107.8602 | 74.7642 | -1.6009 | 1.2885 | 32.1424 | 41.4145 | 24.9462 | 2.0037 | 32.0719 | 64.263 | 16.0062 | 13.6025 | -0.1104 | 105.6775 | 101.7567 | 72.0747 | 94.2577 | 61.1617
2012-10-30 17:17 | 26.0366 | 116.3164 | 86.4227 | 108.8535 | 75.3731 | -1.5968 | 1.2884 | 32.143 | 41.4117 | 24.9488 | 2.0038 | 32.0725 | 64.2662 | 16.006 | 13.6083 | -0.1101 | 105.6779 | 102.7081 | 72.8144 | 95.2452 | 61.7648
2012-10-30 17:17 | 26.0426 | 117.3004 | 87.0998 | 109.7448 | 76.0236 | -1.5968 | 1.2883 | 32.1436 | 41.4093 | 24.9512 | 2.0038 | 32.073 | 64.2683 | 16.006 | 13.6265 | -0.1101 | 105.6776 | 103.6739 | 73.4733 | 96.1183 | 62.3971
2012-10-30 17:17 | 26.0493 | 118.3576 | 87.8342 | 110.8943 | 76.6632 | -1.5974 | 1.2881 | 32.1434 | 41.4045 | 24.9538 | 2.0038 | 32.0729 | 64.2681 | 16.006 | 13.6494 | -0.1101 | 105.6726 | 104.7082 | 74.1848 | 97.2449 | 63.0138
2012-10-30 17:17 | 26.059 | 119.3916 | 88.5942 | 111.9904 | 77.4541 | -1.5962 | 1.288 | 32.1436 | 41.4008 | 24.9563 | 2.0038 | 32.0731 | 64.2694 | 16.0058 | 13.6266 | -0.11 | 105.6701 | 105.765 | 74.9676 | 98.3638 | 63.8275
2012-10-30 17:17 | 26.0635 | 120.1589 | 89.2427 | 112.7104 | 78.1633 | -1.5962 | 1.2879 | 32.144 | 41.3981 | 24.9585 | 2.0038 | 32.0734 | 64.2704 | 16.0059 | 13.6026 | -0.11 | 105.6685 | 106.5563 | 75.6401 | 99.1078 | 64.5607
2012-10-30 17:17 | 26.0725 | 121.0182 | 89.9069 | 113.7797 | 78.5645 | -1.5962 | 1.2878 | 32.1441 | 41.3949 | 24.9607 | 2.0039 | 32.0735 | 64.2713 | 16.0057 | 13.6482 | -0.11 | 105.6663 | 107.37 | 76.2587 | 100.1315 | 64.9163
2012-10-30 17:17 | 26.0755 | 121.9136 | 90.5877 | 114.8698 | 79.2754 | -1.5962 | 1.2876 | 32.1442 | 41.3904 | 24.9635 | 2.0039 | 32.0735 | 64.2717 | 16.0056 | 13.6634 | -0.11 | 105.6621 | 108.2502 | 76.9243 | 101.2064 | 65.612
2012-10-30 17:17 | 26.0807 | 122.9826 | 91.3388 | 115.7238 | 79.9731 | -1.5985 | 1.2875 | 32.1441 | 41.3859 | 24.9661 | 2.0039 | 32.0734 | 64.272 | 16.0055 | 13.6117 | -0.1102 | 105.6579 | 109.3709 | 77.7271 | 102.1121 | 66.3614
2012-10-30 17:17 | 26.0882 | 123.9555 | 92.1071 | 116.857 | 80.5663 | -1.5974 | 1.2874 | 32.1443 | 41.3823 | 24.9686 | 2.0039 | 32.0735 | 64.2719 | 16.0056 | 13.6517 | -0.1101 | 105.6541 | 110.3038 | 78.4554 | 103.2053 | 66.9146
2012-10-30 17:17 | 26.1084 | 124.7513 | 92.7328 | 117.9388 | 81.1264 | -1.5928 | 1.2873 | 32.1442 | 41.3788 | 24.9705 | 2.0039 | 32.0735 | 64.2729 | 16.0054 | 13.6678 | -0.1098 | 105.6517 | 111.0835 | 79.065 | 104.271 | 67.4586
2012-10-30 17:17 | 26.1024 | 125.643 | 93.4344 | 118.6231 | 81.8768 | -1.5997 | 1.2872 | 32.1441 | 41.3745 | 24.9729 | 2.0039 | 32.0734 | 64.2721 | 16.0054 | 13.6293 | -0.1103 | 105.6466 | 112.0137 | 79.8051 | 104.9938 | 68.2475
2012-10-30 17:18 | 26.1084 | 126.6383 | 94.1403 | 119.6535 | 82.5249 | -1.6003 | 1.287 | 32.1441 | 41.3704 | 24.9754 | 2.0039 | 32.0734 | 64.2724 | 16.0053 | 13.6597 | -0.1103 | 105.6428 | 112.9786 | 80.4806 | 105.9938 | 68.8652
2012-10-30 17:18 | 26.1166 | 127.6021 | 94.9013 | 120.7548 | 83.2167 | -1.5968 | 1.2869 | 32.1442 | 41.367 | 24.9776 | 2.0039 | 32.0733 | 64.2726 | 16.0053 | 13.7167 | -0.1101 | 105.6396 | 113.8854 | 81.1846 | 107.0381 | 69.5
2012-10-30 17:18 | 26.1181 | 128.4187 | 95.5496 | 121.8461 | 83.9282 | -1.602 | 1.2868 | 32.1441 | 41.363 | 24.98 | 2.0039 | 32.0733 | 64.2726 | 16.0052 | 13.6653 | -0.1104 | 105.6356 | 114.7534 | 81.8843 | 108.1808 | 70.2629
2012-10-30 17:18 | 26.1279 | 129.2553 | 96.1531 | 122.5624 | 84.4732 | -1.6009 | 1.2867 | 32.1439 | 41.3594 | 24.9817 | 2.0039 | 32.0731 | 64.2721 | 16.0051 | 13.6792 | -0.1104 | 105.6315 | 115.5761 | 82.4739 | 108.8832 | 70.794
2012-10-30 17:18 | 26.1301 | 129.9856 | 96.7907 | 123.5906 | 84.8966 | -1.602 | 1.2866 | 32.1437 | 41.3549 | 24.9842 | 2.0039 | 32.0729 | 64.2717 | 16.005 | 13.6855 | -0.1104 | 105.6266 | 116.3001 | 83.1052 | 109.9051 | 71.2111
2012-10-30 17:18 | 26.1324 | 130.8573 | 97.5039 | 124.357 | 85.578 | -1.5997 | 1.2865 | 32.1436 | 41.3514 | 24.9862 | 2.0039 | 32.0728 | 64.2711 | 16.0051 | 13.7203 | -0.1103 | 105.6225 | 117.137 | 83.7836 | 110.6367 | 71.8577
2012-10-30 17:18 | 26.1369 | 131.8041 | 98.2532 | 125.2856 | 86.3145 | -1.6009 | 1.2864 | 32.1436 | 41.3479 | 24.9882 | 2.0039 | 32.0728 | 64.2713 | 16.0051 | 13.7045 | -0.1104 | 105.6192 | 118.0996 | 84.5487 | 111.5811 | 72.61
2012-10-30 17:18 | 26.1376 | 132.6407 | 98.9156 | 126.3174 | 86.7929 | -1.5991 | 1.2862 | 32.1438 | 41.3444 | 24.9907 | 2.0039 | 32.0728 | 64.2715 | 16.005 | 13.6726 | -0.1102 | 105.6158 | 118.9681 | 85.243 | 112.6448 | 73.1203
2012-10-30 17:18 | 26.1563 | 133.4384 | 99.4976 | 127.1468 | 87.4712 | -1.6009 | 1.2861 | 32.1439 | 41.341 | 24.9929 | 2.0039 | 32.0729 | 64.272 | 16.005 | 13.671 | -0.1104 | 105.613 | 119.7674 | 85.8266 | 113.4758 | 73.8002
2012-10-30 17:18 | 26.1571 | 134.3026 | 100.1789 | 127.9244 | 87.9723 | -1.5951 | 1.286 | 32.1438 | 41.3371 | 24.995 | 2.0039 | 32.073 | 64.2724 | 16.005 | 13.6799 | -0.11 | 105.6095 | 120.6227 | 86.499 | 114.2445 | 74.2924
2012-10-30 17:18 | 26.1623 | 135.1257 | 100.8041 | 128.9391 | 88.58 | -1.5991 | 1.2859 | 32.1443 | 41.3341 | 24.9976 | 2.004 | 32.0734 | 64.2744 | 16.0048 | 13.7095 | -0.1102 | 105.6085 | 121.4162 | 87.0946 | 115.2296 | 74.8705
2012-10-30 17:18 | 26.1773 | 135.936 | 101.5412 | 129.8726 | 89.311 | -1.5997 | 1.2858 | 32.1443 | 41.3313 | 24.9993 | 2.004 | 32.0734 | 64.2738 | 16.005 | 13.7612 | -0.1103 | 105.6051 | 122.1748 | 87.78 | 116.1114 | 75.5498
2012-10-30 17:18 | 26.181 | 136.7201 | 102.1536 | 130.6275 | 89.8677 | -1.5997 | 1.2857 | 32.1442 | 41.3278 | 25.0013 | 2.004 | 32.0732 | 64.2737 | 16.0048 | 13.7404 | -0.1103 | 105.6015 | 122.9797 | 88.4132 | 116.8871 | 76.1273
2012-10-30 17:18 | 26.1885 | 137.493 | 102.8 | 131.539 | 90.3751 | -1.5928 | 1.2856 | 32.144 | 41.3238 | 25.0034 | 2.0039 | 32.0729 | 64.2724 | 16.0049 | 13.7479 | -0.1098 | 105.5962 | 123.7451 | 89.0521 | 117.7911 | 76.6272
2012-10-30 17:18 | 26.1982 | 138.2953 | 103.441 | 132.3387 | 91.0777 | -1.602 | 1.2855 | 32.1437 | 41.32 | 25.0053 | 2.0039 | 32.0727 | 64.2716 | 16.0049 | 13.7617 | -0.1104 | 105.5916 | 124.5336 | 89.6793 | 118.577 | 77.316
2012-10-30 17:18 | 26.2065 | 138.9815 | 104.0353 | 133.0953 | 91.5971 | -1.6026 | 1.2854 | 32.1437 | 41.3165 | 25.0074 | 2.0039 | 32.0726 | 64.2718 | 16.0048 | 13.7659 | -0.1105 | 105.5883 | 125.2156 | 90.2694 | 119.3294 | 77.8312
2012-10-30 17:18 | 26.2132 | 139.7086 | 104.6318 | 134.0454 | 92.1715 | -1.5974 | 1.2853 | 32.1437 | 41.313 | 25.0095 | 2.0039 | 32.0725 | 64.2717 | 16.0047 | 13.7442 | -0.1101 | 105.5847 | 125.9644 | 90.8876 | 120.3012 | 78.4273
2012-10-30 17:18 | 26.2199 | 140.6661 | 105.391 | 134.8584 | 92.867 | -1.6026 | 1.2851 | 32.1433 | 41.3081 | 25.0119 | 2.0039 | 32.0724 | 64.2707 | 16.0047 | 13.7997 | -0.1105 | 105.5788 | 126.8664 | 91.5913 | 121.0587 | 79.0673
2012-10-30 17:18 | 26.2259 | 141.5044 | 106.0663 | 135.7513 | 93.5238 | -1.5985 | 1.285 | 32.1431 | 41.3044 | 25.0138 | 2.0039 | 32.0721 | 64.2697 | 16.0047 | 13.765 | -0.1102 | 105.5741 | 127.7394 | 92.3013 | 121.9863 | 79.7588
2012-10-30 17:18 | 26.2371 | 142.2255 | 106.6069 | 136.6614 | 93.9588 | -1.6003 | 1.2849 | 32.1428 | 41.3009 | 25.0155 | 2.0039 | 32.0718 | 64.2688 | 16.0046 | 13.8169 | -0.1103 | 105.5697 | 128.4086 | 92.79 | 122.8445 | 80.1419
2012-10-30 17:18 | 26.2401 | 142.9732 | 107.2717 | 137.403 | 94.4914 | -1.5974 | 1.2848 | 32.1425 | 41.297 | 25.0174 | 2.0039 | 32.0715 | 64.2679 | 16.0046 | 13.828 | -0.1101 | 105.5649 | 129.1452 | 93.4437 | 123.575 | 80.6634
2012-10-30 17:18 | 26.2454 | 143.7813 | 107.9805 | 138.1586 | 95.083 | -1.5997 | 1.2847 | 32.1426 | 41.2934 | 25.0196 | 2.0039 | 32.0714 | 64.2678 | 16.0045 | 13.8008 | -0.1103 | 105.5612 | 129.9805 | 94.1797 | 124.3578 | 81.2822
2012-10-30 17:18 | 26.2521 | 144.6102 | 108.6113 | 139.1076 | 95.6117 | -1.6003 | 1.2846 | 32.1428 | 41.2894 | 25.0224 | 2.0039 | 32.0715 | 64.2681 | 16.0046 | 13.7953 | -0.1103 | 105.5575 | 130.8149 | 94.816 | 125.3123 | 81.8164
2012-10-30 17:18 | 26.2686 | 145.4494 | 109.2898 | 139.9825 | 96.3544 | -1.5974 | 1.2845 | 32.143 | 41.2875 | 25.0238 | 2.0039 | 32.0719 | 64.2689 | 16.0047 | 13.7629 | -0.1101 | 105.5564 | 131.6865 | 95.5269 | 126.2196 | 82.5915
2012-10-30 17:18 | 26.2648 | 146.1302 | 109.9443 | 140.7033 | 96.9443 | -1.5933 | 1.2844 | 32.1432 | 41.2847 | 25.0259 | 2.0039 | 32.0721 | 64.2705 | 16.0045 | 13.7632 | -0.1098 | 105.5552 | 132.367 | 96.1811 | 126.9401 | 83.1811
2012-10-30 17:18 | 26.2798 | 146.8455 | 110.4512 | 141.5621 | 97.2883 | -1.6003 | 1.2843 | 32.1429 | 41.281 | 25.0277 | 2.0039 | 32.0717 | 64.2689 | 16.0046 | 13.7741 | -0.1103 | 105.5499 | 133.0714 | 96.6771 | 127.788 | 83.5142
2012-10-30 17:18 | 26.2813 | 147.5321 | 111.0658 | 142.1891 | 98.0044 | -1.6003 | 1.2842 | 32.143 | 41.2778 | 25.0297 | 2.0039 | 32.0717 | 64.2687 | 16.0046 | 13.7309 | -0.1103 | 105.5465 | 133.8012 | 97.3349 | 128.4582 | 84.2735
2012-10-30 17:19 | 26.2933 | 148.3066 | 111.7401 | 143.0572 | 98.6743 | -1.5991 | 1.2841 | 32.1428 | 41.2741 | 25.0317 | 2.0039 | 32.0715 | 64.2683 | 16.0045 | 13.751 | -0.1102 | 105.5423 | 134.5556 | 97.9891 | 129.3062 | 84.9233
2012-10-30 17:19 | 26.3015 | 149.1046 | 112.3464 | 143.9334 | 99.1513 | -1.5951 | 1.284 | 32.1427 | 41.2703 | 25.0339 | 2.0039 | 32.0714 | 64.2678 | 16.0046 | 13.7226 | -0.11 | 105.5381 | 135.382 | 98.6238 | 130.2108 | 85.4287
2012-10-30 17:19 | 26.2978 | 149.8717 | 113.0053 | 144.752 | 99.866 | -1.5957 | 1.2839 | 32.1427 | 41.267 | 25.0358 | 2.0039 | 32.0715 | 64.2682 | 16.0045 | 13.7229 | -0.11 | 105.5352 | 136.1488 | 99.2824 | 131.0291 | 86.1431
2012-10-30 17:19 | 26.3052 | 150.6118 | 113.6261 | 145.3861 | 100.5728 | -1.598 | 1.2838 | 32.143 | 41.2653 | 25.0374 | 2.0039 | 32.0717 | 64.2689 | 16.0045 | 13.7507 | -0.1102 | 105.5342 | 136.8611 | 99.8754 | 131.6354 | 86.8221
2012-10-30 17:19 | 26.3247 | 151.3102 | 114.1152 | 146.1709 | 100.9174 | -1.6043 | 1.2837 | 32.1428 | 41.2617 | 25.0393 | 2.0039 | 32.0715 | 64.2677 | 16.0047 | 13.7132 | -0.1106 | 105.5294 | 137.597 | 100.402 | 132.4577 | 87.2042
2012-10-30 17:19 | 26.3269 | 152.0881 | 114.731 | 147.0004 | 101.4678 | -1.5957 | 1.2836 | 32.1425 | 41.2576 | 25.0412 | 2.0039 | 32.0712 | 64.2662 | 16.0047 | 13.7398 | -0.11 | 105.5237 | 138.3483 | 100.9912 | 133.2606 | 87.728
2012-10-30 17:19 | 26.3314 | 152.7103 | 115.3606 | 147.667 | 102.0922 | -1.602 | 1.2835 | 32.1425 | 41.2544 | 25.0432 | 2.0039 | 32.0712 | 64.2668 | 16.0045 | 13.7199 | -0.1104 | 105.5212 | 138.9904 | 101.6407 | 133.9471 | 88.3723
2012-10-30 17:19 | 26.3501 | 153.3547 | 115.925 | 148.3711 | 102.4852 | -1.598 | 1.2834 | 32.142 | 41.2508 | 25.0445 | 2.0038 | 32.0705 | 64.2642 | 16.0046 | 13.702 | -0.1102 | 105.515 | 139.6527 | 102.223 | 134.6691 | 88.7832
2012-10-30 17:19 | 26.3501 | 153.9844 | 116.4328 | 149.1532 | 103.0069 | -1.5968 | 1.2833 | 32.1417 | 41.2469 | 25.0464 | 2.0038 | 32.0703 | 64.2634 | 16.0045 | 13.7183 | -0.1101 | 105.5103 | 140.2661 | 102.7145 | 135.4349 | 89.2886
2012-10-30 17:19 | 26.3524 | 154.7288 | 117.0833 | 149.6334 | 103.6028 | -1.5968 | 1.2832 | 32.1417 | 41.2436 | 25.0484 | 2.0038 | 32.0703 | 64.2631 | 16.0046 | 13.745 | -0.1101 | 105.5067 | 140.9838 | 103.3383 | 135.8884 | 89.8578
2012-10-30 17:19 | 26.3741 | 155.4736 | 117.6623 | 150.5187 | 104.2067 | -1.6003 | 1.2831 | 32.1419 | 41.2413 | 25.0502 | 2.0038 | 32.0705 | 64.264 | 16.0046 | 13.7626 | -0.1103 | 105.5052 | 141.711 | 103.8997 | 136.7561 | 90.4441
2012-10-30 17:19 | 26.3771 | 156.08 | 118.2189 | 151.3973 | 104.6861 | -1.5985 | 1.283 | 32.1417 | 41.2381 | 25.0518 | 2.0038 | 32.0701 | 64.2625 | 16.0046 | 13.7086 | -0.1102 | 105.5007 | 142.3714 | 104.5103 | 137.6887 | 90.9775
2012-10-30 17:19 | 26.3831 | 156.7012 | 118.7255 | 151.7815 | 105.2744 | -1.5974 | 1.283 | 32.1414 | 41.2358 | 25.0527 | 2.0038 | 32.07 | 64.2618 | 16.0046 | 13.7513 | -0.1101 | 105.4976 | 142.9499 | 104.9742 | 138.0302 | 91.5231
2012-10-30 17:19 | 26.3846 | 157.2446 | 119.2505 | 152.4701 | 105.5859 | -1.5997 | 1.2828 | 32.1415 | 41.2325 | 25.0549 | 2.0038 | 32.0701 | 64.2624 | 16.0045 | 13.7365 | -0.1103 | 105.4949 | 143.5081 | 105.514 | 138.7336 | 91.8494
2012-10-30 17:19 | 26.3936 | 157.9428 | 119.7793 | 153.2746 | 106.1848 | -1.5974 | 1.2827 | 32.1412 | 41.2287 | 25.0567 | 2.0038 | 32.0698 | 64.2615 | 16.0045 | 13.7821 | -0.1101 | 105.4901 | 144.1607 | 105.9972 | 139.4925 | 92.4027
2012-10-30 17:19 | 26.4048 | 158.6977 | 120.4584 | 154.0662 | 106.7821 | -1.5991 | 1.2826 | 32.1412 | 41.2258 | 25.0584 | 2.0038 | 32.0698 | 64.2604 | 16.0047 | 13.7567 | -0.1102 | 105.4862 | 144.941 | 106.7017 | 140.3095 | 93.0254
2012-10-30 17:19 | 26.407 | 159.4207 | 121.011 | 154.7138 | 107.3213 | -1.6055 | 1.2826 | 32.1411 | 41.223 | 25.06 | 2.0038 | 32.0697 | 64.2602 | 16.0047 | 13.7589 | -0.1107 | 105.4833 | 145.6618 | 107.2521 | 140.9549 | 93.5624
2012-10-30 17:19 | 26.413 | 160.0508 | 121.5372 | 155.4428 | 107.7358 | -1.6061 | 1.2824 | 32.1407 | 41.2188 | 25.062 | 2.0038 | 32.0692 | 64.2587 | 16.0046 | 13.8056 | -0.1107 | 105.4775 | 146.2452 | 107.7316 | 141.6372 | 93.9302
2012-10-30 17:19 | 26.4242 | 160.7407 | 122.099 | 156.1503 | 108.4051 | -1.5991 | 1.2823 | 32.1403 | 41.215 | 25.0638 | 2.0037 | 32.0688 | 64.2569 | 16.0046 | 13.7761 | -0.1102 | 105.4719 | 146.9646 | 108.3229 | 142.3742 | 94.629
2012-10-30 17:19 | 26.437 | 161.4634 | 122.7462 | 156.9293 | 109.0269 | -1.5951 | 1.2822 | 32.14 | 41.2115 | 25.0653 | 2.0037 | 32.0684 | 64.2562 | 16.0044 | 13.7563 | -0.11 | 105.4677 | 147.7071 | 108.9899 | 143.173 | 95.2706
2012-10-30 17:19 | 26.4385 | 162.078 | 123.2644 | 157.4611 | 109.403 | -1.6026 | 1.2822 | 32.14 | 41.2084 | 25.0673 | 2.0037 | 32.0685 | 64.2559 | 16.0046 | 13.7944 | -0.1105 | 105.4643 | 148.2836 | 109.47 | 143.6667 | 95.6086
2012-10-30 17:19 | 26.4482 | 162.6661 | 123.7719 | 158.1988 | 109.8288 | -1.6014 | 1.2821 | 32.1397 | 41.2052 | 25.0687 | 2.0037 | 32.0683 | 64.2547 | 16.0047 | 13.7919 | -0.1104 | 105.4599 | 148.8742 | 109.98 | 144.4069 | 96.0369
2012-10-30 17:19 | 26.4587 | 163.3018 | 124.2993 | 158.8795 | 110.388 | -1.5997 | 1.282 | 32.1394 | 41.2015 | 25.0705 | 2.0037 | 32.0678 | 64.2528 | 16.0047 | 13.7739 | -0.1103 | 105.4543 | 149.5279 | 110.5254 | 145.1056 | 96.6141
2012-10-30 17:19 | 26.4617 | 164.0002 | 124.865 | 159.4987 | 110.7694 | -1.5968 | 1.2819 | 32.1393 | 41.1983 | 25.0722 | 2.0037 | 32.0676 | 64.2525 | 16.0046 | 13.7891 | -0.1101 | 105.4508 | 150.2111 | 111.0759 | 145.7096 | 96.9803
2012-10-30 17:19 | 26.4624 | 164.6731 | 125.4019 | 160.1816 | 111.1031 | -1.602 | 1.2818 | 32.1396 | 41.1964 | 25.0739 | 2.0037 | 32.068 | 64.2536 | 16.0047 | 13.8265 | -0.1104 | 105.4499 | 150.8466 | 111.5754 | 146.3551 | 97.2766
2012-10-30 17:19 | 26.4789 | 165.3194 | 125.9205 | 160.935 | 111.57 | -1.6003 | 1.2817 | 32.1397 | 41.1944 | 25.0753 | 2.0037 | 32.0681 | 64.2539 | 16.0047 | 13.8063 | -0.1103 | 105.4484 | 151.5131 | 112.1142 | 147.1287 | 97.7637
2012-10-30 17:19 | 26.4864 | 165.8946 | 126.3504 | 161.4779 | 112.1301 | -1.6032 | 1.2817 | 32.1395 | 41.1917 | 25.0765 | 2.0036 | 32.0678 | 64.2522 | 16.0048 | 13.8179 | -0.1105 | 105.4439 | 152.0767 | 112.5325 | 147.66 | 98.3122
2012-10-30 17:20 | 26.4916 | 166.5104 | 126.8541 | 161.9875 | 112.4911 | -1.5957 | 1.2816 | 32.1392 | 41.1886 | 25.078 | 2.0036 | 32.0675 | 64.2514 | 16.0047 | 13.8557 | -0.11 | 105.44 | 152.6547 | 112.9984 | 148.1318 | 98.6354
2012-10-30 17:20 | 26.4946 | 167.1204 | 127.31 | 162.7436 | 113.0544 | -1.6072 | 1.2815 | 32.1389 | 41.1848 | 25.0799 | 2.0036 | 32.0672 | 64.2502 | 16.0047 | 13.8261 | -0.1108 | 105.435 | 153.2943 | 113.4839 | 148.9175 | 99.2283
2012-10-30 17:20 | 26.5036 | 167.8027 | 127.8482 | 163.4555 | 113.5254 | -1.6072 | 1.2814 | 32.1387 | 41.1821 | 25.0812 | 2.0036 | 32.067 | 64.2489 | 16.0048 | 13.8229 | -0.1108 | 105.431 | 153.9798 | 114.0253 | 149.6326 | 99.7025
2012-10-30 17:20 | 26.5148 | 168.3629 | 128.385 | 163.9289 | 113.9834 | -1.5991 | 1.2813 | 32.1383 | 41.1783 | 25.0828 | 2.0036 | 32.0665 | 64.2471 | 16.0048 | 13.8341 | -0.1102 | 105.4254 | 154.5288 | 114.5509 | 150.0948 | 100.1493
2012-10-30 17:20 | 26.5186 | 168.9157 | 128.8878 | 164.6477 | 114.3773 | -1.5997 | 1.2812 | 32.1381 | 41.1749 | 25.0846 | 2.0035 | 32.0664 | 64.2465 | 16.0048 | 13.842 | -0.1103 | 105.4214 | 155.0737 | 115.0458 | 150.8057 | 100.5353
2012-10-30 17:20 | 26.5245 | 169.5254 | 129.3852 | 165.3572 | 114.8832 | -1.6003 | 1.2811 | 32.1378 | 41.1715 | 25.0862 | 2.0035 | 32.0661 | 64.2451 | 16.0048 | 13.8275 | -0.1103 | 105.4166 | 155.6979 | 115.5577 | 151.5297 | 101.0557
2012-10-30 17:20 | 26.5328 | 170.1727 | 129.8927 | 165.9682 | 115.4143 | -1.598 | 1.281 | 32.1375 | 41.1685 | 25.0876 | 2.0035 | 32.0658 | 64.2438 | 16.0049 | 13.8684 | -0.1102 | 105.4123 | 156.3043 | 116.0243 | 152.0998 | 101.5459
2012-10-30 17:20 | 26.547 | 170.7599 | 130.4139 | 166.4764 | 115.8794 | -1.6003 | 1.2809 | 32.1373 | 41.1657 | 25.089 | 2.0035 | 32.0656 | 64.2432 | 16.0049 | 13.8663 | -0.1103 | 105.4089 | 156.8936 | 116.5476 | 152.6101 | 102.0131
2012-10-30 17:20 | 26.556 | 171.3342 | 130.869 | 167.139 | 116.074 | -1.598 | 1.2809 | 32.1374 | 41.1638 | 25.0904 | 2.0035 | 32.0656 | 64.2432 | 16.0048 | 13.8305 | -0.1102 | 105.407 | 157.5037 | 117.0385 | 153.3085 | 102.2435
2012-10-30 17:20 | 26.5552 | 171.8715 | 131.2877 | 167.7492 | 116.5015 | -1.602 | 1.2808 | 32.1372 | 41.1604 | 25.0921 | 2.0035 | 32.0654 | 64.2425 | 16.0049 | 13.8501 | -0.1104 | 105.4028 | 158.0214 | 117.4376 | 153.8991 | 102.6514
2012-10-30 17:20 | 26.5582 | 172.4206 | 131.6989 | 168.1921 | 116.8607 | -1.598 | 1.2807 | 32.137 | 41.1579 | 25.0933 | 2.0035 | 32.0652 | 64.2412 | 16.0049 | 13.8287 | -0.1102 | 105.3991 | 158.5919 | 117.8702 | 154.3634 | 103.032
2012-10-30 17:20 | 26.5702 | 173.0139 | 132.1764 | 168.8972 | 117.3238 | -1.5962 | 1.2806 | 32.1369 | 41.1556 | 25.0946 | 2.0034 | 32.0652 | 64.2411 | 16.005 | 13.8732 | -0.11 | 105.3967 | 159.1407 | 118.3032 | 155.024 | 103.4506
2012-10-30 17:20 | 26.5829 | 173.5144 | 132.5631 | 169.5604 | 117.8711 | -1.5985 | 1.2806 | 32.1368 | 41.1536 | 25.0956 | 2.0034 | 32.065 | 64.24 | 16.0051 | 13.8493 | -0.1102 | 105.3935 | 159.6651 | 118.7138 | 155.7111 | 104.0218
2012-10-30 17:20 | 26.5927 | 174.0785 | 133.0118 | 169.9326 | 118.3511 | -1.602 | 1.2805 | 32.1364 | 41.1497 | 25.0973 | 2.0034 | 32.0646 | 64.2381 | 16.0051 | 13.8265 | -0.1104 | 105.3878 | 160.252 | 119.1853 | 156.1061 | 104.5246
2012-10-30 17:20 | 26.5964 | 174.5976 | 133.4508 | 170.5559 | 118.5899 | -1.5945 | 1.2804 | 32.1362 | 41.1466 | 25.0989 | 2.0034 | 32.0644 | 64.237 | 16.0052 | 13.8587 | -0.1099 | 105.3837 | 160.7389 | 119.5921 | 156.6972 | 104.7312
2012-10-30 17:20 | 26.6091 | 175.0555 | 133.853 | 171.1767 | 118.9515 | -1.5957 | 1.2803 | 32.1359 | 41.1442 | 25.1 | 2.0034 | 32.0641 | 64.2359 | 16.0052 | 13.8796 | -0.11 | 105.3801 | 161.1759 | 119.9734 | 157.2971 | 105.0719
2012-10-30 17:20 | 26.6114 | 175.6359 | 134.3557 | 171.643 | 119.6367 | -1.602 | 1.2802 | 32.1357 | 41.1408 | 25.1018 | 2.0033 | 32.064 | 64.2353 | 16.0052 | 13.8656 | -0.1104 | 105.3761 | 161.7703 | 120.4901 | 157.7774 | 105.7711
2012-10-30 17:20 | 26.6203 | 176.3089 | 134.8881 | 172.2291 | 120.1502 | -1.5985 | 1.2802 | 32.1359 | 41.1393 | 25.1029 | 2.0033 | 32.0641 | 64.2353 | 16.0053 | 13.8867 | -0.1102 | 105.3746 | 162.4222 | 121.0014 | 158.3424 | 106.2635
2012-10-30 17:20 | 26.6226 | 176.8115 | 135.3299 | 172.9337 | 120.5396 | -1.602 | 1.2801 | 32.1356 | 41.1362 | 25.1043 | 2.0033 | 32.0637 | 64.2338 | 16.0053 | 13.8564 | -0.1104 | 105.37 | 162.9551 | 121.4735 | 159.0773 | 106.6832
2012-10-30 17:20 | 26.6421 | 177.3586 | 135.7699 | 173.398 | 121.0449 | -1.6026 | 1.28 | 32.1354 | 41.1336 | 25.1056 | 2.0033 | 32.0635 | 64.2334 | 16.0052 | 13.9248 | -0.1105 | 105.367 | 163.4338 | 121.8451 | 159.4732 | 107.1201
2012-10-30 17:20 | 26.6458 | 177.9097 | 136.2783 | 173.9442 | 121.5042 | -1.5939 | 1.2799 | 32.1352 | 41.1309 | 25.107 | 2.0033 | 32.0632 | 64.232 | 16.0053 | 13.8674 | -0.1099 | 105.3628 | 164.0423 | 122.4109 | 160.0768 | 107.6368
2012-10-30 17:20 | 26.6548 | 178.4192 | 136.7463 | 174.5271 | 121.8208 | -1.5974 | 1.2799 | 32.1351 | 41.1282 | 25.1084 | 2.0033 | 32.0631 | 64.2318 | 16.0052 | 13.8764 | -0.1101 | 105.36 | 164.5428 | 122.8699 | 160.6507 | 107.9444
2012-10-30 17:20 | 26.66 | 178.9094 | 137.157 | 175.0986 | 122.2779 | -1.5991 | 1.2798 | 32.1347 | 41.1253 | 25.1096 | 2.0033 | 32.0629 | 64.2303 | 16.0054 | 13.8775 | -0.1102 | 105.3556 | 165.0319 | 123.2795 | 161.2211 | 108.4004
2012-10-30 17:20 | 26.6817 | 179.4514 | 137.5958 | 175.4835 | 122.6642 | -1.598 | 1.2797 | 32.1346 | 41.1228 | 25.1109 | 2.0032 | 32.0625 | 64.229 | 16.0053 | 13.9033 | -0.1102 | 105.3518 | 165.5481 | 123.6925 | 161.5802 | 108.7609
2012-10-30 17:20 | 26.681 | 179.8342 | 137.9625 | 176.0336 | 122.9028 | -1.6009 | 1.2796 | 32.1345 | 41.1207 | 25.1121 | 2.0032 | 32.0625 | 64.2288 | 16.0054 | 13.8945 | -0.1104 | 105.3495 | 165.9397 | 124.068 | 162.1391 | 109.0083
2012-10-30 17:20 | 26.6705 | 180.3637 | 138.3693 | 176.6053 | 123.3786 | -1.6055 | 1.2796 | 32.1345 | 41.119 | 25.1131 | 2.0032 | 32.0625 | 64.2285 | 16.0054 | 13.888 | -0.1107 | 105.3475 | 166.4757 | 124.4813 | 162.7173 | 109.4906
2012-10-30 17:20 | 26.684 | 180.9369 | 138.8117 | 177.0722 | 123.8177 | -1.6026 | 1.2795 | 32.1342 | 41.1159 | 25.1146 | 2.0032 | 32.0622 | 64.2275 | 16.0054 | 13.9097 | -0.1105 | 105.3435 | 167.0272 | 124.902 | 163.1625 | 109.908
2012-10-30 17:21 | 26.6877 | 181.3791 | 139.2099 | 177.6271 | 124.1064 | -1.6014 | 1.2794 | 32.1339 | 41.1137 | 25.1155 | 2.0032 | 32.0619 | 64.2258 | 16.0055 | 13.9012 | -0.1104 | 105.3395 | 167.4779 | 125.3087 | 163.7259 | 110.2052
2012-10-30 17:21 | 26.6967 | 181.8971 | 139.6671 | 178.2226 | 124.564 | -1.5933 | 1.2793 | 32.1336 | 41.1098 | 25.1173 | 2.0032 | 32.0616 | 64.2245 | 16.0055 | 13.9264 | -0.1098 | 105.3343 | 167.9707 | 125.7407 | 164.2962 | 110.6376
2012-10-30 17:21 | 26.7109 | 182.4008 | 140.053 | 178.6122 | 125.0686 | -1.5985 | 1.2793 | 32.1335 | 41.1077 | 25.1184 | 2.0032 | 32.0614 | 64.2241 | 16.0054 | 13.9569 | -0.1102 | 105.3318 | 168.4439 | 126.0961 | 164.6553 | 111.1117
2012-10-30 17:21 | 26.7154 | 182.8636 | 140.4408 | 179.1839 | 125.2062 | -1.6003 | 1.2792 | 32.1332 | 41.1053 | 25.1195 | 2.0031 | 32.0613 | 64.2232 | 16.0055 | 13.9166 | -0.1103 | 105.3284 | 168.947 | 126.5242 | 165.2673 | 111.2896
2012-10-30 17:21 | 26.7304 | 183.3291 | 140.843 | 179.684 | 125.6319 | -1.6032 | 1.2791 | 32.1327 | 41.1022 | 25.1205 | 2.0031 | 32.0607 | 64.2204 | 16.0056 | 13.9113 | -0.1105 | 105.3226 | 169.4178 | 126.9317 | 165.7727 | 111.7206
2012-10-30 17:21 | 26.7244 | 183.7855 | 141.1821 | 179.9636 | 125.9378 | -1.6003 | 1.2791 | 32.1325 | 41.1 | 25.1215 | 2.0031 | 32.0605 | 64.2194 | 16.0057 | 13.8931 | -0.1103 | 105.3194 | 169.8924 | 127.289 | 166.0705 | 112.0447
2012-10-30 17:21 | 26.7221 | 184.1868 | 141.5364 | 180.5601 | 126.2741 | -1.5951 | 1.279 | 32.1326 | 41.0982 | 25.1229 | 2.0031 | 32.0606 | 64.2192 | 16.0058 | 13.9152 | -0.11 | 105.3174 | 170.2716 | 127.6212 | 166.6449 | 112.3589
2012-10-30 17:21 | 26.7424 | 184.7127 | 141.994 | 181.0814 | 126.7624 | -1.6003 | 1.2789 | 32.1325 | 41.0958 | 25.1242 | 2.0031 | 32.0606 | 64.2192 | 16.0058 | 13.8948 | -0.1103 | 105.3151 | 170.8179 | 128.0992 | 167.1866 | 112.8676
2012-10-30 17:21 | 26.7454 | 185.2292 | 142.4268 | 181.5151 | 127.2487 | -1.5997 | 1.2789 | 32.1325 | 41.0935 | 25.1255 | 2.0031 | 32.0603 | 64.2187 | 16.0057 | 13.9059 | -0.1103 | 105.3122 | 171.3233 | 128.5209 | 167.6092 | 113.3428
2012-10-30 17:21 | 26.7461 | 185.6416 | 142.7498 | 182.0914 | 127.4816 | -1.6037 | 1.2788 | 32.1326 | 41.0922 | 25.1265 | 2.0031 | 32.0605 | 64.219 | 16.0058 | 13.8659 | -0.1106 | 105.3112 | 171.7757 | 128.8839 | 168.2255 | 113.6157
2012-10-30 17:21 | 26.7573 | 186.1076 | 143.1795 | 182.5452 | 127.9154 | -1.6066 | 1.2788 | 32.1325 | 41.0901 | 25.1276 | 2.0031 | 32.0604 | 64.2189 | 16.0057 | 13.8527 | -0.1108 | 105.309 | 172.2549 | 129.3268 | 168.6925 | 114.0627
2012-10-30 17:21 | 26.7528 | 186.5659 | 143.5512 | 182.9754 | 128.2117 | -1.5991 | 1.2787 | 32.1323 | 41.088 | 25.1287 | 2.003 | 32.0601 | 64.2175 | 16.0058 | 13.9052 | -0.1102 | 105.3055 | 172.6607 | 129.646 | 169.0702 | 114.3065
2012-10-30 17:21 | 26.7701 | 187.0813 | 143.9449 | 183.5647 | 128.526 | -1.5997 | 1.2786 | 32.1324 | 41.086 | 25.13 | 2.003 | 32.0603 | 64.218 | 16.0059 | 13.9265 | -0.1103 | 105.3041 | 173.1548 | 130.0184 | 169.6382 | 114.5995
2012-10-30 17:21 | 26.7715 | 187.5731 | 144.3535 | 184.0442 | 128.9502 | -1.5957 | 1.2786 | 32.1322 | 41.0842 | 25.1308 | 2.003 | 32.0601 | 64.2172 | 16.0059 | 13.9198 | -0.11 | 105.3014 | 173.6533 | 130.4337 | 170.1244 | 115.0304
2012-10-30 17:21 | 26.779 | 188.0474 | 144.7452 | 184.3887 | 129.4758 | -1.6043 | 1.2785 | 32.1318 | 41.0815 | 25.1318 | 2.003 | 32.0598 | 64.2158 | 16.0058 | 13.8825 | -0.1106 | 105.2973 | 174.1649 | 130.8627 | 170.5062 | 115.5933
2012-10-30 17:21 | 26.7925 | 188.5043 | 145.0546 | 184.998 | 129.6875 | -1.6009 | 1.2785 | 32.1321 | 41.0797 | 25.1333 | 2.003 | 32.0599 | 64.216 | 16.006 | 13.9042 | -0.1104 | 105.2957 | 174.6001 | 131.1504 | 171.0938 | 115.7833
2012-10-30 17:21 | 26.7918 | 188.8922 | 145.4459 | 185.3266 | 130.1098 | -1.5991 | 1.2784 | 32.132 | 41.0776 | 25.1345 | 2.003 | 32.0598 | 64.2158 | 16.0059 | 13.9035 | -0.1102 | 105.2934 | 174.9887 | 131.5424 | 171.4231 | 116.2063
2012-10-30 17:21 | 26.8097 | 189.3143 | 145.8252 | 185.7797 | 130.3564 | -1.6049 | 1.2784 | 32.1318 | 41.0757 | 25.1354 | 2.003 | 32.0596 | 64.2145 | 16.0061 | 13.844 | -0.1106 | 105.2902 | 175.4703 | 131.9812 | 171.9357 | 116.5124
2012-10-30 17:21 | 26.8112 | 189.8457 | 146.2497 | 186.3145 | 130.8077 | -1.598 | 1.2783 | 32.1314 | 41.0728 | 25.1366 | 2.0029 | 32.0593 | 64.2128 | 16.0061 | 13.9229 | -0.1102 | 105.2856 | 175.9228 | 132.3268 | 172.3916 | 116.8848
2012-10-30 17:21 | 26.821 | 190.3227 | 146.6403 | 186.8851 | 131.2065 | -1.5991 | 1.2782 | 32.1308 | 41.0692 | 25.1377 | 2.0029 | 32.0585 | 64.2098 | 16.0061 | 13.8675 | -0.1102 | 105.279 | 176.4552 | 132.7728 | 173.0176 | 117.339
2012-10-30 17:21 | 26.8157 | 190.7964 | 147.0582 | 187.2328 | 131.5253 | -1.602 | 1.2781 | 32.1305 | 41.0665 | 25.139 | 2.0029 | 32.0583 | 64.2089 | 16.0061 | 13.8663 | -0.1104 | 105.2753 | 176.9301 | 133.1919 | 173.3665 | 117.659
2012-10-30 17:21 | 26.8232 | 191.2555 | 147.3851 | 187.7699 | 131.7655 | -1.602 | 1.278 | 32.1301 | 41.0638 | 25.14 | 2.0028 | 32.0579 | 64.207 | 16.0062 | 13.8575 | -0.1104 | 105.2708 | 177.398 | 133.5276 | 173.9124 | 117.908
2012-10-30 17:21 | 26.8374 | 191.6742 | 147.8034 | 188.3609 | 132.1765 | -1.6009 | 1.278 | 32.1301 | 41.0619 | 25.1412 | 2.0028 | 32.0579 | 64.2068 | 16.0062 | 13.9083 | -0.1104 | 105.2687 | 177.7659 | 133.8951 | 174.4526 | 118.2682
2012-10-30 17:21 | 26.8442 | 192.2191 | 148.2067 | 188.7876 | 132.5572 | -1.6043 | 1.2779 | 32.1299 | 41.0592 | 25.1424 | 2.0028 | 32.0577 | 64.2057 | 16.0063 | 13.8988 | -0.1106 | 105.2649 | 178.3203 | 134.3079 | 174.8888 | 118.6584
2012-10-30 17:21 | 26.8636 | 192.7413 | 148.6073 | 189.2426 | 132.9888 | -1.5968 | 1.2778 | 32.1298 | 41.0569 | 25.1437 | 2.0028 | 32.0576 | 64.2052 | 16.0064 | 13.8816 | -0.1101 | 105.2621 | 178.8597 | 134.7257 | 175.361 | 119.1072
2012-10-30 17:21 | 26.8614 | 193.187 | 149.0403 | 189.7742 | 133.2004 | -1.5985 | 1.2778 | 32.1295 | 41.0543 | 25.1448 | 2.0028 | 32.0573 | 64.2037 | 16.0064 | 13.9079 | -0.1102 | 105.258 | 179.2791 | 135.1324 | 175.8663 | 119.2925
2012-10-30 17:22 | 26.8606 | 193.5673 | 149.34 | 190.272 | 133.4897 | -1.6014 | 1.2777 | 32.1294 | 41.0524 | 25.1459 | 2.0028 | 32.0571 | 64.2027 | 16.0064 | 13.9112 | -0.1104 | 105.2551 | 179.6561 | 135.4288 | 176.3608 | 119.5785
2012-10-30 17:22 | 26.8749 | 194.0924 | 149.7165 | 190.6632 | 133.9355 | -1.6032 | 1.2776 | 32.1287 | 41.0482 | 25.1473 | 2.0027 | 32.0564 | 64.1999 | 16.0065 | 13.8929 | -0.1105 | 105.2481 | 180.1995 | 135.8236 | 176.7703 | 120.0426
2012-10-30 17:22 | 26.8749 | 194.5522 | 150.1135 | 191.1226 | 134.3022 | -1.6014 | 1.2775 | 32.128 | 41.0445 | 25.1485 | 2.0027 | 32.0558 | 64.1973 | 16.0066 | 13.8847 | -0.1104 | 105.2418 | 180.6675 | 136.2288 | 177.2379 | 120.4175
2012-10-30 17:22 | 26.8861 | 195.0233 | 150.4793 | 191.6444 | 134.5502 | -1.6026 | 1.2775 | 32.1276 | 41.0427 | 25.149 | 2.0026 | 32.0554 | 64.1955 | 16.0066 | 13.8878 | -0.1105 | 105.2382 | 181.1355 | 136.5915 | 177.7566 | 120.6624
2012-10-30 17:22 | 26.8973 | 195.3961 | 150.8259 | 192.1308 | 134.9014 | -1.5962 | 1.2774 | 32.1276 | 41.0411 | 25.15 | 2.0026 | 32.0553 | 64.1954 | 16.0065 | 13.9031 | -0.11 | 105.2365 | 181.493 | 136.9228 | 178.2277 | 120.9983
2012-10-30 17:22 | 26.9033 | 195.8148 | 151.1057 | 192.3961 | 135.0641 | -1.6072 | 1.2774 | 32.1277 | 41.0389 | 25.1514 | 2.0026 | 32.0554 | 64.1952 | 16.0067 | 13.8928 | -0.1108 | 105.2342 | 181.922 | 137.2129 | 178.5033 | 121.1713
2012-10-30 17:22 | 26.9115 | 196.2319 | 151.4302 | 192.8992 | 135.378 | -1.6003 | 1.2773 | 32.1276 | 41.037 | 25.1525 | 2.0026 | 32.0553 | 64.1946 | 16.0066 | 13.8847 | -0.1103 | 105.2316 | 182.3472 | 137.5455 | 179.0145 | 121.4933
2012-10-30 17:22 | 26.9168 | 196.6736 | 151.7673 | 193.4035 | 135.7005 | -1.6049 | 1.2772 | 32.1272 | 41.0343 | 25.1535 | 2.0026 | 32.0549 | 64.193 | 16.0067 | 13.8737 | -0.1106 | 105.2273 | 182.7999 | 137.8936 | 179.5298 | 121.8268
2012-10-30 17:22 | 26.9183 | 197.0644 | 152.1006 | 193.6643 | 136.0576 | -1.602 | 1.2772 | 32.1267 | 41.0317 | 25.1543 | 2.0026 | 32.0543 | 64.1909 | 16.0066 | 13.8833 | -0.1104 | 105.2226 | 183.1811 | 138.2173 | 179.781 | 122.1743
2012-10-30 17:22 | 26.9235 | 197.3959 | 152.3528 | 194.0664 | 136.1736 | -1.5962 | 1.2771 | 32.1264 | 41.0293 | 25.1553 | 2.0025 | 32.0541 | 64.1898 | 16.0066 | 13.9171 | -0.11 | 105.2191 | 183.4788 | 138.4357 | 180.1493 | 122.2565
2012-10-30 17:22 | 26.9198 | 197.8051 | 152.6887 | 194.6011 | 136.495 | -1.6003 | 1.2771 | 32.1263 | 41.0273 | 25.1564 | 2.0025 | 32.054 | 64.1893 | 16.0067 | 13.8767 | -0.1103 | 105.2166 | 183.9284 | 138.812 | 180.7244 | 122.6183
2012-10-30 17:22 | 26.934 | 198.2087 | 153.0002 | 194.8629 | 136.8419 | -1.6014 | 1.277 | 32.1264 | 41.0265 | 25.157 | 2.0025 | 32.054 | 64.1889 | 16.0067 | 13.8746 | -0.1104 | 105.2154 | 184.3341 | 139.1256 | 180.9883 | 122.9673
2012-10-30 17:22 | 26.9385 | 198.634 | 153.2865 | 195.3115 | 137.0491 | -1.5974 | 1.277 | 32.1265 | 41.0252 | 25.1581 | 2.0025 | 32.0541 | 64.189 | 16.0068 | 13.8995 | -0.1101 | 105.2141 | 184.7345 | 139.387 | 181.412 | 123.1496
2012-10-30 17:22 | 26.9452 | 199.0537 | 153.6369 | 195.7742 | 137.5417 | -1.6009 | 1.2769 | 32.1261 | 41.0227 | 25.1588 | 2.0025 | 32.0538 | 64.1878 | 16.0069 | 13.9062 | -0.1104 | 105.2105 | 185.1475 | 139.7307 | 181.868 | 123.6355
2012-10-30 17:22 | 26.946 | 199.4085 | 154.0085 | 196.1996 | 137.841 | -1.5985 | 1.2769 | 32.1259 | 41.0208 | 25.1598 | 2.0025 | 32.0535 | 64.1863 | 16.007 | 13.8499 | -0.1102 | 105.2072 | 185.5586 | 140.1586 | 182.3497 | 123.9911
2012-10-30 17:22 | 26.9624 | 199.8144 | 154.2644 | 196.5383 | 138.0449 | -1.6014 | 1.2768 | 32.126 | 41.0195 | 25.1607 | 2.0025 | 32.0536 | 64.187 | 16.0068 | 13.9112 | -0.1104 | 105.2065 | 185.9032 | 140.3532 | 182.6271 | 124.1337
2012-10-30 17:22 | 26.9497 | 200.1636 | 154.5029 | 197.0486 | 138.1202 | -1.6026 | 1.2768 | 32.1256 | 41.0166 | 25.1619 | 2.0025 | 32.0533 | 64.1854 | 16.0069 | 13.8618 | -0.1105 | 105.202 | 186.3018 | 140.6411 | 183.1868 | 124.2584
2012-10-30 17:22 | 26.9647 | 200.4933 | 154.7768 | 197.1894 | 138.5755 | -1.6009 | 1.2767 | 32.1256 | 41.0157 | 25.1624 | 2.0025 | 32.0532 | 64.1852 | 16.007 | 13.8972 | -0.1104 | 105.2009 | 186.5961 | 140.8796 | 183.2922 | 124.6783
2012-10-30 17:22 | 26.9639 | 200.8112 | 155.0581 | 197.6723 | 138.7852 | -1.5968 | 1.2767 | 32.1256 | 41.0139 | 25.1635 | 2.0025 | 32.0532 | 64.1848 | 16.007 | 13.8923 | -0.1101 | 105.1987 | 186.9189 | 141.1658 | 183.78 | 124.8929
2012-10-30 17:22 | 26.9639 | 201.1653 | 155.348 | 198.0931 | 139.1337 | -1.6003 | 1.2766 | 32.1253 | 41.0118 | 25.1643 | 2.0024 | 32.0528 | 64.1834 | 16.007 | 13.9249 | -0.1103 | 105.1952 | 187.2404 | 141.4231 | 184.1682 | 125.2088
2012-10-30 17:22 | 26.9729 | 201.5875 | 155.6586 | 198.3577 | 139.4568 | -1.6003 | 1.2766 | 32.1253 | 41.0103 | 25.1652 | 2.0024 | 32.0528 | 64.1835 | 16.0069 | 13.9013 | -0.1103 | 105.1938 | 187.6862 | 141.7573 | 184.4564 | 125.5555
2012-10-30 17:22 | 26.9819 | 201.9152 | 155.9769 | 198.7482 | 139.5347 | -1.6037 | 1.2765 | 32.1251 | 41.0083 | 25.1662 | 2.0024 | 32.0528 | 64.183 | 16.007 | 13.9307 | -0.1106 | 105.1913 | 187.9845 | 142.0462 | 184.8175 | 125.604
2012-10-30 17:22 | 26.9954 | 202.2475 | 156.2147 | 199.1983 | 139.8202 | -1.6009 | 1.2765 | 32.1248 | 41.0067 | 25.1666 | 2.0024 | 32.0524 | 64.1806 | 16.0072 | 13.9279 | -0.1104 | 105.1874 | 188.3196 | 142.2868 | 185.2704 | 125.8923
2012-10-30 17:22 | 27.0051 | 202.568 | 156.4681 | 199.3614 | 140.1211 | -1.6003 | 1.2764 | 32.1243 | 41.0039 | 25.1676 | 2.0024 | 32.0518 | 64.179 | 16.0071 | 13.905 | -0.1103 | 105.1829 | 188.663 | 142.5631 | 185.4564 | 126.2161
2012-10-30 17:22 | 27.0006 | 202.9256 | 156.7654 | 199.8683 | 140.4002 | -1.6032 | 1.2763 | 32.1238 | 41.0008 | 25.1688 | 2.0023 | 32.0514 | 64.1776 | 16.0071 | 13.9127 | -0.1105 | 105.1784 | 189.0129 | 142.8527 | 185.9556 | 126.4875
2012-10-30 17:22 | 27.0074 | 203.2669 | 157.0972 | 200.119 | 140.7788 | -1.5991 | 1.2763 | 32.1237 | 40.9998 | 25.1693 | 2.0023 | 32.0513 | 64.1765 | 16.0072 | 13.8991 | -0.1102 | 105.1762 | 189.3678 | 143.1981 | 186.2199 | 126.8797
2012-10-30 17:22 | 27.0096 | 203.6234 | 157.3625 | 200.4241 | 140.919 | -1.6026 | 1.2763 | 32.1238 | 40.9991 | 25.1698 | 2.0023 | 32.0514 | 64.177 | 16.0072 | 13.9298 | -0.1105 | 105.1761 | 189.6936 | 143.4327 | 186.4943 | 126.9892
2012-10-30 17:22 | 27.0208 | 203.8943 | 157.6176 | 200.9468 | 141.1219 | -1.5991 | 1.2762 | 32.1237 | 40.9972 | 25.1708 | 2.0023 | 32.0512 | 64.176 | 16.0072 | 13.9451 | -0.1102 | 105.1732 | 189.9492 | 143.6725 | 187.0017 | 127.1768
2012-10-30 17:23 | 27.0201 | 204.2715 | 157.8762 | 201.0563 | 141.4778 | -1.6043 | 1.2762 | 32.1234 | 40.9953 | 25.1715 | 2.0023 | 32.0509 | 64.1747 | 16.0073 | 13.9526 | -0.1106 | 105.17 | 190.3189 | 143.9236 | 187.1037 | 127.5252
2012-10-30 17:23 | 27.0291 | 204.5442 | 158.1407 | 201.4863 | 141.5673 | -1.5974 | 1.2761 | 32.1233 | 40.9935 | 25.1724 | 2.0023 | 32.0509 | 64.1745 | 16.0073 | 13.8882 | -0.1101 | 105.168 | 190.656 | 144.2525 | 187.5981 | 127.6791
2012-10-30 17:23 | 27.0343 | 204.8801 | 158.3899 | 201.8771 | 141.8294 | -1.6026 | 1.2761 | 32.1234 | 40.9925 | 25.1732 | 2.0023 | 32.0508 | 64.1738 | 16.0074 | 13.9586 | -0.1105 | 105.1663 | 190.9215 | 144.4313 | 187.9185 | 127.8708
2012-10-30 17:23 | 27.0328 | 205.162 | 158.6285 | 202.0133 | 142.0696 | -1.5991 | 1.2761 | 32.1233 | 40.9914 | 25.1737 | 2.0022 | 32.0507 | 64.1734 | 16.0074 | 13.9286 | -0.1102 | 105.1648 | 191.2334 | 144.6999 | 188.0847 | 128.141
2012-10-30 17:23 | 27.044 | 205.496 | 158.8798 | 202.5086 | 142.1997 | -1.5991 | 1.276 | 32.1229 | 40.9897 | 25.1741 | 2.0022 | 32.0504 | 64.1722 | 16.0074 | 13.9642 | -0.1102 | 105.1618 | 191.5318 | 144.9156 | 188.5444 | 128.2355
2012-10-30 17:23 | 27.05 | 205.8642 | 159.1763 | 202.9238 | 142.6354 | -1.5974 | 1.276 | 32.1229 | 40.9875 | 25.1754 | 2.0022 | 32.0504 | 64.1718 | 16.0075 | 13.8847 | -0.1101 | 105.1594 | 191.9795 | 145.2916 | 189.0391 | 128.7507
2012-10-30 17:23 | 27.056 | 206.2087 | 159.469 | 203.1148 | 142.8985 | -1.602 | 1.2759 | 32.1227 | 40.9862 | 25.176 | 2.0022 | 32.0502 | 64.1711 | 16.0074 | 13.9355 | -0.1104 | 105.1573 | 192.2732 | 145.5335 | 189.1793 | 128.963
2012-10-30 17:23 | 27.0613 | 206.529 | 159.7107 | 203.5255 | 143.042 | -1.602 | 1.2759 | 32.1224 | 40.9843 | 25.1768 | 2.0022 | 32.05 | 64.1702 | 16.0075 | 13.896 | -0.1104 | 105.1544 | 192.633 | 145.8147 | 189.6295 | 129.146
2012-10-30 17:23 | 27.068 | 206.8349 | 159.9809 | 203.8864 | 143.2982 | -1.5997 | 1.2758 | 32.1222 | 40.9824 | 25.1776 | 2.0022 | 32.0497 | 64.169 | 16.0075 | 13.9149 | -0.1103 | 105.1514 | 192.92 | 146.066 | 189.9715 | 129.3833
2012-10-30 17:23 | 27.0747 | 207.2038 | 160.2512 | 204.1057 | 143.5813 | -1.6026 | 1.2758 | 32.1222 | 40.9807 | 25.1785 | 2.0022 | 32.0497 | 64.1695 | 16.0073 | 13.9542 | -0.1105 | 105.1503 | 193.2496 | 146.297 | 190.1515 | 129.6271
2012-10-30 17:23 | 27.077 | 207.5211 | 160.5055 | 204.5488 | 143.791 | -1.6014 | 1.2757 | 32.122 | 40.9792 | 25.1793 | 2.0022 | 32.0495 | 64.1682 | 16.0075 | 13.9361 | -0.1104 | 105.1474 | 193.585 | 146.5694 | 190.6127 | 129.8549
2012-10-30 17:23 | 27.0867 | 207.8182 | 160.7594 | 205.0151 | 144.0195 | -1.5991 | 1.2757 | 32.1219 | 40.9773 | 25.1802 | 2.0021 | 32.0493 | 64.1673 | 16.0075 | 13.9418 | -0.1102 | 105.1446 | 193.8764 | 146.8176 | 191.0733 | 130.0777
2012-10-30 17:23 | 27.0912 | 208.1456 | 161.0392 | 205.1338 | 144.3889 | -1.5997 | 1.2756 | 32.1219 | 40.9762 | 25.1808 | 2.0021 | 32.0494 | 64.1673 | 16.0076 | 13.9218 | -0.1103 | 105.1435 | 194.2238 | 147.1174 | 191.212 | 130.4671
2012-10-30 17:23 | 27.0987 | 208.5074 | 161.3104 | 205.4798 | 144.6192 | -1.6026 | 1.2756 | 32.1217 | 40.9744 | 25.1817 | 2.0021 | 32.0492 | 64.1665 | 16.0076 | 13.9741 | -0.1105 | 105.1409 | 194.5333 | 147.3363 | 191.5057 | 130.6451
2012-10-30 17:23 | 27.1002 | 208.7491 | 161.5524 | 205.9577 | 144.9204 | -1.5991 | 1.2755 | 32.1214 | 40.9724 | 25.1824 | 2.0021 | 32.0489 | 64.1648 | 16.0077 | 13.9553 | -0.1102 | 105.1372 | 194.7938 | 147.5971 | 192.0024 | 130.9651
2012-10-30 17:23 | 27.1099 | 209.0974 | 161.8569 | 206.18 | 145.2413 | -1.5974 | 1.2755 | 32.1213 | 40.971 | 25.1832 | 2.0021 | 32.0488 | 64.1642 | 16.0078 | 13.9446 | -0.1101 | 105.1352 | 195.1528 | 147.9123 | 192.2354 | 131.2967
2012-10-30 17:23 | 27.1181 | 209.4049 | 162.1212 | 206.4716 | 145.4148 | -1.6032 | 1.2755 | 32.1212 | 40.9693 | 25.184 | 2.0021 | 32.0486 | 64.1635 | 16.0077 | 13.9121 | -0.1105 | 105.1328 | 195.4928 | 148.2091 | 192.5595 | 131.5027
2012-10-30 17:23 | 27.1137 | 209.7706 | 162.4389 | 206.853 | 145.7678 | -1.5962 | 1.2754 | 32.121 | 40.9678 | 25.1846 | 2.002 | 32.0483 | 64.1623 | 16.0078 | 13.9443 | -0.11 | 105.1301 | 195.8263 | 148.4946 | 192.9087 | 131.8235
2012-10-30 17:23 | 27.1196 | 210.0719 | 162.7553 | 207.2605 | 146.0505 | -1.5985 | 1.2753 | 32.12 | 40.9639 | 25.1855 | 2.002 | 32.0475 | 64.1588 | 16.0078 | 13.9502 | -0.1102 | 105.1227 | 196.1217 | 148.8051 | 193.3103 | 132.1003
2012-10-30 17:23 | 27.1249 | 210.3883 | 162.9779 | 207.4391 | 146.2211 | -1.6043 | 1.2753 | 32.1199 | 40.9632 | 25.1858 | 2.002 | 32.0474 | 64.1579 | 16.008 | 13.9759 | -0.1106 | 105.1211 | 196.4124 | 149.002 | 193.4632 | 132.2452
2012-10-30 17:23 | 27.1331 | 210.6642 | 163.2113 | 207.8489 | 146.2678 | -1.6049 | 1.2753 | 32.12 | 40.9621 | 25.1866 | 2.002 | 32.0474 | 64.1586 | 16.0078 | 13.9719 | -0.1106 | 105.1207 | 196.6923 | 149.2394 | 193.877 | 132.2959
2012-10-30 17:23 | 27.1324 | 210.9392 | 163.4202 | 208.1277 | 146.5017 | -1.5928 | 1.2752 | 32.1195 | 40.9594 | 25.1874 | 2.0019 | 32.0468 | 64.1557 | 16.0079 | 13.9915 | -0.1098 | 105.1151 | 196.9477 | 149.4287 | 194.1362 | 132.5102
2012-10-30 17:23 | 27.1429 | 211.1741 | 163.6598 | 208.2839 | 146.6083 | -1.6061 | 1.2752 | 32.1194 | 40.9577 | 25.1883 | 2.0019 | 32.0467 | 64.1551 | 16.008 | 13.9776 | -0.1107 | 105.1129 | 197.1965 | 149.6822 | 194.3063 | 132.6307
2012-10-30 17:23 | 27.1361 | 211.4666 | 163.7928 | 208.7299 | 146.7592 | -1.6049 | 1.2751 | 32.119 | 40.956 | 25.1887 | 2.0019 | 32.0464 | 64.1539 | 16.008 | 13.9871 | -0.1106 | 105.1099 | 197.4795 | 149.8057 | 194.7428 | 132.7721
2012-10-30 17:23 | 27.1488 | 211.8101 | 164.1007 | 208.9277 | 147.1412 | -1.6032 | 1.2751 | 32.1188 | 40.9544 | 25.1894 | 2.0019 | 32.0462 | 64.1532 | 16.0079 | 13.955 | -0.1105 | 105.1076 | 197.8551 | 150.1457 | 194.9727 | 133.1862
2012-10-30 17:23 | 27.1563 | 212.0972 | 164.3412 | 209.2226 | 147.3603 | -1.598 | 1.2751 | 32.1189 | 40.9535 | 25.1902 | 2.0019 | 32.0462 | 64.153 | 16.0079 | 13.9625 | -0.1102 | 105.1065 | 198.1347 | 150.3787 | 195.2601 | 133.3978
2012-10-30 17:24 | 27.1616 | 212.3821 | 164.5563 | 209.5705 | 147.2886 | -1.5991 | 1.275 | 32.119 | 40.9528 | 25.1907 | 2.0019 | 32.0463 | 64.1527 | 16.0082 | 13.9963 | -0.1102 | 105.1054 | 198.3858 | 150.56 | 195.5742 | 133.2923
2012-10-30 17:24 | 27.1601 | 212.6326 | 164.7413 | 209.8525 | 147.5295 | -1.6066 | 1.275 | 32.1189 | 40.9512 | 25.1915 | 2.0019 | 32.0462 | 64.1529 | 16.008 | 13.9744 | -0.1108 | 105.1042 | 198.6582 | 150.7669 | 195.8781 | 133.5551
2012-10-30 17:24 | 27.1706 | 212.9542 | 164.9772 | 210.0993 | 147.8979 | -1.5962 | 1.2749 | 32.1186 | 40.9492 | 25.1923 | 2.0019 | 32.0459 | 64.1515 | 16.0081 | 13.9971 | -0.11 | 105.1007 | 198.9571 | 150.9801 | 196.1022 | 133.9008
2012-10-30 17:24 | 27.1795 | 213.2112 | 165.2115 | 210.3877 | 148.1301 | -1.5933 | 1.2749 | 32.1183 | 40.9477 | 25.1928 | 2.0018 | 32.0457 | 64.15 | 16.0082 | 13.9898 | -0.1098 | 105.0977 | 199.2214 | 151.2217 | 196.3979 | 134.1403
2012-10-30 17:24 | 27.1788 | 213.4192 | 165.4284 | 210.7848 | 148.241 | -1.6032 | 1.2749 | 32.1184 | 40.9469 | 25.1935 | 2.0018 | 32.0458 | 64.1507 | 16.0082 | 13.9809 | -0.1105 | 105.0976 | 199.4383 | 151.4475 | 196.8039 | 134.2601
2012-10-30 17:24 | 27.1795 | 213.8019 | 165.6234 | 210.8563 | 148.5382 | -1.6037 | 1.2748 | 32.1181 | 40.9454 | 25.1939 | 2.0018 | 32.0455 | 64.1498 | 16.008 | 13.9816 | -0.1106 | 105.0952 | 199.8203 | 151.6418 | 196.8747 | 134.5566
2012-10-30 17:24 | 27.1825 | 213.9782 | 165.8441 | 211.2098 | 148.7106 | -1.598 | 1.2748 | 32.1178 | 40.9433 | 25.1947 | 2.0018 | 32.0452 | 64.1482 | 16.0082 | 14.0172 | -0.1102 | 105.0914 | 199.961 | 151.8269 | 197.1926 | 134.6934
2012-10-30 17:24 | 27.1968 | 214.2603 | 166.1228 | 211.5864 | 148.8789 | -1.602 | 1.2747 | 32.1176 | 40.9414 | 25.1954 | 2.0018 | 32.0449 | 64.1465 | 16.0083 | 13.9867 | -0.1104 | 105.0879 | 200.2736 | 152.1361 | 197.5997 | 134.8922
2012-10-30 17:24 | 27.1953 | 214.4911 | 166.3002 | 211.6952 | 149.1316 | -1.6014 | 1.2747 | 32.1174 | 40.9403 | 25.1959 | 2.0018 | 32.0448 | 64.146 | 16.0083 | 13.9567 | -0.1104 | 105.0863 | 200.5344 | 152.3435 | 197.7385 | 135.1749
2012-10-30 17:24 | 27.1953 | 214.7195 | 166.4829 | 212.0732 | 149.1354 | -1.6014 | 1.2747 | 32.1173 | 40.9391 | 25.1965 | 2.0018 | 32.0447 | 64.1462 | 16.0082 | 13.9893 | -0.1104 | 105.0853 | 200.7302 | 152.4936 | 198.0839 | 135.1461
2012-10-30 17:24 | 27.2102 | 214.9904 | 166.6884 | 212.3598 | 149.4084 | -1.6009 | 1.2746 | 32.1173 | 40.9376 | 25.1973 | 2.0018 | 32.0445 | 64.1453 | 16.0082 | 13.9675 | -0.1104 | 105.083 | 201.0229 | 152.7209 | 198.3923 | 135.4409
2012-10-30 17:24 | 27.2215 | 215.3049 | 166.8864 | 212.4893 | 149.6508 | -1.598 | 1.2746 | 32.1172 | 40.9368 | 25.1977 | 2.0017 | 32.0444 | 64.1448 | 16.0083 | 14.0236 | -0.1102 | 105.0816 | 201.2813 | 152.8628 | 198.4657 | 135.6272
2012-10-30 17:24 | 27.2252 | 215.5055 | 167.0768 | 212.7568 | 149.6545 | -1.6032 | 1.2746 | 32.1172 | 40.9355 | 25.1985 | 2.0017 | 32.0444 | 64.1442 | 16.0084 | 14.0721 | -0.1105 | 105.0797 | 201.4334 | 153.0047 | 198.6847 | 135.5824
2012-10-30 17:24 | 27.2237 | 215.7285 | 167.2853 | 213.1137 | 149.8881 | -1.5962 | 1.2745 | 32.1169 | 40.9344 | 25.1987 | 2.0017 | 32.0443 | 64.1442 | 16.0083 | 14.0217 | -0.11 | 105.0786 | 201.7068 | 153.2636 | 199.092 | 135.8664
2012-10-30 17:24 | 27.223 | 215.968 | 167.4286 | 213.2626 | 150.1532 | -1.5985 | 1.2745 | 32.1168 | 40.9329 | 25.1995 | 2.0017 | 32.0442 | 64.1435 | 16.0083 | 13.9803 | -0.1102 | 105.0764 | 201.9877 | 153.4483 | 199.2823 | 136.1729
2012-10-30 17:24 | 27.2402 | 216.2451 | 167.6131 | 213.6067 | 150.3246 | -1.5962 | 1.2744 | 32.1157 | 40.9298 | 25.1997 | 2.0017 | 32.043 | 64.139 | 16.0083 | 13.9975 | -0.11 | 105.0688 | 202.2476 | 153.6156 | 199.6092 | 136.3271
2012-10-30 17:24 | 27.2409 | 216.4742 | 167.89 | 213.824 | 150.4756 | -1.5968 | 1.2744 | 32.1157 | 40.9282 | 25.2006 | 2.0016 | 32.0429 | 64.1379 | 16.0084 | 13.99 | -0.1101 | 105.0661 | 202.4842 | 153.9 | 199.834 | 136.4856
2012-10-30 17:24 | 27.2454 | 216.7071 | 168.0813 | 214.049 | 150.8656 | -1.5951 | 1.2743 | 32.1152 | 40.9259 | 25.2013 | 2.0016 | 32.0425 | 64.1364 | 16.0084 | 13.9456 | -0.11 | 105.0623 | 202.7615 | 154.1357 | 200.1034 | 136.92
2012-10-30 17:24 | 27.2492 | 216.9314 | 168.2758 | 214.3481 | 150.9581 | -1.5991 | 1.2743 | 32.1152 | 40.9249 | 25.2019 | 2.0016 | 32.0425 | 64.1362 | 16.0084 | 13.9943 | -0.1102 | 105.0611 | 202.9371 | 154.2815 | 200.3538 | 136.9638
2012-10-30 17:24 | 27.2551 | 217.1814 | 168.4219 | 214.6494 | 151.0989 | -1.5997 | 1.2743 | 32.1152 | 40.9244 | 25.2022 | 2.0016 | 32.0425 | 64.1363 | 16.0084 | 14.0409 | -0.1103 | 105.0607 | 203.1405 | 154.381 | 200.6085 | 137.058
2012-10-30 17:24 | 27.2686 | 217.4391 | 168.7001 | 214.7219 | 151.3438 | -1.5968 | 1.2743 | 32.1153 | 40.9239 | 25.2027 | 2.0016 | 32.0426 | 64.1371 | 16.0083 | 14.0259 | -0.1101 | 105.061 | 203.4132 | 154.6742 | 200.696 | 137.3179
2012-10-30 17:24 | 27.2619 | 217.6804 | 168.8491 | 215.0382 | 151.5338 | -1.6061 | 1.2743 | 32.1154 | 40.9233 | 25.2032 | 2.0016 | 32.0427 | 64.1374 | 16.0084 | 13.954 | -0.1107 | 105.0608 | 203.7264 | 154.8951 | 201.0842 | 137.5798
2012-10-30 17:24 | 27.2776 | 217.9285 | 169.0906 | 215.3925 | 151.9467 | -1.6026 | 1.2742 | 32.1155 | 40.9222 | 25.204 | 2.0016 | 32.0427 | 64.1375 | 16.0083 | 13.9982 | -0.1105 | 105.0597 | 203.9303 | 155.0924 | 201.3943 | 137.9485
2012-10-30 17:24 | 27.2821 | 218.1298 | 169.2586 | 215.4403 | 152.3718 | -1.5974 | 1.2742 | 32.1155 | 40.9218 | 25.2043 | 2.0016 | 32.0428 | 64.1374 | 16.0085 | 14.0475 | -0.1101 | 105.0593 | 204.0823 | 155.2111 | 201.3928 | 138.3243
2012-10-30 17:24 | 27.2851 | 218.3376 | 169.5226 | 215.7978 | 152.3439 | -1.6009 | 1.2742 | 32.1155 | 40.9206 | 25.205 | 2.0016 | 32.0427 | 64.1374 | 16.0084 | 14.0139 | -0.1104 | 105.058 | 204.3237 | 155.5087 | 201.7839 | 138.33
2012-10-30 17:24 | 27.2746 | 218.5948 | 169.7496 | 215.9448 | 152.5533 | -1.598 | 1.2741 | 32.1153 | 40.9193 | 25.2055 | 2.0016 | 32.0425 | 64.1366 | 16.0084 | 14.0115 | -0.1102 | 105.0559 | 204.5833 | 155.7381 | 201.9333 | 138.5418
2012-10-30 17:24 | 27.2926 | 218.8373 | 169.9079 | 216.1833 | 152.6099 | -1.598 | 1.2741 | 32.115 | 40.9181 | 25.2058 | 2.0016 | 32.0423 | 64.1352 | 16.0085 | 14.0336 | -0.1102 | 105.0533 | 204.8037 | 155.8743 | 202.1497 | 138.5763
2012-10-30 17:24 | 27.2948 | 219.0718 | 170.1243 | 216.5675 | 152.5813 | -1.5939 | 1.2741 | 32.1149 | 40.9168 | 25.2064 | 2.0016 | 32.0421 | 64.1343 | 16.0085 | 14.0236 | -0.1099 | 105.0511 | 205.0482 | 156.1007 | 202.5439 | 138.5577
2012-10-30 17:25 | 27.3031 | 219.2691 | 170.3352 | 216.6703 | 152.8559 | -1.6043 | 1.274 | 32.1144 | 40.9146 | 25.2071 | 2.0015 | 32.0416 | 64.1326 | 16.0085 | 14.0196 | -0.1106 | 105.0473 | 205.2495 | 156.3156 | 202.6507 | 138.8363
2012-10-30 17:25 | 27.3053 | 219.4721 | 170.4737 | 216.8182 | 152.9121 | -1.598 | 1.274 | 32.1143 | 40.9135 | 25.2076 | 2.0015 | 32.0415 | 64.1319 | 16.0086 | 14.03 | -0.1102 | 105.0453 | 205.4421 | 156.4437 | 202.7882 | 138.8821
2012-10-30 17:25 | 27.3158 | 219.6952 | 170.6245 | 217.1988 | 153.2275 | -1.5997 | 1.274 | 32.114 | 40.9126 | 25.2077 | 2.0015 | 32.0412 | 64.1308 | 16.0085 | 13.9672 | -0.1103 | 105.0434 | 205.728 | 156.6573 | 203.2316 | 139.2603
2012-10-30 17:25 | 27.318 | 219.91 | 170.8293 | 217.268 | 153.4073 | -1.6014 | 1.2739 | 32.1142 | 40.9118 | 25.2083 | 2.0015 | 32.0414 | 64.1315 | 16.0085 | 14.0223 | -0.1104 | 105.0433 | 205.8877 | 156.807 | 203.2457 | 139.385
2012-10-30 17:25 | 27.3173 | 220.1692 | 170.9845 | 217.5626 | 153.4645 | -1.5985 | 1.2739 | 32.1141 | 40.9106 | 25.209 | 2.0015 | 32.0413 | 64.131 | 16.0085 | 13.9972 | -0.1102 | 105.0416 | 206.172 | 156.9873 | 203.5654 | 139.4673
2012-10-30 17:25 | 27.3203 | 220.3336 | 171.159 | 217.8727 | 153.6876 | -1.6026 | 1.2739 | 32.1139 | 40.9093 | 25.2094 | 2.0015 | 32.041 | 64.1298 | 16.0085 | 14.0368 | -0.1105 | 105.0391 | 206.2968 | 157.1222 | 203.8359 | 139.6508
2012-10-30 17:25 | 27.3248 | 220.5782 | 171.3662 | 217.9245 | 153.8466 | -1.6055 | 1.2738 | 32.1137 | 40.9079 | 25.21 | 2.0015 | 32.0408 | 64.1285 | 16.0087 | 14.025 | -0.1107 | 105.0363 | 206.5532 | 157.3412 | 203.8995 | 139.8216
2012-10-30 17:25 | 27.3338 | 220.7605 | 171.5194 | 218.2327 | 153.7822 | -1.6014 | 1.2738 | 32.1137 | 40.9074 | 25.2104 | 2.0015 | 32.0409 | 64.1295 | 16.0086 | 14.03 | -0.1104 | 105.0369 | 206.7305 | 157.4894 | 204.2027 | 139.7522
2012-10-30 17:25 | 27.33 | 220.9614 | 171.6754 | 218.4927 | 154.0642 | -1.5957 | 1.2738 | 32.1136 | 40.9063 | 25.2109 | 2.0015 | 32.0408 | 64.1284 | 16.0088 | 13.9977 | -0.11 | 105.0347 | 206.9637 | 157.6777 | 204.495 | 140.0665
2012-10-30 17:25 | 27.348 | 221.1683 | 171.8961 | 218.5895 | 154.1517 | -1.6026 | 1.2738 | 32.1135 | 40.9046 | 25.2117 | 2.0014 | 32.0406 | 64.1276 | 16.0087 | 14.0075 | -0.1105 | 105.0322 | 207.1608 | 157.8886 | 204.582 | 140.1442
2012-10-30 17:25 | 27.342 | 221.3474 | 172.0188 | 218.93 | 154.2616 | -1.6003 | 1.2737 | 32.1132 | 40.9033 | 25.212 | 2.0014 | 32.0404 | 64.1266 | 16.0087 | 14.0422 | -0.1103 | 105.0299 | 207.3052 | 157.9766 | 204.8878 | 140.2194
2012-10-30 17:25 | 27.345 | 221.6141 | 172.2128 | 219.0354 | 154.5273 | -1.5991 | 1.2737 | 32.1131 | 40.9018 | 25.2128 | 2.0014 | 32.0403 | 64.1261 | 16.0088 | 14.0412 | -0.1102 | 105.0279 | 207.5729 | 158.1716 | 204.9942 | 140.4861
2012-10-30 17:25 | 27.3532 | 221.8113 | 172.3731 | 219.2641 | 154.6436 | -1.6003 | 1.2737 | 32.1128 | 40.9005 | 25.2132 | 2.0014 | 32.04 | 64.1255 | 16.0087 | 14.029 | -0.1103 | 105.026 | 207.7823 | 158.3441 | 205.2351 | 140.6146
2012-10-30 17:25 | 27.3577 | 222.0559 | 172.5765 | 219.556 | 154.88 | -1.6014 | 1.2736 | 32.1126 | 40.8992 | 25.2136 | 2.0014 | 32.0397 | 64.1234 | 16.0089 | 14.0091 | -0.1104 | 105.0226 | 208.0468 | 158.5674 | 205.5469 | 140.8709
2012-10-30 17:25 | 27.3577 | 222.3118 | 172.7286 | 219.8552 | 155.0229 | -1.6003 | 1.2736 | 32.1123 | 40.8978 | 25.214 | 2.0014 | 32.0394 | 64.1223 | 16.0088 | 14.0294 | -0.1103 | 105.0202 | 208.2824 | 158.6992 | 205.8258 | 140.9935
2012-10-30 17:25 | 27.3667 | 222.4979 | 172.9482 | 219.9666 | 155.2561 | -1.5962 | 1.2736 | 32.1123 | 40.897 | 25.2146 | 2.0014 | 32.0394 | 64.1222 | 16.0088 | 14.0303 | -0.11 | 105.0192 | 208.4676 | 158.9179 | 205.9363 | 141.2258
2012-10-30 17:25 | 27.3682 | 222.7159 | 173.0238 | 220.2082 | 155.3233 | -1.5991 | 1.2735 | 32.1122 | 40.8961 | 25.2149 | 2.0014 | 32.0393 | 64.122 | 16.0089 | 14.0196 | -0.1102 | 105.018 | 208.6963 | 159.0042 | 206.1886 | 141.3037
2012-10-30 17:25 | 27.3764 | 222.8894 | 173.2427 | 220.5038 | 155.6214 | -1.6037 | 1.2735 | 32.1122 | 40.8947 | 25.2158 | 2.0014 | 32.0393 | 64.1224 | 16.0087 | 14.0074 | -0.1106 | 105.0171 | 208.882 | 159.2353 | 206.4964 | 141.614
2012-10-30 17:25 | 27.3809 | 223.0986 | 173.4227 | 220.5202 | 155.8385 | -1.6003 | 1.2735 | 32.1121 | 40.8944 | 25.2159 | 2.0014 | 32.0393 | 64.122 | 16.0088 | 14.0486 | -0.1103 | 105.0164 | 209.05 | 159.3741 | 206.4716 | 141.7899
2012-10-30 17:25 | 27.3839 | 223.2945 | 173.5428 | 220.8184 | 155.8492 | -1.6037 | 1.2735 | 32.112 | 40.8936 | 25.2163 | 2.0014 | 32.0392 | 64.1217 | 16.0088 | 14.023 | -0.1106 | 105.0153 | 209.2715 | 159.5198 | 206.7954 | 141.8262
2012-10-30 17:25 | 27.3914 | 223.4673 | 173.7454 | 221.0621 | 155.9843 | -1.5974 | 1.2734 | 32.1119 | 40.8921 | 25.217 | 2.0013 | 32.039 | 64.1205 | 16.0089 | 14.0509 | -0.1101 | 105.0126 | 209.4164 | 159.6945 | 207.0112 | 141.9334
2012-10-30 17:25 | 27.3967 | 223.7088 | 173.9185 | 221.0714 | 156.2216 | -1.598 | 1.2734 | 32.1119 | 40.8917 | 25.2172 | 2.0013 | 32.0391 | 64.1213 | 16.0088 | 14.0481 | -0.1102 | 105.013 | 209.6607 | 159.8704 | 207.0233 | 142.1735
2012-10-30 17:25 | 27.3989 | 223.8487 | 174.0221 | 221.42 | 156.2973 | -1.5991 | 1.2734 | 32.1121 | 40.8912 | 25.2178 | 2.0013 | 32.0393 | 64.1213 | 16.0089 | 14.034 | -0.1102 | 105.0125 | 209.8147 | 159.9881 | 207.386 | 142.2633
2012-10-30 17:25 | 27.4071 | 224.0183 | 174.202 | 221.6408 | 156.4988 | -1.5997 | 1.2734 | 32.1119 | 40.8904 | 25.218 | 2.0014 | 32.0391 | 64.1216 | 16.0087 | 14.0218 | -0.1103 | 105.012 | 209.9965 | 160.1802 | 207.619 | 142.477
2012-10-30 17:25 | 27.4079 | 224.2316 | 174.3511 | 221.6731 | 156.6234 | -1.6049 | 1.2733 | 32.1117 | 40.8892 | 25.2184 | 2.0013 | 32.0388 | 64.1201 | 16.0088 | 14.043 | -0.1106 | 105.0093 | 210.1886 | 160.3081 | 207.6301 | 142.5804
2012-10-30 17:25 | 27.4161 | 224.4445 | 174.5389 | 222.008 | 156.7786 | -1.5974 | 1.2733 | 32.1115 | 40.8874 | 25.2192 | 2.0013 | 32.0386 | 64.1192 | 16.0088 | 14.0267 | -0.1101 | 105.0066 | 210.4178 | 160.5122 | 207.9813 | 142.7519
2012-10-30 17:25 | 27.4146 | 224.6122 | 174.6779 | 222.1286 | 157.16 | -1.6072 | 1.2733 | 32.1113 | 40.8867 | 25.2194 | 2.0013 | 32.0384 | 64.1179 | 16.0089 | 14.0497 | -0.1108 | 105.0047 | 210.5625 | 160.6282 | 208.0789 | 143.1103
2012-10-30 17:26 | 27.4191 | 224.8056 | 174.8581 | 222.3299 | 157.122 | -1.6003 | 1.2732 | 32.111 | 40.8852 | 25.2199 | 2.0013 | 32.0381 | 64.1171 | 16.0088 | 14.0338 | -0.1103 | 105.0023 | 210.7718 | 160.8243 | 208.2961 | 143.0882
2012-10-30 17:26 | 27.4273 | 224.9555 | 175.0107 | 222.6176 | 157.2501 | -1.6037 | 1.2732 | 32.111 | 40.8847 | 25.2202 | 2.0013 | 32.0381 | 64.1172 | 16.0088 | 14.0216 | -0.1106 | 105.002 | 210.9339 | 160.9891 | 208.596 | 143.2285
2012-10-30 17:26 | 27.4296 | 225.1467 | 175.2082 | 222.7537 | 157.442 | -1.6043 | 1.2732 | 32.1117 | 40.8853 | 25.2208 | 2.0013 | 32.0387 | 64.119 | 16.009 | 14.0402 | -0.1106 | 105.0043 | 211.1065 | 161.168 | 208.7135 | 143.4018
2012-10-30 17:26 | 27.4341 | 225.344 | 175.322 | 222.8448 | 157.2492 | -1.6043 | 1.2732 | 32.1114 | 40.8839 | 25.2213 | 2.0013 | 32.0386 | 64.1186 | 16.0089 | 14.0406 | -0.1106 | 105.0025 | 211.3034 | 161.2814 | 208.8042 | 143.2086
2012-10-30 17:26 | 27.4416 | 225.5522 | 175.4779 | 223.179 | 157.423 | -1.6089 | 1.2732 | 32.1113 | 40.8835 | 25.2214 | 2.0013 | 32.0384 | 64.1185 | 16.0088 | 14.0563 | -0.1109 | 105.0019 | 211.4959 | 161.4216 | 209.1227 | 143.3667
2012-10-30 17:26 | 27.4423 | 225.7103 | 175.5958 | 223.2191 | 157.6714 | -1.5985 | 1.2732 | 32.1112 | 40.8825 | 25.2218 | 2.0013 | 32.0382 | 64.1176 | 16.0089 | 14.0244 | -0.1102 | 105.0001 | 211.6859 | 161.5714 | 209.1947 | 143.647
2012-10-30 17:26 | 27.4491 | 225.8902 | 175.7197 | 223.4148 | 157.7827 | -1.6066 | 1.2731 | 32.1108 | 40.8809 | 25.2222 | 2.0013 | 32.038 | 64.1162 | 16.0089 | 14.0679 | -0.1108 | 104.9972 | 211.8223 | 161.6518 | 209.3469 | 143.7148
2012-10-30 17:26 | 27.4528 | 226.0434 | 175.8992 | 223.7255 | 157.9257 | -1.6049 | 1.2731 | 32.1108 | 40.8798 | 25.2228 | 2.0013 | 32.0379 | 64.1159 | 16.0089 | 14.0593 | -0.1106 | 104.9957 | 211.9841 | 161.8399 | 209.6662 | 143.8664
2012-10-30 17:26 | 27.4573 | 226.1935 | 176.0169 | 223.7418 | 158.1276 | -1.6009 | 1.2693 | 32.0154 | 40.6378 | 25.2225 | 1.9952 | 31.9419 | 63.7314 | 16.0091 | 14.0231 | -0.1104 | 104.3692 | 212.1704 | 161.9938 | 209.7187 | 144.1045
2012-10-30 17:26 | 27.4678 | 226.928 | 176.6134 | 224.5946 | 158.715 | -1.5997 | 1.2659 | 31.9293 | 40.4188 | 25.2229 | 1.9899 | 31.8569 | 63.3915 | 16.0094 | 14.0499 | -0.1103 | 103.8103 | 212.8781 | 162.5635 | 210.5447 | 144.6651
2012-10-30 17:26 | 27.4842 | 227.7666 | 177.419 | 225.4846 | 159.5563 | -1.6003 | 1.268 | 31.9853 | 40.5563 | 25.2257 | 1.9934 | 31.9131 | 63.6145 | 16.0096 | 14.0663 | -0.1103 | 104.1708 | 213.7003 | 163.3527 | 211.4183 | 145.49
2012-10-30 17:26 | 27.4955 | 228.3129 | 177.8798 | 225.9524 | 159.9783 | -1.6009 | 1.2723 | 32.0977 | 40.8379 | 25.2281 | 2.0003 | 32.0247 | 64.0602 | 16.0096 | 14.0572 | -0.1104 | 104.8982 | 214.2557 | 163.8226 | 211.8952 | 145.9211
2012-10-30 17:26 | 27.5007 | 228.8649 | 178.3228 | 226.5242 | 160.3245 | -1.602 | 1.2732 | 32.1233 | 40.9004 | 25.2298 | 2.0019 | 32.0502 | 64.1629 | 16.0095 | 14.0542 | -0.1104 | 105.0633 | 214.8107 | 164.2686 | 212.47 | 146.2703
2012-10-30 17:27 | 27.5075 | 229.3672 | 178.7558 | 226.9913 | 160.6569 | -1.6026 | 1.2735 | 32.1306 | 40.9171 | 25.2309 | 2.0024 | 32.0574 | 64.1915 | 16.0096 | 14.1018 | -0.1105 | 105.1086 | 215.2654 | 164.654 | 212.8895 | 146.5551
2012-10-30 17:27 | 27.5194 | 229.682 | 179.0101 | 227.3337 | 160.7618 | -1.5985 | 1.2735 | 32.1325 | 40.9204 | 25.2318 | 2.0025 | 32.0593 | 64.1988 | 16.0096 | 14.0852 | -0.1102 | 105.1192 | 215.5968 | 164.9249 | 213.2485 | 146.6766
2012-10-30 17:27 | 27.5187 | 229.8346 | 179.0876 | 227.3448 | 160.796 | -1.5997 | 1.2735 | 32.1335 | 40.9222 | 25.2323 | 2.0026 | 32.0603 | 64.2026 | 16.0097 | 14.113 | -0.1103 | 105.1248 | 215.7216 | 164.9746 | 213.2318 | 146.683
2012-10-30 17:27 | 27.5262 | 229.9837 | 179.1889 | 227.7063 | 160.8151 | -1.5997 | 1.2735 | 32.1344 | 40.924 | 25.2326 | 2.0026 | 32.0612 | 64.2065 | 16.0096 | 14.092 | -0.1103 | 105.1305 | 215.8917 | 165.0969 | 213.6143 | 146.7231
2012-10-30 17:27 | 27.5157 | 230.1462 | 179.3075 | 227.743 | 161.1303 | -1.6032 | 1.2735 | 32.1346 | 40.9238 | 25.233 | 2.0026 | 32.0614 | 64.2068 | 16.0097 | 14.11 | -0.1105 | 105.1306 | 216.0362 | 165.1975 | 213.633 | 147.0203
2012-10-30 17:27 | 27.5254 | 230.3132 | 179.458 | 227.8668 | 161.2718 | -1.5957 | 1.2735 | 32.1347 | 40.9233 | 25.2335 | 2.0026 | 32.0615 | 64.2073 | 16.0097 | 14.0504 | -0.11 | 105.1306 | 216.2628 | 165.4076 | 213.8164 | 147.2214
2012-10-30 17:27 | 27.5344 | 230.4677 | 179.5882 | 228.041 | 161.4049 | -1.6026 | 1.2735 | 32.1347 | 40.9227 | 25.234 | 2.0026 | 32.0615 | 64.2073 | 16.0097 | 14.1042 | -0.1105 | 105.13 | 216.3635 | 165.484 | 213.9368 | 147.3007
2012-10-30 17:27 | 27.5367 | 230.6273 | 179.6688 | 228.3344 | 161.7368 | -1.5962 | 1.2734 | 32.1347 | 40.9216 | 25.2346 | 2.0026 | 32.0615 | 64.2075 | 16.0097 | 14.0739 | -0.11 | 105.1291 | 216.5534 | 165.5949 | 214.2605 | 147.6629
2012-10-30 17:27 | 27.5314 | 230.7675 | 179.816 | 228.3488 | 161.8246 | -1.6037 | 1.2735 | 32.1352 | 40.9226 | 25.2347 | 2.0027 | 32.0619 | 64.2091 | 16.0097 | 14.0931 | -0.1106 | 105.1317 | 216.6744 | 165.7229 | 214.2557 | 147.7315
2012-10-30 17:27 | 27.5419 | 230.9117 | 180.0573 | 228.4891 | 161.8739 | -1.6014 | 1.2735 | 32.1356 | 40.9237 | 25.2347 | 2.0027 | 32.0624 | 64.211 | 16.0097 | 14.0832 | -0.1104 | 105.1347 | 216.8285 | 165.9741 | 214.4059 | 147.7907
2012-10-30 17:27 | 27.5329 | 231.0639 | 180.0796 | 228.6807 | 161.8843 | -1.6032 | 1.2734 | 32.1352 | 40.9218 | 25.2353 | 2.0027 | 32.062 | 64.2098 | 16.0096 | 14.0661 | -0.1105 | 105.1316 | 216.9978 | 166.0135 | 214.6146 | 147.8182
2012-10-30 17:27 | 27.5427 | 231.1563 | 180.2185 | 228.8243 | 162.1113 | -1.6037 | 1.2734 | 32.1356 | 40.9226 | 25.2354 | 2.0027 | 32.0624 | 64.2105 | 16.0098 | 14.0473 | -0.1106 | 105.1331 | 217.109 | 166.1712 | 214.777 | 148.064
2012-10-30 17:27 | 27.5471 | 231.3061 | 180.3833 | 228.8995 | 162.2709 | -1.5962 | 1.2734 | 32.1355 | 40.9219 | 25.2356 | 2.0027 | 32.0623 | 64.2103 | 16.0098 | 14.0803 | -0.11 | 105.1322 | 217.2258 | 166.303 | 214.8192 | 148.1906
2012-10-30 17:27 | 27.5531 | 231.5166 | 180.4637 | 229.0589 | 162.2924 | -1.6037 | 1.2734 | 32.135 | 40.9194 | 25.2364 | 2.0027 | 32.0618 | 64.2085 | 16.0097 | 14.1026 | -0.1106 | 105.1279 | 217.414 | 166.3611 | 214.9563 | 148.1898
2012-10-30 17:27 | 27.5569 | 231.6423 | 180.6396 | 229.2711 | 162.4708 | -1.6014 | 1.2734 | 32.1352 | 40.9194 | 25.2367 | 2.0027 | 32.062 | 64.2094 | 16.0097 | 14.1145 | -0.1104 | 105.1289 | 217.5278 | 166.5251 | 215.1566 | 148.3563
2012-10-30 17:27 | 27.5576 | 231.7688 | 180.746 | 229.4881 | 162.5141 | -1.5957 | 1.2734 | 32.1353 | 40.92 | 25.2365 | 2.0027 | 32.0621 | 64.2096 | 16.0098 | 14.0704 | -0.11 | 105.1296 | 217.6984 | 166.6756 | 215.4177 | 148.4437
2012-10-30 17:27 | 27.5614 | 231.9102 | 180.8515 | 229.4998 | 162.7624 | -1.6014 | 1.2733 | 32.1354 | 40.9193 | 25.237 | 2.0027 | 32.0622 | 64.2108 | 16.0095 | 14.0823 | -0.1104 | 105.1301 | 217.8279 | 166.7692 | 215.4175 | 148.6801
2012-10-30 17:27 | 27.5689 | 232.071 | 180.9647 | 229.7433 | 162.7156 | -1.6037 | 1.2733 | 32.1352 | 40.9183 | 25.2374 | 2.0027 | 32.0621 | 64.2093 | 16.0098 | 14.0939 | -0.1106 | 105.1276 | 217.9771 | 166.8708 | 215.6494 | 148.6217
2012-10-30 17:27 | 27.5726 | 232.1808 | 181.1014 | 229.9043 | 162.9133 | -1.6043 | 1.2733 | 32.1352 | 40.9177 | 25.2377 | 2.0027 | 32.062 | 64.2095 | 16.0097 | 14.0609 | -0.1106 | 105.1272 | 218.1199 | 167.0405 | 215.8434 | 148.8524
2012-10-30 17:27 | 27.5756 | 232.3332 | 181.2179 | 229.9387 | 163.0055 | -1.6014 | 1.2733 | 32.1348 | 40.9165 | 25.2379 | 2.0026 | 32.0616 | 64.2074 | 16.0097 | 14.0639 | -0.1104 | 105.124 | 218.2693 | 167.154 | 215.8748 | 148.9416
2012-10-30 17:27 | 27.5793 | 232.4627 | 181.3038 | 230.0722 | 163.149 | -1.598 | 1.2732 | 32.1349 | 40.9155 | 25.2386 | 2.0026 | 32.0616 | 64.2076 | 16.0098 | 14.0921 | -0.1102 | 105.123 | 218.3706 | 167.2117 | 215.9801 | 149.0569
2012-10-30 17:27 | 27.5838 | 232.6283 | 181.4061 | 230.3362 | 163.2466 | -1.6026 | 1.2732 | 32.1347 | 40.9146 | 25.239 | 2.0026 | 32.0615 | 64.2074 | 16.0097 | 14.0395 | -0.1105 | 105.122 | 218.5888 | 167.3666 | 216.2967 | 149.2071
2012-10-30 17:27 | 27.5861 | 232.7761 | 181.5572 | 230.3346 | 163.245 | -1.6009 | 1.2732 | 32.1346 | 40.9142 | 25.2389 | 2.0026 | 32.0613 | 64.2069 | 16.0096 | 14.0296 | -0.1104 | 105.1211 | 218.7465 | 167.5276 | 216.305 | 149.2154
2012-10-30 17:27 | 27.5943 | 232.8668 | 181.6593 | 230.5119 | 163.164 | -1.598 | 1.2732 | 32.1343 | 40.913 | 25.2392 | 2.0026 | 32.0611 | 64.2059 | 16.0096 | 14.0133 | -0.1102 | 105.1189 | 218.8535 | 167.646 | 216.4986 | 149.1507
2012-10-30 17:28 | 27.6026 | 233.0402 | 181.7537 | 230.6893 | 163.2344 | -1.5968 | 1.2732 | 32.1339 | 40.9115 | 25.2395 | 2.0026 | 32.0606 | 64.2043 | 16.0096 | 14.0624 | -0.1101 | 105.1158 | 218.9778 | 167.6913 | 216.6269 | 149.172
2012-10-30 17:28 | 27.5973 | 233.145 | 181.862 | 230.9002 | 163.2524 | -1.598 | 1.2731 | 32.1336 | 40.9104 | 25.2398 | 2.0026 | 32.0604 | 64.203 | 16.0096 | 14.0449 | -0.1102 | 105.1134 | 219.1001 | 167.8171 | 216.8553 | 149.2075
2012-10-30 17:28 | 27.6055 | 233.3104 | 181.9406 | 230.8613 | 163.3576 | -1.5974 | 1.2731 | 32.1338 | 40.9104 | 25.2401 | 2.0026 | 32.0607 | 64.2038 | 16.0097 | 14.0082 | -0.1101 | 105.1141 | 219.3022 | 167.9324 | 216.8531 | 149.3494
2012-10-30 17:28 | 27.6108 | 233.41 | 182.032 | 231.0828 | 163.3435 | -1.6003 | 1.2731 | 32.1339 | 40.9103 | 25.2403 | 2.0026 | 32.0606 | 64.204 | 16.0096 | 14.0706 | -0.1103 | 105.1143 | 219.3394 | 167.9614 | 217.0122 | 149.2729
2012-10-30 17:28 | 27.6115 | 233.5129 | 182.0641 | 231.2093 | 163.5422 | -1.5985 | 1.2731 | 32.1334 | 40.9079 | 25.241 | 2.0026 | 32.0603 | 64.2029 | 16.0096 | 14.0387 | -0.1102 | 105.1107 | 219.4742 | 168.0254 | 217.1706 | 149.5035
2012-10-30 17:28 | 27.622 | 233.7199 | 182.1959 | 231.3574 | 163.7002 | -1.5991 | 1.273 | 32.1333 | 40.9072 | 25.2412 | 2.0025 | 32.0601 | 64.2018 | 16.0096 | 14.0573 | -0.1102 | 105.1089 | 219.6626 | 168.1386 | 217.3001 | 149.6429
2012-10-30 17:28 | 27.6295 | 233.8571 | 182.3052 | 231.4279 | 163.7931 | -1.5968 | 1.2731 | 32.1335 | 40.9077 | 25.2413 | 2.0026 | 32.0603 | 64.2029 | 16.0096 | 14.0771 | -0.1101 | 105.1105 | 219.78 | 168.2281 | 217.3508 | 149.716
2012-10-30 17:28 | 27.622 | 233.9439 | 182.376 | 231.6562 | 163.7157 | -1.6026 | 1.273 | 32.1333 | 40.9065 | 25.2417 | 2.0025 | 32.06 | 64.2014 | 16.0097 | 14.0247 | -0.1105 | 105.1079 | 219.9192 | 168.3513 | 217.6315 | 149.691
2012-10-30 17:28 | 27.6325 | 234.0173 | 182.4609 | 231.7572 | 163.9786 | -1.5985 | 1.273 | 32.1329 | 40.9053 | 25.2418 | 2.0025 | 32.0596 | 64.2001 | 16.0096 | 14.0352 | -0.1102 | 105.1054 | 219.9821 | 168.4257 | 217.722 | 149.9434
2012-10-30 17:28 | 27.6355 | 234.1932 | 182.5266 | 231.7759 | 163.9389 | -1.6014 | 1.273 | 32.1326 | 40.9037 | 25.2424 | 2.0025 | 32.0594 | 64.1992 | 16.0096 | 14.0423 | -0.1104 | 105.103 | 220.1509 | 168.4843 | 217.7336 | 149.8966
2012-10-30 17:28 | 27.6445 | 234.3122 | 182.6256 | 232.0365 | 163.9479 | -1.5928 | 1.2729 | 32.1323 | 40.9025 | 25.2426 | 2.0025 | 32.059 | 64.1972 | 16.0097 | 14.035 | -0.1098 | 105.0997 | 220.2772 | 168.5906 | 218.0015 | 149.9129
2012-10-30 17:28 | 27.6512 | 234.3857 | 182.6715 | 232.1258 | 164.3002 | -1.6009 | 1.2729 | 32.1315 | 40.9 | 25.2429 | 2.0024 | 32.0582 | 64.1945 | 16.0096 | 14.0539 | -0.1104 | 105.0945 | 220.3318 | 168.6176 | 218.0719 | 150.2463
2012-10-30 17:28 | 27.6505 | 234.5343 | 182.8235 | 232.1486 | 164.2956 | -1.5985 | 1.2729 | 32.1317 | 40.8998 | 25.2433 | 2.0025 | 32.0584 | 64.1954 | 16.0096 | 13.9921 | -0.1102 | 105.0951 | 220.5422 | 168.8314 | 218.1565 | 150.3035
2012-10-30 17:28 | 27.652 | 234.6026 | 182.8876 | 232.4135 | 164.3476 | -1.5974 | 1.2729 | 32.1318 | 40.9001 | 25.2433 | 2.0025 | 32.0585 | 64.1959 | 16.0095 | 14.0099 | -0.1101 | 105.096 | 220.5927 | 168.8777 | 218.4036 | 150.3377
2012-10-30 17:28 | 27.6527 | 234.7762 | 183.0019 | 232.4614 | 164.5424 | -1.6049 | 1.2728 | 32.1313 | 40.8983 | 25.2436 | 2.0024 | 32.058 | 64.1936 | 16.0097 | 14.0351 | -0.1106 | 105.0919 | 220.7411 | 168.9668 | 218.4263 | 150.5073
2012-10-30 17:28 | 27.6565 | 234.9136 | 183.0488 | 232.4731 | 164.5851 | -1.6037 | 1.2728 | 32.1308 | 40.8969 | 25.2436 | 2.0024 | 32.0575 | 64.1917 | 16.0096 | 14.0633 | -0.1106 | 105.0886 | 220.8503 | 168.9855 | 218.4098 | 150.5218
2012-10-30 17:28 | 27.6639 | 234.9799 | 183.1384 | 232.7831 | 164.5303 | -1.5974 | 1.2728 | 32.1307 | 40.8957 | 25.2443 | 2.0024 | 32.0574 | 64.1904 | 16.0098 | 14.0259 | -0.1101 | 105.0862 | 220.954 | 169.1125 | 218.7572 | 150.5044
2012-10-30 17:28 | 27.6737 | 235.0997 | 183.2618 | 232.7733 | 164.68 | -1.6026 | 1.2728 | 32.1306 | 40.8949 | 25.2446 | 2.0024 | 32.0573 | 64.1909 | 16.0096 | 14.0479 | -0.1105 | 105.0857 | 221.0518 | 169.2139 | 218.7254 | 150.6321
2012-10-30 17:28 | 27.6782 | 235.1789 | 183.3681 | 232.8249 | 164.6223 | -1.6032 | 1.2728 | 32.1304 | 40.8941 | 25.2447 | 2.0024 | 32.0572 | 64.1905 | 16.0096 | 14.0687 | -0.1105 | 105.0847 | 221.1102 | 169.2994 | 218.7562 | 150.5536
2012-10-30 17:28 | 27.6857 | 235.2846 | 183.407 | 233.0368 | 164.7347 | -1.5991 | 1.2727 | 32.1302 | 40.8934 | 25.245 | 2.0024 | 32.0569 | 64.1898 | 16.0095 | 14.0517 | -0.1102 | 105.0831 | 221.2329 | 169.3553 | 218.9851 | 150.683
2012-10-30 17:28 | 27.6887 | 235.4133 | 183.5157 | 233.1342 | 164.8465 | -1.6032 | 1.2727 | 32.13 | 40.8927 | 25.245 | 2.0024 | 32.0568 | 64.1895 | 16.0094 | 14.0833 | -0.1105 | 105.0822 | 221.33 | 169.4324 | 219.0509 | 150.7632
2012-10-30 17:28 | 27.6932 | 235.5317 | 183.5829 | 233.174 | 164.9403 | -1.5974 | 1.2727 | 32.13 | 40.8924 | 25.2452 | 2.0024 | 32.0567 | 64.1889 | 16.0095 | 14.0715 | -0.1101 | 105.0813 | 221.4602 | 169.5114 | 219.1025 | 150.8688
2012-10-30 17:28 | 27.7036 | 235.6325 | 183.6129 | 233.3299 | 165.04 | -1.5985 | 1.2727 | 32.1294 | 40.8905 | 25.2454 | 2.0023 | 32.056 | 64.1858 | 16.0096 | 14.086 | -0.1102 | 105.0763 | 221.5465 | 169.5269 | 219.2439 | 150.954
2012-10-30 17:28 | 27.7066 | 235.7494 | 183.7412 | 233.5176 | 165.2956 | -1.5991 | 1.2726 | 32.1288 | 40.8882 | 25.246 | 2.0023 | 32.0556 | 64.1844 | 16.0095 | 14.0808 | -0.1102 | 105.0726 | 221.6686 | 169.6604 | 219.4368 | 151.2148
2012-10-30 17:28 | 27.7044 | 235.8611 | 183.8369 | 233.4761 | 165.4527 | -1.598 | 1.2726 | 32.1287 | 40.888 | 25.2459 | 2.0023 | 32.0555 | 64.1835 | 16.0096 | 14.0582 | -0.1102 | 105.0715 | 221.8029 | 169.7787 | 219.4179 | 151.3945
2012-10-30 17:28 | 27.7194 | 235.9901 | 183.9264 | 233.6994 | 165.4832 | -1.5974 | 1.2726 | 32.1285 | 40.887 | 25.2463 | 2.0023 | 32.0553 | 64.1831 | 16.0096 | 14.0651 | -0.1101 | 105.07 | 221.925 | 169.8613 | 219.6343 | 151.4181
2012-10-30 17:28 | 27.7156 | 236.0491 | 184.0558 | 233.8213 | 165.6543 | -1.5997 | 1.2726 | 32.1282 | 40.8857 | 25.2465 | 2.0022 | 32.0549 | 64.1815 | 16.0095 | 14.0287 | -0.1103 | 105.0672 | 222.0204 | 170.0271 | 219.7926 | 151.6256
2012-10-30 17:29 | 27.7224 | 236.1502 | 184.1448 | 233.8046 | 165.7038 | -1.6026 | 1.2726 | 32.1281 | 40.8852 | 25.2467 | 2.0022 | 32.0548 | 64.1812 | 16.0095 | 14.0314 | -0.1105 | 105.0663 | 222.1188 | 170.1134 | 219.7732 | 151.6724
2012-10-30 17:29 | 27.7216 | 236.3104 | 184.2616 | 234.0042 | 165.6875 | -1.6049 | 1.2726 | 32.1278 | 40.8843 | 25.2468 | 2.0022 | 32.0545 | 64.1804 | 16.0095 | 14.0061 | -0.1106 | 105.0647 | 222.3043 | 170.2555 | 219.9981 | 151.6814
2012-10-30 17:29 | 27.7268 | 236.3078 | 184.3177 | 234.1469 | 165.7121 | -1.6003 | 1.2725 | 32.1275 | 40.8825 | 25.2473 | 2.0022 | 32.0541 | 64.1782 | 16.0095 | 13.995 | -0.1103 | 105.0607 | 222.3128 | 170.3227 | 220.1519 | 151.7171
2012-10-30 17:29 | 27.7291 | 236.4554 | 184.4493 | 234.1256 | 165.7727 | -1.598 | 1.2725 | 32.1271 | 40.8814 | 25.2474 | 2.0022 | 32.0538 | 64.1776 | 16.0094 | 14.0177 | -0.1102 | 105.0591 | 222.4377 | 170.4316 | 220.1079 | 151.755
2012-10-30 17:29 | 27.7328 | 236.5534 | 184.4922 | 234.3376 | 165.7181 | -1.5991 | 1.2725 | 32.1267 | 40.88 | 25.2477 | 2.0022 | 32.0534 | 64.1762 | 16.0094 | 14.0132 | -0.1102 | 105.0562 | 222.5402 | 170.479 | 220.3244 | 151.7049
2012-10-30 17:29 | 27.7441 | 236.6511 | 184.5348 | 234.3568 | 165.9432 | -1.6043 | 1.2724 | 32.1262 | 40.8787 | 25.2476 | 2.0021 | 32.0529 | 64.1738 | 16.0095 | 14.0408 | -0.1106 | 105.0525 | 222.6103 | 170.494 | 220.316 | 151.9024
2012-10-30 17:29 | 27.7486 | 236.7695 | 184.6373 | 234.4517 | 166.1031 | -1.6009 | 1.2724 | 32.1259 | 40.8768 | 25.2484 | 2.0021 | 32.0526 | 64.1721 | 16.0096 | 13.9883 | -0.1104 | 105.049 | 222.7812 | 170.649 | 220.4634 | 152.1148
2012-10-30 17:29 | 27.7478 | 236.859 | 184.7187 | 234.6316 | 165.9741 | -1.598 | 1.2674 | 31.9999 | 40.5581 | 25.2476 | 1.9947 | 31.9338 | 63.697 | 16.0096 | 14.016 | -0.1102 | 104.2551 | 222.843 | 170.7027 | 220.6156 | 151.9581
2012-10-30 17:29 | 27.7523 | 236.9578 | 184.7507 | 234.7226 | 166.1418 | -1.6009 | 1.2661 | 31.9663 | 40.474 | 25.2469 | 1.9921 | 31.8933 | 63.5353 | 16.0097 | 13.992 | -0.1104 | 104.0093 | 222.9658 | 170.7587 | 220.7306 | 152.1498
2012-10-30 17:29 | 27.7635 | 237.0122 | 184.8638 | 234.7023 | 166.2036 | -1.6014 | 1.2656 | 31.9515 | 40.4363 | 25.2471 | 1.9912 | 31.8783 | 63.4754 | 16.0097 | 14.0521 | -0.1104 | 103.9118 | 222.9601 | 170.8117 | 220.6502 | 152.1515
2012-10-30 17:29 | 27.7546 | 237.0936 | 184.9097 | 234.9095 | 166.2258 | -1.5991 | 1.2651 | 31.9401 | 40.4071 | 25.2473 | 1.9905 | 31.8671 | 63.4307 | 16.0098 | 14.0554 | -0.1102 | 103.8378 | 223.0382 | 170.8543 | 220.8541 | 152.1704
2012-10-30 17:29 | 27.7658 | 237.1951 | 184.9797 | 234.9757 | 166.4585 | -1.6003 | 1.2649 | 31.9345 | 40.3925 | 25.2476 | 1.9901 | 31.8616 | 63.4084 | 16.0099 | 14.0258 | -0.1103 | 103.8009 | 223.1693 | 170.9539 | 220.9499 | 152.4327
2012-10-30 17:29 | 27.759 | 237.2197 | 185.0395 | 234.9414 | 166.5606 | -1.5997 | 1.2648 | 31.9337 | 40.39 | 25.2478 | 1.9901 | 31.8608 | 63.4049 | 16.01 | 13.966 | -0.1103 | 103.7948 | 223.2537 | 171.0735 | 220.9754 | 152.5946
2012-10-30 17:29 | 27.7643 | 237.3036 | 185.1545 | 235.1982 | 166.2821 | -1.5957 | 1.2648 | 31.9345 | 40.3919 | 25.2479 | 1.9901 | 31.8617 | 63.4079 | 16.0101 | 14.0243 | -0.11 | 103.7997 | 223.2793 | 171.1302 | 221.1739 | 152.2578
2012-10-30 17:29 | 27.7665 | 237.345 | 185.2077 | 235.1965 | 166.4981 | -1.6026 | 1.2649 | 31.9361 | 40.396 | 25.2479 | 1.9902 | 31.8633 | 63.4144 | 16.0101 | 13.9695 | -0.1105 | 103.8104 | 223.3755 | 171.2382 | 221.227 | 152.5286
2012-10-30 17:29 | 27.768 | 237.4565 | 185.2523 | 235.2608 | 166.243 | -1.5974 | 1.265 | 31.939 | 40.4029 | 25.2482 | 1.9904 | 31.8662 | 63.4257 | 16.0102 | 14.024 | -0.1101 | 103.8286 | 223.4325 | 171.2283 | 221.2368 | 152.219
2012-10-30 17:29 | 27.774 | 237.5646 | 185.2857 | 235.4397 | 166.2918 | -1.6003 | 1.2651 | 31.942 | 40.4099 | 25.2485 | 1.9905 | 31.869 | 63.4363 | 16.0103 | 14.0259 | -0.1103 | 103.8462 | 223.5387 | 171.2598 | 221.4138 | 152.2659
2012-10-30 17:29 | 27.7673 | 237.6286 | 185.3378 | 235.4015 | 166.1801 | -1.6014 | 1.2654 | 31.9496 | 40.4288 | 25.2487 | 1.9913 | 31.8819 | 63.4878 | 16.0102 | 14.0518 | -0.1104 | 103.9165 | 223.5768 | 171.286 | 221.3497 | 152.1283
2012-10-30 17:29 | 27.7688 | 237.6654 | 185.3589 | 235.4502 | 166.3954 | -1.5985 | 1.2701 | 32.0689 | 40.7296 | 25.2498 | 1.9985 | 31.9959 | 63.9433 | 16.0101 | 14.0125 | -0.1102 | 104.6729 | 223.6529 | 171.3464 | 221.4377 | 152.3829
2012-10-30 17:29 | 27.7703 | 237.7533 | 185.3957 | 235.656 | 166.362 | -1.5985 | 1.2706 | 32.0832 | 40.7648 | 25.2506 | 1.9994 | 32.0103 | 64.0009 | 16.0101 | 14.0711 | -0.1102 | 104.7658 | 223.6822 | 171.3246 | 221.5849 | 152.2909
2012-10-30 17:29 | 27.7673 | 237.8367 | 185.4907 | 235.5626 | 166.5183 | -1.6061 | 1.2712 | 32.0978 | 40.8016 | 25.2507 | 2.0003 | 32.0249 | 64.0592 | 16.0101 | 14.06 | -0.1107 | 104.8609 | 223.7767 | 171.4307 | 221.5026 | 152.4583
2012-10-30 17:29 | 27.771 | 237.9621 | 185.5453 | 235.7981 | 166.8331 | -1.5997 | 1.2717 | 32.1107 | 40.8338 | 25.2511 | 2.0011 | 32.0373 | 64.109 | 16.01 | 14.0718 | -0.1103 | 104.9428 | 223.8903 | 171.4735 | 221.7263 | 152.7613
2012-10-30 17:29 | 27.7665 | 237.9183 | 185.6427 | 235.7622 | 167.093 | -1.6043 | 1.272 | 32.1194 | 40.8558 | 25.2512 | 2.0016 | 32.046 | 64.1438 | 16.0101 | 14.0632 | -0.1106 | 104.9996 | 223.8551 | 171.5795 | 221.699 | 153.0298
2012-10-30 17:29 | 27.7703 | 238.0753 | 185.6975 | 235.821 | 167.1085 | -1.6026 | 1.2722 | 32.1253 | 40.87 | 25.2517 | 2.002 | 32.052 | 64.168 | 16.01 | 14.063 | -0.1105 | 105.038 | 224.0123 | 171.6345 | 221.758 | 153.0455
2012-10-30 17:29 | 27.7755 | 238.159 | 185.8007 | 235.9833 | 167.2342 | -1.6003 | 1.2724 | 32.1294 | 40.8803 | 25.2518 | 2.0023 | 32.056 | 64.1844 | 16.0099 | 14.0804 | -0.1103 | 105.0647 | 224.0786 | 171.7203 | 221.9029 | 153.1538
2012-10-30 17:29 | 27.7778 | 238.2555 | 185.9794 | 235.8835 | 167.3182 | -1.5991 | 1.2725 | 32.1327 | 40.8888 | 25.2517 | 2.0025 | 32.0593 | 64.1972 | 16.01 | 14.0378 | -0.1102 | 105.086 | 224.2177 | 171.9416 | 221.8457 | 153.2804
2012-10-30 17:29 | 27.7748 | 238.3507 | 185.9764 | 236.1358 | 167.3619 | -1.6078 | 1.2725 | 32.1349 | 40.8929 | 25.2525 | 2.0026 | 32.0613 | 64.206 | 16.0098 | 14.0756 | -0.1108 | 105.0989 | 224.2751 | 171.9008 | 222.0602 | 153.2863
2012-10-30 17:30 | 27.7778 | 238.3929 | 186.1205 | 236.1899 | 167.5438 | -1.5997 | 1.2726 | 32.1361 | 40.8953 | 25.253 | 2.0027 | 32.0626 | 64.211 | 16.0099 | 14.099 | -0.1103 | 105.1063 | 224.2939 | 172.0215 | 222.0909 | 153.4448
2012-10-30 17:30 | 27.7785 | 238.5703 | 186.2505 | 236.1552 | 167.6184 | -1.6043 | 1.2726 | 32.1371 | 40.8982 | 25.2528 | 2.0027 | 32.0637 | 64.215 | 16.01 | 14.0549 | -0.1106 | 105.1133 | 224.5154 | 172.1956 | 222.1003 | 153.5635
2012-10-30 17:30 | 27.777 | 238.6041 | 186.2804 | 236.413 | 167.4885 | -1.5974 | 1.2727 | 32.1383 | 40.9009 | 25.253 | 2.0028 | 32.0648 | 64.2201 | 16.0098 | 14.0941 | -0.1101 | 105.121 | 224.51 | 172.1863 | 222.3189 | 153.3944
2012-10-30 17:30 | 27.7867 | 238.7316 | 186.3646 | 236.4384 | 167.471 | -1.5997 | 1.2727 | 32.1387 | 40.9014 | 25.2533 | 2.0028 | 32.0652 | 64.2215 | 16.0099 | 14.0998 | -0.1103 | 105.1229 | 224.6318 | 172.2648 | 222.3386 | 153.3712
2012-10-30 17:30 | 27.7763 | 238.7566 | 186.413 | 236.4594 | 167.4645 | -1.5985 | 1.2727 | 32.1392 | 40.9022 | 25.2536 | 2.0029 | 32.0657 | 64.2235 | 16.0099 | 14.0527 | -0.1102 | 105.1256 | 224.7039 | 172.3603 | 222.4067 | 153.4118
2012-10-30 17:30 | 27.774 | 238.8485 | 186.4538 | 236.626 | 167.4194 | -1.6101 | 1.2726 | 32.1392 | 40.9017 | 25.2539 | 2.0029 | 32.0657 | 64.2235 | 16.0098 | 14.0667 | -0.111 | 105.1252 | 224.7818 | 172.3871 | 222.5593 | 153.3527
2012-10-30 17:30 | 27.7845 | 238.9689 | 186.531 | 236.6561 | 167.7839 | -1.5997 | 1.2727 | 32.1395 | 40.9025 | 25.2538 | 2.0029 | 32.0661 | 64.2248 | 16.0099 | 14.0527 | -0.1103 | 105.1273 | 224.9162 | 172.4783 | 222.6034 | 153.7312
2012-10-30 17:30 | 27.7905 | 239.0652 | 186.6232 | 236.6818 | 167.7977 | -1.6014 | 1.2727 | 32.1397 | 40.9028 | 25.254 | 2.0029 | 32.0661 | 64.2252 | 16.0099 | 14.0424 | -0.1104 | 105.128 | 225.0228 | 172.5808 | 222.6394 | 153.7553
2012-10-30 17:30 | 27.7875 | 239.0622 | 186.6163 | 236.8791 | 167.923 | -1.6026 | 1.2727 | 32.1403 | 40.9037 | 25.2545 | 2.0029 | 32.0669 | 64.2281 | 16.0099 | 14.0883 | -0.1105 | 105.1319 | 224.9739 | 172.528 | 222.7908 | 153.8347
2012-10-30 17:30 | 27.7897 | 239.139 | 186.7557 | 236.8145 | 168.1198 | -1.602 | 1.2727 | 32.1406 | 40.9042 | 25.2546 | 2.003 | 32.0671 | 64.2296 | 16.0098 | 14.0661 | -0.1104 | 105.1338 | 225.0729 | 172.6896 | 222.7484 | 154.0537
2012-10-30 17:30 | 27.786 | 239.2648 | 186.7913 | 236.9365 | 167.9605 | -1.6014 | 1.2727 | 32.1408 | 40.9045 | 25.2547 | 2.003 | 32.0673 | 64.2298 | 16.0098 | 14.046 | -0.1104 | 105.1344 | 225.2188 | 172.7453 | 222.8905 | 153.9145
2012-10-30 17:30 | 27.7875 | 239.2702 | 186.8829 | 237.15 | 168.0514 | -1.5974 | 1.2727 | 32.1404 | 40.9035 | 25.2547 | 2.0029 | 32.0669 | 64.2283 | 16.0099 | 14.0557 | -0.1101 | 105.1318 | 225.2145 | 172.8272 | 223.0943 | 153.9957
2012-10-30 17:30 | 27.7957 | 239.3844 | 186.9382 | 237.0758 | 168.2036 | -1.6043 | 1.2726 | 32.1406 | 40.9029 | 25.2553 | 2.003 | 32.067 | 64.2294 | 16.0097 | 14.0354 | -0.1106 | 105.1323 | 225.349 | 172.9028 | 223.0404 | 154.1682
2012-10-30 17:30 | 27.8002 | 239.4792 | 187.0133 | 237.2059 | 168.2353 | -1.6066 | 1.2726 | 32.1405 | 40.9033 | 25.255 | 2.003 | 32.0669 | 64.2289 | 16.0098 | 14.0602 | -0.1108 | 105.1321 | 225.419 | 172.9531 | 223.1457 | 154.1751
2012-10-30 17:30 | 27.798 | 239.5672 | 187.0817 | 237.3962 | 168.4043 | -1.6043 | 1.2726 | 32.141 | 40.9041 | 25.2553 | 2.003 | 32.0674 | 64.2301 | 16.0099 | 14.0866 | -0.1106 | 105.1341 | 225.4806 | 172.9951 | 223.3096 | 154.3177
2012-10-30 17:30 | 27.8047 | 239.625 | 187.1629 | 237.3322 | 168.6329 | -1.6026 | 1.2726 | 32.1409 | 40.9035 | 25.2555 | 2.003 | 32.0674 | 64.2306 | 16.0098 | 14.0607 | -0.1105 | 105.1342 | 225.5643 | 173.1022 | 223.2715 | 154.5722
2012-10-30 17:30 | 27.8032 | 239.6666 | 187.2319 | 237.4524 | 168.3512 | -1.5985 | 1.2726 | 32.1409 | 40.9033 | 25.2556 | 2.003 | 32.0674 | 64.2306 | 16.0098 | 14.0592 | -0.1102 | 105.1338 | 225.6074 | 173.1727 | 223.3932 | 154.292
2012-10-30 17:30 | 27.8025 | 239.7484 | 187.2861 | 237.542 | 168.5528 | -1.6049 | 1.2726 | 32.1403 | 40.901 | 25.256 | 2.003 | 32.0668 | 64.2283 | 16.0098 | 14.0177 | -0.1106 | 105.1293 | 225.7307 | 173.2684 | 223.5243 | 154.5351
2012-10-30 17:30 | 27.8047 | 239.833 | 187.3472 | 237.5207 | 168.4188 | -1.5985 | 1.2726 | 32.1402 | 40.9001 | 25.2565 | 2.0029 | 32.0665 | 64.2275 | 16.0097 | 14.0688 | -0.1102 | 105.1276 | 225.7642 | 173.2784 | 223.4519 | 154.35
2012-10-30 17:30 | 27.801 | 239.9078 | 187.3631 | 237.7133 | 168.2828 | -1.6003 | 1.2725 | 32.1398 | 40.8991 | 25.2564 | 2.0029 | 32.0662 | 64.2258 | 16.0098 | 14.0447 | -0.1103 | 105.125 | 225.8631 | 173.3184 | 223.6686 | 154.2381
2012-10-30 17:30 | 27.81 | 240.0345 | 187.3838 | 237.7105 | 168.5915 | -1.6026 | 1.2725 | 32.1397 | 40.8982 | 25.2568 | 2.0029 | 32.0661 | 64.2256 | 16.0098 | 14.0374 | -0.1105 | 105.1238 | 225.9971 | 173.3464 | 223.6731 | 154.5541
2012-10-30 17:30 | 27.8062 | 240.0032 | 187.4467 | 237.7224 | 168.5098 | -1.5985 | 1.2725 | 32.1396 | 40.8989 | 25.2563 | 2.0029 | 32.0662 | 64.2261 | 16.0097 | 14.0377 | -0.1102 | 105.125 | 225.9655 | 173.409 | 223.6847 | 154.4721
2012-10-30 17:30 | 27.8122 | 240.0603 | 187.4998 | 237.913 | 168.652 | -1.5997 | 1.2725 | 32.1395 | 40.8981 | 25.2566 | 2.0029 | 32.0661 | 64.2257 | 16.0097 | 14.0478 | -0.1103 | 105.1238 | 226.0125 | 173.452 | 223.8652 | 154.6042
2012-10-30 17:30 | 27.8182 | 240.2232 | 187.5411 | 237.8012 | 168.767 | -1.6009 | 1.2725 | 32.1394 | 40.8976 | 25.2568 | 2.0029 | 32.0659 | 64.2255 | 16.0096 | 14.0579 | -0.1104 | 105.123 | 226.1653 | 173.4832 | 223.7433 | 154.7091
2012-10-30 17:30 | 27.8212 | 240.2773 | 187.6225 | 238.0869 | 168.7778 | -1.6009 | 1.2725 | 32.1393 | 40.8967 | 25.2571 | 2.0029 | 32.0657 | 64.2239 | 16.0097 | 14.0568 | -0.1104 | 105.1205 | 226.2205 | 173.5657 | 224.0301 | 154.721
2012-10-30 17:30 | 27.8152 | 240.3105 | 187.691 | 238.0259 | 168.9586 | -1.5933 | 1.2725 | 32.1394 | 40.897 | 25.2572 | 2.0029 | 32.0659 | 64.2253 | 16.0096 | 14.0916 | -0.1098 | 105.1223 | 226.2189 | 173.5994 | 223.9343 | 154.867
2012-10-30 17:30 | 27.8092 | 240.4026 | 187.7517 | 238.1142 | 168.8281 | -1.6026 | 1.2725 | 32.1393 | 40.8967 | 25.2572 | 2.0029 | 32.0658 | 64.225 | 16.0095 | 14.057 | -0.1105 | 105.1217 | 226.3456 | 173.6947 | 224.0572 | 154.7711
2012-10-30 17:31 | 27.8227 | 240.3926 | 187.7652 | 238.2337 | 168.9895 | -1.5968 | 1.2725 | 32.1389 | 40.8953 | 25.2575 | 2.0029 | 32.0654 | 64.2236 | 16.0096 | 14.0909 | -0.1101 | 105.1189 | 226.3017 | 173.6743 | 224.1428 | 154.8986
2012-10-30 17:31 | 27.8264 | 240.5061 | 187.8826 | 238.1392 | 169.0165 | -1.6009 | 1.2724 | 32.1387 | 40.8947 | 25.2575 | 2.0029 | 32.0652 | 64.2222 | 16.0097 | 14.0498 | -0.1104 | 105.117 | 226.4563 | 173.8328 | 224.0894 | 154.9667
2012-10-30 17:31 | 27.8257 | 240.6114 | 187.9016 | 238.3662 | 168.8173 | -1.6043 | 1.2724 | 32.1386 | 40.8942 | 25.2576 | 2.0029 | 32.0651 | 64.2227 | 16.0095 | 14.0898 | -0.1106 | 105.1169 | 226.5216 | 173.8118 | 224.2764 | 154.7275
2012-10-30 17:31 | 27.8272 | 240.5893 | 187.9031 | 238.407 | 168.9628 | -1.5968 | 1.2724 | 32.1385 | 40.8931 | 25.2581 | 2.0028 | 32.0648 | 64.221 | 16.0096 | 14.075 | -0.1101 | 105.114 | 226.5143 | 173.8281 | 224.332 | 154.8878
2012-10-30 17:31 | 27.8287 | 240.6968 | 188.0105 | 238.3378 | 168.8709 | -1.6043 | 1.2724 | 32.1381 | 40.8926 | 25.2578 | 2.0028 | 32.0647 | 64.2205 | 16.0096 | 14.0561 | -0.1106 | 105.113 | 226.6407 | 173.9544 | 224.2817 | 154.8148
2012-10-30 17:31 | 27.8302 | 240.7415 | 188.0473 | 238.5592 | 168.8685 | -1.602 | 1.2724 | 32.1376 | 40.8907 | 25.2582 | 2.0028 | 32.0642 | 64.2191 | 16.0094 | 14.0617 | -0.1104 | 105.1099 | 226.6798 | 173.9856 | 224.4975 | 154.8068
2012-10-30 17:31 | 27.8317 | 240.896 | 188.1076 | 238.6471 | 169.0374 | -1.5997 | 1.2723 | 32.1371 | 40.8887 | 25.2586 | 2.0028 | 32.0635 | 64.2164 | 16.0095 | 14.0999 | -0.1103 | 105.1051 | 226.7961 | 174.0077 | 224.5472 | 154.9375
2012-10-30 17:31 | 27.8362 | 241.081 | 188.1709 | 238.8871 | 169.2326 | -1.6026 | 1.2723 | 32.1362 | 40.8863 | 25.2587 | 2.0027 | 32.0627 | 64.213 | 16.0095 | 14.0555 | -0.1105 | 105.0993 | 227.0255 | 174.1154 | 224.8316 | 155.1771
2012-10-30 17:31 | 27.8444 | 241.301 | 188.391 | 239.0602 | 169.3187 | -1.6003 | 1.2722 | 32.1352 | 40.8828 | 25.2593 | 2.0027 | 32.0616 | 64.2093 | 16.0093 | 14.0922 | -0.1103 | 105.0921 | 227.2088 | 174.2988 | 224.968 | 155.2265
2012-10-30 17:31 | 27.8549 | 241.5548 | 188.5819 | 239.4043 | 169.4849 | -1.5968 | 1.2721 | 32.1337 | 40.878 | 25.26 | 2.0026 | 32.0602 | 64.2046 | 16.0091 | 14.0619 | -0.1101 | 105.0826 | 227.4929 | 174.52 | 225.3424 | 155.423
2012-10-30 17:31 | 27.8571 | 241.5962 | 188.5763 | 239.3477 | 169.5572 | -1.6026 | 1.2721 | 32.1334 | 40.877 | 25.26 | 2.0026 | 32.0598 | 64.202 | 16.0093 | 14.0805 | -0.1105 | 105.079 | 227.5157 | 174.4958 | 225.2672 | 155.4767
2012-10-30 17:31 | 27.8579 | 241.6833 | 188.6398 | 239.4858 | 169.5035 | -1.5985 | 1.2721 | 32.1329 | 40.8751 | 25.2605 | 2.0026 | 32.0594 | 64.2006 | 16.0092 | 14.0649 | -0.1102 | 105.0757 | 227.6184 | 174.5749 | 225.4209 | 155.4386
2012-10-30 17:31 | 27.8624 | 241.7466 | 188.6953 | 239.5845 | 169.5275 | -1.5985 | 1.2721 | 32.1326 | 40.8745 | 25.2604 | 2.0025 | 32.059 | 64.1996 | 16.0092 | 14.0858 | -0.1102 | 105.074 | 227.6608 | 174.6095 | 225.4987 | 155.4417
2012-10-30 17:31 | 27.8646 | 241.7567 | 188.7054 | 239.5082 | 169.6387 | -1.5997 | 1.272 | 32.1321 | 40.8735 | 25.2601 | 2.0025 | 32.0586 | 64.1975 | 16.0092 | 14.0961 | -0.1103 | 105.0709 | 227.6606 | 174.6093 | 225.4121 | 155.5426
2012-10-30 17:31 | 27.8676 | 241.846 | 188.7398 | 239.6368 | 169.6962 | -1.6003 | 1.272 | 32.1318 | 40.8723 | 25.2604 | 2.0025 | 32.0583 | 64.1961 | 16.0092 | 14.0624 | -0.1103 | 105.0684 | 227.7836 | 174.6774 | 225.5744 | 155.6338
2012-10-30 17:31 | 27.8736 | 241.8598 | 188.8321 | 239.6978 | 169.7061 | -1.5945 | 1.272 | 32.1311 | 40.8702 | 25.2606 | 2.0025 | 32.0576 | 64.194 | 16.0092 | 14.0643 | -0.1099 | 105.0642 | 227.7955 | 174.7678 | 225.6335 | 155.6418
2012-10-30 17:31 | 27.8789 | 241.9397 | 188.8295 | 239.7188 | 169.7503 | -1.5985 | 1.272 | 32.1306 | 40.8688 | 25.2607 | 2.0024 | 32.0571 | 64.192 | 16.0092 | 14.0574 | -0.1102 | 105.0609 | 227.8823 | 174.7721 | 225.6614 | 155.6929
2012-10-30 17:31 | 27.8759 | 242.0347 | 188.8579 | 239.9003 | 169.7473 | -1.6043 | 1.2719 | 32.1303 | 40.8677 | 25.2609 | 2.0024 | 32.0567 | 64.1906 | 16.0091 | 14.0544 | -0.1106 | 105.0583 | 227.9803 | 174.8035 | 225.8459 | 155.6929
2012-10-30 17:31 | 27.8796 | 242.0463 | 188.9126 | 239.849 | 169.9574 | -1.5974 | 1.2719 | 32.1303 | 40.8676 | 25.261 | 2.0024 | 32.0569 | 64.1908 | 16.0092 | 14.0581 | -0.1101 | 105.0584 | 227.9882 | 174.8545 | 225.7909 | 155.8993
2012-10-30 17:32 | 27.8893 | 242.0991 | 188.9773 | 239.7881 | 169.9476 | -1.6061 | 1.2719 | 32.1302 | 40.8673 | 25.261 | 2.0024 | 32.0567 | 64.1903 | 16.0091 | 14.0474 | -0.1107 | 105.0577 | 228.0517 | 174.9299 | 225.7407 | 155.9002
2012-10-30 17:32 | 27.8916 | 242.1289 | 189.0149 | 240.0573 | 170.0161 | -1.5968 | 1.2705 | 32.0944 | 40.7753 | 25.2616 | 2.0004 | 32.0244 | 64.0614 | 16.0091 | 14.0701 | -0.1101 | 104.8367 | 228.0588 | 174.9448 | 225.9872 | 155.946
2012-10-30 17:32 | 27.8961 | 242.1726 | 189.0586 | 240.0068 | 170.1803 | -1.5968 | 1.2655 | 31.967 | 40.4549 | 25.26 | 1.9922 | 31.8928 | 63.5353 | 16.0092 | 14.0991 | -0.1101 | 103.9902 | 228.0735 | 174.9595 | 225.9077 | 156.0812
2012-10-30 17:32 | 27.9021 | 242.2257 | 189.1235 | 240.005 | 170.1084 | -1.598 | 1.2655 | 31.9662 | 40.4524 | 25.2602 | 1.9922 | 31.8934 | 63.5372 | 16.0093 | 14.0724 | -0.1102 | 103.9896 | 228.1533 | 175.0511 | 225.9326 | 156.036
2012-10-30 17:32 | 27.8901 | 242.2372 | 189.1507 | 240.1853 | 170.1587 | -1.6003 | 1.265 | 31.9532 | 40.4198 | 25.2601 | 1.9913 | 31.8802 | 63.4846 | 16.0093 | 14.036 | -0.1103 | 103.9044 | 228.2012 | 175.1147 | 226.1493 | 156.1227
2012-10-30 17:32 | 27.8938 | 242.3272 | 189.123 | 240.0045 | 170.3961 | -1.5997 | 1.2647 | 31.9455 | 40.4 | 25.2603 | 1.9909 | 31.8724 | 63.454 | 16.0093 | 14.0315 | -0.1103 | 103.854 | 228.2957 | 175.0915 | 225.973 | 156.3646
2012-10-30 17:32 | 27.8968 | 242.2792 | 189.2162 | 240.2508 | 170.2939 | -1.5985 | 1.2645 | 31.9412 | 40.3889 | 25.2603 | 1.9906 | 31.8682 | 63.4366 | 16.0094 | 14.0631 | -0.1102 | 103.8255 | 228.2161 | 175.1531 | 226.1877 | 156.2308
2012-10-30 17:32 | 27.9088 | 242.3697 | 189.2557 | 240.2589 | 170.341 | -1.5968 | 1.2644 | 31.9402 | 40.3865 | 25.2604 | 1.9905 | 31.8671 | 63.4323 | 16.0094 | 14.0995 | -0.1101 | 103.8188 | 228.2702 | 175.1562 | 226.1594 | 156.2415
2012-10-30 17:32 | 27.9028 | 242.4068 | 189.3243 | 240.2019 | 170.4363 | -1.6026 | 1.2644 | 31.9394 | 40.3843 | 25.2604 | 1.9905 | 31.8663 | 63.4283 | 16.0096 | 14.0528 | -0.1105 | 103.8126 | 228.354 | 175.2715 | 226.1491 | 156.3835
2012-10-30 17:32 | 27.9051 | 242.4287 | 189.3697 | 240.3769 | 170.3996 | -1.6061 | 1.2645 | 31.941 | 40.3884 | 25.2605 | 1.9906 | 31.868 | 63.4352 | 16.0095 | 14.0754 | -0.1107 | 103.8236 | 228.3533 | 175.2943 | 226.3015 | 156.3242
2012-10-30 17:32 | 27.8998 | 242.4352 | 189.4076 | 240.3716 | 170.4956 | -1.6032 | 1.2646 | 31.9433 | 40.3941 | 25.2606 | 1.9907 | 31.8702 | 63.4439 | 16.0096 | 14.0538 | -0.1105 | 103.838 | 228.3814 | 175.3538 | 226.3178 | 156.4418
2012-10-30 17:32 | 27.9021 | 242.5081 | 189.4295 | 240.2836 | 170.381 | -1.5968 | 1.2646 | 31.9449 | 40.3979 | 25.2606 | 1.9908 | 31.8718 | 63.45 | 16.0097 | 14.0602 | -0.1101 | 103.8478 | 228.4479 | 175.3693 | 226.2234 | 156.3208
2012-10-30 17:32 | 27.9111 | 242.5367 | 189.4424 | 240.4417 | 170.3667 | -1.5997 | 1.265 | 31.956 | 40.4256 | 25.261 | 1.9913 | 31.8797 | 63.4816 | 16.0096 | 14.0611 | -0.1103 | 103.9071 | 228.4756 | 175.3813 | 226.3806 | 156.3056
2012-10-30 17:32 | 27.9193 | 242.5802 | 189.5644 | 240.5402 | 170.4021 | -1.6055 | 1.2696 | 32.0722 | 40.7177 | 25.2624 | 1.9987 | 31.9985 | 63.9558 | 16.0095 | 14.0693 | -0.1107 | 104.6736 | 228.5109 | 175.4951 | 226.4709 | 156.3328
2012-10-30 17:32 | 27.9088 | 242.6795 | 189.5461 | 240.4237 | 170.6136 | -1.6026 | 1.2701 | 32.0862 | 40.7532 | 25.2625 | 1.9996 | 32.0127 | 64.0122 | 16.0096 | 14.0261 | -0.1105 | 104.7654 | 228.6534 | 175.52 | 226.3976 | 156.5875
2012-10-30 17:32 | 27.9133 | 242.6958 | 189.5977 | 240.6087 | 170.6999 | -1.6026 | 1.2707 | 32.1006 | 40.7893 | 25.2627 | 2.0005 | 32.0271 | 64.0701 | 16.0096 | 14.0429 | -0.1105 | 104.8594 | 228.6529 | 175.5548 | 226.5658 | 156.657
2012-10-30 17:32 | 27.9103 | 242.7202 | 189.6339 | 240.6293 | 170.8411 | -1.6026 | 1.2711 | 32.1121 | 40.8181 | 25.2629 | 2.0012 | 32.0383 | 64.1153 | 16.0095 | 14.0358 | -0.1105 | 104.9334 | 228.6844 | 175.5981 | 226.5935 | 156.8053
2012-10-30 17:32 | 27.9193 | 242.8194 | 189.7018 | 240.5009 | 170.7878 | -1.5968 | 1.2714 | 32.1209 | 40.8401 | 25.2633 | 2.0018 | 32.0471 | 64.1508 | 16.0094 | 14.0244 | -0.1101 | 104.9909 | 228.795 | 175.6774 | 226.4765 | 156.7634
2012-10-30 17:32 | 27.9133 | 242.8762 | 189.735 | 240.7108 | 170.7077 | -1.6009 | 1.2717 | 32.1273 | 40.8558 | 25.2636 | 2.0022 | 32.0537 | 64.1768 | 16.0095 | 14.0184 | -0.1104 | 105.0326 | 228.8578 | 175.7166 | 226.6924 | 156.6893
2012-10-30 17:32 | 27.9163 | 242.887 | 189.8047 | 240.6588 | 170.8354 | -1.6003 | 1.2718 | 32.1311 | 40.865 | 25.2639 | 2.0024 | 32.0575 | 64.1924 | 16.0095 | 14.0051 | -0.1103 | 105.0574 | 228.8819 | 175.7996 | 226.6537 | 156.8303
2012-10-30 17:32 | 27.929 | 242.935 | 189.8449 | 240.6715 | 170.7273 | -1.6032 | 1.2719 | 32.1341 | 40.8727 | 25.2637 | 2.0026 | 32.0605 | 64.2043 | 16.0094 | 14.0504 | -0.1105 | 105.0771 | 228.8846 | 175.7945 | 226.6211 | 156.6769
2012-10-30 17:32 | 27.9283 | 242.9775 | 189.8638 | 240.8631 | 170.711 | -1.6113 | 1.272 | 32.1358 | 40.8771 | 25.2638 | 2.0027 | 32.0621 | 64.2113 | 16.0093 | 14.0497 | -0.1111 | 105.0884 | 228.9278 | 175.8141 | 226.8134 | 156.6613
2012-10-30 17:32 | 27.9238 | 242.9847 | 189.9103 | 240.839 | 170.9403 | -1.5985 | 1.2721 | 32.1372 | 40.8802 | 25.2641 | 2.0028 | 32.0635 | 64.2169 | 16.0093 | 14.0289 | -0.1102 | 105.0972 | 228.9558 | 175.8814 | 226.8101 | 156.9114
2012-10-30 17:32 | 27.9238 | 243.0788 | 189.9614 | 240.7919 | 170.8312 | -1.5974 | 1.2721 | 32.1383 | 40.8826 | 25.2644 | 2.0029 | 32.0646 | 64.2209 | 16.0094 | 14.0044 | -0.1101 | 105.1034 | 229.0744 | 175.957 | 226.7875 | 156.8268
2012-10-30 17:32 | 27.9298 | 243.1397 | 189.9909 | 240.9941 | 170.7398 | -1.5945 | 1.2721 | 32.139 | 40.8842 | 25.2644 | 2.0029 | 32.0651 | 64.2236 | 16.0093 | 14.0104 | -0.1099 | 105.1078 | 229.1293 | 175.9805 | 226.9837 | 156.7294
2012-10-30 17:32 | 27.9275 | 243.1374 | 190.0239 | 240.9408 | 170.9946 | -1.5991 | 1.2721 | 32.1394 | 40.885 | 25.2645 | 2.0029 | 32.0655 | 64.2249 | 16.0094 | 14 | -0.1102 | 105.1098 | 229.1374 | 176.0239 | 226.9408 | 156.9946
2012-10-30 17:32 | 27.9395 | 243.2004 | 190.0595 | 240.8782 | 170.9638 | -1.598 | 1.2721 | 32.1397 | 40.8852 | 25.2649 | 2.003 | 32.0661 | 64.2274 | 16.0092 | 14.0283 | -0.1102 | 105.1127 | 229.1721 | 176.0312 | 226.8499 | 156.9355
2012-10-30 17:33 | 27.941 | 243.2294 | 190.057 | 241.0602 | 171.0315 | -1.6009 | 1.2721 | 32.1403 | 40.8868 | 25.2649 | 2.003 | 32.0666 | 64.2294 | 16.0093 | 13.9727 | -0.1104 | 105.1162 | 229.2567 | 176.0843 | 227.0875 | 157.0588
2012-10-30 17:33 | 27.9418 | 243.289 | 190.1442 | 241.1316 | 171.157 | -1.6026 | 1.2721 | 32.1407 | 40.8876 | 25.265 | 2.003 | 32.067 | 64.2314 | 16.0092 | 13.9776 | -0.1105 | 105.119 | 229.3114 | 176.1666 | 227.154 | 157.1794
2012-10-30 17:33 | 27.9515 | 243.3614 | 190.2049 | 241.0119 | 171.0966 | -1.5991 | 1.2722 | 32.1409 | 40.8886 | 25.2646 | 2.003 | 32.0672 | 64.2318 | 16.0093 | 13.9709 | -0.1102 | 105.1204 | 229.3905 | 176.234 | 227.041 | 157.1257
2012-10-30 17:33 | 27.956 | 243.4404 | 190.2133 | 241.2125 | 170.9452 | -1.5968 | 1.2722 | 32.1412 | 40.8891 | 25.2648 | 2.0031 | 32.0674 | 64.2338 | 16.009 | 13.9958 | -0.1101 | 105.1229 | 229.4446 | 176.2175 | 227.2167 | 156.9494
2012-10-30 17:33 | 27.9545 | 243.4389 | 190.2315 | 241.2738 | 171.1151 | -1.6084 | 1.2722 | 32.1413 | 40.8886 | 25.2653 | 2.0031 | 32.0675 | 64.2341 | 16.009 | 13.9576 | -0.1109 | 105.1227 | 229.4813 | 176.2739 | 227.3162 | 157.1575
2012-10-30 17:33 | 27.9695 | 243.4539 | 190.2936 | 241.2339 | 171.1224 | -1.602 | 1.2722 | 32.1413 | 40.889 | 25.2651 | 2.0031 | 32.0676 | 64.2338 | 16.0092 | 14.046 | -0.1104 | 105.1228 | 229.4079 | 176.2476 | 227.1879 | 157.0764
2012-10-30 17:33 | 27.9612 | 243.5476 | 190.3245 | 241.3119 | 171.0751 | -1.5968 | 1.2722 | 32.1418 | 40.8901 | 25.2652 | 2.0031 | 32.0681 | 64.2364 | 16.0091 | 13.9888 | -0.1101 | 105.1265 | 229.5588 | 176.3357 | 227.3231 | 157.0863
2012-10-30 17:33 | 27.9672 | 243.5653 | 190.303 | 241.4238 | 171.159 | -1.6009 | 1.2722 | 32.142 | 40.8902 | 25.2655 | 2.0031 | 32.0682 | 64.2362 | 16.0092 | 13.9989 | -0.1104 | 105.1264 | 229.5664 | 176.3041 | 227.4249 | 157.1601
2012-10-30 17:33 | 27.965 | 243.6023 | 190.3676 | 241.3471 | 171.0945 | -1.6026 | 1.2721 | 32.1417 | 40.889 | 25.2656 | 2.0031 | 32.0679 | 64.2352 | 16.0092 | 14.0171 | -0.1105 | 105.1242 | 229.5852 | 176.3505 | 227.33 | 157.0774
2012-10-30 17:33 | 27.9687 | 243.6962 | 190.332 | 241.492 | 171.0164 | -1.5945 | 1.2721 | 32.1417 | 40.8884 | 25.266 | 2.0031 | 32.0681 | 64.2363 | 16.009 | 14.0126 | -0.1099 | 105.1247 | 229.6836 | 176.3194 | 227.4794 | 157.0038
2012-10-30 17:33 | 27.9725 | 243.7392 | 190.3829 | 241.5625 | 171.1604 | -1.6003 | 1.2721 | 32.1416 | 40.8888 | 25.2656 | 2.0031 | 32.0679 | 64.2347 | 16.0092 | 14.0042 | -0.1103 | 105.1234 | 229.735 | 176.3787 | 227.5583 | 157.1562
2012-10-30 17:33 | 27.9784 | 243.7687 | 190.4163 | 241.5056 | 171.2793 | -1.598 | 1.2721 | 32.1416 | 40.8885 | 25.2658 | 2.0031 | 32.0678 | 64.2358 | 16.0089 | 14.0346 | -0.1102 | 105.1243 | 229.7341 | 176.3817 | 227.471 | 157.2447
2012-10-30 17:33 | 27.9837 | 243.7896 | 190.4726 | 241.507 | 171.3236 | -1.6009 | 1.2721 | 32.1415 | 40.8882 | 25.2659 | 2.0031 | 32.068 | 64.2366 | 16.0089 | 14.0439 | -0.1104 | 105.1248 | 229.7457 | 176.4287 | 227.4631 | 157.2797
2012-10-30 17:33 | 27.9799 | 243.8721 | 190.4727 | 241.719 | 171.2224 | -1.602 | 1.2721 | 32.1416 | 40.8879 | 25.2662 | 2.0031 | 32.0679 | 64.2362 | 16.0089 | 14.0646 | -0.1104 | 105.1241 | 229.8075 | 176.4081 | 227.6544 | 157.1578
2012-10-30 17:33 | 27.9897 | 243.8937 | 190.5139 | 241.6896 | 171.3179 | -1.6003 | 1.2721 | 32.1415 | 40.8882 | 25.2659 | 2.0031 | 32.0678 | 64.2355 | 16.0089 | 14.0948 | -0.1103 | 105.1238 | 229.7989 | 176.4191 | 227.5948 | 157.2231
2012-10-30 17:33 | 27.9829 | 243.9496 | 190.5111 | 241.6396 | 171.3072 | -1.6026 | 1.2721 | 32.1412 | 40.8871 | 25.2661 | 2.0031 | 32.0675 | 64.2349 | 16.0088 | 14.0431 | -0.1105 | 105.122 | 229.9065 | 176.468 | 227.5965 | 157.2641
2012-10-30 17:33 | 27.9859 | 243.984 | 190.5494 | 241.827 | 171.2517 | -1.5957 | 1.2721 | 32.141 | 40.8863 | 25.2662 | 2.0031 | 32.0671 | 64.2336 | 16.0088 | 14.0217 | -0.11 | 105.1199 | 229.9623 | 176.5277 | 227.8053 | 157.23
2012-10-30 17:33 | 27.9949 | 244.0165 | 190.6016 | 241.8321 | 171.6192 | -1.5997 | 1.2721 | 32.1409 | 40.8857 | 25.2665 | 2.0031 | 32.0671 | 64.2334 | 16.0088 | 14.0551 | -0.1103 | 105.1191 | 229.9614 | 176.5465 | 227.777 | 157.5641
2012-10-30 17:33 | 27.9889 | 244.0497 | 190.6152 | 241.83 | 171.5548 | -1.602 | 1.2721 | 32.1408 | 40.8852 | 25.2667 | 2.0031 | 32.0671 | 64.2336 | 16.0087 | 14.0083 | -0.1104 | 105.1189 | 230.0414 | 176.6069 | 227.8217 | 157.5465
2012-10-30 17:33 | 28.0002 | 244.108 | 190.6775 | 241.9433 | 171.6479 | -1.602 | 1.272 | 32.1406 | 40.884 | 25.267 | 2.0031 | 32.0669 | 64.2326 | 16.0088 | 14.0196 | -0.1104 | 105.1166 | 230.0884 | 176.6579 | 227.9237 | 157.6283
2012-10-30 17:33 | 28.0009 | 244.144 | 190.7528 | 242.0303 | 171.446 | -1.6049 | 1.2721 | 32.1407 | 40.8851 | 25.2666 | 2.0031 | 32.067 | 64.2325 | 16.0089 | 14.0367 | -0.1106 | 105.1176 | 230.1073 | 176.7161 | 227.9936 | 157.4093
2012-10-30 17:33 | 28.0047 | 244.1792 | 190.7605 | 241.9046 | 171.4965 | -1.6014 | 1.2721 | 32.1408 | 40.8849 | 25.2668 | 2.0031 | 32.0671 | 64.234 | 16.0087 | 14.0119 | -0.1104 | 105.1189 | 230.1673 | 176.7486 | 227.8927 | 157.4846
2012-10-30 17:33 | 28.0226 | 244.2402 | 190.798 | 242.0638 | 171.495 | -1.6089 | 1.272 | 32.1402 | 40.8831 | 25.267 | 2.0031 | 32.0666 | 64.2317 | 16.0087 | 14.0217 | -0.1109 | 105.1148 | 230.2185 | 176.7763 | 228.0421 | 157.4733
2012-10-30 17:33 | 28.0159 | 244.2884 | 190.8463 | 242.1512 | 171.5701 | -1.5974 | 1.272 | 32.14 | 40.8823 | 25.2671 | 2.0031 | 32.0662 | 64.2304 | 16.0086 | 14.0517 | -0.1101 | 105.1127 | 230.2367 | 176.7946 | 228.0995 | 157.5184
2012-10-30 17:33 | 28.0241 | 244.3476 | 190.8388 | 242.1594 | 171.512 | -1.6037 | 1.272 | 32.1399 | 40.8819 | 25.2672 | 2.0031 | 32.0661 | 64.2301 | 16.0086 | 14.0313 | -0.1106 | 105.112 | 230.3163 | 176.8075 | 228.1281 | 157.4807
2012-10-30 17:33 | 28.0166 | 244.3557 | 190.8588 | 242.0656 | 171.6098 | -1.6026 | 1.272 | 32.1397 | 40.8814 | 25.2673 | 2.003 | 32.0659 | 64.2286 | 16.0088 | 14.0442 | -0.1105 | 105.11 | 230.3115 | 176.8146 | 228.0214 | 157.5656
2012-10-30 17:34 | 28.0234 | 244.4213 | 190.8813 | 242.2058 | 171.5152 | -1.6009 | 1.272 | 32.1395 | 40.8802 | 25.2676 | 2.003 | 32.0657 | 64.2282 | 16.0087 | 14.0877 | -0.1104 | 105.1084 | 230.3336 | 176.7936 | 228.1181 | 157.4275
2012-10-30 17:34 | 28.0234 | 244.4331 | 190.8931 | 242.3234 | 171.3711 | -1.6026 | 1.272 | 32.1392 | 40.8797 | 25.2675 | 2.003 | 32.0654 | 64.2272 | 16.0087 | 14.0388 | -0.1105 | 105.1069 | 230.3943 | 176.8543 | 228.2846 | 157.3323
2012-10-30 17:34 | 28.0166 | 244.5125 | 190.8392 | 242.2735 | 171.3837 | -1.5997 | 1.2719 | 32.139 | 40.8789 | 25.2678 | 2.003 | 32.0653 | 64.2267 | 16.0087 | 14.0524 | -0.1103 | 105.1056 | 230.4601 | 176.7868 | 228.2211 | 157.3313
2012-10-30 17:34 | 28.0211 | 244.5406 | 190.8986 | 242.2701 | 171.5129 | -1.6049 | 1.2719 | 32.1384 | 40.8771 | 25.2679 | 2.0029 | 32.0647 | 64.224 | 16.0087 | 14.0691 | -0.1106 | 105.1011 | 230.4715 | 176.8295 | 228.201 | 157.4438
2012-10-30 17:34 | 28.0249 | 244.5992 | 190.9495 | 242.3328 | 171.7038 | -1.6003 | 1.2719 | 32.138 | 40.8761 | 25.2679 | 2.0029 | 32.0644 | 64.2231 | 16.0086 | 14.0729 | -0.1103 | 105.0992 | 230.5263 | 176.8766 | 228.2599 | 157.6309
2012-10-30 17:34 | 28.0099 | 244.5921 | 190.9541 | 242.5061 | 171.5953 | -1.6032 | 1.2719 | 32.1379 | 40.8756 | 25.2679 | 2.0029 | 32.0641 | 64.2219 | 16.0086 | 14.0416 | -0.1105 | 105.0975 | 230.5505 | 176.9125 | 228.4645 | 157.5537
2012-10-30 17:34 | 28.0151 | 244.6522 | 190.9947 | 242.4289 | 171.6784 | -1.6014 | 1.2719 | 32.1374 | 40.8744 | 25.268 | 2.0029 | 32.0637 | 64.2206 | 16.0086 | 14.0509 | -0.1104 | 105.095 | 230.6013 | 176.9438 | 228.378 | 157.6275
2012-10-30 17:34 | 28.0286 | 244.6617 | 191.0396 | 242.4071 | 171.6568 | -1.5991 | 1.2719 | 32.1372 | 40.8739 | 25.2679 | 2.0029 | 32.0634 | 64.22 | 16.0085 | 14.1296 | -0.1102 | 105.0939 | 230.5321 | 176.91 | 228.2775 | 157.5272
2012-10-30 17:34 | 28.0189 | 244.6952 | 191.0024 | 242.519 | 171.6666 | -1.602 | 1.2718 | 32.1372 | 40.8735 | 25.2681 | 2.0029 | 32.0634 | 64.2195 | 16.0086 | 14.0669 | -0.1104 | 105.093 | 230.6283 | 176.9355 | 228.4521 | 157.5997
2012-10-30 17:34 | 28.0204 | 244.6927 | 191.0549 | 242.5675 | 171.8006 | -1.6009 | 1.2718 | 32.1364 | 40.8717 | 25.2682 | 2.0029 | 32.0627 | 64.2174 | 16.0084 | 14.0725 | -0.1104 | 105.089 | 230.6202 | 176.9824 | 228.495 | 157.7281
2012-10-30 17:34 | 28.0294 | 244.788 | 191.0993 | 242.5844 | 171.9655 | -1.6014 | 1.2718 | 32.1358 | 40.8694 | 25.2685 | 2.0028 | 32.0621 | 64.2147 | 16.0085 | 14.0488 | -0.1104 | 105.0841 | 230.7392 | 177.0505 | 228.5356 | 157.9167
2012-10-30 17:34 | 28.0271 | 244.7739 | 191.1833 | 242.4919 | 172.0295 | -1.6009 | 1.2718 | 32.1355 | 40.8687 | 25.2685 | 2.0028 | 32.0618 | 64.2127 | 16.0086 | 14.0465 | -0.1104 | 105.0815 | 230.7274 | 177.1368 | 228.4454 | 157.983
2012-10-30 17:34 | 28.0271 | 244.7974 | 191.1912 | 242.688 | 172.0295 | -1.6009 | 1.2717 | 32.135 | 40.8673 | 25.2686 | 2.0028 | 32.0613 | 64.2109 | 16.0086 | 14.0914 | -0.1104 | 105.0782 | 230.706 | 177.0998 | 228.5966 | 157.9381
2012-10-30 17:34 | 28.0406 | 244.8345 | 191.2636 | 242.7093 | 172.0547 | -1.6043 | 1.2717 | 32.1349 | 40.8673 | 25.2684 | 2.0028 | 32.0612 | 64.2109 | 16.0085 | 14.0315 | -0.1106 | 105.0782 | 230.803 | 177.2321 | 228.6778 | 158.0232
2012-10-30 17:34 | 28.0286 | 244.8969 | 191.2712 | 242.6738 | 171.7621 | -1.6014 | 1.2717 | 32.1347 | 40.8663 | 25.2687 | 2.0028 | 32.061 | 64.2103 | 16.0084 | 14.0766 | -0.1104 | 105.0766 | 230.8203 | 177.1946 | 228.5972 | 157.6855
2012-10-30 17:34 | 28.0331 | 244.9132 | 191.3071 | 242.6744 | 171.8055 | -1.6009 | 1.2717 | 32.1344 | 40.8653 | 25.2689 | 2.0027 | 32.0607 | 64.2092 | 16.0084 | 14.0199 | -0.1104 | 105.0746 | 230.8933 | 177.2872 | 228.6545 | 157.7856
2012-10-30 17:34 | 28.0354 | 244.9155 | 191.2584 | 242.7708 | 171.8546 | -1.6026 | 1.2717 | 32.1337 | 40.8635 | 25.2689 | 2.0027 | 32.06 | 64.2067 | 16.0084 | 14.0671 | -0.1105 | 105.0703 | 230.8484 | 177.1913 | 228.7037 | 157.7875
2012-10-30 17:34 | 28.0346 | 244.9421 | 191.3047 | 242.8406 | 172.0643 | -1.5991 | 1.2717 | 32.1336 | 40.8632 | 25.2689 | 2.0027 | 32.0599 | 64.2065 | 16.0083 | 14.0255 | -0.1102 | 105.0697 | 230.9166 | 177.2792 | 228.8151 | 158.0388
2012-10-30 17:34 | 28.0451 | 245.0467 | 191.3348 | 242.8393 | 172.1059 | -1.6061 | 1.2717 | 32.1335 | 40.8627 | 25.269 | 2.0027 | 32.0598 | 64.2058 | 16.0083 | 14.0564 | -0.1107 | 105.0686 | 230.9903 | 177.2784 | 228.7829 | 158.0495
2012-10-30 17:34 | 28.0414 | 245.0195 | 191.3429 | 242.8199 | 172.0048 | -1.6043 | 1.2716 | 32.1327 | 40.8606 | 25.2691 | 2.0026 | 32.0589 | 64.2025 | 16.0083 | 14.0486 | -0.1106 | 105.0631 | 230.9709 | 177.2943 | 228.7713 | 157.9562
2012-10-30 17:34 | 28.0391 | 245.1269 | 191.3798 | 242.8529 | 172.1038 | -1.5968 | 1.27 | 32.091 | 40.7552 | 25.2688 | 2.0003 | 32.022 | 64.0545 | 16.0084 | 14.0341 | -0.1101 | 104.8097 | 231.0928 | 177.3457 | 228.8188 | 158.0697
2012-10-30 17:34 | 28.0391 | 245.1308 | 191.3838 | 242.9549 | 172.1662 | -1.6032 | 1.2653 | 31.9721 | 40.4551 | 25.2678 | 1.9926 | 31.8982 | 63.5604 | 16.0083 | 14.0626 | -0.1105 | 104.0156 | 231.0682 | 177.3212 | 228.8923 | 158.1036
2012-10-30 17:34 | 28.0496 | 245.1061 | 191.4139 | 243.0399 | 172.243 | -1.6049 | 1.2649 | 31.9615 | 40.4288 | 25.2675 | 1.992 | 31.8883 | 63.5207 | 16.0084 | 14.0854 | -0.1106 | 103.9495 | 231.0207 | 177.3285 | 228.9545 | 158.1576
2012-10-30 17:34 | 28.0414 | 245.0469 | 191.4646 | 243.02 | 172.6558 | -1.5985 | 1.2645 | 31.9508 | 40.4016 | 25.2676 | 1.9913 | 31.8774 | 63.4771 | 16.0084 | 14.0568 | -0.1102 | 103.8787 | 230.9901 | 177.4078 | 228.9632 | 158.599
2012-10-30 17:34 | 28.0481 | 245.132 | 191.5459 | 242.96 | 172.6703 | -1.602 | 1.2643 | 31.9448 | 40.3864 | 25.2677 | 1.9909 | 31.8715 | 63.453 | 16.0086 | 14.0268 | -0.1104 | 103.8394 | 231.1052 | 177.5191 | 228.9332 | 158.6435
2012-10-30 17:34 | 28.0466 | 245.1383 | 191.5719 | 243.033 | 172.1698 | -1.5951 | 1.2642 | 31.9436 | 40.3834 | 25.2676 | 1.9908 | 31.8702 | 63.4474 | 16.0087 | 14.0416 | -0.11 | 103.8308 | 231.0967 | 177.5303 | 228.9914 | 158.1282
2012-10-30 17:35 | 28.0473 | 245.1743 | 191.549 | 243.0612 | 172.1666 | -1.5985 | 1.264 | 31.9396 | 40.3731 | 25.2677 | 1.9906 | 31.8662 | 63.4315 | 16.0086 | 14.1035 | -0.1102 | 103.8046 | 231.0708 | 177.4455 | 228.9577 | 158.0631
2012-10-30 17:35 | 28.0421 | 245.1809 | 191.5988 | 243.1383 | 171.8457 | -1.5991 | 1.2641 | 31.941 | 40.3763 | 25.268 | 1.9907 | 31.8677 | 63.4376 | 16.0087 | 14.0615 | -0.1102 | 103.8139 | 231.1194 | 177.5373 | 229.0768 | 157.7842
2012-10-30 17:35 | 28.0346 | 245.1616 | 191.5442 | 243.0877 | 171.7018 | -1.5991 | 1.2642 | 31.9439 | 40.3838 | 25.2679 | 1.9908 | 31.8706 | 63.448 | 16.0089 | 14.1071 | -0.1102 | 103.8318 | 231.0545 | 177.4371 | 228.9806 | 157.5947
2012-10-30 17:35 | 28.0399 | 245.2375 | 191.5416 | 243.0851 | 171.6409 | -1.5968 | 1.2643 | 31.9457 | 40.3878 | 25.2681 | 1.9909 | 31.8722 | 63.4549 | 16.0088 | 14.0838 | -0.1101 | 103.8428 | 231.1537 | 177.4578 | 229.0013 | 157.5571
2012-10-30 17:35 | 28.0301 | 245.2473 | 191.4847 | 243.2008 | 171.8337 | -1.598 | 1.2654 | 31.9757 | 40.4636 | 25.2682 | 1.9928 | 31.9017 | 63.5721 | 16.0089 | 14.1189 | -0.1102 | 104.0357 | 231.1284 | 177.3658 | 229.0819 | 157.7148
2012-10-30 17:35 | 28.0339 | 245.2432 | 191.5513 | 243.2164 | 171.7985 | -1.6014 | 1.2692 | 32.0735 | 40.7091 | 25.2697 | 1.9989 | 32.0001 | 63.9652 | 16.0088 | 14.1023 | -0.1104 | 104.6744 | 231.1409 | 177.449 | 229.1141 | 157.6962
2012-10-30 17:35 | 28.0271 | 245.2874 | 191.5681 | 243.1547 | 171.8463 | -1.5997 | 1.2699 | 32.0901 | 40.7513 | 25.2697 | 2 | 32.0167 | 64.032 | 16.0087 | 14.0996 | -0.1103 | 104.7833 | 231.1878 | 177.4685 | 229.0551 | 157.7467
2012-10-30 17:35 | 28.0279 | 245.2764 | 191.5021 | 243.1359 | 172.0108 | -1.602 | 1.2705 | 32.1043 | 40.7869 | 25.27 | 2.0008 | 32.0308 | 64.0884 | 16.0087 | 14.1452 | -0.1104 | 104.8753 | 231.1312 | 177.3569 | 228.9907 | 157.8656
2012-10-30 17:35 | 28.0256 | 245.3094 | 191.5548 | 243.2708 | 172.2073 | -1.6055 | 1.2709 | 32.1151 | 40.814 | 25.2702 | 2.0015 | 32.0413 | 64.1306 | 16.0087 | 14.1592 | -0.1107 | 104.9446 | 231.1502 | 177.3956 | 229.1116 | 158.0481
2012-10-30 17:35 | 28.0286 | 245.3751 | 191.6481 | 243.2817 | 172.4091 | -1.6066 | 1.2712 | 32.1231 | 40.8347 | 25.2701 | 2.002 | 32.0493 | 64.163 | 16.0086 | 14.1174 | -0.1108 | 104.9977 | 231.2577 | 177.5307 | 229.1643 | 158.2917
2012-10-30 17:35 | 28.0211 | 245.3676 | 191.7113 | 243.282 | 172.5419 | -1.6061 | 1.2714 | 32.128 | 40.8461 | 25.2707 | 2.0023 | 32.0542 | 64.1827 | 16.0086 | 14.1099 | -0.1107 | 105.0287 | 231.2577 | 177.6014 | 229.1721 | 158.432
2012-10-30 17:35 | 28.0249 | 245.438 | 191.719 | 243.1956 | 172.7641 | -1.602 | 1.2715 | 32.1305 | 40.8529 | 25.2703 | 2.0025 | 32.0568 | 64.1929 | 16.0086 | 14.0647 | -0.1104 | 105.0458 | 231.3733 | 177.6543 | 229.1309 | 158.6994
2012-10-30 17:35 | 28.0256 | 245.474 | 191.8061 | 243.3061 | 172.5854 | -1.5957 | 1.2715 | 32.1324 | 40.8573 | 25.2706 | 2.0026 | 32.0586 | 64.2005 | 16.0085 | 14.0818 | -0.11 | 105.0578 | 231.3922 | 177.7243 | 229.2243 | 158.5036
2012-10-30 17:35 | 28.0264 | 245.4552 | 191.8344 | 243.3383 | 172.3952 | -1.5962 | 1.2716 | 32.134 | 40.8609 | 25.2709 | 2.0027 | 32.0601 | 64.2071 | 16.0084 | 14.107 | -0.11 | 105.068 | 231.3482 | 177.7274 | 229.2313 | 158.2882
2012-10-30 17:35 | 28.0316 | 245.5427 | 191.8356 | 243.3043 | 172.4745 | -1.6003 | 1.2716 | 32.1352 | 40.8642 | 25.2709 | 2.0028 | 32.0615 | 64.2123 | 16.0084 | 14.0347 | -0.1103 | 105.0764 | 231.508 | 177.8009 | 229.2696 | 158.4398
2012-10-30 17:35 | 28.0301 | 245.5216 | 191.8263 | 243.291 | 172.434 | -1.5962 | 1.2717 | 32.1361 | 40.8666 | 25.2707 | 2.0028 | 32.0622 | 64.2146 | 16.0086 | 14.0781 | -0.11 | 105.0812 | 231.4435 | 177.7482 | 229.2129 | 158.3559
2012-10-30 17:35 | 28.0369 | 245.5323 | 191.8252 | 243.4586 | 172.5149 | -1.598 | 1.2717 | 32.1366 | 40.8676 | 25.2709 | 2.0029 | 32.0627 | 64.2177 | 16.0084 | 14.0767 | -0.1102 | 105.0852 | 231.4556 | 177.7485 | 229.3819 | 158.4382
2012-10-30 17:35 | 28.0421 | 245.5924 | 191.9168 | 243.4246 | 172.4772 | -1.6009 | 1.2717 | 32.1371 | 40.8683 | 25.2712 | 2.0029 | 32.0632 | 64.2193 | 16.0084 | 14.0615 | -0.1104 | 105.0876 | 231.5309 | 177.8553 | 229.3631 | 158.4157
2012-10-30 17:35 | 28.0309 | 245.6165 | 191.8978 | 243.4369 | 172.4426 | -1.6026 | 1.2717 | 32.1379 | 40.8704 | 25.2712 | 2.003 | 32.064 | 64.2229 | 16.0083 | 14.0871 | -0.1105 | 105.0933 | 231.5294 | 177.8107 | 229.3498 | 158.3555
2012-10-30 17:35 | 28.0406 | 245.5987 | 191.9153 | 243.5015 | 172.5069 | -1.5962 | 1.2717 | 32.1384 | 40.8715 | 25.2714 | 2.003 | 32.0645 | 64.2253 | 16.0082 | 14.0764 | -0.11 | 105.0969 | 231.5223 | 177.8389 | 229.4251 | 158.4305
2012-10-30 17:35 | 28.0443 | 245.6769 | 191.9347 | 243.4934 | 172.495 | -1.5974 | 1.2717 | 32.1384 | 40.8714 | 25.2714 | 2.003 | 32.0645 | 64.2249 | 16.0083 | 14.1127 | -0.1101 | 105.0963 | 231.5642 | 177.822 | 229.3807 | 158.3823
2012-10-30 17:35 | 28.0391 | 245.6913 | 191.9295 | 243.449 | 172.5872 | -1.6032 | 1.2717 | 32.1383 | 40.8709 | 25.2715 | 2.003 | 32.0644 | 64.2248 | 16.0083 | 14.0871 | -0.1105 | 105.0956 | 231.6042 | 177.8424 | 229.3619 | 158.5001
2012-10-30 17:35 | 28.0518 | 245.7392 | 191.9304 | 243.646 | 172.5531 | -1.5991 | 1.2717 | 32.1384 | 40.8713 | 25.2715 | 2.003 | 32.0646 | 64.2255 | 16.0083 | 14.059 | -0.1102 | 105.0969 | 231.6802 | 177.8714 | 229.587 | 158.4941
2012-10-30 17:35 | 28.0526 | 245.7557 | 191.9862 | 243.6037 | 172.7294 | -1.6014 | 1.2717 | 32.1385 | 40.8715 | 25.2716 | 2.003 | 32.0647 | 64.2263 | 16.0082 | 14.072 | -0.1104 | 105.0977 | 231.6837 | 177.9142 | 229.5317 | 158.6574
2012-10-30 17:35 | 28.0496 | 245.7919 | 192.0264 | 243.5811 | 172.742 | -1.598 | 1.2717 | 32.1384 | 40.8711 | 25.2715 | 2.003 | 32.0646 | 64.2258 | 16.0082 | 14.065 | -0.1102 | 105.0969 | 231.7269 | 177.9614 | 229.5161 | 158.677
2012-10-30 17:36 | 28.0496 | 245.8076 | 192.0539 | 243.6164 | 172.7654 | -1.5991 | 1.2717 | 32.1386 | 40.8715 | 25.2716 | 2.003 | 32.0646 | 64.2267 | 16.008 | 14.065 | -0.1102 | 105.0982 | 231.7426 | 177.9889 | 229.5514 | 158.7004
2012-10-30 17:36 | 28.0556 | 245.841 | 192.0834 | 243.7282 | 172.6661 | -1.5985 | 1.2717 | 32.1389 | 40.8723 | 25.2715 | 2.0031 | 32.065 | 64.2279 | 16.008 | 14.0587 | -0.1102 | 105.1002 | 231.7823 | 178.0247 | 229.6695 | 158.6074
2012-10-30 17:36 | 28.0556 | 245.8253 | 192.1188 | 243.7204 | 172.6622 | -1.6026 | 1.2717 | 32.1387 | 40.8717 | 25.2717 | 2.003 | 32.0648 | 64.2273 | 16.008 | 14.071 | -0.1105 | 105.099 | 231.7543 | 178.0478 | 229.6494 | 158.5912
2012-10-30 17:36 | 28.0541 | 245.8748 | 192.1134 | 243.7816 | 172.7036 | -1.6014 | 1.2717 | 32.1382 | 40.8705 | 25.2717 | 2.003 | 32.0644 | 64.2255 | 16.008 | 14.0368 | -0.1104 | 105.096 | 231.838 | 178.0766 | 229.7448 | 158.6668
2012-10-30 17:36 | 28.0518 | 245.9078 | 192.1385 | 243.7048 | 172.8027 | -1.6009 | 1.2717 | 32.1382 | 40.87 | 25.2719 | 2.003 | 32.0643 | 64.2253 | 16.008 | 14.0549 | -0.1104 | 105.0953 | 231.8529 | 178.0836 | 229.6499 | 158.7478
2012-10-30 17:36 | 28.0548 | 245.9382 | 192.1258 | 243.7588 | 172.7316 | -1.6003 | 1.2717 | 32.1384 | 40.8705 | 25.272 | 2.003 | 32.0646 | 64.2264 | 16.008 | 14.0171 | -0.1103 | 105.0969 | 231.9211 | 178.1087 | 229.7417 | 158.7145
2012-10-30 17:36 | 28.0563 | 245.9671 | 192.1548 | 243.8426 | 172.6473 | -1.5985 | 1.2717 | 32.1382 | 40.87 | 25.2719 | 2.003 | 32.0643 | 64.2257 | 16.0079 | 14.039 | -0.1102 | 105.0958 | 231.9281 | 178.1158 | 229.8036 | 158.6083
2012-10-30 17:36 | 28.0653 | 245.9761 | 192.1481 | 243.883 | 172.6563 | -1.5968 | 1.2717 | 32.1377 | 40.8688 | 25.2719 | 2.003 | 32.0638 | 64.2239 | 16.0079 | 14.0644 | -0.1101 | 105.0927 | 231.9117 | 178.0837 | 229.8186 | 158.5919
2012-10-30 17:36 | 28.0706 | 245.9697 | 192.1652 | 243.8491 | 172.7669 | -1.6043 | 1.2716 | 32.1374 | 40.8672 | 25.2724 | 2.003 | 32.0634 | 64.2226 | 16.0078 | 14.0452 | -0.1106 | 105.0898 | 231.9245 | 178.12 | 229.8039 | 158.7217
2012-10-30 17:36 | 28.0638 | 246.0373 | 192.1623 | 243.8501 | 172.6587 | -1.6043 | 1.2716 | 32.1373 | 40.8673 | 25.2722 | 2.003 | 32.0634 | 64.2226 | 16.0078 | 14.0588 | -0.1106 | 105.0899 | 231.9785 | 178.1035 | 229.7913 | 158.5999
2012-10-30 17:36 | 28.0706 | 246.001 | 192.1495 | 243.9746 | 172.6226 | -1.6003 | 1.2716 | 32.1372 | 40.8667 | 25.2724 | 2.003 | 32.0633 | 64.2222 | 16.0078 | 14.0574 | -0.1103 | 105.0889 | 231.9436 | 178.0921 | 229.9172 | 158.5652
2012-10-30 17:36 | 28.0728 | 246.0502 | 192.1517 | 243.9532 | 172.6833 | -1.6037 | 1.2716 | 32.137 | 40.8662 | 25.2723 | 2.003 | 32.0631 | 64.2214 | 16.0078 | 14.0596 | -0.1106 | 105.0876 | 231.9906 | 178.0921 | 229.8936 | 158.6237
2012-10-30 17:36 | 28.0765 | 246.097 | 192.2025 | 243.8393 | 172.6402 | -1.602 | 1.2716 | 32.1368 | 40.8658 | 25.2723 | 2.003 | 32.0629 | 64.2205 | 16.0078 | 14.047 | -0.1104 | 105.0863 | 232.05 | 178.1555 | 229.7923 | 158.5932
2012-10-30 17:36 | 28.078 | 246.1142 | 192.1608 | 243.9702 | 172.6456 | -1.6009 | 1.2716 | 32.1364 | 40.8643 | 25.2726 | 2.0029 | 32.0625 | 64.2193 | 16.0077 | 14.0567 | -0.1104 | 105.0836 | 232.0575 | 178.1041 | 229.9135 | 158.5889
2012-10-30 17:36 | 28.0751 | 246.1191 | 192.1933 | 244.0222 | 172.7285 | -1.6003 | 1.2716 | 32.1362 | 40.8643 | 25.2723 | 2.0029 | 32.0624 | 64.2187 | 16.0078 | 14.0538 | -0.1103 | 105.083 | 232.0653 | 178.1395 | 229.9684 | 158.6747
2012-10-30 17:36 | 28.0788 | 246.1699 | 192.197 | 244.1004 | 172.8024 | -1.6026 | 1.2716 | 32.1362 | 40.8636 | 25.2727 | 2.0029 | 32.0623 | 64.2182 | 16.0077 | 14.0371 | -0.1105 | 105.0818 | 232.1328 | 178.1599 | 230.0633 | 158.7653
2012-10-30 17:36 | 28.09 | 246.1811 | 192.2317 | 244.0606 | 172.8955 | -1.602 | 1.2716 | 32.1359 | 40.8626 | 25.2728 | 2.0029 | 32.062 | 64.2178 | 16.0076 | 14.0564 | -0.1104 | 105.0805 | 232.1247 | 178.1753 | 230.0042 | 158.8391
2012-10-30 17:36 | 28.0863 | 246.1538 | 192.2594 | 244.0177 | 172.8021 | -1.5997 | 1.2716 | 32.1358 | 40.8626 | 25.2727 | 2.0029 | 32.062 | 64.2173 | 16.0077 | 14.0894 | -0.1103 | 105.0799 | 232.0644 | 178.17 | 229.9283 | 158.7127
2012-10-30 17:36 | 28.0833 | 246.19 | 192.2564 | 244.148 | 172.7133 | -1.5997 | 1.2716 | 32.1357 | 40.8625 | 25.2726 | 2.0029 | 32.0619 | 64.2176 | 16.0075 | 14.0416 | -0.1103 | 105.08 | 232.1484 | 178.2148 | 230.1064 | 158.6717
2012-10-30 17:36 | 28.0938 | 246.2475 | 192.2866 | 244.1154 | 172.8642 | -1.5962 | 1.2715 | 32.1354 | 40.8614 | 25.2729 | 2.0029 | 32.0615 | 64.2157 | 16.0077 | 14.0929 | -0.11 | 105.077 | 232.1546 | 178.1937 | 230.0225 | 158.7713
2012-10-30 17:36 | 28.093 | 246.2311 | 192.3172 | 244.1616 | 173.0856 | -1.6014 | 1.2715 | 32.135 | 40.8605 | 25.2727 | 2.0029 | 32.061 | 64.2139 | 16.0076 | 14.0717 | -0.1104 | 105.0744 | 232.1594 | 178.2455 | 230.0899 | 159.0139
2012-10-30 17:36 | 28.0968 | 246.2701 | 192.3602 | 244.0949 | 173.0114 | -1.6003 | 1.2715 | 32.1343 | 40.8584 | 25.2729 | 2.0028 | 32.0605 | 64.2123 | 16.0074 | 14.0551 | -0.1103 | 105.0706 | 232.215 | 178.3051 | 230.0398 | 158.9563
2012-10-30 17:36 | 28.1005 | 246.3326 | 192.3639 | 244.2397 | 173.019 | -1.6043 | 1.2715 | 32.1338 | 40.8571 | 25.273 | 2.0028 | 32.06 | 64.2099 | 16.0075 | 14.0465 | -0.1106 | 105.067 | 232.2861 | 178.3174 | 230.1932 | 158.9725
2012-10-30 17:36 | 28.099 | 246.2998 | 192.3899 | 244.25 | 173.154 | -1.6037 | 1.2715 | 32.1338 | 40.8571 | 25.273 | 2.0028 | 32.0599 | 64.2103 | 16.0074 | 14.0287 | -0.1106 | 105.0674 | 232.2711 | 178.3612 | 230.2213 | 159.1253
2012-10-30 17:36 | 28.0998 | 246.3593 | 192.4536 | 244.1802 | 173.1782 | -1.602 | 1.2715 | 32.1337 | 40.8572 | 25.2728 | 2.0028 | 32.0599 | 64.2101 | 16.0074 | 14.0377 | -0.1104 | 105.0673 | 232.3216 | 178.4159 | 230.1425 | 159.1405
2012-10-30 17:37 | 28.1035 | 246.316 | 192.4455 | 244.2427 | 173.1624 | -1.6003 | 1.2715 | 32.1336 | 40.8565 | 25.273 | 2.0028 | 32.0598 | 64.2093 | 16.0075 | 14.0577 | -0.1103 | 105.0658 | 232.2583 | 178.3878 | 230.185 | 159.1047
2012-10-30 17:37 | 28.1065 | 246.366 | 192.5427 | 244.2927 | 173.1576 | -1.5962 | 1.2714 | 32.1331 | 40.8552 | 25.2731 | 2.0028 | 32.0592 | 64.2072 | 16.0074 | 14.0281 | -0.11 | 105.0624 | 232.3379 | 178.5146 | 230.2646 | 159.1295
2012-10-30 17:37 | 28.1043 | 246.3873 | 192.5798 | 244.267 | 173.3036 | -1.6032 | 1.2714 | 32.1326 | 40.8534 | 25.2734 | 2.0027 | 32.0588 | 64.2057 | 16.0074 | 14.0462 | -0.1105 | 105.0591 | 232.3411 | 178.5336 | 230.2208 | 159.2574
2012-10-30 17:37 | 28.096 | 246.43 | 192.5519 | 244.2509 | 173.2407 | -1.6026 | 1.2714 | 32.1323 | 40.8527 | 25.2733 | 2.0027 | 32.0585 | 64.2045 | 16.0074 | 14.0094 | -0.1105 | 105.0572 | 232.4206 | 178.5425 | 230.2415 | 159.2313
2012-10-30 17:37 | 28.1058 | 246.4359 | 192.5892 | 244.3979 | 173.2271 | -1.6003 | 1.2714 | 32.1319 | 40.8516 | 25.2733 | 2.0027 | 32.058 | 64.2028 | 16.0073 | 14.0192 | -0.1103 | 105.0544 | 232.4167 | 178.57 | 230.3787 | 159.2079
2012-10-30 17:37 | 28.1058 | 246.4711 | 192.6441 | 244.394 | 173.2661 | -1.6026 | 1.2714 | 32.1318 | 40.851 | 25.2736 | 2.0027 | 32.0579 | 64.2025 | 16.0073 | 14.0437 | -0.1105 | 105.0535 | 232.4274 | 178.6004 | 230.3503 | 159.2224
2012-10-30 17:37 | 28.1102 | 246.4559 | 192.6642 | 244.3278 | 173.4031 | -1.6014 | 1.2714 | 32.1319 | 40.852 | 25.2731 | 2.0027 | 32.0581 | 64.2033 | 16.0073 | 13.995 | -0.1104 | 105.0553 | 232.4609 | 178.6692 | 230.3328 | 159.4081
2012-10-30 17:37 | 28.1095 | 246.5297 | 192.6832 | 244.3036 | 173.3712 | -1.5933 | 1.2714 | 32.1317 | 40.851 | 25.2735 | 2.0027 | 32.0578 | 64.2027 | 16.0072 | 14.0025 | -0.1098 | 105.0537 | 232.5272 | 178.6807 | 230.3011 | 159.3687
2012-10-30 17:37 | 28.108 | 246.5517 | 192.6974 | 244.4119 | 173.2878 | -1.5985 | 1.2713 | 32.1312 | 40.8492 | 25.2737 | 2.0027 | 32.0572 | 64.2004 | 16.0072 | 14.0214 | -0.1102 | 105.0496 | 232.5303 | 178.676 | 230.3905 | 159.2664
2012-10-30 17:37 | 28.1028 | 246.5347 | 192.7158 | 244.489 | 173.3723 | -1.5997 | 1.2713 | 32.1307 | 40.8483 | 25.2735 | 2.0026 | 32.0567 | 64.1981 | 16.0072 | 13.9999 | -0.1103 | 105.0465 | 232.5348 | 178.7159 | 230.4891 | 159.3724
2012-10-30 17:37 | 28.1028 | 246.6209 | 192.7825 | 244.4459 | 173.4152 | -1.6009 | 1.2713 | 32.1302 | 40.8473 | 25.2734 | 2.0026 | 32.0563 | 64.1972 | 16.007 | 14.0203 | -0.1104 | 105.0446 | 232.6006 | 178.7622 | 230.4256 | 159.3949
2012-10-30 17:37 | 28.1005 | 246.6265 | 192.7802 | 244.4828 | 173.3739 | -1.5962 | 1.2713 | 32.1297 | 40.8457 | 25.2736 | 2.0026 | 32.0558 | 64.1951 | 16.0071 | 13.9853 | -0.11 | 105.0408 | 232.6412 | 178.7949 | 230.4975 | 159.3886
2012-10-30 17:37 | 28.099 | 246.625 | 192.7591 | 244.4813 | 173.3529 | -1.5951 | 1.2713 | 32.1295 | 40.8449 | 25.2737 | 2.0026 | 32.0555 | 64.194 | 16.007 | 14.0083 | -0.11 | 105.0389 | 232.6167 | 178.7508 | 230.473 | 159.3446
2012-10-30 17:37 | 28.099 | 246.7503 | 192.7787 | 244.6341 | 173.1696 | -1.5957 | 1.2682 | 32.0517 | 40.6478 | 25.2735 | 1.9975 | 31.9745 | 63.869 | 16.0073 | 14.0409 | -0.11 | 104.5167 | 232.7094 | 178.7378 | 230.5932 | 159.1287
2012-10-30 17:37 | 28.0945 | 246.6988 | 192.7664 | 244.563 | 173.1612 | -1.598 | 1.2652 | 31.9734 | 40.4514 | 25.2722 | 1.9928 | 31.8997 | 63.5707 | 16.0072 | 13.9997 | -0.1102 | 104.0221 | 232.6991 | 178.7667 | 230.5633 | 159.1615
2012-10-30 17:37 | 28.096 | 246.6611 | 192.7404 | 244.5135 | 173.3577 | -1.5985 | 1.2645 | 31.9559 | 40.4074 | 25.2721 | 1.9918 | 31.8826 | 63.5023 | 16.0073 | 14.0216 | -0.1102 | 103.9097 | 232.6395 | 178.7188 | 230.4919 | 159.3361
2012-10-30 17:37 | 28.093 | 246.666 | 192.7374 | 244.5654 | 173.4756 | -1.6032 | 1.264 | 31.9434 | 40.376 | 25.272 | 1.991 | 31.87 | 63.4517 | 16.0074 | 14.0186 | -0.1105 | 103.8277 | 232.6474 | 178.7188 | 230.5468 | 159.457
2012-10-30 17:37 | 28.102 | 246.6632 | 192.7542 | 244.6332 | 173.3754 | -1.602 | 1.2638 | 31.9399 | 40.3672 | 25.272 | 1.9908 | 31.8665 | 63.4383 | 16.0073 | 14.0358 | -0.1104 | 103.8055 | 232.6274 | 178.7184 | 230.5974 | 159.3396
2012-10-30 17:37 | 28.1013 | 246.6743 | 192.8007 | 244.5933 | 173.4137 | -1.6032 | 1.2638 | 31.9391 | 40.3647 | 25.2722 | 1.9907 | 31.8656 | 63.4339 | 16.0074 | 14.0025 | -0.1105 | 103.7986 | 232.6718 | 178.7982 | 230.5908 | 159.4112
2012-10-30 17:37 | 28.1013 | 246.6899 | 192.7928 | 244.6051 | 172.9964 | -1.6043 | 1.2638 | 31.9398 | 40.3671 | 25.2719 | 1.9907 | 31.8663 | 63.4366 | 16.0075 | 14.0392 | -0.1106 | 103.8037 | 232.6507 | 178.7536 | 230.5659 | 158.9572
2012-10-30 17:37 | 28.099 | 246.6994 | 192.7159 | 244.6224 | 172.7875 | -1.5974 | 1.2638 | 31.9404 | 40.3677 | 25.2724 | 1.9907 | 31.867 | 63.4389 | 16.0076 | 14.0083 | -0.1101 | 103.8066 | 232.6911 | 178.7076 | 230.6141 | 158.7792
2012-10-30 17:37 | 28.0998 | 246.708 | 192.701 | 244.729 | 173.0924 | -1.5962 | 1.2639 | 31.9406 | 40.3682 | 25.2724 | 1.9908 | 31.8671 | 63.4395 | 16.0076 | 13.9642 | -0.11 | 103.8077 | 232.7438 | 178.7368 | 230.7648 | 159.1282
2012-10-30 17:37 | 28.0983 | 246.7379 | 192.7034 | 244.5864 | 172.97 | -1.598 | 1.2639 | 31.9427 | 40.3737 | 25.2723 | 1.9909 | 31.8693 | 63.4479 | 16.0077 | 14.0117 | -0.1102 | 103.8216 | 232.7262 | 178.6917 | 230.5747 | 158.9583
2012-10-30 17:37 | 28.1043 | 246.7086 | 192.6269 | 244.6316 | 172.8474 | -1.6037 | 1.2662 | 32.0014 | 40.5216 | 25.2727 | 1.9945 | 31.9275 | 63.6804 | 16.0075 | 14.0218 | -0.1106 | 104.202 | 232.6868 | 178.6051 | 230.6098 | 158.8256
2012-10-30 17:37 | 28.0968 | 246.7207 | 192.5919 | 244.679 | 173.0426 | -1.5985 | 1.2691 | 32.0748 | 40.706 | 25.2737 | 1.9991 | 32.001 | 63.9739 | 16.0075 | 14.0102 | -0.1102 | 104.6799 | 232.7105 | 178.5817 | 230.6688 | 159.0324
2012-10-30 17:38 | 28.1035 | 246.7509 | 192.685 | 244.6465 | 173.2131 | -1.6014 | 1.2697 | 32.09 | 40.7441 | 25.274 | 2.0001 | 32.0165 | 64.0358 | 16.0075 | 13.9883 | -0.1104 | 104.7799 | 232.7626 | 178.6967 | 230.6582 | 159.2248
2012-10-30 17:38 | 28.0983 | 246.7927 | 192.727 | 244.5903 | 173.4068 | -1.5997 | 1.2702 | 32.1041 | 40.7796 | 25.2743 | 2.001 | 32.0305 | 64.0927 | 16.0073 | 13.9791 | -0.1103 | 104.8724 | 232.8136 | 178.7479 | 230.6112 | 159.4277
2012-10-30 17:38 | 28.1013 | 246.7683 | 192.7457 | 244.7227 | 173.4371 | -1.6061 | 1.2707 | 32.1155 | 40.808 | 25.2745 | 2.0017 | 32.0418 | 64.1375 | 16.0074 | 14.0106 | -0.1107 | 104.9455 | 232.7577 | 178.7351 | 230.7121 | 159.4265
2012-10-30 17:38 | 28.1058 | 246.8159 | 192.758 | 244.7233 | 173.5742 | -1.602 | 1.271 | 32.1228 | 40.8268 | 25.2744 | 2.0021 | 32.0489 | 64.1659 | 16.0075 | 14.0314 | -0.1104 | 104.9927 | 232.7845 | 178.7266 | 230.6919 | 159.5428
2012-10-30 17:38 | 28.0998 | 246.8216 | 192.8463 | 244.6663 | 173.6423 | -1.6055 | 1.2711 | 32.1277 | 40.8388 | 25.2747 | 2.0024 | 32.0536 | 64.1848 | 16.0074 | 14.0254 | -0.1107 | 105.0236 | 232.7962 | 178.8209 | 230.6409 | 159.6169
2012-10-30 17:38 | 28.1043 | 246.8653 | 192.9333 | 244.5728 | 173.8458 | -1.6009 | 1.2712 | 32.1304 | 40.8458 | 25.2747 | 2.0026 | 32.0565 | 64.1965 | 16.0074 | 14.0055 | -0.1104 | 105.0423 | 232.8598 | 178.9278 | 230.5673 | 159.8403
2012-10-30 17:38 | 28.0908 | 246.844 | 192.9237 | 244.7043 | 173.8245 | -1.5974 | 1.2713 | 32.1322 | 40.85 | 25.2749 | 2.0027 | 32.0584 | 64.2041 | 16.0074 | 13.9552 | -0.1101 | 105.0542 | 232.8888 | 178.9685 | 230.7491 | 159.8693
2012-10-30 17:38 | 28.1028 | 246.8521 | 192.9907 | 244.7399 | 173.6999 | -1.6032 | 1.2714 | 32.1337 | 40.8534 | 25.2751 | 2.0028 | 32.0598 | 64.2103 | 16.0072 | 14.0203 | -0.1105 | 105.0636 | 232.8318 | 178.9704 | 230.7196 | 159.6796
2012-10-30 17:38 | 28.1035 | 246.8998 | 193.0464 | 244.6543 | 173.7318 | -1.602 | 1.2714 | 32.135 | 40.8565 | 25.2753 | 2.0029 | 32.061 | 64.2148 | 16.0074 | 14.0128 | -0.1104 | 105.0713 | 232.887 | 179.0336 | 230.6415 | 159.719
2012-10-30 17:38 | 28.102 | 246.9257 | 193.0449 | 244.692 | 173.7927 | -1.6003 | 1.2714 | 32.1356 | 40.8579 | 25.2753 | 2.0029 | 32.0614 | 64.2171 | 16.0072 | 13.995 | -0.1103 | 105.075 | 232.9307 | 179.0499 | 230.697 | 159.7977
2012-10-30 17:38 | 28.1073 | 246.9467 | 193.058 | 244.7718 | 173.6654 | -1.6014 | 1.2715 | 32.1364 | 40.8603 | 25.2751 | 2.003 | 32.0624 | 64.221 | 16.0072 | 14.0003 | -0.1104 | 105.0813 | 232.9464 | 179.0577 | 230.7715 | 159.6651
2012-10-30 17:38 | 28.1028 | 246.9304 | 193.0496 | 244.7987 | 173.6375 | -1.598 | 1.2715 | 32.1369 | 40.8608 | 25.2756 | 2.003 | 32.0629 | 64.2227 | 16.0073 | 14.0244 | -0.1102 | 105.0835 | 232.906 | 179.0252 | 230.7743 | 159.6131
2012-10-30 17:38 | 28.1073 | 246.9741 | 193.1012 | 244.7365 | 173.6264 | -1.5974 | 1.2715 | 32.137 | 40.8617 | 25.2751 | 2.003 | 32.063 | 64.2232 | 16.0072 | 14.0166 | -0.1101 | 105.085 | 232.9575 | 179.0846 | 230.7199 | 159.6098
2012-10-30 17:38 | 28.1132 | 246.9643 | 193.0796 | 244.8326 | 173.6089 | -1.5991 | 1.2715 | 32.1378 | 40.8632 | 25.2754 | 2.0031 | 32.0637 | 64.226 | 16.0072 | 14.0225 | -0.1102 | 105.0892 | 232.9418 | 179.0571 | 230.8101 | 159.5864
2012-10-30 17:38 | 28.1147 | 247.0246 | 193.1047 | 244.8733 | 173.4856 | -1.602 | 1.2715 | 32.1383 | 40.8651 | 25.2751 | 2.0031 | 32.0643 | 64.2291 | 16.007 | 14.0444 | -0.1104 | 105.0942 | 232.9802 | 179.0603 | 230.8289 | 159.4412
2012-10-30 17:38 | 28.1147 | 246.9815 | 193.0929 | 244.8458 | 173.7781 | -1.602 | 1.2715 | 32.1385 | 40.8653 | 25.2753 | 2.0032 | 32.0644 | 64.2299 | 16.0069 | 14.0159 | -0.1104 | 105.0952 | 232.9656 | 179.077 | 230.8299 | 159.7622
2012-10-30 17:38 | 28.123 | 247.029 | 193.1444 | 244.8032 | 173.5719 | -1.6009 | 1.2715 | 32.1386 | 40.8651 | 25.2756 | 2.0032 | 32.0646 | 64.2306 | 16.007 | 14.0323 | -0.1104 | 105.0958 | 232.9967 | 179.1121 | 230.7709 | 159.5396
2012-10-30 17:38 | 28.1215 | 247.0784 | 193.1468 | 244.8879 | 173.4339 | -1.6009 | 1.2715 | 32.1385 | 40.8651 | 25.2755 | 2.0032 | 32.0646 | 64.2305 | 16.007 | 14.0063 | -0.1104 | 105.0956 | 233.0721 | 179.1405 | 230.8816 | 159.4276
2012-10-30 17:38 | 28.1177 | 247.0197 | 193.0881 | 244.9429 | 173.5783 | -1.6009 | 1.2715 | 32.1389 | 40.8657 | 25.2756 | 2.0032 | 32.0648 | 64.2319 | 16.0069 | 14.0433 | -0.1104 | 105.0976 | 232.9764 | 179.0448 | 230.8996 | 159.535
2012-10-30 17:38 | 28.1222 | 247.1183 | 193.1397 | 244.8337 | 173.5984 | -1.6003 | 1.2715 | 32.1391 | 40.8662 | 25.2757 | 2.0032 | 32.0649 | 64.2321 | 16.0069 | 14.0193 | -0.1103 | 105.0983 | 233.099 | 179.1204 | 230.8144 | 159.5791
2012-10-30 17:38 | 28.1252 | 247.0507 | 193.0956 | 244.9112 | 173.6014 | -1.5974 | 1.2715 | 32.1393 | 40.8664 | 25.2759 | 2.0032 | 32.0653 | 64.2336 | 16.007 | 14.0386 | -0.1101 | 105.0999 | 233.0121 | 179.057 | 230.8726 | 159.5628
2012-10-30 17:38 | 28.1342 | 247.1224 | 193.1792 | 244.9986 | 173.7352 | -1.6003 | 1.2716 | 32.1394 | 40.867 | 25.2757 | 2.0032 | 32.0654 | 64.2339 | 16.0069 | 14.0313 | -0.1103 | 105.1009 | 233.0911 | 179.1479 | 230.9673 | 159.7039
2012-10-30 17:38 | 28.1185 | 247.1263 | 193.1321 | 244.881 | 173.6883 | -1.5962 | 1.2716 | 32.1393 | 40.867 | 25.2755 | 2.0032 | 32.0654 | 64.2341 | 16.007 | 14.0074 | -0.11 | 105.1011 | 233.1189 | 179.1247 | 230.8736 | 159.6809
2012-10-30 17:38 | 28.1237 | 247.1158 | 193.1962 | 244.9097 | 173.7364 | -1.6003 | 1.2715 | 32.1392 | 40.8657 | 25.2762 | 2.0032 | 32.0652 | 64.234 | 16.0068 | 14.033 | -0.1103 | 105.0996 | 233.0828 | 179.1632 | 230.8767 | 159.7034
2012-10-30 17:38 | 28.1275 | 247.104 | 193.2118 | 245.0703 | 173.7363 | -1.6009 | 1.2715 | 32.1393 | 40.8662 | 25.2759 | 2.0032 | 32.0653 | 64.2342 | 16.0068 | 14.0327 | -0.1104 | 105.1005 | 233.0713 | 179.1791 | 231.0376 | 159.7036
2012-10-30 17:39 | 28.1267 | 247.1541 | 193.207 | 244.9911 | 173.8331 | -1.5997 | 1.2716 | 32.1396 | 40.8674 | 25.2757 | 2.0033 | 32.0655 | 64.2354 | 16.0067 | 14.0197 | -0.1103 | 105.1028 | 233.1344 | 179.1873 | 230.9714 | 159.8134
2012-10-30 17:39 | 28.1207 | 247.1677 | 193.2442 | 244.9498 | 173.8583 | -1.6014 | 1.2715 | 32.1391 | 40.8659 | 25.2759 | 2.0032 | 32.0651 | 64.2339 | 16.0067 | 13.9933 | -0.1104 | 105.0998 | 233.1744 | 179.2509 | 230.9565 | 159.865
2012-10-30 17:39 | 28.1252 | 247.1996 | 193.2212 | 245.0484 | 173.8784 | -1.6037 | 1.2715 | 32.1393 | 40.8663 | 25.2759 | 2.0032 | 32.0653 | 64.2346 | 16.0067 | 14.0182 | -0.1106 | 105.1009 | 233.1814 | 179.203 | 231.0302 | 159.8602
2012-10-30 17:39 | 28.1237 | 247.2177 | 193.2512 | 245.0704 | 173.8964 | -1.6009 | 1.2715 | 32.1394 | 40.8666 | 25.2759 | 2.0033 | 32.0654 | 64.2353 | 16.0066 | 14.0697 | -0.1104 | 105.1019 | 233.148 | 179.1815 | 231.0007 | 159.8267
2012-10-30 17:39 | 28.1237 | 247.2255 | 193.2865 | 245.1018 | 173.8691 | -1.6043 | 1.2715 | 32.1391 | 40.8658 | 25.276 | 2.0032 | 32.065 | 64.2337 | 16.0066 | 14.086 | -0.1106 | 105.0995 | 233.1395 | 179.2005 | 231.0158 | 159.7831
2012-10-30 17:39 | 28.129 | 247.2622 | 193.3586 | 245.0287 | 173.9797 | -1.602 | 1.2715 | 32.1385 | 40.8642 | 25.276 | 2.0032 | 32.0645 | 64.2318 | 16.0066 | 14.0709 | -0.1104 | 105.096 | 233.1913 | 179.2877 | 230.9578 | 159.9088
2012-10-30 17:39 | 28.1237 | 247.2412 | 193.3376 | 245.09 | 173.9315 | -1.6014 | 1.2715 | 32.1381 | 40.8631 | 25.2761 | 2.0032 | 32.064 | 64.2304 | 16.0065 | 14.0942 | -0.1104 | 105.0935 | 233.147 | 179.2434 | 230.9958 | 159.8373
2012-10-30 17:39 | 28.1267 | 247.3225 | 193.3681 | 245.1126 | 173.9852 | -1.5974 | 1.2715 | 32.138 | 40.8628 | 25.2761 | 2.0032 | 32.064 | 64.2304 | 16.0065 | 14.0605 | -0.1101 | 105.0932 | 233.262 | 179.3076 | 231.0521 | 159.9247
2012-10-30 17:39 | 28.129 | 247.3209 | 193.4057 | 245.1737 | 173.9875 | -1.6032 | 1.2715 | 32.1378 | 40.8618 | 25.2763 | 2.0032 | 32.0637 | 64.2289 | 16.0065 | 14.0465 | -0.1105 | 105.0906 | 233.2744 | 179.3592 | 231.1272 | 159.941
2012-10-30 17:39 | 28.126 | 247.3179 | 193.4066 | 245.1746 | 173.9416 | -1.5991 | 1.2714 | 32.1373 | 40.8608 | 25.2763 | 2.0031 | 32.0633 | 64.2272 | 16.0066 | 14.0149 | -0.1102 | 105.0879 | 233.303 | 179.3917 | 231.1597 | 159.9267
2012-10-30 17:39 | 28.126 | 247.3336 | 193.4341 | 245.1668 | 173.9416 | -1.6003 | 1.2714 | 32.1373 | 40.8607 | 25.2762 | 2.0031 | 32.0632 | 64.2273 | 16.0064 | 14.0924 | -0.1103 | 105.0879 | 233.2412 | 179.3417 | 231.0744 | 159.8492
2012-10-30 17:39 | 28.1162 | 247.3512 | 193.4754 | 245.1296 | 174.0293 | -1.6014 | 1.2714 | 32.1368 | 40.8597 | 25.2761 | 2.0031 | 32.0627 | 64.2252 | 16.0065 | 14.0541 | -0.1104 | 105.085 | 233.2971 | 179.4213 | 231.0755 | 159.9752
2012-10-30 17:39 | 28.1237 | 247.3587 | 193.4711 | 245.235 | 174.0602 | -1.6026 | 1.2714 | 32.1365 | 40.8586 | 25.2763 | 2.0031 | 32.0625 | 64.225 | 16.0063 | 14.033 | -0.1105 | 105.0835 | 233.3257 | 179.4381 | 231.202 | 160.0272
2012-10-30 17:39 | 28.1282 | 247.3554 | 193.5031 | 245.2944 | 174.1076 | -1.6009 | 1.2714 | 32.1364 | 40.8586 | 25.2762 | 2.0031 | 32.0624 | 64.2247 | 16.0062 | 14.062 | -0.1104 | 105.0832 | 233.2934 | 179.4411 | 231.2324 | 160.0456
2012-10-30 17:39 | 28.1275 | 247.386 | 193.5063 | 245.2428 | 174.2084 | -1.5991 | 1.2714 | 32.136 | 40.8573 | 25.2763 | 2.0031 | 32.0619 | 64.2228 | 16.0062 | 14.0491 | -0.1102 | 105.0802 | 233.3369 | 179.4572 | 231.1937 | 160.1593
2012-10-30 17:39 | 28.1267 | 247.4322 | 193.5134 | 245.2616 | 173.9657 | -1.6061 | 1.2714 | 32.1354 | 40.8556 | 25.2765 | 2.0031 | 32.0615 | 64.221 | 16.0062 | 14.0564 | -0.1107 | 105.0766 | 233.3758 | 179.457 | 231.2052 | 159.9093
2012-10-30 17:39 | 28.1357 | 247.4412 | 193.5499 | 245.2353 | 173.8928 | -1.602 | 1.2714 | 32.1351 | 40.855 | 25.2763 | 2.003 | 32.0611 | 64.2197 | 16.0063 | 14.0858 | -0.1104 | 105.0747 | 233.3554 | 179.4641 | 231.1495 | 159.807
2012-10-30 17:39 | 28.1312 | 247.4524 | 193.5179 | 245.3523 | 173.9741 | -1.6014 | 1.2713 | 32.1348 | 40.8539 | 25.2766 | 2.003 | 32.0609 | 64.2189 | 16.0062 | 14.0609 | -0.1104 | 105.0728 | 233.3915 | 179.457 | 231.2914 | 159.9132
2012-10-30 17:39 | 28.1312 | 247.4524 | 193.5297 | 245.3288 | 174.0755 | -1.6084 | 1.2713 | 32.1344 | 40.8529 | 25.2764 | 2.003 | 32.0603 | 64.2164 | 16.0062 | 14.0731 | -0.1109 | 105.0693 | 233.3793 | 179.4566 | 231.2557 | 160.0024
2012-10-30 17:39 | 28.1357 | 247.4843 | 193.5892 | 245.2823 | 174.08 | -1.5997 | 1.2713 | 32.1337 | 40.8508 | 25.2768 | 2.0029 | 32.0597 | 64.2141 | 16.0063 | 14.0287 | -0.1103 | 105.0649 | 233.4556 | 179.5605 | 231.2536 | 160.0513
2012-10-30 17:39 | 28.129 | 247.4893 | 193.555 | 245.2835 | 174.3035 | -1.6049 | 1.2712 | 32.133 | 40.8491 | 25.2767 | 2.0029 | 32.0589 | 64.2109 | 16.0062 | 14.0261 | -0.1106 | 105.06 | 233.4632 | 179.5289 | 231.2574 | 160.2774
2012-10-30 17:39 | 28.1252 | 247.5051 | 193.5787 | 245.3933 | 174.1905 | -1.5997 | 1.2712 | 32.1327 | 40.8487 | 25.2764 | 2.0029 | 32.0587 | 64.2102 | 16.0062 | 14.0427 | -0.1103 | 105.0589 | 233.4624 | 179.536 | 231.3506 | 160.1478
2012-10-30 17:39 | 28.1327 | 247.4852 | 193.6294 | 245.3969 | 174.198 | -1.6009 | 1.2712 | 32.1324 | 40.8474 | 25.2768 | 2.0029 | 32.0584 | 64.2086 | 16.0062 | 14.0746 | -0.1104 | 105.0559 | 233.4106 | 179.5548 | 231.3223 | 160.1234
2012-10-30 17:39 | 28.141 | 247.5327 | 193.6534 | 245.3229 | 174.0814 | -1.5997 | 1.2712 | 32.132 | 40.8461 | 25.2769 | 2.0029 | 32.058 | 64.2075 | 16.0061 | 14.0054 | -0.1103 | 105.0537 | 233.5273 | 179.648 | 231.3175 | 160.076
2012-10-30 17:40 | 28.1439 | 247.5356 | 193.6563 | 245.3532 | 173.9322 | -1.5968 | 1.2712 | 32.1319 | 40.8454 | 25.2772 | 2.0029 | 32.0578 | 64.2073 | 16.006 | 14.0573 | -0.1101 | 105.0527 | 233.4783 | 179.599 | 231.2959 | 159.8749
2012-10-30 17:40 | 28.1469 | 247.5464 | 193.5807 | 245.4542 | 173.9274 | -1.6003 | 1.2712 | 32.1315 | 40.8449 | 25.2769 | 2.0028 | 32.0574 | 64.2054 | 16.0061 | 14.0399 | -0.1103 | 105.0503 | 233.5065 | 179.5408 | 231.4143 | 159.8875
2012-10-30 17:40 | 28.138 | 247.5923 | 193.56 | 245.4688 | 173.9809 | -1.6055 | 1.2712 | 32.131 | 40.8436 | 25.2769 | 2.0028 | 32.057 | 64.204 | 16.006 | 14.031 | -0.1107 | 105.0476 | 233.5613 | 179.529 | 231.4378 | 159.9499
2012-10-30 17:40 | 28.1417 | 247.5725 | 193.5912 | 245.449 | 174.0782 | -1.5962 | 1.2711 | 32.1307 | 40.8428 | 25.2769 | 2.0028 | 32.0567 | 64.2029 | 16.006 | 14.051 | -0.11 | 105.0457 | 233.5215 | 179.5402 | 231.398 | 160.0272
2012-10-30 17:40 | 28.1469 | 247.5934 | 193.5886 | 245.4816 | 174.1068 | -1.5974 | 1.2711 | 32.1305 | 40.8418 | 25.2772 | 2.0028 | 32.0565 | 64.2023 | 16.006 | 14.0277 | -0.1101 | 105.0441 | 233.5657 | 179.5609 | 231.4539 | 160.0791
2012-10-30 17:40 | 28.1507 | 247.6011 | 193.6317 | 245.4384 | 174.1574 | -1.5991 | 1.2711 | 32.1302 | 40.8412 | 25.2771 | 2.0028 | 32.0562 | 64.2008 | 16.006 | 13.9988 | -0.1102 | 105.0421 | 233.6023 | 179.6329 | 231.4396 | 160.1586
2012-10-30 17:40 | 28.1507 | 247.6481 | 193.6277 | 245.4227 | 173.939 | -1.5968 | 1.2678 | 32.0456 | 40.6271 | 25.2768 | 1.9976 | 31.9737 | 63.871 | 16.006 | 14.0315 | -0.1101 | 104.4981 | 233.6166 | 179.5962 | 231.3912 | 159.9075
2012-10-30 17:40 | 28.1522 | 247.5909 | 193.5821 | 245.5379 | 173.9678 | -1.6043 | 1.2651 | 31.9757 | 40.4523 | 25.2754 | 1.9931 | 31.9021 | 63.5852 | 16.006 | 14.0003 | -0.1106 | 104.0375 | 233.5906 | 179.5818 | 231.5376 | 159.9675
2012-10-30 17:40 | 28.1462 | 247.5849 | 193.58 | 245.4927 | 174.149 | -1.598 | 1.2644 | 31.9581 | 40.408 | 25.2752 | 1.992 | 31.8844 | 63.5141 | 16.0062 | 14.0025 | -0.1102 | 103.922 | 233.5824 | 179.5775 | 231.4902 | 160.1465
2012-10-30 17:40 | 28.1484 | 247.591 | 193.6019 | 245.5106 | 174.2878 | -1.6009 | 1.2639 | 31.9461 | 40.3778 | 25.2751 | 1.9912 | 31.8725 | 63.4662 | 16.0063 | 14.0047 | -0.1104 | 103.8439 | 233.5863 | 179.5972 | 231.5059 | 160.2831
2012-10-30 17:40 | 28.1469 | 247.5856 | 193.6357 | 245.4425 | 174.3292 | -1.6009 | 1.2637 | 31.9415 | 40.366 | 25.2753 | 1.991 | 31.8679 | 63.4482 | 16.0062 | 14.0644 | -0.1104 | 103.8142 | 233.5212 | 179.5713 | 231.3781 | 160.2648
2012-10-30 17:40 | 28.1425 | 247.5655 | 193.6235 | 245.4968 | 174.317 | -1.6043 | 1.2637 | 31.9404 | 40.3635 | 25.2749 | 1.9909 | 31.8668 | 63.4435 | 16.0063 | 13.9988 | -0.1106 | 103.8071 | 233.5667 | 179.6247 | 231.498 | 160.3182
2012-10-30 17:40 | 28.1439 | 247.5277 | 193.6209 | 245.5296 | 173.8932 | -1.5985 | 1.2638 | 31.9416 | 40.3663 | 25.2752 | 1.991 | 31.868 | 63.4483 | 16.0062 | 13.9879 | -0.1102 | 103.8146 | 233.5398 | 179.633 | 231.5417 | 159.9053
2012-10-30 17:40 | 28.1492 | 247.5761 | 193.6027 | 245.531 | 173.641 | -1.6032 | 1.2638 | 31.9427 | 40.3689 | 25.2753 | 1.991 | 31.8691 | 63.4524 | 16.0063 | 14.0626 | -0.1105 | 103.8213 | 233.5135 | 179.5401 | 231.4684 | 159.5784
2012-10-30 17:40 | 28.1365 | 247.5556 | 193.5036 | 245.5614 | 173.3436 | -1.6003 | 1.2639 | 31.945 | 40.3743 | 25.2755 | 1.9912 | 31.8713 | 63.4607 | 16.0065 | 14.0458 | -0.1103 | 103.835 | 233.5098 | 179.4578 | 231.5156 | 159.2978
2012-10-30 17:40 | 28.1387 | 247.5891 | 193.4429 | 245.5048 | 173.3068 | -1.6026 | 1.2644 | 31.9594 | 40.4106 | 25.2756 | 1.9919 | 31.883 | 63.5073 | 16.0064 | 14.0235 | -0.1105 | 103.9179 | 233.5656 | 179.4194 | 231.4813 | 159.2833
2012-10-30 17:40 | 28.141 | 247.6149 | 193.4374 | 245.4836 | 173.6562 | -1.6014 | 1.2688 | 32.0703 | 40.6898 | 25.2767 | 1.999 | 31.9965 | 63.9602 | 16.0065 | 14.0299 | -0.1104 | 104.65 | 233.585 | 179.4075 | 231.4537 | 159.6263
2012-10-30 17:40 | 28.1305 | 247.624 | 193.4151 | 245.4927 | 173.6184 | -1.6055 | 1.2695 | 32.0892 | 40.7376 | 25.2769 | 2.0002 | 32.0153 | 64.0355 | 16.0064 | 14.0317 | -0.1107 | 104.773 | 233.5923 | 179.3834 | 231.461 | 159.5867
2012-10-30 17:40 | 28.1402 | 247.5554 | 193.4523 | 245.522 | 173.9246 | -1.6026 | 1.2701 | 32.1051 | 40.7774 | 25.2772 | 2.0012 | 32.0311 | 64.0999 | 16.0062 | 14.0699 | -0.1105 | 104.8773 | 233.4855 | 179.3824 | 231.4521 | 159.8547
2012-10-30 17:40 | 28.1342 | 247.6316 | 193.5366 | 245.563 | 174.0005 | -1.5991 | 1.2705 | 32.115 | 40.8029 | 25.2769 | 2.0018 | 32.0409 | 64.1391 | 16.0061 | 14.0313 | -0.1102 | 104.9421 | 233.6003 | 179.5053 | 231.5317 | 159.9692
2012-10-30 17:40 | 28.141 | 247.5875 | 193.567 | 245.5541 | 173.9527 | -1.5974 | 1.2708 | 32.122 | 40.8199 | 25.2775 | 2.0022 | 32.048 | 64.1675 | 16.0062 | 14.0381 | -0.1101 | 104.9873 | 233.5494 | 179.5289 | 231.516 | 159.9146
2012-10-30 17:40 | 28.132 | 247.6373 | 193.554 | 245.4119 | 174.0373 | -1.6014 | 1.271 | 32.1277 | 40.8345 | 25.2774 | 2.0026 | 32.0538 | 64.1909 | 16.0061 | 14.0617 | -0.1104 | 105.0253 | 233.5756 | 179.4923 | 231.3502 | 159.9756
2012-10-30 17:40 | 28.129 | 247.6499 | 193.555 | 245.5108 | 174.0928 | -1.6003 | 1.2711 | 32.1311 | 40.8427 | 25.2777 | 2.0028 | 32.0572 | 64.2047 | 16.006 | 14.0342 | -0.1103 | 105.0475 | 233.6157 | 179.5208 | 231.4766 | 160.0586
2012-10-30 17:40 | 28.1402 | 247.665 | 193.6055 | 245.5181 | 174.0026 | -1.602 | 1.2712 | 32.1333 | 40.848 | 25.2778 | 2.003 | 32.0593 | 64.214 | 16.0059 | 14.0169 | -0.1104 | 105.062 | 233.6481 | 179.5886 | 231.5012 | 159.9857
2012-10-30 17:41 | 28.1372 | 247.6699 | 193.6182 | 245.5582 | 174.1595 | -1.5991 | 1.2713 | 32.1347 | 40.8515 | 25.2779 | 2.0031 | 32.0606 | 64.219 | 16.0059 | 14.0343 | -0.1102 | 105.0705 | 233.6356 | 179.5839 | 231.5239 | 160.1252
2012-10-30 17:41 | 28.1395 | 247.7035 | 193.6165 | 245.5565 | 174.2555 | -1.6055 | 1.2713 | 32.1356 | 40.8535 | 25.2781 | 2.0031 | 32.0615 | 64.2224 | 16.0059 | 14.0733 | -0.1107 | 105.0759 | 233.6302 | 179.5432 | 231.4832 | 160.1822
2012-10-30 17:41 | 28.138 | 247.7216 | 193.6897 | 245.5119 | 174.0823 | -1.6049 | 1.2713 | 32.1368 | 40.8569 | 25.2779 | 2.0032 | 32.0626 | 64.2276 | 16.0058 | 14.0677 | -0.1106 | 105.0845 | 233.6539 | 179.622 | 231.4442 | 160.0146
2012-10-30 17:41 | 28.1417 | 247.694 | 193.6855 | 245.5156 | 174.0431 | -1.6043 | 1.2714 | 32.1376 | 40.8588 | 25.278 | 2.0032 | 32.0635 | 64.2305 | 16.0059 | 14.0714 | -0.1106 | 105.0893 | 233.6226 | 179.6141 | 231.4442 | 159.9717
2012-10-30 17:41 | 28.1372 | 247.709 | 193.6417 | 245.5582 | 174.023 | -1.5991 | 1.2714 | 32.1379 | 40.8597 | 25.2778 | 2.0032 | 32.0637 | 64.2316 | 16.0059 | 14.0261 | -0.1102 | 105.0913 | 233.6829 | 179.6156 | 231.5321 | 159.9969
2012-10-30 17:41 | 28.1425 | 247.7496 | 193.6353 | 245.5556 | 174.1687 | -1.5962 | 1.2714 | 32.1379 | 40.8594 | 25.278 | 2.0033 | 32.0638 | 64.2324 | 16.0058 | 14.0069 | -0.11 | 105.0919 | 233.7427 | 179.6284 | 231.5487 | 160.1618
2012-10-30 17:41 | 28.1499 | 247.7413 | 193.6741 | 245.5905 | 174.1839 | -1.6009 | 1.2714 | 32.1383 | 40.8603 | 25.278 | 2.0033 | 32.0642 | 64.2338 | 16.0058 | 14.0429 | -0.1104 | 105.0941 | 233.6984 | 179.6312 | 231.5476 | 160.141
2012-10-30 17:41 | 28.1454 | 247.7721 | 193.6696 | 245.6565 | 173.9883 | -1.5974 | 1.2714 | 32.1385 | 40.8613 | 25.2779 | 2.0033 | 32.0644 | 64.2352 | 16.0057 | 14.0914 | -0.1101 | 105.0965 | 233.6807 | 179.5782 | 231.5651 | 159.8969
2012-10-30 17:41 | 28.1589 | 247.7738 | 193.687 | 245.6034 | 174.0525 | -1.5968 | 1.2714 | 32.1384 | 40.8603 | 25.2782 | 2.0033 | 32.0642 | 64.2341 | 16.0057 | 14.1008 | -0.1101 | 105.0943 | 233.673 | 179.5862 | 231.5026 | 159.9517
2012-10-30 17:41 | 28.1544 | 247.7497 | 193.6864 | 245.591 | 174.1572 | -1.6009 | 1.2714 | 32.138 | 40.8594 | 25.2782 | 2.0033 | 32.0639 | 64.2332 | 16.0057 | 14.0474 | -0.1104 | 105.0926 | 233.7023 | 179.639 | 231.5436 | 160.1098
2012-10-30 17:41 | 28.1514 | 247.8172 | 193.6717 | 245.592 | 174.1347 | -1.6003 | 1.2714 | 32.1384 | 40.8606 | 25.278 | 2.0033 | 32.0642 | 64.2347 | 16.0056 | 14.0566 | -0.1103 | 105.0953 | 233.7606 | 179.6151 | 231.5354 | 160.0781
2012-10-30 17:41 | 28.1499 | 247.7962 | 193.6584 | 245.6649 | 174.18 | -1.5974 | 1.2714 | 32.1384 | 40.86 | 25.2784 | 2.0033 | 32.0642 | 64.2346 | 16.0056 | 14.0837 | -0.1101 | 105.0947 | 233.7125 | 179.5747 | 231.5812 | 160.0963
2012-10-30 17:41 | 28.1522 | 247.7475 | 193.7157 | 245.7456 | 174.1277 | -1.5945 | 1.2714 | 32.1381 | 40.8598 | 25.2781 | 2.0033 | 32.0638 | 64.234 | 16.0054 | 14.0534 | -0.1099 | 105.0937 | 233.6941 | 179.6623 | 231.6922 | 160.0743
2012-10-30 17:41 | 28.1552 | 247.825 | 193.6715 | 245.6977 | 174.2322 | -1.6061 | 1.2714 | 32.1386 | 40.8607 | 25.2784 | 2.0033 | 32.0645 | 64.2359 | 16.0056 | 14.0849 | -0.1107 | 105.0965 | 233.7401 | 179.5866 | 231.6128 | 160.1473
2012-10-30 17:41 | 28.1529 | 247.8422 | 193.7242 | 245.6758 | 174.304 | -1.6061 | 1.2714 | 32.1385 | 40.8606 | 25.2783 | 2.0033 | 32.0645 | 64.2358 | 16.0055 | 14.0541 | -0.1107 | 105.0964 | 233.7881 | 179.6701 | 231.6217 | 160.2499
2012-10-30 17:41 | 28.1589 | 247.8717 | 193.7302 | 245.6818 | 174.3529 | -1.5991 | 1.2714 | 32.1385 | 40.8604 | 25.2783 | 2.0033 | 32.0645 | 64.2363 | 16.0055 | 14.0437 | -0.1102 | 105.0967 | 233.828 | 179.6865 | 231.6381 | 160.3092
2012-10-30 17:41 | 28.1664 | 247.8557 | 193.8202 | 245.7363 | 174.4384 | -1.5922 | 1.2714 | 32.1383 | 40.8601 | 25.2783 | 2.0033 | 32.0642 | 64.2355 | 16.0053 | 14.0716 | -0.1098 | 105.0956 | 233.7841 | 179.7486 | 231.6647 | 160.3668
2012-10-30 17:41 | 28.1589 | 247.8913 | 193.7656 | 245.7601 | 174.2866 | -1.5962 | 1.2714 | 32.138 | 40.859 | 25.2785 | 2.0033 | 32.0638 | 64.234 | 16.0054 | 14.0723 | -0.11 | 105.093 | 233.819 | 179.6933 | 231.6878 | 160.2143
2012-10-30 17:41 | 28.1634 | 247.8606 | 193.8094 | 245.745 | 174.3574 | -1.6043 | 1.2713 | 32.1378 | 40.8582 | 25.2786 | 2.0033 | 32.0636 | 64.233 | 16.0054 | 14.0686 | -0.1106 | 105.0912 | 233.792 | 179.7408 | 231.6764 | 160.2888
2012-10-30 17:41 | 28.1649 | 247.8973 | 193.8266 | 245.7583 | 174.3745 | -1.6009 | 1.2713 | 32.1379 | 40.8583 | 25.2787 | 2.0033 | 32.0637 | 64.2341 | 16.0052 | 14.0783 | -0.1104 | 105.0924 | 233.819 | 179.7483 | 231.68 | 160.2962
2012-10-30 17:41 | 28.1657 | 247.9686 | 193.8981 | 245.7434 | 174.3714 | -1.602 | 1.2713 | 32.1375 | 40.8577 | 25.2785 | 2.0033 | 32.0635 | 64.2328 | 16.0053 | 14.0505 | -0.1104 | 105.0905 | 233.9181 | 179.8476 | 231.6929 | 160.3209
2012-10-30 17:41 | 28.1634 | 247.9702 | 193.884 | 245.7529 | 174.3808 | -1.5974 | 1.2713 | 32.1378 | 40.8582 | 25.2786 | 2.0033 | 32.0637 | 64.2335 | 16.0053 | 14.1176 | -0.1101 | 105.0917 | 233.8526 | 179.7664 | 231.6353 | 160.2632
2012-10-30 17:41 | 28.1664 | 247.9341 | 193.8909 | 245.8029 | 174.3994 | -1.6014 | 1.2713 | 32.1374 | 40.8577 | 25.2783 | 2.0033 | 32.0632 | 64.2325 | 16.0052 | 14.1083 | -0.1104 | 105.0902 | 233.8258 | 179.7826 | 231.6946 | 160.2911
2012-10-30 17:41 | 28.1664 | 247.938 | 193.8831 | 245.8774 | 174.4618 | -1.6037 | 1.2713 | 32.1372 | 40.8572 | 25.2783 | 2.0033 | 32.0631 | 64.2321 | 16.0052 | 14.1083 | -0.1106 | 105.0893 | 233.8297 | 179.7748 | 231.7691 | 160.3535
2012-10-30 17:42 | 28.1612 | 247.9563 | 193.9093 | 245.7742 | 174.4879 | -1.5957 | 1.2713 | 32.1368 | 40.8556 | 25.2787 | 2.0033 | 32.0626 | 64.2304 | 16.0051 | 14.1113 | -0.11 | 105.0859 | 233.845 | 179.798 | 231.6629 | 160.3766
2012-10-30 17:42 | 28.1657 | 247.9569 | 193.9688 | 245.763 | 174.4026 | -1.598 | 1.2713 | 32.1364 | 40.8549 | 25.2785 | 2.0033 | 32.0623 | 64.2291 | 16.0051 | 14.0873 | -0.1102 | 105.084 | 233.8696 | 179.8815 | 231.6757 | 160.3153
2012-10-30 17:42 | 28.1657 | 247.9412 | 193.9609 | 245.8296 | 174.4221 | -1.5974 | 1.2713 | 32.1357 | 40.8531 | 25.2785 | 2.0032 | 32.0616 | 64.2261 | 16.0051 | 14.1076 | -0.1101 | 105.0792 | 233.8336 | 179.8533 | 231.722 | 160.3145
2012-10-30 17:42 | 28.1717 | 247.9511 | 193.9669 | 245.9297 | 174.4047 | -1.5951 | 1.2713 | 32.1359 | 40.8536 | 25.2785 | 2.0032 | 32.0619 | 64.2275 | 16.0051 | 14.0606 | -0.11 | 105.0812 | 233.8905 | 179.9063 | 231.8691 | 160.3441
2012-10-30 17:42 | 28.1627 | 248.0244 | 193.9383 | 245.8266 | 174.3138 | -1.6049 | 1.2713 | 32.1358 | 40.8533 | 25.2786 | 2.0032 | 32.0617 | 64.2269 | 16.005 | 14.0965 | -0.1106 | 105.0802 | 233.9279 | 179.8418 | 231.7301 | 160.2173
2012-10-30 17:42 | 28.1634 | 247.9859 | 193.9586 | 245.8548 | 174.3301 | -1.5957 | 1.2713 | 32.1355 | 40.8524 | 25.2786 | 2.0032 | 32.0613 | 64.2255 | 16.005 | 14.0727 | -0.11 | 105.0779 | 233.9132 | 179.8859 | 231.7821 | 160.2574
2012-10-30 17:42 | 28.1709 | 248.0051 | 193.9622 | 245.8505 | 174.2869 | -1.6066 | 1.2712 | 32.1352 | 40.8509 | 25.279 | 2.0032 | 32.061 | 64.2242 | 16.005 | 14.0598 | -0.1108 | 105.0751 | 233.9453 | 179.9024 | 231.7907 | 160.2271
2012-10-30 17:42 | 28.1612 | 248.0072 | 193.9211 | 245.9349 | 174.2694 | -1.6009 | 1.2712 | 32.1349 | 40.8512 | 25.2784 | 2.0032 | 32.0608 | 64.2236 | 16.005 | 14.0991 | -0.1104 | 105.0749 | 233.9081 | 179.822 | 231.8358 | 160.1703
2012-10-30 17:42 | 28.1672 | 248.0132 | 193.9035 | 245.9409 | 174.3768 | -1.5974 | 1.2712 | 32.1352 | 40.8513 | 25.2787 | 2.0032 | 32.061 | 64.2249 | 16.0048 | 14.1295 | -0.1101 | 105.0763 | 233.8837 | 179.774 | 231.8114 | 160.2473
2012-10-30 17:42 | 28.1627 | 248.0244 | 193.9854 | 245.9481 | 174.3216 | -1.6014 | 1.2712 | 32.135 | 40.8505 | 25.279 | 2.0032 | 32.061 | 64.2252 | 16.0047 | 14.0843 | -0.1104 | 105.0757 | 233.9401 | 179.9011 | 231.8638 | 160.2373
2012-10-30 17:42 | 28.1597 | 248.0292 | 193.9628 | 245.9059 | 174.4005 | -1.6055 | 1.2712 | 32.135 | 40.8509 | 25.2787 | 2.0032 | 32.0607 | 64.2241 | 16.0047 | 14.0813 | -0.1107 | 105.075 | 233.9479 | 179.8815 | 231.8246 | 160.3192
2012-10-30 17:42 | 28.1657 | 248.0939 | 193.9413 | 245.8571 | 174.4846 | -1.6055 | 1.2712 | 32.1347 | 40.8495 | 25.2791 | 2.0032 | 32.0607 | 64.2239 | 16.0047 | 14.1076 | -0.1107 | 105.0734 | 233.9863 | 179.8337 | 231.7495 | 160.377
2012-10-30 17:42 | 28.1559 | 248.0606 | 193.9511 | 245.9178 | 174.5645 | -1.6003 | 1.2712 | 32.1343 | 40.849 | 25.2788 | 2.0032 | 32.0602 | 64.2224 | 16.0047 | 14.1264 | -0.1103 | 105.0714 | 233.9342 | 179.8247 | 231.7914 | 160.4381
2012-10-30 17:42 | 28.1612 | 248.0189 | 194.0075 | 246.0093 | 174.6361 | -1.6072 | 1.2712 | 32.1339 | 40.8474 | 25.2792 | 2.0031 | 32.0599 | 64.2205 | 16.0048 | 14.0705 | -0.1108 | 105.068 | 233.9484 | 179.937 | 231.9388 | 160.5656
2012-10-30 17:42 | 28.1589 | 248.0989 | 194.0484 | 245.9482 | 174.6846 | -1.6037 | 1.2711 | 32.1334 | 40.846 | 25.2792 | 2.0031 | 32.0593 | 64.218 | 16.0048 | 14.1049 | -0.1106 | 105.064 | 233.994 | 179.9435 | 231.8433 | 160.5797
2012-10-30 17:42 | 28.1499 | 248.0899 | 194.0669 | 245.8452 | 174.8043 | -1.6043 | 1.2711 | 32.1328 | 40.8452 | 25.2787 | 2.0031 | 32.0586 | 64.2158 | 16.0047 | 14.1122 | -0.1106 | 105.061 | 233.9777 | 179.9547 | 231.733 | 160.6921
2012-10-30 17:42 | 28.1619 | 248.0901 | 194.1103 | 245.9552 | 174.7188 | -1.6032 | 1.2711 | 32.1327 | 40.8451 | 25.2788 | 2.0031 | 32.0585 | 64.2155 | 16.0047 | 14.0998 | -0.1105 | 105.0606 | 233.9903 | 180.0105 | 231.8554 | 160.619
2012-10-30 17:42 | 28.1649 | 248.1401 | 194.1212 | 246.0287 | 174.8076 | -1.6084 | 1.2711 | 32.1327 | 40.8444 | 25.2791 | 2.0031 | 32.0586 | 64.2163 | 16.0045 | 14.0946 | -0.1109 | 105.0607 | 234.0455 | 180.0266 | 231.9341 | 160.713
2012-10-30 17:42 | 28.1589 | 248.0754 | 194.1741 | 245.9953 | 174.8172 | -1.5997 | 1.2711 | 32.1324 | 40.8438 | 25.2791 | 2.0031 | 32.0583 | 64.2153 | 16.0045 | 14.0968 | -0.1103 | 105.0591 | 233.9786 | 180.0773 | 231.8985 | 160.7204
2012-10-30 17:42 | 28.1619 | 248.1293 | 194.1928 | 245.9669 | 174.8163 | -1.5991 | 1.2711 | 32.1324 | 40.8436 | 25.2791 | 2.0031 | 32.0583 | 64.215 | 16.0045 | 14.0712 | -0.1102 | 105.0586 | 234.0581 | 180.1216 | 231.8957 | 160.7451
2012-10-30 17:42 | 28.1574 | 248.1561 | 194.1844 | 245.9781 | 174.8079 | -1.6003 | 1.2711 | 32.132 | 40.8431 | 25.2788 | 2.003 | 32.0579 | 64.2134 | 16.0045 | 14.0912 | -0.1103 | 105.0566 | 234.0649 | 180.0932 | 231.8869 | 160.7167
2012-10-30 17:42 | 28.1432 | 248.1263 | 194.2252 | 246.101 | 174.7508 | -1.5997 | 1.2711 | 32.1314 | 40.8417 | 25.2788 | 2.003 | 32.0573 | 64.2113 | 16.0045 | 14.077 | -0.1103 | 105.053 | 234.0493 | 180.1482 | 232.024 | 160.6738
2012-10-30 17:42 | 28.1514 | 248.1305 | 194.2059 | 246.0661 | 174.7863 | -1.6003 | 1.268 | 32.0543 | 40.6457 | 25.2789 | 1.9982 | 31.9806 | 63.9042 | 16.0046 | 14.073 | -0.1103 | 104.5498 | 234.0575 | 180.1329 | 231.9931 | 160.7133
2012-10-30 17:42 | 28.1529 | 248.1242 | 194.2271 | 246.0598 | 174.8073 | -1.6014 | 1.2649 | 31.9725 | 40.4412 | 25.2772 | 1.9931 | 31.8986 | 63.5766 | 16.0047 | 14.0622 | -0.1104 | 104.0177 | 234.062 | 180.1649 | 231.9976 | 160.7451
2012-10-30 17:42 | 28.1507 | 248.1338 | 194.2052 | 246.0262 | 174.9847 | -1.602 | 1.2642 | 31.9544 | 40.3951 | 25.2773 | 1.992 | 31.8806 | 63.5051 | 16.0046 | 14.0763 | -0.1104 | 103.9002 | 234.0575 | 180.1289 | 231.9499 | 160.9084
2012-10-30 17:43 | 28.1529 | 248.1242 | 194.2781 | 246.0206 | 174.9322 | -1.6026 | 1.2638 | 31.9456 | 40.3734 | 25.2771 | 1.9914 | 31.872 | 63.4707 | 16.0046 | 14.0704 | -0.1105 | 103.8441 | 234.0538 | 180.2077 | 231.9502 | 160.8618
2012-10-30 17:43 | 28.1514 | 248.1227 | 194.2648 | 246.0465 | 174.8254 | -1.6009 | 1.2636 | 31.9414 | 40.3628 | 25.2771 | 1.9911 | 31.8678 | 63.4534 | 16.0048 | 14.1382 | -0.1104 | 103.8162 | 233.9845 | 180.1266 | 231.9083 | 160.6872
2012-10-30 17:43 | 28.1507 | 248.1063 | 194.2209 | 246.0537 | 174.6256 | -1.6043 | 1.2637 | 31.9428 | 40.366 | 25.2772 | 1.9912 | 31.8692 | 63.459 | 16.0048 | 14.0723 | -0.1106 | 103.825 | 234.034 | 180.1486 | 231.9814 | 160.5533
2012-10-30 17:43 | 28.1537 | 248.1172 | 194.2514 | 246.0449 | 174.5077 | -1.5968 | 1.2637 | 31.9443 | 40.3694 | 25.2775 | 1.9913 | 31.8706 | 63.464 | 16.0049 | 14.0549 | -0.1101 | 103.8334 | 234.0623 | 180.1965 | 231.99 | 160.4528
2012-10-30 17:43 | 28.1499 | 248.1408 | 194.2123 | 245.9784 | 174.7068 | -1.6026 | 1.2638 | 31.9453 | 40.3727 | 25.2771 | 1.9914 | 31.8716 | 63.4686 | 16.0047 | 14.0837 | -0.1105 | 103.8413 | 234.0571 | 180.1286 | 231.8947 | 160.6231
2012-10-30 17:43 | 28.1552 | 248.0913 | 194.2176 | 246.0543 | 174.4545 | -1.5991 | 1.2639 | 31.9475 | 40.3776 | 25.2775 | 1.9915 | 31.8738 | 63.4773 | 16.0048 | 14.0849 | -0.1102 | 103.8548 | 234.0064 | 180.1327 | 231.9694 | 160.3696
2012-10-30 17:43 | 28.1559 | 248.1311 | 194.2065 | 246.1059 | 174.2914 | -1.6009 | 1.2638 | 31.9468 | 40.3758 | 25.2775 | 1.9915 | 31.8732 | 63.4741 | 16.005 | 14.053 | -0.1104 | 103.8499 | 234.0781 | 180.1535 | 232.0529 | 160.2384
2012-10-30 17:43 | 28.1499 | 248.0977 | 194.1887 | 245.9471 | 174.6404 | -1.6026 | 1.2666 | 32.0175 | 40.554 | 25.2779 | 1.9959 | 31.944 | 63.7568 | 16.0048 | 14.0633 | -0.1105 | 104.3109 | 234.0344 | 180.1254 | 231.8838 | 160.5771
2012-10-30 17:43 | 28.1492 | 248.144 | 194.188 | 246.0208 | 174.6007 | -1.6049 | 1.269 | 32.0795 | 40.7096 | 25.2789 | 1.9997 | 32.0055 | 64.003 | 16.0048 | 14.034 | -0.1106 | 104.7126 | 234.11 | 180.154 | 231.9868 | 160.5667
2012-10-30 17:43 | 28.1559 | 248.1037 | 194.2026 | 246.0354 | 174.7167 | -1.5997 | 1.2697 | 32.0956 | 40.7503 | 25.279 | 2.0007 | 32.0212 | 64.066 | 16.0046 | 14.1142 | -0.1103 | 104.8163 | 233.9895 | 180.0884 | 231.9212 | 160.6025
2012-10-30 17:43 | 28.1454 | 248.1089 | 194.2471 | 246.1071 | 174.7881 | -1.6014 | 1.2702 | 32.1092 | 40.7847 | 25.279 | 2.0016 | 32.035 | 64.1214 | 16.0046 | 14.0833 | -0.1104 | 104.9061 | 234.0256 | 180.1638 | 232.0238 | 160.7048
2012-10-30 17:43 | 28.1529 | 248.136 | 194.2821 | 246.048 | 175.0259 | -1.6014 | 1.2706 | 32.1192 | 40.8101 | 25.2791 | 2.0022 | 32.0452 | 64.1619 | 16.0047 | 14.0989 | -0.1104 | 104.972 | 234.0371 | 180.1832 | 231.9491 | 160.927
2012-10-30 17:43 | 28.1432 | 248.1615 | 194.3234 | 246.0344 | 175.0279 | -1.6009 | 1.2708 | 32.126 | 40.8269 | 25.2794 | 2.0027 | 32.0519 | 64.1897 | 16.0045 | 14.0648 | -0.1104 | 105.0166 | 234.0967 | 180.2586 | 231.9696 | 160.9631
2012-10-30 17:43 | 28.1432 | 248.1458 | 194.3391 | 246.0187 | 175.0474 | -1.5974 | 1.271 | 32.1304 | 40.8381 | 25.2794 | 2.003 | 32.0563 | 64.2073 | 16.0045 | 14.0811 | -0.1101 | 105.0454 | 234.0647 | 180.258 | 231.9376 | 160.9663
2012-10-30 17:43 | 28.1567 | 248.2142 | 194.4076 | 245.997 | 174.9048 | -1.598 | 1.2711 | 32.1332 | 40.8452 | 25.2794 | 2.0031 | 32.059 | 64.2182 | 16.0045 | 14.0783 | -0.1102 | 105.0634 | 234.1359 | 180.3293 | 231.9187 | 160.8265
2012-10-30 17:43 | 28.1439 | 248.17 | 194.3948 | 246.1017 | 174.7827 | -1.6032 | 1.2712 | 32.1349 | 40.8498 | 25.2793 | 2.0032 | 32.0606 | 64.2249 | 16.0044 | 14.0451 | -0.1105 | 105.0747 | 234.1249 | 180.3497 | 232.0566 | 160.7376
2012-10-30 17:43 | 28.1477 | 248.1582 | 194.3554 | 246.0193 | 174.8138 | -1.5985 | 1.2712 | 32.1361 | 40.8529 | 25.2792 | 2.0033 | 32.0618 | 64.2301 | 16.0043 | 14.0285 | -0.1102 | 105.083 | 234.1297 | 180.3269 | 231.9908 | 160.7853
2012-10-30 17:43 | 28.1447 | 248.2022 | 194.3721 | 246.0908 | 174.7406 | -1.6003 | 1.2713 | 32.1371 | 40.8549 | 25.2795 | 2.0034 | 32.0629 | 64.2343 | 16.0044 | 14.0255 | -0.1103 | 105.0892 | 234.1767 | 180.3466 | 232.0653 | 160.7151
2012-10-30 17:43 | 28.1537 | 248.2425 | 194.3614 | 246.0292 | 174.8042 | -1.6055 | 1.2713 | 32.1378 | 40.8571 | 25.2793 | 2.0034 | 32.0637 | 64.2374 | 16.0044 | 14.01 | -0.1107 | 105.0946 | 234.2325 | 180.3514 | 232.0192 | 160.7942
2012-10-30 17:43 | 28.1417 | 248.1992 | 194.3298 | 246.0172 | 174.7961 | -1.6003 | 1.2713 | 32.1384 | 40.8576 | 25.2799 | 2.0035 | 32.0642 | 64.2392 | 16.0045 | 13.998 | -0.1103 | 105.0968 | 234.2012 | 180.3318 | 232.0192 | 160.7981
2012-10-30 17:43 | 28.1432 | 248.2555 | 194.3352 | 246.1089 | 174.7469 | -1.6061 | 1.2713 | 32.1391 | 40.8596 | 25.2798 | 2.0035 | 32.0648 | 64.2424 | 16.0043 | 14.0362 | -0.1107 | 105.102 | 234.2193 | 180.299 | 232.0727 | 160.7107
2012-10-30 17:43 | 28.1514 | 248.2989 | 194.3473 | 246.1327 | 174.8995 | -1.6009 | 1.2714 | 32.1396 | 40.8611 | 25.2796 | 2.0035 | 32.0653 | 64.2443 | 16.0043 | 14.0526 | -0.1104 | 105.1054 | 234.2463 | 180.2947 | 232.0801 | 160.8469
2012-10-30 17:43 | 28.1537 | 248.2425 | 194.3614 | 246.0959 | 174.9096 | -1.602 | 1.2714 | 32.14 | 40.8618 | 25.2798 | 2.0036 | 32.0658 | 64.2464 | 16.0042 | 14.0834 | -0.1104 | 105.1082 | 234.1591 | 180.278 | 232.0125 | 160.8262
2012-10-30 17:43 | 28.1604 | 248.2805 | 194.3996 | 246.1104 | 175.0256 | -1.6009 | 1.2714 | 32.1404 | 40.8624 | 25.28 | 2.0036 | 32.0662 | 64.2482 | 16.0042 | 14.0656 | -0.1104 | 105.1106 | 234.2149 | 180.334 | 232.0448 | 160.96
2012-10-30 17:44 | 28.1589 | 248.2399 | 194.4334 | 246.1559 | 175.0592 | -1.6009 | 1.2714 | 32.1402 | 40.8627 | 25.2796 | 2.0036 | 32.066 | 64.2479 | 16.004 | 14.0519 | -0.1104 | 105.1106 | 234.188 | 180.3815 | 232.104 | 161.0073
2012-10-30 17:44 | 28.1552 | 248.2675 | 194.4179 | 246.1836 | 175.2155 | -1.602 | 1.2714 | 32.1401 | 40.8622 | 25.2798 | 2.0036 | 32.0659 | 64.2478 | 16.0041 | 14.036 | -0.1104 | 105.1099 | 234.2315 | 180.3819 | 232.1476 | 161.1795
2012-10-30 17:44 | 28.1477 | 248.2874 | 194.4812 | 246.1173 | 175.1807 | -1.6003 | 1.2714 | 32.1401 | 40.8618 | 25.2799 | 2.0036 | 32.0657 | 64.2472 | 16.004 | 14.0529 | -0.1103 | 105.109 | 234.2345 | 180.4283 | 232.0644 | 161.1278
2012-10-30 17:44 | 28.1537 | 248.2934 | 194.5068 | 246.0959 | 175.1047 | -1.5997 | 1.2714 | 32.1399 | 40.8617 | 25.2797 | 2.0036 | 32.0656 | 64.2465 | 16.004 | 14.0181 | -0.1103 | 105.1082 | 234.2753 | 180.4887 | 232.0778 | 161.0866
2012-10-30 17:44 | 28.1582 | 248.2901 | 194.5113 | 246.1004 | 175.0624 | -1.6003 | 1.2713 | 32.1398 | 40.8606 | 25.2802 | 2.0036 | 32.0654 | 64.246 | 16.004 | 14.043 | -0.1103 | 105.1066 | 234.2471 | 180.4683 | 232.0574 | 161.0194
2012-10-30 17:44 | 28.1574 | 248.2814 | 194.5223 | 246.1505 | 174.9445 | -1.6009 | 1.2714 | 32.14 | 40.8617 | 25.2798 | 2.0036 | 32.0657 | 64.247 | 16.004 | 14.0463 | -0.1104 | 105.1087 | 234.2351 | 180.476 | 232.1042 | 160.8982
2012-10-30 17:44 | 28.1529 | 248.3396 | 194.4903 | 246.1656 | 175.0103 | -1.5991 | 1.2714 | 32.1399 | 40.8615 | 25.2799 | 2.0036 | 32.0657 | 64.2477 | 16.0039 | 14.0418 | -0.1102 | 105.1092 | 234.2978 | 180.4485 | 232.1238 | 160.9685
2012-10-30 17:44 | 28.1552 | 248.3341 | 194.4965 | 246.2188 | 175.2546 | -1.6014 | 1.2713 | 32.1397 | 40.8606 | 25.2801 | 2.0036 | 32.0655 | 64.2467 | 16.0039 | 14.0482 | -0.1104 | 105.1074 | 234.2859 | 180.4483 | 232.1706 | 161.2064
2012-10-30 17:44 | 28.1619 | 248.3447 | 194.5307 | 246.1001 | 175.2261 | -1.5945 | 1.2713 | 32.1399 | 40.861 | 25.2802 | 2.0036 | 32.0656 | 64.247 | 16.0039 | 14.0875 | -0.1099 | 105.1079 | 234.2572 | 180.4432 | 232.0126 | 161.1386
2012-10-30 17:44 | 28.1567 | 248.3747 | 194.5255 | 246.2125 | 175.057 | -1.602 | 1.2713 | 32.1397 | 40.8609 | 25.28 | 2.0036 | 32.0653 | 64.2461 | 16.0038 | 14.0497 | -0.1104 | 105.107 | 234.325 | 180.4758 | 232.1628 | 161.0073
2012-10-30 17:44 | 28.1567 | 248.3473 | 194.5334 | 246.283 | 174.979 | -1.6032 | 1.2713 | 32.1393 | 40.8594 | 25.2802 | 2.0036 | 32.0649 | 64.2445 | 16.0038 | 14.0375 | -0.1105 | 105.1039 | 234.3098 | 180.4959 | 232.2455 | 160.9415
2012-10-30 17:44 | 28.1664 | 248.3727 | 194.5902 | 246.1987 | 175.0706 | -1.6026 | 1.2713 | 32.139 | 40.8596 | 25.2797 | 2.0036 | 32.0649 | 64.2447 | 16.0037 | 14.0594 | -0.1105 | 105.1043 | 234.3133 | 180.5308 | 232.1393 | 161.0112
2012-10-30 17:44 | 28.1627 | 248.4042 | 194.5433 | 246.2224 | 175.2972 | -1.5974 | 1.2713 | 32.139 | 40.8587 | 25.2801 | 2.0036 | 32.0648 | 64.244 | 16.0038 | 14.0557 | -0.1101 | 105.1027 | 234.3485 | 180.4876 | 232.1667 | 161.2415
2012-10-30 17:44 | 28.1664 | 248.3727 | 194.5902 | 246.1987 | 175.2033 | -1.5968 | 1.2713 | 32.1389 | 40.859 | 25.2799 | 2.0036 | 32.0646 | 64.2434 | 16.0038 | 14.0268 | -0.1101 | 105.1024 | 234.3459 | 180.5634 | 232.1719 | 161.1765
2012-10-30 17:44 | 28.1559 | 248.4013 | 194.5876 | 246.2587 | 175.0523 | -1.5985 | 1.2713 | 32.1384 | 40.8573 | 25.2801 | 2.0035 | 32.0641 | 64.2414 | 16.0038 | 14.053 | -0.1102 | 105.0987 | 234.3483 | 180.5346 | 232.2057 | 160.9993
2012-10-30 17:44 | 28.1612 | 248.4027 | 194.5772 | 246.264 | 174.9249 | -1.598 | 1.2713 | 32.1382 | 40.8571 | 25.2799 | 2.0035 | 32.0639 | 64.2413 | 16.0036 | 14.0175 | -0.1102 | 105.0984 | 234.3852 | 180.5597 | 232.2465 | 160.9074
2012-10-30 17:44 | 28.1694 | 248.4148 | 194.5814 | 246.3192 | 174.9956 | -1.6009 | 1.2713 | 32.1378 | 40.8558 | 25.2801 | 2.0035 | 32.0635 | 64.2399 | 16.0036 | 14.0828 | -0.1104 | 105.0957 | 234.332 | 180.4986 | 232.2364 | 160.9128
2012-10-30 17:44 | 28.1724 | 248.4296 | 194.6001 | 246.1929 | 175.0259 | -1.6014 | 1.2713 | 32.1376 | 40.8552 | 25.2801 | 2.0035 | 32.0633 | 64.239 | 16.0036 | 14.0736 | -0.1104 | 105.0942 | 234.356 | 180.5265 | 232.1193 | 160.9523
2012-10-30 17:44 | 28.1739 | 248.4193 | 194.5663 | 246.2414 | 174.9688 | -1.5991 | 1.2712 | 32.1373 | 40.8545 | 25.2801 | 2.0035 | 32.063 | 64.2382 | 16.0035 | 14.0995 | -0.1102 | 105.0927 | 234.3198 | 180.4668 | 232.1419 | 160.8693
2012-10-30 17:44 | 28.1762 | 248.4529 | 194.5764 | 246.3417 | 174.9204 | -1.5985 | 1.2712 | 32.1371 | 40.8538 | 25.2802 | 2.0035 | 32.0629 | 64.2375 | 16.0036 | 14.1304 | -0.1102 | 105.0913 | 234.3225 | 180.446 | 232.2113 | 160.79
2012-10-30 17:44 | 28.1776 | 248.4387 | 194.5621 | 246.2843 | 175.0428 | -1.6049 | 1.2712 | 32.1367 | 40.8523 | 25.2805 | 2.0035 | 32.0625 | 64.2363 | 16.0034 | 14.0788 | -0.1106 | 105.0886 | 234.3599 | 180.4833 | 232.2055 | 160.964
2012-10-30 17:44 | 28.1806 | 248.4613 | 194.5416 | 246.3108 | 175.1473 | -1.5985 | 1.2712 | 32.1366 | 40.8528 | 25.28 | 2.0035 | 32.0624 | 64.2358 | 16.0035 | 14.0736 | -0.1102 | 105.0886 | 234.3877 | 180.468 | 232.2372 | 161.0737
2012-10-30 17:44 | 28.1829 | 248.4988 | 194.5989 | 246.27 | 175.0325 | -1.6049 | 1.2712 | 32.1364 | 40.8519 | 25.2803 | 2.0034 | 32.062 | 64.2341 | 16.0036 | 14.0677 | -0.1106 | 105.086 | 234.4311 | 180.5312 | 232.2023 | 160.9648
2012-10-30 17:44 | 28.1859 | 248.5057 | 194.5744 | 246.3632 | 174.895 | -1.6003 | 1.2712 | 32.1361 | 40.8515 | 25.2801 | 2.0035 | 32.0619 | 64.2346 | 16.0033 | 14.0667 | -0.1103 | 105.0861 | 234.439 | 180.5077 | 232.2965 | 160.8283
2012-10-30 17:44 | 28.1859 | 248.4509 | 194.5626 | 246.4023 | 175.0238 | -1.6043 | 1.2712 | 32.1356 | 40.8504 | 25.28 | 2.0034 | 32.0614 | 64.2326 | 16.0033 | 14.0707 | -0.1106 | 105.083 | 234.3802 | 180.4919 | 232.3316 | 160.9531
2012-10-30 17:44 | 28.1844 | 248.516 | 194.5611 | 246.3186 | 175.1472 | -1.602 | 1.2712 | 32.1355 | 40.8497 | 25.2802 | 2.0034 | 32.0612 | 64.2312 | 16.0034 | 14.0407 | -0.1104 | 105.0808 | 234.4753 | 180.5204 | 232.2779 | 161.1065
2012-10-30 17:45 | 28.1874 | 248.5151 | 194.5641 | 246.3137 | 175.1424 | -1.5974 | 1.2712 | 32.1351 | 40.8486 | 25.2803 | 2.0034 | 32.0609 | 64.2307 | 16.0033 | 14.0478 | -0.1101 | 105.0793 | 234.4673 | 180.5163 | 232.2659 | 161.0946
2012-10-30 17:45 | 28.1866 | 248.5182 | 194.5436 | 246.3403 | 175.0947 | -1.5968 | 1.2711 | 32.1349 | 40.8479 | 25.2804 | 2.0034 | 32.0606 | 64.2295 | 16.0033 | 14.0633 | -0.1101 | 105.0774 | 234.4549 | 180.4803 | 232.277 | 161.0314
2012-10-30 17:45 | 28.1874 | 248.5307 | 194.5326 | 246.3686 | 175.1345 | -1.5991 | 1.2711 | 32.1348 | 40.8477 | 25.2804 | 2.0034 | 32.0605 | 64.2293 | 16.0032 | 14.0396 | -0.1102 | 105.077 | 234.4911 | 180.493 | 232.329 | 161.0949
2012-10-30 17:45 | 28.1889 | 248.5635 | 194.5852 | 246.4014 | 175.2063 | -1.5974 | 1.2711 | 32.134 | 40.8462 | 25.2801 | 2.0033 | 32.0598 | 64.2265 | 16.0032 | 14.0329 | -0.1101 | 105.0726 | 234.5306 | 180.5523 | 232.3685 | 161.1734
2012-10-30 17:45 | 28.1956 | 248.5507 | 194.5998 | 246.365 | 175.1037 | -1.5991 | 1.2711 | 32.1334 | 40.8446 | 25.2802 | 2.0033 | 32.0593 | 64.2243 | 16.0032 | 14.06 | -0.1102 | 105.0689 | 234.4907 | 180.5398 | 232.305 | 161.0437
2012-10-30 17:45 | 28.1979 | 248.553 | 194.5706 | 246.3203 | 175.028 | -1.598 | 1.2711 | 32.1333 | 40.8442 | 25.2802 | 2.0033 | 32.0591 | 64.2235 | 16.0032 | 14.0705 | -0.1102 | 105.0678 | 234.4825 | 180.5001 | 232.2498 | 160.9575
2012-10-30 17:45 | 28.1956 | 248.5585 | 194.5762 | 246.3258 | 174.8891 | -1.6049 | 1.2711 | 32.1326 | 40.8422 | 25.2803 | 2.0033 | 32.0584 | 64.221 | 16.0032 | 14.0029 | -0.1106 | 105.0633 | 234.5556 | 180.5733 | 232.3229 | 160.8862
2012-10-30 17:45 | 28.1964 | 248.5358 | 194.5534 | 246.4677 | 174.9289 | -1.5991 | 1.2711 | 32.1323 | 40.8419 | 25.28 | 2.0032 | 32.0581 | 64.2202 | 16.0031 | 14.0486 | -0.1102 | 105.0622 | 234.4872 | 180.5048 | 232.4191 | 160.8803
2012-10-30 17:45 | 28.1971 | 248.5482 | 194.562 | 246.3861 | 175.0428 | -1.6043 | 1.271 | 32.1321 | 40.8415 | 25.28 | 2.0032 | 32.0579 | 64.2191 | 16.0032 | 14.0207 | -0.1106 | 105.0606 | 234.5275 | 180.5413 | 232.3654 | 161.0221
2012-10-30 17:45 | 28.2009 | 248.5599 | 194.5579 | 246.3272 | 174.9958 | -1.6003 | 1.271 | 32.1316 | 40.8392 | 25.2806 | 2.0032 | 32.0574 | 64.2173 | 16.0031 | 14.0531 | -0.1103 | 105.0565 | 234.5068 | 180.5048 | 232.2741 | 160.9427
2012-10-30 17:45 | 28.2076 | 248.5705 | 194.5096 | 246.3613 | 174.9869 | -1.5957 | 1.271 | 32.1315 | 40.8392 | 25.2804 | 2.0032 | 32.0571 | 64.2171 | 16.0029 | 13.9945 | -0.11 | 105.0563 | 234.576 | 180.5151 | 232.3668 | 160.9924
2012-10-30 17:45 | 28.2106 | 248.55 | 194.5401 | 246.4427 | 175.0172 | -1.6072 | 1.271 | 32.131 | 40.8382 | 25.2802 | 2.0032 | 32.0569 | 64.2155 | 16.003 | 14.0302 | -0.1108 | 105.0537 | 234.5198 | 180.5099 | 232.4125 | 160.987
2012-10-30 17:45 | 28.2039 | 248.5472 | 194.5374 | 246.4125 | 175.0574 | -1.6009 | 1.271 | 32.1305 | 40.8364 | 25.2805 | 2.0031 | 32.0562 | 64.2134 | 16.0029 | 13.999 | -0.1104 | 105.0498 | 234.5482 | 180.5384 | 232.4135 | 161.0584
2012-10-30 17:45 | 28.2113 | 248.5899 | 194.5605 | 246.3729 | 175.0882 | -1.5974 | 1.2689 | 32.0772 | 40.7021 | 25.2799 | 1.9998 | 32.0026 | 63.9983 | 16.003 | 14.0227 | -0.1101 | 104.7004 | 234.5672 | 180.5378 | 232.3502 | 161.0655
2012-10-30 17:45 | 28.2061 | 248.5847 | 194.5199 | 246.3716 | 175.0166 | -1.6003 | 1.2649 | 31.9742 | 40.4432 | 25.2786 | 1.9934 | 31.901 | 63.5928 | 16.003 | 14.042 | -0.1103 | 104.036 | 234.5427 | 180.4779 | 232.3296 | 160.9746
2012-10-30 17:45 | 28.2136 | 248.5882 | 194.5392 | 246.5202 | 175.0437 | -1.6014 | 1.2646 | 31.9684 | 40.4286 | 25.2786 | 1.993 | 31.8945 | 63.566 | 16.0032 | 14.0454 | -0.1104 | 103.9946 | 234.5428 | 180.4938 | 232.4748 | 160.9983
2012-10-30 17:45 | 28.2076 | 248.5548 | 194.5175 | 246.3966 | 175.0533 | -1.6014 | 1.2641 | 31.9546 | 40.3933 | 25.2788 | 1.9921 | 31.8805 | 63.5102 | 16.0032 | 14.0108 | -0.1104 | 103.9034 | 234.544 | 180.5067 | 232.3858 | 161.0425
2012-10-30 17:45 | 28.2099 | 248.5219 | 194.5512 | 246.4185 | 174.5912 | -1.6003 | 1.2637 | 31.9454 | 40.3709 | 25.2783 | 1.9916 | 31.8715 | 63.4747 | 16.0031 | 14.005 | -0.1103 | 103.8456 | 234.5169 | 180.5462 | 232.4135 | 160.5862
2012-10-30 17:45 | 28.2143 | 248.5459 | 194.4456 | 246.4542 | 174.4707 | -1.6014 | 1.2636 | 31.9407 | 40.3588 | 25.2785 | 1.9913 | 31.867 | 63.4566 | 16.0031 | 14.0543 | -0.1104 | 103.8154 | 234.4916 | 180.3913 | 232.3999 | 160.4164
2012-10-30 17:45 | 28.2248 | 248.5368 | 194.3815 | 246.5353 | 174.809 | -1.6037 | 1.2635 | 31.9391 | 40.3548 | 25.2784 | 1.9912 | 31.8653 | 63.4497 | 16.0032 | 14.0362 | -0.1106 | 103.8045 | 234.5006 | 180.3453 | 232.4991 | 160.7728
2012-10-30 17:45 | 28.2218 | 248.5142 | 194.3902 | 246.4069 | 174.6928 | -1.6061 | 1.2635 | 31.9396 | 40.3562 | 25.2783 | 1.9912 | 31.8657 | 63.4508 | 16.0033 | 14.074 | -0.1107 | 103.807 | 234.4402 | 180.3162 | 232.3329 | 160.6188
2012-10-30 17:45 | 28.2211 | 248.49 | 194.4092 | 246.3278 | 174.8833 | -1.6014 | 1.2635 | 31.9401 | 40.3574 | 25.2784 | 1.9912 | 31.8663 | 63.4533 | 16.0033 | 14.0692 | -0.1104 | 103.8107 | 234.4208 | 180.34 | 232.2586 | 160.8141
2012-10-30 17:45 | 28.2084 | 248.4851 | 194.3886 | 246.417 | 174.9487 | -1.602 | 1.2637 | 31.945 | 40.3698 | 25.2784 | 1.9915 | 31.8713 | 63.4729 | 16.0033 | 14.0647 | -0.1104 | 103.8427 | 234.4204 | 180.3239 | 232.3523 | 160.884
2012-10-30 17:45 | 28.2158 | 248.4573 | 194.4117 | 246.5184 | 174.9873 | -1.6026 | 1.2638 | 31.9468 | 40.3741 | 25.2785 | 1.9916 | 31.873 | 63.4793 | 16.0034 | 14.0762 | -0.1105 | 103.8534 | 234.3811 | 180.3355 | 232.4422 | 160.9111
2012-10-30 17:45 | 28.2248 | 248.4741 | 194.4168 | 246.4295 | 175.0158 | -1.6009 | 1.2681 | 32.0555 | 40.6483 | 25.2792 | 1.9984 | 31.9815 | 63.9131 | 16.0033 | 14.1056 | -0.1104 | 104.5614 | 234.3685 | 180.3112 | 232.3239 | 160.9102
2012-10-30 17:46 | 28.2158 | 248.4965 | 194.451 | 246.3382 | 174.8429 | -1.5939 | 1.2692 | 32.0847 | 40.7211 | 25.2799 | 2.0002 | 32.0103 | 64.0279 | 16.0033 | 14.0721 | -0.1099 | 104.749 | 234.4244 | 180.3789 | 232.2661 | 160.7708
2012-10-30 17:46 | 28.2211 | 248.4939 | 194.5035 | 246.3827 | 174.9341 | -1.6026 | 1.2698 | 32.101 | 40.7624 | 25.2801 | 2.0013 | 32.0268 | 64.0947 | 16.0032 | 14.0733 | -0.1105 | 104.8571 | 234.4206 | 180.4302 | 232.3094 | 160.8608
2012-10-30 17:46 | 28.2248 | 248.5055 | 194.5308 | 246.4334 | 175.0432 | -1.6014 | 1.2703 | 32.1148 | 40.7965 | 25.2806 | 2.0021 | 32.0404 | 64.1493 | 16.0031 | 14.0484 | -0.1104 | 104.9457 | 234.4571 | 180.4824 | 232.385 | 160.9948
2012-10-30 17:46 | 28.2218 | 248.5064 | 194.4727 | 246.403 | 175.048 | -1.6014 | 1.2706 | 32.1231 | 40.817 | 25.2809 | 2.0027 | 32.0487 | 64.1828 | 16.003 | 14.074 | -0.1104 | 104.9999 | 234.4324 | 180.3987 | 232.329 | 160.974
2012-10-30 17:46 | 28.2263 | 248.5383 | 194.5401 | 246.2664 | 175.0407 | -1.5997 | 1.2709 | 32.128 | 40.83 | 25.2807 | 2.003 | 32.0538 | 64.2029 | 16.0031 | 14.0622 | -0.1103 | 105.0329 | 234.4761 | 180.4779 | 232.2042 | 160.9785
2012-10-30 17:46 | 28.2226 | 248.5268 | 194.5403 | 246.3685 | 175.0761 | -1.6055 | 1.271 | 32.1317 | 40.8393 | 25.2808 | 2.0032 | 32.0575 | 64.2183 | 16.0029 | 14.0626 | -0.1107 | 105.0576 | 234.4642 | 180.4777 | 232.3059 | 161.0135
2012-10-30 17:46 | 28.2143 | 248.542 | 194.5635 | 246.4151 | 175.1029 | -1.6026 | 1.2711 | 32.1336 | 40.8445 | 25.2805 | 2.0033 | 32.0594 | 64.2258 | 16.003 | 14.0502 | -0.1105 | 105.0703 | 234.4918 | 180.5133 | 232.3649 | 161.0527
2012-10-30 17:46 | 28.2181 | 248.5732 | 194.5555 | 246.364 | 175.2238 | -1.5985 | 1.2711 | 32.1357 | 40.8491 | 25.2809 | 2.0035 | 32.0613 | 64.2341 | 16.0028 | 14.0744 | -0.1102 | 105.0832 | 234.4988 | 180.4811 | 232.2896 | 161.1494
2012-10-30 17:46 | 28.2084 | 248.5556 | 194.6047 | 246.2681 | 175.2258 | -1.6049 | 1.2712 | 32.1371 | 40.8521 | 25.2812 | 2.0035 | 32.0626 | 64.2385 | 16.003 | 14.0728 | -0.1106 | 105.0907 | 234.4828 | 180.5319 | 232.1953 | 161.153
2012-10-30 17:46 | 28.2263 | 248.5853 | 194.6187 | 246.4153 | 175.1344 | -1.5997 | 1.2712 | 32.1384 | 40.8555 | 25.2812 | 2.0036 | 32.064 | 64.2443 | 16.003 | 14.0703 | -0.1103 | 105.0998 | 234.515 | 180.5484 | 232.345 | 161.0641
2012-10-30 17:46 | 28.2188 | 248.5856 | 194.6426 | 246.4392 | 175.1347 | -1.6003 | 1.2713 | 32.1388 | 40.8571 | 25.2809 | 2.0036 | 32.0643 | 64.2455 | 16.003 | 14.0873 | -0.1103 | 105.1025 | 234.4983 | 180.5553 | 232.3519 | 161.0474
2012-10-30 17:46 | 28.2263 | 248.5853 | 194.654 | 246.3565 | 175.3764 | -1.5991 | 1.2713 | 32.1392 | 40.8579 | 25.281 | 2.0037 | 32.0649 | 64.248 | 16.0029 | 14.0418 | -0.1102 | 105.1059 | 234.5435 | 180.6122 | 232.3147 | 161.3346
2012-10-30 17:46 | 28.2173 | 248.6233 | 194.7393 | 246.4024 | 175.3596 | -1.602 | 1.2713 | 32.1396 | 40.8587 | 25.281 | 2.0037 | 32.0652 | 64.2502 | 16.0027 | 14.0532 | -0.1104 | 105.109 | 234.5701 | 180.6861 | 232.3492 | 161.3064
2012-10-30 17:46 | 28.2256 | 248.6707 | 194.728 | 246.5361 | 175.2664 | -1.6014 | 1.2713 | 32.1401 | 40.8601 | 25.281 | 2.0038 | 32.0656 | 64.2525 | 16.0026 | 14.0696 | -0.1104 | 105.1126 | 234.6011 | 180.6584 | 232.4665 | 161.1968
2012-10-30 17:46 | 28.2151 | 248.6837 | 194.7725 | 246.4237 | 175.2286 | -1.6014 | 1.2713 | 32.1403 | 40.8608 | 25.281 | 2.0038 | 32.0659 | 64.2536 | 16.0026 | 14.0591 | -0.1104 | 105.1143 | 234.6246 | 180.7134 | 232.3646 | 161.1695
2012-10-30 17:46 | 28.2151 | 248.7072 | 194.7607 | 246.4002 | 175.0569 | -1.598 | 1.2713 | 32.1402 | 40.8603 | 25.2811 | 2.0038 | 32.0657 | 64.253 | 16.0025 | 14.0836 | -0.1102 | 105.1133 | 234.6236 | 180.6771 | 232.3166 | 160.9733
2012-10-30 17:46 | 28.2151 | 248.6798 | 194.745 | 246.5217 | 174.9476 | -1.6003 | 1.2713 | 32.1405 | 40.861 | 25.2811 | 2.0038 | 32.0661 | 64.2547 | 16.0025 | 14.0591 | -0.1103 | 105.1157 | 234.6207 | 180.6859 | 232.4626 | 160.8885
2012-10-30 17:46 | 28.2241 | 248.7005 | 194.754 | 246.4915 | 175.1595 | -1.6009 | 1.2713 | 32.1407 | 40.8617 | 25.281 | 2.0038 | 32.0663 | 64.2554 | 16.0025 | 14.0804 | -0.1104 | 105.1171 | 234.6201 | 180.6736 | 232.4111 | 161.0791
2012-10-30 17:46 | 28.2136 | 248.6626 | 194.6885 | 246.4026 | 175.231 | -1.6026 | 1.2713 | 32.1404 | 40.8609 | 25.281 | 2.0038 | 32.0661 | 64.2547 | 16.0025 | 14.0658 | -0.1105 | 105.1156 | 234.5968 | 180.6227 | 232.3368 | 161.1652
2012-10-30 17:46 | 28.2069 | 248.6677 | 194.725 | 246.5017 | 175.154 | -1.6014 | 1.2713 | 32.1405 | 40.8609 | 25.2812 | 2.0038 | 32.0661 | 64.2546 | 16.0025 | 14.0509 | -0.1104 | 105.1155 | 234.6168 | 180.6741 | 232.4508 | 161.1031
2012-10-30 17:46 | 28.2151 | 248.6837 | 194.7254 | 246.5491 | 175.213 | -1.598 | 1.2713 | 32.1406 | 40.8615 | 25.281 | 2.0038 | 32.0662 | 64.255 | 16.0025 | 14.0673 | -0.1102 | 105.1165 | 234.6164 | 180.6581 | 232.4818 | 161.1457
2012-10-30 17:46 | 28.2188 | 248.6952 | 194.7369 | 246.4117 | 175.3104 | -1.6049 | 1.2713 | 32.1405 | 40.8607 | 25.2813 | 2.0038 | 32.0661 | 64.2552 | 16.0024 | 14.0792 | -0.1106 | 105.1159 | 234.616 | 180.6577 | 232.3325 | 161.2312
2012-10-30 17:47 | 28.2136 | 248.7253 | 194.8024 | 246.5202 | 175.3754 | -1.6026 | 1.2713 | 32.1408 | 40.8615 | 25.2813 | 2.0039 | 32.0664 | 64.2564 | 16.0023 | 14.0576 | -0.1105 | 105.1179 | 234.6677 | 180.7448 | 232.4626 | 161.3178
2012-10-30 17:47 | 28.2099 | 248.6707 | 194.7712 | 246.5987 | 175.2117 | -1.6014 | 1.2713 | 32.1401 | 40.8599 | 25.2811 | 2.0038 | 32.0658 | 64.2542 | 16.0023 | 14.0743 | -0.1104 | 105.1141 | 234.5964 | 180.6969 | 232.5244 | 161.1374
2012-10-30 17:47 | 28.2113 | 248.7347 | 194.7412 | 246.5296 | 175.1897 | -1.6009 | 1.2713 | 32.1396 | 40.8587 | 25.2812 | 2.0038 | 32.0652 | 64.2519 | 16.0023 | 14.0676 | -0.1104 | 105.1107 | 234.6671 | 180.6736 | 232.462 | 161.1221
2012-10-30 17:47 | 28.2128 | 248.7676 | 194.7466 | 246.5233 | 175.1599 | -1.5985 | 1.2712 | 32.1387 | 40.8561 | 25.2813 | 2.0038 | 32.0644 | 64.2492 | 16.0021 | 14.0854 | -0.1102 | 105.1053 | 234.6822 | 180.6612 | 232.4379 | 161.0745
2012-10-30 17:47 | 28.2091 | 248.756 | 194.6958 | 246.6175 | 175.2343 | -1.5939 | 1.2713 | 32.1387 | 40.8564 | 25.2811 | 2.0038 | 32.0643 | 64.249 | 16.0021 | 14.0695 | -0.1099 | 105.1054 | 234.6865 | 180.6263 | 232.548 | 161.1648
2012-10-30 17:47 | 28.2099 | 248.7764 | 194.7634 | 246.6144 | 175.3834 | -1.5974 | 1.2712 | 32.1385 | 40.856 | 25.2811 | 2.0037 | 32.0641 | 64.2483 | 16.0021 | 14.0988 | -0.1101 | 105.1042 | 234.6776 | 180.6646 | 232.5156 | 161.2846
2012-10-30 17:47 | 28.2143 | 248.7612 | 194.7992 | 246.5365 | 175.4308 | -1.6003 | 1.2712 | 32.1385 | 40.8555 | 25.2814 | 2.0037 | 32.0642 | 64.2487 | 16.0021 | 14.1073 | -0.1103 | 105.1042 | 234.6539 | 180.6919 | 232.4292 | 161.3235
2012-10-30 17:47 | 28.2061 | 248.7961 | 194.8185 | 246.6028 | 175.3016 | -1.6055 | 1.2712 | 32.1382 | 40.8553 | 25.2811 | 2.0037 | 32.0638 | 64.2471 | 16.0021 | 14.0787 | -0.1107 | 105.1024 | 234.7174 | 180.7398 | 232.5241 | 161.2229
2012-10-30 17:47 | 28.2084 | 248.7671 | 194.8562 | 246.6442 | 175.2882 | -1.5951 | 1.2712 | 32.1378 | 40.8541 | 25.2812 | 2.0037 | 32.0634 | 64.2459 | 16.002 | 14.0892 | -0.11 | 105.0999 | 234.6779 | 180.767 | 232.555 | 161.199
2012-10-30 17:47 | 28.2166 | 248.7596 | 194.8329 | 246.5819 | 175.2106 | -1.6032 | 1.2712 | 32.1377 | 40.8536 | 25.2812 | 2.0037 | 32.0633 | 64.2452 | 16.0021 | 14.0974 | -0.1105 | 105.0988 | 234.6622 | 180.7355 | 232.4845 | 161.1132
2012-10-30 17:47 | 28.2054 | 248.8032 | 194.8178 | 246.5629 | 175.1447 | -1.598 | 1.2712 | 32.1374 | 40.8525 | 25.2814 | 2.0037 | 32.063 | 64.2444 | 16.002 | 14.0739 | -0.1102 | 105.0969 | 234.7293 | 180.7439 | 232.489 | 161.0708
2012-10-30 17:47 | 28.2024 | 248.765 | 194.7519 | 246.744 | 175.0949 | -1.5974 | 1.2712 | 32.1367 | 40.8514 | 25.2811 | 2.0037 | 32.0624 | 64.2425 | 16.0019 | 14.0342 | -0.1101 | 105.0938 | 234.7308 | 180.7177 | 232.7098 | 161.0607
2012-10-30 17:47 | 28.2084 | 248.8141 | 194.7894 | 246.5698 | 175.2531 | -1.6032 | 1.2712 | 32.1366 | 40.8507 | 25.2813 | 2.0037 | 32.0623 | 64.242 | 16.0019 | 14.0647 | -0.1105 | 105.0927 | 234.7494 | 180.7247 | 232.5051 | 161.1884
2012-10-30 17:47 | 28.2084 | 248.8101 | 194.813 | 246.5933 | 175.2375 | -1.6032 | 1.2711 | 32.1359 | 40.8491 | 25.2813 | 2.0036 | 32.0615 | 64.2387 | 16.0019 | 14.0851 | -0.1105 | 105.0878 | 234.725 | 180.7279 | 232.5082 | 161.1524
2012-10-30 17:47 | 28.2046 | 248.7946 | 194.817 | 246.6718 | 175.1127 | -1.5968 | 1.2711 | 32.1354 | 40.8481 | 25.2811 | 2.0036 | 32.0611 | 64.2367 | 16.002 | 14.0813 | -0.1101 | 105.0849 | 234.7133 | 180.7357 | 232.5905 | 161.0314
2012-10-30 17:47 | 28.2113 | 248.7935 | 194.7962 | 246.7216 | 175.1936 | -1.602 | 1.2711 | 32.1349 | 40.8463 | 25.2814 | 2.0036 | 32.0607 | 64.2356 | 16.0018 | 14.0594 | -0.1104 | 105.0819 | 234.7341 | 180.7368 | 232.6622 | 161.1342
2012-10-30 17:47 | 28.2091 | 248.7873 | 194.7901 | 246.5823 | 175.125 | -1.5997 | 1.2711 | 32.1342 | 40.8449 | 25.2812 | 2.0035 | 32.0599 | 64.2325 | 16.0018 | 14.1225 | -0.1103 | 105.0773 | 234.6648 | 180.6676 | 232.4598 | 161.0025
2012-10-30 17:47 | 28.2121 | 248.7825 | 194.7577 | 246.6323 | 175.0305 | -1.6014 | 1.2711 | 32.1342 | 40.8447 | 25.2813 | 2.0035 | 32.0599 | 64.2322 | 16.0019 | 14.0684 | -0.1104 | 105.0768 | 234.7141 | 180.6893 | 232.5639 | 160.9621
2012-10-30 17:47 | 28.2158 | 248.7862 | 194.7693 | 246.7339 | 175.1005 | -1.6032 | 1.2711 | 32.1339 | 40.8439 | 25.2814 | 2.0035 | 32.0596 | 64.2315 | 16.0018 | 14.0721 | -0.1105 | 105.0754 | 234.7141 | 180.6972 | 232.6618 | 161.0284
2012-10-30 17:47 | 28.2106 | 248.7888 | 194.7484 | 246.7405 | 175.1656 | -1.602 | 1.2711 | 32.1336 | 40.8437 | 25.281 | 2.0035 | 32.0593 | 64.2302 | 16.0018 | 14.1281 | -0.1104 | 105.0739 | 234.6607 | 180.6203 | 232.6124 | 161.0375
2012-10-30 17:47 | 28.2128 | 248.8145 | 194.7623 | 246.586 | 175.156 | -1.6003 | 1.271 | 32.1336 | 40.8427 | 25.2815 | 2.0035 | 32.0594 | 64.2305 | 16.0018 | 14.0609 | -0.1103 | 105.0732 | 234.7536 | 180.7014 | 232.5251 | 161.0951
2012-10-30 17:47 | 28.2158 | 248.8136 | 194.7182 | 246.7183 | 175.0771 | -1.6009 | 1.271 | 32.1331 | 40.8415 | 25.2815 | 2.0035 | 32.0589 | 64.2289 | 16.0017 | 14.0925 | -0.1104 | 105.0705 | 234.7211 | 180.6257 | 232.6258 | 160.9846
2012-10-30 17:47 | 28.2099 | 248.8625 | 194.7241 | 246.6888 | 174.9619 | -1.6003 | 1.271 | 32.1329 | 40.8412 | 25.2815 | 2.0035 | 32.0586 | 64.2282 | 16.0016 | 14.0784 | -0.1103 | 105.0693 | 234.7841 | 180.6457 | 232.6104 | 160.8835
2012-10-30 17:47 | 28.2136 | 248.8232 | 194.7317 | 246.6494 | 175.1256 | -1.598 | 1.271 | 32.1324 | 40.8401 | 25.2813 | 2.0034 | 32.0583 | 64.2267 | 16.0017 | 14.1107 | -0.1102 | 105.0669 | 234.7125 | 180.621 | 232.5387 | 161.0149
2012-10-30 17:48 | 28.2113 | 248.8444 | 194.7058 | 246.604 | 175.0531 | -1.6043 | 1.271 | 32.1319 | 40.8387 | 25.2814 | 2.0034 | 32.0577 | 64.2247 | 16.0016 | 14.0635 | -0.1106 | 105.0634 | 234.7809 | 180.6423 | 232.5405 | 160.9896
2012-10-30 17:48 | 28.2091 | 248.8461 | 194.6604 | 246.7742 | 175.0548 | -1.6014 | 1.271 | 32.1318 | 40.8386 | 25.2812 | 2.0034 | 32.0574 | 64.2234 | 16.0016 | 14.0899 | -0.1104 | 105.062 | 234.7562 | 180.5705 | 232.6843 | 160.9649
2012-10-30 17:48 | 28.2106 | 248.8319 | 194.7366 | 246.6856 | 175.2592 | -1.5939 | 1.271 | 32.1315 | 40.8376 | 25.2814 | 2.0034 | 32.0571 | 64.2227 | 16.0015 | 14.1118 | -0.1099 | 105.0603 | 234.7201 | 180.6248 | 232.5738 | 161.1474
2012-10-30 17:48 | 28.2151 | 248.8521 | 194.7254 | 246.7058 | 175.174 | -1.6072 | 1.2709 | 32.1316 | 40.8375 | 25.2816 | 2.0034 | 32.0571 | 64.2227 | 16.0015 | 14.104 | -0.1108 | 105.0603 | 234.7481 | 180.6214 | 232.6018 | 161.07
2012-10-30 17:48 | 28.2069 | 248.8439 | 194.7329 | 246.5957 | 175.3063 | -1.6061 | 1.2652 | 31.9839 | 40.4653 | 25.2802 | 1.9944 | 31.9132 | 63.6473 | 16.0015 | 14.104 | -0.1107 | 104.1127 | 234.7399 | 180.6289 | 232.4917 | 161.2023
2012-10-30 17:48 | 28.2061 | 248.8548 | 194.7517 | 246.736 | 175.2781 | -1.6037 | 1.2649 | 31.9755 | 40.4447 | 25.2798 | 1.9936 | 31.9013 | 63.599 | 16.0017 | 14.0787 | -0.1106 | 104.0437 | 234.7761 | 180.673 | 232.6573 | 161.1994
2012-10-30 17:48 | 28.2024 | 248.8159 | 194.7362 | 246.7754 | 175.3798 | -1.6014 | 1.2643 | 31.96 | 40.4056 | 25.2797 | 1.9927 | 31.8861 | 63.5384 | 16.0017 | 14.128 | -0.1104 | 103.9439 | 234.6879 | 180.6082 | 232.6474 | 161.2518
2012-10-30 17:48 | 28.2001 | 248.8058 | 194.7457 | 246.6712 | 175.3424 | -1.6003 | 1.2638 | 31.9486 | 40.3774 | 25.2794 | 1.992 | 31.8749 | 63.4942 | 16.0016 | 14.138 | -0.1103 | 103.8716 | 234.6678 | 180.6077 | 232.5332 | 161.2044
2012-10-30 17:48 | 28.2031 | 248.7813 | 194.7762 | 246.6938 | 174.9824 | -1.6009 | 1.2636 | 31.944 | 40.3652 | 25.2796 | 1.9917 | 31.8701 | 63.4748 | 16.0017 | 14.1083 | -0.1104 | 103.8401 | 234.673 | 180.6679 | 232.5855 | 160.8741
2012-10-30 17:48 | 28.1956 | 248.7778 | 194.7255 | 246.7882 | 175.1662 | -1.6043 | 1.2636 | 31.9423 | 40.3608 | 25.2798 | 1.9915 | 31.8684 | 63.4673 | 16.0019 | 14.1212 | -0.1106 | 103.8281 | 234.6566 | 180.6043 | 232.667 | 161.045
2012-10-30 17:48 | 28.2016 | 248.7759 | 194.7826 | 246.6688 | 175.34 | -1.5997 | 1.2636 | 31.942 | 40.3605 | 25.2795 | 1.9915 | 31.8681 | 63.4662 | 16.0019 | 14.0864 | -0.1103 | 103.8267 | 234.6895 | 180.6962 | 232.5824 | 161.2536
2012-10-30 17:48 | 28.2099 | 248.7803 | 194.7477 | 246.6614 | 175.1531 | -1.5997 | 1.2635 | 31.9419 | 40.3598 | 25.2797 | 1.9915 | 31.8681 | 63.4658 | 16.0019 | 14.0866 | -0.1103 | 103.8256 | 234.6937 | 180.6611 | 232.5748 | 161.0665
2012-10-30 17:48 | 28.1949 | 248.7731 | 194.7562 | 246.6425 | 175.2513 | -1.5985 | 1.2636 | 31.9433 | 40.3635 | 25.2797 | 1.9916 | 31.8694 | 63.4714 | 16.0018 | 14.1165 | -0.1102 | 103.8348 | 234.6566 | 180.6397 | 232.526 | 161.1348
2012-10-30 17:48 | 28.2054 | 248.7719 | 194.7903 | 246.794 | 175.2033 | -1.6014 | 1.2637 | 31.9453 | 40.3682 | 25.2799 | 1.9917 | 31.8716 | 63.4794 | 16.002 | 14.1351 | -0.1104 | 103.8476 | 234.6368 | 180.6552 | 232.6589 | 161.0682
2012-10-30 17:48 | 28.2039 | 248.696 | 194.7377 | 246.6907 | 175.1237 | -1.602 | 1.2657 | 31.9973 | 40.4992 | 25.2801 | 1.9949 | 31.9228 | 63.6837 | 16.002 | 14.1051 | -0.1104 | 104.1829 | 234.5909 | 180.6326 | 232.5856 | 161.0186
2012-10-30 17:48 | 28.1911 | 248.7145 | 194.7721 | 246.5564 | 175.1031 | -1.6037 | 1.2687 | 32.0737 | 40.6909 | 25.2814 | 1.9997 | 31.9995 | 63.9904 | 16.0019 | 14.08 | -0.1106 | 104.6813 | 234.6345 | 180.6921 | 232.4764 | 161.0231
2012-10-30 17:48 | 28.1926 | 248.7199 | 194.7343 | 246.652 | 174.8236 | -1.5957 | 1.2693 | 32.0903 | 40.7326 | 25.2816 | 2.0008 | 32.0162 | 64.0579 | 16.0017 | 14.0774 | -0.11 | 104.7906 | 234.6425 | 180.6569 | 232.5746 | 160.7462
2012-10-30 17:48 | 28.1941 | 248.7449 | 194.7594 | 246.7436 | 174.9305 | -1.598 | 1.27 | 32.1066 | 40.774 | 25.2817 | 2.0018 | 32.0322 | 64.1213 | 16.0019 | 14.0626 | -0.1102 | 104.8952 | 234.6823 | 180.6968 | 232.681 | 160.8679
2012-10-30 17:48 | 28.1896 | 248.7639 | 194.7077 | 246.6529 | 175.0782 | -1.6037 | 1.2704 | 32.1183 | 40.8039 | 25.2816 | 2.0025 | 32.0437 | 64.1676 | 16.0018 | 14.0867 | -0.1106 | 104.9715 | 234.6772 | 180.621 | 232.5662 | 160.9915
2012-10-30 17:48 | 28.1971 | 248.7949 | 194.7584 | 246.5742 | 175.2457 | -1.5997 | 1.2707 | 32.1248 | 40.8201 | 25.2817 | 2.0029 | 32.0503 | 64.1945 | 16.0017 | 14.0901 | -0.1103 | 105.0146 | 234.7048 | 180.6683 | 232.4841 | 161.1556
2012-10-30 17:48 | 28.1956 | 248.7738 | 194.7373 | 246.7059 | 175.1896 | -1.6026 | 1.2708 | 32.1291 | 40.831 | 25.2817 | 2.0032 | 32.0548 | 64.2126 | 16.0017 | 14.109 | -0.1105 | 105.0436 | 234.6648 | 180.6283 | 232.5969 | 161.0806
2012-10-30 17:48 | 28.1986 | 248.7455 | 194.7757 | 246.6658 | 175.3838 | -1.5974 | 1.271 | 32.1324 | 40.839 | 25.282 | 2.0034 | 32.058 | 64.2253 | 16.0017 | 14.0875 | -0.1101 | 105.0643 | 234.658 | 180.6882 | 232.5783 | 161.2963
2012-10-30 17:48 | 28.1941 | 248.788 | 194.834 | 246.6339 | 175.4067 | -1.6014 | 1.2711 | 32.1347 | 40.8449 | 25.2819 | 2.0036 | 32.0602 | 64.2349 | 16.0016 | 14.0626 | -0.1104 | 105.0798 | 234.7254 | 180.7714 | 232.5713 | 161.3441
2012-10-30 17:48 | 28.1881 | 248.7898 | 194.8516 | 246.5652 | 175.3265 | -1.6032 | 1.2711 | 32.1359 | 40.8483 | 25.2817 | 2.0037 | 32.0615 | 64.2401 | 16.0015 | 14.0485 | -0.1105 | 105.0884 | 234.7413 | 180.8031 | 232.5167 | 161.278
2012-10-30 17:48 | 28.1896 | 248.7835 | 194.7784 | 246.649 | 175.2109 | -1.6014 | 1.2711 | 32.1368 | 40.8501 | 25.282 | 2.0037 | 32.0624 | 64.2435 | 16.0015 | 14.103 | -0.1104 | 105.0936 | 234.6805 | 180.6754 | 232.546 | 161.1079
2012-10-30 17:49 | 28.1919 | 248.7388 | 194.824 | 246.7139 | 175.342 | -1.6078 | 1.2712 | 32.1373 | 40.8517 | 25.2818 | 2.0037 | 32.063 | 64.2461 | 16.0015 | 14.0849 | -0.1108 | 105.0978 | 234.6539 | 180.7391 | 232.629 | 161.2571
2012-10-30 17:49 | 28.1866 | 248.7844 | 194.8187 | 246.5637 | 175.3289 | -1.6032 | 1.2712 | 32.1376 | 40.8527 | 25.2817 | 2.0038 | 32.0631 | 64.2468 | 16.0014 | 14.0674 | -0.1105 | 105.0995 | 234.717 | 180.7513 | 232.4963 | 161.2615
2012-10-30 17:49 | 28.1911 | 248.8046 | 194.8114 | 246.6074 | 175.31 | -1.6055 | 1.2712 | 32.1386 | 40.8543 | 25.2823 | 2.0038 | 32.0642 | 64.2514 | 16.0014 | 14.0719 | -0.1107 | 105.1057 | 234.7327 | 180.7395 | 232.5355 | 161.2381
2012-10-30 17:49 | 28.1896 | 248.7522 | 194.8138 | 246.6764 | 175.1406 | -1.5997 | 1.2712 | 32.1395 | 40.8567 | 25.2822 | 2.0039 | 32.065 | 64.2552 | 16.0013 | 14.0622 | -0.1103 | 105.1118 | 234.69 | 180.7516 | 232.6142 | 161.0784
2012-10-30 17:49 | 28.1896 | 248.8031 | 194.7902 | 246.6176 | 175.2109 | -1.6003 | 1.2712 | 32.1398 | 40.8577 | 25.282 | 2.0039 | 32.0653 | 64.2563 | 16.0013 | 14.0785 | -0.1103 | 105.114 | 234.7246 | 180.7117 | 232.5391 | 161.1324
2012-10-30 17:49 | 28.2031 | 248.8479 | 194.8509 | 246.6468 | 175.0019 | -1.6032 | 1.2713 | 32.1399 | 40.8582 | 25.2819 | 2.0039 | 32.0654 | 64.257 | 16.0013 | 14.0635 | -0.1105 | 105.1152 | 234.7844 | 180.7874 | 232.5833 | 160.9384
2012-10-30 17:49 | 28.1949 | 248.8006 | 194.7916 | 246.7365 | 174.9157 | -1.5997 | 1.2713 | 32.1404 | 40.8593 | 25.282 | 2.004 | 32.0659 | 64.2588 | 16.0013 | 14.0716 | -0.1103 | 105.1181 | 234.729 | 180.72 | 232.6649 | 160.8441
2012-10-30 17:49 | 28.1986 | 248.8395 | 194.7364 | 246.6266 | 175.0443 | -1.5997 | 1.2713 | 32.1404 | 40.8594 | 25.282 | 2.004 | 32.066 | 64.2586 | 16.0014 | 14.1079 | -0.1103 | 105.118 | 234.7316 | 180.6285 | 232.5187 | 160.9364
2012-10-30 17:49 | 28.1971 | 248.8341 | 194.7152 | 246.6173 | 175.0896 | -1.602 | 1.2713 | 32.1406 | 40.8599 | 25.2819 | 2.004 | 32.0661 | 64.26 | 16.0012 | 14.1023 | -0.1104 | 105.1199 | 234.7318 | 180.6129 | 232.515 | 160.9873
2012-10-30 17:49 | 28.2061 | 248.8196 | 194.681 | 246.7242 | 175.1884 | -1.5928 | 1.2713 | 32.141 | 40.8606 | 25.2822 | 2.004 | 32.0665 | 64.2615 | 16.0012 | 14.095 | -0.1098 | 105.122 | 234.7246 | 180.586 | 232.6292 | 161.0934
2012-10-30 17:49 | 28.2001 | 248.7901 | 194.7575 | 246.6751 | 175.3463 | -1.6043 | 1.2713 | 32.1409 | 40.8605 | 25.282 | 2.004 | 32.0663 | 64.261 | 16.0012 | 14.0931 | -0.1106 | 105.1215 | 234.697 | 180.6644 | 232.582 | 161.2532
2012-10-30 17:49 | 28.1979 | 248.8231 | 194.7553 | 246.669 | 175.2934 | -1.6014 | 1.2713 | 32.1411 | 40.8612 | 25.2818 | 2.004 | 32.0665 | 64.2619 | 16.0011 | 14.0868 | -0.1104 | 105.1232 | 234.7363 | 180.6685 | 232.5822 | 161.2066
2012-10-30 17:49 | 28.2039 | 248.8213 | 194.8202 | 246.7024 | 175.1862 | -1.6032 | 1.2713 | 32.1411 | 40.8607 | 25.2822 | 2.004 | 32.0666 | 64.2621 | 16.0011 | 14.1377 | -0.1105 | 105.1228 | 234.6836 | 180.6825 | 232.5647 | 161.0485
2012-10-30 17:49 | 28.2039 | 248.8487 | 194.8202 | 246.769 | 175.2096 | -1.6014 | 1.2713 | 32.1408 | 40.8604 | 25.2819 | 2.004 | 32.0662 | 64.2611 | 16.001 | 14.1173 | -0.1104 | 105.1216 | 234.7314 | 180.7029 | 232.6517 | 161.0923
2012-10-30 17:49 | 28.2084 | 248.8415 | 194.7815 | 246.6442 | 175.1126 | -1.6061 | 1.2713 | 32.1406 | 40.8599 | 25.282 | 2.004 | 32.0662 | 64.2608 | 16.0011 | 14.1014 | -0.1107 | 105.1207 | 234.7401 | 180.6801 | 232.5428 | 161.0112
2012-10-30 17:49 | 28.2076 | 248.8524 | 194.7493 | 246.6865 | 175.0572 | -1.6026 | 1.2713 | 32.1403 | 40.859 | 25.282 | 2.004 | 32.0658 | 64.2598 | 16.001 | 14.0884 | -0.1105 | 105.1188 | 234.764 | 180.6609 | 232.5981 | 160.9688
2012-10-30 17:49 | 28.2016 | 248.866 | 194.759 | 246.7863 | 175.0629 | -1.5974 | 1.2713 | 32.1401 | 40.8587 | 25.2819 | 2.004 | 32.0657 | 64.259 | 16.001 | 14.1068 | -0.1101 | 105.1177 | 234.7592 | 180.6522 | 232.6795 | 160.9561
2012-10-30 17:49 | 28.2076 | 248.8289 | 194.7689 | 246.6865 | 175.1899 | -1.6032 | 1.2713 | 32.1401 | 40.8584 | 25.2821 | 2.004 | 32.0656 | 64.2586 | 16.001 | 14.068 | -0.1105 | 105.117 | 234.7609 | 180.7009 | 232.6185 | 161.1219
2012-10-30 17:49 | 28.2076 | 248.8563 | 194.7179 | 246.7179 | 175.022 | -1.6014 | 1.2713 | 32.1399 | 40.8585 | 25.2817 | 2.004 | 32.0654 | 64.2577 | 16.001 | 14.0598 | -0.1104 | 105.1162 | 234.7965 | 180.6581 | 232.6581 | 160.9622
2012-10-30 17:49 | 28.2054 | 248.8502 | 194.7432 | 246.798 | 175.012 | -1.6066 | 1.2712 | 32.1392 | 40.8562 | 25.2821 | 2.0039 | 32.0648 | 64.2561 | 16.0008 | 14.0576 | -0.1108 | 105.1123 | 234.7926 | 180.6856 | 232.7404 | 160.9544
2012-10-30 17:49 | 28.2099 | 248.8899 | 194.728 | 246.6653 | 175.1531 | -1.6003 | 1.2712 | 32.1394 | 40.8569 | 25.282 | 2.004 | 32.0649 | 64.2565 | 16.0009 | 14.0866 | -0.1103 | 105.1135 | 234.8033 | 180.6414 | 232.5787 | 161.0665
2012-10-30 17:49 | 28.2054 | 248.8854 | 194.696 | 246.7118 | 175.0901 | -1.598 | 1.2712 | 32.1392 | 40.8557 | 25.2823 | 2.0039 | 32.0647 | 64.2558 | 16.0009 | 14.0739 | -0.1102 | 105.1115 | 234.8115 | 180.6221 | 232.6379 | 161.0162
2012-10-30 17:49 | 28.2158 | 248.8998 | 194.7025 | 246.8671 | 175.202 | -1.5991 | 1.2712 | 32.1387 | 40.8546 | 25.2823 | 2.0039 | 32.0643 | 64.2539 | 16.0009 | 14.1088 | -0.1102 | 105.1085 | 234.791 | 180.5937 | 232.7583 | 161.0932
2012-10-30 17:49 | 28.2076 | 248.8681 | 194.7179 | 246.714 | 175.1899 | -1.5991 | 1.2712 | 32.1381 | 40.8529 | 25.2823 | 2.0039 | 32.0637 | 64.2512 | 16.001 | 14.0965 | -0.1102 | 105.1042 | 234.7716 | 180.6214 | 232.6175 | 161.0934
2012-10-30 17:49 | 28.2166 | 248.9123 | 194.7583 | 246.7621 | 174.9803 | -1.5974 | 1.2712 | 32.1378 | 40.8526 | 25.282 | 2.0039 | 32.0633 | 64.2506 | 16.0007 | 14.1055 | -0.1101 | 105.1033 | 234.8068 | 180.6528 | 232.6566 | 160.8748
2012-10-30 17:49 | 28.2113 | 248.9266 | 194.7451 | 246.8117 | 175.0726 | -1.6043 | 1.2712 | 32.1378 | 40.8525 | 25.2821 | 2.0039 | 32.0634 | 64.2517 | 16.0005 | 14.1451 | -0.1106 | 105.1042 | 234.7815 | 180.6 | 232.6666 | 160.9275
2012-10-30 17:49 | 28.2121 | 248.9196 | 194.7381 | 246.7733 | 175.1553 | -1.6061 | 1.2712 | 32.1378 | 40.8531 | 25.2818 | 2.0039 | 32.0635 | 64.2513 | 16.0007 | 14.1255 | -0.1107 | 105.1044 | 234.7941 | 180.6126 | 232.6478 | 161.0298
2012-10-30 17:50 | 28.2061 | 248.9018 | 194.7399 | 246.7321 | 175.1454 | -1.6084 | 1.2712 | 32.1377 | 40.8523 | 25.2821 | 2.0039 | 32.0632 | 64.2504 | 16.0007 | 14.1113 | -0.1109 | 105.1027 | 234.7905 | 180.6286 | 232.6208 | 161.0341
2012-10-30 17:50 | 28.2143 | 248.8904 | 194.7128 | 246.8735 | 175.0756 | -1.6072 | 1.2711 | 32.1372 | 40.8511 | 25.2821 | 2.0038 | 32.0628 | 64.2491 | 16.0006 | 14.0869 | -0.1108 | 105.1002 | 234.8035 | 180.6259 | 232.7866 | 160.9887
2012-10-30 17:50 | 28.2099 | 248.9174 | 194.7634 | 246.6653 | 175.1024 | -1.5991 | 1.2711 | 32.1368 | 40.85 | 25.2821 | 2.0038 | 32.0623 | 64.2472 | 16.0006 | 14.1151 | -0.1102 | 105.0972 | 234.8023 | 180.6483 | 232.5502 | 160.9873
2012-10-30 17:50 | 28.2061 | 248.9175 | 194.7321 | 246.7595 | 175.2001 | -1.602 | 1.2711 | 32.1367 | 40.8496 | 25.2821 | 2.0038 | 32.0623 | 64.2466 | 16.0007 | 14.1113 | -0.1104 | 105.0962 | 234.8062 | 180.6208 | 232.6482 | 161.0888
2012-10-30 17:50 | 28.2076 | 248.8876 | 194.7179 | 246.8903 | 175.1391 | -1.6026 | 1.2711 | 32.1364 | 40.8491 | 25.282 | 2.0038 | 32.062 | 64.2461 | 16.0006 | 14.1006 | -0.1105 | 105.0952 | 234.787 | 180.6173 | 232.7897 | 161.0385
2012-10-30 17:50 | 28.2143 | 248.9296 | 194.7324 | 246.6815 | 175.0717 | -1.6032 | 1.2711 | 32.1363 | 40.8486 | 25.2823 | 2.0038 | 32.0619 | 64.2462 | 16.0004 | 14.1155 | -0.1105 | 105.0948 | 234.8141 | 180.6169 | 232.566 | 160.9562
2012-10-30 17:50 | 28.2031 | 248.9066 | 194.7173 | 246.7839 | 174.9863 | -1.602 | 1.2711 | 32.1359 | 40.8475 | 25.2822 | 2.0038 | 32.0614 | 64.2438 | 16.0005 | 14.092 | -0.1104 | 105.0913 | 234.8146 | 180.6253 | 232.6919 | 160.8943
2012-10-30 17:50 | 28.2136 | 248.9563 | 194.7278 | 246.8336 | 175.0788 | -1.5997 | 1.2711 | 32.1359 | 40.8479 | 25.282 | 2.0038 | 32.0615 | 64.244 | 16.0006 | 14.0903 | -0.1103 | 105.0919 | 234.866 | 180.6375 | 232.7433 | 160.9885
2012-10-30 17:50 | 28.2046 | 248.9395 | 194.7109 | 246.7697 | 174.9878 | -1.5991 | 1.2711 | 32.1357 | 40.8472 | 25.2822 | 2.0038 | 32.0614 | 64.2437 | 16.0006 | 14.0894 | -0.1102 | 105.0909 | 234.8501 | 180.6215 | 232.6803 | 160.8984
2012-10-30 17:50 | 28.2158 | 248.9193 | 194.7182 | 246.7378 | 175.1434 | -1.6009 | 1.2711 | 32.1356 | 40.8469 | 25.2821 | 2.0038 | 32.0613 | 64.2433 | 16.0005 | 14.1047 | -0.1104 | 105.0902 | 234.8146 | 180.6135 | 232.6331 | 161.0387
2012-10-30 17:50 | 28.2069 | 248.93 | 194.6661 | 246.8661 | 175.2087 | -1.6043 | 1.2711 | 32.1354 | 40.8461 | 25.2822 | 2.0038 | 32.0609 | 64.2424 | 16.0004 | 14.0958 | -0.1106 | 105.0885 | 234.8342 | 180.5703 | 232.7703 | 161.1129
2012-10-30 17:50 | 28.2128 | 248.9085 | 194.7231 | 246.7975 | 175.3824 | -1.5991 | 1.2711 | 32.1348 | 40.8451 | 25.2821 | 2.0037 | 32.0604 | 64.2402 | 16.0004 | 14.114 | -0.1102 | 105.0853 | 234.7945 | 180.6091 | 232.6835 | 161.2684
2012-10-30 17:50 | 28.2099 | 248.9291 | 194.7398 | 246.7515 | 175.2312 | -1.6055 | 1.2711 | 32.1346 | 40.8449 | 25.2817 | 2.0037 | 32.0601 | 64.2389 | 16.0005 | 14.0743 | -0.1107 | 105.0838 | 234.8548 | 180.6655 | 232.6772 | 161.1569
2012-10-30 17:50 | 28.2121 | 248.943 | 194.7106 | 246.883 | 174.9797 | -1.6014 | 1.271 | 32.1343 | 40.8428 | 25.2826 | 2.0037 | 32.0598 | 64.2382 | 16.0003 | 14.1092 | -0.1104 | 105.081 | 234.8338 | 180.6014 | 232.7738 | 160.8705
2012-10-30 17:50 | 28.2196 | 248.9779 | 194.7416 | 246.8592 | 175.1394 | -1.6055 | 1.271 | 32.1338 | 40.8419 | 25.2824 | 2.0037 | 32.0594 | 64.236 | 16.0004 | 14.0963 | -0.1107 | 105.0779 | 234.8816 | 180.6453 | 232.7629 | 161.0431
2012-10-30 17:50 | 28.2226 | 248.9692 | 194.7643 | 246.7564 | 175.0644 | -1.602 | 1.271 | 32.1332 | 40.8411 | 25.2819 | 2.0036 | 32.0587 | 64.2337 | 16.0003 | 14.0993 | -0.1104 | 105.0748 | 234.8699 | 180.665 | 232.6571 | 160.9651
2012-10-30 17:50 | 28.2203 | 248.9512 | 194.7423 | 246.8089 | 174.9879 | -1.5968 | 1.271 | 32.1328 | 40.8393 | 25.2824 | 2.0036 | 32.0584 | 64.2327 | 16.0003 | 14.0643 | -0.1101 | 105.072 | 234.8869 | 180.678 | 232.7446 | 160.9236
2012-10-30 17:50 | 28.2158 | 248.9663 | 194.6828 | 246.871 | 175.0888 | -1.5997 | 1.271 | 32.1327 | 40.8401 | 25.2818 | 2.0036 | 32.0583 | 64.2323 | 16.0002 | 14.0925 | -0.1103 | 105.0723 | 234.8738 | 180.5903 | 232.7785 | 160.9963
2012-10-30 17:50 | 28.2263 | 248.9416 | 194.7601 | 246.7483 | 175.0837 | -1.5968 | 1.2709 | 32.1324 | 40.8384 | 25.2824 | 2.0036 | 32.058 | 64.2311 | 16.0003 | 14.0785 | -0.1101 | 105.0695 | 234.8631 | 180.6816 | 232.6698 | 161.0052
2012-10-30 17:50 | 28.2218 | 248.941 | 194.6849 | 246.8575 | 174.9231 | -1.598 | 1.2709 | 32.1322 | 40.8384 | 25.2821 | 2.0036 | 32.0578 | 64.2305 | 16.0002 | 14.0618 | -0.1102 | 105.0689 | 234.8792 | 180.6231 | 232.7957 | 160.8613
2012-10-30 17:50 | 28.2233 | 248.9464 | 194.71 | 246.8629 | 175.1548 | -1.6084 | 1.2692 | 32.0889 | 40.7285 | 25.282 | 2.0009 | 32.015 | 64.0595 | 16.0002 | 14.1 | -0.1109 | 104.788 | 234.8464 | 180.61 | 232.7629 | 161.0548
2012-10-30 17:50 | 28.2248 | 248.9479 | 194.739 | 246.7273 | 175.3554 | -1.6043 | 1.2647 | 31.9722 | 40.4349 | 25.2807 | 1.9936 | 31.8987 | 63.5946 | 16.0003 | 14.0729 | -0.1106 | 104.0295 | 234.875 | 180.6661 | 232.6544 | 161.2825
2012-10-30 17:50 | 28.2308 | 248.9578 | 194.7214 | 246.8469 | 175.1545 | -1.5991 | 1.2645 | 31.9683 | 40.4251 | 25.2806 | 1.9934 | 31.8944 | 63.5772 | 16.0003 | 14.136 | -0.1102 | 104.0023 | 234.8218 | 180.5854 | 232.7109 | 161.0185
2012-10-30 17:50 | 28.2308 | 248.8756 | 194.7136 | 246.8586 | 175.2482 | -1.6009 | 1.264 | 31.9534 | 40.388 | 25.2803 | 1.9924 | 31.8795 | 63.5177 | 16.0003 | 14.0952 | -0.1104 | 103.9057 | 234.7804 | 180.6184 | 232.7634 | 161.153
2012-10-30 17:50 | 28.2346 | 248.8677 | 194.7488 | 246.8428 | 175.2637 | -1.602 | 1.2637 | 31.9455 | 40.3681 | 25.2802 | 1.9919 | 31.8715 | 63.4853 | 16.0004 | 14.1235 | -0.1104 | 103.8534 | 234.7442 | 180.6253 | 232.7193 | 161.1402
2012-10-30 17:50 | 28.2293 | 248.8663 | 194.7631 | 246.7278 | 175.274 | -1.5957 | 1.2635 | 31.9427 | 40.3612 | 25.2802 | 1.9917 | 31.8687 | 63.4741 | 16.0005 | 14.0856 | -0.11 | 103.8352 | 234.7807 | 180.6775 | 232.6422 | 161.1884
2012-10-30 17:50 | 28.2263 | 248.8868 | 194.7444 | 246.862 | 175.2398 | -1.6049 | 1.2635 | 31.9425 | 40.3603 | 25.2805 | 1.9917 | 31.8686 | 63.4734 | 16.0005 | 14.103 | -0.1106 | 103.8337 | 234.7838 | 180.6414 | 232.759 | 161.1368
2012-10-30 17:51 | 28.2368 | 248.8542 | 194.7431 | 246.896 | 175.301 | -1.6003 | 1.2636 | 31.9431 | 40.3619 | 25.2804 | 1.9918 | 31.8691 | 63.4758 | 16.0005 | 14.1339 | -0.1103 | 103.8377 | 234.7203 | 180.6092 | 232.7621 | 161.1671
2012-10-30 17:51 | 28.2323 | 248.838 | 194.7661 | 246.7348 | 175.3551 | -1.5985 | 1.2636 | 31.9449 | 40.3658 | 25.2807 | 1.9919 | 31.8707 | 63.482 | 16.0005 | 14.0845 | -0.1102 | 103.8478 | 234.7535 | 180.6816 | 232.6503 | 161.2706
2012-10-30 17:51 | 28.2293 | 248.8271 | 194.7749 | 246.7318 | 175.3521 | -1.6014 | 1.2637 | 31.9464 | 40.3705 | 25.2801 | 1.992 | 31.8724 | 63.4889 | 16.0004 | 14.0897 | -0.1104 | 103.8594 | 234.7374 | 180.6852 | 232.6421 | 161.2624
2012-10-30 17:51 | 28.2308 | 248.8169 | 194.8118 | 246.89 | 175.3146 | -1.6055 | 1.2638 | 31.9488 | 40.3764 | 25.2803 | 1.9921 | 31.8749 | 63.4982 | 16.0006 | 14.1034 | -0.1107 | 103.8746 | 234.7135 | 180.7084 | 232.7866 | 161.2112
2012-10-30 17:51 | 28.2256 | 248.7999 | 194.8144 | 246.779 | 175.4069 | -1.6003 | 1.2652 | 31.9855 | 40.4688 | 25.2806 | 1.9945 | 31.9139 | 63.6536 | 16.0006 | 14.1308 | -0.1103 | 104.1224 | 234.6691 | 180.6836 | 232.6482 | 161.2761
2012-10-30 17:51 | 28.2241 | 248.8063 | 194.7776 | 246.703 | 175.4445 | -1.6003 | 1.2686 | 32.0729 | 40.6882 | 25.2818 | 1.9999 | 31.9988 | 63.9933 | 16.0005 | 14.1049 | -0.1103 | 104.6815 | 234.7014 | 180.6727 | 232.5981 | 161.3396
2012-10-30 17:51 | 28.2308 | 248.8365 | 194.8707 | 246.839 | 175.4356 | -1.5991 | 1.2692 | 32.0881 | 40.7268 | 25.2818 | 2.0008 | 32.0138 | 64.0533 | 16.0005 | 14.1156 | -0.1102 | 104.7801 | 234.7209 | 180.7551 | 232.7234 | 161.32
2012-10-30 17:51 | 28.2211 | 248.8385 | 194.9278 | 246.7353 | 175.4844 | -1.602 | 1.2698 | 32.1026 | 40.763 | 25.2821 | 2.0017 | 32.0281 | 64.1108 | 16.0004 | 14.163 | -0.1104 | 104.8739 | 234.6755 | 180.7648 | 232.5723 | 161.3214
2012-10-30 17:51 | 28.2076 | 248.8642 | 194.9457 | 246.7257 | 175.4084 | -1.5997 | 1.2709 | 32.1316 | 40.8363 | 25.2824 | 2.0036 | 32.0571 | 64.2281 | 16.0002 | 14.1251 | -0.1103 | 105.0644 | 234.7391 | 180.8206 | 232.6006 | 161.2833
2012-10-30 17:51 | 28.2106 | 248.9807 | 195.0273 | 246.7561 | 175.5754 | -1.5991 | 1.2712 | 32.1393 | 40.8549 | 25.2829 | 2.004 | 32.0647 | 64.2589 | 16 | 14.0914 | -0.1102 | 105.1139 | 234.8893 | 180.9359 | 232.6647 | 161.484
2012-10-30 17:51 | 28.2009 | 248.9005 | 194.9823 | 246.7112 | 175.5657 | -1.602 | 1.2712 | 32.1394 | 40.8554 | 25.2828 | 2.004 | 32.0648 | 64.2594 | 16.0001 | 14.1184 | -0.1104 | 105.1148 | 234.7821 | 180.8639 | 232.5928 | 161.4473
2012-10-30 17:51 | 28.2061 | 248.9057 | 194.9796 | 246.8535 | 175.4226 | -1.5991 | 1.2712 | 32.1396 | 40.8563 | 25.2827 | 2.0041 | 32.0652 | 64.2608 | 16.0001 | 14.1195 | -0.1102 | 105.1171 | 234.7862 | 180.8601 | 232.734 | 161.3031
2012-10-30 17:51 | 28.2046 | 248.9551 | 194.982 | 246.7697 | 175.464 | -1.602 | 1.2712 | 32.1402 | 40.8575 | 25.2828 | 2.0041 | 32.0657 | 64.2627 | 16.0001 | 14.0976 | -0.1104 | 105.1203 | 234.8575 | 180.8844 | 232.6721 | 161.3664
2012-10-30 17:51 | 28.2046 | 248.9512 | 194.982 | 246.7462 | 175.4367 | -1.6032 | 1.2712 | 32.1401 | 40.8569 | 25.283 | 2.0041 | 32.0656 | 64.2629 | 15.9999 | 14.1139 | -0.1105 | 105.1198 | 234.8373 | 180.8681 | 232.6323 | 161.3228
2012-10-30 17:51 | 28.2009 | 248.9592 | 195.0334 | 246.7151 | 175.3666 | -1.598 | 1.2712 | 32.1401 | 40.8573 | 25.2828 | 2.0041 | 32.0656 | 64.263 | 15.9999 | 14.0817 | -0.1102 | 105.1203 | 234.8775 | 180.9517 | 232.6334 | 161.2849
2012-10-30 17:51 | 28.2061 | 248.9605 | 194.9639 | 246.8496 | 175.1376 | -1.5939 | 1.2712 | 32.1401 | 40.8568 | 25.2831 | 2.0041 | 32.0657 | 64.2629 | 16 | 14.0991 | -0.1099 | 105.1196 | 234.8614 | 180.8648 | 232.7505 | 161.0385
2012-10-30 17:51 | 28.2113 | 248.9736 | 194.9455 | 246.7843 | 175.3458 | -1.6014 | 1.2712 | 32.1402 | 40.8571 | 25.283 | 2.0041 | 32.0656 | 64.2631 | 15.9999 | 14.1247 | -0.1104 | 105.1201 | 234.8489 | 180.8208 | 232.6596 | 161.2211
2012-10-30 17:51 | 28.2001 | 248.9545 | 194.9775 | 246.6751 | 175.2409 | -1.5985 | 1.2712 | 32.1402 | 40.8577 | 25.2828 | 2.0041 | 32.0657 | 64.2638 | 15.9998 | 14.1053 | -0.1102 | 105.1215 | 234.8492 | 180.8722 | 232.5698 | 161.1356
2012-10-30 17:52 | 28.1889 | 248.9747 | 194.931 | 246.7893 | 175.1634 | -1.5991 | 1.2712 | 32.1399 | 40.8564 | 25.283 | 2.0041 | 32.0654 | 64.2618 | 16 | 14.1023 | -0.1102 | 105.1182 | 234.8724 | 180.8287 | 232.687 | 161.0611
2012-10-30 17:52 | 28.1994 | 249.0008 | 194.9651 | 246.8194 | 175.2792 | -1.6061 | 1.2712 | 32.1397 | 40.856 | 25.283 | 2.0041 | 32.0652 | 64.2619 | 15.9998 | 14.1413 | -0.1107 | 105.1179 | 234.8595 | 180.8238 | 232.6781 | 161.1379
2012-10-30 17:52 | 28.2016 | 248.9717 | 194.9201 | 246.8608 | 175.2932 | -1.602 | 1.2712 | 32.1395 | 40.8552 | 25.2831 | 2.0041 | 32.0649 | 64.2605 | 15.9998 | 14.1639 | -0.1104 | 105.1157 | 234.8078 | 180.7562 | 232.6969 | 161.1293
2012-10-30 17:52 | 28.1904 | 248.9644 | 194.9285 | 246.8104 | 175.3873 | -1.6066 | 1.2712 | 32.1394 | 40.8552 | 25.283 | 2.0041 | 32.0649 | 64.2606 | 15.9998 | 14.1568 | -0.1108 | 105.1158 | 234.8076 | 180.7717 | 232.6536 | 161.2305
2012-10-30 17:52 | 28.1971 | 248.9672 | 194.9156 | 246.7936 | 175.3511 | -1.5997 | 1.2712 | 32.1394 | 40.8552 | 25.283 | 2.0041 | 32.0648 | 64.2608 | 15.9997 | 14.1187 | -0.1103 | 105.116 | 234.8485 | 180.7969 | 232.6749 | 161.2324
2012-10-30 17:52 | 28.2009 | 249.0062 | 194.939 | 246.8326 | 175.269 | -1.6026 | 1.2712 | 32.1392 | 40.8544 | 25.2831 | 2.0041 | 32.0647 | 64.2597 | 15.9998 | 14.1143 | -0.1105 | 105.1142 | 234.8919 | 180.8247 | 232.7183 | 161.1547
2012-10-30 17:52 | 28.1986 | 248.9844 | 194.921 | 246.9048 | 175.3331 | -1.5991 | 1.2712 | 32.139 | 40.8545 | 25.2828 | 2.0041 | 32.0645 | 64.2595 | 15.9996 | 14.112 | -0.1102 | 105.114 | 234.8724 | 180.809 | 232.7928 | 161.2211
2012-10-30 17:52 | 28.2024 | 249.0586 | 194.9405 | 246.8067 | 175.2042 | -1.6037 | 1.2711 | 32.1387 | 40.8531 | 25.2832 | 2.0041 | 32.0642 | 64.2587 | 15.9996 | 14.1158 | -0.1106 | 105.1118 | 234.9428 | 180.8247 | 232.6909 | 161.0884
2012-10-30 17:52 | 28.2009 | 249.0062 | 194.9037 | 246.7739 | 175.1012 | -1.6014 | 1.2712 | 32.1385 | 40.8532 | 25.2829 | 2.0041 | 32.0641 | 64.2583 | 15.9996 | 14.1061 | -0.1104 | 105.1115 | 234.9001 | 180.7976 | 232.6678 | 160.9951
2012-10-30 17:52 | 28.2009 | 249.0023 | 194.8723 | 246.9032 | 175.0583 | -1.602 | 1.2712 | 32.1384 | 40.8528 | 25.2829 | 2.0041 | 32.0639 | 64.258 | 15.9995 | 14.1469 | -0.1104 | 105.1108 | 234.8554 | 180.7254 | 232.7563 | 160.9114
2012-10-30 17:52 | 28.1971 | 248.9985 | 194.8488 | 246.8602 | 175.1169 | -1.5974 | 1.2711 | 32.1382 | 40.852 | 25.283 | 2.004 | 32.0636 | 64.2568 | 15.9995 | 14.0819 | -0.1101 | 105.1088 | 234.9166 | 180.7669 | 232.7783 | 161.035
2012-10-30 17:52 | 28.2031 | 249.0202 | 194.8587 | 246.8192 | 175.1034 | -1.6009 | 1.2711 | 32.138 | 40.8516 | 25.2829 | 2.004 | 32.0635 | 64.2562 | 15.9995 | 14.0839 | -0.1104 | 105.1078 | 234.9363 | 180.7748 | 232.7353 | 161.0195
2012-10-30 17:52 | 28.2069 | 248.9848 | 194.8272 | 246.8308 | 175.1384 | -1.602 | 1.2711 | 32.1381 | 40.8522 | 25.2828 | 2.0041 | 32.0636 | 64.2572 | 15.9994 | 14.1203 | -0.1104 | 105.1095 | 234.8645 | 180.7069 | 232.7105 | 161.0181
2012-10-30 17:52 | 28.2084 | 248.9863 | 194.8719 | 246.8911 | 174.9994 | -1.602 | 1.2711 | 32.1379 | 40.8515 | 25.283 | 2.004 | 32.0634 | 64.2564 | 15.9994 | 14.1381 | -0.1104 | 105.1079 | 234.8482 | 180.7338 | 232.753 | 160.8613
2012-10-30 17:52 | 28.2001 | 248.9663 | 194.7968 | 246.9102 | 174.9989 | -1.6009 | 1.2711 | 32.1376 | 40.8507 | 25.283 | 2.004 | 32.0631 | 64.2558 | 15.9992 | 14.0931 | -0.1104 | 105.1066 | 234.8732 | 180.7037 | 232.8171 | 160.9058
2012-10-30 17:52 | 28.2076 | 248.9659 | 194.82 | 246.8237 | 175.104 | -1.6014 | 1.2711 | 32.1373 | 40.8498 | 25.2829 | 2.004 | 32.0627 | 64.2537 | 15.9993 | 14.121 | -0.1104 | 105.1035 | 234.8449 | 180.699 | 232.7027 | 160.983
2012-10-30 17:52 | 28.2039 | 249.0327 | 194.8006 | 246.8082 | 175.034 | -1.6037 | 1.2711 | 32.1369 | 40.8494 | 25.2827 | 2.004 | 32.0623 | 64.2521 | 15.9994 | 14.1132 | -0.1106 | 105.1015 | 234.9195 | 180.6874 | 232.695 | 160.9208
2012-10-30 17:52 | 28.2151 | 249.0126 | 194.8197 | 246.9369 | 175.1154 | -1.5985 | 1.2711 | 32.1364 | 40.8479 | 25.2828 | 2.0039 | 32.0619 | 64.2505 | 15.9994 | 14.1367 | -0.1102 | 105.0984 | 234.8759 | 180.683 | 232.8002 | 160.9787
2012-10-30 17:52 | 28.2031 | 249.0084 | 194.7919 | 246.8936 | 175.2986 | -1.5957 | 1.2711 | 32.1363 | 40.8477 | 25.2827 | 2.0039 | 32.0618 | 64.25 | 15.9993 | 14.1369 | -0.11 | 105.0977 | 234.8715 | 180.655 | 232.7567 | 161.1617
2012-10-30 17:52 | 28.2076 | 249.009 | 194.8475 | 246.8628 | 175.4163 | -1.598 | 1.2711 | 32.1361 | 40.8466 | 25.283 | 2.0039 | 32.0616 | 64.2495 | 15.9993 | 14.1128 | -0.1102 | 105.0961 | 234.8962 | 180.7347 | 232.75 | 161.3035
2012-10-30 17:52 | 28.2166 | 249.0376 | 194.8526 | 246.7974 | 175.3043 | -1.6003 | 1.271 | 32.1358 | 40.846 | 25.2831 | 2.0039 | 32.0613 | 64.2487 | 15.9992 | 14.1382 | -0.1103 | 105.0947 | 234.8994 | 180.7144 | 232.6592 | 161.1661
2012-10-30 17:52 | 28.2158 | 249.0094 | 194.8754 | 246.969 | 175.0615 | -1.6026 | 1.2711 | 32.1358 | 40.8465 | 25.2827 | 2.0039 | 32.0612 | 64.2483 | 15.9992 | 14.1088 | -0.1105 | 105.0948 | 234.9006 | 180.7666 | 232.8602 | 160.9527
2012-10-30 17:52 | 28.2173 | 248.9952 | 194.9083 | 246.9235 | 175.1957 | -1.6014 | 1.271 | 32.1353 | 40.8449 | 25.2829 | 2.0039 | 32.0608 | 64.2466 | 15.9992 | 14.1185 | -0.1104 | 105.0915 | 234.8767 | 180.7898 | 232.805 | 161.0772
2012-10-30 17:52 | 28.2158 | 248.9976 | 194.9147 | 246.8593 | 175.2722 | -1.6026 | 1.271 | 32.1347 | 40.843 | 25.2831 | 2.0038 | 32.0601 | 64.2431 | 15.9994 | 14.117 | -0.1105 | 105.0861 | 234.8806 | 180.7977 | 232.7423 | 161.1552
2012-10-30 17:52 | 28.2173 | 249.0226 | 194.8611 | 246.8843 | 175.2854 | -1.6026 | 1.271 | 32.1346 | 40.843 | 25.2829 | 2.0038 | 32.06 | 64.2431 | 15.9993 | 14.1185 | -0.1105 | 105.0861 | 234.9041 | 180.7426 | 232.7658 | 161.1669
2012-10-30 17:53 | 28.2181 | 248.996 | 194.9012 | 246.9086 | 175.3487 | -1.5991 | 1.271 | 32.1343 | 40.842 | 25.2832 | 2.0038 | 32.0599 | 64.2423 | 15.9993 | 14.1274 | -0.1102 | 105.0843 | 234.8686 | 180.7738 | 232.7812 | 161.2213
2012-10-30 17:53 | 28.2173 | 249.007 | 194.9044 | 246.9705 | 175.3401 | -1.5997 | 1.271 | 32.1342 | 40.842 | 25.283 | 2.0038 | 32.0597 | 64.2425 | 15.9991 | 14.1348 | -0.1103 | 105.0845 | 234.8722 | 180.7696 | 232.8357 | 161.2053
2012-10-30 17:53 | 28.2173 | 249.0579 | 194.9397 | 246.9587 | 175.3986 | -1.6026 | 1.271 | 32.134 | 40.8413 | 25.283 | 2.0038 | 32.0594 | 64.2415 | 15.9991 | 14.1756 | -0.1105 | 105.0828 | 234.8823 | 180.7641 | 232.7831 | 161.223
2012-10-30 17:53 | 28.2203 | 249.0374 | 194.9781 | 246.8873 | 175.4836 | -1.6043 | 1.271 | 32.1335 | 40.8405 | 25.2828 | 2.0038 | 32.059 | 64.2399 | 15.9991 | 14.15 | -0.1106 | 105.0804 | 234.8874 | 180.8281 | 232.7373 | 161.3336
2012-10-30 17:53 | 28.2241 | 249.0216 | 195.0015 | 246.8362 | 175.5733 | -1.602 | 1.2709 | 32.1332 | 40.8392 | 25.2831 | 2.0038 | 32.0587 | 64.2388 | 15.9991 | 14.1375 | -0.1104 | 105.0781 | 234.8841 | 180.864 | 232.6987 | 161.4358
2012-10-30 17:53 | 28.2226 | 249.0436 | 195.0158 | 246.9405 | 175.4391 | -1.6009 | 1.2709 | 32.1329 | 40.8386 | 25.283 | 2.0038 | 32.0584 | 64.2375 | 15.9991 | 14.1401 | -0.1104 | 105.0761 | 234.9035 | 180.8757 | 232.8004 | 161.299
2012-10-30 17:53 | 28.2099 | 249.0074 | 195.0148 | 246.9827 | 175.4732 | -1.6037 | 1.2709 | 32.1326 | 40.8379 | 25.283 | 2.0038 | 32.058 | 64.2365 | 15.999 | 14.1355 | -0.1106 | 105.0743 | 234.8719 | 180.8793 | 232.8472 | 161.3377
2012-10-30 17:53 | 28.2143 | 249.0431 | 195.0507 | 246.893 | 175.423 | -1.598 | 1.2709 | 32.1323 | 40.8373 | 25.2829 | 2.0037 | 32.0579 | 64.2359 | 15.9989 | 14.1562 | -0.1102 | 105.0733 | 234.8869 | 180.8945 | 232.7368 | 161.2668
2012-10-30 17:53 | 28.2158 | 249.0564 | 195.0168 | 246.922 | 175.2878 | -1.6026 | 1.2709 | 32.1323 | 40.8375 | 25.2827 | 2.0038 | 32.0578 | 64.2359 | 15.9989 | 14.1496 | -0.1105 | 105.0734 | 234.9068 | 180.8672 | 232.7724 | 161.1382
2012-10-30 17:53 | 28.2151 | 249.0283 | 195.0161 | 246.8821 | 175.2949 | -1.5985 | 1.2655 | 31.994 | 40.4879 | 25.2819 | 1.9951 | 31.9195 | 63.6823 | 15.999 | 14.1448 | -0.1102 | 104.1703 | 234.8835 | 180.8713 | 232.7373 | 161.1501
2012-10-30 17:53 | 28.2106 | 249.0316 | 195.0313 | 246.9364 | 175.2631 | -1.6032 | 1.2647 | 31.9736 | 40.4368 | 25.2817 | 1.9939 | 31.8998 | 63.6038 | 15.999 | 14.1729 | -0.1105 | 104.0405 | 234.8587 | 180.8584 | 232.7635 | 161.0902
2012-10-30 17:53 | 28.2061 | 249.0075 | 194.9718 | 246.9632 | 175.3874 | -1.6032 | 1.264 | 31.9558 | 40.3927 | 25.2811 | 1.9927 | 31.8821 | 63.5325 | 15.9992 | 14.1644 | -0.1105 | 103.9253 | 234.8431 | 180.8074 | 232.7988 | 161.223
2012-10-30 17:53 | 28.2069 | 248.9927 | 194.9765 | 246.9287 | 175.5249 | -1.6061 | 1.2636 | 31.9465 | 40.369 | 25.2813 | 1.9921 | 31.8725 | 63.4946 | 15.9991 | 14.1488 | -0.1107 | 103.8636 | 234.8439 | 180.8277 | 232.7799 | 161.3761
2012-10-30 17:53 | 28.2113 | 249.0362 | 195.0202 | 246.8626 | 175.3731 | -1.6014 | 1.2635 | 31.9426 | 40.3592 | 25.2812 | 1.9919 | 31.8686 | 63.4785 | 15.9992 | 14.1655 | -0.1104 | 103.8377 | 234.8707 | 180.8547 | 232.6971 | 161.2076
2012-10-30 17:53 | 28.2046 | 248.9747 | 194.9938 | 246.9108 | 175.5265 | -1.5997 | 1.2635 | 31.9421 | 40.3581 | 25.2811 | 1.9918 | 31.8681 | 63.4759 | 15.9994 | 14.1751 | -0.1103 | 103.834 | 234.7996 | 180.8187 | 232.7357 | 161.3514
2012-10-30 17:53 | 28.2024 | 248.9412 | 194.9956 | 246.9673 | 175.536 | -1.6032 | 1.2634 | 31.9403 | 40.3537 | 25.281 | 1.9917 | 31.8662 | 63.4692 | 15.9992 | 14.1566 | -0.1105 | 103.8229 | 234.7846 | 180.839 | 232.8107 | 161.3794
2012-10-30 17:53 | 28.2009 | 248.9123 | 195.0491 | 246.9032 | 175.2456 | -1.5939 | 1.2635 | 31.943 | 40.3603 | 25.2811 | 1.9919 | 31.8687 | 63.4787 | 15.9993 | 14.1551 | -0.1099 | 103.839 | 234.7572 | 180.894 | 232.7481 | 161.0905
2012-10-30 17:53 | 28.1986 | 248.9295 | 194.9643 | 246.8969 | 175.2199 | -1.5991 | 1.2636 | 31.9455 | 40.3662 | 25.2814 | 1.9921 | 31.8714 | 63.4895 | 15.9993 | 14.1895 | -0.1102 | 103.8556 | 234.74 | 180.7748 | 232.7074 | 161.0304
2012-10-30 17:53 | 28.1956 | 248.9148 | 195.0202 | 246.8626 | 175.5097 | -1.6037 | 1.2647 | 31.9728 | 40.4356 | 25.2811 | 1.9939 | 31.9006 | 63.606 | 15.9993 | 14.1579 | -0.1106 | 104.0415 | 234.7569 | 180.8623 | 232.7047 | 161.3518
2012-10-30 17:53 | 28.1836 | 248.8519 | 194.9964 | 246.9094 | 175.5484 | -1.6009 | 1.2686 | 32.0723 | 40.686 | 25.2822 | 2 | 31.9979 | 63.9948 | 15.9992 | 14.1623 | -0.1104 | 104.6808 | 234.6896 | 180.8341 | 232.7471 | 161.3861
2012-10-30 17:53 | 28.1866 | 248.9215 | 195.0466 | 246.9202 | 175.5826 | -1.6049 | 1.2692 | 32.0887 | 40.7272 | 25.2825 | 2.001 | 32.0139 | 64.0591 | 15.9991 | 14.1612 | -0.1106 | 104.7864 | 234.7603 | 180.8854 | 232.759 | 161.4214
2012-10-30 17:53 | 28.1836 | 248.9302 | 195.075 | 246.831 | 175.525 | -1.6009 | 1.2699 | 32.106 | 40.7704 | 25.2829 | 2.0021 | 32.0312 | 64.1285 | 15.9991 | 14.1704 | -0.1104 | 104.8988 | 234.7598 | 180.9046 | 232.6606 | 161.3546
2012-10-30 17:53 | 28.1851 | 248.916 | 195.0804 | 246.8403 | 175.5226 | -1.6037 | 1.2704 | 32.1188 | 40.8027 | 25.283 | 2.0029 | 32.044 | 64.1795 | 15.9992 | 14.1393 | -0.1106 | 104.9822 | 234.7767 | 180.9411 | 232.701 | 161.3833
2012-10-30 17:53 | 28.1859 | 248.9286 | 195.0773 | 246.802 | 175.5351 | -1.6014 | 1.2706 | 32.1251 | 40.8185 | 25.2832 | 2.0033 | 32.0505 | 64.206 | 15.9991 | 14.1686 | -0.1104 | 105.0245 | 234.76 | 180.9087 | 232.6334 | 161.3665
2012-10-30 17:53 | 28.1784 | 248.9211 | 195.0226 | 246.8258 | 175.5705 | -1.5997 | 1.2708 | 32.1293 | 40.8293 | 25.2832 | 2.0035 | 32.0548 | 64.2233 | 15.9991 | 14.1326 | -0.1103 | 105.0525 | 234.7885 | 180.89 | 232.6932 | 161.4379
2012-10-30 17:54 | 28.1769 | 248.9274 | 195.0565 | 246.8439 | 175.7057 | -1.5991 | 1.2709 | 32.1324 | 40.8372 | 25.2831 | 2.0037 | 32.0579 | 64.2355 | 15.999 | 14.1188 | -0.1102 | 105.0727 | 234.8086 | 180.9377 | 232.7251 | 161.5869
2012-10-30 17:54 | 28.1776 | 248.9594 | 195.1161 | 246.8485 | 175.7493 | -1.6043 | 1.2709 | 32.134 | 40.8407 | 25.2835 | 2.0038 | 32.0595 | 64.2423 | 15.999 | 14.1522 | -0.1106 | 105.083 | 234.8072 | 180.9639 | 232.6963 | 161.5971
2012-10-30 17:54 | 28.1754 | 248.9572 | 195.1415 | 246.7915 | 175.751 | -1.602 | 1.271 | 32.1359 | 40.8458 | 25.2833 | 2.004 | 32.0613 | 64.2504 | 15.9988 | 14.1092 | -0.1104 | 105.0963 | 234.848 | 181.0323 | 232.6823 | 161.6418
2012-10-30 17:54 | 28.1679 | 248.9262 | 195.1104 | 246.7683 | 175.642 | -1.6014 | 1.2711 | 32.1369 | 40.848 | 25.2835 | 2.004 | 32.0623 | 64.2544 | 15.9988 | 14.1058 | -0.1104 | 105.1024 | 234.8204 | 181.0046 | 232.6625 | 161.5362
2012-10-30 17:54 | 28.1791 | 248.9766 | 195.1412 | 246.7912 | 175.6142 | -1.6009 | 1.2711 | 32.1377 | 40.8502 | 25.2834 | 2.0041 | 32.0631 | 64.2575 | 15.9988 | 14.1088 | -0.1104 | 105.1077 | 234.8678 | 181.0324 | 232.6824 | 161.5054
2012-10-30 17:54 | 28.1739 | 248.9636 | 195.14 | 246.8526 | 175.648 | -1.6014 | 1.2711 | 32.1382 | 40.8512 | 25.2835 | 2.0041 | 32.0637 | 64.2598 | 15.9988 | 14.1281 | -0.1104 | 105.111 | 234.8355 | 181.0119 | 232.7245 | 161.5199
2012-10-30 17:54 | 28.1702 | 248.9755 | 195.1598 | 246.8489 | 175.6404 | -1.602 | 1.271 | 32.1362 | 40.8466 | 25.2833 | 2.004 | 32.0617 | 64.2522 | 15.9987 | 14.0795 | -0.1104 | 105.0988 | 234.896 | 181.0803 | 232.7694 | 161.5609
2012-10-30 17:54 | 28.1762 | 248.9972 | 195.1894 | 246.8158 | 175.6347 | -1.6049 | 1.271 | 32.136 | 40.8458 | 25.2835 | 2.004 | 32.0614 | 64.2513 | 15.9986 | 14.1263 | -0.1106 | 105.097 | 234.8709 | 181.0631 | 232.6895 | 161.5084
2012-10-30 17:54 | 28.1754 | 248.9964 | 195.2043 | 246.7601 | 175.3606 | -1.6037 | 1.2711 | 32.137 | 40.8485 | 25.2834 | 2.0041 | 32.0625 | 64.2556 | 15.9986 | 14.0806 | -0.1106 | 105.1041 | 234.9158 | 181.1237 | 232.6795 | 161.28
2012-10-30 17:54 | 28.1739 | 249.0145 | 195.1124 | 246.8526 | 175.2889 | -1.5985 | 1.2711 | 32.1377 | 40.8505 | 25.2832 | 2.0041 | 32.0632 | 64.2581 | 15.9987 | 14.071 | -0.1102 | 105.1086 | 234.9435 | 181.0414 | 232.7816 | 161.2179
2012-10-30 17:54 | 28.1702 | 248.999 | 195.0577 | 246.8725 | 175.414 | -1.6009 | 1.2711 | 32.1383 | 40.8516 | 25.2835 | 2.0042 | 32.0637 | 64.2609 | 15.9985 | 14.0836 | -0.1104 | 105.1124 | 234.9154 | 180.9741 | 232.7889 | 161.3304
2012-10-30 17:54 | 28.1754 | 248.9846 | 195.1061 | 246.7915 | 175.6495 | -1.598 | 1.2711 | 32.1384 | 40.8521 | 25.2833 | 2.0042 | 32.0638 | 64.261 | 15.9987 | 14.0929 | -0.1102 | 105.1131 | 234.8917 | 181.0132 | 232.6986 | 161.5566
2012-10-30 17:54 | 28.1687 | 249.0367 | 195.0797 | 246.7769 | 175.5686 | -1.6032 | 1.2711 | 32.1382 | 40.8514 | 25.2834 | 2.0042 | 32.0636 | 64.2606 | 15.9985 | 14.0984 | -0.1105 | 105.1121 | 234.9383 | 180.9813 | 232.6785 | 161.4702
2012-10-30 17:54 | 28.1732 | 249.002 | 195.1196 | 246.8872 | 175.5809 | -1.5991 | 1.2711 | 32.1384 | 40.8517 | 25.2836 | 2.0042 | 32.0638 | 64.2617 | 15.9984 | 14.0866 | -0.1102 | 105.1134 | 234.9154 | 181.033 | 232.8006 | 161.4943
2012-10-30 17:54 | 28.1799 | 249.04 | 195.1342 | 246.8351 | 175.5408 | -1.6014 | 1.2711 | 32.1379 | 40.8507 | 25.2835 | 2.0042 | 32.0634 | 64.2603 | 15.9984 | 14.0729 | -0.1104 | 105.1109 | 234.9671 | 181.0613 | 232.7622 | 161.4679
2012-10-30 17:54 | 28.1717 | 249.0162 | 195.1063 | 246.9876 | 175.4779 | -1.6026 | 1.2711 | 32.1378 | 40.8501 | 25.2836 | 2.0042 | 32.0633 | 64.26 | 15.9983 | 14.028 | -0.1105 | 105.1101 | 234.9882 | 181.0783 | 232.9596 | 161.4499
2012-10-30 17:54 | 28.1604 | 249.001 | 195.1068 | 246.7921 | 175.4745 | -1.5974 | 1.2711 | 32.1375 | 40.8493 | 25.2837 | 2.0041 | 32.063 | 64.259 | 15.9983 | 14.0126 | -0.1101 | 105.1083 | 234.9884 | 181.0942 | 232.7795 | 161.4619
2012-10-30 17:54 | 28.1739 | 249.0771 | 195.1007 | 246.8801 | 175.3006 | -1.6032 | 1.2711 | 32.1373 | 40.8492 | 25.2834 | 2.0041 | 32.0628 | 64.2581 | 15.9984 | 13.9935 | -0.1105 | 105.1073 | 235.0836 | 181.1072 | 232.8866 | 161.3071
2012-10-30 17:54 | 28.1724 | 249.0599 | 195.0992 | 246.9217 | 175.4513 | -1.5974 | 1.2711 | 32.1373 | 40.849 | 25.2835 | 2.0041 | 32.0627 | 64.258 | 15.9982 | 14.0328 | -0.1101 | 105.107 | 235.0271 | 181.0664 | 232.8889 | 161.4185
2012-10-30 17:54 | 28.1739 | 249.0301 | 195.1085 | 246.9584 | 175.5816 | -1.5991 | 1.2711 | 32.1373 | 40.8489 | 25.2836 | 2.0041 | 32.0627 | 64.2581 | 15.9982 | 14.0343 | -0.1102 | 105.107 | 234.9958 | 181.0742 | 232.9241 | 161.5473
2012-10-30 17:54 | 28.1814 | 249.0689 | 195.116 | 246.8797 | 175.4681 | -1.6014 | 1.2711 | 32.1378 | 40.8501 | 25.2836 | 2.0042 | 32.0631 | 64.2598 | 15.9983 | 14.0091 | -0.1104 | 105.1099 | 235.0598 | 181.1069 | 232.8706 | 161.459
2012-10-30 17:54 | 28.1799 | 249.0674 | 195.0674 | 246.9958 | 175.3534 | -1.6037 | 1.2711 | 32.1378 | 40.8504 | 25.2834 | 2.0042 | 32.0633 | 64.2601 | 15.9983 | 14.0321 | -0.1106 | 105.1105 | 235.0353 | 181.0353 | 232.9637 | 161.3213
2012-10-30 17:54 | 28.1702 | 249.0382 | 195.0773 | 246.9273 | 175.5779 | -1.6003 | 1.2711 | 32.1373 | 40.8489 | 25.2836 | 2.0042 | 32.0627 | 64.2584 | 15.9981 | 14.0469 | -0.1103 | 105.1074 | 234.9913 | 181.0304 | 232.8804 | 161.531
2012-10-30 17:55 | 28.1724 | 249.0678 | 195.107 | 246.8394 | 175.6387 | -1.602 | 1.2711 | 32.137 | 40.8485 | 25.2833 | 2.0041 | 32.0623 | 64.2567 | 15.9982 | 14.0817 | -0.1104 | 105.1052 | 234.9861 | 181.0253 | 232.7577 | 161.557
2012-10-30 17:55 | 28.1597 | 249.0238 | 195.1336 | 246.9716 | 175.747 | -1.6032 | 1.271 | 32.1365 | 40.847 | 25.2836 | 2.0041 | 32.062 | 64.2546 | 15.9984 | 14.0323 | -0.1105 | 105.1016 | 234.9915 | 181.1013 | 232.9393 | 161.7147
2012-10-30 17:55 | 28.1657 | 249.0846 | 195.1946 | 246.9032 | 175.7296 | -1.6061 | 1.271 | 32.1353 | 40.8436 | 25.2837 | 2.004 | 32.0607 | 64.2504 | 15.9981 | 14.0505 | -0.1107 | 105.094 | 235.0341 | 181.1441 | 232.8527 | 161.6791
2012-10-30 17:55 | 28.1679 | 249.0946 | 195.1497 | 246.8271 | 175.5639 | -1.6084 | 1.271 | 32.1345 | 40.8419 | 25.2834 | 2.004 | 32.0599 | 64.2474 | 15.9982 | 14.0038 | -0.1109 | 105.0893 | 235.0908 | 181.1459 | 232.8233 | 161.5601
2012-10-30 17:55 | 28.1604 | 249.0558 | 195.1304 | 247.0233 | 175.5291 | -1.5985 | 1.271 | 32.1346 | 40.8425 | 25.2834 | 2.004 | 32.0601 | 64.2479 | 15.9981 | 14.0534 | -0.1102 | 105.0904 | 235.0024 | 181.077 | 232.9699 | 161.4757
2012-10-30 17:55 | 28.1604 | 249.0754 | 195.1579 | 246.9449 | 175.7555 | -1.6032 | 1.271 | 32.1351 | 40.8435 | 25.2835 | 2.004 | 32.0605 | 64.2497 | 15.9982 | 14.1023 | -0.1105 | 105.0932 | 234.9731 | 181.0556 | 232.8426 | 161.6532
2012-10-30 17:55 | 28.1679 | 249.0789 | 195.1575 | 246.8897 | 175.6888 | -1.6026 | 1.271 | 32.1351 | 40.8432 | 25.2837 | 2.004 | 32.0606 | 64.2495 | 15.9983 | 14.0731 | -0.1105 | 105.0927 | 235.0058 | 181.0844 | 232.8166 | 161.6157
2012-10-30 17:55 | 28.1679 | 249.0672 | 195.1575 | 247.0269 | 175.724 | -1.6032 | 1.271 | 32.1347 | 40.8426 | 25.2834 | 2.004 | 32.0601 | 64.2484 | 15.9981 | 14.0976 | -0.1105 | 105.0909 | 234.9696 | 181.0599 | 232.9293 | 161.6264
2012-10-30 17:55 | 28.1567 | 249.0599 | 195.2013 | 246.855 | 175.7479 | -1.5991 | 1.271 | 32.1347 | 40.8422 | 25.2837 | 2.004 | 32.06 | 64.2478 | 15.9981 | 14.0823 | -0.1102 | 105.09 | 234.9776 | 181.119 | 232.7727 | 161.6656
2012-10-30 17:55 | 28.1657 | 249.0611 | 195.1789 | 246.9855 | 175.5891 | -1.6061 | 1.2709 | 32.1342 | 40.8408 | 25.2836 | 2.004 | 32.0596 | 64.2469 | 15.9979 | 14.0832 | -0.1107 | 105.0877 | 234.9779 | 181.0957 | 232.9023 | 161.5059
2012-10-30 17:55 | 28.1597 | 249.0668 | 195.1808 | 247.003 | 175.7236 | -1.6026 | 1.2709 | 32.1339 | 40.8399 | 25.2838 | 2.004 | 32.0593 | 64.2456 | 15.9979 | 14.1139 | -0.1105 | 105.0855 | 234.9529 | 181.0669 | 232.8891 | 161.6097
2012-10-30 17:55 | 28.1642 | 249.0557 | 195.2363 | 246.9761 | 175.8374 | -1.6009 | 1.2709 | 32.1337 | 40.8394 | 25.2839 | 2.0039 | 32.0591 | 64.2447 | 15.998 | 14.1755 | -0.1104 | 105.0841 | 234.8802 | 181.0608 | 232.8006 | 161.6619
2012-10-30 17:55 | 28.1791 | 249.1489 | 195.3102 | 246.9284 | 175.946 | -1.5974 | 1.2709 | 32.1337 | 40.8396 | 25.2837 | 2.004 | 32.0592 | 64.2451 | 15.998 | 14.1618 | -0.1101 | 105.0847 | 234.9871 | 181.1484 | 232.7666 | 161.7842
2012-10-30 17:55 | 28.1687 | 249.1071 | 195.2723 | 246.9963 | 176.0254 | -1.6009 | 1.2709 | 32.1332 | 40.8381 | 25.2838 | 2.0039 | 32.0587 | 64.2431 | 15.9979 | 14.1841 | -0.1104 | 105.0812 | 234.923 | 181.0882 | 232.8122 | 161.8413
2012-10-30 17:55 | 28.1717 | 249.1375 | 195.3028 | 246.9484 | 175.8996 | -1.5974 | 1.2709 | 32.1329 | 40.8375 | 25.2837 | 2.0039 | 32.0582 | 64.2416 | 15.9979 | 14.1871 | -0.1101 | 105.0791 | 234.9504 | 181.1157 | 232.7613 | 161.7125
2012-10-30 17:55 | 28.1702 | 249.1282 | 195.3327 | 246.9547 | 175.8005 | -1.602 | 1.2709 | 32.1325 | 40.8361 | 25.2839 | 2.0039 | 32.058 | 64.2401 | 15.998 | 14.1489 | -0.1104 | 105.0762 | 234.9793 | 181.1838 | 232.8058 | 161.6516
2012-10-30 17:55 | 28.1859 | 249.1048 | 195.2777 | 247.0997 | 175.5351 | -1.6049 | 1.2709 | 32.1321 | 40.8358 | 25.2835 | 2.0039 | 32.0575 | 64.2391 | 15.9978 | 14.1686 | -0.1106 | 105.0749 | 234.9362 | 181.1091 | 232.9311 | 161.3665
2012-10-30 17:55 | 28.1844 | 249.1385 | 195.2919 | 246.9102 | 175.5024 | -1.5985 | 1.2709 | 32.132 | 40.8351 | 25.2837 | 2.0039 | 32.0574 | 64.2383 | 15.9979 | 14.2038 | -0.1102 | 105.0734 | 234.9347 | 181.0881 | 232.7064 | 161.2986
2012-10-30 17:55 | 28.1844 | 249.1502 | 195.2408 | 246.9885 | 175.5336 | -1.602 | 1.2709 | 32.132 | 40.8354 | 25.2836 | 2.0039 | 32.0574 | 64.2389 | 15.9978 | 14.2365 | -0.1104 | 105.0743 | 234.9137 | 181.0043 | 232.752 | 161.2971
2012-10-30 17:55 | 28.1956 | 249.138 | 195.252 | 247.0898 | 175.7283 | -1.6026 | 1.2709 | 32.1319 | 40.8356 | 25.2833 | 2.0038 | 32.0572 | 64.2378 | 15.9978 | 14.2354 | -0.1105 | 105.0734 | 234.9026 | 181.0166 | 232.8544 | 161.4929
2012-10-30 17:55 | 28.1821 | 249.1401 | 195.2542 | 247.0136 | 175.867 | -1.6014 | 1.2708 | 32.1312 | 40.8333 | 25.2837 | 2.0038 | 32.0567 | 64.2362 | 15.9977 | 14.2301 | -0.1104 | 105.0695 | 234.91 | 181.0241 | 232.7835 | 161.6369
2012-10-30 17:55 | 28.1949 | 249.1568 | 195.2828 | 246.9598 | 175.9032 | -1.6026 | 1.2709 | 32.1317 | 40.8345 | 25.2836 | 2.0039 | 32.057 | 64.2377 | 15.9976 | 14.2307 | -0.1105 | 105.0723 | 234.9261 | 181.0521 | 232.7291 | 161.6725
2012-10-30 17:55 | 28.1896 | 249.1202 | 195.3364 | 247.1582 | 175.8394 | -1.6078 | 1.2708 | 32.1311 | 40.8329 | 25.2837 | 2.0038 | 32.0566 | 64.236 | 15.9976 | 14.258 | -0.1108 | 105.0689 | 234.8622 | 181.0784 | 232.9002 | 161.5814
2012-10-30 17:55 | 28.1934 | 249.1318 | 195.297 | 246.9623 | 176.015 | -1.6061 | 1.2708 | 32.1308 | 40.8325 | 25.2834 | 2.0038 | 32.0562 | 64.2345 | 15.9976 | 14.2659 | -0.1107 | 105.067 | 234.8659 | 181.0311 | 232.6964 | 161.7491
2012-10-30 17:55 | 28.1911 | 249.1883 | 195.3261 | 246.9874 | 175.8136 | -1.6003 | 1.2708 | 32.1305 | 40.8317 | 25.2835 | 2.0038 | 32.0559 | 64.2332 | 15.9977 | 14.2513 | -0.1103 | 105.0649 | 234.937 | 181.0748 | 232.7361 | 161.5623
2012-10-30 17:55 | 28.1926 | 249.1467 | 195.3433 | 247.0946 | 175.8268 | -1.6078 | 1.2671 | 32.0351 | 40.5906 | 25.2829 | 1.9982 | 31.9662 | 63.8744 | 15.9976 | 14.3059 | -0.1108 | 104.465 | 234.8408 | 181.0374 | 232.7887 | 161.5209
2012-10-30 17:55 | 28.1964 | 249.1583 | 195.3786 | 247.0358 | 176.0102 | -1.6003 | 1.2648 | 31.9754 | 40.4409 | 25.282 | 1.9941 | 31.9014 | 63.6158 | 15.9976 | 14.3341 | -0.1103 | 104.0567 | 234.8242 | 181.0445 | 232.7017 | 161.6761
2012-10-30 17:56 | 28.1866 | 249.1329 | 195.3373 | 246.9202 | 175.852 | -1.5991 | 1.2641 | 31.9603 | 40.4024 | 25.2822 | 1.9932 | 31.8866 | 63.5557 | 15.9978 | 14.3162 | -0.1102 | 103.9581 | 234.8167 | 181.0211 | 232.604 | 161.5358
2012-10-30 17:56 | 28.1889 | 249.1469 | 195.3514 | 247.0909 | 175.745 | -1.6032 | 1.2637 | 31.9496 | 40.3761 | 25.2817 | 1.9925 | 31.8755 | 63.5116 | 15.9978 | 14.3144 | -0.1105 | 103.8876 | 234.8325 | 181.037 | 232.7765 | 161.4306
2012-10-30 17:56 | 28.1919 | 249.1343 | 195.2955 | 247.0313 | 176.115 | -1.602 | 1.2635 | 31.944 | 40.3617 | 25.2819 | 1.9921 | 31.87 | 63.4897 | 15.9978 | 14.3704 | -0.1104 | 103.8514 | 234.7639 | 180.9251 | 232.6609 | 161.7446
2012-10-30 17:56 | 28.1926 | 249.088 | 195.363 | 246.9615 | 176.0259 | -1.6032 | 1.2634 | 31.9413 | 40.3553 | 25.2816 | 1.992 | 31.8672 | 63.4786 | 15.9978 | 14.3344 | -0.1105 | 103.8339 | 234.7536 | 181.0286 | 232.6271 | 161.6915
2012-10-30 17:56 | 28.1971 | 249.0964 | 195.3911 | 247.1501 | 175.9094 | -1.5997 | 1.2634 | 31.9406 | 40.3531 | 25.2819 | 1.9919 | 31.8665 | 63.4756 | 15.9978 | 14.343 | -0.1103 | 103.8287 | 234.7534 | 181.0481 | 232.8071 | 161.5664
2012-10-30 17:56 | 28.1859 | 249.1087 | 195.3641 | 247.0174 | 176.0114 | -1.5991 | 1.2635 | 31.9428 | 40.3589 | 25.2817 | 1.9921 | 31.8687 | 63.4842 | 15.9979 | 14.3644 | -0.1102 | 103.8432 | 234.7443 | 180.9997 | 232.653 | 161.647
2012-10-30 17:56 | 28.1829 | 249.0861 | 195.3769 | 246.94 | 175.7429 | -1.6037 | 1.2635 | 31.9438 | 40.3608 | 25.2821 | 1.9921 | 31.8696 | 63.4878 | 15.9979 | 14.4022 | -0.1106 | 103.8486 | 234.6839 | 180.9747 | 232.5378 | 161.3407
2012-10-30 17:56 | 28.1866 | 249.0663 | 195.3884 | 247.0886 | 175.7115 | -1.5991 | 1.2635 | 31.943 | 40.3593 | 25.2819 | 1.992 | 31.8689 | 63.484 | 15.9981 | 14.4018 | -0.1102 | 103.8433 | 234.6645 | 180.9866 | 232.6868 | 161.3097
2012-10-30 17:56 | 28.1836 | 249.0555 | 195.3579 | 247.023 | 175.7202 | -1.5997 | 1.2636 | 31.946 | 40.3669 | 25.2817 | 1.9922 | 31.8719 | 63.496 | 15.9981 | 14.3988 | -0.1103 | 103.8629 | 234.6567 | 180.9591 | 232.6242 | 161.3214
2012-10-30 17:56 | 28.1814 | 249.0259 | 195.3479 | 246.9032 | 175.8624 | -1.5974 | 1.2655 | 31.9942 | 40.488 | 25.2822 | 1.9954 | 31.9222 | 63.6976 | 15.9979 | 14.3721 | -0.1101 | 104.1857 | 234.6538 | 180.9758 | 232.5311 | 161.4903
2012-10-30 17:56 | 28.1821 | 249.007 | 195.3643 | 247.0215 | 175.867 | -1.6014 | 1.2686 | 32.0749 | 40.6911 | 25.2831 | 2.0003 | 32.0008 | 64.0111 | 15.998 | 14.4381 | -0.1104 | 104.7022 | 234.5689 | 180.9262 | 232.5834 | 161.4289
2012-10-30 17:56 | 28.1821 | 249.0344 | 195.4547 | 247.0489 | 175.9217 | -1.6032 | 1.2692 | 32.0894 | 40.7277 | 25.2833 | 2.0012 | 32.0148 | 64.0677 | 15.9979 | 14.4707 | -0.1105 | 104.7954 | 234.5637 | 180.984 | 232.5782 | 161.451
2012-10-30 17:56 | 28.1702 | 249.046 | 195.4231 | 246.8646 | 175.9059 | -1.5974 | 1.2696 | 32.1012 | 40.7571 | 25.2836 | 2.0019 | 32.0267 | 64.1151 | 15.9979 | 14.4588 | -0.1101 | 104.8723 | 234.5872 | 180.9643 | 232.4058 | 161.4471
2012-10-30 17:56 | 28.1784 | 249.0385 | 195.392 | 247.0374 | 175.8399 | -1.6026 | 1.2701 | 32.1125 | 40.7862 | 25.2834 | 2.0027 | 32.0381 | 64.1617 | 15.9977 | 14.4384 | -0.1105 | 104.948 | 234.6001 | 180.9536 | 232.599 | 161.4015
2012-10-30 17:56 | 28.1679 | 249.0437 | 195.3972 | 246.9877 | 175.8528 | -1.5916 | 1.2704 | 32.1212 | 40.8079 | 25.2836 | 2.0032 | 32.0467 | 64.1956 | 15.9978 | 14.5136 | -0.1097 | 105.0035 | 234.5301 | 180.8836 | 232.4741 | 161.3392
2012-10-30 17:56 | 28.1717 | 248.9888 | 195.4325 | 246.9092 | 175.9191 | -1.6072 | 1.2706 | 32.1267 | 40.8217 | 25.2838 | 2.0035 | 32.0521 | 64.218 | 15.9977 | 14.4766 | -0.1108 | 105.0396 | 234.5122 | 180.9559 | 232.4326 | 161.4425
2012-10-30 17:56 | 28.1784 | 249.0777 | 195.392 | 247.0961 | 175.8985 | -1.6043 | 1.2708 | 32.1296 | 40.829 | 25.2838 | 2.0037 | 32.055 | 64.2297 | 15.9977 | 14.4874 | -0.1106 | 105.0587 | 234.5903 | 180.9046 | 232.6087 | 161.4111
2012-10-30 17:56 | 28.1724 | 249.056 | 195.3978 | 246.9177 | 176.0955 | -1.5997 | 1.2708 | 32.1318 | 40.8341 | 25.2841 | 2.0039 | 32.057 | 64.2378 | 15.9976 | 14.4569 | -0.1103 | 105.0719 | 234.5991 | 180.9409 | 232.4608 | 161.6386
2012-10-30 17:56 | 28.1814 | 249.0533 | 195.45 | 246.9738 | 176.0342 | -1.6003 | 1.2709 | 32.1334 | 40.8383 | 25.284 | 2.004 | 32.0587 | 64.2443 | 15.9977 | 14.4374 | -0.1103 | 105.0826 | 234.6159 | 181.0126 | 232.5364 | 161.5968
2012-10-30 17:56 | 28.1754 | 249.0747 | 195.4283 | 247.003 | 175.9384 | -1.602 | 1.271 | 32.1346 | 40.8416 | 25.2837 | 2.0041 | 32.06 | 64.2499 | 15.9976 | 14.4762 | -0.1104 | 105.0915 | 234.5985 | 180.9521 | 232.5268 | 161.4622
2012-10-30 17:56 | 28.1754 | 249.0434 | 195.4715 | 246.9443 | 176.0516 | -1.5974 | 1.271 | 32.1354 | 40.8428 | 25.2843 | 2.0041 | 32.0607 | 64.2529 | 15.9976 | 14.464 | -0.1101 | 105.0956 | 234.5794 | 181.0075 | 232.4803 | 161.5876
2012-10-30 17:56 | 28.1784 | 249.0973 | 195.5099 | 246.8963 | 176.0976 | -1.6055 | 1.271 | 32.1363 | 40.8459 | 25.2839 | 2.0042 | 32.0618 | 64.2568 | 15.9977 | 14.4507 | -0.1107 | 105.1028 | 234.6466 | 181.0592 | 232.4456 | 161.6469
2012-10-30 17:56 | 28.1874 | 249.0789 | 195.46 | 247.0268 | 176.0168 | -1.5985 | 1.271 | 32.1371 | 40.8478 | 25.284 | 2.0042 | 32.0624 | 64.2598 | 15.9975 | 14.4474 | -0.1102 | 105.1076 | 234.6315 | 181.0126 | 232.5794 | 161.5694
2012-10-30 17:56 | 28.1776 | 249.0886 | 195.5012 | 247.0522 | 176.0538 | -1.5957 | 1.2711 | 32.138 | 40.8499 | 25.284 | 2.0043 | 32.0632 | 64.2633 | 15.9975 | 14.4825 | -0.11 | 105.1132 | 234.6061 | 181.0187 | 232.5697 | 161.5713
2012-10-30 17:56 | 28.1739 | 249.1241 | 195.4818 | 246.9114 | 176.0501 | -1.6037 | 1.2711 | 32.1382 | 40.8503 | 25.2841 | 2.0043 | 32.0635 | 64.2646 | 15.9974 | 14.491 | -0.1106 | 105.1149 | 234.6331 | 180.9908 | 232.4204 | 161.5591
2012-10-30 17:56 | 28.1881 | 249.1344 | 195.5314 | 247.0157 | 176.1073 | -1.6043 | 1.2711 | 32.1386 | 40.8518 | 25.2838 | 2.0043 | 32.064 | 64.2662 | 15.9975 | 14.5297 | -0.1106 | 105.1181 | 234.6047 | 181.0017 | 232.486 | 161.5776
2012-10-30 17:56 | 28.1784 | 249.1208 | 195.506 | 247.0256 | 176.2421 | -1.6009 | 1.2711 | 32.1391 | 40.853 | 25.2839 | 2.0043 | 32.0644 | 64.2681 | 15.9974 | 14.4751 | -0.1104 | 105.1211 | 234.6457 | 181.0309 | 232.5505 | 161.767
2012-10-30 17:56 | 28.1859 | 249.1517 | 195.5253 | 246.9508 | 176.2339 | -1.5997 | 1.2711 | 32.1393 | 40.8527 | 25.2844 | 2.0044 | 32.0646 | 64.2699 | 15.9972 | 14.5397 | -0.1103 | 105.1226 | 234.612 | 180.9856 | 232.4111 | 161.6942
2012-10-30 17:57 | 28.1859 | 249.1048 | 195.5528 | 247.0566 | 176.066 | -1.602 | 1.2711 | 32.1393 | 40.8536 | 25.2839 | 2.0044 | 32.0647 | 64.2702 | 15.9972 | 14.5642 | -0.1104 | 105.1238 | 234.5406 | 180.9886 | 232.4924 | 161.5018
2012-10-30 17:57 | 28.1784 | 249.109 | 195.5453 | 247.0648 | 176.1796 | -1.5991 | 1.2711 | 32.1395 | 40.8538 | 25.284 | 2.0044 | 32.0648 | 64.27 | 15.9974 | 14.4996 | -0.1102 | 105.1238 | 234.6094 | 181.0457 | 232.5652 | 161.68
2012-10-30 17:57 | 28.1859 | 249.1283 | 195.5764 | 246.9508 | 176.109 | -1.6009 | 1.2711 | 32.1393 | 40.8531 | 25.2841 | 2.0044 | 32.0646 | 64.2693 | 15.9974 | 14.5071 | -0.1104 | 105.1224 | 234.6212 | 181.0693 | 232.4437 | 161.6019
2012-10-30 17:57 | 28.1911 | 249.1178 | 195.6012 | 247.1245 | 176.1376 | -1.6014 | 1.2711 | 32.1392 | 40.8527 | 25.2842 | 2.0044 | 32.0645 | 64.269 | 15.9973 | 14.5775 | -0.1104 | 105.1217 | 234.5403 | 181.0237 | 232.547 | 161.5601
2012-10-30 17:57 | 28.1971 | 249.1395 | 195.5993 | 246.9268 | 176.167 | -1.5974 | 1.2711 | 32.1395 | 40.8531 | 25.2844 | 2.0044 | 32.0648 | 64.2696 | 15.9974 | 14.5713 | -0.1101 | 105.1227 | 234.5682 | 181.028 | 232.3555 | 161.5957
2012-10-30 17:57 | 28.1881 | 249.1696 | 195.5589 | 247.0549 | 175.873 | -1.5985 | 1.2711 | 32.1393 | 40.8527 | 25.2844 | 2.0044 | 32.0645 | 64.2698 | 15.9972 | 14.5541 | -0.1102 | 105.1225 | 234.6155 | 181.0048 | 232.5008 | 161.3189
2012-10-30 17:57 | 28.2009 | 249.1824 | 195.6149 | 247.1147 | 176.1201 | -1.6003 | 1.2711 | 32.1392 | 40.8524 | 25.2843 | 2.0044 | 32.0645 | 64.2695 | 15.9972 | 14.5547 | -0.1103 | 105.122 | 234.6277 | 181.0602 | 232.56 | 161.5654
2012-10-30 17:57 | 28.1949 | 249.1451 | 195.5578 | 247.0303 | 176.321 | -1.5985 | 1.2711 | 32.1389 | 40.8519 | 25.2842 | 2.0044 | 32.0643 | 64.2689 | 15.9971 | 14.5773 | -0.1102 | 105.1208 | 234.5678 | 180.9805 | 232.453 | 161.7437
2012-10-30 17:57 | 28.1926 | 249.1584 | 195.5752 | 247.0555 | 176.0493 | -1.6026 | 1.2711 | 32.139 | 40.8526 | 25.284 | 2.0044 | 32.0643 | 64.269 | 15.9971 | 14.5872 | -0.1105 | 105.1216 | 234.5712 | 180.988 | 232.4683 | 161.4621
2012-10-30 17:57 | 28.2091 | 249.1749 | 195.5838 | 247.1268 | 176.101 | -1.6043 | 1.2711 | 32.139 | 40.8515 | 25.2846 | 2.0044 | 32.0643 | 64.2688 | 15.9971 | 14.5792 | -0.1106 | 105.1203 | 234.5957 | 181.0046 | 232.5476 | 161.5218
2012-10-30 17:57 | 28.2039 | 249.1737 | 195.5983 | 247.0472 | 176.166 | -1.6009 | 1.2711 | 32.1388 | 40.8516 | 25.2843 | 2.0044 | 32.0641 | 64.2686 | 15.9971 | 14.5618 | -0.1104 | 105.1202 | 234.6119 | 181.0365 | 232.4854 | 161.6042
2012-10-30 17:57 | 28.2054 | 249.2261 | 195.5526 | 247.08 | 176.0621 | -1.6061 | 1.2711 | 32.1385 | 40.851 | 25.2842 | 2.0044 | 32.0638 | 64.2674 | 15.997 | 14.5633 | -0.1107 | 105.1184 | 234.6628 | 180.9893 | 232.5167 | 161.4988
2012-10-30 17:57 | 28.2166 | 249.2333 | 195.6424 | 247.1578 | 176.1358 | -1.5991 | 1.2711 | 32.1388 | 40.8517 | 25.2843 | 2.0044 | 32.0641 | 64.2687 | 15.997 | 14.5867 | -0.1102 | 105.1203 | 234.6466 | 181.0557 | 232.5711 | 161.5491
2012-10-30 17:57 | 28.2188 | 249.212 | 195.6289 | 246.9837 | 176.1419 | -1.6043 | 1.2711 | 32.1389 | 40.852 | 25.2842 | 2.0044 | 32.0643 | 64.2696 | 15.9969 | 14.5889 | -0.1106 | 105.1217 | 234.6231 | 181.04 | 232.3948 | 161.553
2012-10-30 17:57 | 28.2203 | 249.2331 | 195.599 | 247.1929 | 175.9833 | -1.5991 | 1.2711 | 32.1388 | 40.8519 | 25.284 | 2.0044 | 32.064 | 64.2686 | 15.997 | 14.5986 | -0.1102 | 105.1205 | 234.6345 | 181.0004 | 232.5943 | 161.3847
2012-10-30 17:57 | 28.2226 | 249.2002 | 195.5973 | 247.1481 | 175.9817 | -1.6049 | 1.2711 | 32.1384 | 40.8508 | 25.2841 | 2.0044 | 32.0637 | 64.2677 | 15.9969 | 14.5927 | -0.1106 | 105.1184 | 234.6075 | 181.0046 | 232.5554 | 161.389
2012-10-30 17:57 | 28.2248 | 249.2611 | 195.6153 | 247.0407 | 176.0151 | -1.6049 | 1.2711 | 32.1385 | 40.8514 | 25.2839 | 2.0044 | 32.0638 | 64.2681 | 15.9968 | 14.6031 | -0.1106 | 105.1195 | 234.658 | 181.0122 | 232.4376 | 161.412
2012-10-30 17:57 | 28.2368 | 249.2066 | 195.5801 | 247.2054 | 175.9334 | -1.6037 | 1.2711 | 32.1385 | 40.8505 | 25.2845 | 2.0044 | 32.0639 | 64.2685 | 15.9969 | 14.5825 | -0.1106 | 105.1191 | 234.6241 | 180.9976 | 232.6229 | 161.3509
2012-10-30 17:57 | 28.2353 | 249.2364 | 195.5865 | 247.1452 | 175.9436 | -1.5951 | 1.2711 | 32.1381 | 40.8497 | 25.2844 | 2.0044 | 32.0634 | 64.2672 | 15.9967 | 14.5973 | -0.11 | 105.1169 | 234.6391 | 180.9892 | 232.5479 | 161.3463
2012-10-30 17:57 | 28.2451 | 249.2266 | 195.5727 | 247.1393 | 175.8871 | -1.5962 | 1.2711 | 32.1382 | 40.8505 | 25.2839 | 2.0044 | 32.0634 | 64.2666 | 15.9969 | 14.5785 | -0.11 | 105.1171 | 234.6481 | 180.9942 | 232.5608 | 161.3086
2012-10-30 17:57 | 28.2413 | 249.2228 | 195.5492 | 247.2413 | 175.8325 | -1.6003 | 1.2711 | 32.1375 | 40.8485 | 25.2841 | 2.0043 | 32.0628 | 64.2649 | 15.9967 | 14.6318 | -0.1103 | 105.1134 | 234.591 | 180.9174 | 232.6095 | 161.2007
2012-10-30 17:57 | 28.248 | 249.2138 | 195.5716 | 247.0756 | 175.8978 | -1.5957 | 1.2711 | 32.1374 | 40.8485 | 25.284 | 2.0043 | 32.0627 | 64.2642 | 15.9967 | 14.6059 | -0.11 | 105.1127 | 234.6079 | 180.9657 | 232.4697 | 161.2919
2012-10-30 17:57 | 28.2548 | 249.2637 | 195.4881 | 247.1882 | 175.8382 | -1.5968 | 1.2711 | 32.1375 | 40.8484 | 25.2842 | 2.0043 | 32.0628 | 64.2648 | 15.9967 | 14.5964 | -0.1101 | 105.1132 | 234.6673 | 180.8917 | 232.5918 | 161.2418
2012-10-30 17:57 | 28.2563 | 249.2417 | 195.5328 | 247.2484 | 176.0037 | -1.5997 | 1.2711 | 32.1373 | 40.8482 | 25.284 | 2.0043 | 32.0627 | 64.2642 | 15.9967 | 14.6387 | -0.1103 | 105.1124 | 234.603 | 180.8941 | 232.6097 | 161.365
2012-10-30 17:57 | 28.2578 | 249.2863 | 195.5421 | 247.0697 | 176.0755 | -1.6066 | 1.271 | 32.1372 | 40.8476 | 25.2842 | 2.0043 | 32.0625 | 64.2638 | 15.9966 | 14.6116 | -0.1108 | 105.1114 | 234.6747 | 180.9305 | 232.4581 | 161.4639
2012-10-30 17:57 | 28.2578 | 249.2863 | 195.5382 | 247.2421 | 176.1067 | -1.602 | 1.271 | 32.1369 | 40.8473 | 25.284 | 2.0043 | 32.0623 | 64.2626 | 15.9967 | 14.6402 | -0.1104 | 105.1099 | 234.6461 | 180.898 | 232.6019 | 161.4665
2012-10-30 17:57 | 28.2638 | 249.2766 | 195.5324 | 247.2246 | 176.0619 | -1.6014 | 1.271 | 32.1363 | 40.8454 | 25.2841 | 2.0043 | 32.0616 | 64.2606 | 15.9966 | 14.6217 | -0.1104 | 105.1061 | 234.6549 | 180.9107 | 232.6029 | 161.4402
2012-10-30 17:58 | 28.266 | 249.2827 | 195.5739 | 247.0897 | 176.111 | -1.6003 | 1.271 | 32.1361 | 40.8456 | 25.2837 | 2.0043 | 32.0614 | 64.2594 | 15.9966 | 14.6565 | -0.1103 | 105.105 | 234.6262 | 180.9174 | 232.4332 | 161.4545
2012-10-30 17:58 | 28.2675 | 249.3195 | 195.5872 | 247.2831 | 175.9837 | -1.6055 | 1.271 | 32.1355 | 40.8436 | 25.2841 | 2.0042 | 32.0608 | 64.2576 | 15.9965 | 14.6825 | -0.1107 | 105.1012 | 234.637 | 180.9047 | 232.6006 | 161.3012
2012-10-30 17:58 | 28.263 | 249.2719 | 195.6102 | 247.1768 | 176.108 | -1.6037 | 1.2709 | 32.1349 | 40.8418 | 25.2842 | 2.0042 | 32.0602 | 64.2553 | 15.9965 | 14.6657 | -0.1106 | 105.097 | 234.6062 | 180.9445 | 232.5111 | 161.4423
2012-10-30 17:58 | 28.2668 | 249.2718 | 195.5551 | 247.1258 | 176.0298 | -1.602 | 1.2709 | 32.1342 | 40.8404 | 25.284 | 2.0042 | 32.0596 | 64.2528 | 15.9964 | 14.6532 | -0.1104 | 105.0932 | 234.6186 | 180.9019 | 232.4726 | 161.3766
2012-10-30 17:58 | 28.2802 | 249.2539 | 195.5921 | 247.3076 | 175.9222 | -1.6026 | 1.2709 | 32.1337 | 40.8392 | 25.284 | 2.0041 | 32.059 | 64.2507 | 15.9964 | 14.6748 | -0.1105 | 105.0899 | 234.5791 | 180.9173 | 232.6328 | 161.2474
2012-10-30 17:58 | 28.269 | 249.2662 | 195.5809 | 247.0966 | 175.868 | -1.5985 | 1.2709 | 32.1338 | 40.8391 | 25.2841 | 2.0041 | 32.0591 | 64.2511 | 15.9964 | 14.6473 | -0.1102 | 105.0902 | 234.6189 | 180.9336 | 232.4493 | 161.2207
2012-10-30 17:58 | 28.2758 | 249.2965 | 195.5012 | 247.2405 | 175.8514 | -1.6037 | 1.2709 | 32.1338 | 40.8392 | 25.2841 | 2.0041 | 32.0591 | 64.2513 | 15.9964 | 14.6255 | -0.1106 | 105.0904 | 234.671 | 180.8757 | 232.615 | 161.2259
2012-10-30 17:58 | 28.2817 | 249.2123 | 195.5268 | 247.2778 | 175.8651 | -1.6043 | 1.2709 | 32.1333 | 40.8378 | 25.2841 | 2.0041 | 32.0586 | 64.2497 | 15.9963 | 14.6477 | -0.1106 | 105.0875 | 234.5646 | 180.8791 | 232.6301 | 161.2174
2012-10-30 17:58 | 28.281 | 249.2977 | 195.4671 | 247.136 | 175.8019 | -1.6009 | 1.2709 | 32.1331 | 40.8379 | 25.2837 | 2.0041 | 32.0584 | 64.2488 | 15.9963 | 14.6511 | -0.1104 | 105.0868 | 234.6466 | 180.816 | 232.4849 | 161.1508
2012-10-30 17:58 | 28.2885 | 249.2974 | 195.5493 | 247.2611 | 175.8016 | -1.5974 | 1.2709 | 32.1328 | 40.8369 | 25.2839 | 2.0041 | 32.0581 | 64.2477 | 15.9962 | 14.6179 | -0.1101 | 105.0846 | 234.6795 | 180.9314 | 232.6432 | 161.1837
2012-10-30 17:58 | 28.293 | 249.2471 | 195.5027 | 247.2616 | 175.9935 | -1.6055 | 1.2709 | 32.1327 | 40.8371 | 25.2837 | 2.0041 | 32.058 | 64.2471 | 15.9963 | 14.6835 | -0.1107 | 105.0842 | 234.5636 | 180.8192 | 232.5781 | 161.31
2012-10-30 17:58 | 28.29 | 249.2832 | 195.4958 | 247.2038 | 175.9359 | -1.6026 | 1.2709 | 32.132 | 40.8351 | 25.2838 | 2.0041 | 32.0575 | 64.2453 | 15.9962 | 14.656 | -0.1105 | 105.0804 | 234.6272 | 180.8398 | 232.5478 | 161.2799
2012-10-30 17:58 | 28.2877 | 249.2575 | 195.462 | 247.2916 | 175.8594 | -1.5939 | 1.2708 | 32.1319 | 40.8345 | 25.2841 | 2.0041 | 32.0574 | 64.2446 | 15.9963 | 14.6456 | -0.1099 | 105.0791 | 234.6119 | 180.8164 | 232.646 | 161.2138
2012-10-30 17:58 | 28.2982 | 249.2993 | 195.5197 | 247.1297 | 176.0105 | -1.5985 | 1.2708 | 32.1316 | 40.8333 | 25.2842 | 2.004 | 32.0568 | 64.2425 | 15.9963 | 14.6887 | -0.1102 | 105.0757 | 234.6106 | 180.831 | 232.441 | 161.3218
2012-10-30 17:58 | 28.3042 | 249.2896 | 195.5178 | 247.2493 | 175.9423 | -1.6055 | 1.2708 | 32.1306 | 40.8314 | 25.2839 | 2.004 | 32.056 | 64.2392 | 15.9962 | 14.6825 | -0.1107 | 105.0706 | 234.6071 | 180.8353 | 232.5668 | 161.2598
2012-10-30 17:58 | 28.3012 | 249.2905 | 195.5148 | 247.3443 | 175.9119 | -1.6009 | 1.2708 | 32.1304 | 40.8307 | 25.284 | 2.004 | 32.0556 | 64.2379 | 15.9962 | 14.6387 | -0.1104 | 105.0686 | 234.6518 | 180.8761 | 232.7056 | 161.2732
2012-10-30 17:58 | 28.305 | 249.3022 | 195.5383 | 247.1522 | 176.107 | -1.6014 | 1.2708 | 32.1299 | 40.8294 | 25.284 | 2.0039 | 32.0552 | 64.236 | 15.9963 | 14.7485 | -0.1104 | 105.0654 | 234.5537 | 180.7898 | 232.4037 | 161.3585
2012-10-30 17:58 | 28.3162 | 249.3212 | 195.4748 | 247.2888 | 175.9777 | -1.5962 | 1.2684 | 32.0694 | 40.6762 | 25.2838 | 2.0005 | 32.0008 | 64.0182 | 15.9962 | 14.7189 | -0.11 | 104.6944 | 234.6023 | 180.7559 | 232.5699 | 161.2588
2012-10-30 17:58 | 28.3199 | 249.321 | 195.5414 | 247.3277 | 176.1688 | -1.5974 | 1.2647 | 31.9744 | 40.4381 | 25.2821 | 1.9942 | 31.8997 | 63.6142 | 15.9963 | 14.7226 | -0.1101 | 104.0523 | 234.5984 | 180.8188 | 232.6051 | 161.4462
2012-10-30 17:58 | 28.3117 | 249.2971 | 195.5528 | 247.1981 | 176.2192 | -1.602 | 1.2642 | 31.9627 | 40.4082 | 25.2824 | 1.9935 | 31.8886 | 63.5697 | 15.9963 | 14.7144 | -0.1104 | 103.978 | 234.5827 | 180.8384 | 232.4837 | 161.5048
2012-10-30 17:58 | 28.3199 | 249.2975 | 195.561 | 247.3042 | 176.1219 | -1.5997 | 1.2637 | 31.9497 | 40.3757 | 25.2822 | 1.9927 | 31.8757 | 63.5188 | 15.9962 | 14.7756 | -0.1103 | 103.8945 | 234.5219 | 180.7854 | 232.5286 | 161.3463
2012-10-30 17:58 | 28.3147 | 249.2649 | 195.5165 | 247.3186 | 176.1441 | -1.6003 | 1.2635 | 31.9448 | 40.3635 | 25.282 | 1.9923 | 31.8706 | 63.4972 | 15.9965 | 14.7827 | -0.1103 | 103.8607 | 234.4822 | 180.7338 | 232.5359 | 161.3614
2012-10-30 17:58 | 28.3237 | 249.2582 | 195.5216 | 247.1787 | 175.5909 | -1.6049 | 1.2634 | 31.9422 | 40.3568 | 25.2821 | 1.9922 | 31.8679 | 63.4868 | 15.9964 | 14.7998 | -0.1106 | 103.8436 | 234.4584 | 180.7218 | 232.3789 | 160.7911
2012-10-30 17:58 | 28.3222 | 249.2293 | 195.4808 | 247.2752 | 175.8119 | -1.6049 | 1.2634 | 31.9423 | 40.3571 | 25.282 | 1.9922 | 31.8681 | 63.4875 | 15.9965 | 14.782 | -0.1106 | 103.8445 | 234.4473 | 180.6988 | 232.4932 | 161.0299
2012-10-30 17:58 | 28.3229 | 249.183 | 195.454 | 247.2955 | 175.9922 | -1.5962 | 1.2635 | 31.9428 | 40.3582 | 25.2821 | 1.9922 | 31.8687 | 63.4891 | 15.9966 | 14.8031 | -0.11 | 103.8473 | 234.3799 | 180.6509 | 232.4924 | 161.1891
2012-10-30 17:58 | 28.3214 | 249.1855 | 195.4879 | 247.1255 | 175.9829 | -1.6003 | 1.2635 | 31.9428 | 40.3588 | 25.2818 | 1.9922 | 31.8687 | 63.4899 | 15.9965 | 14.7812 | -0.1103 | 103.8486 | 234.4043 | 180.7067 | 232.3443 | 161.2017
2012-10-30 17:58 | 28.3289 | 249.1969 | 195.4718 | 247.2466 | 175.2799 | -1.5974 | 1.2635 | 31.9447 | 40.3627 | 25.2824 | 1.9923 | 31.8705 | 63.4971 | 15.9965 | 14.805 | -0.1101 | 103.8598 | 234.3919 | 180.6668 | 232.4416 | 160.4749
2012-10-30 17:58 | 28.3207 | 249.1691 | 195.3653 | 247.2619 | 175.2249 | -1.5991 | 1.2636 | 31.9475 | 40.3701 | 25.2821 | 1.9925 | 31.8731 | 63.5071 | 15.9966 | 14.809 | -0.1102 | 103.8772 | 234.3601 | 180.5563 | 232.4529 | 160.4159
2012-10-30 17:59 | 28.3289 | 249.1773 | 195.3264 | 247.1056 | 175.1863 | -1.6072 | 1.2657 | 32.0003 | 40.5028 | 25.2827 | 1.9958 | 31.9261 | 63.7185 | 15.9966 | 14.8295 | -0.1108 | 104.2213 | 234.3478 | 180.4969 | 232.2761 | 160.3568
2012-10-30 17:59 | 28.3259 | 249.2056 | 195.288 | 247.2084 | 175.0155 | -1.598 | 1.2687 | 32.0767 | 40.6952 | 25.2835 | 2.0006 | 32.0021 | 64.0226 | 15.9965 | 14.802 | -0.1102 | 104.7178 | 234.4036 | 180.486 | 232.4064 | 160.2135
2012-10-30 17:59 | 28.3259 | 249.1626 | 195.2094 | 247.2632 | 175.4838 | -1.6049 | 1.2692 | 32.0907 | 40.7309 | 25.2834 | 2.0015 | 32.0159 | 64.0783 | 15.9963 | 14.8183 | -0.1106 | 104.8092 | 234.3443 | 180.3911 | 232.4449 | 160.6655
2012-10-30 17:59 | 28.3289 | 249.189 | 195.2792 | 247.133 | 175.6859 | -1.5985 | 1.2698 | 32.1042 | 40.7645 | 25.2837 | 2.0023 | 32.0295 | 64.1325 | 15.9964 | 14.8824 | -0.1102 | 104.8971 | 234.3066 | 180.3968 | 232.2506 | 160.8035
2012-10-30 17:59 | 28.3207 | 249.1848 | 195.2357 | 247.0935 | 175.5098 | -1.6009 | 1.2702 | 32.115 | 40.7915 | 25.284 | 2.003 | 32.0402 | 64.1762 | 15.9962 | 14.8376 | -0.1104 | 104.9676 | 234.3472 | 180.3981 | 232.2559 | 160.6722
2012-10-30 17:59 | 28.3364 | 249.2279 | 195.2789 | 247.2149 | 175.5685 | -1.6049 | 1.2704 | 32.1218 | 40.8088 | 25.284 | 2.0034 | 32.0469 | 64.2028 | 15.9963 | 14.8981 | -0.1106 | 105.0116 | 234.3298 | 180.3808 | 232.3168 | 160.6704
2012-10-30 17:59 | 28.3237 | 249.1721 | 195.2465 | 247.1474 | 175.7471 | -1.598 | 1.2706 | 32.1266 | 40.821 | 25.284 | 2.0037 | 32.0518 | 64.2231 | 15.9961 | 14.865 | -0.1102 | 105.0441 | 234.3071 | 180.3815 | 232.2824 | 160.8821
2012-10-30 17:59 | 28.3289 | 249.2008 | 195.2635 | 247.0155 | 175.8303 | -1.6055 | 1.2707 | 32.1297 | 40.8287 | 25.2841 | 2.0039 | 32.0549 | 64.2355 | 15.9961 | 14.8702 | -0.1107 | 105.0642 | 234.3306 | 180.3933 | 232.1453 | 160.9601
2012-10-30 17:59 | 28.3252 | 249.1814 | 195.3148 | 247.2077 | 175.8774 | -1.6043 | 1.2708 | 32.1317 | 40.8338 | 25.2842 | 2.004 | 32.057 | 64.2431 | 15.9963 | 14.8787 | -0.1106 | 105.077 | 234.3027 | 180.4361 | 232.329 | 160.9987
2012-10-30 17:59 | 28.3199 | 249.2074 | 195.3999 | 247.0692 | 175.8565 | -1.5997 | 1.2709 | 32.1331 | 40.8366 | 25.2845 | 2.0041 | 32.0582 | 64.2483 | 15.9962 | 14.8816 | -0.1103 | 105.085 | 234.3258 | 180.5183 | 232.1876 | 160.9749
2012-10-30 17:59 | 28.3207 | 249.1691 | 195.3693 | 247.0935 | 175.7909 | -1.602 | 1.2709 | 32.1341 | 40.8399 | 25.2841 | 2.0042 | 32.0595 | 64.253 | 15.9963 | 14.9354 | -0.1104 | 105.093 | 234.2337 | 180.4339 | 232.1581 | 160.8555
2012-10-30 17:59 | 28.3267 | 249.2103 | 195.3871 | 247.17 | 175.7618 | -1.598 | 1.2709 | 32.1349 | 40.8417 | 25.2843 | 2.0042 | 32.0602 | 64.2562 | 15.9962 | 14.8599 | -0.1102 | 105.0979 | 234.3504 | 180.5272 | 232.3101 | 160.9019
2012-10-30 17:59 | 28.3304 | 249.1984 | 195.3947 | 247.0562 | 175.9255 | -1.6009 | 1.271 | 32.1355 | 40.8433 | 25.2841 | 2.0043 | 32.0607 | 64.2582 | 15.9962 | 14.888 | -0.1104 | 105.1015 | 234.3104 | 180.5067 | 232.1682 | 161.0375
2012-10-30 17:59 | 28.3237 | 249.1956 | 195.388 | 247.073 | 175.9305 | -1.6014 | 1.271 | 32.1357 | 40.8436 | 25.2844 | 2.0043 | 32.061 | 64.2593 | 15.9962 | 14.8446 | -0.1104 | 105.1029 | 234.351 | 180.5434 | 232.2284 | 161.0859
2012-10-30 17:59 | 28.3364 | 249.1965 | 195.4203 | 247.2228 | 175.9003 | -1.6043 | 1.271 | 32.1363 | 40.8447 | 25.2846 | 2.0043 | 32.0616 | 64.262 | 15.9961 | 14.8614 | -0.1106 | 105.1067 | 234.3351 | 180.5589 | 232.3614 | 161.0389
2012-10-30 17:59 | 28.3327 | 249.2085 | 195.3852 | 247.1211 | 176.045 | -1.6009 | 1.271 | 32.137 | 40.8464 | 25.2846 | 2.0044 | 32.0622 | 64.2646 | 15.9961 | 14.8536 | -0.1104 | 105.111 | 234.3549 | 180.5316 | 232.2675 | 161.1914
2012-10-30 17:59 | 28.3282 | 249.2236 | 195.4554 | 247.0187 | 176.0444 | -1.6037 | 1.271 | 32.1369 | 40.8466 | 25.2843 | 2.0044 | 32.0621 | 64.2644 | 15.996 | 14.8654 | -0.1106 | 105.111 | 234.3582 | 180.59 | 232.1533 | 161.179
2012-10-30 17:59 | 28.3387 | 249.1754 | 195.4423 | 247.1428 | 175.8753 | -1.6037 | 1.271 | 32.1373 | 40.8475 | 25.2844 | 2.0044 | 32.0624 | 64.2662 | 15.996 | 14.8433 | -0.1106 | 105.1137 | 234.3321 | 180.599 | 232.2995 | 161.032
2012-10-30 17:59 | 28.3387 | 249.1871 | 195.4266 | 247.2251 | 176.0041 | -1.5991 | 1.2711 | 32.1378 | 40.8488 | 25.2844 | 2.0044 | 32.0631 | 64.2687 | 15.996 | 14.8433 | -0.1102 | 105.1175 | 234.3438 | 180.5833 | 232.3818 | 161.1608
2012-10-30 17:59 | 28.3409 | 249.2128 | 195.4798 | 247.0627 | 176.0141 | -1.6049 | 1.2711 | 32.1384 | 40.8505 | 25.2843 | 2.0045 | 32.0636 | 64.2711 | 15.9959 | 14.8496 | -0.1106 | 105.1216 | 234.3632 | 180.6302 | 232.2131 | 161.1645
2012-10-30 17:59 | 28.3454 | 249.2016 | 195.4608 | 247.1025 | 175.9054 | -1.6014 | 1.2711 | 32.1384 | 40.8505 | 25.2843 | 2.0045 | 32.0637 | 64.2717 | 15.9959 | 14.8337 | -0.1104 | 105.1222 | 234.3679 | 180.6271 | 232.2688 | 161.0717
2012-10-30 17:59 | 28.3477 | 249.2078 | 195.4788 | 247.2145 | 175.8686 | -1.6009 | 1.2711 | 32.1386 | 40.8507 | 25.2844 | 2.0045 | 32.0637 | 64.2717 | 15.9958 | 14.8238 | -0.1104 | 105.1223 | 234.384 | 180.655 | 232.3907 | 161.0448
2012-10-30 17:59 | 28.3521 | 249.224 | 195.4557 | 247.164 | 175.8809 | -1.5933 | 1.2711 | 32.1387 | 40.8519 | 25.284 | 2.0045 | 32.0639 | 64.2725 | 15.9958 | 14.8323 | -0.1098 | 105.1244 | 234.3917 | 180.6234 | 232.3317 | 161.0486
2012-10-30 17:59 | 28.3596 | 249.208 | 195.475 | 247.0462 | 176.0055 | -1.6003 | 1.2711 | 32.1387 | 40.8509 | 25.2845 | 2.0045 | 32.0639 | 64.273 | 15.9958 | 14.8724 | -0.1103 | 105.124 | 234.3356 | 180.6026 | 232.1738 | 161.1331
2012-10-30 17:59 | 28.3536 | 249.2607 | 195.469 | 247.0989 | 175.9448 | -1.6037 | 1.2711 | 32.1386 | 40.8507 | 25.2846 | 2.0045 | 32.064 | 64.2735 | 15.9957 | 14.8053 | -0.1106 | 105.1242 | 234.4554 | 180.6637 | 232.2936 | 161.1395
2012-10-30 17:59 | 28.3626 | 249.2227 | 195.4544 | 247.155 | 175.9109 | -1.602 | 1.2711 | 32.1384 | 40.8501 | 25.2846 | 2.0045 | 32.0636 | 64.272 | 15.9957 | 14.8346 | -0.1104 | 105.1221 | 234.3881 | 180.6198 | 232.3204 | 161.0763
2012-10-30 18:00 | 28.3664 | 249.2422 | 195.4739 | 247.1548 | 175.8405 | -1.5939 | 1.271 | 32.1376 | 40.8479 | 25.2846 | 2.0045 | 32.0628 | 64.2685 | 15.9958 | 14.8344 | -0.1099 | 105.1164 | 234.4078 | 180.6395 | 232.3204 | 161.0061
2012-10-30 18:00 | 28.3701 | 249.2068 | 195.4422 | 247.139 | 176.0433 | -1.6003 | 1.271 | 32.1372 | 40.8467 | 25.2849 | 2.0044 | 32.0624 | 64.2672 | 15.9957 | 14.8544 | -0.1103 | 105.1139 | 234.3524 | 180.5878 | 232.2846 | 161.1889
2012-10-30 18:00 | 28.3701 | 249.2498 | 195.4737 | 247.0253 | 175.9535 | -1.6055 | 1.271 | 32.137 | 40.8465 | 25.2845 | 2.0044 | 32.0622 | 64.2665 | 15.9957 | 14.8421 | -0.1107 | 105.113 | 234.4077 | 180.6316 | 232.1832 | 161.1114
2012-10-30 18:00 | 28.3724 | 249.2404 | 195.4445 | 247.1648 | 175.9714 | -1.598 | 1.271 | 32.1371 | 40.8472 | 25.2843 | 2.0044 | 32.0624 | 64.2672 | 15.9957 | 14.873 | -0.1102 | 105.1144 | 234.3674 | 180.5715 | 232.2918 | 161.0984
2012-10-30 18:00 | 28.3806 | 249.1898 | 195.4567 | 247.2239 | 176.0343 | -1.6055 | 1.271 | 32.1369 | 40.8464 | 25.2845 | 2.0044 | 32.0622 | 64.2667 | 15.9956 | 14.8649 | -0.1107 | 105.1131 | 234.3249 | 180.5918 | 232.359 | 161.1694
2012-10-30 18:00 | 28.3814 | 249.2572 | 195.4535 | 247.1542 | 176.0078 | -1.6049 | 1.271 | 32.1369 | 40.8464 | 25.2845 | 2.0044 | 32.0621 | 64.2662 | 15.9957 | 14.8575 | -0.1106 | 105.1126 | 234.3997 | 180.596 | 232.2967 | 161.1503
2012-10-30 18:00 | 28.3911 | 249.2747 | 195.4986 | 247.1051 | 176.0448 | -1.6009 | 1.271 | 32.1366 | 40.8458 | 25.2844 | 2.0044 | 32.0618 | 64.2651 | 15.9956 | 14.8387 | -0.1104 | 105.1109 | 234.436 | 180.6599 | 232.2664 | 161.2061
2012-10-30 18:00 | 28.3896 | 249.2497 | 195.4617 | 247.1154 | 176.0121 | -1.6066 | 1.271 | 32.1364 | 40.8452 | 25.2845 | 2.0044 | 32.0617 | 64.2647 | 15.9956 | 14.8657 | -0.1108 | 105.1098 | 234.384 | 180.596 | 232.2497 | 161.1464
2012-10-30 18:00 | 28.3918 | 249.2559 | 195.4797 | 247.192 | 175.9284 | -1.598 | 1.271 | 32.136 | 40.8439 | 25.2847 | 2.0044 | 32.0613 | 64.2635 | 15.9955 | 14.8761 | -0.1102 | 105.1075 | 234.3798 | 180.6036 | 232.3159 | 161.0523
2012-10-30 18:00 | 28.3971 | 249.2533 | 195.4968 | 247.2365 | 176 | -1.6003 | 1.271 | 32.1362 | 40.8451 | 25.2841 | 2.0044 | 32.0614 | 64.2638 | 15.9955 | 14.8732 | -0.1103 | 105.1088 | 234.3801 | 180.6236 | 232.3633 | 161.1268
2012-10-30 18:00 | 28.4053 | 249.2733 | 195.5403 | 247.1193 | 176.016 | -1.6037 | 1.271 | 32.136 | 40.8441 | 25.2845 | 2.0044 | 32.0612 | 64.2634 | 15.9954 | 14.8447 | -0.1106 | 105.1075 | 234.4286 | 180.6956 | 232.2746 | 161.1713
2012-10-30 18:00 | 28.4008 | 249.2883 | 195.469 | 247.107 | 175.9764 | -1.6003 | 1.271 | 32.1358 | 40.8438 | 25.2844 | 2.0044 | 32.061 | 64.2629 | 15.9954 | 14.8932 | -0.1103 | 105.1066 | 234.3951 | 180.5758 | 232.2138 | 161.0832
2012-10-30 18:00 | 28.4038 | 249.2835 | 195.5035 | 247.2353 | 176.0184 | -1.602 | 1.271 | 32.1361 | 40.8448 | 25.2842 | 2.0044 | 32.0613 | 64.2643 | 15.9953 | 14.8555 | -0.1104 | 105.1091 | 234.428 | 180.648 | 232.3798 | 161.1629
2012-10-30 18:00 | 28.4128 | 249.2964 | 195.536 | 247.2052 | 176.0079 | -1.5974 | 1.271 | 32.1356 | 40.8433 | 25.2843 | 2.0044 | 32.0609 | 64.2626 | 15.9953 | 14.8441 | -0.1101 | 105.106 | 234.4523 | 180.6919 | 232.3611 | 161.1638
2012-10-30 18:00 | 28.4121 | 249.2957 | 195.5039 | 247.2084 | 176.179 | -1.5991 | 1.2709 | 32.135 | 40.8417 | 25.2845 | 2.0044 | 32.0603 | 64.2609 | 15.9952 | 14.8923 | -0.1102 | 105.1026 | 234.4034 | 180.6116 | 232.3161 | 161.2867
2012-10-30 18:00 | 28.4068 | 249.3413 | 195.5104 | 247.1051 | 176.1425 | -1.6026 | 1.2709 | 32.1349 | 40.8413 | 25.2844 | 2.0044 | 32.0601 | 64.2601 | 15.9952 | 14.8748 | -0.1105 | 105.1014 | 234.4665 | 180.6356 | 232.2303 | 161.2677
2012-10-30 18:00 | 28.4188 | 249.3103 | 195.5224 | 247.1406 | 176.0725 | -1.5968 | 1.271 | 32.1352 | 40.8424 | 25.2843 | 2.0044 | 32.0604 | 64.2614 | 15.9951 | 14.9275 | -0.1101 | 105.1038 | 234.3828 | 180.5949 | 232.2131 | 161.145
2012-10-30 18:00 | 28.4211 | 249.3165 | 195.5718 | 247.2448 | 176.0631 | -1.5951 | 1.2709 | 32.135 | 40.8417 | 25.2844 | 2.0044 | 32.0602 | 64.2606 | 15.9951 | 14.9013 | -0.11 | 105.1023 | 234.4152 | 180.6705 | 232.3435 | 161.1618
2012-10-30 18:00 | 28.427 | 249.3537 | 195.5306 | 247.2585 | 176.0651 | -1.5997 | 1.2709 | 32.1344 | 40.8404 | 25.2844 | 2.0043 | 32.0596 | 64.2583 | 15.9951 | 14.8868 | -0.1103 | 105.0987 | 234.4669 | 180.6438 | 232.3717 | 161.1783
2012-10-30 18:00 | 28.4323 | 249.3159 | 195.5555 | 247.162 | 176.3242 | -1.6061 | 1.2709 | 32.134 | 40.8386 | 25.2848 | 2.0043 | 32.0593 | 64.2571 | 15.9951 | 14.9247 | -0.1107 | 105.0957 | 234.3912 | 180.6308 | 232.2373 | 161.3995
2012-10-30 18:00 | 28.4248 | 249.375 | 195.6463 | 247.1506 | 176.2503 | -1.6032 | 1.2709 | 32.1337 | 40.8381 | 25.2847 | 2.0043 | 32.0589 | 64.2557 | 15.9951 | 14.9254 | -0.1105 | 105.0938 | 234.4496 | 180.7209 | 232.2252 | 161.3249
2012-10-30 18:00 | 28.4315 | 249.3621 | 195.6294 | 247.2709 | 176.2023 | -1.6061 | 1.2709 | 32.1334 | 40.8376 | 25.2844 | 2.0043 | 32.0587 | 64.2546 | 15.9951 | 14.9239 | -0.1107 | 105.0921 | 234.4382 | 180.7055 | 232.347 | 161.2784
2012-10-30 18:00 | 28.4338 | 249.3527 | 195.6513 | 247.2222 | 176.2944 | -1.6061 | 1.2709 | 32.1328 | 40.8361 | 25.2843 | 2.0042 | 32.058 | 64.2522 | 15.995 | 14.9018 | -0.1107 | 105.0884 | 234.4509 | 180.7495 | 232.3204 | 161.3926
2012-10-30 18:00 | 28.439 | 249.3148 | 195.6487 | 247.153 | 176.3543 | -1.5985 | 1.2708 | 32.1324 | 40.8352 | 25.2844 | 2.0042 | 32.0576 | 64.2505 | 15.9951 | 14.964 | -0.1102 | 105.0857 | 234.3508 | 180.6847 | 232.189 | 161.3903
2012-10-30 18:00 | 28.445 | 249.3952 | 195.6508 | 247.1903 | 176.329 | -1.6009 | 1.2708 | 32.1318 | 40.8336 | 25.2844 | 2.0042 | 32.0571 | 64.2487 | 15.995 | 14.9456 | -0.1104 | 105.0823 | 234.4496 | 180.7052 | 232.2447 | 161.3834
2012-10-30 18:01 | 28.442 | 249.3726 | 195.6831 | 247.2461 | 176.1582 | -1.6003 | 1.2708 | 32.1313 | 40.8318 | 25.2848 | 2.0042 | 32.0565 | 64.2468 | 15.9949 | 14.9548 | -0.1103 | 105.0787 | 234.4178 | 180.7283 | 232.2913 | 161.2034
2012-10-30 18:01 | 28.4533 | 249.3839 | 195.6512 | 247.2887 | 176.1187 | -1.602 | 1.2708 | 32.1309 | 40.8313 | 25.2843 | 2.0041 | 32.0562 | 64.2453 | 15.9949 | 14.9661 | -0.1104 | 105.0766 | 234.4178 | 180.6851 | 232.3226 | 161.1526
2012-10-30 18:01 | 28.4525 | 249.4223 | 195.7054 | 247.19 | 176.0554 | -1.5997 | 1.2708 | 32.1304 | 40.83 | 25.2844 | 2.0041 | 32.0556 | 64.2432 | 15.9948 | 15.002 | -0.1103 | 105.0732 | 234.4203 | 180.7034 | 232.188 | 161.0534
2012-10-30 18:01 | 28.463 | 249.3936 | 195.6255 | 247.1966 | 175.902 | -1.5985 | 1.2707 | 32.1299 | 40.8287 | 25.2845 | 2.0041 | 32.0552 | 64.2418 | 15.9948 | 14.931 | -0.1102 | 105.0705 | 234.4626 | 180.6945 | 232.2656 | 160.971
2012-10-30 18:01 | 28.4592 | 249.3546 | 195.6374 | 247.2829 | 175.824 | -1.6003 | 1.2707 | 32.1297 | 40.8281 | 25.2845 | 2.0041 | 32.055 | 64.241 | 15.9948 | 14.9353 | -0.1103 | 105.069 | 234.4193 | 180.7021 | 232.3476 | 160.8887
2012-10-30 18:01 | 28.4682 | 249.4184 | 195.5639 | 247.3272 | 175.9072 | -1.6009 | 1.2707 | 32.1293 | 40.8274 | 25.2843 | 2.0041 | 32.0546 | 64.2397 | 15.9948 | 14.9402 | -0.1104 | 105.0672 | 234.4782 | 180.6237 | 232.387 | 160.967
2012-10-30 18:01 | 28.46 | 249.3593 | 195.5361 | 247.2289 | 176.0668 | -1.5997 | 1.2643 | 31.9659 | 40.4153 | 25.2829 | 1.994 | 31.8943 | 63.5984 | 15.9948 | 14.875 | -0.1103 | 104.0138 | 234.4843 | 180.6611 | 232.3539 | 161.1918
2012-10-30 18:01 | 28.4675 | 249.3981 | 195.5475 | 247.2285 | 175.9963 | -1.6055 | 1.2645 | 31.9703 | 40.4264 | 25.283 | 1.9941 | 31.896 | 63.605 | 15.9948 | 14.8825 | -0.1107 | 104.0314 | 234.5156 | 180.665 | 232.346 | 161.1138
2012-10-30 18:01 | 28.4615 | 249.3765 | 195.5101 | 247.2029 | 175.9083 | -1.5991 | 1.2639 | 31.9547 | 40.3878 | 25.2825 | 1.9932 | 31.8803 | 63.5424 | 15.9949 | 14.9009 | -0.1102 | 103.9302 | 234.4756 | 180.6092 | 232.302 | 161.0074
2012-10-30 18:01 | 28.4712 | 249.3313 | 195.4765 | 247.3419 | 176.0156 | -1.5991 | 1.2636 | 31.9483 | 40.3714 | 25.2827 | 1.9928 | 31.874 | 63.517 | 15.9949 | 14.8862 | -0.1102 | 103.8884 | 234.4451 | 180.5903 | 232.4557 | 161.1294
2012-10-30 18:01 | 28.4735 | 249.3376 | 195.5142 | 247.258 | 175.8812 | -1.6037 | 1.2635 | 31.9437 | 40.3595 | 25.2828 | 1.9924 | 31.8694 | 63.498 | 15.9951 | 14.9211 | -0.1106 | 103.8575 | 234.4165 | 180.5931 | 232.3369 | 160.9601
2012-10-30 18:01 | 28.484 | 249.3324 | 195.4854 | 247.2254 | 175.9815 | -1.5997 | 1.2634 | 31.9427 | 40.3578 | 25.2823 | 1.9924 | 31.8684 | 63.4948 | 15.995 | 14.8908 | -0.1103 | 103.8526 | 234.4416 | 180.5946 | 232.3346 | 161.0907
2012-10-30 18:01 | 28.4802 | 249.2777 | 195.4777 | 247.1707 | 176.0597 | -1.6003 | 1.2634 | 31.9426 | 40.3572 | 25.2824 | 1.9924 | 31.8681 | 63.4934 | 15.995 | 14.9237 | -0.1103 | 103.8505 | 234.354 | 180.554 | 232.247 | 161.136
2012-10-30 18:01 | 28.4765 | 249.2662 | 195.4779 | 247.2454 | 175.9974 | -1.6043 | 1.2636 | 31.9463 | 40.3666 | 25.2824 | 1.9926 | 31.872 | 63.5082 | 15.9952 | 14.9037 | -0.1106 | 103.8748 | 234.3625 | 180.5742 | 232.3417 | 161.0937
2012-10-30 18:01 | 28.4855 | 249.2634 | 195.483 | 247.3092 | 175.9869 | -1.5997 | 1.2637 | 31.9497 | 40.3746 | 25.2829 | 1.9928 | 31.8755 | 63.5219 | 15.9952 | 14.9453 | -0.1103 | 103.8965 | 234.3181 | 180.5377 | 232.3639 | 161.0416
2012-10-30 18:01 | 28.4817 | 249.2283 | 195.432 | 247.2035 | 175.6357 | -1.6014 | 1.2648 | 31.9781 | 40.4469 | 25.2825 | 1.9947 | 31.9065 | 63.6455 | 15.9953 | 14.9456 | -0.1104 | 104.0924 | 234.2827 | 180.4864 | 232.2579 | 160.6901
2012-10-30 18:01 | 28.481 | 249.2393 | 195.4078 | 247.1402 | 175.393 | -1.602 | 1.2685 | 32.0729 | 40.6851 | 25.2838 | 2.0005 | 31.9986 | 64.0132 | 15.9953 | 14.9245 | -0.1104 | 104.6983 | 234.3148 | 180.4833 | 232.2157 | 160.4685
2012-10-30 18:01 | 28.481 | 249.2198 | 195.3842 | 247.1793 | 175.2096 | -1.6078 | 1.2692 | 32.0911 | 40.731 | 25.2839 | 2.0016 | 32.0165 | 64.0857 | 15.9951 | 14.9367 | -0.1108 | 104.8166 | 234.2831 | 180.4475 | 232.2426 | 160.2729
2012-10-30 18:01 | 28.487 | 249.1827 | 195.343 | 247.2754 | 175.3795 | -1.6003 | 1.2698 | 32.1066 | 40.7698 | 25.2842 | 2.0026 | 32.0318 | 64.1468 | 15.9951 | 14.9183 | -0.1103 | 104.9166 | 234.2644 | 180.4247 | 232.3571 | 160.4612
2012-10-30 18:01 | 28.4885 | 249.2703 | 195.3524 | 247.1359 | 175.7362 | -1.5974 | 1.2703 | 32.1178 | 40.7982 | 25.2842 | 2.0033 | 32.043 | 64.1919 | 15.9951 | 14.8953 | -0.1101 | 104.9901 | 234.375 | 180.4571 | 232.2406 | 160.8409
2012-10-30 18:01 | 28.4907 | 249.249 | 195.3624 | 247.0989 | 175.8711 | -1.6026 | 1.2705 | 32.125 | 40.8162 | 25.2845 | 2.0038 | 32.0502 | 64.2216 | 15.9949 | 14.9464 | -0.1105 | 105.0377 | 234.3026 | 180.416 | 232.1525 | 160.9247
2012-10-30 18:01 | 28.4937 | 249.2168 | 195.3694 | 247.1098 | 175.7141 | -1.6014 | 1.2707 | 32.1297 | 40.8276 | 25.2847 | 2.0041 | 32.0548 | 64.2401 | 15.9949 | 14.8924 | -0.1104 | 105.0678 | 234.3244 | 180.477 | 232.2174 | 160.8217
2012-10-30 18:01 | 28.5004 | 249.2313 | 195.3289 | 247.2262 | 175.7871 | -1.6026 | 1.2708 | 32.1326 | 40.8351 | 25.2847 | 2.0043 | 32.0578 | 64.2521 | 15.9948 | 14.8746 | -0.1105 | 105.0872 | 234.3567 | 180.4543 | 232.3516 | 160.9125
2012-10-30 18:02 | 28.493 | 249.22 | 195.353 | 247.1992 | 175.9944 | -1.5985 | 1.2709 | 32.1343 | 40.8399 | 25.2844 | 2.0044 | 32.0595 | 64.2593 | 15.9948 | 14.8468 | -0.1102 | 105.0991 | 234.3732 | 180.5062 | 232.3524 | 161.1476
2012-10-30 18:02 | 28.4892 | 249.2397 | 195.4081 | 247.0426 | 176.018 | -1.6078 | 1.271 | 32.1357 | 40.8431 | 25.2846 | 2.0045 | 32.0609 | 64.2648 | 15.9948 | 14.8104 | -0.1108 | 105.1078 | 234.4293 | 180.5977 | 232.2322 | 161.2076
2012-10-30 18:02 | 28.4974 | 249.2714 | 195.3927 | 247.1997 | 175.8466 | -1.598 | 1.271 | 32.137 | 40.8467 | 25.2845 | 2.0045 | 32.0622 | 64.2696 | 15.9949 | 14.8227 | -0.1102 | 105.1162 | 234.4487 | 180.57 | 232.377 | 161.0239
2012-10-30 18:02 | 28.4982 | 249.19 | 195.4721 | 247.1769 | 175.7147 | -1.5951 | 1.2711 | 32.138 | 40.8491 | 25.2846 | 2.0046 | 32.0632 | 64.2741 | 15.9947 | 14.7786 | -0.11 | 105.1232 | 234.4114 | 180.6935 | 232.3983 | 160.9361
2012-10-30 18:02 | 28.5019 | 249.2563 | 195.4012 | 247.0945 | 175.7886 | -1.5997 | 1.2711 | 32.1387 | 40.8509 | 25.2845 | 2.0047 | 32.0638 | 64.2769 | 15.9947 | 14.7823 | -0.1103 | 105.1278 | 234.474 | 180.6189 | 232.3122 | 161.0063
2012-10-30 18:02 | 28.4952 | 249.2457 | 195.3316 | 247.1191 | 175.6336 | -1.6043 | 1.2711 | 32.1392 | 40.8523 | 25.2845 | 2.0047 | 32.0643 | 64.2792 | 15.9946 | 14.7675 | -0.1106 | 105.1315 | 234.4782 | 180.5641 | 232.3516 | 160.8661
2012-10-30 18:02 | 28.5132 | 249.252 | 195.381 | 247.1567 | 175.718 | -1.6009 | 1.2711 | 32.1389 | 40.8514 | 25.2846 | 2.0047 | 32.0641 | 64.2787 | 15.9945 | 14.7936 | -0.1104 | 105.1302 | 234.4584 | 180.5874 | 232.3631 | 160.9244
2012-10-30 18:02 | 28.5124 | 249.2668 | 195.3841 | 247.1911 | 175.7367 | -1.6032 | 1.2711 | 32.139 | 40.8519 | 25.2844 | 2.0047 | 32.0643 | 64.2788 | 15.9947 | 14.748 | -0.1105 | 105.1307 | 234.5188 | 180.6361 | 232.4431 | 160.9887
2012-10-30 18:02 | 28.5064 | 249.2452 | 195.3192 | 247.1499 | 175.7541 | -1.5974 | 1.2711 | 32.1391 | 40.8515 | 25.2848 | 2.0047 | 32.0641 | 64.279 | 15.9945 | 14.742 | -0.1101 | 105.1305 | 234.5032 | 180.5772 | 232.4079 | 161.0121
2012-10-30 18:02 | 28.5072 | 249.246 | 195.32 | 247.1115 | 175.6378 | -1.6032 | 1.2711 | 32.1388 | 40.8511 | 25.2846 | 2.0047 | 32.0639 | 64.2778 | 15.9946 | 14.7061 | -0.1105 | 105.1289 | 234.5399 | 180.6139 | 232.4054 | 160.9317
2012-10-30 18:02 | 28.5199 | 249.2391 | 195.2816 | 247.1673 | 175.4944 | -1.6009 | 1.2711 | 32.1386 | 40.8504 | 25.2847 | 2.0047 | 32.0638 | 64.2772 | 15.9946 | 14.7229 | -0.1104 | 105.1276 | 234.5162 | 180.5587 | 232.4444 | 160.7715
2012-10-30 18:02 | 28.5229 | 249.2029 | 195.2925 | 247.1703 | 175.6808 | -1.5997 | 1.2711 | 32.1387 | 40.8508 | 25.2847 | 2.0047 | 32.0639 | 64.278 | 15.9945 | 14.7503 | -0.1103 | 105.1288 | 234.4526 | 180.5422 | 232.42 | 160.9305
2012-10-30 18:02 | 28.5177 | 249.2839 | 195.2755 | 247.0672 | 175.8161 | -1.5997 | 1.2711 | 32.1383 | 40.8498 | 25.2847 | 2.0047 | 32.0635 | 64.2771 | 15.9943 | 14.6677 | -0.1103 | 105.1269 | 234.6162 | 180.6078 | 232.3995 | 161.1484
2012-10-30 18:02 | 28.5222 | 249.2766 | 195.2761 | 247.1069 | 175.7894 | -1.6003 | 1.2711 | 32.1389 | 40.8512 | 25.2847 | 2.0047 | 32.064 | 64.2789 | 15.9944 | 14.6722 | -0.1103 | 105.1301 | 234.6044 | 180.6039 | 232.4347 | 161.1172
2012-10-30 18:02 | 28.5311 | 249.2503 | 195.2692 | 247.1707 | 175.7476 | -1.602 | 1.2711 | 32.1392 | 40.8522 | 25.2845 | 2.0047 | 32.0643 | 64.28 | 15.9944 | 14.6566 | -0.1104 | 105.1321 | 234.5937 | 180.6126 | 232.5141 | 161.091
2012-10-30 18:02 | 28.5371 | 249.2485 | 195.2792 | 247.2472 | 175.8433 | -1.5997 | 1.2711 | 32.1394 | 40.8524 | 25.2846 | 2.0048 | 32.0645 | 64.2815 | 15.9943 | 14.6626 | -0.1103 | 105.1339 | 234.5859 | 180.6166 | 232.5846 | 161.1807
2012-10-30 18:02 | 28.5304 | 249.2496 | 195.3196 | 247.1347 | 176.0045 | -1.5991 | 1.2711 | 32.1394 | 40.8528 | 25.2844 | 2.0047 | 32.0645 | 64.2809 | 15.9944 | 14.6355 | -0.1102 | 105.1337 | 234.6141 | 180.6841 | 232.4992 | 161.369
2012-10-30 18:02 | 28.5431 | 249.2897 | 195.352 | 247.1239 | 175.7986 | -1.6049 | 1.2711 | 32.1394 | 40.8524 | 25.2847 | 2.0048 | 32.0644 | 64.2814 | 15.9942 | 14.689 | -0.1106 | 105.1338 | 234.6007 | 180.663 | 232.4349 | 161.1096
2012-10-30 18:02 | 28.5424 | 249.2773 | 195.2923 | 247.1545 | 175.8096 | -1.6009 | 1.2711 | 32.1398 | 40.8535 | 25.2846 | 2.0048 | 32.0649 | 64.2833 | 15.9941 | 14.6638 | -0.1104 | 105.1369 | 234.6135 | 180.6285 | 232.4907 | 161.1458
2012-10-30 18:02 | 28.5506 | 249.2894 | 195.2927 | 247.2568 | 175.7475 | -1.602 | 1.2711 | 32.1397 | 40.8532 | 25.2847 | 2.0048 | 32.0649 | 64.283 | 15.9942 | 14.6149 | -0.1104 | 105.1362 | 234.6745 | 180.6778 | 232.6419 | 161.1326
2012-10-30 18:02 | 28.5521 | 249.287 | 195.3138 | 247.1995 | 175.8349 | -1.6032 | 1.2711 | 32.1394 | 40.8525 | 25.2847 | 2.0048 | 32.0646 | 64.2818 | 15.9943 | 14.6572 | -0.1105 | 105.1343 | 234.6298 | 180.6566 | 232.5423 | 161.1777
2012-10-30 18:02 | 28.5499 | 249.3004 | 195.3156 | 247.1111 | 175.7859 | -1.5968 | 1.2711 | 32.1389 | 40.8517 | 25.2844 | 2.0047 | 32.0641 | 64.2799 | 15.9942 | 14.6264 | -0.1101 | 105.1316 | 234.674 | 180.6892 | 232.4847 | 161.1595
2012-10-30 18:02 | 28.5499 | 249.2769 | 195.2723 | 247.2443 | 175.782 | -1.5985 | 1.2711 | 32.1386 | 40.8507 | 25.2846 | 2.0047 | 32.0637 | 64.2785 | 15.9942 | 14.6224 | -0.1102 | 105.1291 | 234.6545 | 180.6499 | 232.6219 | 161.1596
2012-10-30 18:02 | 28.5491 | 249.3035 | 195.3423 | 247.1808 | 175.7968 | -1.602 | 1.2711 | 32.1384 | 40.8504 | 25.2844 | 2.0047 | 32.0636 | 64.2784 | 15.9941 | 14.6134 | -0.1104 | 105.1288 | 234.6901 | 180.7289 | 232.5674 | 161.1834
2012-10-30 18:02 | 28.5506 | 249.305 | 195.332 | 247.1823 | 175.689 | -1.6032 | 1.2711 | 32.1381 | 40.8492 | 25.2847 | 2.0047 | 32.0632 | 64.2769 | 15.9941 | 14.6516 | -0.1105 | 105.1261 | 234.6534 | 180.6804 | 232.5307 | 161.0374
2012-10-30 18:03 | 28.5611 | 249.3116 | 195.3032 | 247.1262 | 175.8088 | -1.5957 | 1.271 | 32.1374 | 40.847 | 25.2849 | 2.0046 | 32.0626 | 64.274 | 15.9942 | 14.654 | -0.11 | 105.121 | 234.6576 | 180.6492 | 232.4722 | 161.1548
2012-10-30 18:03 | 28.5604 | 249.2679 | 195.3182 | 247.1921 | 175.7808 | -1.6032 | 1.271 | 32.1371 | 40.8461 | 25.2849 | 2.0046 | 32.0622 | 64.2727 | 15.9941 | 14.6125 | -0.1105 | 105.1188 | 234.6554 | 180.7057 | 232.5796 | 161.1683
2012-10-30 18:03 | 28.5656 | 249.3279 | 195.3155 | 247.2013 | 175.6845 | -1.5974 | 1.271 | 32.1367 | 40.8457 | 25.2845 | 2.0046 | 32.0618 | 64.2715 | 15.994 | 14.6095 | -0.1101 | 105.1172 | 234.7184 | 180.706 | 232.5918 | 161.075
2012-10-30 18:03 | 28.5716 | 249.2908 | 195.3176 | 247.2817 | 175.7217 | -1.5945 | 1.271 | 32.1362 | 40.8441 | 25.2848 | 2.0046 | 32.0613 | 64.2695 | 15.994 | 14.6074 | -0.1099 | 105.1136 | 234.6834 | 180.7102 | 232.6743 | 161.1143
2012-10-30 18:03 | 28.5783 | 249.3484 | 195.3361 | 247.1865 | 175.826 | -1.6026 | 1.271 | 32.136 | 40.8438 | 25.2847 | 2.0046 | 32.0612 | 64.2697 | 15.9939 | 14.6304 | -0.1105 | 105.1135 | 234.718 | 180.7057 | 232.5561 | 161.1956
2012-10-30 18:03 | 28.5671 | 249.3137 | 195.2856 | 247.191 | 175.7758 | -1.5997 | 1.271 | 32.1358 | 40.8437 | 25.2845 | 2.0046 | 32.0611 | 64.2686 | 15.994 | 14.6233 | -0.1103 | 105.1123 | 234.6904 | 180.6623 | 232.5677 | 161.1525
2012-10-30 18:03 | 28.5731 | 249.3549 | 195.2916 | 247.291 | 175.6178 | -1.6043 | 1.271 | 32.1357 | 40.8431 | 25.2846 | 2.0046 | 32.0609 | 64.2684 | 15.9939 | 14.617 | -0.1106 | 105.1115 | 234.7379 | 180.6746 | 232.674 | 161.0008
2012-10-30 18:03 | 28.5701 | 249.3598 | 195.2925 | 247.2136 | 175.6382 | -1.5974 | 1.2709 | 32.1354 | 40.8423 | 25.2847 | 2.0045 | 32.0606 | 64.2668 | 15.994 | 14.5936 | -0.1101 | 105.1091 | 234.7662 | 180.6989 | 232.62 | 161.0446
2012-10-30 18:03 | 28.5716 | 249.3534 | 195.2783 | 247.1916 | 175.6593 | -1.5991 | 1.271 | 32.1356 | 40.8429 | 25.2847 | 2.0046 | 32.0607 | 64.2678 | 15.9938 | 14.5747 | -0.1102 | 105.1107 | 234.7787 | 180.7036 | 232.6169 | 161.0846
2012-10-30 18:03 | 28.5693 | 249.3511 | 195.2249 | 247.1344 | 175.657 | -1.5991 | 1.2709 | 32.135 | 40.8405 | 25.2851 | 2.0045 | 32.0602 | 64.2655 | 15.9939 | 14.6132 | -0.1102 | 105.106 | 234.7379 | 180.6117 | 232.5212 | 161.0438
2012-10-30 18:03 | 28.5716 | 249.3417 | 195.2233 | 247.3169 | 175.6827 | -1.5945 | 1.2709 | 32.1346 | 40.8404 | 25.2845 | 2.0045 | 32.0598 | 64.2644 | 15.9938 | 14.5339 | -0.1099 | 105.1048 | 234.8078 | 180.6894 | 232.783 | 161.1488
2012-10-30 18:03 | 28.5738 | 249.3596 | 195.2923 | 247.233 | 175.7981 | -1.5985 | 1.2709 | 32.1342 | 40.8392 | 25.2848 | 2.0045 | 32.0594 | 64.2623 | 15.9939 | 14.5973 | -0.1102 | 105.1015 | 234.7623 | 180.695 | 232.6357 | 161.2008
2012-10-30 18:03 | 28.5813 | 249.4258 | 195.2841 | 247.1739 | 175.7353 | -1.5985 | 1.2709 | 32.1341 | 40.8388 | 25.2847 | 2.0045 | 32.0592 | 64.2621 | 15.9938 | 14.5804 | -0.1102 | 105.1009 | 234.8454 | 180.7037 | 232.5935 | 161.1549
2012-10-30 18:03 | 28.5634 | 249.3452 | 195.2465 | 247.2382 | 175.6706 | -1.6032 | 1.2709 | 32.1337 | 40.838 | 25.2847 | 2.0045 | 32.059 | 64.2612 | 15.9938 | 14.5339 | -0.1105 | 105.0992 | 234.8113 | 180.7126 | 232.7043 | 161.1367
2012-10-30 18:03 | 28.5693 | 249.359 | 195.2878 | 247.3577 | 175.7194 | -1.6003 | 1.2709 | 32.1332 | 40.8366 | 25.2848 | 2.0044 | 32.0585 | 64.2593 | 15.9937 | 14.5684 | -0.1103 | 105.0959 | 234.7906 | 180.7194 | 232.7893 | 161.151
2012-10-30 18:03 | 28.5708 | 249.4349 | 195.3129 | 247.23 | 175.8224 | -1.6055 | 1.2708 | 32.133 | 40.8361 | 25.2847 | 2.0044 | 32.0582 | 64.2587 | 15.9936 | 14.5495 | -0.1107 | 105.0948 | 234.8854 | 180.7634 | 232.6805 | 161.2729
2012-10-30 18:03 | 28.5738 | 249.4379 | 195.3119 | 247.2173 | 175.7278 | -1.6026 | 1.2708 | 32.1329 | 40.836 | 25.2846 | 2.0044 | 32.0581 | 64.2582 | 15.9937 | 14.5973 | -0.1105 | 105.0942 | 234.8406 | 180.7146 | 232.62 | 161.1305
2012-10-30 18:03 | 28.5723 | 249.4129 | 195.2515 | 247.2941 | 175.3399 | -1.6032 | 1.2708 | 32.1327 | 40.8356 | 25.2845 | 2.0044 | 32.0578 | 64.2572 | 15.9936 | 14.5673 | -0.1105 | 105.0928 | 234.8456 | 180.6842 | 232.7268 | 160.7726
2012-10-30 18:03 | 28.5768 | 249.3665 | 195.2599 | 247.3143 | 175.5161 | -1.598 | 1.2708 | 32.1323 | 40.8342 | 25.2848 | 2.0044 | 32.0574 | 64.256 | 15.9935 | 14.584 | -0.1102 | 105.0902 | 234.7825 | 180.6759 | 232.7303 | 160.9321
2012-10-30 18:03 | 28.5768 | 249.433 | 195.2796 | 247.2203 | 175.8167 | -1.6003 | 1.2708 | 32.132 | 40.8336 | 25.2848 | 2.0044 | 32.0571 | 64.2544 | 15.9936 | 14.5595 | -0.1103 | 105.088 | 234.8735 | 180.7201 | 232.6608 | 161.2572
2012-10-30 18:03 | 28.5783 | 249.3993 | 195.2221 | 247.2296 | 175.4708 | -1.6061 | 1.2708 | 32.132 | 40.8339 | 25.2846 | 2.0044 | 32.0572 | 64.2552 | 15.9935 | 14.561 | -0.1107 | 105.0891 | 234.8383 | 180.6611 | 232.6686 | 160.9098
2012-10-30 18:03 | 28.5753 | 249.4276 | 195.2309 | 247.3285 | 175.4756 | -1.6049 | 1.2666 | 32.0243 | 40.5615 | 25.2839 | 1.9978 | 31.9519 | 63.8335 | 15.9935 | 14.5784 | -0.1106 | 104.3949 | 234.8492 | 180.6525 | 232.7501 | 160.8972
2012-10-30 18:03 | 28.5731 | 249.3941 | 195.2012 | 247.3263 | 175.3017 | -1.6066 | 1.2648 | 31.9777 | 40.4446 | 25.2833 | 1.9948 | 31.9033 | 63.6396 | 15.9935 | 14.5395 | -0.1108 | 104.0842 | 234.8546 | 180.6617 | 232.7868 | 160.7622
2012-10-30 18:03 | 28.5821 | 249.4344 | 195.1316 | 247.2295 | 175.3926 | -1.6009 | 1.2642 | 31.9616 | 40.4046 | 25.2828 | 1.9937 | 31.8873 | 63.5751 | 15.9937 | 14.5322 | -0.1104 | 103.9797 | 234.9022 | 180.5994 | 232.6973 | 160.8604
2012-10-30 18:04 | 28.5731 | 249.3549 | 195.1109 | 247.3733 | 175.4929 | -1.6032 | 1.2638 | 31.9514 | 40.3788 | 25.2829 | 1.9931 | 31.877 | 63.534 | 15.9937 | 14.5273 | -0.1105 | 103.9128 | 234.8276 | 180.5836 | 232.846 | 160.9656
2012-10-30 18:04 | 28.5783 | 249.3249 | 195.1475 | 247.3393 | 175.705 | -1.6014 | 1.2636 | 31.947 | 40.368 | 25.2827 | 1.9928 | 31.8727 | 63.5167 | 15.9938 | 14.5447 | -0.1104 | 103.8847 | 234.7802 | 180.6028 | 232.7946 | 161.1603
2012-10-30 18:04 | 28.5738 | 249.3165 | 195.1784 | 247.1781 | 175.8527 | -1.602 | 1.2635 | 31.9444 | 40.3614 | 25.2827 | 1.9927 | 31.87 | 63.5058 | 15.9938 | 14.5851 | -0.1104 | 103.8673 | 234.7314 | 180.5933 | 232.593 | 161.2676
2012-10-30 18:04 | 28.5738 | 249.3478 | 195.198 | 247.2525 | 175.6771 | -1.6072 | 1.2635 | 31.9451 | 40.3628 | 25.2828 | 1.9927 | 31.8708 | 63.5087 | 15.9938 | 14.5361 | -0.1108 | 103.8716 | 234.8117 | 180.6619 | 232.7164 | 161.141
2012-10-30 18:04 | 28.5664 | 249.2699 | 195.1631 | 247.3392 | 175.5643 | -1.5985 | 1.2636 | 31.9468 | 40.3673 | 25.2827 | 1.9928 | 31.8723 | 63.5144 | 15.9939 | 14.5043 | -0.1102 | 103.8817 | 234.7656 | 180.6588 | 232.8349 | 161.06
2012-10-30 18:04 | 28.5768 | 249.3117 | 195.2167 | 247.2242 | 175.6488 | -1.602 | 1.2636 | 31.9485 | 40.3713 | 25.283 | 1.9929 | 31.8741 | 63.522 | 15.9938 | 14.5595 | -0.1104 | 103.8933 | 234.7522 | 180.6572 | 232.6647 | 161.0893
2012-10-30 18:04 | 28.5776 | 249.2694 | 195.1782 | 247.1545 | 175.7745 | -1.5951 | 1.2637 | 31.9494 | 40.3735 | 25.283 | 1.993 | 31.875 | 63.5255 | 15.9938 | 14.5155 | -0.11 | 103.899 | 234.7539 | 180.6627 | 232.639 | 161.259
2012-10-30 18:04 | 28.5723 | 249.2602 | 195.1179 | 247.2001 | 175.7419 | -1.6003 | 1.2646 | 31.9717 | 40.43 | 25.283 | 1.9944 | 31.8987 | 63.6195 | 15.9939 | 14.5142 | -0.1103 | 104.0495 | 234.746 | 180.6037 | 232.6859 | 161.2277
2012-10-30 18:04 | 28.5723 | 249.2367 | 195.224 | 247.2706 | 175.8122 | -1.5957 | 1.2686 | 32.0757 | 40.6921 | 25.2838 | 2.0008 | 32.0013 | 64.0292 | 15.994 | 14.4939 | -0.11 | 104.7213 | 234.7428 | 180.7301 | 232.7767 | 161.3183
2012-10-30 18:04 | 28.5701 | 249.2893 | 195.2729 | 247.1274 | 175.7592 | -1.5991 | 1.2692 | 32.0908 | 40.7292 | 25.2845 | 2.0018 | 32.0163 | 64.09 | 15.9938 | 14.4549 | -0.1102 | 104.8192 | 234.8344 | 180.818 | 232.6725 | 161.3043
2012-10-30 18:04 | 28.5671 | 249.2354 | 195.2424 | 247.2145 | 175.6665 | -1.6014 | 1.2698 | 32.1064 | 40.7689 | 25.2845 | 2.0028 | 32.0316 | 64.1518 | 15.9937 | 14.4846 | -0.1104 | 104.9207 | 234.7508 | 180.7578 | 232.7299 | 161.1819
2012-10-30 18:04 | 28.5723 | 249.3072 | 195.279 | 247.2824 | 175.7692 | -1.6037 | 1.2702 | 32.1178 | 40.7976 | 25.2847 | 2.0034 | 32.0429 | 64.1962 | 15.9939 | 14.4816 | -0.1106 | 104.9938 | 234.8256 | 180.7974 | 232.8008 | 161.2876
2012-10-30 18:04 | 28.5611 | 249.2372 | 195.315 | 247.1145 | 175.8127 | -1.6032 | 1.2705 | 32.1257 | 40.8171 | 25.285 | 2.004 | 32.0509 | 64.2291 | 15.9937 | 14.4704 | -0.1105 | 105.0462 | 234.7668 | 180.8446 | 232.6441 | 161.3423
2012-10-30 18:04 | 28.5671 | 249.2902 | 195.2542 | 247.0735 | 175.686 | -1.6009 | 1.2707 | 32.1302 | 40.8285 | 25.285 | 2.0042 | 32.0555 | 64.2469 | 15.9938 | 14.4927 | -0.1104 | 105.0754 | 234.7975 | 180.7615 | 232.5808 | 161.1933
2012-10-30 18:04 | 28.5649 | 249.2802 | 195.2402 | 247.2476 | 175.5511 | -1.6032 | 1.2708 | 32.133 | 40.8356 | 25.285 | 2.0044 | 32.058 | 64.2573 | 15.9937 | 14.4701 | -0.1105 | 105.0929 | 234.8101 | 180.7701 | 232.7775 | 161.081
2012-10-30 18:04 | 28.5678 | 249.2948 | 195.3059 | 247.1369 | 175.7882 | -1.5968 | 1.2709 | 32.135 | 40.841 | 25.2848 | 2.0046 | 32.0601 | 64.2662 | 15.9937 | 14.4445 | -0.1101 | 105.1072 | 234.8503 | 180.8614 | 232.6924 | 161.3437
2012-10-30 18:04 | 28.5776 | 249.3164 | 195.2804 | 247.0605 | 175.7667 | -1.602 | 1.271 | 32.1367 | 40.8449 | 25.285 | 2.0046 | 32.0618 | 64.2725 | 15.9937 | 14.4788 | -0.1104 | 105.1174 | 234.8376 | 180.8016 | 232.5817 | 161.2879
2012-10-30 18:04 | 28.5656 | 249.3161 | 195.2487 | 247.2091 | 175.6962 | -1.598 | 1.271 | 32.1374 | 40.8461 | 25.2855 | 2.0047 | 32.0625 | 64.276 | 15.9936 | 14.426 | -0.1102 | 105.1221 | 234.8901 | 180.8227 | 232.7831 | 161.2702
2012-10-30 18:04 | 28.5738 | 249.3596 | 195.3159 | 247.1193 | 175.802 | -1.6026 | 1.271 | 32.1382 | 40.8482 | 25.2854 | 2.0048 | 32.0632 | 64.2791 | 15.9935 | 14.4709 | -0.1105 | 105.1274 | 234.8887 | 180.845 | 232.6484 | 161.3311
2012-10-30 18:04 | 28.5723 | 249.3346 | 195.2672 | 247.0473 | 175.699 | -1.6003 | 1.271 | 32.1383 | 40.8488 | 25.2853 | 2.0048 | 32.0635 | 64.2801 | 15.9936 | 14.4531 | -0.1103 | 105.1289 | 234.8815 | 180.8141 | 232.5942 | 161.2459
2012-10-30 18:04 | 28.5678 | 249.3379 | 195.2627 | 247.176 | 175.7023 | -1.5991 | 1.2711 | 32.1394 | 40.8512 | 25.2854 | 2.0048 | 32.0644 | 64.2837 | 15.9936 | 14.4445 | -0.1102 | 105.1349 | 234.8934 | 180.8182 | 232.7315 | 161.2578
2012-10-30 18:04 | 28.5619 | 249.3398 | 195.2568 | 247.1936 | 175.7706 | -1.6026 | 1.2711 | 32.1397 | 40.8522 | 25.2853 | 2.0048 | 32.0647 | 64.2845 | 15.9936 | 14.4508 | -0.1105 | 105.1366 | 234.889 | 180.806 | 232.7428 | 161.3198
2012-10-30 18:04 | 28.5746 | 249.3995 | 195.2656 | 247.1123 | 175.7637 | -1.6037 | 1.2711 | 32.1394 | 40.8514 | 25.2854 | 2.0049 | 32.0645 | 64.2846 | 15.9935 | 14.4635 | -0.1106 | 105.136 | 234.936 | 180.8021 | 232.6488 | 161.3002
2012-10-30 18:04 | 28.5723 | 249.362 | 195.279 | 247.1649 | 175.6404 | -1.6009 | 1.2711 | 32.14 | 40.8531 | 25.2852 | 2.0049 | 32.0651 | 64.2865 | 15.9935 | 14.4449 | -0.1104 | 105.1396 | 234.9171 | 180.8341 | 232.72 | 161.1955
2012-10-30 18:04 | 28.5664 | 249.3443 | 195.2613 | 247.2491 | 175.6775 | -1.598 | 1.2711 | 32.1402 | 40.8537 | 25.2852 | 2.0049 | 32.0654 | 64.2882 | 15.9934 | 14.4757 | -0.1102 | 105.1419 | 234.8686 | 180.7856 | 232.7734 | 161.2018
2012-10-30 18:04 | 28.5678 | 249.4201 | 195.3177 | 247.1094 | 175.7179 | -1.5928 | 1.2711 | 32.1409 | 40.8552 | 25.2853 | 2.0049 | 32.0659 | 64.2905 | 15.9934 | 14.4853 | -0.1098 | 105.1457 | 234.9348 | 180.8324 | 232.6241 | 161.2326
2012-10-30 18:05 | 28.5626 | 249.3953 | 195.2772 | 247.1591 | 175.5995 | -1.5957 | 1.2711 | 32.141 | 40.8545 | 25.2859 | 2.005 | 32.066 | 64.2909 | 15.9934 | 14.4597 | -0.11 | 105.1454 | 234.9356 | 180.8175 | 232.6994 | 161.1398
2012-10-30 18:05 | 28.5701 | 249.3911 | 195.3082 | 247.2841 | 175.7046 | -1.6032 | 1.2711 | 32.141 | 40.8552 | 25.2854 | 2.005 | 32.0661 | 64.291 | 15.9934 | 14.4753 | -0.1105 | 105.1462 | 234.9158 | 180.8329 | 232.8088 | 161.2293
2012-10-30 18:05 | 28.5604 | 249.4088 | 195.3143 | 247.0707 | 175.5778 | -1.6037 | 1.2711 | 32.141 | 40.8559 | 25.2852 | 2.005 | 32.0662 | 64.2919 | 15.9933 | 14.4697 | -0.1106 | 105.1478 | 234.9391 | 180.8446 | 232.601 | 161.1081
2012-10-30 18:05 | 28.5596 | 249.3767 | 195.2349 | 247.2188 | 175.6004 | -1.6061 | 1.2711 | 32.141 | 40.855 | 25.2855 | 2.005 | 32.066 | 64.2915 | 15.9933 | 14.4852 | -0.1107 | 105.1465 | 234.8915 | 180.7497 | 232.7336 | 161.1152
2012-10-30 18:05 | 28.5581 | 249.4065 | 195.2609 | 247.2564 | 175.595 | -1.6026 | 1.2712 | 32.1413 | 40.8566 | 25.2851 | 2.005 | 32.0663 | 64.2923 | 15.9933 | 14.4511 | -0.1105 | 105.1489 | 234.9554 | 180.8098 | 232.8053 | 161.1439
2012-10-30 18:05 | 28.5506 | 249.4264 | 195.2652 | 247.1432 | 175.607 | -1.602 | 1.2711 | 32.1411 | 40.8557 | 25.2853 | 2.005 | 32.0662 | 64.2927 | 15.9931 | 14.4395 | -0.1104 | 105.1484 | 234.9869 | 180.8257 | 232.7037 | 161.1675
2012-10-30 18:05 | 28.5589 | 249.4034 | 195.1988 | 247.1946 | 175.428 | -1.6014 | 1.2711 | 32.1411 | 40.8557 | 25.2853 | 2.005 | 32.0661 | 64.2918 | 15.9933 | 14.4764 | -0.1104 | 105.1474 | 234.927 | 180.7224 | 232.7182 | 160.9516
2012-10-30 18:05 | 28.5589 | 249.3799 | 195.2303 | 247.2416 | 175.5412 | -1.6003 | 1.2712 | 32.1411 | 40.8563 | 25.285 | 2.005 | 32.0662 | 64.2925 | 15.9932 | 14.4437 | -0.1103 | 105.1487 | 234.9362 | 180.7866 | 232.7979 | 161.0975
2012-10-30 18:05 | 28.5544 | 249.4263 | 195.2218 | 247.0686 | 175.5523 | -1.5997 | 1.2711 | 32.1406 | 40.8549 | 25.285 | 2.005 | 32.0658 | 64.2908 | 15.9932 | 14.4229 | -0.1103 | 105.1457 | 235.0034 | 180.7989 | 232.6457 | 161.1294
2012-10-30 18:05 | 28.5536 | 249.4137 | 195.1621 | 247.1893 | 175.4968 | -1.6014 | 1.2711 | 32.1402 | 40.8539 | 25.2851 | 2.0049 | 32.0653 | 64.289 | 15.9932 | 14.4262 | -0.1104 | 105.1428 | 234.9875 | 180.7359 | 232.7631 | 161.0706
2012-10-30 18:05 | 28.5559 | 249.3847 | 195.2273 | 247.2973 | 175.4523 | -1.6003 | 1.2711 | 32.1401 | 40.8538 | 25.2849 | 2.005 | 32.0651 | 64.289 | 15.993 | 14.4611 | -0.1103 | 105.1428 | 234.9236 | 180.7662 | 232.8362 | 160.9912
2012-10-30 18:05 | 28.5559 | 249.4239 | 195.2233 | 247.1406 | 175.5069 | -1.598 | 1.2711 | 32.1399 | 40.8528 | 25.2853 | 2.0049 | 32.065 | 64.2875 | 15.9932 | 14.506 | -0.1102 | 105.1403 | 234.9179 | 180.7173 | 232.6346 | 161.0009
2012-10-30 18:05 | 28.5506 | 249.399 | 195.1709 | 247.2842 | 175.4938 | -1.598 | 1.271 | 32.1394 | 40.8507 | 25.2857 | 2.0049 | 32.0644 | 64.2856 | 15.9931 | 14.4681 | -0.1102 | 105.1362 | 234.9309 | 180.7028 | 232.8161 | 161.0257
2012-10-30 18:05 | 28.5581 | 249.4378 | 195.2648 | 247.182 | 175.7082 | -1.6014 | 1.271 | 32.1385 | 40.8495 | 25.2851 | 2.0048 | 32.0638 | 64.2829 | 15.9931 | 14.4756 | -0.1104 | 105.1324 | 234.9622 | 180.7892 | 232.7064 | 161.2326
2012-10-30 18:05 | 28.5574 | 249.4371 | 195.2052 | 247.1931 | 175.6021 | -1.6026 | 1.271 | 32.138 | 40.848 | 25.2852 | 2.0048 | 32.0632 | 64.281 | 15.993 | 14.4463 | -0.1105 | 105.129 | 234.9908 | 180.7589 | 232.7468 | 161.1558
2012-10-30 18:05 | 28.5529 | 249.4444 | 195.2203 | 247.3178 | 175.6327 | -1.6032 | 1.271 | 32.1383 | 40.8485 | 25.2854 | 2.0048 | 32.0633 | 64.2816 | 15.993 | 14.4663 | -0.1105 | 105.1301 | 234.9781 | 180.754 | 232.8515 | 161.1664
2012-10-30 18:05 | 28.5641 | 249.4595 | 195.2787 | 247.2076 | 175.6596 | -1.6014 | 1.271 | 32.138 | 40.8473 | 25.2856 | 2.0048 | 32.063 | 64.2805 | 15.993 | 14.4612 | -0.1104 | 105.1278 | 234.9983 | 180.8175 | 232.7464 | 161.1984
2012-10-30 18:05 | 28.5544 | 249.4145 | 195.1668 | 247.2331 | 175.3962 | -1.6026 | 1.271 | 32.1377 | 40.8472 | 25.2852 | 2.0048 | 32.0628 | 64.2799 | 15.9929 | 14.4025 | -0.1105 | 105.1271 | 235.012 | 180.7643 | 232.8306 | 160.9937
2012-10-30 18:05 | 28.5484 | 249.4164 | 195.2394 | 247.2428 | 175.519 | -1.5991 | 1.271 | 32.138 | 40.8477 | 25.2854 | 2.0048 | 32.0629 | 64.2806 | 15.9929 | 14.4373 | -0.1102 | 105.1283 | 234.9791 | 180.8021 | 232.8055 | 161.0817
2012-10-30 18:05 | 28.5536 | 249.4411 | 195.2525 | 247.2049 | 175.4188 | -1.602 | 1.271 | 32.1381 | 40.8485 | 25.2851 | 2.0048 | 32.0632 | 64.2813 | 15.993 | 14.4425 | -0.1104 | 105.1298 | 234.9986 | 180.81 | 232.7624 | 160.9763
2012-10-30 18:05 | 28.5469 | 249.3875 | 195.1554 | 247.2766 | 175.4511 | -1.6026 | 1.271 | 32.1379 | 40.8479 | 25.2851 | 2.0048 | 32.063 | 64.2809 | 15.9929 | 14.4236 | -0.1105 | 105.1288 | 234.9639 | 180.7318 | 232.853 | 161.0275
2012-10-30 18:05 | 28.5521 | 249.4122 | 195.1802 | 247.3092 | 175.5266 | -1.5997 | 1.271 | 32.1373 | 40.8462 | 25.2853 | 2.0048 | 32.0624 | 64.2787 | 15.9928 | 14.4573 | -0.1103 | 105.1249 | 234.9549 | 180.7229 | 232.8519 | 161.0693
2012-10-30 18:05 | 28.5491 | 249.4484 | 195.2047 | 247.1652 | 175.5236 | -1.6055 | 1.271 | 32.1372 | 40.8466 | 25.2848 | 2.0048 | 32.0623 | 64.2781 | 15.9928 | 14.4625 | -0.1107 | 105.1247 | 234.9859 | 180.7422 | 232.7027 | 161.0611
2012-10-30 18:05 | 28.5439 | 249.4197 | 195.1288 | 247.2814 | 175.3818 | -1.602 | 1.271 | 32.1371 | 40.8457 | 25.2853 | 2.0048 | 32.0622 | 64.2778 | 15.9928 | 14.4491 | -0.1104 | 105.1235 | 234.9706 | 180.6797 | 232.8323 | 160.9327
2012-10-30 18:05 | 28.5401 | 249.4707 | 195.18 | 247.2619 | 175.5614 | -1.5997 | 1.271 | 32.1365 | 40.8439 | 25.2854 | 2.0047 | 32.0616 | 64.2754 | 15.9929 | 14.4331 | -0.1103 | 105.1193 | 235.0376 | 180.7469 | 232.8288 | 161.1283
2012-10-30 18:05 | 28.5499 | 249.4179 | 195.178 | 247.1699 | 175.5907 | -1.6026 | 1.2709 | 32.1362 | 40.8433 | 25.2853 | 2.0047 | 32.0613 | 64.2749 | 15.9927 | 14.4674 | -0.1105 | 105.1182 | 234.9505 | 180.7106 | 232.7025 | 161.1233
2012-10-30 18:06 | 28.5386 | 249.4379 | 195.0999 | 247.2722 | 175.3335 | -1.5991 | 1.2709 | 32.136 | 40.8431 | 25.2852 | 2.0047 | 32.0612 | 64.2741 | 15.9928 | 14.4153 | -0.1102 | 105.1172 | 235.0226 | 180.6846 | 232.8569 | 160.9182
2012-10-30 18:06 | 28.5521 | 249.4553 | 195.1409 | 247.3405 | 175.4407 | -1.6037 | 1.271 | 32.1361 | 40.8434 | 25.2851 | 2.0048 | 32.0613 | 64.2749 | 15.9927 | 14.4492 | -0.1106 | 105.1183 | 235.0061 | 180.6917 | 232.8913 | 160.9915
2012-10-30 18:06 | 28.5469 | 249.454 | 195.1868 | 247.2217 | 175.5448 | -1.6009 | 1.2709 | 32.1358 | 40.8426 | 25.2851 | 2.0047 | 32.061 | 64.2735 | 15.9927 | 14.4195 | -0.1104 | 105.1161 | 235.0345 | 180.7673 | 232.8022 | 161.1253
2012-10-30 18:06 | 28.5499 | 249.4374 | 195.1505 | 247.2286 | 175.5166 | -1.6026 | 1.2709 | 32.1356 | 40.8419 | 25.2853 | 2.0047 | 32.0607 | 64.2725 | 15.9927 | 14.3899 | -0.1105 | 105.1144 | 235.0475 | 180.7606 | 232.8387 | 161.1267
2012-10-30 18:06 | 28.5476 | 249.4195 | 195.1639 | 247.3635 | 175.5065 | -1.6055 | 1.2709 | 32.1352 | 40.8411 | 25.2851 | 2.0047 | 32.0602 | 64.2706 | 15.9926 | 14.4202 | -0.1107 | 105.1117 | 234.9993 | 180.7437 | 232.9433 | 161.0863
2012-10-30 18:06 | 28.5506 | 249.489 | 195.1512 | 247.1863 | 175.4548 | -1.5974 | 1.2709 | 32.135 | 40.8403 | 25.2853 | 2.0047 | 32.0602 | 64.2708 | 15.9925 | 14.415 | -0.1101 | 105.1111 | 235.074 | 180.7362 | 232.7713 | 161.0398
2012-10-30 18:06 | 28.5506 | 249.4342 | 195.1355 | 247.2724 | 175.3494 | -1.6009 | 1.2709 | 32.1347 | 40.8391 | 25.2855 | 2.0047 | 32.0597 | 64.2692 | 15.9925 | 14.4681 | -0.1104 | 105.1084 | 234.9661 | 180.6674 | 232.8043 | 160.8813
2012-10-30 18:06 | 28.5551 | 249.4387 | 195.1557 | 247.3475 | 175.3735 | -1.5997 | 1.2709 | 32.1342 | 40.8391 | 25.2848 | 2.0047 | 32.0593 | 64.2678 | 15.9924 | 14.4359 | -0.1103 | 105.1069 | 235.0028 | 180.7198 | 232.9116 | 160.9376
2012-10-30 18:06 | 28.5589 | 249.513 | 195.1477 | 247.2024 | 175.467 | -1.6049 | 1.2709 | 32.134 | 40.8381 | 25.2851 | 2.0046 | 32.0592 | 64.2674 | 15.9924 | 14.4274 | -0.1106 | 105.1055 | 235.0856 | 180.7203 | 232.775 | 161.0396
2012-10-30 18:06 | 28.5491 | 249.4641 | 195.0947 | 247.2396 | 175.3323 | -1.6084 | 1.2708 | 32.1336 | 40.8367 | 25.2853 | 2.0046 | 32.0587 | 64.2657 | 15.9924 | 14.4013 | -0.1109 | 105.1024 | 235.0628 | 180.6934 | 232.8383 | 160.931
2012-10-30 18:06 | 28.5536 | 249.4372 | 195.0835 | 247.3146 | 175.3173 | -1.5945 | 1.2708 | 32.1332 | 40.8359 | 25.2851 | 2.0046 | 32.0582 | 64.2633 | 15.9925 | 14.4262 | -0.1099 | 105.0992 | 235.011 | 180.6573 | 232.8884 | 160.8911
2012-10-30 18:06 | 28.5566 | 249.4833 | 195.114 | 247.3215 | 175.3008 | -1.5997 | 1.2708 | 32.1325 | 40.8342 | 25.2851 | 2.0045 | 32.0576 | 64.2609 | 15.9925 | 14.4455 | -0.1103 | 105.0951 | 235.0378 | 180.6685 | 232.876 | 160.8553
2012-10-30 18:06 | 28.5589 | 249.5247 | 195.132 | 247.2376 | 175.3694 | -1.6061 | 1.2708 | 32.1319 | 40.8327 | 25.2851 | 2.0045 | 32.0571 | 64.259 | 15.9924 | 14.4519 | -0.1107 | 105.0917 | 235.0728 | 180.6801 | 232.7857 | 160.9175
2012-10-30 18:06 | 28.5469 | 249.4423 | 195.0768 | 247.2727 | 175.5409 | -1.602 | 1.2643 | 31.9662 | 40.415 | 25.2836 | 1.9942 | 31.8924 | 63.5997 | 15.9926 | 14.4154 | -0.1104 | 104.0147 | 235.0269 | 180.6614 | 232.8573 | 161.1255
2012-10-30 18:06 | 28.5364 | 249.3535 | 195.1252 | 247.3914 | 175.5889 | -1.6026 | 1.2634 | 31.9438 | 40.3587 | 25.2835 | 1.9927 | 31.8693 | 63.5075 | 15.9927 | 14.3723 | -0.1105 | 103.8662 | 234.9812 | 180.7529 | 233.0191 | 161.2166
2012-10-30 18:06 | 28.5289 | 249.4243 | 195.1099 | 247.2076 | 175.679 | -1.5997 | 1.2634 | 31.9436 | 40.3584 | 25.2833 | 1.9927 | 31.8691 | 63.5065 | 15.9927 | 14.4464 | -0.1103 | 103.8648 | 234.9779 | 180.6635 | 232.7612 | 161.2326
2012-10-30 18:06 | 28.5222 | 249.4019 | 195.0953 | 247.2205 | 175.7035 | -1.602 | 1.2635 | 31.9454 | 40.3632 | 25.2832 | 1.9928 | 31.871 | 63.5137 | 15.9928 | 14.3866 | -0.1104 | 103.8769 | 235.0153 | 180.7087 | 232.8339 | 161.3169
2012-10-30 18:06 | 28.5319 | 249.2942 | 195.1757 | 247.336 | 175.7835 | -1.5985 | 1.2636 | 31.9491 | 40.3719 | 25.2836 | 1.9931 | 31.8748 | 63.5285 | 15.9928 | 14.4249 | -0.1102 | 103.9004 | 234.8693 | 180.7508 | 232.9111 | 161.3586
2012-10-30 18:06 | 28.5207 | 249.3339 | 195.2471 | 247.2308 | 175.9558 | -1.6014 | 1.2638 | 31.9541 | 40.3846 | 25.2836 | 1.9933 | 31.8782 | 63.5421 | 15.9928 | 14.4055 | -0.1104 | 103.9267 | 234.9284 | 180.8416 | 232.8253 | 161.5503
2012-10-30 18:06 | 28.5124 | 249.3569 | 195.2388 | 247.1089 | 175.5688 | -1.6009 | 1.2684 | 32.0694 | 40.6756 | 25.2841 | 2.0005 | 31.9944 | 64.0058 | 15.993 | 14.4299 | -0.1104 | 104.6814 | 234.927 | 180.8089 | 232.679 | 161.1389
2012-10-30 18:06 | 28.5237 | 249.3251 | 195.1951 | 247.3356 | 175.4513 | -1.6061 | 1.2689 | 32.0849 | 40.714 | 25.2847 | 2.0015 | 32.0097 | 64.0677 | 15.9928 | 14.4167 | -0.1107 | 104.7817 | 234.9084 | 180.7784 | 232.9189 | 161.0346
2012-10-30 18:07 | 28.5094 | 249.3695 | 195.2161 | 247.1333 | 175.7337 | -1.6009 | 1.2695 | 32.1004 | 40.7529 | 25.2849 | 2.0025 | 32.0253 | 64.13 | 15.9928 | 14.3779 | -0.1104 | 104.8829 | 234.9916 | 180.8382 | 232.7554 | 161.3558
2012-10-30 18:07 | 28.5049 | 249.3494 | 195.2116 | 247.1249 | 175.6355 | -1.5991 | 1.2701 | 32.1132 | 40.7854 | 25.285 | 2.0033 | 32.0381 | 64.1821 | 15.9926 | 14.3775 | -0.1102 | 104.9675 | 234.9719 | 180.8341 | 232.7474 | 161.258
2012-10-30 18:07 | 28.5109 | 249.3632 | 195.2176 | 247.2954 | 175.7313 | -1.6049 | 1.2704 | 32.1214 | 40.8057 | 25.2854 | 2.0038 | 32.0465 | 64.2163 | 15.9925 | 14.4243 | -0.1106 | 105.022 | 234.9389 | 180.7933 | 232.8711 | 161.307
2012-10-30 18:07 | 28.5139 | 249.4249 | 195.311 | 247.1104 | 175.8319 | -1.6026 | 1.2706 | 32.127 | 40.8201 | 25.2852 | 2.0042 | 32.0521 | 64.2386 | 15.9926 | 14.4191 | -0.1105 | 105.0587 | 235.0058 | 180.8919 | 232.6913 | 161.4128
2012-10-30 18:07 | 28.5057 | 249.3698 | 195.2635 | 247.1766 | 175.687 | -1.5991 | 1.2707 | 32.131 | 40.83 | 25.2853 | 2.0044 | 32.056 | 64.2543 | 15.9925 | 14.415 | -0.1102 | 105.0843 | 234.9548 | 180.8485 | 232.7616 | 161.272
2012-10-30 18:07 | 28.4974 | 249.4084 | 195.2945 | 247.1957 | 175.8973 | -1.598 | 1.2709 | 32.1354 | 40.8409 | 25.2855 | 2.0047 | 32.0604 | 64.2721 | 15.9925 | 14.3945 | -0.1102 | 105.1129 | 235.0139 | 180.9 | 232.8012 | 161.5028
2012-10-30 18:07 | 28.4982 | 249.3975 | 195.4014 | 247.1495 | 176.0816 | -1.6003 | 1.271 | 32.1381 | 40.8477 | 25.2855 | 2.0049 | 32.0631 | 64.2834 | 15.9923 | 14.4116 | -0.1103 | 105.1312 | 234.9859 | 180.9898 | 232.7379 | 161.67
2012-10-30 18:07 | 28.4982 | 249.4601 | 195.3582 | 247.0947 | 175.9996 | -1.5997 | 1.271 | 32.1389 | 40.8501 | 25.2854 | 2.0049 | 32.0638 | 64.286 | 15.9925 | 14.4442 | -0.1103 | 105.1361 | 235.0159 | 180.914 | 232.6505 | 161.5554
2012-10-30 18:07 | 28.4982 | 249.4132 | 195.3778 | 247.2592 | 175.9411 | -1.6003 | 1.2711 | 32.1396 | 40.8512 | 25.2857 | 2.005 | 32.0645 | 64.289 | 15.9924 | 14.3994 | -0.1103 | 105.1402 | 235.0138 | 180.9784 | 232.8598 | 161.5417
2012-10-30 18:07 | 28.4952 | 249.4493 | 195.4141 | 247.1426 | 176.0318 | -1.6049 | 1.2711 | 32.1401 | 40.8528 | 25.2856 | 2.005 | 32.0651 | 64.2916 | 15.9923 | 14.4208 | -0.1106 | 105.1444 | 235.0285 | 180.9933 | 232.7218 | 161.611
2012-10-30 18:07 | 28.4907 | 249.4017 | 195.3664 | 247.1499 | 175.8165 | -1.6014 | 1.2711 | 32.1405 | 40.8543 | 25.2853 | 2.0051 | 32.0655 | 64.2934 | 15.9923 | 14.3755 | -0.1104 | 105.1477 | 235.0262 | 180.9909 | 232.7744 | 161.441
2012-10-30 18:07 | 28.4922 | 249.4502 | 195.4386 | 247.2415 | 176.021 | -1.5985 | 1.2711 | 32.1406 | 40.8539 | 25.2856 | 2.0051 | 32.0656 | 64.2935 | 15.9923 | 14.3689 | -0.1102 | 105.1474 | 235.0813 | 181.0697 | 232.8726 | 161.6521
2012-10-30 18:07 | 28.4855 | 249.4592 | 195.4594 | 247.0506 | 176.0299 | -1.6061 | 1.2711 | 32.1409 | 40.8549 | 25.2855 | 2.0051 | 32.0658 | 64.2948 | 15.9922 | 14.3173 | -0.1107 | 105.1497 | 235.1419 | 181.1421 | 232.7333 | 161.7126
2012-10-30 18:07 | 28.4944 | 249.4602 | 195.4408 | 247.1771 | 175.9529 | -1.6043 | 1.2711 | 32.1412 | 40.8552 | 25.2858 | 2.0051 | 32.0662 | 64.2963 | 15.9922 | 14.3466 | -0.1106 | 105.1516 | 235.1136 | 181.0942 | 232.8305 | 161.6063
2012-10-30 18:07 | 28.4937 | 249.4909 | 195.4165 | 247.3135 | 175.6633 | -1.5997 | 1.2711 | 32.1412 | 40.8549 | 25.2859 | 2.0051 | 32.0662 | 64.2969 | 15.9921 | 14.3785 | -0.1103 | 105.1518 | 235.1124 | 181.038 | 232.935 | 161.2848
2012-10-30 18:07 | 28.4982 | 249.4758 | 195.4485 | 247.0398 | 175.8825 | -1.5951 | 1.2711 | 32.1414 | 40.8565 | 25.2853 | 2.0051 | 32.0664 | 64.2974 | 15.9921 | 14.3586 | -0.11 | 105.154 | 235.1172 | 181.0899 | 232.6812 | 161.5239
2012-10-30 18:07 | 28.49 | 249.4715 | 195.401 | 247.1923 | 175.7572 | -1.602 | 1.2711 | 32.1416 | 40.8561 | 25.2858 | 2.0052 | 32.0666 | 64.2986 | 15.992 | 14.3504 | -0.1104 | 105.1547 | 235.1211 | 181.0506 | 232.8419 | 161.4068
2012-10-30 18:07 | 28.4922 | 249.4463 | 195.415 | 247.2415 | 175.7204 | -1.598 | 1.2711 | 32.1415 | 40.8561 | 25.2857 | 2.0052 | 32.0666 | 64.2985 | 15.9921 | 14.3607 | -0.1102 | 105.1546 | 235.0856 | 181.0543 | 232.8808 | 161.3597
2012-10-30 18:07 | 28.4967 | 249.5213 | 195.4313 | 247.2342 | 175.7717 | -1.6037 | 1.2711 | 32.1414 | 40.8564 | 25.2855 | 2.0051 | 32.0665 | 64.2982 | 15.9921 | 14.3734 | -0.1106 | 105.1545 | 235.1479 | 181.0579 | 232.8608 | 161.3983
2012-10-30 18:07 | 28.5004 | 249.5015 | 195.435 | 247.1752 | 175.6505 | -1.6026 | 1.2712 | 32.1417 | 40.8573 | 25.2853 | 2.0052 | 32.0667 | 64.2991 | 15.992 | 14.3689 | -0.1105 | 105.1564 | 235.1326 | 181.0661 | 232.8063 | 161.2816
2012-10-30 18:07 | 28.4959 | 249.4931 | 195.3598 | 247.3157 | 175.6304 | -1.6003 | 1.2711 | 32.1416 | 40.8562 | 25.2858 | 2.0051 | 32.0666 | 64.2984 | 15.9921 | 14.3685 | -0.1103 | 105.1546 | 235.1246 | 180.9913 | 232.9472 | 161.2619
2012-10-30 18:07 | 28.4967 | 249.5213 | 195.4077 | 247.2577 | 175.7171 | -1.6026 | 1.2711 | 32.1414 | 40.8562 | 25.2855 | 2.0052 | 32.0664 | 64.2981 | 15.992 | 14.3489 | -0.1105 | 105.1544 | 235.1724 | 181.0588 | 232.9088 | 161.3682
2012-10-30 18:07 | 28.4952 | 249.5041 | 195.3787 | 247.1465 | 175.5985 | -1.6009 | 1.2711 | 32.1413 | 40.8556 | 25.2858 | 2.0051 | 32.0664 | 64.298 | 15.992 | 14.4086 | -0.1104 | 105.1536 | 235.0955 | 180.9701 | 232.7379 | 161.1899
2012-10-30 18:07 | 28.5042 | 249.5405 | 195.3367 | 247.2887 | 175.5919 | -1.6032 | 1.2711 | 32.1413 | 40.8555 | 25.2857 | 2.0051 | 32.0662 | 64.2971 | 15.9921 | 14.3809 | -0.1105 | 105.1526 | 235.1596 | 180.9558 | 232.9078 | 161.211
2012-10-30 18:07 | 28.5012 | 249.5258 | 195.4005 | 247.2936 | 175.706 | -1.6003 | 1.2711 | 32.1412 | 40.8554 | 25.2857 | 2.0051 | 32.0662 | 64.2967 | 15.9921 | 14.382 | -0.1103 | 105.1521 | 235.1438 | 181.0185 | 232.9116 | 161.324
2012-10-30 18:07 | 28.5072 | 249.5435 | 195.3711 | 247.1664 | 175.7939 | -1.602 | 1.2711 | 32.1411 | 40.8547 | 25.2859 | 2.0051 | 32.0661 | 64.2969 | 15.992 | 14.3676 | -0.1104 | 105.1515 | 235.1759 | 181.0035 | 232.7988 | 161.4263
2012-10-30 18:08 | 28.4974 | 249.5063 | 195.3259 | 247.2819 | 175.7568 | -1.5991 | 1.2711 | 32.141 | 40.8551 | 25.2856 | 2.0051 | 32.066 | 64.297 | 15.9918 | 14.3863 | -0.1102 | 105.1521 | 235.12 | 180.9396 | 232.8956 | 161.3705
2012-10-30 18:08 | 28.4855 | 249.4866 | 195.3769 | 247.3053 | 175.78 | -1.5916 | 1.2711 | 32.1409 | 40.855 | 25.2855 | 2.0051 | 32.0658 | 64.2968 | 15.9917 | 14.4071 | -0.1097 | 105.1517 | 235.0795 | 180.9698 | 232.8982 | 161.3729
2012-10-30 18:08 | 28.4982 | 249.5463 | 195.3739 | 247.2631 | 175.8435 | -1.5985 | 1.2711 | 32.1408 | 40.8547 | 25.2856 | 2.0051 | 32.0658 | 64.2963 | 15.9918 | 14.379 | -0.1102 | 105.1509 | 235.1673 | 180.9949 | 232.8841 | 161.4645
2012-10-30 18:08 | 28.4922 | 249.5363 | 195.3836 | 247.1749 | 175.7048 | -1.6032 | 1.2711 | 32.1407 | 40.854 | 25.2857 | 2.0051 | 32.0656 | 64.2956 | 15.9918 | 14.4056 | -0.1105 | 105.1496 | 235.1307 | 180.978 | 232.7693 | 161.2992
2012-10-30 18:08 | 28.4937 | 249.5065 | 195.334 | 247.3683 | 175.5501 | -1.6009 | 1.2711 | 32.1405 | 40.8532 | 25.2859 | 2.0051 | 32.0654 | 64.2951 | 15.9918 | 14.35 | -0.1104 | 105.1483 | 235.1565 | 180.984 | 233.0183 | 161.2001
2012-10-30 18:08 | 28.5004 | 249.5759 | 195.3879 | 247.1909 | 175.8027 | -1.598 | 1.2711 | 32.14 | 40.8528 | 25.2855 | 2.0051 | 32.065 | 64.2932 | 15.9918 | 14.3526 | -0.1102 | 105.146 | 235.2233 | 181.0353 | 232.8383 | 161.4501
2012-10-30 18:08 | 28.4915 | 249.52 | 195.3554 | 247.2447 | 175.6221 | -1.6009 | 1.271 | 32.1393 | 40.8503 | 25.2858 | 2.005 | 32.0643 | 64.2904 | 15.9918 | 14.3641 | -0.1104 | 105.1407 | 235.1559 | 180.9913 | 232.8806 | 161.258
2012-10-30 18:08 | 28.4989 | 249.5156 | 195.3667 | 247.35 | 175.6646 | -1.6009 | 1.271 | 32.1389 | 40.8493 | 25.2858 | 2.005 | 32.0639 | 64.289 | 15.9918 | 14.3674 | -0.1104 | 105.1384 | 235.1482 | 180.9993 | 232.9826 | 161.2972
2012-10-30 18:08 | 28.4974 | 249.5376 | 195.3731 | 247.2937 | 175.7919 | -1.5997 | 1.271 | 32.1387 | 40.8487 | 25.2859 | 2.005 | 32.0636 | 64.2876 | 15.9918 | 14.419 | -0.1103 | 105.1363 | 235.1186 | 180.9541 | 232.8747 | 161.3729
2012-10-30 18:08 | 28.5034 | 249.5867 | 195.3948 | 247.2017 | 175.9268 | -1.6043 | 1.271 | 32.1388 | 40.8498 | 25.2854 | 2.005 | 32.0639 | 64.2886 | 15.9918 | 14.4046 | -0.1106 | 105.1384 | 235.1821 | 180.9902 | 232.7971 | 161.5222
2012-10-30 18:08 | 28.4937 | 249.53 | 195.3222 | 247.337 | 175.7258 | -1.602 | 1.271 | 32.1385 | 40.8485 | 25.2858 | 2.005 | 32.0636 | 64.2886 | 15.9915 | 14.3745 | -0.1104 | 105.1372 | 235.1555 | 180.9477 | 232.9625 | 161.3513
2012-10-30 18:08 | 28.4937 | 249.5496 | 195.3851 | 247.29 | 175.8312 | -1.602 | 1.271 | 32.1382 | 40.8478 | 25.2856 | 2.005 | 32.0632 | 64.2865 | 15.9916 | 14.4316 | -0.1104 | 105.1344 | 235.118 | 180.9535 | 232.8584 | 161.3996
2012-10-30 18:08 | 28.4937 | 249.5535 | 195.4165 | 247.2782 | 175.8468 | -1.5957 | 1.271 | 32.1379 | 40.8474 | 25.2855 | 2.005 | 32.0631 | 64.2863 | 15.9916 | 14.4153 | -0.11 | 105.1338 | 235.1382 | 181.0012 | 232.8629 | 161.4315
2012-10-30 18:08 | 28.4967 | 249.486 | 195.4235 | 247.4732 | 175.6078 | -1.5968 | 1.271 | 32.1377 | 40.8468 | 25.2856 | 2.005 | 32.0628 | 64.2854 | 15.9915 | 14.4019 | -0.1101 | 105.1322 | 235.0841 | 181.0216 | 233.0713 | 161.2059
2012-10-30 18:08 | 28.4952 | 249.5041 | 195.4023 | 247.2876 | 175.7429 | -1.5968 | 1.271 | 32.1377 | 40.8464 | 25.2857 | 2.005 | 32.0626 | 64.2849 | 15.9915 | 14.4249 | -0.1101 | 105.1314 | 235.0792 | 180.9774 | 232.8627 | 161.318
2012-10-30 18:08 | 28.4997 | 249.5595 | 195.4068 | 247.3116 | 175.642 | -1.5997 | 1.271 | 32.1375 | 40.8462 | 25.2856 | 2.005 | 32.0626 | 64.2846 | 15.9915 | 14.409 | -0.1103 | 105.1308 | 235.1505 | 180.9978 | 232.9026 | 161.233
2012-10-30 18:08 | 28.5027 | 249.5429 | 195.3666 | 247.3852 | 175.645 | -1.5997 | 1.271 | 32.137 | 40.8449 | 25.2856 | 2.0049 | 32.0621 | 64.2824 | 15.9916 | 14.4161 | -0.1103 | 105.1273 | 235.1268 | 180.9505 | 232.9691 | 161.2289
2012-10-30 18:08 | 28.5102 | 249.5935 | 195.3977 | 247.3026 | 175.7696 | -1.6009 | 1.2709 | 32.1365 | 40.8439 | 25.2855 | 2.0049 | 32.0616 | 64.2808 | 15.9915 | 14.4399 | -0.1104 | 105.1247 | 235.1536 | 180.9578 | 232.8627 | 161.3297
2012-10-30 18:08 | 28.4959 | 249.5361 | 195.3598 | 247.2217 | 175.6889 | -1.6003 | 1.2709 | 32.1361 | 40.843 | 25.2853 | 2.0049 | 32.0612 | 64.2794 | 15.9914 | 14.3767 | -0.1103 | 105.1224 | 235.1594 | 180.9831 | 232.845 | 161.3122
2012-10-30 18:08 | 28.5057 | 249.5264 | 195.3539 | 247.4626 | 175.5699 | -1.6026 | 1.2709 | 32.1361 | 40.8425 | 25.2856 | 2.0049 | 32.061 | 64.2789 | 15.9914 | 14.3905 | -0.1105 | 105.1214 | 235.1359 | 180.9634 | 233.0721 | 161.1794
2012-10-30 18:08 | 28.4989 | 249.5313 | 195.3235 | 247.2599 | 175.6451 | -1.6014 | 1.2709 | 32.1357 | 40.8417 | 25.2855 | 2.0049 | 32.0608 | 64.278 | 15.9913 | 14.3797 | -0.1104 | 105.1197 | 235.1516 | 180.9438 | 232.8802 | 161.2654
2012-10-30 18:08 | 28.5162 | 249.5682 | 195.2897 | 247.3673 | 175.6975 | -1.5968 | 1.2709 | 32.1354 | 40.8408 | 25.2855 | 2.0049 | 32.0603 | 64.2762 | 15.9914 | 14.4092 | -0.1101 | 105.117 | 235.159 | 180.8805 | 232.9581 | 161.2883
2012-10-30 18:08 | 28.5094 | 249.5222 | 195.334 | 247.3958 | 175.6868 | -1.6014 | 1.2709 | 32.1351 | 40.8399 | 25.2856 | 2.0049 | 32.0601 | 64.2757 | 15.9913 | 14.4024 | -0.1104 | 105.1156 | 235.1198 | 180.9316 | 232.9934 | 161.2844
2012-10-30 18:08 | 28.5094 | 249.5575 | 195.3497 | 247.3566 | 175.4644 | -1.6049 | 1.2709 | 32.1344 | 40.8381 | 25.2857 | 2.0048 | 32.0594 | 64.272 | 15.9914 | 14.3983 | -0.1106 | 105.1101 | 235.1592 | 180.9514 | 232.9583 | 161.0661
2012-10-30 18:08 | 28.5109 | 249.5511 | 195.2883 | 247.2758 | 175.4971 | -1.6084 | 1.2708 | 32.134 | 40.837 | 25.2857 | 2.0048 | 32.059 | 64.271 | 15.9913 | 14.3631 | -0.1109 | 105.108 | 235.188 | 180.9252 | 232.9127 | 161.134
2012-10-30 18:08 | 28.5057 | 249.5146 | 195.2792 | 247.4234 | 175.4802 | -1.5945 | 1.2708 | 32.1336 | 40.836 | 25.2858 | 2.0048 | 32.0588 | 64.2702 | 15.9913 | 14.3375 | -0.1099 | 105.1062 | 235.1771 | 180.9417 | 233.0859 | 161.1427
2012-10-30 18:09 | 28.5207 | 249.5414 | 195.3335 | 247.3914 | 175.5654 | -1.602 | 1.2708 | 32.1333 | 40.8358 | 25.2855 | 2.0048 | 32.0584 | 64.2693 | 15.9911 | 14.377 | -0.1104 | 105.1051 | 235.1644 | 180.9565 | 233.0144 | 161.1884
2012-10-30 18:09 | 28.5094 | 249.5731 | 195.2947 | 247.2783 | 175.6127 | -1.6014 | 1.2708 | 32.1325 | 40.834 | 25.2853 | 2.0047 | 32.0577 | 64.2662 | 15.9912 | 14.3698 | -0.1104 | 105.1002 | 235.2033 | 180.9249 | 232.9085 | 161.2429
2012-10-30 18:09 | 28.5072 | 249.5239 | 195.2453 | 247.3309 | 175.3724 | -1.6026 | 1.2654 | 31.996 | 40.4889 | 25.2845 | 1.9965 | 31.9268 | 63.7427 | 15.9912 | 14.3594 | -0.1105 | 104.2316 | 235.1645 | 180.8859 | 232.9715 | 161.013
2012-10-30 18:09 | 28.5079 | 249.5481 | 195.2146 | 247.4178 | 175.4707 | -1.5997 | 1.2647 | 31.9779 | 40.4435 | 25.2843 | 1.995 | 31.9034 | 63.6488 | 15.9913 | 14.3642 | -0.1103 | 104.0923 | 235.1839 | 180.8504 | 233.0536 | 161.1065
2012-10-30 18:09 | 28.5094 | 249.5301 | 195.2083 | 247.2626 | 175.6244 | -1.5974 | 1.2641 | 31.9605 | 40.4 | 25.284 | 1.9939 | 31.8859 | 63.5785 | 15.9914 | 14.3738 | -0.1101 | 103.9784 | 235.1563 | 180.8345 | 232.8888 | 161.2506
2012-10-30 18:09 | 28.5072 | 249.5161 | 195.1786 | 247.3035 | 175.5636 | -1.6037 | 1.2637 | 31.9505 | 40.3753 | 25.2836 | 1.9933 | 31.8761 | 63.5388 | 15.9916 | 14.3961 | -0.1106 | 103.9141 | 235.12 | 180.7825 | 232.9074 | 161.1675
2012-10-30 18:09 | 28.5057 | 249.4637 | 195.1181 | 247.3842 | 175.2538 | -1.6003 | 1.2636 | 31.9475 | 40.3676 | 25.2837 | 1.9931 | 31.873 | 63.5263 | 15.9916 | 14.3987 | -0.1103 | 103.8939 | 235.065 | 180.7194 | 232.9855 | 160.8551
2012-10-30 18:09 | 28.5184 | 249.4608 | 195.1976 | 247.3734 | 175.2704 | -1.602 | 1.2635 | 31.9452 | 40.362 | 25.2836 | 1.993 | 31.8707 | 63.5172 | 15.9916 | 14.3869 | -0.1104 | 103.8792 | 235.0739 | 180.8107 | 232.9865 | 160.8835
2012-10-30 18:09 | 28.5049 | 249.4747 | 195.137 | 247.2659 | 175.7409 | -1.6026 | 1.2635 | 31.9452 | 40.3622 | 25.2834 | 1.993 | 31.8707 | 63.5168 | 15.9917 | 14.3489 | -0.1105 | 103.879 | 235.1258 | 180.7881 | 232.917 | 161.392
2012-10-30 18:09 | 28.5034 | 249.4144 | 195.1315 | 247.3859 | 175.4779 | -1.6043 | 1.2635 | 31.9465 | 40.3647 | 25.284 | 1.993 | 31.8721 | 63.522 | 15.9918 | 14.4127 | -0.1106 | 103.8867 | 235.0017 | 180.7188 | 232.9732 | 161.0652
2012-10-30 18:09 | 28.5057 | 249.3776 | 195.181 | 247.3333 | 175.0158 | -1.6009 | 1.2636 | 31.9474 | 40.3672 | 25.2838 | 1.9931 | 31.873 | 63.526 | 15.9917 | 14.4028 | -0.1104 | 103.8933 | 234.9748 | 180.7782 | 232.9305 | 160.613
2012-10-30 18:09 | 28.5012 | 249.3927 | 195.0822 | 247.227 | 175.2376 | -1.5933 | 1.2637 | 31.9503 | 40.3743 | 25.2839 | 1.9933 | 31.8757 | 63.5371 | 15.9916 | 14.3697 | -0.1098 | 103.9114 | 235.023 | 180.7125 | 232.8573 | 160.8679
2012-10-30 18:09 | 28.5049 | 249.3807 | 195.0427 | 247.3168 | 175.2374 | -1.5962 | 1.2662 | 32.0154 | 40.5385 | 25.2843 | 1.997 | 31.9358 | 63.7765 | 15.9917 | 14.3897 | -0.11 | 104.315 | 234.991 | 180.653 | 232.9271 | 160.8477
2012-10-30 18:09 | 28.4907 | 249.3313 | 195.1031 | 247.3301 | 175.4847 | -1.5997 | 1.2687 | 32.0801 | 40.7012 | 25.285 | 2.0013 | 32.0048 | 64.0524 | 15.9917 | 14.3633 | -0.1103 | 104.7536 | 234.968 | 180.7398 | 232.9668 | 161.1214
2012-10-30 18:09 | 28.4944 | 249.3741 | 195.1304 | 247.2084 | 175.7226 | -1.6049 | 1.2694 | 32.0977 | 40.7461 | 25.285 | 2.0025 | 32.0225 | 64.1241 | 15.9915 | 14.3792 | -0.1106 | 104.8701 | 234.9949 | 180.7512 | 232.8292 | 161.3434
2012-10-30 18:09 | 28.4967 | 249.3999 | 195.1681 | 247.199 | 175.6663 | -1.5991 | 1.27 | 32.1113 | 40.7799 | 25.2854 | 2.0033 | 32.0362 | 64.1792 | 15.9914 | 14.3979 | -0.1102 | 104.9591 | 235.002 | 180.7702 | 232.8011 | 161.2684
2012-10-30 18:09 | 28.4959 | 249.3404 | 195.1751 | 247.3313 | 175.6928 | -1.6043 | 1.2703 | 32.1199 | 40.8015 | 25.2855 | 2.0039 | 32.045 | 64.2141 | 15.9916 | 14.4215 | -0.1106 | 105.0156 | 234.9189 | 180.7536 | 232.9098 | 161.2713
2012-10-30 18:09 | 28.4892 | 249.3611 | 195.2313 | 247.2385 | 175.6627 | -1.6037 | 1.2705 | 32.1256 | 40.8162 | 25.2854 | 2.0042 | 32.0508 | 64.2372 | 15.9915 | 14.3944 | -0.1106 | 105.0535 | 234.9667 | 180.8369 | 232.8441 | 161.2683
2012-10-30 18:09 | 28.4982 | 249.3857 | 195.2128 | 247.1378 | 175.7498 | -1.5928 | 1.2707 | 32.1294 | 40.8254 | 25.2857 | 2.0045 | 32.0545 | 64.2524 | 15.9915 | 14.3586 | -0.1098 | 105.0779 | 235.0271 | 180.8542 | 232.7792 | 161.3912
2012-10-30 18:09 | 28.4922 | 249.3445 | 195.2343 | 247.2728 | 175.5916 | -1.5962 | 1.2708 | 32.1325 | 40.8336 | 25.2855 | 2.0047 | 32.0574 | 64.2644 | 15.9914 | 14.3811 | -0.11 | 105.098 | 234.9634 | 180.8532 | 232.8917 | 161.2105
2012-10-30 18:09 | 28.4885 | 249.4035 | 195.262 | 247.2613 | 175.7323 | -1.6026 | 1.2709 | 32.1351 | 40.8397 | 25.2858 | 2.0048 | 32.06 | 64.2751 | 15.9914 | 14.3652 | -0.1105 | 105.1148 | 235.0383 | 180.8968 | 232.8961 | 161.3671
2012-10-30 18:09 | 28.4862 | 249.3933 | 195.3265 | 247.1219 | 175.8666 | -1.6043 | 1.2709 | 32.1364 | 40.8426 | 25.2861 | 2.0049 | 32.0613 | 64.2804 | 15.9913 | 14.3955 | -0.1106 | 105.123 | 234.9978 | 180.931 | 232.7264 | 161.4711
2012-10-30 18:09 | 28.4907 | 249.3822 | 195.276 | 247.2008 | 175.7111 | -1.6009 | 1.271 | 32.1373 | 40.845 | 25.2859 | 2.005 | 32.0624 | 64.2844 | 15.9914 | 14.3878 | -0.1104 | 105.1294 | 234.9944 | 180.8882 | 232.813 | 161.3233
2012-10-30 18:09 | 28.4915 | 249.3673 | 195.324 | 247.2799 | 175.6923 | -1.598 | 1.271 | 32.1378 | 40.8467 | 25.2858 | 2.005 | 32.0628 | 64.2862 | 15.9914 | 14.4008 | -0.1102 | 105.1329 | 234.9665 | 180.9232 | 232.8791 | 161.2915
2012-10-30 18:09 | 28.49 | 249.3971 | 195.3146 | 247.1413 | 175.7728 | -1.6003 | 1.271 | 32.1388 | 40.849 | 25.2858 | 2.0051 | 32.0637 | 64.2907 | 15.9912 | 14.4156 | -0.1103 | 105.1397 | 234.9815 | 180.899 | 232.7257 | 161.3572
2012-10-30 18:09 | 28.4915 | 249.4182 | 195.2886 | 247.1781 | 175.6767 | -1.6003 | 1.271 | 32.1395 | 40.8507 | 25.286 | 2.0051 | 32.0645 | 64.2942 | 15.9911 | 14.4212 | -0.1103 | 105.1449 | 234.997 | 180.8674 | 232.7569 | 161.2555
2012-10-30 18:09 | 28.4892 | 249.3376 | 195.3413 | 247.2581 | 175.6861 | -1.6049 | 1.2711 | 32.14 | 40.8521 | 25.2859 | 2.0052 | 32.0649 | 64.2956 | 15.9911 | 14.4148 | -0.1106 | 105.1477 | 234.9228 | 180.9265 | 232.8433 | 161.2713
2012-10-30 18:10 | 28.4952 | 249.4102 | 195.3277 | 247.1857 | 175.7429 | -1.6043 | 1.2711 | 32.1401 | 40.853 | 25.2855 | 2.0052 | 32.0651 | 64.2964 | 15.9911 | 14.4086 | -0.1106 | 105.1494 | 235.0016 | 180.9191 | 232.7771 | 161.3343
2012-10-30 18:10 | 28.487 | 249.398 | 195.2919 | 247.107 | 175.6371 | -1.5968 | 1.2711 | 32.1406 | 40.8536 | 25.2859 | 2.0052 | 32.0655 | 64.298 | 15.9911 | 14.3637 | -0.1101 | 105.1516 | 235.0343 | 180.9282 | 232.7433 | 161.2734
2012-10-30 18:10 | 28.4877 | 249.3557 | 195.2926 | 247.2526 | 175.7393 | -1.6026 | 1.2711 | 32.1408 | 40.8542 | 25.2859 | 2.0052 | 32.0658 | 64.2994 | 15.9911 | 14.3848 | -0.1105 | 105.1536 | 234.9709 | 180.9078 | 232.8678 | 161.3545
2012-10-30 18:10 | 28.49 | 249.3501 | 195.3696 | 247.1805 | 175.847 | -1.6037 | 1.2711 | 32.1415 | 40.8561 | 25.2857 | 2.0053 | 32.0664 | 64.302 | 15.9911 | 14.3952 | -0.1106 | 105.1581 | 234.9549 | 180.9744 | 232.7853 | 161.4518
2012-10-30 18:10 | 28.4922 | 249.458 | 195.3482 | 247.1435 | 175.8609 | -1.6014 | 1.2711 | 32.1414 | 40.8558 | 25.2858 | 2.0053 | 32.0663 | 64.3017 | 15.991 | 14.4178 | -0.1104 | 105.1575 | 235.0402 | 180.9304 | 232.7257 | 161.4431
2012-10-30 18:10 | 28.493 | 249.3375 | 195.3412 | 247.2893 | 175.7056 | -1.6043 | 1.2711 | 32.1413 | 40.8552 | 25.2859 | 2.0053 | 32.0662 | 64.3014 | 15.9909 | 14.3901 | -0.1106 | 105.1566 | 234.9474 | 180.9511 | 232.8992 | 161.3155
2012-10-30 18:10 | 28.4847 | 249.3879 | 195.3957 | 247.2026 | 175.8573 | -1.5985 | 1.2711 | 32.1412 | 40.8551 | 25.2858 | 2.0053 | 32.0662 | 64.3012 | 15.991 | 14.3899 | -0.1102 | 105.1563 | 234.998 | 181.0058 | 232.8127 | 161.4674
2012-10-30 18:10 | 28.4832 | 249.3551 | 195.3432 | 247.158 | 175.8753 | -1.6095 | 1.2711 | 32.1406 | 40.854 | 25.2856 | 2.0052 | 32.0656 | 64.2989 | 15.991 | 14.3558 | -0.111 | 105.153 | 234.9993 | 180.9874 | 232.8022 | 161.5195
2012-10-30 18:10 | 28.4885 | 249.3526 | 195.3602 | 247.273 | 175.8806 | -1.6055 | 1.2711 | 32.1408 | 40.854 | 25.2859 | 2.0052 | 32.0657 | 64.2994 | 15.9909 | 14.3815 | -0.1107 | 105.1534 | 234.9711 | 180.9787 | 232.8915 | 161.4991
2012-10-30 18:10 | 28.4855 | 249.3809 | 195.4437 | 247.2113 | 175.9596 | -1.602 | 1.2711 | 32.1409 | 40.854 | 25.2861 | 2.0053 | 32.0659 | 64.3004 | 15.9909 | 14.3581 | -0.1104 | 105.1544 | 235.0228 | 181.0856 | 232.8532 | 161.6015
2012-10-30 18:10 | 28.4862 | 249.4051 | 195.4208 | 247.1924 | 175.8783 | -1.6009 | 1.2711 | 32.1408 | 40.8538 | 25.286 | 2.0052 | 32.0657 | 64.2992 | 15.991 | 14.3547 | -0.1104 | 105.153 | 235.0504 | 181.0661 | 232.8377 | 161.5236
2012-10-30 18:10 | 28.4892 | 249.3885 | 195.4042 | 247.262 | 175.7915 | -1.5985 | 1.2711 | 32.1406 | 40.8534 | 25.286 | 2.0052 | 32.0656 | 64.2984 | 15.9911 | 14.3577 | -0.1102 | 105.1518 | 235.0308 | 181.0465 | 232.9043 | 161.4338
2012-10-30 18:10 | 28.4915 | 249.4182 | 195.4458 | 247.2564 | 175.8446 | -1.6026 | 1.2711 | 32.1404 | 40.8526 | 25.2862 | 2.0052 | 32.0654 | 64.2978 | 15.991 | 14.3804 | -0.1105 | 105.1505 | 235.0378 | 181.0654 | 232.876 | 161.4642
2012-10-30 18:10 | 28.4974 | 249.4515 | 195.4477 | 247.1801 | 175.8934 | -1.598 | 1.2711 | 32.1405 | 40.8528 | 25.2861 | 2.0052 | 32.0654 | 64.2983 | 15.9909 | 14.3537 | -0.1102 | 105.1511 | 235.0978 | 181.094 | 232.8264 | 161.5397
2012-10-30 18:10 | 28.4959 | 249.3795 | 195.4266 | 247.2177 | 175.9114 | -1.6003 | 1.2711 | 32.1405 | 40.8532 | 25.286 | 2.0052 | 32.0654 | 64.2982 | 15.991 | 14.3726 | -0.1103 | 105.1514 | 235.0069 | 181.054 | 232.8451 | 161.5388
2012-10-30 18:10 | 28.493 | 249.3962 | 195.4512 | 247.301 | 175.9944 | -1.602 | 1.2711 | 32.1405 | 40.8528 | 25.2861 | 2.0052 | 32.0655 | 64.299 | 15.9908 | 14.4023 | -0.1104 | 105.1518 | 234.9939 | 181.0489 | 232.8987 | 161.5921
2012-10-30 18:10 | 28.487 | 249.4098 | 195.5316 | 247.1618 | 175.9104 | -1.5957 | 1.2711 | 32.1401 | 40.8519 | 25.2861 | 2.0052 | 32.0649 | 64.2966 | 15.9909 | 14.3514 | -0.11 | 105.1485 | 235.0584 | 181.1802 | 232.8104 | 161.559
2012-10-30 18:10 | 28.4937 | 249.4126 | 195.4637 | 247.2469 | 175.718 | -1.6084 | 1.271 | 32.1398 | 40.8509 | 25.2863 | 2.0052 | 32.0648 | 64.2962 | 15.9909 | 14.3622 | -0.1109 | 105.1471 | 235.0504 | 181.1015 | 232.8847 | 161.3558
2012-10-30 18:10 | 28.4959 | 249.4422 | 195.4462 | 247.347 | 175.7358 | -1.5991 | 1.271 | 32.1395 | 40.8508 | 25.2859 | 2.0052 | 32.0645 | 64.2954 | 15.9907 | 14.3644 | -0.1102 | 105.1462 | 235.0778 | 181.0818 | 232.9826 | 161.3714
2012-10-30 18:10 | 28.4922 | 249.4267 | 195.4779 | 247.2219 | 175.9077 | -1.6014 | 1.2711 | 32.1398 | 40.8516 | 25.2858 | 2.0052 | 32.0647 | 64.296 | 15.9908 | 14.3566 | -0.1104 | 105.1476 | 235.0701 | 181.1213 | 232.8653 | 161.5511
2012-10-30 18:10 | 28.4959 | 249.4069 | 195.4384 | 247.2138 | 175.8412 | -1.6032 | 1.271 | 32.1397 | 40.8509 | 25.2861 | 2.0052 | 32.0647 | 64.296 | 15.9908 | 14.3481 | -0.1105 | 105.1469 | 235.0588 | 181.0903 | 232.8657 | 161.4931
2012-10-30 18:10 | 28.4915 | 249.4065 | 195.4693 | 247.3387 | 175.8719 | -1.5997 | 1.271 | 32.1393 | 40.8503 | 25.2859 | 2.0052 | 32.0643 | 64.2949 | 15.9907 | 14.3396 | -0.1103 | 105.1453 | 235.0669 | 181.1297 | 232.9991 | 161.5323
2012-10-30 18:10 | 28.4847 | 249.3997 | 195.5215 | 247.2457 | 175.8729 | -1.6049 | 1.271 | 32.1392 | 40.8491 | 25.2864 | 2.0052 | 32.0641 | 64.2942 | 15.9907 | 14.3247 | -0.1106 | 105.1434 | 235.075 | 181.1968 | 232.921 | 161.5482
2012-10-30 18:10 | 28.4967 | 249.4273 | 195.4745 | 247.2264 | 175.7171 | -1.6032 | 1.271 | 32.1388 | 40.8489 | 25.2859 | 2.0052 | 32.0637 | 64.2927 | 15.9906 | 14.3815 | -0.1105 | 105.1416 | 235.0458 | 181.093 | 232.8449 | 161.3356
2012-10-30 18:10 | 28.4915 | 249.4652 | 195.3593 | 247.2643 | 175.6143 | -1.5997 | 1.271 | 32.1385 | 40.8478 | 25.2862 | 2.0051 | 32.0635 | 64.2915 | 15.9908 | 14.3682 | -0.1103 | 105.1393 | 235.097 | 180.9911 | 232.8961 | 161.2461
2012-10-30 18:10 | 28.4937 | 249.4361 | 195.444 | 247.3252 | 175.6594 | -1.6037 | 1.271 | 32.1385 | 40.8478 | 25.2862 | 2.0051 | 32.0636 | 64.2923 | 15.9906 | 14.3989 | -0.1106 | 105.14 | 235.0372 | 181.0451 | 232.9263 | 161.2605
2012-10-30 18:11 | 28.5012 | 249.471 | 195.4633 | 247.227 | 175.6591 | -1.5951 | 1.271 | 32.1383 | 40.8483 | 25.2856 | 2.0051 | 32.0633 | 64.2912 | 15.9906 | 14.3738 | -0.11 | 105.1394 | 235.0972 | 181.0895 | 232.8532 | 161.2853
2012-10-30 18:11 | 28.5012 | 249.4827 | 195.3729 | 247.2936 | 175.5537 | -1.6032 | 1.271 | 32.1382 | 40.847 | 25.2862 | 2.0051 | 32.0633 | 64.2912 | 15.9906 | 14.386 | -0.1105 | 105.1381 | 235.0967 | 180.9869 | 232.9076 | 161.1677
2012-10-30 18:11 | 28.5027 | 249.4177 | 195.3784 | 247.3342 | 175.6333 | -1.6037 | 1.271 | 32.1378 | 40.8466 | 25.2858 | 2.0051 | 32.0628 | 64.2893 | 15.9906 | 14.3916 | -0.1106 | 105.1359 | 235.0261 | 180.9868 | 232.9426 | 161.2417
2012-10-30 18:11 | 28.4907 | 249.4683 | 195.3978 | 247.236 | 175.7774 | -1.6014 | 1.271 | 32.1375 | 40.846 | 25.2857 | 2.0051 | 32.0624 | 64.2876 | 15.9907 | 14.3347 | -0.1104 | 105.1336 | 235.1336 | 181.0631 | 232.9013 | 161.4427
2012-10-30 18:11 | 28.5057 | 249.4402 | 195.3499 | 247.2393 | 175.5973 | -1.5985 | 1.2709 | 32.1373 | 40.8443 | 25.2864 | 2.0051 | 32.0623 | 64.2877 | 15.9904 | 14.4028 | -0.1102 | 105.132 | 235.0374 | 180.9471 | 232.8365 | 161.1945
2012-10-30 18:11 | 28.5034 | 249.4418 | 195.3987 | 247.3702 | 175.6535 | -1.6043 | 1.2709 | 32.1368 | 40.8443 | 25.2857 | 2.0051 | 32.0619 | 64.2858 | 15.9905 | 14.3964 | -0.1106 | 105.1301 | 235.0454 | 181.0023 | 232.9738 | 161.2571
2012-10-30 18:11 | 28.4907 | 249.437 | 195.3821 | 247.2047 | 175.6057 | -1.6037 | 1.2709 | 32.1363 | 40.8424 | 25.286 | 2.005 | 32.0613 | 64.2841 | 15.9903 | 14.3266 | -0.1106 | 105.1265 | 235.1104 | 181.0555 | 232.8781 | 161.2791
2012-10-30 18:11 | 28.5019 | 249.4403 | 195.3265 | 247.2237 | 175.4178 | -1.5991 | 1.2709 | 32.1362 | 40.8428 | 25.2857 | 2.005 | 32.0613 | 64.2839 | 15.9905 | 14.3459 | -0.1102 | 105.1267 | 235.0944 | 180.9806 | 232.8778 | 161.0719
2012-10-30 18:11 | 28.4885 | 249.4426 | 195.3052 | 247.3357 | 175.4278 | -1.5991 | 1.2709 | 32.1363 | 40.8428 | 25.2858 | 2.0051 | 32.0613 | 64.2848 | 15.9902 | 14.3529 | -0.1102 | 105.1275 | 235.0897 | 180.9523 | 232.9828 | 161.0749
2012-10-30 18:11 | 28.4952 | 249.4728 | 195.3277 | 247.2484 | 175.7234 | -1.6026 | 1.2709 | 32.136 | 40.8417 | 25.2861 | 2.005 | 32.061 | 64.2835 | 15.9902 | 14.3474 | -0.1105 | 105.1252 | 235.1254 | 180.9803 | 232.901 | 161.376
2012-10-30 18:11 | 28.4892 | 249.4785 | 195.247 | 247.2933 | 175.5808 | -1.5997 | 1.2709 | 32.1358 | 40.8413 | 25.2859 | 2.005 | 32.0607 | 64.2819 | 15.9903 | 14.374 | -0.1103 | 105.1231 | 235.1045 | 180.873 | 232.9193 | 161.2068
2012-10-30 18:11 | 28.4922 | 249.4698 | 195.2343 | 247.3237 | 175.4393 | -1.6003 | 1.2683 | 32.0695 | 40.674 | 25.2853 | 2.0009 | 31.9947 | 64.0177 | 15.9903 | 14.3526 | -0.1103 | 104.6917 | 235.1172 | 180.8817 | 232.9711 | 161.0867
2012-10-30 18:11 | 28.49 | 249.3893 | 195.1771 | 247.4155 | 175.1093 | -1.5985 | 1.2635 | 31.9472 | 40.3659 | 25.2843 | 1.9933 | 31.8727 | 63.5304 | 15.9903 | 14.3585 | -0.1102 | 103.8963 | 235.0308 | 180.8186 | 233.057 | 160.7508
2012-10-30 18:11 | 28.484 | 249.3637 | 195.1043 | 247.2058 | 175.2321 | -1.602 | 1.2635 | 31.9458 | 40.3633 | 25.2838 | 1.9931 | 31.8713 | 63.5238 | 15.9906 | 14.3648 | -0.1104 | 103.8871 | 234.9989 | 180.7395 | 232.841 | 160.8673
2012-10-30 18:11 | 28.4982 | 249.3897 | 195.091 | 247.3023 | 175.0746 | -1.6037 | 1.2635 | 31.9455 | 40.3625 | 25.2838 | 1.9931 | 31.8711 | 63.5233 | 15.9904 | 14.3463 | -0.1106 | 103.8858 | 235.0434 | 180.7447 | 232.956 | 160.7283
2012-10-30 18:11 | 28.484 | 249.3755 | 195.065 | 247.339 | 175.0994 | -1.602 | 1.2635 | 31.9473 | 40.3665 | 25.2841 | 1.9932 | 31.8728 | 63.5295 | 15.9906 | 14.3484 | -0.1104 | 103.896 | 235.0271 | 180.7166 | 232.9906 | 160.751
2012-10-30 18:11 | 28.4907 | 249.343 | 195.0874 | 247.1734 | 175.2388 | -1.6003 | 1.2635 | 31.9472 | 40.3666 | 25.2839 | 1.9932 | 31.8725 | 63.5291 | 15.9904 | 14.3347 | -0.1103 | 103.8957 | 235.0083 | 180.7527 | 232.8387 | 160.9041
2012-10-30 18:11 | 28.5004 | 249.3566 | 195.0499 | 247.2967 | 175.0534 | -1.6026 | 1.2635 | 31.9478 | 40.3676 | 25.2842 | 1.9932 | 31.8733 | 63.5312 | 15.9907 | 14.3404 | -0.1105 | 103.8988 | 235.0162 | 180.7095 | 232.9563 | 160.713
2012-10-30 18:12 | 28.4922 | 249.3249 | 194.9946 | 247.2846 | 175.1818 | -1.6066 | 1.2636 | 31.9493 | 40.3707 | 25.2846 | 1.9933 | 31.8745 | 63.5365 | 15.9906 | 14.3281 | -0.1108 | 103.9072 | 234.9968 | 180.6665 | 232.9565 | 160.8537
2012-10-30 18:12 | 28.4952 | 249.3084 | 195.0408 | 247.2131 | 175.3409 | -1.6043 | 1.268 | 32.0615 | 40.6542 | 25.2849 | 2.0002 | 31.9841 | 63.974 | 15.9906 | 14.3025 | -0.1106 | 104.6283 | 235.0059 | 180.7383 | 232.9106 | 161.0384
2012-10-30 18:12 | 28.4877 | 249.3518 | 195.0097 | 247.2252 | 175.1812 | -1.5945 | 1.2691 | 32.0898 | 40.7252 | 25.2854 | 2.0021 | 32.0149 | 64.098 | 15.9904 | 14.2991 | -0.1099 | 104.8232 | 235.0527 | 180.7106 | 232.9261 | 160.8821
2012-10-30 18:12 | 28.4892 | 249.2984 | 195.0152 | 247.2737 | 175.4051 | -1.6014 | 1.2696 | 32.1038 | 40.7605 | 25.2857 | 2.003 | 32.0288 | 64.1533 | 15.9905 | 14.3047 | -0.1104 | 104.9138 | 234.9937 | 180.7105 | 232.969 | 161.1004
2012-10-30 18:12 | 28.4989 | 249.3669 | 195.0838 | 247.1581 | 175.6763 | -1.5962 | 1.2701 | 32.1154 | 40.7902 | 25.2855 | 2.0037 | 32.0404 | 64.1999 | 15.9905 | 14.3021 | -0.11 | 104.9901 | 235.0648 | 180.7817 | 232.856 | 161.3742
2012-10-30 18:12 | 28.4937 | 249.3617 | 195.0393 | 247.1254 | 175.6048 | -1.5957 | 1.2704 | 32.1229 | 40.8088 | 25.2858 | 2.0042 | 32.0479 | 64.2305 | 15.9903 | 14.2806 | -0.11 | 105.0392 | 235.0811 | 180.7587 | 232.8448 | 161.3242
2012-10-30 18:12 | 28.4877 | 249.2969 | 195.1119 | 247.2722 | 175.6144 | -1.602 | 1.2706 | 32.1284 | 40.8226 | 25.2859 | 2.0045 | 32.0534 | 64.2525 | 15.9903 | 14.3073 | -0.1104 | 105.075 | 234.9896 | 180.8046 | 232.9649 | 161.3071
2012-10-30 18:12 | 28.4817 | 249.3692 | 195.1491 | 247.1604 | 175.7879 | -1.6032 | 1.2708 | 32.1323 | 40.8324 | 25.286 | 2.0048 | 32.0573 | 64.2679 | 15.9904 | 14.3094 | -0.1105 | 105.1003 | 235.0598 | 180.8397 | 232.851 | 161.4785
2012-10-30 18:12 | 28.49 | 249.3736 | 195.1653 | 247.0982 | 175.644 | -1.6032 | 1.2708 | 32.1346 | 40.8375 | 25.2863 | 2.0049 | 32.0595 | 64.2772 | 15.9903 | 14.2932 | -0.1105 | 105.1147 | 235.0804 | 180.8721 | 232.805 | 161.3508
2012-10-30 18:12 | 28.4885 | 249.3604 | 195.2188 | 247.2143 | 175.5918 | -1.6009 | 1.2709 | 32.1363 | 40.842 | 25.2862 | 2.005 | 32.0613 | 64.2842 | 15.9903 | 14.304 | -0.1104 | 105.1262 | 235.0564 | 180.9148 | 232.9103 | 161.2878
2012-10-30 18:12 | 28.4772 | 249.3373 | 195.1839 | 247.1873 | 175.5805 | -1.6003 | 1.271 | 32.1379 | 40.8467 | 25.2859 | 2.0051 | 32.0628 | 64.2902 | 15.9903 | 14.2968 | -0.1103 | 105.1369 | 235.0405 | 180.8871 | 232.8905 | 161.2837
2012-10-30 18:12 | 28.4817 | 249.3379 | 195.2198 | 247.1174 | 175.706 | -1.5997 | 1.271 | 32.1389 | 40.849 | 25.286 | 2.0052 | 32.0637 | 64.2946 | 15.9901 | 14.3217 | -0.1103 | 105.1436 | 235.0162 | 180.8981 | 232.7957 | 161.3843
2012-10-30 18:12 | 28.4787 | 249.3506 | 195.1736 | 247.1026 | 175.7225 | -1.6014 | 1.271 | 32.1396 | 40.8504 | 25.2863 | 2.0053 | 32.0645 | 64.2983 | 15.9901 | 14.3146 | -0.1104 | 105.1487 | 235.036 | 180.859 | 232.788 | 161.4079
2012-10-30 18:12 | 28.4847 | 249.3448 | 195.2582 | 247.2888 | 175.7051 | -1.6009 | 1.2711 | 32.1403 | 40.8523 | 25.2861 | 2.0053 | 32.0651 | 64.3006 | 15.99 | 14.3043 | -0.1104 | 105.153 | 235.0405 | 180.9539 | 232.9845 | 161.4008
2012-10-30 18:12 | 28.4825 | 249.3583 | 195.3189 | 247.173 | 175.8707 | -1.6037 | 1.2711 | 32.1405 | 40.8535 | 25.2858 | 2.0053 | 32.0654 | 64.3018 | 15.9901 | 14.3225 | -0.1106 | 105.1553 | 235.0358 | 180.9964 | 232.8505 | 161.5482
2012-10-30 18:12 | 28.4922 | 249.4111 | 195.2696 | 247.0887 | 175.8882 | -1.598 | 1.2711 | 32.1409 | 40.8533 | 25.2865 | 2.0053 | 32.0657 | 64.303 | 15.9901 | 14.2954 | -0.1102 | 105.1563 | 235.1157 | 180.9742 | 232.7933 | 161.5928
2012-10-30 18:12 | 28.4877 | 249.3752 | 195.3005 | 247.1273 | 175.8681 | -1.6009 | 1.2711 | 32.141 | 40.8542 | 25.2861 | 2.0054 | 32.0658 | 64.3038 | 15.99 | 14.3073 | -0.1104 | 105.158 | 235.0679 | 180.9932 | 232.82 | 161.5608
2012-10-30 18:12 | 28.4862 | 249.3894 | 195.3226 | 247.1885 | 175.9486 | -1.598 | 1.2711 | 32.1414 | 40.8549 | 25.2863 | 2.0054 | 32.0663 | 64.306 | 15.99 | 14.3098 | -0.1102 | 105.1609 | 235.0796 | 181.0128 | 232.8787 | 161.6388
2012-10-30 18:12 | 28.4944 | 249.3898 | 195.374 | 247.1026 | 176.0622 | -1.6032 | 1.2711 | 32.1418 | 40.8567 | 25.2858 | 2.0054 | 32.0667 | 64.3071 | 15.99 | 14.3548 | -0.1105 | 105.1638 | 235.035 | 181.0192 | 232.7478 | 161.7074
2012-10-30 18:12 | 28.4959 | 249.4343 | 195.3637 | 247.1903 | 175.9427 | -1.598 | 1.2711 | 32.1415 | 40.8553 | 25.2862 | 2.0054 | 32.0664 | 64.3057 | 15.99 | 14.3277 | -0.1102 | 105.161 | 235.1066 | 181.036 | 232.8626 | 161.615
2012-10-30 18:12 | 28.4944 | 249.3585 | 195.3897 | 247.2437 | 175.8709 | -1.598 | 1.2711 | 32.1416 | 40.8554 | 25.2863 | 2.0054 | 32.0664 | 64.3062 | 15.99 | 14.318 | -0.1102 | 105.1616 | 235.0405 | 181.0717 | 232.9257 | 161.5529
2012-10-30 18:12 | 28.4974 | 249.3928 | 195.4517 | 247.1722 | 175.8115 | -1.6032 | 1.2711 | 32.1416 | 40.8555 | 25.2862 | 2.0054 | 32.0664 | 64.3059 | 15.99 | 14.3292 | -0.1105 | 105.1614 | 235.0636 | 181.1225 | 232.843 | 161.4823
2012-10-30 18:12 | 28.49 | 249.405 | 195.4207 | 247.0943 | 175.9641 | -1.5985 | 1.2711 | 32.1416 | 40.8555 | 25.2862 | 2.0054 | 32.0665 | 64.3075 | 15.9898 | 14.33 | -0.1102 | 105.163 | 235.075 | 181.0907 | 232.7643 | 161.6341
2012-10-30 18:12 | 28.4974 | 249.3928 | 195.4124 | 247.231 | 175.7841 | -1.6026 | 1.2711 | 32.1416 | 40.856 | 25.286 | 2.0054 | 32.0665 | 64.3071 | 15.9898 | 14.3578 | -0.1105 | 105.1631 | 235.035 | 181.0546 | 232.8732 | 161.4263
2012-10-30 18:12 | 28.4967 | 249.4469 | 195.4352 | 247.2068 | 175.8966 | -1.6026 | 1.2711 | 32.1416 | 40.8556 | 25.2862 | 2.0054 | 32.0665 | 64.3071 | 15.9898 | 14.2469 | -0.1105 | 105.1626 | 235.2 | 181.1883 | 232.9599 | 161.6497
2012-10-30 18:12 | 28.4997 | 249.4029 | 195.4225 | 247.1432 | 175.9387 | -1.6009 | 1.2711 | 32.1411 | 40.8545 | 25.286 | 2.0054 | 32.0659 | 64.305 | 15.9898 | 14.2703 | -0.1104 | 105.1595 | 235.1326 | 181.1522 | 232.8729 | 161.6684
2012-10-30 18:13 | 28.5012 | 249.4318 | 195.4358 | 247.2113 | 175.8426 | -1.5991 | 1.2711 | 32.1413 | 40.8551 | 25.2861 | 2.0054 | 32.0662 | 64.3057 | 15.9899 | 14.2922 | -0.1102 | 105.1607 | 235.1396 | 181.1436 | 232.9191 | 161.5504
2012-10-30 18:13 | 28.5012 | 249.4162 | 195.424 | 247.3053 | 175.9363 | -1.6049 | 1.2711 | 32.1409 | 40.8544 | 25.2859 | 2.0054 | 32.0658 | 64.3042 | 15.9899 | 14.3085 | -0.1106 | 105.1586 | 235.1077 | 181.1155 | 232.9968 | 161.6278
2012-10-30 18:13 | 28.4989 | 249.4334 | 195.4925 | 247.1189 | 176.0159 | -1.598 | 1.2711 | 32.1408 | 40.8532 | 25.2864 | 2.0054 | 32.0658 | 64.3044 | 15.9898 | 14.3225 | -0.1102 | 105.1576 | 235.1109 | 181.17 | 232.7964 | 161.6934
2012-10-30 18:13 | 28.4997 | 249.399 | 195.4304 | 247.2059 | 175.6966 | -1.5991 | 1.2711 | 32.1407 | 40.8529 | 25.2865 | 2.0054 | 32.0656 | 64.3039 | 15.9897 | 14.3397 | -0.1102 | 105.1568 | 235.0593 | 181.0907 | 232.8662 | 161.3569
2012-10-30 18:13 | 28.5004 | 249.4271 | 195.4075 | 247.3241 | 175.7325 | -1.5985 | 1.2711 | 32.1404 | 40.8527 | 25.2861 | 2.0054 | 32.0653 | 64.3032 | 15.9896 | 14.324 | -0.1102 | 105.1559 | 235.1031 | 181.0835 | 233.0001 | 161.4085
2012-10-30 18:13 | 28.4952 | 249.4376 | 195.3984 | 247.2523 | 175.7273 | -1.6043 | 1.2711 | 32.1404 | 40.8528 | 25.286 | 2.0054 | 32.0653 | 64.3031 | 15.9896 | 14.2944 | -0.1106 | 105.1559 | 235.1432 | 181.104 | 232.9579 | 161.4329
2012-10-30 18:13 | 28.5019 | 249.4247 | 195.4129 | 247.1532 | 175.7379 | -1.5997 | 1.2711 | 32.1407 | 40.8536 | 25.2859 | 2.0054 | 32.0655 | 64.3034 | 15.9898 | 14.3174 | -0.1103 | 105.1571 | 235.1073 | 181.0955 | 232.8358 | 161.4205
2012-10-30 18:13 | 28.5087 | 249.4511 | 195.3922 | 247.2854 | 175.7564 | -1.5997 | 1.2711 | 32.1404 | 40.8529 | 25.2859 | 2.0054 | 32.0652 | 64.303 | 15.9896 | 14.2915 | -0.1103 | 105.1559 | 235.1596 | 181.1007 | 232.9939 | 161.4649
2012-10-30 18:13 | 28.4997 | 249.4303 | 195.4382 | 247.3038 | 175.9309 | -1.6037 | 1.2711 | 32.1401 | 40.8521 | 25.286 | 2.0054 | 32.065 | 64.3021 | 15.9896 | 14.3233 | -0.1106 | 105.1542 | 235.107 | 181.1149 | 232.9805 | 161.6076
2012-10-30 18:13 | 28.5034 | 249.481 | 195.4105 | 247.1743 | 175.7589 | -1.6049 | 1.2711 | 32.1401 | 40.8522 | 25.2859 | 2.0054 | 32.0651 | 64.3027 | 15.9895 | 14.3515 | -0.1106 | 105.155 | 235.1295 | 181.059 | 232.8228 | 161.4074
2012-10-30 18:13 | 28.5109 | 249.4259 | 195.4141 | 247.2406 | 175.7157 | -1.5997 | 1.2711 | 32.1398 | 40.8514 | 25.286 | 2.0054 | 32.0648 | 64.302 | 15.9894 | 14.3509 | -0.1103 | 105.1534 | 235.075 | 181.0632 | 232.8897 | 161.3648
2012-10-30 18:13 | 28.5004 | 249.4075 | 195.3879 | 247.3554 | 175.6973 | -1.5968 | 1.271 | 32.1397 | 40.851 | 25.286 | 2.0054 | 32.0646 | 64.3009 | 15.9895 | 14.324 | -0.1101 | 105.1519 | 235.0835 | 181.0639 | 233.0314 | 161.3733
2012-10-30 18:13 | 28.5034 | 249.4653 | 195.4459 | 247.284 | 175.8565 | -1.6072 | 1.271 | 32.1394 | 40.85 | 25.2862 | 2.0053 | 32.0643 | 64.2999 | 15.9894 | 14.3719 | -0.1108 | 105.15 | 235.0934 | 181.074 | 232.9121 | 161.4846
2012-10-30 18:13 | 28.4989 | 249.4413 | 195.4139 | 247.1659 | 175.6451 | -1.6049 | 1.271 | 32.1388 | 40.8493 | 25.2857 | 2.0053 | 32.0639 | 64.2983 | 15.9894 | 14.3511 | -0.1106 | 105.1476 | 235.0902 | 181.0628 | 232.8148 | 161.294
2012-10-30 18:13 | 28.4974 | 249.4398 | 195.4006 | 247.3054 | 175.5382 | -1.6014 | 1.271 | 32.1386 | 40.8484 | 25.2859 | 2.0053 | 32.0635 | 64.2973 | 15.9893 | 14.3374 | -0.1104 | 105.1457 | 235.1024 | 181.0632 | 232.968 | 161.2008
2012-10-30 18:13 | 28.5087 | 249.4785 | 195.3844 | 247.3677 | 175.4402 | -1.602 | 1.271 | 32.1384 | 40.848 | 25.2859 | 2.0053 | 32.0633 | 64.2961 | 15.9894 | 14.3772 | -0.1104 | 105.1441 | 235.1013 | 181.0072 | 232.9905 | 161.063
2012-10-30 18:13 | 28.5019 | 249.4912 | 195.354 | 247.2198 | 175.5583 | -1.6014 | 1.271 | 32.1382 | 40.8472 | 25.2861 | 2.0053 | 32.0631 | 64.2958 | 15.9892 | 14.3745 | -0.1104 | 105.143 | 235.1167 | 180.9795 | 232.8453 | 161.1838
2012-10-30 18:13 | 28.4967 | 249.486 | 195.2977 | 247.293 | 175.5219 | -1.6032 | 1.271 | 32.1379 | 40.8465 | 25.286 | 2.0053 | 32.0627 | 64.2941 | 15.9893 | 14.3244 | -0.1105 | 105.1407 | 235.1616 | 180.9733 | 232.9686 | 161.1975
2012-10-30 18:13 | 28.4937 | 249.4752 | 195.334 | 247.3566 | 175.6828 | -1.6072 | 1.271 | 32.1377 | 40.8464 | 25.2857 | 2.0053 | 32.0626 | 64.2939 | 15.9892 | 14.3459 | -0.1108 | 105.1403 | 235.1293 | 180.9881 | 233.0107 | 161.3369
2012-10-30 18:13 | 28.4959 | 249.497 | 195.3559 | 247.2961 | 175.728 | -1.6043 | 1.271 | 32.1373 | 40.8451 | 25.2859 | 2.0052 | 32.0622 | 64.2923 | 15.9892 | 14.3399 | -0.1106 | 105.1374 | 235.1571 | 181.016 | 232.9562 | 161.3881
2012-10-30 18:13 | 28.5004 | 249.4741 | 195.3564 | 247.2653 | 175.6037 | -1.6032 | 1.2709 | 32.1369 | 40.8437 | 25.2861 | 2.0052 | 32.0618 | 64.2907 | 15.9892 | 14.3608 | -0.1105 | 105.1344 | 235.1133 | 180.9956 | 232.9045 | 161.2429
2012-10-30 18:13 | 28.4982 | 249.4758 | 195.3464 | 247.412 | 175.6561 | -1.5974 | 1.2709 | 32.1363 | 40.8429 | 25.2858 | 2.0052 | 32.0613 | 64.2884 | 15.9893 | 14.3586 | -0.1101 | 105.1313 | 235.1172 | 180.9878 | 233.0534 | 161.2975
2012-10-30 18:13 | 28.4959 | 249.5322 | 195.3676 | 247.2608 | 175.6226 | -1.598 | 1.2709 | 32.136 | 40.8414 | 25.2862 | 2.0052 | 32.061 | 64.288 | 15.9891 | 14.3767 | -0.1102 | 105.1294 | 235.1555 | 180.9909 | 232.8841 | 161.2459
2012-10-30 18:13 | 28.4967 | 249.4939 | 195.2584 | 247.3243 | 175.6078 | -1.6009 | 1.2709 | 32.136 | 40.8414 | 25.2861 | 2.0052 | 32.0609 | 64.2872 | 15.9893 | 14.3652 | -0.1104 | 105.1286 | 235.1287 | 180.8932 | 232.9591 | 161.2426
2012-10-30 18:13 | 28.4974 | 249.4398 | 195.3416 | 247.3524 | 175.7022 | -1.6032 | 1.2709 | 32.1359 | 40.8418 | 25.2858 | 2.0052 | 32.0608 | 64.287 | 15.9892 | 14.3578 | -0.1105 | 105.1288 | 235.082 | 180.9838 | 232.9946 | 161.3444
2012-10-30 18:13 | 28.493 | 249.4902 | 195.353 | 247.2579 | 175.5846 | -1.6014 | 1.2709 | 32.1353 | 40.8399 | 25.2859 | 2.0051 | 32.0602 | 64.2847 | 15.9891 | 14.3778 | -0.1104 | 105.1246 | 235.1124 | 180.9752 | 232.8801 | 161.2068
2012-10-30 18:14 | 28.4944 | 249.4759 | 195.2993 | 247.2789 | 175.465 | -1.6032 | 1.2709 | 32.1353 | 40.84 | 25.286 | 2.0051 | 32.0602 | 64.2852 | 15.989 | 14.3384 | -0.1105 | 105.1251 | 235.1375 | 180.9609 | 232.9405 | 161.1266
2012-10-30 18:14 | 28.5034 | 249.5084 | 195.2926 | 247.4211 | 175.4037 | -1.5968 | 1.2709 | 32.1352 | 40.8394 | 25.2861 | 2.0051 | 32.06 | 64.2838 | 15.9892 | 14.3638 | -0.1101 | 105.1232 | 235.1446 | 180.9288 | 233.0573 | 161.0399
2012-10-30 18:14 | 28.493 | 249.4471 | 195.3019 | 247.2188 | 175.5768 | -1.602 | 1.2708 | 32.1347 | 40.8381 | 25.2862 | 2.0051 | 32.0595 | 64.2823 | 15.989 | 14.3574 | -0.1104 | 105.1205 | 235.0897 | 180.9445 | 232.8614 | 161.2194
2012-10-30 18:14 | 28.5057 | 249.4676 | 195.2949 | 247.349 | 175.5816 | -1.6037 | 1.2708 | 32.1342 | 40.8369 | 25.2861 | 2.005 | 32.0589 | 64.2796 | 15.9892 | 14.3538 | -0.1106 | 105.1165 | 235.1138 | 180.9411 | 232.9952 | 161.2278
2012-10-30 18:14 | 28.493 | 249.4471 | 195.2587 | 247.3715 | 175.5533 | -1.6003 | 1.2708 | 32.134 | 40.8361 | 25.2863 | 2.005 | 32.0589 | 64.2791 | 15.9892 | 14.3411 | -0.1103 | 105.1152 | 235.106 | 180.9176 | 233.0304 | 161.2122
2012-10-30 18:14 | 28.4937 | 249.5339 | 195.2908 | 247.2782 | 175.5696 | -1.6061 | 1.2708 | 32.1336 | 40.8356 | 25.286 | 2.005 | 32.0586 | 64.2786 | 15.9891 | 14.3255 | -0.1107 | 105.1142 | 235.2084 | 180.9653 | 232.9527 | 161.2441
2012-10-30 18:14 | 28.4892 | 249.4668 | 195.2273 | 247.356 | 175.2646 | -1.6049 | 1.2695 | 32.1011 | 40.7528 | 25.2861 | 2.0035 | 32.0345 | 64.1816 | 15.9891 | 14.3496 | -0.1106 | 104.9345 | 235.1172 | 180.8777 | 233.0064 | 160.915
2012-10-30 18:14 | 28.5004 | 249.4976 | 195.2896 | 247.4064 | 175.4085 | -1.5985 | 1.2645 | 31.9716 | 40.4274 | 25.2845 | 1.9948 | 31.8956 | 63.6267 | 15.989 | 14.3608 | -0.1102 | 104.0541 | 235.1368 | 180.9288 | 233.0456 | 161.0477
2012-10-30 18:14 | 28.49 | 249.4558 | 195.236 | 247.2824 | 175.3903 | -1.5945 | 1.2644 | 31.97 | 40.4237 | 25.2843 | 1.9948 | 31.8958 | 63.6265 | 15.9892 | 14.3463 | -0.1099 | 104.0502 | 235.1095 | 180.8897 | 232.9361 | 161.044
2012-10-30 18:14 | 28.4944 | 249.425 | 195.1265 | 247.3181 | 175.2074 | -1.598 | 1.2638 | 31.9549 | 40.3858 | 25.284 | 1.9939 | 31.8805 | 63.5654 | 15.9893 | 14.3874 | -0.1102 | 103.9512 | 235.0376 | 180.7391 | 232.9307 | 160.82
2012-10-30 18:14 | 28.4907 | 249.3861 | 195.1581 | 247.4162 | 175.2583 | -1.6037 | 1.2636 | 31.9481 | 40.3685 | 25.2841 | 1.9934 | 31.8736 | 63.5375 | 15.9893 | 14.3551 | -0.1106 | 103.906 | 235.031 | 180.803 | 233.0611 | 160.9032
2012-10-30 18:14 | 28.4974 | 249.4163 | 195.1137 | 247.2623 | 175.5538 | -1.5991 | 1.2635 | 31.9457 | 40.3623 | 25.2842 | 1.9933 | 31.871 | 63.5275 | 15.9893 | 14.419 | -0.1102 | 103.8898 | 234.9973 | 180.6947 | 232.8433 | 161.1348
2012-10-30 18:14 | 28.49 | 249.4245 | 195.0749 | 247.3529 | 175.3942 | -1.6049 | 1.2635 | 31.9455 | 40.362 | 25.2841 | 1.9933 | 31.8709 | 63.5271 | 15.9893 | 14.3667 | -0.1106 | 103.8891 | 235.0578 | 180.7082 | 232.9862 | 161.0275
2012-10-30 18:14 | 28.4855 | 249.3456 | 195.149 | 247.3601 | 175.2843 | -1.6009 | 1.2635 | 31.9457 | 40.3625 | 25.284 | 1.9933 | 31.8711 | 63.5277 | 15.9894 | 14.3622 | -0.1104 | 103.8902 | 234.9834 | 180.7868 | 232.9979 | 160.9221
2012-10-30 18:14 | 28.4825 | 249.3426 | 195.1224 | 247.2043 | 175.0472 | -1.602 | 1.2635 | 31.9474 | 40.3666 | 25.2842 | 1.9933 | 31.8727 | 63.5333 | 15.9896 | 14.3265 | -0.1104 | 103.8999 | 235.0161 | 180.7959 | 232.8778 | 160.7207
2012-10-30 18:14 | 28.475 | 249.3508 | 195.0678 | 247.3614 | 174.6456 | -1.6003 | 1.2636 | 31.9497 | 40.3722 | 25.2843 | 1.9935 | 31.8752 | 63.5434 | 15.9895 | 14.3027 | -0.1103 | 103.9156 | 235.0481 | 180.7651 | 233.0587 | 160.3429
2012-10-30 18:14 | 28.4825 | 249.37 | 195.0753 | 247.2709 | 175.0628 | -1.6014 | 1.2637 | 31.9509 | 40.3755 | 25.2842 | 1.9936 | 31.8762 | 63.5475 | 15.9895 | 14.3184 | -0.1104 | 103.923 | 235.0516 | 180.7569 | 232.9525 | 160.7444
2012-10-30 18:14 | 28.4765 | 249.3523 | 194.9868 | 247.1983 | 175.2363 | -1.5962 | 1.2676 | 32.0526 | 40.6313 | 25.2851 | 1.9998 | 31.9759 | 63.9458 | 15.9894 | 14.3165 | -0.11 | 104.5771 | 235.0358 | 180.6703 | 232.8818 | 160.9198
2012-10-30 18:14 | 28.4825 | 249.3231 | 195.0635 | 247.3219 | 175.3126 | -1.5991 | 1.2689 | 32.0863 | 40.7158 | 25.2857 | 2.002 | 32.0109 | 64.0858 | 15.9895 | 14.3347 | -0.1102 | 104.8016 | 234.9884 | 180.7288 | 232.9872 | 160.9779
2012-10-30 18:14 | 28.4795 | 249.371 | 195.0841 | 247.1622 | 175.6452 | -1.602 | 1.2694 | 32.0991 | 40.7481 | 25.2859 | 2.0028 | 32.0241 | 64.1391 | 15.9893 | 14.3358 | -0.1104 | 104.8872 | 235.0352 | 180.7483 | 232.8264 | 161.3094
2012-10-30 18:14 | 28.478 | 249.2912 | 195.0786 | 247.3056 | 175.5188 | -1.5957 | 1.2699 | 32.1107 | 40.7773 | 25.286 | 2.0036 | 32.0359 | 64.1867 | 15.9893 | 14.322 | -0.11 | 104.964 | 234.9692 | 180.7566 | 232.9836 | 161.1968
2012-10-30 18:14 | 28.4712 | 249.2883 | 195.2447 | 247.2244 | 175.6408 | -1.598 | 1.2702 | 32.1197 | 40.7997 | 25.2863 | 2.0041 | 32.0446 | 64.2214 | 15.9893 | 14.3152 | -0.1102 | 105.0211 | 234.9731 | 180.9295 | 232.9092 | 161.3256
2012-10-30 18:14 | 28.4742 | 249.3069 | 195.1377 | 247.2039 | 175.4799 | -1.6066 | 1.2705 | 32.1259 | 40.8159 | 25.2861 | 2.0045 | 32.0508 | 64.2463 | 15.9893 | 14.2978 | -0.1108 | 105.0622 | 235.0091 | 180.8399 | 232.9061 | 161.1821
2012-10-30 18:14 | 28.4705 | 249.2915 | 195.1851 | 247.2785 | 175.3864 | -1.5997 | 1.2707 | 32.1306 | 40.8276 | 25.2862 | 2.0048 | 32.0554 | 64.2647 | 15.9893 | 14.3145 | -0.1103 | 105.0922 | 234.977 | 180.8706 | 232.964 | 161.0719
2012-10-30 18:14 | 28.4652 | 249.3841 | 195.1758 | 247.0617 | 175.4865 | -1.6009 | 1.2708 | 32.1338 | 40.8359 | 25.2861 | 2.005 | 32.0586 | 64.2782 | 15.9892 | 14.2684 | -0.1104 | 105.1141 | 235.1157 | 180.9074 | 232.7933 | 161.2181
2012-10-30 18:14 | 28.4675 | 249.3707 | 195.1349 | 247.2364 | 175.2898 | -1.5991 | 1.2709 | 32.1357 | 40.8409 | 25.286 | 2.0052 | 32.0606 | 64.2865 | 15.9891 | 14.2993 | -0.1102 | 105.1274 | 235.0714 | 180.8356 | 232.9371 | 160.9905
2012-10-30 18:14 | 28.472 | 249.3478 | 195.1984 | 247.1821 | 175.5245 | -1.6014 | 1.2709 | 32.1372 | 40.8441 | 25.2864 | 2.0052 | 32.062 | 64.2917 | 15.9892 | 14.3038 | -0.1104 | 105.1358 | 235.044 | 180.8946 | 232.8783 | 161.2207
2012-10-30 18:14 | 28.4622 | 249.2989 | 195.1414 | 247.2115 | 175.4289 | -1.6003 | 1.271 | 32.1383 | 40.8471 | 25.2862 | 2.0053 | 32.0632 | 64.2965 | 15.9891 | 14.2858 | -0.1103 | 105.1435 | 235.0131 | 180.8556 | 232.9257 | 161.1431
2012-10-30 18:15 | 28.469 | 249.3448 | 195.1836 | 247.2418 | 175.4474 | -1.598 | 1.271 | 32.1391 | 40.849 | 25.2864 | 2.0054 | 32.064 | 64.2998 | 15.9891 | 14.2926 | -0.1102 | 105.1488 | 235.0522 | 180.891 | 232.9492 | 161.1548
2012-10-30 18:15 | 28.466 | 249.4084 | 195.2277 | 247.1134 | 175.6825 | -1.602 | 1.271 | 32.1397 | 40.8503 | 25.2865 | 2.0054 | 32.0644 | 64.302 | 15.989 | 14.3304 | -0.1104 | 105.1523 | 235.078 | 180.8973 | 232.783 | 161.3521
2012-10-30 18:15 | 28.466 | 249.3379 | 195.1963 | 247.2231 | 175.4444 | -1.6003 | 1.271 | 32.1401 | 40.8514 | 25.2864 | 2.0054 | 32.0648 | 64.3034 | 15.9891 | 14.359 | -0.1103 | 105.1548 | 234.9789 | 180.8373 | 232.8641 | 161.0854
2012-10-30 18:15 | 28.4682 | 249.3753 | 195.2417 | 247.241 | 175.548 | -1.5957 | 1.271 | 32.1397 | 40.8504 | 25.2864 | 2.0054 | 32.0647 | 64.3026 | 15.9891 | 14.3408 | -0.11 | 105.1531 | 235.0345 | 180.9009 | 232.9002 | 161.2072
2012-10-30 18:15 | 28.4712 | 249.3666 | 195.1858 | 247.1147 | 175.4964 | -1.5985 | 1.2711 | 32.14 | 40.8516 | 25.2862 | 2.0054 | 32.0649 | 64.3039 | 15.989 | 14.3316 | -0.1102 | 105.1556 | 235.035 | 180.8542 | 232.7831 | 161.1648
2012-10-30 18:15 | 28.469 | 249.3644 | 195.1796 | 247.1791 | 175.3888 | -1.6037 | 1.2711 | 32.1403 | 40.8523 | 25.2862 | 2.0055 | 32.0652 | 64.3058 | 15.9889 | 14.2967 | -0.1106 | 105.1582 | 235.0677 | 180.8829 | 232.8824 | 161.0921
2012-10-30 18:15 | 28.463 | 249.3427 | 195.1658 | 247.2554 | 175.2814 | -1.6014 | 1.2711 | 32.1406 | 40.8529 | 25.2863 | 2.0055 | 32.0655 | 64.3068 | 15.9889 | 14.2948 | -0.1104 | 105.1597 | 235.0479 | 180.871 | 232.9606 | 160.9866
2012-10-30 18:15 | 28.4705 | 249.3659 | 195.1733 | 247.1375 | 175.4294 | -1.5957 | 1.2711 | 32.1409 | 40.8534 | 25.2865 | 2.0055 | 32.0657 | 64.3074 | 15.989 | 14.2901 | -0.11 | 105.1608 | 235.0758 | 180.8832 | 232.8474 | 161.1393
2012-10-30 18:15 | 28.4645 | 249.3403 | 195.1319 | 247.1746 | 175.3726 | -1.5991 | 1.2711 | 32.1411 | 40.8544 | 25.2862 | 2.0055 | 32.066 | 64.309 | 15.9889 | 14.3167 | -0.1102 | 105.1633 | 235.0236 | 180.8152 | 232.8579 | 161.0559
2012-10-30 18:15 | 28.469 | 249.3487 | 195.1443 | 247.3319 | 175.4786 | -1.6032 | 1.2711 | 32.1414 | 40.8549 | 25.2863 | 2.0055 | 32.0662 | 64.3097 | 15.9889 | 14.3212 | -0.1105 | 105.1646 | 235.0275 | 180.8231 | 233.0107 | 161.1574
2012-10-30 18:15 | 28.478 | 249.4008 | 195.1494 | 247.1411 | 175.4993 | -1.5997 | 1.2711 | 32.1415 | 40.8552 | 25.2864 | 2.0055 | 32.0664 | 64.3105 | 15.9889 | 14.2976 | -0.1103 | 105.1656 | 235.1032 | 180.8518 | 232.8435 | 161.2017
2012-10-30 18:15 | 28.4742 | 249.397 | 195.1456 | 247.2195 | 175.3433 | -1.6055 | 1.2711 | 32.1418 | 40.8558 | 25.2864 | 2.0056 | 32.0666 | 64.3119 | 15.9888 | 14.2611 | -0.1107 | 105.1677 | 235.1359 | 180.8845 | 232.9584 | 161.0822
2012-10-30 18:15 | 28.475 | 249.3391 | 195.0953 | 247.2439 | 175.3558 | -1.5951 | 1.2711 | 32.1416 | 40.8553 | 25.2863 | 2.0056 | 32.0664 | 64.311 | 15.9888 | 14.3109 | -0.11 | 105.1663 | 235.0282 | 180.7844 | 232.933 | 161.0449
2012-10-30 18:15 | 28.4757 | 249.4181 | 195.1274 | 247.0918 | 175.4541 | -1.5997 | 1.2711 | 32.1415 | 40.8551 | 25.2864 | 2.0056 | 32.0663 | 64.3114 | 15.9885 | 14.3279 | -0.1103 | 105.1665 | 235.0902 | 180.7995 | 232.7639 | 161.1262
2012-10-30 18:15 | 28.478 | 249.3577 | 195.0904 | 247.2508 | 175.3042 | -1.6014 | 1.2711 | 32.1415 | 40.8553 | 25.2862 | 2.0056 | 32.0664 | 64.3113 | 15.9887 | 14.2772 | -0.1104 | 105.1667 | 235.0805 | 180.8132 | 232.9736 | 161.027
2012-10-30 18:15 | 28.481 | 249.3529 | 195.1524 | 247.2851 | 175.3579 | -1.6037 | 1.2711 | 32.1415 | 40.8548 | 25.2865 | 2.0055 | 32.0663 | 64.3106 | 15.9888 | 14.2842 | -0.1106 | 105.1655 | 235.0687 | 180.8682 | 233.0009 | 161.0737
2012-10-30 18:15 | 28.4855 | 249.4161 | 195.1293 | 247.1447 | 175.5146 | -1.6014 | 1.2711 | 32.1417 | 40.8558 | 25.2862 | 2.0056 | 32.0666 | 64.312 | 15.9887 | 14.2806 | -0.1104 | 105.1678 | 235.1355 | 180.8487 | 232.8641 | 161.234
2012-10-30 18:15 | 28.4817 | 249.3536 | 195.0588 | 247.1957 | 175.3937 | -1.6032 | 1.2711 | 32.1412 | 40.8546 | 25.2862 | 2.0056 | 32.066 | 64.3104 | 15.9886 | 14.2768 | -0.1105 | 105.165 | 235.0768 | 180.782 | 232.9189 | 161.1169
2012-10-30 18:15 | 28.49 | 249.3971 | 195.1653 | 247.2784 | 175.441 | -1.602 | 1.2711 | 32.1408 | 40.8537 | 25.2862 | 2.0055 | 32.0658 | 64.309 | 15.9887 | 14.3177 | -0.1104 | 105.1627 | 235.0794 | 180.8476 | 232.9607 | 161.1233
2012-10-30 18:15 | 28.4892 | 249.4394 | 195.1055 | 247.1679 | 175.4598 | -1.6043 | 1.2711 | 32.1411 | 40.8541 | 25.2863 | 2.0055 | 32.066 | 64.3099 | 15.9886 | 14.3373 | -0.1106 | 105.1641 | 235.1021 | 180.7682 | 232.8306 | 161.1225
2012-10-30 18:15 | 28.481 | 249.3803 | 195.1406 | 247.2381 | 175.4867 | -1.6026 | 1.2711 | 32.1414 | 40.8546 | 25.2865 | 2.0056 | 32.0662 | 64.3111 | 15.9886 | 14.2924 | -0.1105 | 105.1657 | 235.0879 | 180.8482 | 232.9457 | 161.1943
2012-10-30 18:15 | 28.49 | 249.3854 | 195.126 | 247.3411 | 175.4957 | -1.602 | 1.2711 | 32.141 | 40.8536 | 25.2865 | 2.0056 | 32.0658 | 64.3098 | 15.9885 | 14.3463 | -0.1104 | 105.1634 | 235.0391 | 180.7797 | 232.9948 | 161.1494
2012-10-30 18:15 | 28.4915 | 249.4299 | 195.1943 | 247.2251 | 175.5753 | -1.5985 | 1.2711 | 32.1407 | 40.8531 | 25.2864 | 2.0055 | 32.0655 | 64.3079 | 15.9886 | 14.3274 | -0.1102 | 105.161 | 235.1025 | 180.8669 | 232.8977 | 161.2479
2012-10-30 18:15 | 28.4922 | 249.4346 | 195.1793 | 247.1984 | 175.4745 | -1.5997 | 1.2711 | 32.1405 | 40.8526 | 25.2863 | 2.0055 | 32.0653 | 64.3078 | 15.9885 | 14.3362 | -0.1103 | 105.1604 | 235.0984 | 180.8431 | 232.8622 | 161.1383
2012-10-30 18:15 | 28.4832 | 249.3786 | 195.1821 | 247.3461 | 175.4928 | -1.6055 | 1.271 | 32.1402 | 40.8516 | 25.2865 | 2.0055 | 32.0651 | 64.3067 | 15.9886 | 14.2905 | -0.1107 | 105.1582 | 235.0881 | 180.8916 | 233.0556 | 161.2023
2012-10-30 18:16 | 28.484 | 249.4498 | 195.23 | 247.1627 | 175.5131 | -1.6026 | 1.271 | 32.1398 | 40.8508 | 25.2863 | 2.0055 | 32.0647 | 64.3053 | 15.9885 | 14.324 | -0.1105 | 105.1561 | 235.1258 | 180.906 | 232.8387 | 161.1891
2012-10-30 18:16 | 28.4937 | 249.3852 | 195.1965 | 247.2273 | 175.3394 | -1.5957 | 1.271 | 32.1394 | 40.8497 | 25.2864 | 2.0055 | 32.0642 | 64.304 | 15.9884 | 14.3337 | -0.11 | 105.1536 | 235.0515 | 180.8628 | 232.8936 | 161.0057
2012-10-30 18:16 | 28.4735 | 249.3689 | 195.1802 | 247.352 | 175.3465 | -1.5968 | 1.271 | 32.1391 | 40.8493 | 25.2861 | 2.0054 | 32.0638 | 64.3022 | 15.9884 | 14.3094 | -0.1101 | 105.1515 | 235.0595 | 180.8708 | 233.0426 | 161.0371
2012-10-30 18:16 | 28.4825 | 249.4405 | 195.1892 | 247.1456 | 175.3789 | -1.6061 | 1.271 | 32.1387 | 40.848 | 25.2863 | 2.0054 | 32.0635 | 64.3012 | 15.9884 | 14.2653 | -0.1107 | 105.1492 | 235.1752 | 180.9239 | 232.8803 | 161.1136
2012-10-30 18:16 | 28.487 | 249.3863 | 195.1466 | 247.3224 | 175.2624 | -1.5985 | 1.271 | 32.1384 | 40.8471 | 25.2864 | 2.0054 | 32.0632 | 64.2995 | 15.9885 | 14.3066 | -0.1102 | 105.1466 | 235.0797 | 180.84 | 233.0158 | 160.9558
2012-10-30 18:16 | 28.487 | 249.4176 | 195.1308 | 247.2559 | 175.2312 | -1.6003 | 1.271 | 32.1378 | 40.8457 | 25.2863 | 2.0054 | 32.0626 | 64.2974 | 15.9884 | 14.3147 | -0.1103 | 105.1431 | 235.1029 | 180.8161 | 232.9412 | 160.9165
2012-10-30 18:16 | 28.4892 | 249.4511 | 195.133 | 247.1915 | 175.3778 | -1.6003 | 1.2709 | 32.1375 | 40.8447 | 25.2865 | 2.0053 | 32.0624 | 64.2961 | 15.9885 | 14.321 | -0.1103 | 105.1408 | 235.1301 | 180.812 | 232.8705 | 161.0568
2012-10-30 18:16 | 28.4847 | 249.3527 | 195.0814 | 247.2966 | 175.3499 | -1.6049 | 1.2709 | 32.1376 | 40.8449 | 25.2865 | 2.0054 | 32.0624 | 64.2966 | 15.9883 | 14.2879 | -0.1106 | 105.1415 | 235.0648 | 180.7935 | 233.0087 | 161.062
2012-10-30 18:16 | 28.4862 | 249.4442 | 195.1615 | 247.306 | 175.3397 | -1.6037 | 1.271 | 32.1375 | 40.8453 | 25.2861 | 2.0054 | 32.0623 | 64.2965 | 15.9883 | 14.2976 | -0.1106 | 105.1417 | 235.1466 | 180.8639 | 233.0084 | 161.0421
2012-10-30 18:16 | 28.4922 | 249.4267 | 195.1282 | 247.265 | 175.3574 | -1.5957 | 1.2709 | 32.1372 | 40.8442 | 25.2863 | 2.0053 | 32.0621 | 64.2954 | 15.9883 | 14.3077 | -0.11 | 105.1396 | 235.119 | 180.8205 | 232.9573 | 161.0497
2012-10-30 18:16 | 28.4862 | 249.4794 | 195.13 | 247.2002 | 175.3436 | -1.5997 | 1.2709 | 32.1371 | 40.8436 | 25.2866 | 2.0054 | 32.062 | 64.2956 | 15.9882 | 14.3466 | -0.1103 | 105.1391 | 235.1328 | 180.7834 | 232.8536 | 160.997
2012-10-30 18:16 | 28.4907 | 249.4644 | 195.0953 | 247.2478 | 175.1764 | -1.6043 | 1.2709 | 32.1366 | 40.8428 | 25.2863 | 2.0053 | 32.0614 | 64.2928 | 15.9883 | 14.3511 | -0.1106 | 105.1356 | 235.1133 | 180.7442 | 232.8967 | 160.8253
2012-10-30 18:16 | 28.4832 | 249.4334 | 195.0681 | 247.3304 | 175.2391 | -1.6009 | 1.2709 | 32.1362 | 40.8413 | 25.2865 | 2.0053 | 32.061 | 64.2912 | 15.9883 | 14.3272 | -0.1104 | 105.1325 | 235.1062 | 180.7409 | 233.0032 | 160.9119
2012-10-30 18:16 | 28.4802 | 249.45 | 195.0965 | 247.3156 | 175.4195 | -1.598 | 1.2709 | 32.1359 | 40.8404 | 25.2866 | 2.0052 | 32.0607 | 64.2896 | 15.9884 | 14.3365 | -0.1102 | 105.13 | 235.1135 | 180.76 | 232.9791 | 161.083
2012-10-30 18:16 | 28.4847 | 249.4584 | 195.0853 | 247.2771 | 175.5294 | -1.6009 | 1.2709 | 32.1353 | 40.8395 | 25.2863 | 2.0052 | 32.0602 | 64.2882 | 15.9882 | 14.3655 | -0.1104 | 105.1277 | 235.0929 | 180.7198 | 232.9116 | 161.1639
2012-10-30 18:16 | 28.487 | 249.5037 | 195.1348 | 247.2519 | 175.4927 | -1.5997 | 1.2708 | 32.135 | 40.8386 | 25.2864 | 2.0052 | 32.0599 | 64.2865 | 15.9883 | 14.3229 | -0.1103 | 105.1251 | 235.1808 | 180.8119 | 232.929 | 161.1698
2012-10-30 18:16 | 28.487 | 249.4724 | 195.0641 | 247.3107 | 175.4302 | -1.5974 | 1.2708 | 32.1347 | 40.8372 | 25.2868 | 2.0052 | 32.0595 | 64.286 | 15.9881 | 14.2984 | -0.1101 | 105.1232 | 235.174 | 180.7657 | 233.0123 | 161.1318
2012-10-30 18:16 | 28.4832 | 249.4647 | 195.1113 | 247.3422 | 175.3679 | -1.5985 | 1.2708 | 32.1345 | 40.8367 | 25.2867 | 2.0052 | 32.0594 | 64.2849 | 15.9883 | 14.3068 | -0.1102 | 105.1216 | 235.1579 | 180.8045 | 233.0354 | 161.0611
2012-10-30 18:16 | 28.4765 | 249.4423 | 195.0811 | 247.3237 | 175.2714 | -1.5985 | 1.2708 | 32.1343 | 40.8365 | 25.2866 | 2.0052 | 32.0592 | 64.2844 | 15.9882 | 14.3205 | -0.1102 | 105.1209 | 235.1218 | 180.7606 | 233.0032 | 160.9509
2012-10-30 18:16 | 28.4795 | 249.4493 | 195.1273 | 247.2797 | 175.4266 | -1.6003 | 1.2708 | 32.1343 | 40.8361 | 25.2868 | 2.0052 | 32.0592 | 64.2841 | 15.9883 | 14.2827 | -0.1103 | 105.1202 | 235.1666 | 180.8446 | 232.997 | 161.1439
2012-10-30 18:16 | 28.4855 | 249.4787 | 195.1097 | 247.2583 | 175.4639 | -1.5968 | 1.2708 | 32.1339 | 40.8353 | 25.2867 | 2.0052 | 32.0588 | 64.2833 | 15.9881 | 14.3255 | -0.1101 | 105.1186 | 235.1532 | 180.7842 | 232.9328 | 161.1384
2012-10-30 18:16 | 28.4832 | 249.4647 | 195.1153 | 247.2912 | 175.3016 | -1.5968 | 1.265 | 31.985 | 40.4601 | 25.2852 | 1.9958 | 31.9089 | 63.6832 | 15.9882 | 14.2946 | -0.1101 | 104.1433 | 235.1701 | 180.8207 | 232.9966 | 161.007
2012-10-30 18:16 | 28.4825 | 249.4327 | 195.032 | 247.314 | 175.3789 | -1.5962 | 1.2644 | 31.9689 | 40.4199 | 25.2848 | 1.9948 | 31.8941 | 63.6236 | 15.9883 | 14.2817 | -0.11 | 104.0434 | 235.151 | 180.7503 | 233.0323 | 161.0972
2012-10-30 18:16 | 28.4772 | 249.4039 | 195.0228 | 247.3166 | 175.3814 | -1.6037 | 1.2637 | 31.9514 | 40.3758 | 25.2847 | 1.9937 | 31.8764 | 63.5528 | 15.9884 | 14.2845 | -0.1106 | 103.9286 | 235.1194 | 180.7383 | 233.0321 | 161.0969
2012-10-30 18:16 | 28.4787 | 249.3976 | 195.0479 | 247.3259 | 175.6405 | -1.6026 | 1.2635 | 31.9473 | 40.3655 | 25.2846 | 1.9935 | 31.8725 | 63.5375 | 15.9883 | 14.2697 | -0.1105 | 103.903 | 235.1279 | 180.7782 | 233.0562 | 161.3708
2012-10-30 18:17 | 28.4772 | 249.3647 | 195.0621 | 247.3087 | 175.678 | -1.5997 | 1.2634 | 31.9449 | 40.36 | 25.2844 | 1.9933 | 31.8703 | 63.5277 | 15.9885 | 14.256 | -0.1103 | 103.8877 | 235.1087 | 180.8061 | 233.0527 | 161.422
2012-10-30 18:17 | 28.4757 | 249.3515 | 195.1117 | 247.1897 | 175.4385 | -1.5951 | 1.2634 | 31.9438 | 40.3567 | 25.2846 | 1.9932 | 31.869 | 63.5228 | 15.9885 | 14.3279 | -0.11 | 103.8796 | 235.0236 | 180.7838 | 232.8618 | 161.1106
2012-10-30 18:17 | 28.4772 | 249.3647 | 195.07 | 247.2421 | 175.3619 | -1.5974 | 1.2635 | 31.9462 | 40.3624 | 25.2849 | 1.9934 | 31.8716 | 63.5332 | 15.9885 | 14.26 | -0.1101 | 103.8956 | 235.1047 | 180.81 | 232.9821 | 161.1019
2012-10-30 18:17 | 28.4765 | 249.3445 | 195.0378 | 247.2806 | 175.4432 | -1.5974 | 1.2636 | 31.9499 | 40.3722 | 25.2846 | 1.9936 | 31.8752 | 63.5475 | 15.9885 | 14.2961 | -0.1101 | 103.9197 | 235.0484 | 180.7417 | 232.9845 | 161.1471
2012-10-30 18:17 | 28.478 | 249.3029 | 195.0786 | 247.3291 | 175.433 | -1.6032 | 1.2637 | 31.9527 | 40.3791 | 25.2847 | 1.9938 | 31.8779 | 63.5576 | 15.9887 | 14.269 | -0.1105 | 103.9367 | 235.0339 | 180.8096 | 233.0601 | 161.164
2012-10-30 18:17 | 28.4727 | 249.3446 | 195.1558 | 247.2376 | 175.1545 | -1.6003 | 1.2671 | 32.0389 | 40.5963 | 25.2853 | 1.9993 | 31.9658 | 63.9087 | 15.9886 | 14.2719 | -0.1103 | 104.505 | 235.0727 | 180.8839 | 232.9657 | 160.8826
2012-10-30 18:17 | 28.4637 | 249.3278 | 195.0683 | 247.209 | 175.0245 | -1.598 | 1.269 | 32.0889 | 40.7218 | 25.2861 | 2.0023 | 32.0141 | 64.1027 | 15.9884 | 14.2465 | -0.1102 | 104.8245 | 235.0813 | 180.8218 | 232.9625 | 160.778
2012-10-30 18:17 | 28.466 | 249.3261 | 195.0823 | 247.1918 | 175.3117 | -1.5974 | 1.2697 | 32.1057 | 40.7645 | 25.2862 | 2.0034 | 32.0308 | 64.17 | 15.9884 | 14.2815 | -0.1101 | 104.9344 | 235.0446 | 180.8008 | 232.9103 | 161.0302
2012-10-30 18:17 | 28.4697 | 249.3416 | 195.09 | 247.1328 | 175.3778 | -1.598 | 1.2702 | 32.1182 | 40.796 | 25.2863 | 2.0042 | 32.0432 | 64.2196 | 15.9884 | 14.2566 | -0.1102 | 105.0156 | 235.085 | 180.8334 | 232.8762 | 161.1212
2012-10-30 18:17 | 28.4645 | 249.3442 | 195.1005 | 247.1863 | 175.5561 | -1.6009 | 1.2705 | 32.1263 | 40.816 | 25.2867 | 2.0046 | 32.0511 | 64.2512 | 15.9884 | 14.2881 | -0.1104 | 105.0672 | 235.0561 | 180.8124 | 232.8982 | 161.268
2012-10-30 18:17 | 28.4667 | 249.3229 | 195.1027 | 247.2042 | 175.6988 | -1.6009 | 1.2707 | 32.1314 | 40.829 | 25.2866 | 2.005 | 32.0562 | 64.2714 | 15.9884 | 14.274 | -0.1104 | 105.1004 | 235.0489 | 180.8287 | 232.9302 | 161.4248
2012-10-30 18:17 | 28.4615 | 249.3725 | 195.1446 | 247.2147 | 175.6741 | -1.6026 | 1.2708 | 32.1346 | 40.8371 | 25.2866 | 2.0052 | 32.0593 | 64.284 | 15.9884 | 14.2403 | -0.1105 | 105.1211 | 235.1322 | 180.9043 | 232.9744 | 161.4338
2012-10-30 18:17 | 28.463 | 249.3505 | 195.1894 | 247.1496 | 175.5702 | -1.6049 | 1.2709 | 32.1364 | 40.8417 | 25.2866 | 2.0053 | 32.0613 | 64.2924 | 15.9883 | 14.2785 | -0.1106 | 105.1341 | 235.072 | 180.9109 | 232.8711 | 161.2917
2012-10-30 18:17 | 28.4637 | 249.3708 | 195.1704 | 247.115 | 175.5201 | -1.5974 | 1.2709 | 32.138 | 40.8456 | 25.2867 | 2.0054 | 32.0627 | 64.2983 | 15.9882 | 14.2792 | -0.1101 | 105.1438 | 235.0916 | 180.8912 | 232.8358 | 161.2409
2012-10-30 18:17 | 28.4652 | 249.3802 | 195.1562 | 247.1988 | 175.5021 | -1.5985 | 1.271 | 32.139 | 40.8485 | 25.2866 | 2.0054 | 32.0637 | 64.302 | 15.9883 | 14.2807 | -0.1102 | 105.1504 | 235.0995 | 180.8755 | 232.9181 | 161.2214
2012-10-30 18:17 | 28.466 | 249.3653 | 195.2317 | 247.2349 | 175.5771 | -1.6066 | 1.271 | 32.1399 | 40.8504 | 25.2868 | 2.0055 | 32.0646 | 64.3062 | 15.9882 | 14.2856 | -0.1108 | 105.1566 | 235.0797 | 180.9461 | 232.9493 | 161.2915
2012-10-30 18:17 | 28.4727 | 249.3681 | 195.1951 | 247.2024 | 175.7126 | -1.6049 | 1.271 | 32.1402 | 40.8509 | 25.2868 | 2.0056 | 32.065 | 64.308 | 15.9881 | 14.2596 | -0.1106 | 105.1589 | 235.1085 | 180.9355 | 232.9428 | 161.453
2012-10-30 18:17 | 28.466 | 249.3692 | 195.2513 | 247.1369 | 175.581 | -1.6026 | 1.271 | 32.1401 | 40.8508 | 25.2868 | 2.0055 | 32.0648 | 64.3075 | 15.988 | 14.2652 | -0.1105 | 105.1583 | 235.104 | 180.9861 | 232.8717 | 161.3158
2012-10-30 18:17 | 28.4742 | 249.3813 | 195.2163 | 247.1138 | 175.636 | -1.5985 | 1.271 | 32.1405 | 40.8514 | 25.287 | 2.0056 | 32.0652 | 64.3088 | 15.9881 | 14.2489 | -0.1102 | 105.1602 | 235.1324 | 180.9674 | 232.8649 | 161.3871
2012-10-30 18:17 | 28.472 | 249.3674 | 195.2494 | 247.2095 | 175.5792 | -1.6003 | 1.271 | 32.1405 | 40.8519 | 25.2867 | 2.0056 | 32.0652 | 64.3086 | 15.9882 | 14.2834 | -0.1103 | 105.1605 | 235.084 | 180.966 | 232.9261 | 161.2958
2012-10-30 18:17 | 28.4787 | 249.3819 | 195.2404 | 247.1849 | 175.6913 | -1.5991 | 1.271 | 32.1405 | 40.8518 | 25.2868 | 2.0056 | 32.0653 | 64.3093 | 15.9881 | 14.2738 | -0.1102 | 105.1611 | 235.1081 | 180.9666 | 232.9111 | 161.4175
2012-10-30 18:17 | 28.487 | 249.3354 | 195.2605 | 247.1853 | 175.5864 | -1.6043 | 1.271 | 32.1406 | 40.8516 | 25.287 | 2.0056 | 32.0653 | 64.3087 | 15.9882 | 14.3147 | -0.1106 | 105.1603 | 235.0207 | 180.9458 | 232.8706 | 161.2717
2012-10-30 18:17 | 28.478 | 249.3655 | 195.2437 | 247.2155 | 175.7296 | -1.6003 | 1.271 | 32.1406 | 40.8521 | 25.2868 | 2.0056 | 32.0654 | 64.31 | 15.988 | 14.2731 | -0.1103 | 105.1621 | 235.0924 | 180.9706 | 232.9424 | 161.4565
2012-10-30 18:18 | 28.4937 | 249.4478 | 195.2672 | 247.1803 | 175.7921 | -1.6014 | 1.271 | 32.1406 | 40.8517 | 25.287 | 2.0056 | 32.0653 | 64.309 | 15.9882 | 14.3133 | -0.1104 | 105.1608 | 235.1345 | 180.9539 | 232.867 | 161.4788
2012-10-30 18:18 | 28.4885 | 249.4387 | 195.2777 | 247.1829 | 175.662 | -1.5991 | 1.271 | 32.1409 | 40.8523 | 25.2871 | 2.0056 | 32.0656 | 64.3107 | 15.9881 | 14.255 | -0.1102 | 105.1629 | 235.1837 | 181.0227 | 232.9279 | 161.407
2012-10-30 18:18 | 28.4915 | 249.3908 | 195.2572 | 247.1859 | 175.6377 | -1.6032 | 1.2711 | 32.141 | 40.8532 | 25.2867 | 2.0056 | 32.0657 | 64.3109 | 15.9882 | 14.2825 | -0.1105 | 105.164 | 235.1083 | 180.9747 | 232.9034 | 161.3552
2012-10-30 18:18 | 28.4952 | 249.3984 | 195.2923 | 247.2053 | 175.6063 | -1.6037 | 1.2711 | 32.1411 | 40.8531 | 25.287 | 2.0056 | 32.0659 | 64.3119 | 15.988 | 14.2495 | -0.1106 | 105.165 | 235.1489 | 181.0428 | 232.9558 | 161.3568
2012-10-30 18:18 | 28.4952 | 249.3867 | 195.3159 | 247.2366 | 175.6609 | -1.5951 | 1.2711 | 32.1416 | 40.8551 | 25.2865 | 2.0057 | 32.0663 | 64.3138 | 15.988 | 14.2944 | -0.11 | 105.1689 | 235.0923 | 181.0215 | 232.9422 | 161.3665
2012-10-30 18:18 | 28.5019 | 249.4169 | 195.2833 | 247.2081 | 175.7223 | -1.6043 | 1.2711 | 32.1413 | 40.8534 | 25.287 | 2.0056 | 32.066 | 64.3124 | 15.988 | 14.2521 | -0.1106 | 105.1658 | 235.1648 | 181.0312 | 232.956 | 161.4702
2012-10-30 18:18 | 28.5117 | 249.4501 | 195.2891 | 247.1513 | 175.9507 | -1.5991 | 1.271 | 32.1408 | 40.8524 | 25.2869 | 2.0056 | 32.0655 | 64.3111 | 15.9879 | 14.2986 | -0.1102 | 105.1634 | 235.1515 | 180.9905 | 232.8527 | 161.6521
2012-10-30 18:18 | 28.5124 | 249.4274 | 195.3016 | 247.2264 | 175.7796 | -1.6009 | 1.271 | 32.1404 | 40.8514 | 25.2869 | 2.0056 | 32.0652 | 64.3098 | 15.9879 | 14.2952 | -0.1104 | 105.1612 | 235.1322 | 181.0064 | 232.9312 | 161.4844
2012-10-30 18:18 | 28.5057 | 249.4324 | 195.346 | 247.2471 | 175.6714 | -1.602 | 1.271 | 32.1405 | 40.8513 | 25.2871 | 2.0056 | 32.0652 | 64.3098 | 15.9879 | 14.3089 | -0.1104 | 105.1611 | 235.1235 | 181.0371 | 232.9382 | 161.3625
2012-10-30 18:18 | 28.5132 | 249.4282 | 195.326 | 247.286 | 175.7765 | -1.598 | 1.271 | 32.1403 | 40.8514 | 25.2868 | 2.0056 | 32.0651 | 64.3097 | 15.9878 | 14.2552 | -0.1102 | 105.1612 | 235.173 | 181.0708 | 233.0308 | 161.5213
2012-10-30 18:18 | 28.5117 | 249.4501 | 195.3442 | 247.2727 | 175.8297 | -1.6043 | 1.271 | 32.1402 | 40.8512 | 25.2867 | 2.0056 | 32.0649 | 64.3093 | 15.9878 | 14.2619 | -0.1106 | 105.1605 | 235.1882 | 181.0823 | 233.0108 | 161.5678
2012-10-30 18:18 | 28.5079 | 249.4463 | 195.3836 | 247.2415 | 175.9117 | -1.6003 | 1.271 | 32.1398 | 40.8499 | 25.287 | 2.0056 | 32.0646 | 64.3084 | 15.9877 | 14.3356 | -0.1103 | 105.1583 | 235.1107 | 181.048 | 232.9059 | 161.5761
2012-10-30 18:18 | 28.5027 | 249.4451 | 195.4059 | 247.2951 | 175.5006 | -1.6055 | 1.271 | 32.1398 | 40.8496 | 25.287 | 2.0056 | 32.0646 | 64.3084 | 15.9876 | 14.3059 | -0.1107 | 105.158 | 235.1392 | 181.1 | 232.9892 | 161.1947
2012-10-30 18:18 | 28.5034 | 249.5006 | 195.3476 | 247.2723 | 175.6262 | -1.598 | 1.271 | 32.1396 | 40.8494 | 25.2868 | 2.0056 | 32.0644 | 64.3072 | 15.9877 | 14.3189 | -0.1102 | 105.1566 | 235.1817 | 181.0287 | 232.9534 | 161.3073
2012-10-30 18:18 | 28.5079 | 249.4542 | 195.3718 | 247.3394 | 175.7829 | -1.6003 | 1.271 | 32.1394 | 40.8491 | 25.2868 | 2.0056 | 32.0643 | 64.3071 | 15.9877 | 14.303 | -0.1103 | 105.1562 | 235.1512 | 181.0688 | 233.0364 | 161.4799
2012-10-30 18:18 | 28.5049 | 249.4433 | 195.3845 | 247.4109 | 175.6628 | -1.5991 | 1.271 | 32.1387 | 40.8475 | 25.2867 | 2.0055 | 32.0634 | 64.3036 | 15.9876 | 14.3041 | -0.1102 | 105.151 | 235.1392 | 181.0804 | 233.1068 | 161.3587
2012-10-30 18:18 | 28.5019 | 249.503 | 195.4169 | 247.2629 | 175.6325 | -1.6009 | 1.2709 | 32.1383 | 40.846 | 25.287 | 2.0055 | 32.063 | 64.302 | 15.9876 | 14.3337 | -0.1104 | 105.148 | 235.1693 | 181.0832 | 232.9292 | 161.2988
2012-10-30 18:18 | 28.5064 | 249.4683 | 195.3349 | 247.2126 | 175.6448 | -1.6009 | 1.2709 | 32.1378 | 40.8454 | 25.2865 | 2.0055 | 32.0626 | 64.3004 | 15.9876 | 14.2974 | -0.1104 | 105.1458 | 235.1709 | 181.0375 | 232.9152 | 161.3474
2012-10-30 18:18 | 28.5079 | 249.4777 | 195.3168 | 247.3786 | 175.498 | -1.6009 | 1.2709 | 32.1377 | 40.845 | 25.2866 | 2.0055 | 32.0624 | 64.3001 | 15.9875 | 14.3111 | -0.1104 | 105.1451 | 235.1666 | 181.0057 | 233.0675 | 161.1869
2012-10-30 18:18 | 28.5117 | 249.5167 | 195.3638 | 247.3432 | 175.5798 | -1.6061 | 1.2709 | 32.1374 | 40.844 | 25.2867 | 2.0055 | 32.0622 | 64.2994 | 15.9875 | 14.3272 | -0.1107 | 105.1434 | 235.1895 | 181.0366 | 233.016 | 161.2526
2012-10-30 18:18 | 28.5139 | 249.5111 | 195.3424 | 247.2906 | 175.7265 | -1.5968 | 1.2709 | 32.1374 | 40.8439 | 25.2868 | 2.0055 | 32.0621 | 64.2993 | 15.9874 | 14.3212 | -0.1101 | 105.1433 | 235.1899 | 181.0212 | 232.9694 | 161.4053
2012-10-30 18:18 | 28.5169 | 249.5258 | 195.3415 | 247.2427 | 175.6475 | -1.5985 | 1.2709 | 32.1373 | 40.8435 | 25.2869 | 2.0054 | 32.062 | 64.2983 | 15.9875 | 14.3446 | -0.1102 | 105.1419 | 235.1812 | 180.9969 | 232.8981 | 161.3029
2012-10-30 18:18 | 28.5169 | 249.4867 | 195.3061 | 247.3484 | 175.5421 | -1.5974 | 1.2709 | 32.1371 | 40.8436 | 25.2865 | 2.0055 | 32.0619 | 64.2986 | 15.9874 | 14.3283 | -0.1101 | 105.1423 | 235.1584 | 180.9778 | 233.0201 | 161.2138
2012-10-30 18:18 | 28.5087 | 249.4745 | 195.2861 | 247.3246 | 175.4988 | -1.6009 | 1.2709 | 32.137 | 40.843 | 25.2868 | 2.0055 | 32.0618 | 64.2987 | 15.9872 | 14.2956 | -0.1104 | 105.1417 | 235.1789 | 180.9905 | 233.029 | 161.2032
2012-10-30 18:19 | 28.5169 | 249.4945 | 195.314 | 247.3171 | 175.4953 | -1.6032 | 1.2709 | 32.1363 | 40.8416 | 25.2865 | 2.0054 | 32.061 | 64.2956 | 15.9872 | 14.3161 | -0.1105 | 105.1372 | 235.1784 | 180.9979 | 233.001 | 161.1792
2012-10-30 18:19 | 28.5147 | 249.5275 | 195.2843 | 247.2757 | 175.6999 | -1.6014 | 1.2709 | 32.136 | 40.8405 | 25.2868 | 2.0054 | 32.0607 | 64.2948 | 15.9872 | 14.3139 | -0.1104 | 105.1353 | 235.2136 | 180.9704 | 232.9618 | 161.386
2012-10-30 18:19 | 28.5102 | 249.5074 | 195.2798 | 247.2791 | 175.473 | -1.6072 | 1.2709 | 32.1355 | 40.8394 | 25.2866 | 2.0054 | 32.0603 | 64.2927 | 15.9873 | 14.342 | -0.1108 | 105.132 | 235.1654 | 180.9378 | 232.9371 | 161.131
2012-10-30 18:19 | 28.5177 | 249.507 | 195.2912 | 247.3688 | 175.5039 | -1.6026 | 1.2708 | 32.135 | 40.8386 | 25.2863 | 2.0054 | 32.0598 | 64.2912 | 15.9871 | 14.325 | -0.1105 | 105.1298 | 235.182 | 180.9662 | 233.0438 | 161.1789
2012-10-30 18:19 | 28.5207 | 249.5374 | 195.2824 | 247.3405 | 175.4561 | -1.6003 | 1.2708 | 32.1351 | 40.8385 | 25.2866 | 2.0053 | 32.0598 | 64.2909 | 15.9872 | 14.3688 | -0.1103 | 105.1294 | 235.1686 | 180.9136 | 232.9717 | 161.0873
2012-10-30 18:19 | 28.5207 | 249.5805 | 195.2981 | 247.2621 | 175.5147 | -1.6003 | 1.2708 | 32.1348 | 40.8379 | 25.2864 | 2.0053 | 32.0597 | 64.2903 | 15.9873 | 14.3443 | -0.1103 | 105.1282 | 235.2362 | 180.9538 | 232.9178 | 161.1704
2012-10-30 18:19 | 28.5282 | 249.494 | 195.1917 | 247.3441 | 175.3544 | -1.598 | 1.2708 | 32.1341 | 40.8356 | 25.2868 | 2.0053 | 32.0589 | 64.287 | 15.9873 | 14.3233 | -0.1102 | 105.1227 | 235.1707 | 180.8684 | 233.0208 | 161.0311
2012-10-30 18:19 | 28.5229 | 249.4496 | 195.2414 | 247.3388 | 175.6145 | -1.6026 | 1.2708 | 32.1337 | 40.8343 | 25.2869 | 2.0053 | 32.0585 | 64.2858 | 15.9872 | 14.3384 | -0.1105 | 105.1201 | 235.1112 | 180.903 | 233.0004 | 161.2761
2012-10-30 18:19 | 28.5207 | 249.5531 | 195.2981 | 247.2778 | 175.6981 | -1.6026 | 1.2708 | 32.1335 | 40.8341 | 25.2867 | 2.0053 | 32.0582 | 64.285 | 15.9871 | 14.3321 | -0.1105 | 105.1191 | 235.221 | 180.966 | 232.9457 | 161.366
2012-10-30 18:19 | 28.5334 | 249.5384 | 195.3187 | 247.3101 | 175.5781 | -1.6037 | 1.2671 | 32.0407 | 40.5997 | 25.2861 | 1.9996 | 31.9678 | 63.9217 | 15.9874 | 14.3448 | -0.1106 | 104.5214 | 235.1936 | 180.9739 | 232.9653 | 161.2333
2012-10-30 18:19 | 28.5214 | 249.4912 | 195.2399 | 247.3608 | 175.3788 | -1.598 | 1.2647 | 31.9785 | 40.4436 | 25.2852 | 1.9956 | 31.9038 | 63.6669 | 15.9872 | 14.3246 | -0.1102 | 104.1105 | 235.1666 | 180.9153 | 233.0362 | 161.0542
2012-10-30 18:19 | 28.5274 | 249.458 | 195.2616 | 247.3628 | 175.5604 | -1.602 | 1.264 | 31.9614 | 40.4004 | 25.2851 | 1.9945 | 31.8869 | 63.5991 | 15.9873 | 14.3755 | -0.1104 | 103.9995 | 235.0825 | 180.8861 | 232.9873 | 161.1849
2012-10-30 18:19 | 28.5259 | 249.4409 | 195.2326 | 247.3378 | 175.516 | -1.5997 | 1.2636 | 31.9496 | 40.3708 | 25.2851 | 1.9938 | 31.875 | 63.5511 | 15.9873 | 14.3618 | -0.1103 | 103.9219 | 235.0791 | 180.8708 | 232.976 | 161.1542
2012-10-30 18:19 | 28.5237 | 249.3956 | 195.2304 | 247.2612 | 175.3108 | -1.6026 | 1.2634 | 31.9454 | 40.3614 | 25.2843 | 1.9935 | 31.8708 | 63.534 | 15.9874 | 14.3841 | -0.1105 | 103.8954 | 235.0115 | 180.8463 | 232.8771 | 160.9267
2012-10-30 18:19 | 28.5252 | 249.4362 | 195.1494 | 247.3097 | 175.1562 | -1.5962 | 1.2634 | 31.9458 | 40.3615 | 25.2848 | 1.9935 | 31.871 | 63.5351 | 15.9874 | 14.4019 | -0.11 | 103.8966 | 235.0343 | 180.7475 | 232.9078 | 160.7543
2012-10-30 18:19 | 28.5297 | 249.3703 | 195.1068 | 247.3651 | 175.0398 | -1.6026 | 1.2634 | 31.9457 | 40.3612 | 25.2849 | 1.9935 | 31.8709 | 63.5342 | 15.9876 | 14.3737 | -0.1105 | 103.8954 | 234.9966 | 180.7331 | 232.9914 | 160.6661
2012-10-30 18:19 | 28.5207 | 249.373 | 195.1331 | 247.2856 | 175.2454 | -1.6003 | 1.2635 | 31.9471 | 40.3649 | 25.2848 | 1.9936 | 31.8724 | 63.5398 | 15.9876 | 14.3729 | -0.1103 | 103.9047 | 235.0001 | 180.7602 | 232.9127 | 160.8725
2012-10-30 18:19 | 28.5229 | 249.3361 | 195.0842 | 247.2291 | 175.2398 | -1.5985 | 1.2636 | 31.9488 | 40.369 | 25.2849 | 1.9937 | 31.8741 | 63.5463 | 15.9876 | 14.3588 | -0.1102 | 103.9154 | 234.9773 | 180.7254 | 232.8703 | 160.881
2012-10-30 18:19 | 28.5267 | 249.3046 | 195.0998 | 247.3073 | 175.3333 | -1.5991 | 1.2654 | 31.9964 | 40.4889 | 25.2852 | 1.9966 | 31.9209 | 63.7331 | 15.9877 | 14.3626 | -0.1102 | 104.2221 | 234.942 | 180.7372 | 232.9447 | 160.9707
2012-10-30 18:19 | 28.5102 | 249.3038 | 195.099 | 247.3378 | 175.4144 | -1.5962 | 1.2686 | 32.0771 | 40.6919 | 25.286 | 2.0017 | 32.0023 | 64.0587 | 15.9876 | 14.3665 | -0.11 | 104.7507 | 234.9373 | 180.7325 | 232.9713 | 161.0479
2012-10-30 18:19 | 28.5139 | 249.331 | 195.1145 | 247.1848 | 175.5586 | -1.602 | 1.2692 | 32.0933 | 40.7329 | 25.2862 | 2.0027 | 32.0181 | 64.1223 | 15.9876 | 14.3824 | -0.1104 | 104.8552 | 234.9486 | 180.7321 | 232.8024 | 161.1762
2012-10-30 18:19 | 28.5132 | 249.299 | 195.1453 | 247.2076 | 175.6126 | -1.6026 | 1.2698 | 32.1093 | 40.7732 | 25.2863 | 2.0037 | 32.0339 | 64.1861 | 15.9874 | 14.3695 | -0.1105 | 104.9593 | 234.9295 | 180.7758 | 232.8381 | 161.2431
2012-10-30 18:19 | 28.5072 | 249.3086 | 195.1825 | 247.2878 | 175.5441 | -1.6043 | 1.2703 | 32.1205 | 40.8015 | 25.2865 | 2.0044 | 32.0453 | 64.2321 | 15.9874 | 14.3472 | -0.1106 | 105.0336 | 234.9614 | 180.8353 | 232.9406 | 161.1969
2012-10-30 18:20 | 28.5124 | 249.3412 | 195.223 | 247.2264 | 175.5962 | -1.6049 | 1.2705 | 32.1277 | 40.8197 | 25.2865 | 2.0049 | 32.0525 | 64.2611 | 15.9874 | 14.3646 | -0.1106 | 105.0808 | 234.9766 | 180.8584 | 232.8618 | 161.2316
2012-10-30 18:20 | 28.5139 | 249.378 | 195.2088 | 247.2005 | 175.6952 | -1.6026 | 1.2707 | 32.132 | 40.8303 | 25.2867 | 2.0051 | 32.0567 | 64.2782 | 15.9873 | 14.3865 | -0.1105 | 105.1084 | 234.9915 | 180.8223 | 232.814 | 161.3087
2012-10-30 18:20 | 28.5079 | 249.325 | 195.2303 | 247.1396 | 175.6151 | -1.6009 | 1.2708 | 32.1347 | 40.837 | 25.2869 | 2.0053 | 32.0595 | 64.2894 | 15.9873 | 14.3764 | -0.1104 | 105.1264 | 234.9486 | 180.8539 | 232.7632 | 161.2387
2012-10-30 18:20 | 28.5072 | 249.3282 | 195.2532 | 247.1977 | 175.6378 | -1.6014 | 1.2709 | 32.1364 | 40.8411 | 25.287 | 2.0054 | 32.0612 | 64.2965 | 15.9871 | 14.3594 | -0.1104 | 105.1376 | 234.9688 | 180.8938 | 232.8383 | 161.2784
2012-10-30 18:20 | 28.5079 | 249.3328 | 195.2814 | 247.2532 | 175.7439 | -1.602 | 1.2709 | 32.1373 | 40.844 | 25.2866 | 2.0055 | 32.0621 | 64.2998 | 15.9872 | 14.356 | -0.1104 | 105.1438 | 234.9768 | 180.9254 | 232.8972 | 161.3879
2012-10-30 18:20 | 28.5072 | 249.3791 | 195.3357 | 247.2212 | 175.7354 | -1.5968 | 1.2709 | 32.1379 | 40.8454 | 25.2867 | 2.0055 | 32.0626 | 64.3025 | 15.9871 | 14.3839 | -0.1101 | 105.1479 | 234.9952 | 180.9518 | 232.8373 | 161.3515
2012-10-30 18:20 | 28.5117 | 249.3992 | 195.3402 | 247.0925 | 175.5876 | -1.5997 | 1.271 | 32.139 | 40.8485 | 25.2865 | 2.0056 | 32.0637 | 64.3067 | 15.9871 | 14.3884 | -0.1103 | 105.1552 | 235.0108 | 180.9518 | 232.7041 | 161.1992
2012-10-30 18:20 | 28.5162 | 249.3489 | 195.3054 | 247.2106 | 175.4594 | -1.6003 | 1.271 | 32.1397 | 40.8496 | 25.2868 | 2.0056 | 32.0644 | 64.3097 | 15.9871 | 14.3766 | -0.1103 | 105.1593 | 234.9723 | 180.9288 | 232.834 | 161.0828
2012-10-30 18:20 | 28.5049 | 249.3455 | 195.2195 | 247.2973 | 175.4872 | -1.6003 | 1.271 | 32.1396 | 40.8499 | 25.2866 | 2.0056 | 32.0643 | 64.3096 | 15.987 | 14.4142 | -0.1103 | 105.1595 | 234.9313 | 180.8053 | 232.8831 | 161.073
2012-10-30 18:20 | 28.5034 | 249.3479 | 195.3162 | 247.1077 | 175.716 | -1.6043 | 1.271 | 32.1399 | 40.8504 | 25.2868 | 2.0056 | 32.0646 | 64.3101 | 15.9872 | 14.3311 | -0.1106 | 105.1605 | 235.0168 | 180.9851 | 232.7766 | 161.3849
2012-10-30 18:20 | 28.5064 | 249.3665 | 195.2838 | 247.146 | 175.4184 | -1.6014 | 1.271 | 32.14 | 40.8503 | 25.2869 | 2.0057 | 32.0647 | 64.3115 | 15.987 | 14.4157 | -0.1104 | 105.1618 | 234.9508 | 180.8681 | 232.7303 | 161.0027
2012-10-30 18:20 | 28.5042 | 249.3917 | 195.2581 | 247.277 | 175.3499 | -1.5968 | 1.271 | 32.1399 | 40.8505 | 25.2867 | 2.0057 | 32.0648 | 64.3111 | 15.9871 | 14.3197 | -0.1101 | 105.1616 | 235.072 | 180.9384 | 232.9573 | 161.0302
2012-10-30 18:20 | 28.5034 | 249.3283 | 195.2415 | 247.2331 | 175.6145 | -1.598 | 1.271 | 32.1401 | 40.8508 | 25.2868 | 2.0057 | 32.0648 | 64.3114 | 15.987 | 14.4209 | -0.1102 | 105.1622 | 234.9074 | 180.8206 | 232.8122 | 161.1936
2012-10-30 18:20 | 28.5087 | 249.3884 | 195.3058 | 247.0973 | 175.854 | -1.6026 | 1.271 | 32.1402 | 40.8511 | 25.2867 | 2.0057 | 32.065 | 64.3131 | 15.9869 | 14.3609 | -0.1105 | 105.1642 | 235.0275 | 180.9449 | 232.7364 | 161.4931
2012-10-30 18:20 | 28.4997 | 249.3598 | 195.2418 | 247.2686 | 175.6498 | -1.5991 | 1.271 | 32.1407 | 40.8519 | 25.287 | 2.0057 | 32.0654 | 64.3147 | 15.9868 | 14.3519 | -0.1102 | 105.1667 | 235.0079 | 180.8899 | 232.9167 | 161.2979
2012-10-30 18:20 | 28.5042 | 249.3917 | 195.3052 | 247.2613 | 175.748 | -1.5991 | 1.271 | 32.1404 | 40.8512 | 25.287 | 2.0057 | 32.0652 | 64.314 | 15.9868 | 14.3605 | -0.1102 | 105.1652 | 235.0312 | 180.9447 | 232.9008 | 161.3875
2012-10-30 18:20 | 28.5004 | 249.3331 | 195.3093 | 247.1752 | 175.7091 | -1.6003 | 1.271 | 32.1397 | 40.8498 | 25.2868 | 2.0057 | 32.0645 | 64.3111 | 15.9869 | 14.3812 | -0.1103 | 105.161 | 234.9519 | 180.9281 | 232.794 | 161.3279
2012-10-30 18:20 | 28.4907 | 249.4057 | 195.2839 | 247.1851 | 175.5549 | -1.6055 | 1.271 | 32.14 | 40.8504 | 25.2869 | 2.0057 | 32.0647 | 64.3123 | 15.9868 | 14.3429 | -0.1107 | 105.1627 | 235.0628 | 180.941 | 232.8422 | 161.212
2012-10-30 18:20 | 28.49 | 249.3658 | 195.3107 | 247.3294 | 175.4879 | -1.6043 | 1.271 | 32.1402 | 40.8508 | 25.2869 | 2.0057 | 32.065 | 64.313 | 15.9869 | 14.3136 | -0.1106 | 105.1638 | 235.0522 | 180.9971 | 233.0158 | 161.1743
2012-10-30 18:20 | 28.4997 | 249.4068 | 195.3007 | 247.202 | 175.5835 | -1.5968 | 1.271 | 32.1402 | 40.851 | 25.2868 | 2.0057 | 32.0649 | 64.3131 | 15.9868 | 14.3315 | -0.1101 | 105.164 | 235.0753 | 180.9692 | 232.8705 | 161.252
2012-10-30 18:20 | 28.4974 | 249.4006 | 195.2748 | 247.2663 | 175.386 | -1.6043 | 1.271 | 32.1403 | 40.8513 | 25.2867 | 2.0057 | 32.065 | 64.3134 | 15.9868 | 14.3578 | -0.1106 | 105.1647 | 235.0428 | 180.917 | 232.9085 | 161.0282
2012-10-30 18:20 | 28.493 | 249.404 | 195.2587 | 247.2893 | 175.5143 | -1.6014 | 1.271 | 32.1396 | 40.8494 | 25.2869 | 2.0057 | 32.0643 | 64.3109 | 15.9867 | 14.3574 | -0.1104 | 105.1603 | 235.0466 | 180.9013 | 232.9319 | 161.1569
2012-10-30 18:20 | 28.5004 | 249.4349 | 195.27 | 247.2497 | 175.7169 | -1.6009 | 1.271 | 32.1397 | 40.8497 | 25.2869 | 2.0057 | 32.0644 | 64.3114 | 15.9867 | 14.3689 | -0.1104 | 105.1611 | 235.066 | 180.9011 | 232.8808 | 161.348
2012-10-30 18:20 | 28.4907 | 249.4174 | 195.2956 | 247.193 | 175.7111 | -1.6014 | 1.271 | 32.1399 | 40.8504 | 25.2867 | 2.0057 | 32.0647 | 64.3127 | 15.9867 | 14.347 | -0.1104 | 105.1631 | 235.0704 | 180.9486 | 232.846 | 161.3641
2012-10-30 18:21 | 28.4915 | 249.3438 | 195.2611 | 247.374 | 175.4777 | -1.6014 | 1.271 | 32.1398 | 40.8503 | 25.2866 | 2.0057 | 32.0645 | 64.3121 | 15.9866 | 14.3641 | -0.1104 | 105.1624 | 234.9797 | 180.897 | 233.0099 | 161.1136
2012-10-30 18:21 | 28.4877 | 249.4066 | 195.273 | 247.237 | 175.4817 | -1.6043 | 1.271 | 32.1396 | 40.8496 | 25.2868 | 2.0057 | 32.0643 | 64.3114 | 15.9866 | 14.3481 | -0.1106 | 105.161 | 235.0585 | 180.9249 | 232.8889 | 161.1336
2012-10-30 18:21 | 28.487 | 249.402 | 195.2409 | 247.1853 | 175.5083 | -1.598 | 1.271 | 32.1388 | 40.8475 | 25.2868 | 2.0057 | 32.0635 | 64.3087 | 15.9865 | 14.3555 | -0.1102 | 105.1562 | 235.0465 | 180.8854 | 232.8298 | 161.1528
2012-10-30 18:21 | 28.4877 | 249.4222 | 195.2376 | 247.3545 | 175.4036 | -1.6014 | 1.271 | 32.1386 | 40.8469 | 25.2869 | 2.0056 | 32.0633 | 64.3075 | 15.9865 | 14.3929 | -0.1104 | 105.1544 | 235.0293 | 180.8447 | 232.9616 | 161.0107
2012-10-30 18:21 | 28.49 | 249.448 | 195.2792 | 247.3059 | 175.4645 | -1.5985 | 1.2709 | 32.1384 | 40.8462 | 25.2869 | 2.0056 | 32.0632 | 64.3068 | 15.9866 | 14.3789 | -0.1102 | 105.153 | 235.0691 | 180.9003 | 232.927 | 161.0856
2012-10-30 18:21 | 28.4855 | 249.4357 | 195.2394 | 247.1564 | 175.4053 | -1.6032 | 1.271 | 32.1385 | 40.8468 | 25.2868 | 2.0056 | 32.0631 | 64.307 | 15.9865 | 14.3336 | -0.1105 | 105.1537 | 235.1021 | 180.9058 | 232.8228 | 161.0717
2012-10-30 18:21 | 28.4952 | 249.3827 | 195.2098 | 247.3111 | 175.3487 | -1.6009 | 1.2709 | 32.1382 | 40.8458 | 25.2868 | 2.0056 | 32.0628 | 64.3056 | 15.9865 | 14.3474 | -0.1104 | 105.1515 | 235.0353 | 180.8624 | 232.9637 | 161.0013
2012-10-30 18:21 | 28.4922 | 249.4502 | 195.2107 | 247.312 | 175.3262 | -1.6066 | 1.2709 | 32.138 | 40.8454 | 25.2868 | 2.0056 | 32.0626 | 64.3052 | 15.9864 | 14.3526 | -0.1108 | 105.1506 | 235.0976 | 180.8581 | 232.9594 | 160.9736
2012-10-30 18:21 | 28.4885 | 249.4113 | 195.1638 | 247.2064 | 175.3966 | -1.6009 | 1.2709 | 32.1375 | 40.8444 | 25.2867 | 2.0056 | 32.0622 | 64.3034 | 15.9865 | 14.3693 | -0.1104 | 105.1478 | 235.042 | 180.7945 | 232.8371 | 161.0273
2012-10-30 18:21 | 28.4885 | 249.4113 | 195.1599 | 247.3827 | 175.4513 | -1.5974 | 1.2709 | 32.1376 | 40.8447 | 25.2866 | 2.0056 | 32.0623 | 64.3042 | 15.9864 | 14.3366 | -0.1101 | 105.1489 | 235.0747 | 180.8233 | 233.0461 | 161.1147
2012-10-30 18:21 | 28.4915 | 249.4495 | 195.21 | 247.2878 | 175.4972 | -1.5985 | 1.2709 | 32.1372 | 40.8438 | 25.2866 | 2.0056 | 32.0619 | 64.3024 | 15.9864 | 14.3274 | -0.1102 | 105.1462 | 235.1221 | 180.8826 | 232.9604 | 161.1698
2012-10-30 18:21 | 28.493 | 249.4236 | 195.1683 | 247.2344 | 175.4519 | -1.6066 | 1.2709 | 32.137 | 40.8428 | 25.287 | 2.0056 | 32.0619 | 64.302 | 15.9865 | 14.3615 | -0.1108 | 105.1447 | 235.0621 | 180.8068 | 232.8729 | 161.0904
2012-10-30 18:21 | 28.4952 | 249.4336 | 195.1312 | 247.3071 | 175.3916 | -1.6043 | 1.2709 | 32.137 | 40.8428 | 25.2869 | 2.0056 | 32.0617 | 64.3016 | 15.9864 | 14.3107 | -0.1106 | 105.1445 | 235.1229 | 180.8205 | 232.9964 | 161.0809
2012-10-30 18:21 | 28.4922 | 249.4228 | 195.1714 | 247.3394 | 175.455 | -1.6009 | 1.2709 | 32.1365 | 40.8416 | 25.2869 | 2.0055 | 32.0612 | 64.2993 | 15.9864 | 14.3281 | -0.1104 | 105.1409 | 235.0947 | 180.8433 | 233.0113 | 161.1269
2012-10-30 18:21 | 28.4952 | 249.4258 | 195.139 | 247.1739 | 175.5048 | -1.5997 | 1.2709 | 32.136 | 40.8404 | 25.2867 | 2.0055 | 32.0608 | 64.2986 | 15.9863 | 14.3107 | -0.1103 | 105.1391 | 235.1151 | 180.8283 | 232.8632 | 161.1941
2012-10-30 18:21 | 28.4974 | 249.4398 | 195.1098 | 247.2858 | 175.4367 | -1.6043 | 1.2709 | 32.1359 | 40.8404 | 25.2867 | 2.0055 | 32.0607 | 64.2979 | 15.9863 | 14.3292 | -0.1106 | 105.1382 | 235.1106 | 180.7806 | 232.9566 | 161.1075
2012-10-30 18:21 | 28.4944 | 249.3624 | 195.1225 | 247.3455 | 175.2893 | -1.5997 | 1.2708 | 32.1354 | 40.8387 | 25.2868 | 2.0055 | 32.0601 | 64.296 | 15.9862 | 14.3017 | -0.1103 | 105.1347 | 235.0607 | 180.8208 | 233.0438 | 160.9876
2012-10-30 18:21 | 28.4974 | 249.4946 | 195.1059 | 247.2976 | 175.3938 | -1.6032 | 1.2708 | 32.1349 | 40.8375 | 25.2868 | 2.0055 | 32.0596 | 64.294 | 15.9862 | 14.3047 | -0.1105 | 105.1315 | 235.1899 | 180.8012 | 232.9929 | 161.0891
2012-10-30 18:21 | 28.5019 | 249.4325 | 195.0554 | 247.2355 | 175.3593 | -1.6003 | 1.2708 | 32.1352 | 40.8384 | 25.2867 | 2.0055 | 32.06 | 64.2955 | 15.9862 | 14.3092 | -0.1103 | 105.1339 | 235.1233 | 180.7462 | 232.9263 | 161.0501
2012-10-30 18:21 | 28.5049 | 249.4433 | 195.0937 | 247.3795 | 175.4286 | -1.5968 | 1.2708 | 32.135 | 40.8378 | 25.2869 | 2.0055 | 32.0598 | 64.2943 | 15.9863 | 14.3122 | -0.1101 | 105.1321 | 235.1311 | 180.7815 | 233.0673 | 161.1164
2012-10-30 18:21 | 28.5034 | 249.4144 | 195.1197 | 247.2801 | 175.4584 | -1.602 | 1.2708 | 32.1341 | 40.836 | 25.2865 | 2.0054 | 32.0588 | 64.2909 | 15.9862 | 14.3148 | -0.1104 | 105.1269 | 235.0996 | 180.8049 | 232.9653 | 161.1436
2012-10-30 18:21 | 28.5034 | 249.4418 | 195.1315 | 247.2135 | 175.4584 | -1.6014 | 1.2708 | 32.1338 | 40.8347 | 25.2868 | 2.0054 | 32.0585 | 64.2899 | 15.9862 | 14.3311 | -0.1104 | 105.1245 | 235.1107 | 180.8004 | 232.8824 | 161.1273
2012-10-30 18:21 | 28.5094 | 249.4713 | 195.0668 | 247.3018 | 175.3317 | -1.6009 | 1.2708 | 32.1332 | 40.8338 | 25.2865 | 2.0054 | 32.058 | 64.2879 | 15.9862 | 14.3208 | -0.1104 | 105.1217 | 235.1505 | 180.746 | 232.981 | 161.0109
2012-10-30 18:21 | 28.5147 | 249.4218 | 195.1468 | 247.3932 | 175.5282 | -1.5991 | 1.2707 | 32.1325 | 40.832 | 25.2865 | 2.0053 | 32.0572 | 64.2847 | 15.9862 | 14.3424 | -0.1102 | 105.1167 | 235.0794 | 180.8044 | 233.0508 | 161.1858
2012-10-30 18:22 | 28.5087 | 249.4471 | 195.129 | 247.2971 | 175.491 | -1.6003 | 1.2701 | 32.117 | 40.7923 | 25.2867 | 2.0048 | 32.0488 | 64.251 | 15.9861 | 14.2834 | -0.1103 | 105.0433 | 235.1637 | 180.8456 | 233.0137 | 161.2076
2012-10-30 18:22 | 28.5117 | 249.501 | 195.1005 | 247.2766 | 175.3105 | -1.6009 | 1.2643 | 31.9674 | 40.4154 | 25.2853 | 1.9949 | 31.8916 | 63.6219 | 15.9862 | 14.3435 | -0.1104 | 104.0372 | 235.1575 | 180.757 | 232.9331 | 160.967
2012-10-30 18:22 | 28.5094 | 249.4087 | 195.1179 | 247.3919 | 175.3629 | -1.6009 | 1.2644 | 31.9715 | 40.4259 | 25.2853 | 1.9952 | 31.8967 | 63.6417 | 15.9863 | 14.3045 | -0.1104 | 104.0676 | 235.1042 | 180.8134 | 233.0874 | 161.0584
2012-10-30 18:22 | 28.5094 | 249.4126 | 195.11 | 247.2783 | 175.4878 | -1.6014 | 1.2639 | 31.9566 | 40.3885 | 25.285 | 1.9943 | 31.8821 | 63.5834 | 15.9863 | 14.329 | -0.1104 | 103.9719 | 235.0836 | 180.781 | 232.9493 | 161.1588
2012-10-30 18:22 | 28.5072 | 249.3673 | 195.0135 | 247.2721 | 174.8651 | -1.6066 | 1.2635 | 31.9483 | 40.3678 | 25.2849 | 1.9938 | 31.8737 | 63.5497 | 15.9864 | 14.2982 | -0.1108 | 103.9175 | 235.0691 | 180.7153 | 232.9739 | 160.5669
2012-10-30 18:22 | 28.5117 | 249.3562 | 194.963 | 247.3667 | 174.3351 | -1.5997 | 1.2634 | 31.9441 | 40.357 | 25.285 | 1.9935 | 31.8693 | 63.5322 | 15.9864 | 14.3272 | -0.1103 | 103.8892 | 235.029 | 180.6358 | 233.0395 | 160.0079
2012-10-30 18:22 | 28.5147 | 249.3592 | 194.8717 | 247.2796 | 174.8804 | -1.602 | 1.2634 | 31.9444 | 40.3585 | 25.2845 | 1.9935 | 31.8696 | 63.533 | 15.9865 | 14.3424 | -0.1104 | 103.8915 | 235.0168 | 180.5293 | 232.9372 | 160.538
2012-10-30 18:22 | 28.5147 | 249.3474 | 194.8914 | 247.1778 | 175.2745 | -1.6003 | 1.2634 | 31.9437 | 40.3563 | 25.2848 | 1.9935 | 31.8689 | 63.5301 | 15.9866 | 14.3587 | -0.1103 | 103.8864 | 234.9887 | 180.5327 | 232.8191 | 160.9158
2012-10-30 18:22 | 28.5207 | 249.3299 | 194.8817 | 247.3836 | 175.1088 | -1.6043 | 1.2634 | 31.9439 | 40.3567 | 25.2849 | 1.9935 | 31.8692 | 63.5315 | 15.9865 | 14.3158 | -0.1106 | 103.8881 | 235.0141 | 180.5659 | 233.0678 | 160.793
2012-10-30 18:22 | 28.5124 | 249.3373 | 194.9598 | 247.1755 | 174.8157 | -1.5974 | 1.2635 | 31.9469 | 40.3644 | 25.2848 | 1.9937 | 31.8721 | 63.5429 | 15.9865 | 14.3238 | -0.1101 | 103.9073 | 235.0135 | 180.636 | 232.8517 | 160.4919
2012-10-30 18:22 | 28.5109 | 249.2849 | 194.864 | 247.2053 | 174.9039 | -1.5951 | 1.2641 | 31.9632 | 40.4057 | 25.2847 | 1.9948 | 31.8907 | 63.617 | 15.9865 | 14.3305 | -0.11 | 104.0227 | 234.9544 | 180.5335 | 232.8748 | 160.5734
2012-10-30 18:22 | 28.5132 | 249.3029 | 194.8467 | 247.333 | 174.7033 | -1.6043 | 1.2685 | 32.0747 | 40.6861 | 25.286 | 2.0017 | 31.9999 | 64.0534 | 15.9866 | 14.2838 | -0.1106 | 104.7395 | 235.0191 | 180.5629 | 233.0492 | 160.4195
2012-10-30 18:22 | 28.5132 | 249.2911 | 194.8859 | 247.2507 | 174.9686 | -1.6032 | 1.2691 | 32.0906 | 40.7254 | 25.2865 | 2.0027 | 32.0157 | 64.1176 | 15.9864 | 14.3409 | -0.1105 | 104.843 | 234.9502 | 180.545 | 232.9098 | 160.6277
2012-10-30 18:22 | 28.5214 | 249.3111 | 194.9295 | 247.1179 | 175.2032 | -1.6026 | 1.2697 | 32.1056 | 40.7636 | 25.2865 | 2.0036 | 32.0306 | 64.1767 | 15.9864 | 14.3287 | -0.1105 | 104.9402 | 234.9824 | 180.6008 | 232.7892 | 160.8745
2012-10-30 18:22 | 28.5117 | 249.3014 | 194.9355 | 247.2335 | 175.2013 | -1.602 | 1.2702 | 32.1187 | 40.7967 | 25.2866 | 2.0044 | 32.0435 | 64.2289 | 15.9864 | 14.3109 | -0.1104 | 105.0256 | 234.9905 | 180.6246 | 232.9226 | 160.8904
2012-10-30 18:22 | 28.5162 | 249.2785 | 194.9243 | 247.2341 | 175.4438 | -1.5951 | 1.2705 | 32.1259 | 40.8152 | 25.2865 | 2.0049 | 32.0507 | 64.2576 | 15.9864 | 14.3358 | -0.11 | 105.0728 | 234.9427 | 180.5885 | 232.8983 | 161.108
2012-10-30 18:22 | 28.5064 | 249.3 | 195.052 | 247.1107 | 175.6838 | -1.6032 | 1.2706 | 32.13 | 40.8252 | 25.2867 | 2.0051 | 32.0546 | 64.2739 | 15.9863 | 14.3464 | -0.1105 | 105.0991 | 234.9536 | 180.7056 | 232.7643 | 161.3374
2012-10-30 18:22 | 28.5124 | 249.2786 | 195.0541 | 247.1481 | 175.4869 | -1.5997 | 1.2708 | 32.1331 | 40.8333 | 25.2867 | 2.0053 | 32.0578 | 64.2865 | 15.9863 | 14.3524 | -0.1103 | 105.1198 | 234.9262 | 180.7017 | 232.7957 | 161.1345
2012-10-30 18:22 | 28.5117 | 249.2857 | 195.073 | 247.2453 | 175.4784 | -1.6009 | 1.2708 | 32.1354 | 40.8389 | 25.2868 | 2.0055 | 32.06 | 64.2959 | 15.9862 | 14.3598 | -0.1104 | 105.1347 | 234.9259 | 180.7132 | 232.8855 | 161.1186
2012-10-30 18:22 | 28.5109 | 249.3632 | 195.1076 | 247.127 | 175.5556 | -1.598 | 1.2709 | 32.1366 | 40.842 | 25.2867 | 2.0056 | 32.0614 | 64.3013 | 15.9862 | 14.3549 | -0.1102 | 105.1433 | 235.0083 | 180.7527 | 232.7721 | 161.2007
2012-10-30 18:22 | 28.5214 | 249.3933 | 195.0788 | 247.1335 | 175.5466 | -1.5997 | 1.2709 | 32.1377 | 40.8445 | 25.2869 | 2.0056 | 32.0623 | 64.3052 | 15.9861 | 14.3695 | -0.1103 | 105.1497 | 235.0238 | 180.7093 | 232.764 | 161.1771
2012-10-30 18:22 | 28.5154 | 249.2933 | 195.12 | 247.2725 | 175.4899 | -1.6049 | 1.2709 | 32.1382 | 40.8457 | 25.2869 | 2.0057 | 32.0628 | 64.3073 | 15.9861 | 14.4002 | -0.1106 | 105.1529 | 234.8931 | 180.7198 | 232.8723 | 161.0897
2012-10-30 18:22 | 28.5064 | 249.3822 | 195.1542 | 247.099 | 175.477 | -1.5991 | 1.271 | 32.1387 | 40.847 | 25.287 | 2.0057 | 32.0635 | 64.3093 | 15.9863 | 14.3994 | -0.1102 | 105.1564 | 234.9828 | 180.7548 | 232.6996 | 161.0776
2012-10-30 18:22 | 28.5162 | 249.3763 | 195.1168 | 247.1558 | 175.4399 | -1.6032 | 1.271 | 32.139 | 40.8478 | 25.2869 | 2.0057 | 32.0637 | 64.3111 | 15.9861 | 14.3562 | -0.1105 | 105.1589 | 235.0201 | 180.7606 | 232.7996 | 161.0837
2012-10-30 18:22 | 28.5109 | 249.3397 | 195.143 | 247.2327 | 175.5088 | -1.6026 | 1.271 | 32.1395 | 40.8492 | 25.2869 | 2.0057 | 32.0642 | 64.3127 | 15.9862 | 14.3509 | -0.1105 | 105.1619 | 234.9888 | 180.7921 | 232.8818 | 161.1579
2012-10-30 18:22 | 28.5072 | 249.2812 | 195.1432 | 247.2212 | 175.4934 | -1.6009 | 1.271 | 32.1403 | 40.8508 | 25.2871 | 2.0058 | 32.0649 | 64.3163 | 15.986 | 14.3431 | -0.1104 | 105.1671 | 234.9381 | 180.8001 | 232.8781 | 161.1503
2012-10-30 18:23 | 28.5079 | 249.3641 | 195.1871 | 247.0887 | 175.5292 | -1.6009 | 1.271 | 32.1401 | 40.8504 | 25.287 | 2.0058 | 32.0648 | 64.3154 | 15.9861 | 14.3519 | -0.1104 | 105.1658 | 235.0122 | 180.8352 | 232.7368 | 161.1773
2012-10-30 18:23 | 28.5109 | 249.3358 | 195.1272 | 247.1583 | 175.4659 | -1.6072 | 1.271 | 32.1404 | 40.8518 | 25.2867 | 2.0058 | 32.0651 | 64.3165 | 15.9861 | 14.3631 | -0.1108 | 105.1684 | 234.9727 | 180.7641 | 232.7952 | 161.1028
2012-10-30 18:23 | 28.5087 | 249.3219 | 195.1722 | 247.3206 | 175.4988 | -1.6055 | 1.271 | 32.1405 | 40.8514 | 25.287 | 2.0058 | 32.0651 | 64.3172 | 15.986 | 14.3446 | -0.1107 | 105.1686 | 234.9773 | 180.8276 | 232.976 | 161.1542
2012-10-30 18:23 | 28.5117 | 249.3601 | 195.1438 | 247.1944 | 175.4979 | -1.5997 | 1.271 | 32.1407 | 40.8518 | 25.287 | 2.0058 | 32.0653 | 64.3178 | 15.986 | 14.319 | -0.1103 | 105.1696 | 235.0411 | 180.8248 | 232.8754 | 161.1789
2012-10-30 18:23 | 28.5169 | 249.3966 | 195.2 | 247.1055 | 175.4953 | -1.602 | 1.2711 | 32.1408 | 40.8528 | 25.2866 | 2.0059 | 32.0655 | 64.319 | 15.9859 | 14.3161 | -0.1104 | 105.1718 | 235.0805 | 180.8839 | 232.7894 | 161.1792
2012-10-30 18:23 | 28.5192 | 249.3754 | 195.1238 | 247.2215 | 175.3571 | -1.5991 | 1.2711 | 32.1411 | 40.8532 | 25.2869 | 2.0059 | 32.0657 | 64.3202 | 15.9858 | 14.3347 | -0.1102 | 105.1734 | 235.0407 | 180.7891 | 232.8868 | 161.0224
2012-10-30 18:23 | 28.5147 | 249.3748 | 195.1075 | 247.2483 | 175.3252 | -1.6026 | 1.271 | 32.141 | 40.8527 | 25.2871 | 2.0059 | 32.0656 | 64.3199 | 15.9858 | 14.3179 | -0.1105 | 105.1726 | 235.0569 | 180.7896 | 232.9304 | 161.0073
2012-10-30 18:23 | 28.5154 | 249.3951 | 195.0925 | 247.202 | 175.4704 | -1.6003 | 1.2711 | 32.1411 | 40.8533 | 25.2867 | 2.0059 | 32.0657 | 64.3205 | 15.9857 | 14.3146 | -0.1103 | 105.1738 | 235.0805 | 180.7779 | 232.8874 | 161.1558
2012-10-30 18:23 | 28.5139 | 249.4093 | 195.1381 | 247.0594 | 175.5118 | -1.5939 | 1.271 | 32.1406 | 40.852 | 25.2869 | 2.0059 | 32.0653 | 64.3185 | 15.9858 | 14.3335 | -0.1099 | 105.1704 | 235.0758 | 180.8046 | 232.7259 | 161.1783
2012-10-30 18:23 | 28.5117 | 249.3758 | 195.0927 | 247.2492 | 175.3769 | -1.5991 | 1.271 | 32.1406 | 40.8522 | 25.2868 | 2.0059 | 32.0653 | 64.319 | 15.9857 | 14.3272 | -0.1102 | 105.1711 | 235.0486 | 180.7655 | 232.922 | 161.0497
2012-10-30 18:23 | 28.5124 | 249.4234 | 195.1562 | 247.2342 | 175.4283 | -1.6037 | 1.271 | 32.1404 | 40.8512 | 25.287 | 2.0059 | 32.0651 | 64.3185 | 15.9856 | 14.3646 | -0.1106 | 105.1696 | 235.0588 | 180.7916 | 232.8696 | 161.0637
2012-10-30 18:23 | 28.5124 | 249.4195 | 195.1484 | 247.1951 | 175.3815 | -1.5985 | 1.271 | 32.1403 | 40.8514 | 25.2867 | 2.0059 | 32.065 | 64.318 | 15.9857 | 14.3442 | -0.1102 | 105.1693 | 235.0753 | 180.8042 | 232.8509 | 161.0373
2012-10-30 18:23 | 28.5094 | 249.4087 | 195.165 | 247.2195 | 175.4683 | -1.6009 | 1.271 | 32.1401 | 40.851 | 25.2867 | 2.0059 | 32.0648 | 64.3172 | 15.9856 | 14.3534 | -0.1104 | 105.1682 | 235.0553 | 180.8116 | 232.8661 | 161.1149
2012-10-30 18:23 | 28.5177 | 249.4092 | 195.1026 | 247.1847 | 175.3048 | -1.5974 | 1.271 | 32.1395 | 40.8492 | 25.2869 | 2.0058 | 32.0643 | 64.315 | 15.9857 | 14.3658 | -0.1101 | 105.1642 | 235.0434 | 180.7368 | 232.8189 | 160.939
2012-10-30 18:23 | 28.5132 | 249.4047 | 195.161 | 247.2821 | 175.3316 | -1.5939 | 1.271 | 32.1393 | 40.8489 | 25.2867 | 2.0058 | 32.064 | 64.3139 | 15.9857 | 14.3613 | -0.1099 | 105.1628 | 235.0434 | 180.7997 | 232.9208 | 160.9703
2012-10-30 18:23 | 28.5102 | 249.4134 | 195.1383 | 247.1419 | 175.5627 | -1.5957 | 1.271 | 32.1395 | 40.8494 | 25.2867 | 2.0058 | 32.0641 | 64.3145 | 15.9857 | 14.3338 | -0.11 | 105.1639 | 235.0796 | 180.8045 | 232.8081 | 161.2289
2012-10-30 18:23 | 28.5154 | 249.4617 | 195.171 | 247.2177 | 175.6616 | -1.6009 | 1.271 | 32.1393 | 40.8488 | 25.2868 | 2.0058 | 32.064 | 64.3149 | 15.9854 | 14.3472 | -0.1104 | 105.1637 | 235.1145 | 180.8238 | 232.8705 | 161.3144
2012-10-30 18:23 | 28.5169 | 249.4319 | 195.1647 | 247.2074 | 175.5928 | -1.598 | 1.271 | 32.139 | 40.8484 | 25.2866 | 2.0058 | 32.0638 | 64.314 | 15.9854 | 14.3405 | -0.1102 | 105.1624 | 235.0914 | 180.8242 | 232.8669 | 161.2523
2012-10-30 18:23 | 28.5124 | 249.4352 | 195.2034 | 247.3048 | 175.5805 | -1.6055 | 1.271 | 32.1389 | 40.848 | 25.2867 | 2.0058 | 32.0636 | 64.3126 | 15.9856 | 14.3401 | -0.1107 | 105.1606 | 235.0951 | 180.8633 | 232.9647 | 161.2404
2012-10-30 18:23 | 28.5139 | 249.4602 | 195.256 | 247.2201 | 175.7655 | -1.5997 | 1.271 | 32.1387 | 40.847 | 25.287 | 2.0058 | 32.0634 | 64.3125 | 15.9854 | 14.3294 | -0.1103 | 105.1594 | 235.1308 | 180.9266 | 232.8907 | 161.4361
2012-10-30 18:23 | 28.5102 | 249.4447 | 195.2366 | 247.2007 | 175.6993 | -1.6014 | 1.2709 | 32.138 | 40.8456 | 25.2867 | 2.0057 | 32.0626 | 64.3095 | 15.9854 | 14.3461 | -0.1104 | 105.155 | 235.0986 | 180.8905 | 232.8546 | 161.3532
2012-10-30 18:23 | 28.5154 | 249.4343 | 195.2418 | 247.2372 | 175.6031 | -1.6032 | 1.2709 | 32.1378 | 40.8449 | 25.2869 | 2.0058 | 32.0625 | 64.3096 | 15.9853 | 14.3472 | -0.1105 | 105.1545 | 235.0871 | 180.8946 | 232.89 | 161.2559
2012-10-30 18:23 | 28.5042 | 249.4348 | 195.2423 | 247.2926 | 175.6621 | -1.598 | 1.2709 | 32.1378 | 40.8446 | 25.2871 | 2.0057 | 32.0625 | 64.3091 | 15.9853 | 14.3319 | -0.1102 | 105.1536 | 235.1029 | 180.9104 | 232.9607 | 161.3302
2012-10-30 18:23 | 28.5184 | 249.4803 | 195.3391 | 247.252 | 175.7075 | -1.6003 | 1.2709 | 32.1374 | 40.8437 | 25.287 | 2.0057 | 32.0621 | 64.3078 | 15.9853 | 14.4032 | -0.1103 | 105.1514 | 235.0771 | 180.9359 | 232.8488 | 161.3043
2012-10-30 18:24 | 28.5064 | 249.4918 | 195.2681 | 247.1969 | 175.6136 | -1.6043 | 1.2709 | 32.1375 | 40.8437 | 25.2871 | 2.0057 | 32.0621 | 64.3074 | 15.9854 | 14.326 | -0.1106 | 105.1511 | 235.1658 | 180.9421 | 232.8709 | 161.2876
2012-10-30 18:24 | 28.5079 | 249.4503 | 195.3011 | 247.3512 | 175.6892 | -1.5928 | 1.2709 | 32.1371 | 40.843 | 25.2868 | 2.0057 | 32.0618 | 64.306 | 15.9854 | 14.3479 | -0.1098 | 105.149 | 235.1024 | 180.9532 | 233.0033 | 161.3413
2012-10-30 18:24 | 28.5102 | 249.4565 | 195.3073 | 247.2555 | 175.6408 | -1.6003 | 1.2709 | 32.1366 | 40.8416 | 25.287 | 2.0057 | 32.0613 | 64.3043 | 15.9853 | 14.3461 | -0.1103 | 105.1459 | 235.1104 | 180.9612 | 232.9094 | 161.2947
2012-10-30 18:24 | 28.5124 | 249.4743 | 195.3606 | 247.199 | 175.7093 | -1.5997 | 1.2709 | 32.1362 | 40.8406 | 25.287 | 2.0056 | 32.0608 | 64.3025 | 15.9853 | 14.4136 | -0.1103 | 105.1431 | 235.0607 | 180.947 | 232.7854 | 161.2957
2012-10-30 18:24 | 28.5049 | 249.4981 | 195.302 | 247.2267 | 175.6004 | -1.602 | 1.2709 | 32.1359 | 40.8404 | 25.2867 | 2.0056 | 32.0606 | 64.3017 | 15.9853 | 14.3857 | -0.1104 | 105.1421 | 235.1124 | 180.9163 | 232.841 | 161.2147
2012-10-30 18:24 | 28.5019 | 249.4717 | 195.2911 | 247.2825 | 175.4725 | -1.5991 | 1.2709 | 32.1358 | 40.84 | 25.2867 | 2.0056 | 32.0606 | 64.302 | 15.9852 | 14.3296 | -0.1102 | 105.142 | 235.1421 | 180.9615 | 232.9529 | 161.1429
2012-10-30 18:24 | 28.5094 | 249.4831 | 195.279 | 247.2861 | 175.4956 | -1.6026 | 1.2708 | 32.1354 | 40.8389 | 25.2868 | 2.0056 | 32.0601 | 64.2996 | 15.9853 | 14.3371 | -0.1105 | 105.1384 | 235.146 | 180.9419 | 232.949 | 161.1585
2012-10-30 18:24 | 28.5019 | 249.4638 | 195.2793 | 247.2394 | 175.5271 | -1.598 | 1.2708 | 32.1348 | 40.8376 | 25.2867 | 2.0056 | 32.0595 | 64.2976 | 15.9852 | 14.3582 | -0.1102 | 105.1352 | 235.1056 | 180.9211 | 232.8812 | 161.1689
2012-10-30 18:24 | 28.5147 | 249.4923 | 195.2646 | 247.2013 | 175.4892 | -1.6043 | 1.2708 | 32.1347 | 40.8369 | 25.2869 | 2.0056 | 32.0594 | 64.2969 | 15.9852 | 14.3669 | -0.1106 | 105.1338 | 235.1254 | 180.8977 | 232.8344 | 161.1223
2012-10-30 18:24 | 28.5064 | 249.4918 | 195.2721 | 247.3262 | 175.3208 | -1.602 | 1.2708 | 32.1345 | 40.8366 | 25.2868 | 2.0055 | 32.0593 | 64.296 | 15.9854 | 14.3464 | -0.1104 | 105.1326 | 235.1454 | 180.9257 | 232.9798 | 160.9744
2012-10-30 18:24 | 28.5207 | 249.5453 | 195.2824 | 247.2856 | 175.542 | -1.6043 | 1.2708 | 32.1344 | 40.8367 | 25.2866 | 2.0055 | 32.059 | 64.2957 | 15.9852 | 14.3607 | -0.1106 | 105.1324 | 235.1846 | 180.9217 | 232.9249 | 161.1813
2012-10-30 18:24 | 28.5147 | 249.5158 | 195.245 | 247.3031 | 175.5438 | -1.5962 | 1.2708 | 32.1341 | 40.8357 | 25.2867 | 2.0055 | 32.0588 | 64.295 | 15.9851 | 14.3506 | -0.11 | 105.1308 | 235.1652 | 180.8944 | 232.9525 | 161.1932
2012-10-30 18:24 | 28.5184 | 249.4999 | 195.1819 | 247.2559 | 175.489 | -1.6026 | 1.2708 | 32.1334 | 40.8337 | 25.2868 | 2.0055 | 32.0581 | 64.2926 | 15.9851 | 14.3461 | -0.1105 | 105.1263 | 235.1538 | 180.8358 | 232.9098 | 161.1429
2012-10-30 18:24 | 28.5192 | 249.489 | 195.2063 | 247.4056 | 175.4507 | -1.5951 | 1.2708 | 32.1328 | 40.8328 | 25.2864 | 2.0055 | 32.0575 | 64.2903 | 15.9851 | 14.3388 | -0.11 | 105.1231 | 235.1502 | 180.8675 | 233.0668 | 161.1119
2012-10-30 18:24 | 28.5154 | 249.5126 | 195.2457 | 247.1863 | 175.4235 | -1.5985 | 1.2707 | 32.1327 | 40.8322 | 25.2866 | 2.0055 | 32.0574 | 64.2903 | 15.985 | 14.3472 | -0.1102 | 105.1225 | 235.1654 | 180.8985 | 232.8391 | 161.0763
2012-10-30 18:24 | 28.5132 | 249.5025 | 195.2317 | 247.2781 | 175.3823 | -1.602 | 1.2682 | 32.0679 | 40.6682 | 25.2863 | 2.0015 | 31.9946 | 64.0384 | 15.985 | 14.394 | -0.1104 | 104.7066 | 235.1085 | 180.8377 | 232.8841 | 160.9883
2012-10-30 18:24 | 28.5229 | 249.5005 | 195.206 | 247.3662 | 175.5286 | -1.6037 | 1.2648 | 31.9795 | 40.4467 | 25.2848 | 1.9959 | 31.9043 | 63.677 | 15.9851 | 14.3588 | -0.1106 | 104.1237 | 235.1417 | 180.8472 | 233.0074 | 161.1698
2012-10-30 18:24 | 28.5139 | 249.468 | 195.2245 | 247.2553 | 175.7147 | -1.6032 | 1.2642 | 31.9652 | 40.4103 | 25.2851 | 1.995 | 31.8899 | 63.6197 | 15.9851 | 14.3375 | -0.1105 | 104.03 | 235.1305 | 180.887 | 232.9178 | 161.3772
2012-10-30 18:24 | 28.5207 | 249.467 | 195.1803 | 247.2347 | 175.6669 | -1.6014 | 1.2637 | 31.9522 | 40.3776 | 25.2849 | 1.9942 | 31.8772 | 63.5685 | 15.9852 | 14.3362 | -0.1104 | 103.9461 | 235.1308 | 180.8441 | 232.8985 | 161.3307
2012-10-30 18:24 | 28.5229 | 249.4574 | 195.1589 | 247.3701 | 175.5325 | -1.5997 | 1.2635 | 31.9466 | 40.3636 | 25.2848 | 1.9938 | 31.8717 | 63.5465 | 15.9853 | 14.3914 | -0.1103 | 103.91 | 235.066 | 180.7675 | 232.9787 | 161.1411
2012-10-30 18:24 | 28.5184 | 249.4138 | 195.1937 | 247.3343 | 175.528 | -1.6009 | 1.2634 | 31.9442 | 40.3573 | 25.2849 | 1.9937 | 31.8693 | 63.5365 | 15.9853 | 14.3339 | -0.1104 | 103.8938 | 235.0799 | 180.8598 | 233.0004 | 161.1941
2012-10-30 18:24 | 28.5214 | 249.4011 | 195.2399 | 247.2041 | 175.6364 | -1.5985 | 1.2634 | 31.9453 | 40.3604 | 25.2847 | 1.9937 | 31.8705 | 63.5405 | 15.9855 | 14.3491 | -0.1102 | 103.9009 | 235.052 | 180.8908 | 232.855 | 161.2873
2012-10-30 18:24 | 28.5117 | 249.3718 | 195.2027 | 247.2923 | 175.2559 | -1.6049 | 1.2635 | 31.9465 | 40.3638 | 25.2846 | 1.9938 | 31.8718 | 63.5459 | 15.9854 | 14.3517 | -0.1106 | 103.9098 | 235.0201 | 180.851 | 232.9406 | 160.9042
2012-10-30 18:24 | 28.5214 | 249.4089 | 195.1456 | 247.3216 | 175.3398 | -1.6009 | 1.2636 | 31.9495 | 40.3707 | 25.2849 | 1.994 | 31.8745 | 63.5572 | 15.9854 | 14.3858 | -0.1104 | 103.928 | 235.0231 | 180.7598 | 232.9358 | 160.954
2012-10-30 18:24 | 28.5177 | 249.3974 | 195.1969 | 247.1612 | 175.4805 | -1.5991 | 1.2636 | 31.9506 | 40.3736 | 25.2848 | 1.994 | 31.8757 | 63.5615 | 15.9854 | 14.3985 | -0.1102 | 103.9351 | 234.9989 | 180.7984 | 232.7627 | 161.082
2012-10-30 18:25 | 28.5177 | 249.3583 | 195.1891 | 247.2435 | 175.4531 | -1.5991 | 1.2661 | 32.0132 | 40.5312 | 25.2853 | 1.9979 | 31.9374 | 63.8083 | 15.9853 | 14.3903 | -0.1102 | 104.3395 | 234.968 | 180.7988 | 232.8532 | 161.0628
2012-10-30 18:25 | 28.5154 | 249.399 | 195.1317 | 247.2843 | 175.607 | -1.6003 | 1.2688 | 32.0833 | 40.7073 | 25.2864 | 2.0023 | 32.0081 | 64.0904 | 15.9855 | 14.3921 | -0.1103 | 104.7976 | 235.0069 | 180.7396 | 232.8922 | 161.2149
2012-10-30 18:25 | 28.5154 | 249.3716 | 195.2535 | 247.2921 | 175.7787 | -1.6014 | 1.2694 | 32.0991 | 40.7473 | 25.2864 | 2.0033 | 32.0234 | 64.1529 | 15.9852 | 14.3431 | -0.1104 | 104.9002 | 235.0285 | 180.9104 | 232.949 | 161.4356
2012-10-30 18:25 | 28.5094 | 249.4165 | 195.2475 | 247.1098 | 175.7259 | -1.598 | 1.2699 | 32.1117 | 40.7795 | 25.2862 | 2.0041 | 32.0363 | 64.2038 | 15.9854 | 14.4065 | -0.1102 | 104.9834 | 235.01 | 180.841 | 232.7033 | 161.3194
2012-10-30 18:25 | 28.5057 | 249.3658 | 195.2438 | 247.302 | 175.6753 | -1.6014 | 1.2703 | 32.1205 | 40.8017 | 25.2863 | 2.0047 | 32.0449 | 64.2397 | 15.9851 | 14.3457 | -0.1104 | 105.0414 | 235.0201 | 180.8981 | 232.9563 | 161.3296
2012-10-30 18:25 | 28.5124 | 249.3608 | 195.3134 | 247.1637 | 175.6625 | -1.5985 | 1.2705 | 32.1262 | 40.8156 | 25.2868 | 2.005 | 32.0509 | 64.263 | 15.9853 | 14.4054 | -0.1102 | 105.0786 | 234.9554 | 180.908 | 232.7583 | 161.2571
2012-10-30 18:25 | 28.5064 | 249.4292 | 195.2721 | 247.1146 | 175.6955 | -1.5974 | 1.2707 | 32.1304 | 40.8266 | 25.2865 | 2.0053 | 32.055 | 64.2802 | 15.9851 | 14.3912 | -0.1101 | 105.1068 | 235.038 | 180.8809 | 232.7234 | 161.3043
2012-10-30 18:25 | 28.5072 | 249.4026 | 195.2886 | 247.2643 | 175.4739 | -1.6078 | 1.2708 | 32.1337 | 40.8356 | 25.286 | 2.0055 | 32.0582 | 64.2926 | 15.9851 | 14.3635 | -0.1108 | 105.1283 | 235.0391 | 180.9251 | 232.9008 | 161.1104
2012-10-30 18:25 | 28.4974 | 249.4006 | 195.3102 | 247.2153 | 175.6592 | -1.6026 | 1.2708 | 32.1355 | 40.8387 | 25.2871 | 2.0056 | 32.0601 | 64.3003 | 15.9852 | 14.3047 | -0.1105 | 105.139 | 235.0959 | 181.0055 | 232.9106 | 161.3545
2012-10-30 18:25 | 28.5012 | 249.4436 | 195.3061 | 247.1251 | 175.6162 | -1.602 | 1.2709 | 32.1369 | 40.8426 | 25.2869 | 2.0057 | 32.0615 | 64.3061 | 15.9851 | 14.3493 | -0.1104 | 105.1488 | 235.0943 | 180.9568 | 232.7758 | 161.2669
2012-10-30 18:25 | 28.5019 | 249.4129 | 195.2676 | 247.2081 | 175.4608 | -1.5985 | 1.2709 | 32.1379 | 40.8452 | 25.2868 | 2.0058 | 32.0625 | 64.3101 | 15.9851 | 14.3623 | -0.1102 | 105.1553 | 235.0506 | 180.9053 | 232.8458 | 161.0985
2012-10-30 18:25 | 28.4922 | 249.4267 | 195.2421 | 247.265 | 175.4667 | -1.5991 | 1.271 | 32.139 | 40.8481 | 25.2868 | 2.0058 | 32.0636 | 64.3142 | 15.9851 | 14.324 | -0.1102 | 105.1623 | 235.1027 | 180.9181 | 232.941 | 161.1427
2012-10-30 18:25 | 28.4967 | 249.4077 | 195.2545 | 247.1441 | 175.5375 | -1.5968 | 1.271 | 32.1397 | 40.8495 | 25.287 | 2.0059 | 32.0644 | 64.3177 | 15.9851 | 14.3611 | -0.1101 | 105.1672 | 235.0466 | 180.8934 | 232.783 | 161.1764
2012-10-30 18:25 | 28.4959 | 249.4265 | 195.2026 | 247.1786 | 175.4899 | -1.602 | 1.271 | 32.1404 | 40.8515 | 25.2869 | 2.0059 | 32.065 | 64.3198 | 15.9852 | 14.3399 | -0.1104 | 105.1713 | 235.0866 | 180.8627 | 232.8387 | 161.15
2012-10-30 18:25 | 28.4974 | 249.4515 | 195.2277 | 247.3289 | 175.4211 | -1.602 | 1.2711 | 32.1408 | 40.8528 | 25.2866 | 2.006 | 32.0655 | 64.3222 | 15.9851 | 14.3578 | -0.1104 | 105.175 | 235.0937 | 180.8699 | 232.9711 | 161.0633
2012-10-30 18:25 | 28.4915 | 249.4182 | 195.2296 | 247.2055 | 175.6182 | -1.5997 | 1.2711 | 32.141 | 40.8529 | 25.2869 | 2.006 | 32.0656 | 64.323 | 15.985 | 14.3233 | -0.1103 | 105.1759 | 235.0949 | 180.9063 | 232.8822 | 161.2949
2012-10-30 18:25 | 28.4855 | 249.4435 | 195.2236 | 247.1251 | 175.6122 | -1.6066 | 1.271 | 32.1408 | 40.8523 | 25.287 | 2.006 | 32.0655 | 64.3231 | 15.9848 | 14.3499 | -0.1108 | 105.1754 | 235.0936 | 180.8737 | 232.7752 | 161.2623
2012-10-30 18:25 | 28.4855 | 249.4318 | 195.1883 | 247.2191 | 175.538 | -1.6009 | 1.2711 | 32.1409 | 40.8529 | 25.2867 | 2.006 | 32.0655 | 64.3225 | 15.985 | 14.3418 | -0.1104 | 105.1755 | 235.09 | 180.8465 | 232.8773 | 161.1962
2012-10-30 18:25 | 28.4877 | 249.4731 | 195.1905 | 247.2683 | 175.7081 | -1.5933 | 1.271 | 32.1409 | 40.8524 | 25.287 | 2.006 | 32.0656 | 64.3236 | 15.9848 | 14.3277 | -0.1098 | 105.176 | 235.1454 | 180.8628 | 232.9406 | 161.3804
2012-10-30 18:25 | 28.4877 | 249.4301 | 195.2298 | 247.2252 | 175.7042 | -1.6014 | 1.271 | 32.1409 | 40.8525 | 25.287 | 2.006 | 32.0654 | 64.3234 | 15.9847 | 14.295 | -0.1104 | 105.1759 | 235.1351 | 180.9348 | 232.9302 | 161.4092
2012-10-30 18:25 | 28.4825 | 249.464 | 195.2128 | 247.1299 | 175.4648 | -1.6072 | 1.2711 | 32.141 | 40.8529 | 25.2869 | 2.006 | 32.0656 | 64.3241 | 15.9847 | 14.3102 | -0.1108 | 105.177 | 235.1538 | 180.9026 | 232.8197 | 161.1546
2012-10-30 18:25 | 28.487 | 249.4568 | 195.233 | 247.2754 | 175.3951 | -1.5997 | 1.2711 | 32.1412 | 40.8538 | 25.2866 | 2.006 | 32.0657 | 64.324 | 15.9849 | 14.3678 | -0.1103 | 105.1778 | 235.089 | 180.8652 | 232.9076 | 161.0273
2012-10-30 18:25 | 28.4802 | 249.4304 | 195.2223 | 247.2373 | 175.5796 | -1.6037 | 1.2711 | 32.1414 | 40.8534 | 25.2872 | 2.006 | 32.0659 | 64.3249 | 15.9848 | 14.2875 | -0.1106 | 105.1783 | 235.1429 | 180.9348 | 232.9498 | 161.2921
2012-10-30 18:25 | 28.4862 | 249.4325 | 195.2086 | 247.1689 | 175.4021 | -1.6055 | 1.2711 | 32.1414 | 40.854 | 25.2868 | 2.006 | 32.066 | 64.3254 | 15.9848 | 14.3425 | -0.1107 | 105.1794 | 235.09 | 180.8661 | 232.8264 | 161.0596
2012-10-30 18:25 | 28.4825 | 249.4797 | 195.201 | 247.2944 | 175.4414 | -1.6014 | 1.2711 | 32.1411 | 40.8534 | 25.2868 | 2.006 | 32.0658 | 64.3239 | 15.9849 | 14.3306 | -0.1104 | 105.1773 | 235.1491 | 180.8704 | 232.9638 | 161.1108
2012-10-30 18:25 | 28.49 | 249.4676 | 195.2321 | 247.3059 | 175.5542 | -1.6014 | 1.271 | 32.1408 | 40.8522 | 25.287 | 2.006 | 32.0654 | 64.3238 | 15.9846 | 14.3544 | -0.1104 | 105.176 | 235.1132 | 180.8777 | 232.9515 | 161.1998
2012-10-30 18:25 | 28.4832 | 249.5039 | 195.1585 | 247.0953 | 175.6021 | -1.5991 | 1.271 | 32.1406 | 40.8519 | 25.287 | 2.006 | 32.0653 | 64.323 | 15.9847 | 14.3028 | -0.1102 | 105.1749 | 235.2011 | 180.8557 | 232.7925 | 161.2993
2012-10-30 18:26 | 28.49 | 249.4637 | 195.181 | 247.3098 | 175.4801 | -1.6037 | 1.2711 | 32.1409 | 40.8528 | 25.2869 | 2.006 | 32.0657 | 64.3247 | 15.9846 | 14.2932 | -0.1106 | 105.1775 | 235.1705 | 180.8878 | 233.0166 | 161.1869
2012-10-30 18:26 | 28.4847 | 249.4192 | 195.1757 | 247.2849 | 175.6075 | -1.6014 | 1.271 | 32.1409 | 40.8524 | 25.2871 | 2.006 | 32.0655 | 64.324 | 15.9846 | 14.3002 | -0.1104 | 105.1764 | 235.119 | 180.8755 | 232.9847 | 161.3073
2012-10-30 18:26 | 28.4937 | 249.483 | 195.1886 | 247.2273 | 175.6204 | -1.5991 | 1.271 | 32.1408 | 40.8524 | 25.287 | 2.006 | 32.0655 | 64.3242 | 15.9846 | 14.3133 | -0.1102 | 105.1766 | 235.1697 | 180.8753 | 232.914 | 161.3071
2012-10-30 18:26 | 28.49 | 249.495 | 195.1889 | 247.1923 | 175.5269 | -1.5974 | 1.271 | 32.1408 | 40.8522 | 25.287 | 2.006 | 32.0654 | 64.3237 | 15.9846 | 14.33 | -0.1101 | 105.1759 | 235.165 | 180.8589 | 232.8623 | 161.1969
2012-10-30 18:26 | 28.4817 | 249.4515 | 195.2002 | 247.3602 | 175.4874 | -1.6003 | 1.271 | 32.1407 | 40.8524 | 25.2868 | 2.006 | 32.0654 | 64.324 | 15.9845 | 14.2645 | -0.1103 | 105.1764 | 235.187 | 180.9357 | 233.0957 | 161.2229
2012-10-30 18:26 | 28.4982 | 249.4875 | 195.2403 | 247.173 | 175.6912 | -1.6003 | 1.271 | 32.1404 | 40.8512 | 25.2871 | 2.006 | 32.0651 | 64.3227 | 15.9845 | 14.3218 | -0.1103 | 105.1739 | 235.1657 | 180.9185 | 232.8512 | 161.3694
2012-10-30 18:26 | 28.4885 | 249.4426 | 195.2031 | 247.2378 | 175.6152 | -1.5951 | 1.271 | 32.1404 | 40.8513 | 25.287 | 2.006 | 32.0649 | 64.3222 | 15.9845 | 14.3285 | -0.11 | 105.1734 | 235.1141 | 180.8746 | 232.9093 | 161.2867
2012-10-30 18:26 | 28.4982 | 249.4601 | 195.2246 | 247.2866 | 175.6132 | -1.5985 | 1.271 | 32.1403 | 40.8512 | 25.2868 | 2.006 | 32.0648 | 64.3217 | 15.9845 | 14.379 | -0.1102 | 105.1729 | 235.0811 | 180.8456 | 232.9076 | 161.2342
2012-10-30 18:26 | 28.5012 | 249.471 | 195.2983 | 247.274 | 175.706 | -1.6032 | 1.271 | 32.1402 | 40.8506 | 25.287 | 2.006 | 32.0649 | 64.3225 | 15.9844 | 14.3738 | -0.1105 | 105.1731 | 235.0972 | 180.9245 | 232.9002 | 161.3322
2012-10-30 18:26 | 28.4892 | 249.502 | 195.2588 | 247.1444 | 175.6081 | -1.6084 | 1.271 | 32.1395 | 40.8488 | 25.2871 | 2.006 | 32.0642 | 64.3195 | 15.9845 | 14.3781 | -0.1109 | 105.1684 | 235.1239 | 180.8807 | 232.7663 | 161.23
2012-10-30 18:26 | 28.4974 | 249.4789 | 195.2513 | 247.325 | 175.4406 | -1.6009 | 1.271 | 32.1393 | 40.8488 | 25.2868 | 2.006 | 32.0639 | 64.3187 | 15.9844 | 14.3333 | -0.1104 | 105.1675 | 235.1456 | 180.918 | 232.9917 | 161.1073
2012-10-30 18:26 | 28.4937 | 249.5065 | 195.2829 | 247.2743 | 175.6009 | -1.5991 | 1.271 | 32.1393 | 40.8484 | 25.2871 | 2.006 | 32.0639 | 64.319 | 15.9843 | 14.3173 | -0.1102 | 105.1675 | 235.1892 | 180.9656 | 232.957 | 161.2836
2012-10-30 18:26 | 28.4989 | 249.4804 | 195.2881 | 247.2403 | 175.5475 | -1.6037 | 1.271 | 32.1389 | 40.8475 | 25.287 | 2.0059 | 32.0635 | 64.317 | 15.9844 | 14.3021 | -0.1106 | 105.1645 | 235.1783 | 180.986 | 232.9382 | 161.2454
2012-10-30 18:26 | 28.5012 | 249.4905 | 195.2354 | 247.2583 | 175.5303 | -1.6055 | 1.271 | 32.1388 | 40.8477 | 25.2867 | 2.0059 | 32.0634 | 64.3169 | 15.9843 | 14.3248 | -0.1107 | 105.1646 | 235.1657 | 180.9106 | 232.9335 | 161.2055
2012-10-30 18:26 | 28.4997 | 249.5047 | 195.2378 | 247.3038 | 175.4898 | -1.602 | 1.2709 | 32.1376 | 40.8442 | 25.2869 | 2.0059 | 32.0622 | 64.3121 | 15.9843 | 14.3356 | -0.1104 | 105.1563 | 235.1691 | 180.9022 | 232.9682 | 161.1542
2012-10-30 18:26 | 28.5019 | 249.4951 | 195.1772 | 247.2668 | 175.6442 | -1.5974 | 1.2709 | 32.1366 | 40.8423 | 25.2866 | 2.0058 | 32.0611 | 64.308 | 15.9843 | 14.3623 | -0.1101 | 105.1503 | 235.1328 | 180.8149 | 232.9045 | 161.2819
2012-10-30 18:26 | 28.5102 | 249.5269 | 195.1933 | 247.2203 | 175.4417 | -1.6026 | 1.2709 | 32.1364 | 40.842 | 25.2865 | 2.0058 | 32.0611 | 64.3081 | 15.9842 | 14.3461 | -0.1105 | 105.1501 | 235.1808 | 180.8472 | 232.8742 | 161.0956
2012-10-30 18:26 | 28.5072 | 249.477 | 195.21 | 247.3309 | 175.388 | -1.5985 | 1.2709 | 32.1365 | 40.8414 | 25.2869 | 2.0058 | 32.061 | 64.3076 | 15.9842 | 14.3268 | -0.1102 | 105.149 | 235.1502 | 180.8832 | 233.0041 | 161.0612
2012-10-30 18:26 | 28.5147 | 249.5314 | 195.2214 | 247.3697 | 175.4111 | -1.5962 | 1.2709 | 32.1362 | 40.8408 | 25.2868 | 2.0058 | 32.0607 | 64.3064 | 15.9842 | 14.3343 | -0.11 | 105.1472 | 235.1971 | 180.8871 | 233.0354 | 161.0768
2012-10-30 18:26 | 28.5207 | 249.5335 | 195.196 | 247.317 | 175.4444 | -1.6043 | 1.2708 | 32.1355 | 40.8393 | 25.2867 | 2.0057 | 32.0602 | 64.3046 | 15.9841 | 14.3321 | -0.1106 | 105.1439 | 235.2014 | 180.8639 | 232.9849 | 161.1123
2012-10-30 18:26 | 28.5154 | 249.5635 | 195.1592 | 247.249 | 175.5055 | -1.6026 | 1.2708 | 32.1353 | 40.8385 | 25.2868 | 2.0057 | 32.0598 | 64.3032 | 15.9841 | 14.339 | -0.1105 | 105.1417 | 235.2245 | 180.8202 | 232.91 | 161.1665
2012-10-30 18:27 | 28.5124 | 249.5135 | 195.1641 | 247.3008 | 175.5103 | -1.5951 | 1.2708 | 32.1347 | 40.8374 | 25.2866 | 2.0057 | 32.0593 | 64.3012 | 15.9841 | 14.3401 | -0.11 | 105.1386 | 235.1734 | 180.824 | 232.9607 | 161.1702
2012-10-30 18:27 | 28.5124 | 249.4978 | 195.1759 | 247.3792 | 175.4869 | -1.5991 | 1.2708 | 32.1343 | 40.8366 | 25.2865 | 2.0057 | 32.059 | 64.2998 | 15.9842 | 14.3401 | -0.1102 | 105.1364 | 235.1577 | 180.8358 | 233.0391 | 161.1468
2012-10-30 18:27 | 28.5244 | 249.5607 | 195.1918 | 247.2345 | 175.4247 | -1.602 | 1.2708 | 32.1341 | 40.8354 | 25.2869 | 2.0057 | 32.0587 | 64.2989 | 15.9841 | 14.3195 | -0.1104 | 105.1342 | 235.2412 | 180.8723 | 232.915 | 161.1052
2012-10-30 18:27 | 28.5154 | 249.5361 | 195.1475 | 247.3195 | 175.4157 | -1.6055 | 1.2708 | 32.1336 | 40.8346 | 25.2866 | 2.0056 | 32.0582 | 64.2968 | 15.9841 | 14.3105 | -0.1107 | 105.1314 | 235.2256 | 180.837 | 233.009 | 161.1052
2012-10-30 18:27 | 28.5244 | 249.5451 | 195.1604 | 247.2697 | 175.3896 | -1.6049 | 1.2665 | 32.0242 | 40.5582 | 25.2859 | 1.9988 | 31.9493 | 63.8613 | 15.984 | 14.3236 | -0.1106 | 104.4195 | 235.2215 | 180.8368 | 232.9461 | 161.066
2012-10-30 18:27 | 28.5237 | 249.4935 | 195.1793 | 247.3278 | 175.4708 | -1.5997 | 1.2647 | 31.9782 | 40.4428 | 25.2853 | 1.9959 | 31.9035 | 63.6777 | 15.9842 | 14.3514 | -0.1103 | 104.1204 | 235.1421 | 180.8279 | 232.9764 | 161.1194
2012-10-30 18:27 | 28.5229 | 249.5005 | 195.1471 | 247.3348 | 175.5676 | -1.6014 | 1.2641 | 31.9624 | 40.4028 | 25.2853 | 1.9949 | 31.8875 | 63.6133 | 15.9843 | 14.3302 | -0.1104 | 104.0161 | 235.1703 | 180.8169 | 233.0046 | 161.2374
2012-10-30 18:27 | 28.5334 | 249.4679 | 195.1222 | 247.2944 | 175.2503 | -1.6009 | 1.2635 | 31.9476 | 40.3658 | 25.2851 | 1.994 | 31.8727 | 63.5543 | 15.9843 | 14.3366 | -0.1104 | 103.9201 | 235.1313 | 180.7856 | 232.9578 | 160.9137
2012-10-30 18:27 | 28.5259 | 249.4487 | 195.0283 | 247.236 | 174.7628 | -1.5957 | 1.2634 | 31.944 | 40.3567 | 25.285 | 1.9938 | 31.8691 | 63.5395 | 15.9844 | 14.3332 | -0.11 | 103.8963 | 235.1155 | 180.6951 | 232.9028 | 160.4296
2012-10-30 18:27 | 28.5297 | 249.4642 | 194.9496 | 247.373 | 174.8837 | -1.6009 | 1.2633 | 31.9425 | 40.3524 | 25.2854 | 1.9937 | 31.8676 | 63.5335 | 15.9844 | 14.3166 | -0.1104 | 103.8859 | 235.1476 | 180.633 | 233.0564 | 160.5671
2012-10-30 18:27 | 28.5326 | 249.4319 | 194.9643 | 247.2936 | 175.1324 | -1.5985 | 1.2634 | 31.9446 | 40.3582 | 25.285 | 1.9938 | 31.8697 | 63.5411 | 15.9846 | 14.3114 | -0.1102 | 103.8993 | 235.1205 | 180.6529 | 232.9822 | 160.821
2012-10-30 18:27 | 28.5341 | 249.3903 | 194.9658 | 247.2324 | 174.9934 | -1.6055 | 1.2635 | 31.9472 | 40.3645 | 25.2852 | 1.994 | 31.8722 | 63.5516 | 15.9844 | 14.3047 | -0.1107 | 103.9161 | 235.0856 | 180.6611 | 232.9277 | 160.6887
2012-10-30 18:27 | 28.5297 | 249.3703 | 194.9182 | 247.2946 | 174.5598 | -1.598 | 1.2636 | 31.9502 | 40.3724 | 25.285 | 1.9941 | 31.8752 | 63.5628 | 15.9847 | 14.3125 | -0.1102 | 103.9353 | 235.0578 | 180.6057 | 232.9821 | 160.2473
2012-10-30 18:27 | 28.5341 | 249.3982 | 194.844 | 247.2912 | 174.615 | -1.6032 | 1.2652 | 31.9924 | 40.4783 | 25.2856 | 1.9965 | 31.9134 | 63.7159 | 15.9845 | 14.3292 | -0.1105 | 104.1941 | 235.069 | 180.5148 | 232.962 | 160.2858
2012-10-30 18:27 | 28.5154 | 249.3834 | 194.7742 | 247.1941 | 175.0645 | -1.6032 | 1.2686 | 32.0786 | 40.6956 | 25.2862 | 2.0021 | 32.0034 | 64.0756 | 15.9845 | 14.3227 | -0.1105 | 104.7712 | 235.0607 | 180.4515 | 232.8714 | 160.7418
2012-10-30 18:27 | 28.5267 | 249.3868 | 194.8602 | 247.1428 | 175.2046 | -1.5997 | 1.2692 | 32.0946 | 40.7358 | 25.2864 | 2.0032 | 32.0195 | 64.1404 | 15.9845 | 14.3422 | -0.1103 | 104.8762 | 235.0446 | 180.518 | 232.8006 | 160.8624
2012-10-30 18:27 | 28.5222 | 249.3667 | 194.946 | 247.2205 | 175.4108 | -1.5968 | 1.2699 | 32.1109 | 40.7766 | 25.2868 | 2.0042 | 32.0357 | 64.2056 | 15.9844 | 14.3132 | -0.1101 | 104.9822 | 235.0535 | 180.6328 | 232.9073 | 161.0976
2012-10-30 18:27 | 28.5244 | 249.3571 | 194.9954 | 247.2384 | 175.3701 | -1.598 | 1.2703 | 32.1224 | 40.8056 | 25.2869 | 2.0049 | 32.0471 | 64.2511 | 15.9844 | 14.3195 | -0.1102 | 105.0566 | 235.0376 | 180.6759 | 232.9189 | 161.0506
2012-10-30 18:27 | 28.5244 | 249.3454 | 194.9914 | 247.2345 | 175.3076 | -1.5985 | 1.2706 | 32.1284 | 40.8209 | 25.2868 | 2.0053 | 32.0529 | 64.2751 | 15.9842 | 14.3603 | -0.1102 | 105.096 | 234.9851 | 180.6311 | 232.8742 | 160.9473
2012-10-30 18:27 | 28.5177 | 249.3935 | 195.0437 | 247.1534 | 175.379 | -1.602 | 1.2707 | 32.1319 | 40.8296 | 25.2869 | 2.0055 | 32.0565 | 64.2894 | 15.9843 | 14.3005 | -0.1104 | 105.119 | 235.093 | 180.7432 | 232.8529 | 161.0785
2012-10-30 18:27 | 28.5192 | 249.4067 | 194.9823 | 247.1235 | 175.3454 | -1.5974 | 1.2708 | 32.1345 | 40.8367 | 25.2868 | 2.0057 | 32.059 | 64.2996 | 15.9843 | 14.3224 | -0.1101 | 105.1363 | 235.0843 | 180.6599 | 232.8011 | 161.023
2012-10-30 18:27 | 28.5259 | 249.386 | 195.0597 | 247.2282 | 175.2584 | -1.6026 | 1.2709 | 32.1367 | 40.8417 | 25.2871 | 2.0058 | 32.0612 | 64.3082 | 15.9843 | 14.3455 | -0.1105 | 105.1499 | 235.0405 | 180.7142 | 232.8827 | 160.9129
2012-10-30 18:27 | 28.5199 | 249.3683 | 195.0537 | 247.2613 | 175.2563 | -1.6009 | 1.2709 | 32.1385 | 40.8462 | 25.2872 | 2.0059 | 32.063 | 64.3155 | 15.9842 | 14.3068 | -0.1104 | 105.1616 | 235.0615 | 180.7469 | 232.9545 | 160.9495
2012-10-30 18:27 | 28.5267 | 249.4103 | 195.0684 | 247.1193 | 175.3333 | -1.6032 | 1.271 | 32.1399 | 40.8498 | 25.287 | 2.006 | 32.0643 | 64.3211 | 15.9842 | 14.2891 | -0.1105 | 105.1709 | 235.1212 | 180.7793 | 232.8302 | 161.0442
2012-10-30 18:27 | 28.5237 | 249.423 | 195.0182 | 247.1398 | 175.2991 | -1.5991 | 1.271 | 32.1406 | 40.8514 | 25.2872 | 2.0061 | 32.0652 | 64.3245 | 15.9842 | 14.2943 | -0.1102 | 105.1759 | 235.1287 | 180.7239 | 232.8455 | 161.0048
2012-10-30 18:28 | 28.5297 | 249.4133 | 195.0478 | 247.1967 | 175.2388 | -1.602 | 1.2711 | 32.1412 | 40.8533 | 25.287 | 2.0061 | 32.0657 | 64.3267 | 15.9842 | 14.3248 | -0.1104 | 105.18 | 235.0885 | 180.723 | 232.8719 | 160.914
2012-10-30 18:28 | 28.5311 | 249.3638 | 195.0335 | 247.202 | 175.4861 | -1.5997 | 1.2711 | 32.1417 | 40.8541 | 25.2872 | 2.0061 | 32.0661 | 64.3287 | 15.9841 | 14.3262 | -0.1103 | 105.1828 | 235.0376 | 180.7073 | 232.8758 | 161.1599
2012-10-30 18:28 | 28.5371 | 249.4286 | 195.1142 | 247.2002 | 175.5116 | -1.5991 | 1.2711 | 32.1424 | 40.8563 | 25.287 | 2.0062 | 32.0668 | 64.3315 | 15.9841 | 14.3036 | -0.1102 | 105.1878 | 235.125 | 180.8106 | 232.8966 | 161.208
2012-10-30 18:28 | 28.5424 | 249.4378 | 195.0919 | 247.1232 | 175.4154 | -1.5968 | 1.2711 | 32.1424 | 40.8564 | 25.287 | 2.0062 | 32.0668 | 64.3323 | 15.9839 | 14.3171 | -0.1101 | 105.1887 | 235.1207 | 180.7748 | 232.8061 | 161.0983
2012-10-30 18:28 | 28.5394 | 249.3839 | 195.0732 | 247.1829 | 175.3617 | -1.6014 | 1.2711 | 32.1429 | 40.8575 | 25.287 | 2.0062 | 32.0674 | 64.3338 | 15.9841 | 14.3141 | -0.1104 | 105.1912 | 235.0698 | 180.7591 | 232.8688 | 161.0476
2012-10-30 18:28 | 28.5394 | 249.3761 | 195.0732 | 247.2808 | 175.4358 | -1.6043 | 1.2711 | 32.1428 | 40.8571 | 25.2871 | 2.0062 | 32.0673 | 64.3339 | 15.984 | 14.3182 | -0.1106 | 105.1911 | 235.0579 | 180.755 | 232.9626 | 161.1176
2012-10-30 18:28 | 28.5319 | 249.4077 | 195.0814 | 247.1793 | 175.4634 | -1.5997 | 1.2711 | 32.1427 | 40.8566 | 25.2874 | 2.0062 | 32.0672 | 64.3332 | 15.984 | 14.2821 | -0.1103 | 105.1898 | 235.1256 | 180.7993 | 232.8972 | 161.1813
2012-10-30 18:28 | 28.5364 | 249.3848 | 195.1331 | 247.1094 | 175.5616 | -1.6049 | 1.2711 | 32.1429 | 40.8571 | 25.2873 | 2.0062 | 32.0674 | 64.3342 | 15.984 | 14.3192 | -0.1106 | 105.1914 | 235.0656 | 180.8139 | 232.7902 | 161.2424
2012-10-30 18:28 | 28.5356 | 249.4075 | 195.1048 | 247.2261 | 175.5023 | -1.5991 | 1.2711 | 32.1429 | 40.8573 | 25.2872 | 2.0062 | 32.0674 | 64.3342 | 15.984 | 14.3348 | -0.1102 | 105.1915 | 235.0727 | 180.77 | 232.8913 | 161.1675
2012-10-30 18:28 | 28.5394 | 249.3956 | 195.1125 | 247.2612 | 175.5295 | -1.5974 | 1.2711 | 32.1427 | 40.8562 | 25.2875 | 2.0062 | 32.0672 | 64.3338 | 15.9839 | 14.3222 | -0.1101 | 105.19 | 235.0734 | 180.7903 | 232.939 | 161.2073
2012-10-30 18:28 | 28.5379 | 249.4059 | 195.166 | 247.1814 | 175.5163 | -1.6032 | 1.2711 | 32.1427 | 40.8569 | 25.2871 | 2.0062 | 32.0673 | 64.3342 | 15.9839 | 14.3615 | -0.1105 | 105.1912 | 235.0444 | 180.8045 | 232.8199 | 161.1548
2012-10-30 18:28 | 28.5439 | 249.4863 | 195.1367 | 247.1286 | 175.5457 | -1.5991 | 1.2711 | 32.1427 | 40.8569 | 25.2871 | 2.0062 | 32.0672 | 64.3335 | 15.984 | 14.3431 | -0.1102 | 105.1904 | 235.1432 | 180.7936 | 232.7855 | 161.2026
2012-10-30 18:28 | 28.5424 | 249.4378 | 195.1587 | 247.229 | 175.4974 | -1.5905 | 1.2711 | 32.1426 | 40.8568 | 25.287 | 2.0062 | 32.0671 | 64.3336 | 15.9839 | 14.3416 | -0.1096 | 105.1903 | 235.0962 | 180.8171 | 232.8874 | 161.1558
2012-10-30 18:28 | 28.5409 | 249.401 | 195.1494 | 247.2902 | 175.5817 | -1.6055 | 1.2711 | 32.1429 | 40.857 | 25.2874 | 2.0062 | 32.0674 | 64.3348 | 15.9838 | 14.336 | -0.1107 | 105.1919 | 235.065 | 180.8134 | 232.9542 | 161.2457
2012-10-30 18:28 | 28.5386 | 249.4379 | 195.1824 | 247.1586 | 175.474 | -1.5997 | 1.2711 | 32.1427 | 40.8567 | 25.2872 | 2.0062 | 32.0671 | 64.3337 | 15.9838 | 14.3582 | -0.1103 | 105.1904 | 235.0797 | 180.8242 | 232.8004 | 161.1158
2012-10-30 18:28 | 28.5521 | 249.4122 | 195.1331 | 247.2348 | 175.3509 | -1.5951 | 1.2711 | 32.1425 | 40.8566 | 25.287 | 2.0062 | 32.067 | 64.3332 | 15.9839 | 14.3227 | -0.11 | 105.1898 | 235.0895 | 180.8104 | 232.9121 | 161.0282
2012-10-30 18:28 | 28.5446 | 249.4204 | 195.1767 | 247.2625 | 175.363 | -1.5997 | 1.2711 | 32.1424 | 40.8565 | 25.2868 | 2.0062 | 32.067 | 64.3334 | 15.9838 | 14.3315 | -0.1103 | 105.1899 | 235.0889 | 180.8452 | 232.931 | 161.0315
2012-10-30 18:28 | 28.5446 | 249.4126 | 195.1727 | 247.2508 | 175.601 | -1.5968 | 1.2711 | 32.1422 | 40.8557 | 25.2871 | 2.0062 | 32.0667 | 64.3328 | 15.9837 | 14.3519 | -0.1101 | 105.1885 | 235.0607 | 180.8208 | 232.8989 | 161.2491
2012-10-30 18:28 | 28.5536 | 249.4686 | 195.1503 | 247.1618 | 175.6452 | -1.6032 | 1.2711 | 32.1419 | 40.8556 | 25.2866 | 2.0062 | 32.0664 | 64.3315 | 15.9836 | 14.3487 | -0.1105 | 105.1871 | 235.1199 | 180.8016 | 232.8131 | 161.2965
2012-10-30 18:28 | 28.5371 | 249.4325 | 195.1456 | 247.2237 | 175.4725 | -1.602 | 1.2711 | 32.1418 | 40.8551 | 25.2869 | 2.0062 | 32.0664 | 64.3315 | 15.9837 | 14.3159 | -0.1104 | 105.1866 | 235.1166 | 180.8297 | 232.9078 | 161.1566
2012-10-30 18:28 | 28.5499 | 249.4257 | 195.1859 | 247.3188 | 175.4502 | -1.6032 | 1.2711 | 32.1415 | 40.8543 | 25.2869 | 2.0062 | 32.0661 | 64.3305 | 15.9836 | 14.2838 | -0.1105 | 105.1847 | 235.1419 | 180.9021 | 233.035 | 161.1664
2012-10-30 18:28 | 28.5431 | 249.462 | 195.1634 | 247.214 | 175.4512 | -1.6014 | 1.2711 | 32.1412 | 40.8536 | 25.2869 | 2.0062 | 32.0658 | 64.3292 | 15.9836 | 14.3259 | -0.1104 | 105.1828 | 235.1361 | 180.8375 | 232.8881 | 161.1253
2012-10-30 18:28 | 28.5484 | 249.4438 | 195.1805 | 247.2428 | 175.6439 | -1.6003 | 1.271 | 32.1408 | 40.8523 | 25.287 | 2.0062 | 32.0655 | 64.3285 | 15.9835 | 14.3516 | -0.1103 | 105.1808 | 235.0922 | 180.8289 | 232.8912 | 161.2923
2012-10-30 18:28 | 28.5506 | 249.4773 | 195.218 | 247.2293 | 175.5719 | -1.6037 | 1.271 | 32.1408 | 40.8525 | 25.2869 | 2.0061 | 32.0653 | 64.3278 | 15.9835 | 14.3702 | -0.1106 | 105.1803 | 235.1071 | 180.8478 | 232.8591 | 161.2017
2012-10-30 18:28 | 28.5476 | 249.4469 | 195.1993 | 247.2303 | 175.4362 | -1.6003 | 1.271 | 32.1404 | 40.8518 | 25.2866 | 2.0061 | 32.0649 | 64.3262 | 15.9835 | 14.3427 | -0.1103 | 105.1779 | 235.1042 | 180.8566 | 232.8876 | 161.0935
2012-10-30 18:29 | 28.5416 | 249.4409 | 195.1776 | 247.2713 | 175.3678 | -1.6003 | 1.271 | 32.14 | 40.8501 | 25.2871 | 2.0061 | 32.0646 | 64.3248 | 15.9836 | 14.304 | -0.1103 | 105.1749 | 235.1369 | 180.8736 | 232.9673 | 161.0638
2012-10-30 18:29 | 28.5484 | 249.4555 | 195.1922 | 247.3368 | 175.3433 | -1.598 | 1.271 | 32.1401 | 40.8509 | 25.2868 | 2.0061 | 32.0647 | 64.3252 | 15.9835 | 14.3639 | -0.1102 | 105.1761 | 235.0916 | 180.8283 | 232.9729 | 160.9794
2012-10-30 18:29 | 28.5446 | 249.5104 | 195.1688 | 247.1803 | 175.3395 | -1.5985 | 1.271 | 32.1401 | 40.8503 | 25.2871 | 2.0061 | 32.0647 | 64.3259 | 15.9834 | 14.3234 | -0.1102 | 105.1762 | 235.187 | 180.8454 | 232.8569 | 161.0161
2012-10-30 18:29 | 28.5514 | 249.4859 | 195.1402 | 247.2654 | 175.2175 | -1.5945 | 1.271 | 32.1399 | 40.8498 | 25.2871 | 2.0061 | 32.0645 | 64.3247 | 15.9835 | 14.3546 | -0.1099 | 105.1744 | 235.1313 | 180.7856 | 232.9108 | 160.8629
2012-10-30 18:29 | 28.5454 | 249.4643 | 195.0832 | 247.3299 | 175.2779 | -1.6014 | 1.271 | 32.1391 | 40.848 | 25.287 | 2.0061 | 32.0638 | 64.3219 | 15.9835 | 14.3364 | -0.1104 | 105.1698 | 235.1279 | 180.7468 | 232.9935 | 160.9415
2012-10-30 18:29 | 28.5454 | 249.476 | 195.1107 | 247.2359 | 175.2779 | -1.6032 | 1.271 | 32.139 | 40.8475 | 25.2872 | 2.006 | 32.0635 | 64.3209 | 15.9834 | 14.3201 | -0.1105 | 105.1684 | 235.1559 | 180.7906 | 232.9158 | 160.9578
2012-10-30 18:29 | 28.5454 | 249.4799 | 195.1146 | 247.2672 | 175.2272 | -1.6009 | 1.271 | 32.1384 | 40.8463 | 25.2869 | 2.006 | 32.063 | 64.3191 | 15.9833 | 14.3568 | -0.1104 | 105.1655 | 235.1231 | 180.7578 | 232.9104 | 160.8704
2012-10-30 18:29 | 28.5499 | 249.4923 | 195.0837 | 247.1856 | 175.3136 | -1.5933 | 1.2709 | 32.1381 | 40.8454 | 25.2871 | 2.006 | 32.0626 | 64.3178 | 15.9833 | 14.3287 | -0.1098 | 105.1632 | 235.1636 | 180.755 | 232.8569 | 160.9849
2012-10-30 18:29 | 28.5476 | 249.4782 | 195.0775 | 247.2694 | 175.1045 | -1.6026 | 1.2709 | 32.1377 | 40.8445 | 25.2869 | 2.006 | 32.0622 | 64.3157 | 15.9834 | 14.3264 | -0.1105 | 105.1602 | 235.1518 | 180.7511 | 232.943 | 160.7781
2012-10-30 18:29 | 28.5581 | 249.4809 | 195.0409 | 247.3152 | 175.2164 | -1.5962 | 1.2709 | 32.1373 | 40.8433 | 25.287 | 2.006 | 32.0618 | 64.3144 | 15.9833 | 14.3328 | -0.11 | 105.1577 | 235.1481 | 180.7081 | 232.9824 | 160.8836
2012-10-30 18:29 | 28.5566 | 249.5068 | 195.0786 | 247.2863 | 175.375 | -1.6037 | 1.2709 | 32.1369 | 40.8421 | 25.2872 | 2.0059 | 32.0614 | 64.3127 | 15.9834 | 14.2986 | -0.1106 | 105.1548 | 235.2082 | 180.78 | 232.9877 | 161.0764
2012-10-30 18:29 | 28.5491 | 249.5267 | 195.0319 | 247.2004 | 175.3089 | -1.602 | 1.2709 | 32.1366 | 40.8418 | 25.2868 | 2.0059 | 32.0612 | 64.3121 | 15.9833 | 14.3034 | -0.1104 | 105.1539 | 235.2233 | 180.7285 | 232.897 | 161.0055
2012-10-30 18:29 | 28.5521 | 249.5101 | 195.0231 | 247.2387 | 175.3158 | -1.6032 | 1.2709 | 32.1363 | 40.841 | 25.287 | 2.0059 | 32.0609 | 64.3108 | 15.9833 | 14.3227 | -0.1105 | 105.1518 | 235.1874 | 180.7004 | 232.916 | 160.9931
2012-10-30 18:29 | 28.5619 | 249.5003 | 195.0604 | 247.3621 | 175.2671 | -1.6049 | 1.2709 | 32.1362 | 40.8407 | 25.2869 | 2.0059 | 32.0607 | 64.3104 | 15.9833 | 14.3203 | -0.1106 | 105.1511 | 235.18 | 180.7401 | 233.0418 | 160.9468
2012-10-30 18:29 | 28.5559 | 249.5491 | 195.0308 | 247.2072 | 175.4133 | -1.602 | 1.2708 | 32.1356 | 40.8388 | 25.2871 | 2.0059 | 32.0601 | 64.3082 | 15.9832 | 14.3143 | -0.1104 | 105.147 | 235.2348 | 180.7165 | 232.8929 | 161.099
2012-10-30 18:29 | 28.5521 | 249.5062 | 195.0074 | 247.27 | 175.3002 | -1.6009 | 1.2708 | 32.1356 | 40.8389 | 25.2871 | 2.0059 | 32.0602 | 64.3084 | 15.9832 | 14.3023 | -0.1104 | 105.1473 | 235.2039 | 180.7051 | 232.9677 | 160.9979
2012-10-30 18:29 | 28.5529 | 249.4992 | 195.0553 | 247.2943 | 175.3049 | -1.6049 | 1.2708 | 32.1354 | 40.8387 | 25.2869 | 2.0059 | 32.06 | 64.3079 | 15.9831 | 14.348 | -0.1106 | 105.1466 | 235.1512 | 180.7073 | 232.9463 | 160.9569
2012-10-30 18:29 | 28.5566 | 249.5342 | 195.0786 | 247.3646 | 175.2501 | -1.6049 | 1.2708 | 32.135 | 40.8374 | 25.287 | 2.0058 | 32.0596 | 64.3063 | 15.9832 | 14.3394 | -0.1106 | 105.1437 | 235.1948 | 180.7392 | 233.0252 | 160.9107
2012-10-30 18:29 | 28.5521 | 249.5023 | 195.0427 | 247.2465 | 175.4251 | -1.6037 | 1.2708 | 32.1347 | 40.8369 | 25.2869 | 2.0058 | 32.0592 | 64.3047 | 15.9832 | 14.3186 | -0.1106 | 105.1416 | 235.1837 | 180.7241 | 232.9279 | 161.1065
2012-10-30 18:29 | 28.5641 | 249.5378 | 195.0429 | 247.2585 | 175.2966 | -1.598 | 1.2708 | 32.1348 | 40.8371 | 25.2869 | 2.0058 | 32.0593 | 64.305 | 15.9832 | 14.3429 | -0.1102 | 105.1421 | 235.1949 | 180.7 | 232.9156 | 160.9537
2012-10-30 18:29 | 28.5604 | 249.4949 | 195.051 | 247.3293 | 175.3827 | -1.6078 | 1.2696 | 32.1043 | 40.7599 | 25.2867 | 2.0041 | 32.0313 | 64.1928 | 15.9832 | 14.331 | -0.1108 | 104.9527 | 235.1639 | 180.72 | 232.9983 | 161.0517
2012-10-30 18:29 | 28.5596 | 249.4667 | 195.0934 | 247.348 | 175.4989 | -1.6026 | 1.2646 | 31.9754 | 40.4356 | 25.2853 | 1.9958 | 31.8994 | 63.6656 | 15.9831 | 14.3424 | -0.1105 | 104.1012 | 235.1243 | 180.751 | 233.0056 | 161.1565
2012-10-30 18:29 | 28.5566 | 249.5068 | 195.0786 | 247.2432 | 175.5389 | -1.5997 | 1.2646 | 31.9751 | 40.4348 | 25.2853 | 1.9959 | 31.9005 | 63.6693 | 15.9832 | 14.2823 | -0.1103 | 104.1042 | 235.2245 | 180.7963 | 232.9609 | 161.2566
2012-10-30 18:29 | 28.5604 | 249.4832 | 195.0942 | 247.2901 | 175.4997 | -1.6032 | 1.264 | 31.9603 | 40.3979 | 25.285 | 1.9949 | 31.8855 | 63.6097 | 15.9832 | 14.3269 | -0.1105 | 104.0076 | 235.1563 | 180.7673 | 232.9632 | 161.1728
2012-10-30 18:29 | 28.5544 | 249.4537 | 195.1118 | 247.2997 | 175.6342 | -1.6009 | 1.2636 | 31.949 | 40.3694 | 25.285 | 1.9942 | 31.874 | 63.5633 | 15.9833 | 14.374 | -0.1104 | 103.9328 | 235.0797 | 180.7378 | 232.9257 | 161.2602
2012-10-30 18:29 | 28.5529 | 249.4287 | 195.1103 | 247.2904 | 175.6874 | -1.6009 | 1.2634 | 31.9458 | 40.3614 | 25.2848 | 1.994 | 31.8708 | 63.5509 | 15.9833 | 14.3725 | -0.1104 | 103.9122 | 235.0562 | 180.7378 | 232.9179 | 161.3149
2012-10-30 18:29 | 28.5356 | 249.4506 | 195.1519 | 247.2653 | 175.5023 | -1.6066 | 1.2635 | 31.9464 | 40.3628 | 25.285 | 1.994 | 31.8714 | 63.5527 | 15.9834 | 14.3511 | -0.1108 | 103.9155 | 235.0995 | 180.8008 | 232.9142 | 161.1512
2012-10-30 18:30 | 28.5544 | 249.4341 | 195.1197 | 247.2175 | 175.5484 | -1.602 | 1.2634 | 31.9462 | 40.3623 | 25.285 | 1.994 | 31.8713 | 63.5519 | 15.9834 | 14.3862 | -0.1104 | 103.9142 | 235.0479 | 180.7335 | 232.8313 | 161.1622
2012-10-30 18:30 | 28.5566 | 249.4363 | 195.1376 | 247.3098 | 175.6169 | -1.6037 | 1.2635 | 31.9471 | 40.3647 | 25.2849 | 1.9941 | 31.872 | 63.5547 | 15.9835 | 14.3476 | -0.1106 | 103.9193 | 235.0887 | 180.79 | 232.9622 | 161.2693
2012-10-30 18:30 | 28.5431 | 249.3993 | 195.1712 | 247.2963 | 175.6971 | -1.6014 | 1.2636 | 31.9501 | 40.3721 | 25.285 | 1.9942 | 31.875 | 63.5664 | 15.9836 | 14.3463 | -0.1104 | 103.9385 | 235.053 | 180.8249 | 232.95 | 161.3508
2012-10-30 18:30 | 28.5506 | 249.399 | 195.1669 | 247.245 | 175.6031 | -1.5951 | 1.2648 | 31.9814 | 40.4508 | 25.2854 | 1.9962 | 31.9066 | 63.6924 | 15.9836 | 14.3702 | -0.11 | 104.1431 | 235.0288 | 180.7967 | 232.8748 | 161.2329
2012-10-30 18:30 | 28.5371 | 249.3855 | 195.1967 | 247.1767 | 175.3125 | -1.5928 | 1.2686 | 32.0779 | 40.6944 | 25.2858 | 2.0022 | 32.0024 | 64.0757 | 15.9835 | 14.3444 | -0.1098 | 104.7701 | 235.0411 | 180.8523 | 232.8323 | 160.9681
2012-10-30 18:30 | 28.5304 | 249.3788 | 195.1703 | 247.1817 | 175.2122 | -1.6014 | 1.2692 | 32.0932 | 40.732 | 25.2866 | 2.0032 | 32.0178 | 64.1375 | 15.9835 | 14.3663 | -0.1104 | 104.8695 | 235.0125 | 180.804 | 232.8154 | 160.8459
2012-10-30 18:30 | 28.5356 | 249.4153 | 195.1873 | 247.2418 | 175.4047 | -1.5962 | 1.2697 | 32.1078 | 40.7686 | 25.2869 | 2.0041 | 32.0327 | 64.1976 | 15.9834 | 14.396 | -0.11 | 104.9662 | 235.0193 | 180.7913 | 232.8458 | 161.0087
2012-10-30 18:30 | 28.5401 | 249.3924 | 195.1839 | 247.2776 | 175.7488 | -1.5968 | 1.2702 | 32.1193 | 40.7981 | 25.2867 | 2.0048 | 32.0439 | 64.2426 | 15.9833 | 14.376 | -0.1101 | 105.0407 | 235.0164 | 180.8079 | 232.9016 | 161.3728
2012-10-30 18:30 | 28.5326 | 249.4123 | 195.255 | 247.2231 | 175.8388 | -1.6049 | 1.2705 | 32.1267 | 40.8164 | 25.287 | 2.0053 | 32.0512 | 64.2719 | 15.9834 | 14.3644 | -0.1106 | 105.0883 | 235.0479 | 180.8906 | 232.8587 | 161.4744
2012-10-30 18:30 | 28.5244 | 249.4002 | 195.2665 | 247.1483 | 175.8072 | -1.6026 | 1.2707 | 32.1312 | 40.828 | 25.287 | 2.0056 | 32.0558 | 64.2904 | 15.9833 | 14.3766 | -0.1105 | 105.1184 | 235.0236 | 180.8899 | 232.7717 | 161.4306
2012-10-30 18:30 | 28.5341 | 249.406 | 195.2722 | 247.1032 | 175.7896 | -1.6043 | 1.2708 | 32.1345 | 40.8361 | 25.2871 | 2.0058 | 32.0591 | 64.3036 | 15.9833 | 14.4108 | -0.1106 | 105.1397 | 234.9952 | 180.8614 | 232.6924 | 161.3788
2012-10-30 18:30 | 28.5304 | 249.4023 | 195.2803 | 247.2679 | 175.7508 | -1.6043 | 1.2709 | 32.1366 | 40.8416 | 25.287 | 2.0059 | 32.0611 | 64.3119 | 15.9832 | 14.3908 | -0.1106 | 105.1536 | 235.0115 | 180.8895 | 232.8771 | 161.36
2012-10-30 18:30 | 28.5244 | 249.4276 | 195.3293 | 247.2384 | 175.7994 | -1.5985 | 1.2709 | 32.1379 | 40.8451 | 25.2869 | 2.006 | 32.0625 | 64.3173 | 15.9833 | 14.3521 | -0.1102 | 105.1624 | 235.0755 | 180.9772 | 232.8863 | 161.4473
2012-10-30 18:30 | 28.5289 | 249.4752 | 195.3653 | 247.1293 | 175.9483 | -1.6061 | 1.2709 | 32.1386 | 40.8463 | 25.2872 | 2.006 | 32.0631 | 64.32 | 15.9832 | 14.3525 | -0.1107 | 105.1663 | 235.1227 | 181.0128 | 232.7768 | 161.5958
2012-10-30 18:30 | 28.5297 | 249.4564 | 195.3582 | 247.2163 | 175.8242 | -1.6032 | 1.271 | 32.1392 | 40.8478 | 25.2872 | 2.0061 | 32.0637 | 64.3227 | 15.9832 | 14.3656 | -0.1105 | 105.1705 | 235.0908 | 180.9926 | 232.8507 | 161.4586
2012-10-30 18:30 | 28.5192 | 249.4224 | 195.3556 | 247.2058 | 175.7864 | -1.6032 | 1.271 | 32.1406 | 40.8513 | 25.2873 | 2.0062 | 32.065 | 64.3277 | 15.9833 | 14.3714 | -0.1105 | 105.179 | 235.051 | 180.9842 | 232.8344 | 161.415
2012-10-30 18:30 | 28.5214 | 249.4246 | 195.4089 | 247.2472 | 175.8198 | -1.6009 | 1.271 | 32.1409 | 40.8521 | 25.2872 | 2.0062 | 32.0654 | 64.3294 | 15.9832 | 14.3695 | -0.1104 | 105.1816 | 235.0551 | 181.0394 | 232.8777 | 161.4503
2012-10-30 18:30 | 28.5139 | 249.4876 | 195.3699 | 247.1652 | 175.7616 | -1.6043 | 1.271 | 32.1414 | 40.853 | 25.2874 | 2.0062 | 32.0659 | 64.3313 | 15.9832 | 14.3661 | -0.1106 | 105.1843 | 235.1215 | 181.0038 | 232.7991 | 161.3955
2012-10-30 18:30 | 28.5199 | 249.4701 | 195.4074 | 247.1399 | 175.7832 | -1.6061 | 1.2711 | 32.1417 | 40.8541 | 25.2873 | 2.0062 | 32.0661 | 64.3323 | 15.9832 | 14.3762 | -0.1107 | 105.1864 | 235.0939 | 181.0312 | 232.7637 | 161.407
2012-10-30 18:30 | 28.5282 | 249.451 | 195.3724 | 247.2344 | 175.612 | -1.6032 | 1.2711 | 32.1419 | 40.8548 | 25.2871 | 2.0063 | 32.0663 | 64.3332 | 15.9832 | 14.3763 | -0.1105 | 105.188 | 235.0747 | 180.9961 | 232.8581 | 161.2357
2012-10-30 18:30 | 28.5184 | 249.4177 | 195.398 | 247.2716 | 175.6646 | -1.602 | 1.2711 | 32.142 | 40.855 | 25.2872 | 2.0063 | 32.0664 | 64.3343 | 15.983 | 14.3543 | -0.1104 | 105.1894 | 235.0634 | 181.0437 | 232.9173 | 161.3103
2012-10-30 18:30 | 28.5192 | 249.4772 | 195.3674 | 247.1392 | 175.6888 | -1.5922 | 1.2711 | 32.142 | 40.8552 | 25.287 | 2.0063 | 32.0664 | 64.3344 | 15.9829 | 14.3632 | -0.1098 | 105.1896 | 235.114 | 181.0042 | 232.776 | 161.3256
2012-10-30 18:30 | 28.5184 | 249.5116 | 195.3273 | 247.2011 | 175.5475 | -1.6061 | 1.2711 | 32.1424 | 40.8558 | 25.2874 | 2.0063 | 32.0669 | 64.3359 | 15.9831 | 14.3706 | -0.1107 | 105.1916 | 235.141 | 180.9567 | 232.8305 | 161.1769
2012-10-30 18:30 | 28.5177 | 249.5031 | 195.3148 | 247.3061 | 175.5897 | -1.6072 | 1.2711 | 32.1425 | 40.8561 | 25.2873 | 2.0063 | 32.067 | 64.3357 | 15.9832 | 14.3577 | -0.1108 | 105.1918 | 235.1454 | 180.9571 | 232.9484 | 161.232
2012-10-30 18:30 | 28.5282 | 249.4901 | 195.3331 | 247.2461 | 175.6744 | -1.5997 | 1.2711 | 32.1427 | 40.8561 | 25.2876 | 2.0063 | 32.0671 | 64.3372 | 15.983 | 14.36 | -0.1103 | 105.1932 | 235.1301 | 180.9731 | 232.8861 | 161.3144
2012-10-30 18:31 | 28.5214 | 249.5264 | 195.3578 | 247.1414 | 175.7301 | -1.602 | 1.2711 | 32.1429 | 40.8568 | 25.2875 | 2.0063 | 32.0673 | 64.3377 | 15.983 | 14.3491 | -0.1104 | 105.1945 | 235.1773 | 181.0087 | 232.7923 | 161.381
2012-10-30 18:31 | 28.5341 | 249.4921 | 195.3155 | 247.299 | 175.731 | -1.602 | 1.2711 | 32.1425 | 40.8557 | 25.2876 | 2.0063 | 32.067 | 64.3369 | 15.9829 | 14.3618 | -0.1104 | 105.1926 | 235.1303 | 180.9537 | 232.9372 | 161.3692
2012-10-30 18:31 | 28.5184 | 249.496 | 195.3587 | 247.2598 | 175.7817 | -1.5974 | 1.2711 | 32.1421 | 40.8551 | 25.2873 | 2.0063 | 32.0666 | 64.3357 | 15.9828 | 14.3257 | -0.1101 | 105.1907 | 235.1703 | 181.033 | 232.9341 | 161.456
2012-10-30 18:31 | 28.5252 | 249.4989 | 195.4402 | 247.2823 | 175.8393 | -1.6026 | 1.2711 | 32.1421 | 40.8546 | 25.2875 | 2.0063 | 32.0665 | 64.335 | 15.9829 | 14.3366 | -0.1105 | 105.1896 | 235.1623 | 181.1036 | 232.9457 | 161.5027
2012-10-30 18:31 | 28.5192 | 249.5438 | 195.3792 | 247.1157 | 175.7942 | -1.5985 | 1.2711 | 32.1422 | 40.8556 | 25.2871 | 2.0063 | 32.0667 | 64.3362 | 15.9828 | 14.3388 | -0.1102 | 105.1918 | 235.205 | 181.0404 | 232.7769 | 161.4554
2012-10-30 18:31 | 28.5237 | 249.5169 | 195.3679 | 247.2965 | 175.7363 | -1.5997 | 1.2711 | 32.1423 | 40.8551 | 25.2876 | 2.0063 | 32.0667 | 64.3364 | 15.9828 | 14.3759 | -0.1103 | 105.1915 | 235.141 | 180.992 | 232.9206 | 161.3604
2012-10-30 18:31 | 28.5199 | 249.4897 | 195.427 | 247.2809 | 175.8808 | -1.5968 | 1.2711 | 32.1421 | 40.8551 | 25.2873 | 2.0063 | 32.0666 | 64.3361 | 15.9827 | 14.3476 | -0.1101 | 105.1912 | 235.1421 | 181.0794 | 232.9333 | 161.5332
2012-10-30 18:31 | 28.5229 | 249.5592 | 195.4457 | 247.2526 | 175.8604 | -1.6014 | 1.2711 | 32.1417 | 40.8543 | 25.2872 | 2.0063 | 32.0662 | 64.3344 | 15.9828 | 14.3669 | -0.1104 | 105.1887 | 235.1923 | 181.0788 | 232.8857 | 161.4935
2012-10-30 18:31 | 28.5267 | 249.5356 | 195.4535 | 247.2211 | 175.708 | -1.5991 | 1.2711 | 32.1416 | 40.8543 | 25.2871 | 2.0063 | 32.0662 | 64.3344 | 15.9828 | 14.3626 | -0.1102 | 105.1886 | 235.173 | 181.0909 | 232.8585 | 161.3454
2012-10-30 18:31 | 28.5289 | 249.5104 | 195.381 | 247.3291 | 175.6907 | -1.6032 | 1.2711 | 32.1414 | 40.8536 | 25.287 | 2.0063 | 32.0659 | 64.3333 | 15.9827 | 14.3689 | -0.1105 | 105.1869 | 235.1415 | 181.0121 | 232.9602 | 161.3218
2012-10-30 18:31 | 28.5229 | 249.524 | 195.4575 | 247.2604 | 175.8213 | -1.6009 | 1.271 | 32.1411 | 40.8528 | 25.2872 | 2.0063 | 32.0657 | 64.3328 | 15.9827 | 14.3751 | -0.1104 | 105.1856 | 235.1489 | 181.0824 | 232.8853 | 161.4462
2012-10-30 18:31 | 28.5274 | 249.5598 | 195.4149 | 247.2375 | 175.7439 | -1.6043 | 1.271 | 32.1409 | 40.8525 | 25.287 | 2.0063 | 32.0654 | 64.3314 | 15.9827 | 14.4245 | -0.1106 | 105.1839 | 235.1353 | 180.9904 | 232.813 | 161.3194
2012-10-30 18:31 | 28.5297 | 249.5308 | 195.4093 | 247.3495 | 175.8164 | -1.6032 | 1.271 | 32.1406 | 40.8515 | 25.2872 | 2.0062 | 32.0652 | 64.3305 | 15.9827 | 14.3982 | -0.1105 | 105.182 | 235.1326 | 181.0111 | 232.9513 | 161.4182
2012-10-30 18:31 | 28.5319 | 249.5565 | 195.4744 | 247.3164 | 175.803 | -1.6026 | 1.271 | 32.1389 | 40.8471 | 25.2872 | 2.0062 | 32.0634 | 64.3241 | 15.9825 | 14.4208 | -0.1105 | 105.1712 | 235.1357 | 181.0536 | 232.8956 | 161.3822
2012-10-30 18:31 | 28.5356 | 249.5758 | 195.4388 | 247.2809 | 175.756 | -1.5991 | 1.2709 | 32.1372 | 40.8426 | 25.2872 | 2.0061 | 32.0616 | 64.3178 | 15.9823 | 14.4531 | -0.1102 | 105.1604 | 235.1227 | 180.9857 | 232.8278 | 161.3029
2012-10-30 18:31 | 28.5401 | 249.5647 | 195.4276 | 247.3207 | 175.5341 | -1.6032 | 1.2709 | 32.1365 | 40.8408 | 25.2874 | 2.006 | 32.061 | 64.3155 | 15.9823 | 14.4086 | -0.1105 | 105.1563 | 235.1561 | 181.019 | 232.9121 | 161.1255
2012-10-30 18:32 | 28.5364 | 249.5492 | 195.3689 | 247.3758 | 175.6475 | -1.6032 | 1.2709 | 32.1362 | 40.8404 | 25.2871 | 2.006 | 32.0607 | 64.3139 | 15.9824 | 14.4008 | -0.1105 | 105.1543 | 235.1484 | 180.9681 | 232.975 | 161.2467
2012-10-30 18:32 | 28.5446 | 249.5809 | 195.4203 | 247.3017 | 175.8508 | -1.5951 | 1.2708 | 32.1358 | 40.8393 | 25.2872 | 2.006 | 32.0603 | 64.3125 | 15.9823 | 14.405 | -0.11 | 105.1517 | 235.1759 | 181.0153 | 232.8967 | 161.4458
2012-10-30 18:32 | 28.5461 | 249.5942 | 195.4021 | 247.3345 | 175.7548 | -1.6037 | 1.2708 | 32.1358 | 40.8397 | 25.2869 | 2.006 | 32.0605 | 64.3131 | 15.9824 | 14.4146 | -0.1106 | 105.1528 | 235.1796 | 180.9875 | 232.9199 | 161.3402
2012-10-30 18:32 | 28.5521 | 249.5923 | 195.4278 | 247.4698 | 175.7295 | -1.598 | 1.2708 | 32.1356 | 40.8385 | 25.2873 | 2.006 | 32.06 | 64.3115 | 15.9823 | 14.388 | -0.1102 | 105.15 | 235.2043 | 181.0398 | 233.0818 | 161.3415
2012-10-30 18:32 | 28.5499 | 249.6058 | 195.4688 | 247.3697 | 175.7625 | -1.6014 | 1.2708 | 32.135 | 40.8371 | 25.2873 | 2.0059 | 32.0595 | 64.3097 | 15.9822 | 14.4021 | -0.1104 | 105.1468 | 235.2037 | 181.0667 | 232.9676 | 161.3604
2012-10-30 18:32 | 28.5506 | 249.5556 | 195.4027 | 247.3234 | 175.7788 | -1.6049 | 1.2708 | 32.1348 | 40.8369 | 25.2871 | 2.0059 | 32.0593 | 64.3085 | 15.9823 | 14.3742 | -0.1106 | 105.1454 | 235.1814 | 181.0285 | 232.9492 | 161.4046
2012-10-30 18:32 | 28.5446 | 249.5496 | 195.4163 | 247.4506 | 175.8587 | -1.6072 | 1.2708 | 32.1344 | 40.836 | 25.287 | 2.0059 | 32.0589 | 64.3073 | 15.9822 | 14.405 | -0.1108 | 105.1433 | 235.1446 | 181.0113 | 233.0456 | 161.4537
2012-10-30 18:32 | 28.5454 | 249.5778 | 195.4604 | 247.3378 | 175.8595 | -1.6026 | 1.2708 | 32.1339 | 40.8351 | 25.2868 | 2.0059 | 32.0585 | 64.306 | 15.9822 | 14.3772 | -0.1105 | 105.141 | 235.2006 | 181.0832 | 232.9606 | 161.4823
2012-10-30 18:32 | 28.5581 | 249.5944 | 195.4456 | 247.3034 | 175.7121 | -1.5968 | 1.2707 | 32.1334 | 40.8332 | 25.2871 | 2.0059 | 32.0579 | 64.3039 | 15.9821 | 14.3899 | -0.1101 | 105.1371 | 235.2045 | 181.0557 | 232.9135 | 161.3222
2012-10-30 18:32 | 28.5536 | 249.6212 | 195.4411 | 247.4635 | 175.8013 | -1.6037 | 1.2694 | 32.0998 | 40.7487 | 25.2867 | 2.0034 | 32.0193 | 64.1492 | 15.9821 | 14.3936 | -0.1106 | 104.8979 | 235.2276 | 181.0475 | 233.0699 | 161.4077
2012-10-30 18:32 | 28.5589 | 249.6109 | 195.4149 | 247.3121 | 175.8339 | -1.6014 | 1.2645 | 31.973 | 40.4292 | 25.2855 | 1.9959 | 31.8991 | 63.6676 | 15.9822 | 14.407 | -0.1104 | 104.0968 | 235.2039 | 181.0079 | 232.9051 | 161.4269
2012-10-30 18:32 | 28.5566 | 249.5498 | 195.4008 | 247.4861 | 175.6521 | -1.5985 | 1.2644 | 31.9713 | 40.4255 | 25.2852 | 1.9957 | 31.8959 | 63.6551 | 15.9822 | 14.3966 | -0.1102 | 104.0806 | 235.1532 | 181.0042 | 233.0895 | 161.2555
2012-10-30 18:32 | 28.5559 | 249.5413 | 195.3726 | 247.3483 | 175.6006 | -1.5974 | 1.2638 | 31.9547 | 40.3835 | 25.2852 | 1.9947 | 31.8795 | 63.5891 | 15.9823 | 14.3836 | -0.1101 | 103.9726 | 235.1577 | 180.989 | 232.9647 | 161.217
2012-10-30 18:32 | 28.5416 | 249.5662 | 195.3465 | 247.2674 | 175.8049 | -1.6026 | 1.2635 | 31.9483 | 40.3675 | 25.285 | 1.9943 | 31.8732 | 63.5641 | 15.9824 | 14.3408 | -0.1105 | 103.9316 | 235.2254 | 181.0057 | 232.9266 | 161.4641
2012-10-30 18:32 | 28.5514 | 249.4938 | 195.3485 | 247.3986 | 175.5415 | -1.6026 | 1.2635 | 31.9469 | 40.3637 | 25.2852 | 1.9942 | 31.8718 | 63.5582 | 15.9823 | 14.3873 | -0.1105 | 103.9219 | 235.1065 | 180.9612 | 233.0113 | 161.1542
2012-10-30 18:32 | 28.5484 | 249.4751 | 195.3573 | 247.3956 | 175.5463 | -1.6009 | 1.2633 | 31.9437 | 40.3558 | 25.2851 | 1.994 | 31.8687 | 63.5455 | 15.9824 | 14.372 | -0.1104 | 103.9013 | 235.1031 | 180.9853 | 233.0236 | 161.1743
2012-10-30 18:32 | 28.5491 | 249.4993 | 195.3776 | 247.1926 | 175.9764 | -1.6014 | 1.2634 | 31.945 | 40.3591 | 25.285 | 1.9941 | 31.87 | 63.5504 | 15.9825 | 14.3972 | -0.1104 | 103.9096 | 235.1021 | 180.9804 | 232.7954 | 161.5792
2012-10-30 18:32 | 28.5491 | 249.4562 | 195.358 | 247.3885 | 175.586 | -1.5997 | 1.2635 | 31.947 | 40.3644 | 25.2849 | 1.9942 | 31.872 | 63.5581 | 15.9826 | 14.4095 | -0.1103 | 103.9225 | 235.0467 | 180.9485 | 232.979 | 161.1765
2012-10-30 18:32 | 28.5581 | 249.4574 | 195.4141 | 247.3505 | 175.716 | -1.6026 | 1.2636 | 31.9499 | 40.3711 | 25.2853 | 1.9944 | 31.8748 | 63.5696 | 15.9825 | 14.3858 | -0.1105 | 103.9407 | 235.0716 | 181.0283 | 232.9647 | 161.3302
2012-10-30 18:32 | 28.5484 | 249.4712 | 195.2905 | 247.2076 | 175.718 | -1.6014 | 1.2647 | 31.9784 | 40.4433 | 25.2852 | 1.9962 | 31.9043 | 63.6873 | 15.9826 | 14.4169 | -0.1104 | 104.1305 | 235.0543 | 180.8736 | 232.7907 | 161.3011
2012-10-30 18:32 | 28.5476 | 249.4391 | 195.3997 | 247.3987 | 175.8695 | -1.5991 | 1.2686 | 32.0769 | 40.6912 | 25.2862 | 2.0023 | 32.0018 | 64.0771 | 15.9825 | 14.3916 | -0.1102 | 104.7684 | 235.0475 | 181.0081 | 233.0071 | 161.4779
2012-10-30 18:32 | 28.5386 | 249.4183 | 195.4575 | 247.2761 | 175.9776 | -1.6026 | 1.2692 | 32.0941 | 40.7345 | 25.2865 | 2.0034 | 32.019 | 64.1461 | 15.9825 | 14.403 | -0.1105 | 104.8806 | 235.0153 | 181.0545 | 232.8731 | 161.5746
2012-10-30 18:32 | 28.5484 | 249.479 | 195.4201 | 247.235 | 175.9366 | -1.6014 | 1.2698 | 32.109 | 40.7715 | 25.2869 | 2.0043 | 32.0338 | 64.2063 | 15.9823 | 14.4006 | -0.1104 | 104.9778 | 235.0784 | 181.0195 | 232.8344 | 161.536
2012-10-30 18:32 | 28.5484 | 249.4751 | 195.4909 | 247.376 | 176.0498 | -1.5974 | 1.2702 | 32.1189 | 40.7973 | 25.2866 | 2.0049 | 32.0436 | 64.2453 | 15.9824 | 14.3761 | -0.1101 | 105.0427 | 235.099 | 181.1148 | 232.9999 | 161.6737
2012-10-30 18:32 | 28.5476 | 249.49 | 195.5097 | 247.1989 | 176.131 | -1.6014 | 1.2705 | 32.1263 | 40.8159 | 25.2867 | 2.0054 | 32.0507 | 64.2743 | 15.9823 | 14.3957 | -0.1104 | 105.0902 | 235.0943 | 181.114 | 232.8032 | 161.7353
2012-10-30 18:32 | 28.5454 | 249.5191 | 195.5272 | 247.2829 | 176.164 | -1.602 | 1.2707 | 32.1319 | 40.8293 | 25.2871 | 2.0057 | 32.0563 | 64.2966 | 15.9823 | 14.3772 | -0.1104 | 105.1259 | 235.1419 | 181.15 | 232.9057 | 161.7868
2012-10-30 18:33 | 28.5416 | 249.4722 | 195.5705 | 247.2987 | 176.1797 | -1.6014 | 1.2708 | 32.1348 | 40.8365 | 25.2874 | 2.0059 | 32.0594 | 64.3091 | 15.9822 | 14.3979 | -0.1104 | 105.1456 | 235.0743 | 181.1726 | 232.9008 | 161.7818
2012-10-30 18:33 | 28.5409 | 249.5067 | 195.617 | 247.2549 | 176.261 | -1.6009 | 1.2709 | 32.1368 | 40.8417 | 25.2873 | 2.006 | 32.0612 | 64.316 | 15.9824 | 14.389 | -0.1104 | 105.1577 | 235.1177 | 181.228 | 232.8659 | 161.872
2012-10-30 18:33 | 28.5297 | 249.476 | 195.6097 | 247.228 | 176.2537 | -1.5962 | 1.2709 | 32.1381 | 40.8454 | 25.2871 | 2.0061 | 32.0626 | 64.322 | 15.9822 | 14.3982 | -0.11 | 105.1674 | 235.0778 | 181.2115 | 232.8298 | 161.8555
2012-10-30 18:33 | 28.5394 | 249.4778 | 195.6469 | 247.3318 | 176.0682 | -1.602 | 1.2709 | 32.1387 | 40.8465 | 25.2873 | 2.0062 | 32.0632 | 64.3246 | 15.9822 | 14.3916 | -0.1104 | 105.1711 | 235.0862 | 181.2553 | 232.9402 | 161.6766
2012-10-30 18:33 | 28.5282 | 249.4784 | 195.6514 | 247.2696 | 176.2014 | -1.6009 | 1.271 | 32.1398 | 40.8494 | 25.2872 | 2.0062 | 32.0642 | 64.3285 | 15.9822 | 14.3804 | -0.1104 | 105.1779 | 235.098 | 181.271 | 232.8892 | 161.821
2012-10-30 18:33 | 28.5319 | 249.533 | 195.6787 | 247.2185 | 176.3379 | -1.598 | 1.271 | 32.1404 | 40.8506 | 25.2874 | 2.0063 | 32.0648 | 64.3309 | 15.9822 | 14.4208 | -0.1102 | 105.1815 | 235.1122 | 181.2579 | 232.7977 | 161.9171
2012-10-30 18:33 | 28.5319 | 249.5134 | 195.6669 | 247.3086 | 176.2051 | -1.5985 | 1.271 | 32.1403 | 40.8505 | 25.2873 | 2.0063 | 32.0648 | 64.3306 | 15.9823 | 14.4004 | -0.1102 | 105.1811 | 235.113 | 181.2665 | 232.9082 | 161.8047
2012-10-30 18:33 | 28.5252 | 249.5145 | 195.6917 | 247.3215 | 176.3468 | -1.5962 | 1.271 | 32.1407 | 40.8513 | 25.2874 | 2.0063 | 32.0651 | 64.332 | 15.9822 | 14.406 | -0.11 | 105.1833 | 235.1085 | 181.2857 | 232.9155 | 161.9408
2012-10-30 18:33 | 28.5319 | 249.5369 | 195.718 | 247.1519 | 176.373 | -1.5997 | 1.271 | 32.141 | 40.852 | 25.2874 | 2.0063 | 32.0656 | 64.3347 | 15.982 | 14.4167 | -0.1103 | 105.1868 | 235.1202 | 181.3013 | 232.7352 | 161.9563
2012-10-30 18:33 | 28.5311 | 249.5243 | 195.7211 | 247.2764 | 176.0833 | -1.6003 | 1.271 | 32.1413 | 40.8532 | 25.2873 | 2.0063 | 32.0658 | 64.3351 | 15.9822 | 14.3751 | -0.1103 | 105.1882 | 235.1492 | 181.346 | 232.9013 | 161.7082
2012-10-30 18:33 | 28.5386 | 249.5553 | 195.7404 | 247.3584 | 175.9932 | -1.6014 | 1.271 | 32.1416 | 40.8535 | 25.2875 | 2.0064 | 32.066 | 64.3363 | 15.982 | 14.3745 | -0.1104 | 105.1898 | 235.1808 | 181.3659 | 232.9839 | 161.6187
2012-10-30 18:33 | 28.5379 | 249.5625 | 195.7083 | 247.2049 | 175.9183 | -1.5997 | 1.2711 | 32.1415 | 40.8537 | 25.2872 | 2.0064 | 32.066 | 64.3362 | 15.9821 | 14.3983 | -0.1103 | 105.1899 | 235.1642 | 181.31 | 232.8066 | 161.52
2012-10-30 18:33 | 28.5326 | 249.475 | 195.6205 | 247.2349 | 175.8076 | -1.6026 | 1.2711 | 32.1416 | 40.854 | 25.2873 | 2.0064 | 32.066 | 64.3363 | 15.9821 | 14.393 | -0.1105 | 105.1903 | 235.082 | 181.2275 | 232.8419 | 161.4146
2012-10-30 18:33 | 28.5326 | 249.5337 | 195.6165 | 247.3602 | 175.8974 | -1.5991 | 1.2711 | 32.1417 | 40.8539 | 25.2874 | 2.0064 | 32.0661 | 64.3375 | 15.9819 | 14.4011 | -0.1102 | 105.1913 | 235.1326 | 181.2154 | 232.9591 | 161.4963
2012-10-30 18:33 | 28.5454 | 249.5621 | 195.6175 | 247.3103 | 175.8399 | -1.602 | 1.2711 | 32.1415 | 40.8536 | 25.2874 | 2.0064 | 32.066 | 64.3363 | 15.982 | 14.365 | -0.1104 | 105.1899 | 235.1971 | 181.2525 | 232.9453 | 161.4749
2012-10-30 18:33 | 28.5409 | 249.6281 | 195.5973 | 247.1844 | 175.8042 | -1.6009 | 1.271 | 32.1413 | 40.8529 | 25.2875 | 2.0064 | 32.0657 | 64.3352 | 15.982 | 14.3809 | -0.1104 | 105.1881 | 235.2472 | 181.2164 | 232.8035 | 161.4233
2012-10-30 18:33 | 28.5364 | 249.5688 | 195.5535 | 247.3444 | 175.8153 | -1.6032 | 1.271 | 32.1415 | 40.8531 | 25.2876 | 2.0064 | 32.0659 | 64.3363 | 15.982 | 14.3804 | -0.1105 | 105.1895 | 235.1884 | 181.1731 | 232.964 | 161.4349
2012-10-30 18:33 | 28.5431 | 249.5638 | 195.5838 | 247.3511 | 175.9001 | -1.602 | 1.271 | 32.1411 | 40.8525 | 25.2872 | 2.0064 | 32.0655 | 64.3351 | 15.9819 | 14.4035 | -0.1104 | 105.1876 | 235.1603 | 181.1803 | 232.9476 | 161.4966
2012-10-30 18:33 | 28.5326 | 249.5728 | 195.5497 | 247.1918 | 175.8818 | -1.598 | 1.271 | 32.1409 | 40.8516 | 25.2876 | 2.0063 | 32.0654 | 64.3343 | 15.982 | 14.3562 | -0.1102 | 105.1859 | 235.2166 | 181.1935 | 232.8356 | 161.5256
2012-10-30 18:33 | 28.5364 | 249.5375 | 195.5457 | 247.3209 | 175.9637 | -1.602 | 1.271 | 32.1408 | 40.8514 | 25.2875 | 2.0063 | 32.0653 | 64.3341 | 15.9819 | 14.3315 | -0.1104 | 105.1856 | 235.206 | 181.2142 | 232.9894 | 161.6322
2012-10-30 18:33 | 28.5476 | 249.5604 | 195.5137 | 247.3713 | 176.049 | -1.5991 | 1.271 | 32.1405 | 40.8507 | 25.2875 | 2.0063 | 32.065 | 64.333 | 15.9819 | 14.3672 | -0.1102 | 105.1837 | 235.1932 | 181.1465 | 233.0041 | 161.6818
2012-10-30 18:33 | 28.5446 | 249.5692 | 195.5932 | 247.384 | 175.9875 | -1.6009 | 1.271 | 32.1405 | 40.8506 | 25.2876 | 2.0063 | 32.0649 | 64.3328 | 15.9819 | 14.3642 | -0.1104 | 105.1834 | 235.205 | 181.229 | 233.0198 | 161.6233
2012-10-30 18:33 | 28.5469 | 249.5754 | 195.5601 | 247.2805 | 175.9898 | -1.6055 | 1.271 | 32.1405 | 40.8507 | 25.2875 | 2.0063 | 32.0648 | 64.3326 | 15.9819 | 14.3338 | -0.1107 | 105.1833 | 235.2416 | 181.2263 | 232.9467 | 161.656
2012-10-30 18:33 | 28.5491 | 249.5737 | 195.5702 | 247.3415 | 175.9607 | -1.6043 | 1.271 | 32.1399 | 40.8492 | 25.2874 | 2.0063 | 32.0643 | 64.3306 | 15.9818 | 14.3564 | -0.1106 | 105.1798 | 235.2173 | 181.2138 | 232.9851 | 161.6043
2012-10-30 18:34 | 28.5506 | 249.5947 | 195.5049 | 247.3234 | 175.7749 | -1.6003 | 1.271 | 32.1397 | 40.8484 | 25.2877 | 2.0063 | 32.0641 | 64.3297 | 15.9819 | 14.3253 | -0.1103 | 105.1781 | 235.2694 | 181.1796 | 232.9981 | 161.4496
2012-10-30 18:34 | 28.5536 | 249.5703 | 195.4961 | 247.3969 | 175.8598 | -1.5945 | 1.271 | 32.1395 | 40.8486 | 25.2873 | 2.0063 | 32.064 | 64.3293 | 15.9818 | 14.3405 | -0.1099 | 105.1778 | 235.2298 | 181.1556 | 233.0564 | 161.5193
2012-10-30 18:34 | 28.5693 | 249.59 | 195.5236 | 247.3499 | 175.9809 | -1.6084 | 1.271 | 32.1396 | 40.8479 | 25.2878 | 2.0063 | 32.064 | 64.3286 | 15.982 | 14.3685 | -0.1109 | 105.1765 | 235.2215 | 181.1551 | 232.9814 | 161.6124
2012-10-30 18:34 | 28.5581 | 249.5866 | 195.5634 | 247.3191 | 175.9151 | -1.5939 | 1.271 | 32.1393 | 40.848 | 25.2872 | 2.0063 | 32.0637 | 64.3285 | 15.9818 | 14.3369 | -0.1099 | 105.1765 | 235.2497 | 181.2265 | 232.9822 | 161.5782
2012-10-30 18:34 | 28.5589 | 249.5913 | 195.4935 | 247.3748 | 175.8534 | -1.6014 | 1.271 | 32.1389 | 40.847 | 25.2873 | 2.0063 | 32.0634 | 64.3271 | 15.9817 | 14.354 | -0.1104 | 105.1741 | 235.2373 | 181.1395 | 233.0208 | 161.4994
2012-10-30 18:34 | 28.5664 | 249.591 | 195.4774 | 247.4097 | 175.7087 | -1.6003 | 1.2709 | 32.1386 | 40.8456 | 25.2876 | 2.0062 | 32.063 | 64.326 | 15.9817 | 14.3248 | -0.1103 | 105.1716 | 235.2662 | 181.1526 | 233.0849 | 161.3839
2012-10-30 18:34 | 28.5634 | 249.5801 | 195.4941 | 247.4145 | 175.7916 | -1.598 | 1.2709 | 32.1387 | 40.8462 | 25.2874 | 2.0062 | 32.0631 | 64.3259 | 15.9818 | 14.334 | -0.1102 | 105.172 | 235.2461 | 181.1601 | 233.0805 | 161.4576
2012-10-30 18:34 | 28.5731 | 249.6603 | 195.4841 | 247.3811 | 175.8052 | -1.5991 | 1.2709 | 32.1381 | 40.8452 | 25.2872 | 2.0062 | 32.0625 | 64.3239 | 15.9817 | 14.3192 | -0.1102 | 105.1691 | 235.3411 | 181.1649 | 233.0619 | 161.486
2012-10-30 18:34 | 28.5753 | 249.5764 | 195.4667 | 247.3716 | 175.7801 | -1.602 | 1.2709 | 32.1379 | 40.8441 | 25.2874 | 2.0062 | 32.0622 | 64.323 | 15.9816 | 14.3377 | -0.1104 | 105.1671 | 235.2387 | 181.129 | 233.0339 | 161.4424
2012-10-30 18:34 | 28.5821 | 249.6028 | 195.4617 | 247.3392 | 175.658 | -1.6037 | 1.2709 | 32.1376 | 40.844 | 25.2871 | 2.0062 | 32.0622 | 64.3231 | 15.9816 | 14.3568 | -0.1106 | 105.1671 | 235.246 | 181.1049 | 232.9824 | 161.3012
2012-10-30 18:34 | 28.5791 | 249.6076 | 195.4469 | 247.391 | 175.5965 | -1.6061 | 1.2709 | 32.1371 | 40.8419 | 25.2875 | 2.0061 | 32.0615 | 64.32 | 15.9817 | 14.3497 | -0.1107 | 105.162 | 235.2579 | 181.0972 | 233.0413 | 161.2468
2012-10-30 18:34 | 28.5911 | 249.6352 | 195.4275 | 247.4265 | 175.5851 | -1.6014 | 1.2709 | 32.1369 | 40.8418 | 25.2873 | 2.0061 | 32.0613 | 64.3195 | 15.9816 | 14.3984 | -0.1104 | 105.1613 | 235.2368 | 181.0291 | 233.0281 | 161.1867
2012-10-30 18:34 | 28.5866 | 249.5955 | 195.3955 | 247.3594 | 175.5611 | -1.6009 | 1.2708 | 32.1363 | 40.8404 | 25.2873 | 2.0061 | 32.0608 | 64.3181 | 15.9814 | 14.3898 | -0.1104 | 105.1585 | 235.2057 | 181.0057 | 232.9696 | 161.1713
2012-10-30 18:34 | 28.5888 | 249.5781 | 195.4409 | 247.3459 | 175.8287 | -1.6003 | 1.2708 | 32.1357 | 40.8384 | 25.2876 | 2.0061 | 32.0603 | 64.3159 | 15.9815 | 14.3594 | -0.1103 | 105.1542 | 235.2187 | 181.0815 | 232.9865 | 161.4693
2012-10-30 18:34 | 28.5963 | 249.617 | 195.4012 | 247.3495 | 175.7854 | -1.5997 | 1.2708 | 32.1355 | 40.8378 | 25.2876 | 2.0061 | 32.06 | 64.3143 | 15.9816 | 14.4159 | -0.1103 | 105.1521 | 235.2011 | 180.9853 | 232.9336 | 161.3695
2012-10-30 18:34 | 28.5903 | 249.5992 | 195.3952 | 247.4414 | 175.6272 | -1.5962 | 1.2708 | 32.1353 | 40.8374 | 25.2875 | 2.006 | 32.0597 | 64.3133 | 15.9815 | 14.3609 | -0.11 | 105.1507 | 235.2383 | 181.0343 | 233.0805 | 161.2663
2012-10-30 18:34 | 28.5896 | 249.5868 | 195.371 | 247.3898 | 175.6577 | -1.598 | 1.2708 | 32.1352 | 40.8371 | 25.2876 | 2.006 | 32.0597 | 64.313 | 15.9816 | 14.348 | -0.1102 | 105.1501 | 235.2388 | 181.023 | 233.0418 | 161.3097
2012-10-30 18:34 | 28.5956 | 249.6397 | 195.3534 | 247.3527 | 175.5584 | -1.6032 | 1.2708 | 32.1349 | 40.8363 | 25.2876 | 2.006 | 32.0594 | 64.3121 | 15.9815 | 14.3499 | -0.1105 | 105.1484 | 235.2898 | 181.0035 | 233.0028 | 161.2085
2012-10-30 18:34 | 28.5971 | 249.6217 | 195.3549 | 247.3777 | 175.5716 | -1.602 | 1.2708 | 32.1346 | 40.8362 | 25.2872 | 2.006 | 32.0592 | 64.3114 | 15.9815 | 14.3922 | -0.1104 | 105.1476 | 235.2295 | 180.9627 | 232.9855 | 161.1794
2012-10-30 18:34 | 28.6046 | 249.5978 | 195.3624 | 247.3852 | 175.5439 | -1.5945 | 1.2708 | 32.1344 | 40.8359 | 25.2871 | 2.006 | 32.0588 | 64.3104 | 15.9814 | 14.3834 | -0.1099 | 105.1463 | 235.2144 | 180.979 | 233.0018 | 161.1605
2012-10-30 18:34 | 28.6038 | 249.5736 | 195.3537 | 247.4236 | 175.3987 | -1.5997 | 1.2707 | 32.1336 | 40.8336 | 25.2872 | 2.0059 | 32.058 | 64.3067 | 15.9815 | 14.4274 | -0.1103 | 105.1403 | 235.1462 | 180.9263 | 232.9962 | 160.9713
2012-10-30 18:34 | 28.6068 | 249.6157 | 195.3057 | 247.3992 | 175.5266 | -1.6032 | 1.2707 | 32.133 | 40.8319 | 25.2874 | 2.0059 | 32.0575 | 64.3051 | 15.9814 | 14.41 | -0.1105 | 105.137 | 235.2057 | 180.8957 | 232.9892 | 161.1166
2012-10-30 18:34 | 28.606 | 249.5718 | 195.3363 | 247.4375 | 175.639 | -1.6009 | 1.2692 | 32.096 | 40.7377 | 25.2874 | 2.0031 | 32.0123 | 64.1237 | 15.9815 | 14.3807 | -0.1104 | 104.8614 | 235.1911 | 180.9556 | 233.0568 | 161.2583
2012-10-30 18:35 | 28.6105 | 249.5724 | 195.3094 | 247.3558 | 175.6318 | -1.6084 | 1.2645 | 31.9744 | 40.4323 | 25.2858 | 1.9961 | 31.8999 | 63.6741 | 15.9815 | 14.3607 | -0.1109 | 104.1064 | 235.2117 | 180.9487 | 232.9951 | 161.2711
2012-10-30 18:35 | 28.6218 | 249.5994 | 195.2932 | 247.4102 | 175.405 | -1.5991 | 1.2643 | 31.9695 | 40.4204 | 25.2855 | 1.9957 | 31.8945 | 63.6521 | 15.9815 | 14.3802 | -0.1102 | 104.0725 | 235.2192 | 180.913 | 233.03 | 161.0248
2012-10-30 18:35 | 28.615 | 249.526 | 195.2274 | 247.4152 | 175.5973 | -1.6003 | 1.2638 | 31.9564 | 40.3875 | 25.2853 | 1.9949 | 31.8812 | 63.599 | 15.9816 | 14.3611 | -0.1103 | 103.9865 | 235.1649 | 180.8663 | 233.0541 | 161.2362
2012-10-30 18:35 | 28.6263 | 249.5452 | 195.2662 | 247.3991 | 175.4564 | -1.602 | 1.2636 | 31.9499 | 40.3706 | 25.2857 | 1.9945 | 31.875 | 63.5734 | 15.9817 | 14.401 | -0.1104 | 103.944 | 235.1442 | 180.8652 | 232.9981 | 161.0554
2012-10-30 18:35 | 28.621 | 249.4968 | 195.1981 | 247.3115 | 175.014 | -1.602 | 1.2635 | 31.948 | 40.3656 | 25.2857 | 1.9943 | 31.8727 | 63.5647 | 15.9817 | 14.363 | -0.1104 | 103.9304 | 235.1338 | 180.8351 | 232.9485 | 160.651
2012-10-30 18:35 | 28.6233 | 249.4678 | 195.1257 | 247.3608 | 174.8485 | -1.6003 | 1.2635 | 31.947 | 40.3636 | 25.2854 | 1.9943 | 31.8719 | 63.5611 | 15.9817 | 14.3939 | -0.1103 | 103.9248 | 235.0739 | 180.7318 | 232.9669 | 160.4546
2012-10-30 18:35 | 28.6233 | 249.46 | 195.0864 | 247.4274 | 175.0943 | -1.6032 | 1.2635 | 31.9476 | 40.3648 | 25.2856 | 1.9943 | 31.8725 | 63.5633 | 15.9818 | 14.3939 | -0.1105 | 103.9281 | 235.0661 | 180.6925 | 233.0335 | 160.7004
2012-10-30 18:35 | 28.6225 | 249.4513 | 195.0817 | 247.313 | 175.2652 | -1.6009 | 1.2635 | 31.9477 | 40.3648 | 25.2858 | 1.9943 | 31.8727 | 63.5638 | 15.9819 | 14.3686 | -0.1104 | 103.9287 | 235.0827 | 180.7131 | 232.9444 | 160.8966
2012-10-30 18:35 | 28.6188 | 249.432 | 195.0662 | 247.278 | 175.402 | -1.5945 | 1.2636 | 31.95 | 40.3708 | 25.2857 | 1.9944 | 31.8749 | 63.5724 | 15.9819 | 14.4016 | -0.1099 | 103.9432 | 235.0304 | 180.6646 | 232.8764 | 161.0004
2012-10-30 18:35 | 28.621 | 249.4459 | 195.0487 | 247.3507 | 175.4433 | -1.6032 | 1.2651 | 31.988 | 40.4671 | 25.2856 | 1.9969 | 31.9146 | 63.7307 | 15.9819 | 14.3753 | -0.1105 | 104.1978 | 235.0706 | 180.6734 | 232.9754 | 161.068
2012-10-30 18:35 | 28.6188 | 249.4006 | 195.0858 | 247.3524 | 175.6557 | -1.6003 | 1.2686 | 32.0791 | 40.6956 | 25.287 | 2.0025 | 32.0039 | 64.0886 | 15.9818 | 14.3608 | -0.1103 | 104.7843 | 235.0398 | 180.725 | 232.9916 | 161.2949
2012-10-30 18:35 | 28.624 | 249.4372 | 195.1853 | 247.2832 | 175.6609 | -1.6009 | 1.2693 | 32.0983 | 40.744 | 25.2872 | 2.0037 | 32.0232 | 64.1659 | 15.9818 | 14.4068 | -0.1104 | 104.9098 | 235.0304 | 180.7785 | 232.8764 | 161.2541
2012-10-30 18:35 | 28.6285 | 249.4417 | 195.2056 | 247.225 | 175.5366 | -1.5957 | 1.2699 | 32.113 | 40.7809 | 25.2875 | 2.0046 | 32.0376 | 64.2241 | 15.9817 | 14.3909 | -0.11 | 105.005 | 235.0508 | 180.8147 | 232.8341 | 161.1457
2012-10-30 18:35 | 28.6195 | 249.4562 | 195.1887 | 247.3061 | 175.352 | -1.6009 | 1.2703 | 32.1219 | 40.8038 | 25.2873 | 2.0052 | 32.0463 | 64.2586 | 15.9817 | 14.3575 | -0.1104 | 105.0624 | 235.0987 | 180.8312 | 232.9486 | 160.9945
2012-10-30 18:35 | 28.6225 | 249.4317 | 195.1956 | 247.3326 | 175.8507 | -1.6026 | 1.2705 | 32.1281 | 40.8189 | 25.2877 | 2.0056 | 32.0526 | 64.2845 | 15.9816 | 14.3727 | -0.1105 | 105.1034 | 235.059 | 180.8229 | 232.9599 | 161.478
2012-10-30 18:35 | 28.6248 | 249.4419 | 195.2765 | 247.2409 | 175.9154 | -1.5968 | 1.2707 | 32.1324 | 40.8299 | 25.2875 | 2.0058 | 32.0568 | 64.3009 | 15.9817 | 14.3668 | -0.1101 | 105.1309 | 235.0751 | 180.9097 | 232.8741 | 161.5486
2012-10-30 18:35 | 28.621 | 249.4577 | 195.2688 | 247.2371 | 175.7165 | -1.6026 | 1.2708 | 32.1354 | 40.8376 | 25.2875 | 2.006 | 32.0597 | 64.3128 | 15.9816 | 14.3794 | -0.1105 | 105.1504 | 235.0783 | 180.8894 | 232.8577 | 161.3371
2012-10-30 18:35 | 28.6278 | 249.4136 | 195.2795 | 247.2752 | 175.5086 | -1.6003 | 1.2709 | 32.1374 | 40.8428 | 25.2875 | 2.0062 | 32.0617 | 64.321 | 15.9816 | 14.378 | -0.1103 | 105.1639 | 235.0356 | 180.9015 | 232.8972 | 161.1306
2012-10-30 18:35 | 28.6218 | 249.4428 | 195.246 | 247.328 | 175.5065 | -1.602 | 1.2709 | 32.1386 | 40.8454 | 25.2878 | 2.0063 | 32.063 | 64.327 | 15.9814 | 14.3475 | -0.1104 | 105.1724 | 235.0953 | 180.8985 | 232.9805 | 161.159
2012-10-30 18:35 | 28.621 | 249.4302 | 195.257 | 247.2919 | 175.494 | -1.6026 | 1.271 | 32.1398 | 40.8487 | 25.2876 | 2.0063 | 32.0641 | 64.3312 | 15.9815 | 14.3998 | -0.1105 | 105.1799 | 235.0304 | 180.8572 | 232.8921 | 161.0942
2012-10-30 18:35 | 28.624 | 249.4411 | 195.2325 | 247.3067 | 175.4228 | -1.5985 | 1.271 | 32.1402 | 40.8499 | 25.2875 | 2.0064 | 32.0646 | 64.3332 | 15.9814 | 14.4109 | -0.1102 | 105.1832 | 235.0302 | 180.8216 | 232.8958 | 161.0119
2012-10-30 18:35 | 28.6173 | 249.4618 | 195.2415 | 247.202 | 175.5762 | -1.6049 | 1.271 | 32.1411 | 40.8524 | 25.2874 | 2.0064 | 32.0655 | 64.3366 | 15.9815 | 14.3961 | -0.1106 | 105.1891 | 235.0657 | 180.8454 | 232.8059 | 161.1801
2012-10-30 18:35 | 28.6233 | 249.4639 | 195.2396 | 247.2354 | 175.4143 | -1.6032 | 1.2711 | 32.142 | 40.8541 | 25.2877 | 2.0065 | 32.0664 | 64.3404 | 15.9815 | 14.3898 | -0.1105 | 105.1945 | 235.0741 | 180.8498 | 232.8456 | 161.0245
2012-10-30 18:35 | 28.6173 | 249.4383 | 195.2061 | 247.2765 | 175.4942 | -1.5997 | 1.2711 | 32.1431 | 40.8577 | 25.2873 | 2.0066 | 32.0674 | 64.3451 | 15.9813 | 14.4328 | -0.1103 | 105.2027 | 235.0055 | 180.7733 | 232.8437 | 161.0614
2012-10-30 18:36 | 28.6218 | 249.4585 | 195.2224 | 247.3045 | 175.4714 | -1.6032 | 1.2711 | 32.1432 | 40.8575 | 25.2875 | 2.0066 | 32.0675 | 64.3454 | 15.9814 | 14.4087 | -0.1105 | 105.2029 | 235.0498 | 180.8137 | 232.8958 | 161.0627
2012-10-30 18:36 | 28.6158 | 249.4877 | 195.1968 | 247.3298 | 175.6215 | -1.6055 | 1.2711 | 32.1432 | 40.8572 | 25.2878 | 2.0066 | 32.0675 | 64.3454 | 15.9813 | 14.4109 | -0.1107 | 105.2026 | 235.0768 | 180.7859 | 232.9189 | 161.2106
2012-10-30 18:36 | 28.6105 | 249.4824 | 195.2386 | 247.1952 | 175.6825 | -1.5974 | 1.2711 | 32.1432 | 40.8575 | 25.2876 | 2.0066 | 32.0676 | 64.3461 | 15.9813 | 14.3729 | -0.1101 | 105.2036 | 235.1095 | 180.8657 | 232.8223 | 161.3096
2012-10-30 18:36 | 28.6188 | 249.4672 | 195.2076 | 247.2466 | 175.4762 | -1.5991 | 1.2711 | 32.1437 | 40.8585 | 25.2876 | 2.0066 | 32.068 | 64.3478 | 15.9813 | 14.3649 | -0.1102 | 105.2062 | 235.1023 | 180.8427 | 232.8817 | 161.1113
2012-10-30 18:36 | 28.6248 | 249.4928 | 195.1979 | 247.2879 | 175.2714 | -1.6032 | 1.2711 | 32.1437 | 40.8586 | 25.2876 | 2.0066 | 32.068 | 64.3481 | 15.9812 | 14.3913 | -0.1105 | 105.2067 | 235.1015 | 180.8066 | 232.8966 | 160.8801
2012-10-30 18:36 | 28.615 | 249.4908 | 195.1921 | 247.3173 | 175.4763 | -1.6037 | 1.2711 | 32.143 | 40.857 | 25.2876 | 2.0066 | 32.0674 | 64.3459 | 15.9812 | 14.3815 | -0.1106 | 105.2028 | 235.1093 | 180.8106 | 232.9358 | 161.0948
2012-10-30 18:36 | 28.6225 | 249.4357 | 195.1681 | 247.2699 | 175.4135 | -1.6014 | 1.2711 | 32.1431 | 40.8573 | 25.2876 | 2.0066 | 32.0676 | 64.3466 | 15.9811 | 14.3401 | -0.1104 | 105.2039 | 235.0956 | 180.828 | 232.9298 | 161.0734
2012-10-30 18:36 | 28.627 | 249.4871 | 195.1923 | 247.2783 | 175.4258 | -1.6043 | 1.2711 | 32.1434 | 40.8586 | 25.2871 | 2.0066 | 32.0677 | 64.3472 | 15.9811 | 14.3976 | -0.1106 | 105.2058 | 235.0895 | 180.7947 | 232.8807 | 161.0282
2012-10-30 18:36 | 28.6255 | 249.5091 | 195.1986 | 247.2298 | 175.3814 | -1.6009 | 1.2711 | 32.1431 | 40.8573 | 25.2874 | 2.0066 | 32.0675 | 64.3467 | 15.981 | 14.3757 | -0.1104 | 105.2041 | 235.1334 | 180.8229 | 232.8541 | 161.0057
2012-10-30 18:36 | 28.6188 | 249.4907 | 195.1408 | 247.2505 | 175.445 | -1.602 | 1.2711 | 32.1428 | 40.8565 | 25.2876 | 2.0066 | 32.0672 | 64.3448 | 15.9811 | 14.3976 | -0.1104 | 105.2013 | 235.0931 | 180.7432 | 232.8529 | 161.0474
2012-10-30 18:36 | 28.6195 | 249.4796 | 195.1376 | 247.3335 | 175.3949 | -1.5991 | 1.2711 | 32.1424 | 40.8553 | 25.2877 | 2.0065 | 32.0669 | 64.3437 | 15.9811 | 14.3901 | -0.1102 | 105.199 | 235.0895 | 180.7475 | 232.9434 | 161.0048
2012-10-30 18:36 | 28.6225 | 249.5179 | 195.1603 | 247.3287 | 175.394 | -1.6026 | 1.2711 | 32.1422 | 40.8549 | 25.2876 | 2.0065 | 32.0665 | 64.3424 | 15.9811 | 14.3401 | -0.1105 | 105.1973 | 235.1778 | 180.8202 | 232.9886 | 161.0539
2012-10-30 18:36 | 28.6165 | 249.4962 | 195.1857 | 247.2835 | 175.6378 | -1.6014 | 1.2711 | 32.142 | 40.8546 | 25.2874 | 2.0065 | 32.0663 | 64.3422 | 15.981 | 14.3218 | -0.1104 | 105.1968 | 235.1744 | 180.8639 | 232.9617 | 161.316
2012-10-30 18:36 | 28.6278 | 249.5153 | 195.1773 | 247.2674 | 175.5203 | -1.598 | 1.2711 | 32.1423 | 40.8549 | 25.2877 | 2.0065 | 32.0667 | 64.3434 | 15.981 | 14.3862 | -0.1102 | 105.1983 | 235.1291 | 180.7911 | 232.8812 | 161.1341
2012-10-30 18:36 | 28.6263 | 249.5413 | 195.2112 | 247.3599 | 175.632 | -1.6009 | 1.2711 | 32.142 | 40.8544 | 25.2875 | 2.0065 | 32.0663 | 64.3422 | 15.9809 | 14.401 | -0.1104 | 105.1966 | 235.1403 | 180.8102 | 232.9589 | 161.231
2012-10-30 18:36 | 28.6173 | 249.4657 | 195.1983 | 247.3548 | 175.541 | -1.6043 | 1.2711 | 32.1421 | 40.8545 | 25.2877 | 2.0065 | 32.0665 | 64.343 | 15.981 | 14.3757 | -0.1106 | 105.1975 | 235.09 | 180.8226 | 232.9791 | 161.1653
2012-10-30 18:36 | 28.618 | 249.5134 | 195.2618 | 247.3673 | 175.6159 | -1.6009 | 1.2711 | 32.1427 | 40.8565 | 25.2873 | 2.0066 | 32.0671 | 64.3455 | 15.9809 | 14.3927 | -0.1104 | 105.202 | 235.1207 | 180.8691 | 232.9746 | 161.2232
2012-10-30 18:36 | 28.6285 | 249.5395 | 195.2723 | 247.2798 | 175.763 | -1.5997 | 1.2711 | 32.1424 | 40.8551 | 25.2877 | 2.0066 | 32.0667 | 64.3436 | 15.981 | 14.4032 | -0.1103 | 105.1988 | 235.1363 | 180.8691 | 232.8766 | 161.3598
2012-10-30 18:36 | 28.6255 | 249.5405 | 195.2536 | 247.3003 | 175.7249 | -1.6009 | 1.2711 | 32.142 | 40.8547 | 25.2873 | 2.0065 | 32.0663 | 64.3425 | 15.9809 | 14.3839 | -0.1104 | 105.1972 | 235.1566 | 180.8697 | 232.9164 | 161.341
2012-10-30 18:36 | 28.6233 | 249.4991 | 195.3104 | 247.3765 | 175.8398 | -1.6009 | 1.2711 | 32.1417 | 40.854 | 25.2874 | 2.0065 | 32.0662 | 64.342 | 15.9808 | 14.4551 | -0.1104 | 105.196 | 235.044 | 180.8553 | 232.9214 | 161.3847
2012-10-30 18:36 | 28.615 | 249.4673 | 195.3335 | 247.3839 | 175.6948 | -1.6014 | 1.271 | 32.1411 | 40.8528 | 25.2872 | 2.0065 | 32.0655 | 64.3388 | 15.981 | 14.4346 | -0.1104 | 105.1916 | 235.0327 | 180.8989 | 232.9493 | 161.2602
2012-10-30 18:36 | 28.6105 | 249.5333 | 195.3329 | 247.4185 | 175.7879 | -1.6014 | 1.271 | 32.1407 | 40.8511 | 25.2875 | 2.0065 | 32.065 | 64.3374 | 15.9809 | 14.3974 | -0.1104 | 105.1885 | 235.1359 | 180.9355 | 233.0211 | 161.3905
2012-10-30 18:36 | 28.6158 | 249.5151 | 195.3225 | 247.3337 | 175.5825 | -1.5985 | 1.271 | 32.1399 | 40.8487 | 25.2877 | 2.0064 | 32.0643 | 64.3344 | 15.9809 | 14.415 | -0.1102 | 105.1831 | 235.1001 | 180.9075 | 232.9187 | 161.1675
2012-10-30 18:36 | 28.6158 | 249.5542 | 195.3225 | 247.3494 | 175.6566 | -1.5962 | 1.271 | 32.1396 | 40.8484 | 25.2875 | 2.0064 | 32.064 | 64.3331 | 15.9809 | 14.3823 | -0.11 | 105.1815 | 235.1719 | 180.9402 | 232.9671 | 161.2743
2012-10-30 18:37 | 28.615 | 249.5769 | 195.3178 | 247.2742 | 175.7729 | -1.6049 | 1.271 | 32.1391 | 40.8478 | 25.2872 | 2.0064 | 32.0635 | 64.3313 | 15.9808 | 14.4101 | -0.1106 | 105.1791 | 235.1668 | 180.9077 | 232.8641 | 161.3628
2012-10-30 18:37 | 28.6165 | 249.5549 | 195.3271 | 247.3579 | 175.7432 | -1.598 | 1.2709 | 32.1388 | 40.8465 | 25.2875 | 2.0064 | 32.0633 | 64.3313 | 15.9806 | 14.4157 | -0.1102 | 105.1777 | 235.1392 | 180.9114 | 232.9422 | 161.3275
2012-10-30 18:37 | 28.6188 | 249.5651 | 195.3177 | 247.4268 | 175.7845 | -1.5974 | 1.2709 | 32.1385 | 40.8454 | 25.2876 | 2.0064 | 32.0628 | 64.3297 | 15.9806 | 14.418 | -0.1101 | 105.1751 | 235.1471 | 180.8997 | 233.0088 | 161.3665
2012-10-30 18:37 | 28.618 | 249.5525 | 195.3915 | 247.2967 | 175.7954 | -1.5991 | 1.2709 | 32.1386 | 40.846 | 25.2873 | 2.0064 | 32.0629 | 64.3297 | 15.9806 | 14.36 | -0.1102 | 105.1758 | 235.1925 | 181.0315 | 232.9367 | 161.4354
2012-10-30 18:37 | 28.6128 | 249.5591 | 195.3313 | 247.3621 | 175.599 | -1.6009 | 1.2709 | 32.1385 | 40.8461 | 25.2872 | 2.0064 | 32.0629 | 64.3298 | 15.9806 | 14.3834 | -0.1104 | 105.1759 | 235.1757 | 180.9479 | 232.9787 | 161.2156
2012-10-30 18:37 | 28.6135 | 249.5402 | 195.3477 | 247.3549 | 175.5255 | -1.6026 | 1.2709 | 32.1382 | 40.8445 | 25.2876 | 2.0063 | 32.0625 | 64.328 | 15.9807 | 14.4004 | -0.1105 | 105.1726 | 235.1398 | 180.9473 | 232.9545 | 161.1251
2012-10-30 18:37 | 28.6128 | 249.5473 | 195.3392 | 247.3895 | 175.4663 | -1.6014 | 1.2709 | 32.1375 | 40.8435 | 25.2873 | 2.0063 | 32.0619 | 64.3258 | 15.9806 | 14.3712 | -0.1104 | 105.1693 | 235.1761 | 180.968 | 233.0183 | 161.0951
2012-10-30 18:37 | 28.6165 | 249.5589 | 195.3271 | 247.3854 | 175.7471 | -1.602 | 1.2709 | 32.1374 | 40.8432 | 25.2872 | 2.0063 | 32.0618 | 64.3258 | 15.9805 | 14.4157 | -0.1104 | 105.1691 | 235.1432 | 180.9114 | 232.9697 | 161.3314
2012-10-30 18:37 | 28.6158 | 249.5582 | 195.3775 | 247.4082 | 175.6956 | -1.6043 | 1.2709 | 32.1367 | 40.8411 | 25.2874 | 2.0063 | 32.0611 | 64.3226 | 15.9806 | 14.3905 | -0.1106 | 105.1637 | 235.1677 | 180.987 | 233.0177 | 161.3051
2012-10-30 18:37 | 28.6068 | 249.5805 | 195.3567 | 247.2934 | 175.9169 | -1.5951 | 1.2708 | 32.1362 | 40.8397 | 25.2876 | 2.0063 | 32.0607 | 64.3218 | 15.9804 | 14.4141 | -0.11 | 105.1615 | 235.1664 | 180.9426 | 232.8793 | 161.5028
2012-10-30 18:37 | 28.5993 | 249.5769 | 195.3217 | 247.3956 | 175.6479 | -1.5991 | 1.2708 | 32.1362 | 40.84 | 25.2873 | 2.0062 | 32.0606 | 64.3208 | 15.9806 | 14.3862 | -0.1102 | 105.1608 | 235.1907 | 180.9355 | 233.0094 | 161.2617
2012-10-30 18:37 | 28.6053 | 249.579 | 195.3867 | 247.4603 | 175.7515 | -1.6014 | 1.2708 | 32.1358 | 40.8388 | 25.2875 | 2.0062 | 32.0602 | 64.3194 | 15.9805 | 14.4167 | -0.1104 | 105.1582 | 235.1623 | 180.97 | 233.0436 | 161.3348
2012-10-30 18:37 | 28.612 | 249.6092 | 195.3816 | 247.4044 | 175.8753 | -1.6003 | 1.2708 | 32.1358 | 40.8388 | 25.2874 | 2.0062 | 32.0602 | 64.3192 | 15.9805 | 14.3867 | -0.1103 | 105.158 | 235.2225 | 180.9949 | 233.0177 | 161.4886
2012-10-30 18:37 | 28.6083 | 249.6055 | 195.4368 | 247.3536 | 175.8482 | -1.5985 | 1.2708 | 32.135 | 40.8369 | 25.2874 | 2.0062 | 32.0595 | 64.3169 | 15.9804 | 14.3871 | -0.1102 | 105.1538 | 235.2184 | 181.0497 | 232.9665 | 161.4611
2012-10-30 18:37 | 28.609 | 249.5709 | 195.41 | 247.4092 | 175.7396 | -1.6003 | 1.2645 | 31.9727 | 40.4281 | 25.2858 | 1.9959 | 31.8951 | 63.6586 | 15.9805 | 14.4082 | -0.1103 | 104.0867 | 235.1627 | 181.0018 | 233.001 | 161.3314
2012-10-30 18:37 | 28.6008 | 249.5627 | 195.4215 | 247.4049 | 175.9929 | -1.6032 | 1.2635 | 31.9471 | 40.3638 | 25.2855 | 1.9944 | 31.8719 | 63.5657 | 15.9806 | 14.3714 | -0.1105 | 103.9295 | 235.1913 | 181.0501 | 233.0335 | 161.6215
2012-10-30 18:37 | 28.6008 | 249.5236 | 195.4018 | 247.2325 | 176.0124 | -1.6009 | 1.2634 | 31.9457 | 40.3604 | 25.2854 | 1.9943 | 31.8705 | 63.5596 | 15.9808 | 14.4652 | -0.1104 | 103.92 | 235.0584 | 180.9366 | 232.7673 | 161.5472
2012-10-30 18:37 | 28.5993 | 249.4947 | 195.4593 | 247.4778 | 176.05 | -1.5991 | 1.2635 | 31.9473 | 40.3645 | 25.2853 | 1.9944 | 31.8722 | 63.5666 | 15.9807 | 14.427 | -0.1102 | 103.9312 | 235.0677 | 181.0323 | 233.0508 | 161.623
2012-10-30 18:37 | 28.5986 | 249.5253 | 195.4428 | 247.3283 | 176.1078 | -1.5968 | 1.2635 | 31.9484 | 40.3667 | 25.2856 | 1.9945 | 31.8733 | 63.5708 | 15.9808 | 14.3896 | -0.1101 | 103.9375 | 235.1357 | 181.0532 | 232.9387 | 161.7182
2012-10-30 18:37 | 28.6016 | 249.4735 | 195.4498 | 247.3783 | 176.1694 | -1.5951 | 1.2635 | 31.9492 | 40.369 | 25.2856 | 1.9945 | 31.8741 | 63.5735 | 15.9809 | 14.4048 | -0.11 | 103.9425 | 235.0687 | 181.045 | 232.9735 | 161.7646
2012-10-30 18:37 | 28.5948 | 249.4667 | 195.5216 | 247.4694 | 176.1821 | -1.6003 | 1.2636 | 31.9504 | 40.3718 | 25.2857 | 1.9946 | 31.8752 | 63.5788 | 15.9806 | 14.398 | -0.1103 | 103.9506 | 235.0687 | 181.1236 | 233.0714 | 161.7841
2012-10-30 18:37 | 28.6001 | 249.519 | 195.5897 | 247.2318 | 176.1562 | -1.6003 | 1.2636 | 31.9509 | 40.3736 | 25.2854 | 1.9946 | 31.8758 | 63.5805 | 15.9807 | 14.4237 | -0.1103 | 103.9541 | 235.0953 | 181.166 | 232.8081 | 161.7325
2012-10-30 18:37 | 28.5971 | 249.469 | 195.516 | 247.4717 | 175.9736 | -1.6043 | 1.2637 | 31.9532 | 40.3788 | 25.2858 | 1.9948 | 31.878 | 63.5891 | 15.9808 | 14.4126 | -0.1106 | 103.9679 | 235.0564 | 181.1034 | 233.0591 | 161.561
2012-10-30 18:38 | 28.5933 | 249.4652 | 195.5515 | 247.3034 | 175.9191 | -1.6003 | 1.2637 | 31.9547 | 40.3825 | 25.2858 | 1.9949 | 31.8795 | 63.5952 | 15.9808 | 14.3639 | -0.1103 | 103.9776 | 235.1013 | 181.1876 | 232.9395 | 161.5552
2012-10-30 18:38 | 28.5933 | 249.4456 | 195.4926 | 247.3739 | 175.6692 | -1.6037 | 1.2682 | 32.0672 | 40.6661 | 25.2865 | 2.0021 | 31.9947 | 64.055 | 15.9809 | 14.421 | -0.1106 | 104.7211 | 235.0246 | 181.0716 | 232.9529 | 161.2482
2012-10-30 18:38 | 28.5963 | 249.4838 | 195.5348 | 247.3495 | 175.8947 | -1.6066 | 1.269 | 32.0884 | 40.7191 | 25.287 | 2.0032 | 32.0129 | 64.1282 | 15.9809 | 14.3506 | -0.1108 | 104.8474 | 235.1332 | 181.1842 | 232.9989 | 161.5441
2012-10-30 18:38 | 28.5881 | 249.4522 | 195.5188 | 247.2982 | 175.9997 | -1.6014 | 1.2696 | 32.1039 | 40.7585 | 25.287 | 2.0042 | 32.0284 | 64.1911 | 15.9807 | 14.3873 | -0.1104 | 104.9497 | 235.0649 | 181.1315 | 232.9109 | 161.6124
2012-10-30 18:38 | 28.5881 | 249.4678 | 195.5502 | 247.2864 | 176.0193 | -1.6032 | 1.2701 | 32.1165 | 40.7896 | 25.2876 | 2.005 | 32.0409 | 64.2411 | 15.9807 | 14.4036 | -0.1105 | 105.0307 | 235.0642 | 181.1466 | 232.8828 | 161.6157
2012-10-30 18:38 | 28.5866 | 249.4428 | 195.5487 | 247.4064 | 176.1036 | -1.6032 | 1.2703 | 32.1235 | 40.8075 | 25.2875 | 2.0054 | 32.0477 | 64.2689 | 15.9806 | 14.3082 | -0.1105 | 105.0764 | 235.1346 | 181.2405 | 233.0982 | 161.7954
2012-10-30 18:38 | 28.5911 | 249.5178 | 195.5768 | 247.1993 | 176.0223 | -1.5968 | 1.2705 | 32.1288 | 40.8212 | 25.2874 | 2.0057 | 32.0532 | 64.2901 | 15.9808 | 14.3658 | -0.1101 | 105.1113 | 235.152 | 181.211 | 232.8335 | 161.6565
2012-10-30 18:38 | 28.5941 | 249.4934 | 195.5444 | 247.3904 | 176.0565 | -1.6043 | 1.2707 | 32.1325 | 40.83 | 25.2877 | 2.006 | 32.0568 | 64.3053 | 15.9806 | 14.3606 | -0.1106 | 105.1353 | 235.1328 | 181.1838 | 233.0298 | 161.6959
2012-10-30 18:38 | 28.5843 | 249.5345 | 195.57 | 247.3101 | 176.1599 | -1.6061 | 1.2708 | 32.1347 | 40.8362 | 25.2873 | 2.0061 | 32.059 | 64.3136 | 15.9807 | 14.3549 | -0.1107 | 105.1499 | 235.1796 | 181.2151 | 232.9552 | 161.805
2012-10-30 18:38 | 28.5896 | 249.5828 | 195.6225 | 247.1782 | 176.1184 | -1.5968 | 1.2708 | 32.1369 | 40.8404 | 25.2881 | 2.0063 | 32.0612 | 64.323 | 15.9806 | 14.3765 | -0.1101 | 105.1634 | 235.2063 | 181.246 | 232.8017 | 161.7419
2012-10-30 18:38 | 28.5903 | 249.5561 | 195.6232 | 247.3748 | 176.1112 | -1.6032 | 1.2709 | 32.1384 | 40.8456 | 25.2874 | 2.0064 | 32.0628 | 64.33 | 15.9804 | 14.4139 | -0.1105 | 105.1756 | 235.1422 | 181.2093 | 232.9609 | 161.6973
2012-10-30 18:38 | 28.5933 | 249.52 | 195.6458 | 247.2838 | 176.2236 | -1.6014 | 1.271 | 32.1399 | 40.8489 | 25.2877 | 2.0065 | 32.0642 | 64.3356 | 15.9805 | 14.3761 | -0.1104 | 105.1845 | 235.1439 | 181.2697 | 232.9077 | 161.8475
2012-10-30 18:38 | 28.5858 | 249.5438 | 195.6226 | 247.2998 | 176.0638 | -1.602 | 1.271 | 32.1408 | 40.8512 | 25.2877 | 2.0065 | 32.0651 | 64.3391 | 15.9805 | 14.3768 | -0.1104 | 105.1903 | 235.167 | 181.2458 | 232.923 | 161.687
2012-10-30 18:38 | 28.5828 | 249.5956 | 195.6314 | 247.3321 | 176.256 | -1.6037 | 1.271 | 32.1416 | 40.8534 | 25.2876 | 2.0066 | 32.066 | 64.3429 | 15.9805 | 14.3942 | -0.1106 | 105.1963 | 235.2014 | 181.2372 | 232.9379 | 161.8618
2012-10-30 18:38 | 28.5858 | 249.5595 | 195.6855 | 247.2528 | 176.2434 | -1.6049 | 1.2711 | 32.1422 | 40.8547 | 25.2877 | 2.0066 | 32.0666 | 64.3455 | 15.9804 | 14.3931 | -0.1106 | 105.2002 | 235.1664 | 181.2924 | 232.8597 | 161.8503
2012-10-30 18:38 | 28.5866 | 249.572 | 195.6784 | 247.2536 | 176.1036 | -1.5997 | 1.2711 | 32.1426 | 40.856 | 25.2876 | 2.0067 | 32.067 | 64.3473 | 15.9803 | 14.4143 | -0.1103 | 105.2033 | 235.1577 | 181.2641 | 232.8393 | 161.6893
2012-10-30 18:38 | 28.5978 | 249.5402 | 195.6778 | 247.3471 | 176.1383 | -1.6014 | 1.2711 | 32.1428 | 40.8562 | 25.2876 | 2.0066 | 32.0672 | 64.3475 | 15.9805 | 14.3806 | -0.1104 | 105.2038 | 235.1596 | 181.2972 | 232.9665 | 161.7577
2012-10-30 18:38 | 28.5896 | 249.5946 | 195.6971 | 247.2409 | 176.1652 | -1.6014 | 1.2711 | 32.1431 | 40.8567 | 25.2878 | 2.0067 | 32.0674 | 64.3489 | 15.9803 | 14.4173 | -0.1104 | 105.2056 | 235.1773 | 181.2798 | 232.8236 | 161.7479
2012-10-30 18:38 | 28.5956 | 249.5614 | 195.6874 | 247.3057 | 176.1126 | -1.6037 | 1.2711 | 32.143 | 40.8567 | 25.2877 | 2.0067 | 32.0673 | 64.3485 | 15.9804 | 14.3336 | -0.1106 | 105.2052 | 235.2278 | 181.3538 | 232.9721 | 161.779
2012-10-30 18:38 | 28.5963 | 249.5856 | 195.6842 | 247.3573 | 176.0938 | -1.6066 | 1.2711 | 32.1432 | 40.8571 | 25.2877 | 2.0067 | 32.0675 | 64.3495 | 15.9803 | 14.4077 | -0.1108 | 105.2066 | 235.1779 | 181.2765 | 232.9496 | 161.6861
2012-10-30 18:38 | 28.5911 | 249.5922 | 195.7222 | 247.3325 | 175.8778 | -1.598 | 1.2711 | 32.1437 | 40.8585 | 25.2878 | 2.0067 | 32.0679 | 64.3513 | 15.9803 | 14.3699 | -0.1102 | 105.2097 | 235.2223 | 181.3523 | 232.9626 | 161.5079
2012-10-30 18:38 | 28.5918 | 249.6085 | 195.6639 | 247.2275 | 175.8902 | -1.6009 | 1.2711 | 32.1439 | 40.8591 | 25.2877 | 2.0067 | 32.0682 | 64.3525 | 15.9803 | 14.3991 | -0.1104 | 105.2116 | 235.2094 | 181.2648 | 232.8284 | 161.4911
2012-10-30 18:38 | 28.5896 | 249.5476 | 195.646 | 247.425 | 175.8217 | -1.6032 | 1.2711 | 32.144 | 40.8587 | 25.288 | 2.0067 | 32.0683 | 64.3526 | 15.9803 | 14.3684 | -0.1105 | 105.2113 | 235.1792 | 181.2776 | 233.0566 | 161.4533
2012-10-30 18:38 | 28.5903 | 249.6188 | 195.5996 | 247.2847 | 175.998 | -1.6043 | 1.2711 | 32.1442 | 40.8598 | 25.2877 | 2.0068 | 32.0685 | 64.3539 | 15.9802 | 14.3568 | -0.1106 | 105.2137 | 235.262 | 181.2428 | 232.9279 | 161.6412
2012-10-30 18:38 | 28.5888 | 249.629 | 195.5981 | 247.3381 | 175.9029 | -1.6009 | 1.2711 | 32.1441 | 40.8595 | 25.2877 | 2.0068 | 32.0684 | 64.3535 | 15.9802 | 14.3512 | -0.1104 | 105.213 | 235.2778 | 181.2469 | 232.9869 | 161.5517
2012-10-30 18:38 | 28.5881 | 249.5892 | 195.5699 | 247.3295 | 175.7811 | -1.5997 | 1.2711 | 32.1439 | 40.8591 | 25.2877 | 2.0068 | 32.0681 | 64.3529 | 15.9801 | 14.3873 | -0.1103 | 105.2119 | 235.2019 | 181.1826 | 232.9422 | 161.3938
2012-10-30 18:38 | 28.5813 | 249.602 | 195.5788 | 247.4011 | 175.868 | -1.6032 | 1.2711 | 32.1437 | 40.8582 | 25.2879 | 2.0067 | 32.068 | 64.3522 | 15.9802 | 14.3682 | -0.1105 | 105.2104 | 235.2338 | 181.2106 | 233.0329 | 161.4998
2012-10-30 18:38 | 28.5873 | 249.6001 | 195.5101 | 247.223 | 175.7921 | -1.6061 | 1.2711 | 32.1435 | 40.858 | 25.2876 | 2.0067 | 32.0678 | 64.3515 | 15.9801 | 14.3538 | -0.1107 | 105.2095 | 235.2463 | 181.1563 | 232.8692 | 161.4383
2012-10-30 18:38 | 28.5896 | 249.5711 | 195.5242 | 247.4211 | 175.7358 | -1.6003 | 1.2711 | 32.143 | 40.8572 | 25.2875 | 2.0067 | 32.0674 | 64.35 | 15.9801 | 14.352 | -0.1103 | 105.2072 | 235.2191 | 181.1722 | 233.0691 | 161.3838
2012-10-30 18:39 | 28.5933 | 249.6414 | 195.5318 | 247.3504 | 175.8722 | -1.5957 | 1.2711 | 32.1432 | 40.8573 | 25.2877 | 2.0067 | 32.0676 | 64.3512 | 15.98 | 14.368 | -0.11 | 105.2085 | 235.2734 | 181.1638 | 232.9824 | 161.5042
2012-10-30 18:39 | 28.5993 | 249.62 | 195.5418 | 247.2467 | 175.8158 | -1.5968 | 1.2711 | 32.1433 | 40.8577 | 25.2877 | 2.0067 | 32.0677 | 64.3514 | 15.98 | 14.3862 | -0.1101 | 105.209 | 235.2338 | 181.1556 | 232.8605 | 161.4296
2012-10-30 18:39 | 28.6046 | 249.6253 | 195.4881 | 247.4714 | 175.7859 | -1.5997 | 1.2711 | 32.1434 | 40.8577 | 25.2878 | 2.0067 | 32.0678 | 64.3516 | 15.9801 | 14.4038 | -0.1103 | 105.2093 | 235.2215 | 181.0843 | 233.0676 | 161.3821
2012-10-30 18:39 | 28.6023 | 249.6386 | 195.5251 | 247.2497 | 175.932 | -1.602 | 1.2711 | 32.1432 | 40.8572 | 25.2878 | 2.0067 | 32.0675 | 64.3506 | 15.98 | 14.3892 | -0.1104 | 105.2078 | 235.2494 | 181.1359 | 232.8605 | 161.5428
2012-10-30 18:39 | 28.5941 | 249.6187 | 195.4619 | 247.4217 | 175.6739 | -1.6026 | 1.2711 | 32.143 | 40.8565 | 25.2878 | 2.0067 | 32.0673 | 64.3491 | 15.9802 | 14.3525 | -0.1105 | 105.2056 | 235.2662 | 181.1094 | 233.0692 | 161.3214
2012-10-30 18:39 | 28.5971 | 249.6412 | 195.4571 | 247.4325 | 175.7316 | -1.598 | 1.2711 | 32.1427 | 40.8565 | 25.2874 | 2.0067 | 32.0672 | 64.3492 | 15.9801 | 14.3595 | -0.1102 | 105.2057 | 235.2817 | 181.0976 | 233.073 | 161.3721
2012-10-30 18:39 | 28.6023 | 249.6504 | 195.478 | 247.3281 | 175.8773 | -1.6037 | 1.2711 | 32.1426 | 40.8556 | 25.2878 | 2.0067 | 32.0669 | 64.3482 | 15.98 | 14.3484 | -0.1106 | 105.2038 | 235.302 | 181.1296 | 232.9797 | 161.5289
2012-10-30 18:39 | 28.6068 | 249.6509 | 195.4589 | 247.4227 | 175.8428 | -1.5997 | 1.2711 | 32.1423 | 40.8551 | 25.2876 | 2.0067 | 32.0667 | 64.3478 | 15.9799 | 14.3652 | -0.1103 | 105.2029 | 235.2857 | 181.0937 | 233.0575 | 161.4776
2012-10-30 18:39 | 28.6053 | 249.622 | 195.4849 | 247.4368 | 175.8374 | -1.6009 | 1.2711 | 32.1424 | 40.8547 | 25.2879 | 2.0067 | 32.0667 | 64.3474 | 15.98 | 14.3473 | -0.1104 | 105.2021 | 235.2747 | 181.1376 | 233.0895 | 161.4901
2012-10-30 18:39 | 28.6031 | 249.6551 | 195.467 | 247.2858 | 175.9757 | -1.6014 | 1.2711 | 32.1422 | 40.8547 | 25.2877 | 2.0067 | 32.0665 | 64.3469 | 15.98 | 14.3411 | -0.1104 | 105.2015 | 235.314 | 181.1259 | 232.9447 | 161.6346
2012-10-30 18:39 | 28.6083 | 249.629 | 195.4918 | 247.4281 | 175.8287 | -1.598 | 1.271 | 32.1417 | 40.8534 | 25.2876 | 2.0067 | 32.066 | 64.3453 | 15.9798 | 14.3259 | -0.1102 | 105.1987 | 235.3031 | 181.1659 | 233.1022 | 161.5028
2012-10-30 18:39 | 28.6008 | 249.6684 | 195.4804 | 247.4558 | 175.9383 | -1.5968 | 1.271 | 32.1416 | 40.8534 | 25.2875 | 2.0067 | 32.066 | 64.3455 | 15.9798 | 14.3551 | -0.1101 | 105.1989 | 235.3133 | 181.1253 | 233.1007 | 161.5832
2012-10-30 18:39 | 28.6038 | 249.6988 | 195.5384 | 247.3217 | 175.8827 | -1.6061 | 1.271 | 32.1413 | 40.8527 | 25.2876 | 2.0066 | 32.0657 | 64.3443 | 15.9798 | 14.3866 | -0.1107 | 105.197 | 235.3122 | 181.1518 | 232.9351 | 161.4961
2012-10-30 18:39 | 28.5978 | 249.6302 | 195.4931 | 247.4881 | 175.8533 | -1.6009 | 1.271 | 32.1411 | 40.8518 | 25.2878 | 2.0066 | 32.0655 | 64.3434 | 15.9798 | 14.3602 | -0.1104 | 105.1952 | 235.27 | 181.1329 | 233.1279 | 161.4931
2012-10-30 18:39 | 28.5978 | 249.6654 | 195.5285 | 247.3902 | 175.9626 | -1.6037 | 1.271 | 32.141 | 40.8516 | 25.2877 | 2.0066 | 32.0653 | 64.3426 | 15.9799 | 14.3521 | -0.1106 | 105.1942 | 235.3133 | 181.1764 | 233.0381 | 161.6105
2012-10-30 18:39 | 28.6038 | 249.6714 | 195.5109 | 247.3766 | 175.7734 | -1.6003 | 1.271 | 32.1404 | 40.8498 | 25.2878 | 2.0066 | 32.0646 | 64.3396 | 15.9799 | 14.354 | -0.1103 | 105.1894 | 235.3174 | 181.1569 | 233.0226 | 161.4194
2012-10-30 18:39 | 28.5956 | 249.6437 | 195.4752 | 247.4663 | 175.7145 | -1.6032 | 1.271 | 32.1399 | 40.849 | 25.2876 | 2.0065 | 32.0642 | 64.338 | 15.9799 | 14.3458 | -0.1105 | 105.187 | 235.2979 | 181.1294 | 233.1205 | 161.3687
2012-10-30 18:39 | 28.5971 | 249.6334 | 195.5121 | 247.4443 | 175.8331 | -1.6055 | 1.271 | 32.14 | 40.8497 | 25.2874 | 2.0066 | 32.0644 | 64.3393 | 15.9797 | 14.3759 | -0.1107 | 105.189 | 235.2575 | 181.1362 | 233.0684 | 161.4572
2012-10-30 18:39 | 28.6023 | 249.6778 | 195.474 | 247.332 | 175.7329 | -1.5985 | 1.271 | 32.1399 | 40.8493 | 25.2875 | 2.0065 | 32.0643 | 64.3384 | 15.9799 | 14.3443 | -0.1102 | 105.1877 | 235.3335 | 181.1297 | 232.9877 | 161.3886
2012-10-30 18:39 | 28.6038 | 249.6988 | 195.4638 | 247.5411 | 175.6056 | -1.6032 | 1.271 | 32.1396 | 40.8483 | 25.2876 | 2.0065 | 32.064 | 64.3378 | 15.9797 | 14.3377 | -0.1105 | 105.186 | 235.3611 | 181.1261 | 233.2034 | 161.2679
2012-10-30 18:39 | 28.6038 | 249.6988 | 195.4598 | 247.4001 | 175.6797 | -1.5974 | 1.2709 | 32.1394 | 40.8471 | 25.2879 | 2.0065 | 32.0637 | 64.3369 | 15.9797 | 14.3703 | -0.1101 | 105.184 | 235.3285 | 181.0895 | 233.0298 | 161.3094
2012-10-30 18:39 | 28.5963 | 249.7305 | 195.4445 | 247.42 | 175.7698 | -1.602 | 1.2709 | 32.1388 | 40.8461 | 25.2876 | 2.0065 | 32.0632 | 64.3348 | 15.9797 | 14.3343 | -0.1104 | 105.1809 | 235.3962 | 181.1102 | 233.0857 | 161.4355
2012-10-30 18:39 | 28.5993 | 249.6748 | 195.4042 | 247.5209 | 175.6011 | -1.5997 | 1.2709 | 32.1385 | 40.8455 | 25.2875 | 2.0065 | 32.0627 | 64.3328 | 15.9797 | 14.3617 | -0.1103 | 105.1783 | 235.3131 | 181.0425 | 233.1592 | 161.2394
2012-10-30 18:39 | 28.5986 | 249.7171 | 195.4389 | 247.4458 | 175.6355 | -1.5974 | 1.2709 | 32.1384 | 40.8451 | 25.2877 | 2.0065 | 32.0628 | 64.3334 | 15.9796 | 14.3325 | -0.1101 | 105.1785 | 235.3846 | 181.1064 | 233.1133 | 161.303
2012-10-30 18:39 | 28.5993 | 249.6748 | 195.3885 | 247.376 | 175.6206 | -1.598 | 1.2709 | 32.1383 | 40.8453 | 25.2874 | 2.0065 | 32.0627 | 64.3331 | 15.9796 | 14.3495 | -0.1102 | 105.1783 | 235.3253 | 181.039 | 233.0265 | 161.2711
2012-10-30 18:39 | 28.5993 | 249.6787 | 195.3728 | 247.5444 | 175.5308 | -1.5985 | 1.2709 | 32.1381 | 40.8442 | 25.2877 | 2.0065 | 32.0625 | 64.332 | 15.9797 | 14.374 | -0.1102 | 105.1762 | 235.3047 | 180.9988 | 233.1704 | 161.1568
2012-10-30 18:39 | 28.5986 | 249.6506 | 195.3957 | 247.434 | 175.6355 | -1.602 | 1.2709 | 32.1382 | 40.845 | 25.2874 | 2.0065 | 32.0626 | 64.3328 | 15.9796 | 14.3692 | -0.1104 | 105.1778 | 235.2814 | 181.0265 | 233.0648 | 161.2663
2012-10-30 18:39 | 28.6083 | 249.7033 | 195.37 | 247.4163 | 175.6413 | -1.6072 | 1.2709 | 32.138 | 40.8444 | 25.2874 | 2.0064 | 32.0622 | 64.3307 | 15.9797 | 14.334 | -0.1108 | 105.1751 | 235.3693 | 181.036 | 233.0823 | 161.3073
2012-10-30 18:40 | 28.606 | 249.6697 | 195.3088 | 247.4963 | 175.561 | -1.5991 | 1.2709 | 32.1377 | 40.8438 | 25.2873 | 2.0064 | 32.062 | 64.3305 | 15.9795 | 14.3317 | -0.1102 | 105.1743 | 235.338 | 180.9771 | 233.1646 | 161.2293
2012-10-30 18:40 | 28.606 | 249.6619 | 195.407 | 247.4885 | 175.6858 | -1.5985 | 1.2709 | 32.137 | 40.8413 | 25.2877 | 2.0064 | 32.0613 | 64.3283 | 15.9794 | 14.3725 | -0.1102 | 105.1696 | 235.2894 | 181.0345 | 233.116 | 161.3133
2012-10-30 18:40 | 28.5986 | 249.7132 | 195.3878 | 247.3714 | 175.6745 | -1.5945 | 1.2708 | 32.1366 | 40.8407 | 25.2875 | 2.0064 | 32.0609 | 64.326 | 15.9796 | 14.3937 | -0.1099 | 105.1666 | 235.3195 | 180.9941 | 232.9777 | 161.2808
2012-10-30 18:40 | 28.6001 | 249.6325 | 195.3304 | 247.557 | 175.7229 | -1.5985 | 1.2709 | 32.1366 | 40.8409 | 25.2873 | 2.0064 | 32.0609 | 64.3262 | 15.9795 | 14.334 | -0.1102 | 105.1671 | 235.2985 | 180.9964 | 233.223 | 161.3889
2012-10-30 18:40 | 28.6046 | 249.6957 | 195.382 | 247.4205 | 175.9147 | -1.6101 | 1.2684 | 32.075 | 40.6852 | 25.287 | 2.0025 | 31.9989 | 64.077 | 15.9796 | 14.367 | -0.111 | 104.7623 | 235.3287 | 181.015 | 233.0535 | 161.5477
2012-10-30 18:40 | 28.5903 | 249.6814 | 195.3402 | 247.461 | 175.7014 | -1.5997 | 1.2648 | 31.9828 | 40.4535 | 25.2859 | 1.9967 | 31.9071 | 63.7101 | 15.9796 | 14.3201 | -0.1103 | 104.1636 | 235.3613 | 181.0201 | 233.1409 | 161.3813
2012-10-30 18:40 | 28.6075 | 249.6438 | 195.3653 | 247.533 | 175.6991 | -1.6055 | 1.2645 | 31.9746 | 40.4323 | 25.2861 | 1.9962 | 31.8993 | 63.6788 | 15.9796 | 14.3781 | -0.1107 | 104.1111 | 235.2657 | 180.9872 | 233.1549 | 161.321
2012-10-30 18:40 | 28.6046 | 249.6761 | 195.3663 | 247.3069 | 175.8445 | -1.6003 | 1.264 | 31.9619 | 40.401 | 25.2857 | 1.9955 | 31.8868 | 63.6286 | 15.9797 | 14.4119 | -0.1103 | 104.0295 | 235.2642 | 180.9544 | 232.895 | 161.4326
2012-10-30 18:40 | 28.5978 | 249.6459 | 195.3202 | 247.5038 | 175.7557 | -1.5991 | 1.2637 | 31.9543 | 40.3819 | 25.2855 | 1.995 | 31.8792 | 63.5982 | 15.9797 | 14.4174 | -0.1102 | 103.9801 | 235.2285 | 180.9028 | 233.0864 | 161.3383
2012-10-30 18:40 | 28.5941 | 249.5952 | 195.3833 | 247.5118 | 175.8535 | -1.602 | 1.2635 | 31.949 | 40.3683 | 25.2856 | 1.9946 | 31.8739 | 63.5767 | 15.9798 | 14.3688 | -0.1104 | 103.9451 | 235.2264 | 181.0145 | 233.143 | 161.4847
2012-10-30 18:40 | 28.5956 | 249.6006 | 195.4241 | 247.4546 | 176.0541 | -1.5997 | 1.2635 | 31.9478 | 40.3653 | 25.2856 | 1.9945 | 31.8726 | 63.5713 | 15.9798 | 14.3825 | -0.1103 | 103.9366 | 235.2181 | 181.0416 | 233.0721 | 161.6716
2012-10-30 18:40 | 28.5873 | 249.5923 | 195.3215 | 247.3679 | 175.8974 | -1.598 | 1.2635 | 31.9486 | 40.3673 | 25.2856 | 1.9946 | 31.8734 | 63.5744 | 15.9799 | 14.3824 | -0.1102 | 103.9417 | 235.2099 | 180.9391 | 232.9855 | 161.515
2012-10-30 18:40 | 28.5941 | 249.5795 | 195.4187 | 247.5196 | 175.9082 | -1.602 | 1.2635 | 31.9483 | 40.3661 | 25.2859 | 1.9946 | 31.8731 | 63.5734 | 15.9799 | 14.3688 | -0.1104 | 103.9395 | 235.2107 | 181.0499 | 233.1508 | 161.5394
2012-10-30 18:40 | 28.5933 | 249.5905 | 195.4218 | 247.3034 | 176.0206 | -1.6037 | 1.2636 | 31.9514 | 40.3746 | 25.2856 | 1.9948 | 31.8761 | 63.5849 | 15.98 | 14.3761 | -0.1106 | 103.9595 | 235.2144 | 181.0457 | 232.9273 | 161.6445
2012-10-30 18:40 | 28.5896 | 249.5711 | 195.4024 | 247.4094 | 175.8256 | -1.5991 | 1.2637 | 31.9542 | 40.3816 | 25.2855 | 1.9949 | 31.879 | 63.5964 | 15.98 | 14.3806 | -0.1102 | 103.978 | 235.1905 | 181.0218 | 233.0288 | 161.445
2012-10-30 18:40 | 28.5926 | 249.5858 | 195.4133 | 247.5181 | 175.8832 | -1.6032 | 1.2651 | 31.9899 | 40.4717 | 25.2857 | 1.9971 | 31.9131 | 63.7326 | 15.98 | 14.351 | -0.1105 | 104.2043 | 235.2348 | 181.0623 | 233.1671 | 161.5322
2012-10-30 18:40 | 28.606 | 249.5718 | 195.4463 | 247.367 | 175.8225 | -1.602 | 1.2687 | 32.081 | 40.7004 | 25.287 | 2.0029 | 32.0053 | 64.102 | 15.9798 | 14.3929 | -0.1104 | 104.8024 | 235.1789 | 181.0534 | 232.9741 | 161.4296
2012-10-30 18:40 | 28.5933 | 249.6022 | 195.3943 | 247.3386 | 175.7902 | -1.5957 | 1.2692 | 32.0936 | 40.7322 | 25.2871 | 2.0037 | 32.0183 | 64.1538 | 15.9799 | 14.3843 | -0.11 | 104.886 | 235.2179 | 181.01 | 232.9543 | 161.4059
2012-10-30 18:40 | 28.5933 | 249.567 | 195.4179 | 247.4366 | 175.9034 | -1.6026 | 1.2697 | 32.1056 | 40.7635 | 25.2866 | 2.0044 | 32.0299 | 64.2001 | 15.98 | 14.3394 | -0.1105 | 104.9636 | 235.2276 | 181.0785 | 233.0972 | 161.564
2012-10-30 18:40 | 28.5948 | 249.588 | 195.4941 | 247.2344 | 175.8893 | -1.5985 | 1.2701 | 32.1163 | 40.7896 | 25.2872 | 2.0051 | 32.0406 | 64.2438 | 15.9798 | 14.3572 | -0.1102 | 105.0334 | 235.2308 | 181.1369 | 232.8772 | 161.5321
2012-10-30 18:40 | 28.5926 | 249.5937 | 195.4604 | 247.3771 | 175.8793 | -1.6003 | 1.2704 | 32.1254 | 40.8128 | 25.2873 | 2.0057 | 32.0498 | 64.281 | 15.9797 | 14.3958 | -0.1103 | 105.0937 | 235.1979 | 181.0646 | 232.9813 | 161.4835
2012-10-30 18:40 | 28.6016 | 249.544 | 195.4576 | 247.4684 | 175.8961 | -1.5997 | 1.2706 | 32.1313 | 40.8272 | 25.2877 | 2.006 | 32.0556 | 64.3043 | 15.9797 | 14.3763 | -0.1103 | 105.1314 | 235.1677 | 181.0813 | 233.0921 | 161.5198
2012-10-30 18:40 | 28.5948 | 249.6194 | 195.4941 | 247.2383 | 175.9362 | -1.5997 | 1.2708 | 32.1343 | 40.8351 | 25.2875 | 2.0062 | 32.0585 | 64.3164 | 15.9796 | 14.3572 | -0.1103 | 105.1515 | 235.2622 | 181.1369 | 232.8811 | 161.579
2012-10-30 18:40 | 28.5963 | 249.6326 | 195.4602 | 247.467 | 175.6996 | -1.5991 | 1.2708 | 32.1361 | 40.8395 | 25.2875 | 2.0063 | 32.0605 | 64.3233 | 15.9798 | 14.4281 | -0.1102 | 105.1627 | 235.2045 | 181.0321 | 233.0389 | 161.2715
2012-10-30 18:40 | 28.5873 | 249.6119 | 195.463 | 247.27 | 175.9287 | -1.6003 | 1.2709 | 32.1383 | 40.8447 | 25.2877 | 2.0065 | 32.0626 | 64.3322 | 15.9798 | 14.3661 | -0.1103 | 105.1769 | 235.2458 | 181.0969 | 232.9039 | 161.5626
2012-10-30 18:40 | 28.5873 | 249.5884 | 195.4826 | 247.317 | 175.7764 | -1.6003 | 1.271 | 32.1398 | 40.8485 | 25.2877 | 2.0066 | 32.064 | 64.3383 | 15.9796 | 14.3538 | -0.1103 | 105.1867 | 235.2346 | 181.1288 | 232.9632 | 161.4226
2012-10-30 18:40 | 28.5858 | 249.5751 | 195.5086 | 247.4526 | 175.6852 | -1.6026 | 1.271 | 32.141 | 40.8515 | 25.2877 | 2.0066 | 32.0652 | 64.3434 | 15.9795 | 14.389 | -0.1105 | 105.195 | 235.1861 | 181.1196 | 233.0636 | 161.2962
2012-10-30 18:40 | 28.5851 | 249.5705 | 195.4844 | 247.2521 | 175.8445 | -1.6084 | 1.271 | 32.1413 | 40.8528 | 25.2874 | 2.0067 | 32.0657 | 64.3451 | 15.9796 | 14.3516 | -0.1109 | 105.198 | 235.2189 | 181.1328 | 232.9005 | 161.4929
2012-10-30 18:40 | 28.5881 | 249.5813 | 195.452 | 247.3413 | 175.7343 | -1.6049 | 1.2711 | 32.142 | 40.8543 | 25.2877 | 2.0067 | 32.0663 | 64.3481 | 15.9795 | 14.3791 | -0.1106 | 105.2023 | 235.2022 | 181.0729 | 232.9622 | 161.3552
2012-10-30 18:41 | 28.5821 | 249.6262 | 195.4263 | 247.4332 | 175.779 | -1.5997 | 1.2711 | 32.1426 | 40.8555 | 25.2879 | 2.0068 | 32.0669 | 64.3504 | 15.9795 | 14.3813 | -0.1103 | 105.2059 | 235.2449 | 181.045 | 233.0519 | 161.3977
2012-10-30 18:41 | 28.5806 | 249.613 | 195.4523 | 247.3377 | 175.8439 | -1.6055 | 1.2711 | 32.1432 | 40.8571 | 25.2878 | 2.0068 | 32.0675 | 64.3523 | 15.9796 | 14.339 | -0.1107 | 105.2093 | 235.274 | 181.1133 | 232.9987 | 161.5049
2012-10-30 18:41 | 28.5821 | 249.6497 | 195.4381 | 247.2569 | 175.7322 | -1.598 | 1.2711 | 32.1435 | 40.8584 | 25.2875 | 2.0068 | 32.0679 | 64.3542 | 15.9795 | 14.3405 | -0.1102 | 105.2126 | 235.3092 | 181.0976 | 232.9164 | 161.3917
2012-10-30 18:41 | 28.5836 | 249.6003 | 195.3964 | 247.4739 | 175.5698 | -1.6043 | 1.2711 | 32.1437 | 40.8586 | 25.2876 | 2.0068 | 32.068 | 64.3545 | 15.9796 | 14.4072 | -0.1106 | 105.2131 | 235.1931 | 180.9892 | 233.0667 | 161.1626
2012-10-30 18:41 | 28.5761 | 249.6046 | 195.4636 | 247.2666 | 175.7418 | -1.5997 | 1.2711 | 32.1443 | 40.8595 | 25.2879 | 2.0068 | 32.0685 | 64.3563 | 15.9796 | 14.3222 | -0.1103 | 105.2159 | 235.2824 | 181.1414 | 232.9444 | 161.4196
2012-10-30 18:41 | 28.5821 | 249.6145 | 195.3988 | 247.4097 | 175.6815 | -1.5985 | 1.2711 | 32.1444 | 40.8603 | 25.2877 | 2.0069 | 32.0687 | 64.3588 | 15.9792 | 14.3609 | -0.1102 | 105.2191 | 235.2536 | 181.0379 | 233.0488 | 161.3206
2012-10-30 18:41 | 28.5828 | 249.6504 | 195.4388 | 247.3517 | 175.7876 | -1.6061 | 1.2711 | 32.1442 | 40.8592 | 25.2881 | 2.0069 | 32.0685 | 64.3573 | 15.9794 | 14.386 | -0.1107 | 105.2165 | 235.2644 | 181.0528 | 232.9657 | 161.4016
2012-10-30 18:41 | 28.5798 | 249.6318 | 195.4044 | 247.3291 | 175.8119 | -1.5968 | 1.2711 | 32.1443 | 40.8601 | 25.2876 | 2.0069 | 32.0684 | 64.3572 | 15.9793 | 14.3382 | -0.1101 | 105.2174 | 235.2936 | 181.0662 | 232.9909 | 161.4737
2012-10-30 18:41 | 28.5873 | 249.6158 | 195.404 | 247.4384 | 175.6906 | -1.5997 | 1.2711 | 32.1445 | 40.8604 | 25.2877 | 2.0069 | 32.0688 | 64.3585 | 15.9793 | 14.3824 | -0.1103 | 105.2189 | 235.2334 | 181.0216 | 233.056 | 161.3082
2012-10-30 18:41 | 28.5851 | 249.5979 | 195.4804 | 247.3579 | 175.946 | -1.6043 | 1.2712 | 32.1446 | 40.861 | 25.2876 | 2.0069 | 32.0688 | 64.3585 | 15.9794 | 14.3639 | -0.1106 | 105.2195 | 235.234 | 181.1165 | 232.994 | 161.5821
2012-10-30 18:41 | 28.5896 | 249.6455 | 195.4692 | 247.3036 | 175.8217 | -1.6003 | 1.2711 | 32.1442 | 40.8599 | 25.2876 | 2.0069 | 32.0684 | 64.3574 | 15.9793 | 14.3928 | -0.1103 | 105.2174 | 235.2527 | 181.0764 | 232.9108 | 161.4289
2012-10-30 18:41 | 28.5836 | 249.6082 | 195.475 | 247.47 | 175.8703 | -1.6066 | 1.2711 | 32.1443 | 40.8598 | 25.2877 | 2.0069 | 32.0685 | 64.3572 | 15.9794 | 14.3868 | -0.1108 | 105.2171 | 235.2214 | 181.0882 | 233.0832 | 161.4835
2012-10-30 18:41 | 28.5753 | 249.6194 | 195.4981 | 247.3637 | 175.9791 | -1.6026 | 1.2711 | 32.1441 | 40.8594 | 25.2879 | 2.0069 | 32.0684 | 64.3569 | 15.9793 | 14.3622 | -0.1105 | 105.2163 | 235.2572 | 181.1359 | 233.0015 | 161.6169
2012-10-30 18:41 | 28.5708 | 249.6345 | 195.4465 | 247.2926 | 175.8146 | -1.5997 | 1.2711 | 32.144 | 40.8593 | 25.2877 | 2.0069 | 32.0682 | 64.357 | 15.9792 | 14.423 | -0.1103 | 105.2163 | 235.2115 | 181.0235 | 232.8696 | 161.3916
2012-10-30 18:41 | 28.5828 | 249.6387 | 195.5096 | 247.4574 | 175.8422 | -1.602 | 1.2711 | 32.1443 | 40.8599 | 25.2878 | 2.0069 | 32.0686 | 64.3581 | 15.9793 | 14.4024 | -0.1104 | 105.218 | 235.2363 | 181.1072 | 233.055 | 161.4398
2012-10-30 18:41 | 28.5828 | 249.6857 | 195.4899 | 247.2576 | 175.8812 | -1.6026 | 1.2711 | 32.1442 | 40.8594 | 25.2879 | 2.0069 | 32.0684 | 64.3573 | 15.9793 | 14.3738 | -0.1105 | 105.2167 | 235.3119 | 181.1161 | 232.8838 | 161.5074
2012-10-30 18:41 | 28.5851 | 249.6018 | 195.5236 | 247.5146 | 175.8484 | -1.602 | 1.2711 | 32.1442 | 40.8595 | 25.2879 | 2.0069 | 32.0684 | 64.3575 | 15.9792 | 14.3679 | -0.1104 | 105.217 | 235.2339 | 181.1557 | 233.1467 | 161.4805
2012-10-30 18:41 | 28.5881 | 249.6518 | 195.5384 | 247.3883 | 175.9958 | -1.5968 | 1.2711 | 32.144 | 40.8594 | 25.2877 | 2.0069 | 32.0683 | 64.3568 | 15.9793 | 14.3709 | -0.1101 | 105.2162 | 235.2809 | 181.1675 | 233.0174 | 161.6249
2012-10-30 18:41 | 28.5873 | 249.6784 | 195.5023 | 247.3091 | 175.9794 | -1.6049 | 1.2711 | 32.1439 | 40.8587 | 25.2879 | 2.0069 | 32.0682 | 64.3565 | 15.9792 | 14.362 | -0.1106 | 105.2152 | 235.3164 | 181.1403 | 232.9471 | 161.6174
2012-10-30 18:41 | 28.5851 | 249.6449 | 195.5433 | 247.5028 | 175.8055 | -1.598 | 1.2711 | 32.1437 | 40.8586 | 25.2876 | 2.0069 | 32.0679 | 64.3556 | 15.9792 | 14.3924 | -0.1102 | 105.2142 | 235.2525 | 181.1509 | 233.1104 | 161.4131
2012-10-30 18:41 | 28.5896 | 249.6925 | 195.5242 | 247.3075 | 175.8334 | -1.602 | 1.2711 | 32.1433 | 40.8573 | 25.2878 | 2.0068 | 32.0675 | 64.3538 | 15.9792 | 14.3643 | -0.1104 | 105.2112 | 235.3282 | 181.1599 | 232.9432 | 161.4691
2012-10-30 18:41 | 28.5911 | 249.647 | 195.4628 | 247.4422 | 175.6358 | -1.6066 | 1.2711 | 32.1431 | 40.857 | 25.2876 | 2.0068 | 32.0674 | 64.3535 | 15.9792 | 14.4107 | -0.1108 | 105.2105 | 235.2363 | 181.0521 | 233.0315 | 161.2251
2012-10-30 18:41 | 28.5971 | 249.6334 | 195.5396 | 247.507 | 175.8409 | -1.5991 | 1.2711 | 32.1427 | 40.8559 | 25.2877 | 2.0068 | 32.0669 | 64.3523 | 15.979 | 14.3881 | -0.1102 | 105.2082 | 235.2453 | 181.1515 | 233.1189 | 161.4528
2012-10-30 18:41 | 28.5903 | 249.6932 | 195.4581 | 247.3396 | 175.8263 | -1.5997 | 1.2711 | 32.1428 | 40.856 | 25.2878 | 2.0068 | 32.0671 | 64.3524 | 15.9791 | 14.3895 | -0.1103 | 105.2085 | 235.3037 | 181.0686 | 232.9501 | 161.4368
2012-10-30 18:41 | 28.5941 | 249.6539 | 195.5051 | 247.4531 | 175.7286 | -1.6055 | 1.2711 | 32.1427 | 40.8561 | 25.2876 | 2.0068 | 32.0669 | 64.3525 | 15.979 | 14.3443 | -0.1107 | 105.2087 | 235.3096 | 181.1608 | 233.1088 | 161.3843
2012-10-30 18:41 | 28.5918 | 249.7221 | 195.5028 | 247.4625 | 175.7809 | -1.5985 | 1.2711 | 32.1424 | 40.8555 | 25.2876 | 2.0068 | 32.0668 | 64.3516 | 15.979 | 14.3583 | -0.1102 | 105.2071 | 235.3638 | 181.1445 | 233.1042 | 161.4226
2012-10-30 18:41 | 28.5963 | 249.6483 | 195.4759 | 247.3965 | 175.6957 | -1.5905 | 1.271 | 32.1416 | 40.8534 | 25.2876 | 2.0068 | 32.0659 | 64.3484 | 15.979 | 14.3995 | -0.1096 | 105.2018 | 235.2488 | 181.0764 | 232.997 | 161.2962
2012-10-30 18:41 | 28.5836 | 249.6708 | 195.4946 | 247.5405 | 175.7688 | -1.6014 | 1.271 | 32.1415 | 40.8535 | 25.2874 | 2.0067 | 32.0658 | 64.3478 | 15.979 | 14.4154 | -0.1104 | 105.2013 | 235.2554 | 181.0792 | 233.1251 | 161.3534
2012-10-30 18:42 | 28.5963 | 249.6678 | 195.4602 | 247.4396 | 175.9299 | -1.598 | 1.271 | 32.1415 | 40.853 | 25.2876 | 2.0067 | 32.0658 | 64.3475 | 15.9791 | 14.3873 | -0.1102 | 105.2005 | 235.2805 | 181.0729 | 233.0523 | 161.5426
2012-10-30 18:42 | 28.6016 | 249.7201 | 195.4576 | 247.3626 | 175.9976 | -1.602 | 1.271 | 32.1413 | 40.8529 | 25.2874 | 2.0067 | 32.0656 | 64.3474 | 15.9789 | 14.4008 | -0.1104 | 105.2003 | 235.3193 | 181.0568 | 232.9618 | 161.5968
2012-10-30 18:42 | 28.6113 | 249.6789 | 195.4948 | 247.5643 | 175.9566 | -1.5997 | 1.271 | 32.1414 | 40.8527 | 25.2877 | 2.0067 | 32.0656 | 64.3472 | 15.979 | 14.3901 | -0.1103 | 105.1999 | 235.2888 | 181.1047 | 233.1742 | 161.5665
2012-10-30 18:42 | 28.5948 | 249.7212 | 195.5294 | 247.3166 | 176.0689 | -1.598 | 1.271 | 32.1412 | 40.8525 | 25.2875 | 2.0067 | 32.0654 | 64.3471 | 15.9789 | 14.4103 | -0.1102 | 105.1996 | 235.3109 | 181.1191 | 232.9063 | 161.6586
2012-10-30 18:42 | 28.5948 | 249.6546 | 195.5098 | 247.5713 | 175.8113 | -1.6055 | 1.271 | 32.1407 | 40.8512 | 25.2875 | 2.0067 | 32.065 | 64.3454 | 15.9789 | 14.394 | -0.1107 | 105.1966 | 235.2606 | 181.1158 | 233.1773 | 161.4173
2012-10-30 18:42 | 28.6023 | 249.6856 | 195.533 | 247.4338 | 175.9242 | -1.6009 | 1.271 | 32.1405 | 40.8502 | 25.2877 | 2.0067 | 32.0647 | 64.3441 | 15.9788 | 14.3851 | -0.1104 | 105.1942 | 235.3005 | 181.1479 | 233.0487 | 161.5391
2012-10-30 18:42 | 28.5903 | 249.7323 | 195.5642 | 247.4453 | 176.123 | -1.6014 | 1.271 | 32.1398 | 40.8482 | 25.2879 | 2.0067 | 32.064 | 64.3414 | 15.9788 | 14.4017 | -0.1104 | 105.1895 | 235.3306 | 181.1625 | 233.0436 | 161.7213
2012-10-30 18:42 | 28.6023 | 249.7208 | 195.5723 | 247.4299 | 175.9164 | -1.5957 | 1.271 | 32.1397 | 40.8484 | 25.2877 | 2.0067 | 32.0639 | 64.3413 | 15.9788 | 14.3892 | -0.11 | 105.1897 | 235.3316 | 181.1831 | 233.0407 | 161.5272
2012-10-30 18:42 | 28.5986 | 249.6741 | 195.5843 | 247.6103 | 176.018 | -1.5991 | 1.271 | 32.1395 | 40.8478 | 25.2877 | 2.0066 | 32.0638 | 64.3401 | 15.9789 | 14.3814 | -0.1102 | 105.1879 | 235.2927 | 181.2029 | 233.2289 | 161.6366
2012-10-30 18:42 | 28.5963 | 249.7305 | 195.6095 | 247.4239 | 175.9064 | -1.602 | 1.271 | 32.1393 | 40.8478 | 25.2873 | 2.0066 | 32.0636 | 64.3402 | 15.9787 | 14.3914 | -0.1104 | 105.188 | 235.3391 | 181.2181 | 233.0325 | 161.515
2012-10-30 18:42 | 28.5978 | 249.7085 | 195.5914 | 247.4646 | 175.5918 | -1.6009 | 1.2709 | 32.139 | 40.8467 | 25.2875 | 2.0066 | 32.0632 | 64.3387 | 15.9787 | 14.3725 | -0.1104 | 105.1854 | 235.336 | 181.2189 | 233.0921 | 161.2193
2012-10-30 18:42 | 28.5956 | 249.6437 | 195.5734 | 247.5447 | 175.7691 | -1.6009 | 1.2709 | 32.1385 | 40.8455 | 25.2876 | 2.0066 | 32.0629 | 64.3371 | 15.9788 | 14.3662 | -0.1104 | 105.1826 | 235.2775 | 181.2072 | 233.1785 | 161.4029
2012-10-30 18:42 | 28.5948 | 249.7172 | 195.5255 | 247.4028 | 175.8269 | -1.5974 | 1.2709 | 32.1381 | 40.8447 | 25.2874 | 2.0066 | 32.0624 | 64.3354 | 15.9787 | 14.3858 | -0.1101 | 105.1802 | 235.3314 | 181.1397 | 233.017 | 161.4411
2012-10-30 18:42 | 28.5978 | 249.6967 | 195.5128 | 247.5782 | 175.826 | -1.5985 | 1.2709 | 32.1377 | 40.8432 | 25.2877 | 2.0066 | 32.062 | 64.3344 | 15.9786 | 14.397 | -0.1102 | 105.1776 | 235.2997 | 181.1158 | 233.1812 | 161.429
2012-10-30 18:42 | 28.5926 | 249.7229 | 195.594 | 247.4633 | 176.0238 | -1.5985 | 1.2709 | 32.137 | 40.842 | 25.2875 | 2.0065 | 32.0614 | 64.3312 | 15.9787 | 14.3795 | -0.1102 | 105.1732 | 235.3434 | 181.2145 | 233.0838 | 161.6443
2012-10-30 18:42 | 28.5956 | 249.7259 | 195.5027 | 247.4937 | 175.7808 | -1.5962 | 1.2709 | 32.1368 | 40.8415 | 25.2874 | 2.0065 | 32.061 | 64.3302 | 15.9786 | 14.3988 | -0.11 | 105.1717 | 235.3271 | 181.1039 | 233.0949 | 161.382
2012-10-30 18:42 | 28.5933 | 249.6844 | 195.4729 | 247.5776 | 175.7707 | -1.6037 | 1.2709 | 32.1366 | 40.8411 | 25.2873 | 2.0065 | 32.0609 | 64.3297 | 15.9786 | 14.4047 | -0.1106 | 105.1708 | 235.2797 | 181.0682 | 233.1729 | 161.366
2012-10-30 18:42 | 28.5971 | 249.7117 | 195.5356 | 247.4639 | 175.9619 | -1.5985 | 1.2708 | 32.1362 | 40.8398 | 25.2875 | 2.0065 | 32.0605 | 64.328 | 15.9787 | 14.4248 | -0.1102 | 105.1678 | 235.2869 | 181.1108 | 233.0391 | 161.5371
2012-10-30 18:42 | 28.6031 | 249.7882 | 195.5534 | 247.4895 | 175.761 | -1.5991 | 1.2708 | 32.1358 | 40.8386 | 25.2876 | 2.0064 | 32.0602 | 64.3266 | 15.9787 | 14.4104 | -0.1102 | 105.1652 | 235.3778 | 181.143 | 233.0791 | 161.3506
2012-10-30 18:42 | 28.5978 | 249.7007 | 195.5246 | 247.6213 | 175.8767 | -1.602 | 1.2708 | 32.1356 | 40.8386 | 25.2873 | 2.0064 | 32.0599 | 64.3255 | 15.9787 | 14.401 | -0.1104 | 105.164 | 235.2997 | 181.1236 | 233.2203 | 161.4757
2012-10-30 18:42 | 28.5956 | 249.7102 | 195.487 | 247.4389 | 175.8433 | -1.5997 | 1.266 | 32.0114 | 40.525 | 25.2863 | 1.9985 | 31.9341 | 63.8215 | 15.9787 | 14.4192 | -0.1103 | 104.3465 | 235.291 | 181.0678 | 233.0197 | 161.4241
2012-10-30 18:42 | 28.5971 | 249.7117 | 195.461 | 247.5383 | 175.8058 | -1.6009 | 1.2647 | 31.9778 | 40.4408 | 25.2858 | 1.9965 | 31.9022 | 63.6939 | 15.9788 | 14.4003 | -0.1104 | 104.1346 | 235.3114 | 181.0607 | 233.138 | 161.4055
2012-10-30 18:42 | 28.5993 | 249.7296 | 195.4553 | 247.5797 | 175.8041 | -1.6061 | 1.264 | 31.9612 | 40.3995 | 25.2854 | 1.9955 | 31.8856 | 63.6279 | 15.9787 | 14.4148 | -0.1107 | 104.0274 | 235.3148 | 181.0405 | 233.1649 | 161.3893
2012-10-30 18:42 | 28.5896 | 249.6729 | 195.4653 | 247.3898 | 176.1262 | -1.6061 | 1.2637 | 31.9542 | 40.3811 | 25.2859 | 1.9951 | 31.879 | 63.6011 | 15.9788 | 14.3398 | -0.1107 | 103.9822 | 235.3331 | 181.1255 | 233.05 | 161.7864
2012-10-30 18:42 | 28.5971 | 249.6843 | 195.4492 | 247.5697 | 175.8838 | -1.6009 | 1.2636 | 31.9501 | 40.3707 | 25.2859 | 1.9948 | 31.8749 | 63.5841 | 15.979 | 14.3636 | -0.1104 | 103.9548 | 235.3207 | 181.0856 | 233.2061 | 161.5202
2012-10-30 18:42 | 28.5993 | 249.6591 | 195.5261 | 247.5562 | 175.8938 | -1.598 | 1.2635 | 31.9484 | 40.3668 | 25.2857 | 1.9947 | 31.8733 | 63.5777 | 15.979 | 14.3862 | -0.1102 | 103.9444 | 235.2729 | 181.1399 | 233.17 | 161.5076
2012-10-30 18:42 | 28.6008 | 249.6332 | 195.5354 | 247.3657 | 175.9227 | -1.6043 | 1.2635 | 31.9487 | 40.3674 | 25.2857 | 1.9947 | 31.8735 | 63.578 | 15.9791 | 14.3796 | -0.1106 | 103.9453 | 235.2536 | 181.1558 | 232.9861 | 161.5431
2012-10-30 18:43 | 28.6016 | 249.634 | 195.4576 | 247.5311 | 175.5409 | -1.598 | 1.2636 | 31.9498 | 40.3705 | 25.2855 | 1.9948 | 31.8746 | 63.5827 | 15.979 | 14.3885 | -0.1102 | 103.9532 | 235.2455 | 181.0691 | 233.1426 | 161.1524
2012-10-30 18:43 | 28.5978 | 249.6263 | 195.4813 | 247.5194 | 175.5918 | -1.6078 | 1.2635 | 31.9485 | 40.367 | 25.2857 | 1.9947 | 31.8733 | 63.5769 | 15.9792 | 14.3847 | -0.1108 | 103.9439 | 235.2416 | 181.0966 | 233.1347 | 161.2071
2012-10-30 18:43 | 28.5903 | 249.5875 | 195.4542 | 247.3513 | 175.3072 | -1.602 | 1.2636 | 31.9503 | 40.3715 | 25.2857 | 1.9948 | 31.8751 | 63.5838 | 15.9793 | 14.4058 | -0.1104 | 103.9553 | 235.1817 | 181.0484 | 232.9455 | 160.9014
2012-10-30 18:43 | 28.5926 | 249.5898 | 195.374 | 247.4202 | 175.5632 | -1.5962 | 1.2653 | 31.9941 | 40.4821 | 25.2858 | 1.9976 | 31.9191 | 63.7601 | 15.9791 | 14.3795 | -0.11 | 104.2422 | 235.2103 | 180.9945 | 233.0407 | 161.1837
2012-10-30 18:43 | 28.5903 | 249.5405 | 195.3638 | 247.5785 | 175.8302 | -1.6009 | 1.2685 | 32.0757 | 40.6873 | 25.2868 | 2.0026 | 32.0001 | 64.084 | 15.9791 | 14.4139 | -0.1104 | 104.7713 | 235.1266 | 180.9499 | 233.1646 | 161.4163
2012-10-30 18:43 | 28.5903 | 249.6188 | 195.4738 | 247.3435 | 175.9473 | -1.6037 | 1.269 | 32.09 | 40.7227 | 25.2872 | 2.0035 | 32.0147 | 64.1428 | 15.979 | 14.4139 | -0.1106 | 104.8655 | 235.2049 | 181.0599 | 232.9296 | 161.5334
2012-10-30 18:43 | 28.5858 | 249.583 | 195.4183 | 247.48 | 175.8881 | -1.6089 | 1.2696 | 32.1039 | 40.7582 | 25.2872 | 2.0044 | 32.0282 | 64.1971 | 15.979 | 14.3686 | -0.1109 | 104.9554 | 235.2144 | 181.0497 | 233.1114 | 161.5195
2012-10-30 18:43 | 28.5828 | 249.576 | 195.5292 | 247.4418 | 176.0686 | -1.5997 | 1.2701 | 32.1161 | 40.7892 | 25.2871 | 2.0052 | 32.0401 | 64.2453 | 15.9789 | 14.3738 | -0.1103 | 105.0345 | 235.2022 | 181.1554 | 233.068 | 161.6948
2012-10-30 18:43 | 28.5858 | 249.6025 | 195.5361 | 247.3625 | 176.056 | -1.5968 | 1.2704 | 32.1247 | 40.8105 | 25.2875 | 2.0057 | 32.0489 | 64.2805 | 15.9789 | 14.3972 | -0.1101 | 105.091 | 235.2053 | 181.1389 | 232.9653 | 161.6588
2012-10-30 18:43 | 28.5858 | 249.583 | 195.5165 | 247.3547 | 175.853 | -1.5991 | 1.2706 | 32.1298 | 40.8236 | 25.2875 | 2.006 | 32.0541 | 64.3011 | 15.9789 | 14.4421 | -0.1102 | 105.1247 | 235.1409 | 181.0744 | 232.9126 | 161.4109
2012-10-30 18:43 | 28.5858 | 249.6025 | 195.5833 | 247.4643 | 176.017 | -1.6014 | 1.2707 | 32.1328 | 40.8308 | 25.2877 | 2.0062 | 32.0572 | 64.3135 | 15.9789 | 14.4054 | -0.1104 | 105.1444 | 235.1971 | 181.1779 | 233.0589 | 161.6116
2012-10-30 18:43 | 28.5716 | 249.6236 | 195.6359 | 247.2464 | 176.1277 | -1.6089 | 1.2708 | 32.135 | 40.8367 | 25.2876 | 2.0064 | 32.0593 | 64.3221 | 15.9789 | 14.3544 | -0.1109 | 105.1588 | 235.2692 | 181.2815 | 232.892 | 161.7733
2012-10-30 18:43 | 28.5866 | 249.6229 | 195.588 | 247.426 | 175.9787 | -1.6003 | 1.2709 | 32.1368 | 40.8414 | 25.2875 | 2.0065 | 32.0609 | 64.3294 | 15.9788 | 14.4143 | -0.1103 | 105.1708 | 235.2086 | 181.1737 | 233.0117 | 161.5644
2012-10-30 18:43 | 28.5798 | 249.5965 | 195.5969 | 247.4544 | 176.1007 | -1.6014 | 1.2709 | 32.1381 | 40.8445 | 25.2876 | 2.0066 | 32.0624 | 64.3352 | 15.9788 | 14.4034 | -0.1104 | 105.1797 | 235.1931 | 181.1935 | 233.051 | 161.6973
2012-10-30 18:43 | 28.5806 | 249.6326 | 195.6292 | 247.3455 | 176.207 | -1.6043 | 1.2709 | 32.1391 | 40.8471 | 25.2875 | 2.0066 | 32.0633 | 64.3392 | 15.9787 | 14.3675 | -0.1106 | 105.1863 | 235.2651 | 181.2617 | 232.978 | 161.8395
2012-10-30 18:43 | 28.5806 | 249.6521 | 195.6567 | 247.2633 | 176.2031 | -1.6003 | 1.271 | 32.1397 | 40.8484 | 25.2877 | 2.0067 | 32.0639 | 64.342 | 15.9786 | 14.392 | -0.1103 | 105.1904 | 235.2601 | 181.2647 | 232.8713 | 161.8111
2012-10-30 18:43 | 28.5888 | 249.6486 | 195.6374 | 247.4125 | 175.9965 | -1.6014 | 1.271 | 32.1403 | 40.8499 | 25.2876 | 2.0067 | 32.0644 | 64.3431 | 15.9788 | 14.4002 | -0.1104 | 105.193 | 235.2484 | 181.2372 | 233.0123 | 161.5963
2012-10-30 18:43 | 28.5888 | 249.6134 | 195.6885 | 247.4399 | 176.0122 | -1.5997 | 1.271 | 32.1408 | 40.8508 | 25.2879 | 2.0067 | 32.065 | 64.3458 | 15.9787 | 14.3635 | -0.1103 | 105.1966 | 235.2499 | 181.325 | 233.0764 | 161.6487
2012-10-30 18:43 | 28.5881 | 249.6401 | 195.6327 | 247.3217 | 175.9997 | -1.6003 | 1.271 | 32.1413 | 40.8521 | 25.2879 | 2.0068 | 32.0655 | 64.348 | 15.9787 | 14.4036 | -0.1103 | 105.2001 | 235.2365 | 181.2291 | 232.9181 | 161.5961
2012-10-30 18:43 | 28.5926 | 249.6602 | 195.6019 | 247.381 | 175.9301 | -1.6043 | 1.271 | 32.1414 | 40.8523 | 25.2878 | 2.0068 | 32.0656 | 64.3487 | 15.9787 | 14.3673 | -0.1106 | 105.201 | 235.2929 | 181.2346 | 233.0137 | 161.5628
2012-10-30 18:43 | 28.5843 | 249.6402 | 195.6289 | 247.4668 | 175.8398 | -1.5968 | 1.271 | 32.1414 | 40.8523 | 25.288 | 2.0068 | 32.0657 | 64.3488 | 15.9786 | 14.3794 | -0.1101 | 105.2011 | 235.2608 | 181.2495 | 233.0874 | 161.4604
2012-10-30 18:43 | 28.5858 | 249.626 | 195.5951 | 247.4056 | 175.9311 | -1.5985 | 1.271 | 32.1415 | 40.8527 | 25.2879 | 2.0068 | 32.0658 | 64.3497 | 15.9785 | 14.4298 | -0.1102 | 105.2025 | 235.1962 | 181.1653 | 232.9758 | 161.5013
2012-10-30 18:43 | 28.5821 | 249.6497 | 195.611 | 247.2883 | 176.025 | -1.5974 | 1.2711 | 32.1415 | 40.8536 | 25.2873 | 2.0068 | 32.0658 | 64.3495 | 15.9786 | 14.3813 | -0.1101 | 105.203 | 235.2684 | 181.2297 | 232.907 | 161.6437
2012-10-30 18:43 | 28.5761 | 249.632 | 195.5697 | 247.3959 | 175.9448 | -1.6009 | 1.271 | 32.1417 | 40.8536 | 25.2876 | 2.0068 | 32.0659 | 64.3497 | 15.9787 | 14.3916 | -0.1104 | 105.2032 | 235.2404 | 181.1781 | 233.0043 | 161.5532
2012-10-30 18:43 | 28.5768 | 249.6288 | 195.6175 | 247.471 | 176.0002 | -1.6014 | 1.271 | 32.1415 | 40.8526 | 25.2878 | 2.0068 | 32.0656 | 64.3493 | 15.9784 | 14.3841 | -0.1104 | 105.2019 | 235.2447 | 181.2334 | 233.0869 | 161.6161
2012-10-30 18:43 | 28.5828 | 249.6739 | 195.651 | 247.3947 | 176.0452 | -1.6026 | 1.271 | 32.1417 | 40.8535 | 25.2877 | 2.0068 | 32.0659 | 64.3504 | 15.9785 | 14.3942 | -0.1105 | 105.2038 | 235.2797 | 181.2568 | 233.0005 | 161.651
2012-10-30 18:43 | 28.5911 | 249.6626 | 195.6279 | 247.3795 | 175.8856 | -1.5968 | 1.271 | 32.1416 | 40.8533 | 25.2877 | 2.0068 | 32.0659 | 64.3502 | 15.9785 | 14.4433 | -0.1101 | 105.2035 | 235.2193 | 181.1846 | 232.9362 | 161.4423
2012-10-30 18:44 | 28.5806 | 249.6404 | 195.5977 | 247.4709 | 175.8283 | -1.5991 | 1.2711 | 32.1419 | 40.854 | 25.2876 | 2.0068 | 32.0662 | 64.3519 | 15.9784 | 14.4002 | -0.1102 | 105.2059 | 235.2402 | 181.1975 | 233.0707 | 161.4281
2012-10-30 18:44 | 28.5858 | 249.6847 | 195.5794 | 247.4604 | 175.8608 | -1.5974 | 1.271 | 32.1415 | 40.8527 | 25.2879 | 2.0068 | 32.0657 | 64.3499 | 15.9784 | 14.3972 | -0.1101 | 105.2026 | 235.2875 | 181.1822 | 233.0632 | 161.4636
2012-10-30 18:44 | 28.5903 | 249.6736 | 195.6271 | 247.2769 | 176.0176 | -1.6026 | 1.271 | 32.1413 | 40.8519 | 25.2881 | 2.0068 | 32.0655 | 64.3489 | 15.9784 | 14.3609 | -0.1105 | 105.2008 | 235.3127 | 181.2662 | 232.916 | 161.6567
2012-10-30 18:44 | 28.5963 | 249.6952 | 195.5506 | 247.4278 | 175.8206 | -1.6009 | 1.271 | 32.1409 | 40.8517 | 25.2876 | 2.0068 | 32.0652 | 64.3481 | 15.9783 | 14.3628 | -0.1104 | 105.1998 | 235.3324 | 181.1878 | 233.065 | 161.4578
2012-10-30 18:44 | 28.6001 | 249.6873 | 195.5662 | 247.51 | 175.719 | -1.5985 | 1.271 | 32.1409 | 40.8519 | 25.2874 | 2.0068 | 32.0653 | 64.348 | 15.9785 | 14.3952 | -0.1102 | 105.1998 | 235.2921 | 181.171 | 233.1148 | 161.3238
2012-10-30 18:44 | 28.5956 | 249.6358 | 195.5302 | 247.4193 | 175.7886 | -1.6014 | 1.271 | 32.1409 | 40.8512 | 25.2878 | 2.0068 | 32.0652 | 64.3479 | 15.9784 | 14.3417 | -0.1104 | 105.1991 | 235.2941 | 181.1885 | 233.0776 | 161.4469
2012-10-30 18:44 | 28.5993 | 249.6865 | 195.5025 | 247.423 | 175.8704 | -1.6014 | 1.271 | 32.1407 | 40.8508 | 25.2878 | 2.0068 | 32.0649 | 64.347 | 15.9783 | 14.4066 | -0.1104 | 105.1978 | 235.2799 | 181.0959 | 233.0164 | 161.4638
2012-10-30 18:44 | 28.6008 | 249.6802 | 195.5079 | 247.3775 | 175.8095 | -1.5974 | 1.271 | 32.1411 | 40.8517 | 25.2879 | 2.0068 | 32.0653 | 64.3495 | 15.9781 | 14.3755 | -0.1101 | 105.2013 | 235.3047 | 181.1324 | 233.002 | 161.434
2012-10-30 18:44 | 28.6001 | 249.6638 | 195.4836 | 247.4669 | 175.7853 | -1.602 | 1.271 | 32.141 | 40.8517 | 25.2877 | 2.0068 | 32.0653 | 64.3487 | 15.9783 | 14.4033 | -0.1104 | 105.2004 | 235.2605 | 181.0803 | 233.0636 | 161.382
2012-10-30 18:44 | 28.6113 | 249.628 | 195.5263 | 247.5016 | 175.9527 | -1.5997 | 1.271 | 32.1411 | 40.852 | 25.2876 | 2.0068 | 32.0653 | 64.3493 | 15.9782 | 14.4309 | -0.1103 | 105.2013 | 235.1971 | 181.0954 | 233.0707 | 161.5218
2012-10-30 18:44 | 28.609 | 249.6884 | 195.5358 | 247.4327 | 175.8723 | -1.6049 | 1.271 | 32.1407 | 40.8511 | 25.2876 | 2.0068 | 32.0649 | 64.3472 | 15.9783 | 14.3837 | -0.1106 | 105.1983 | 235.3047 | 181.1521 | 233.049 | 161.4886
2012-10-30 18:44 | 28.6038 | 249.6832 | 195.4991 | 247.404 | 175.6641 | -1.602 | 1.271 | 32.1404 | 40.8506 | 25.2874 | 2.0068 | 32.0645 | 64.3457 | 15.9783 | 14.403 | -0.1104 | 105.1963 | 235.2802 | 181.0961 | 233.001 | 161.2611
2012-10-30 18:44 | 28.6113 | 249.6868 | 195.5145 | 247.5016 | 175.6248 | -1.6014 | 1.271 | 32.1402 | 40.85 | 25.2875 | 2.0068 | 32.0645 | 64.3463 | 15.9781 | 14.4023 | -0.1104 | 105.1963 | 235.2845 | 181.1122 | 233.0993 | 161.2225
2012-10-30 18:44 | 28.609 | 249.6923 | 195.4611 | 247.5502 | 175.6108 | -1.6003 | 1.271 | 32.1399 | 40.8488 | 25.2877 | 2.0068 | 32.0641 | 64.3448 | 15.9781 | 14.4041 | -0.1103 | 105.1936 | 235.2882 | 181.057 | 233.1461 | 161.2067
2012-10-30 18:44 | 28.6105 | 249.6781 | 195.4665 | 247.4577 | 175.7372 | -1.602 | 1.271 | 32.1394 | 40.8479 | 25.2875 | 2.0067 | 32.0637 | 64.3429 | 15.9781 | 14.4382 | -0.1104 | 105.1908 | 235.2399 | 181.0283 | 233.0195 | 161.299
2012-10-30 18:44 | 28.6113 | 249.6985 | 195.4477 | 247.4428 | 175.5702 | -1.6003 | 1.2709 | 32.1388 | 40.8465 | 25.2874 | 2.0067 | 32.0631 | 64.3408 | 15.9781 | 14.4268 | -0.1103 | 105.1873 | 235.2717 | 181.0209 | 233.016 | 161.1434
2012-10-30 18:44 | 28.6068 | 249.6509 | 195.4314 | 247.5088 | 175.4876 | -1.6043 | 1.2709 | 32.1389 | 40.8465 | 25.2875 | 2.0067 | 32.0633 | 64.3411 | 15.9782 | 14.4019 | -0.1106 | 105.1876 | 235.249 | 181.0295 | 233.1069 | 161.0857
2012-10-30 18:44 | 28.606 | 249.6775 | 195.3834 | 247.4924 | 175.4946 | -1.6043 | 1.2709 | 32.1381 | 40.8448 | 25.2873 | 2.0066 | 32.0624 | 64.3377 | 15.9781 | 14.4052 | -0.1106 | 105.1825 | 235.2723 | 180.9782 | 233.0872 | 161.0894
2012-10-30 18:44 | 28.6068 | 249.6901 | 195.3685 | 247.454 | 175.6788 | -1.6072 | 1.2709 | 32.1382 | 40.8448 | 25.2876 | 2.0067 | 32.0626 | 64.3387 | 15.9781 | 14.3733 | -0.1108 | 105.1835 | 235.3168 | 180.9952 | 233.0807 | 161.3055
2012-10-30 18:44 | 28.6053 | 249.7003 | 195.3867 | 247.4447 | 175.6266 | -1.5997 | 1.2709 | 32.1378 | 40.8434 | 25.2877 | 2.0066 | 32.0621 | 64.3371 | 15.978 | 14.38 | -0.1103 | 105.1805 | 235.3203 | 181.0067 | 233.0647 | 161.2466
2012-10-30 18:44 | 28.6143 | 249.7211 | 195.3957 | 247.4341 | 175.5224 | -1.6026 | 1.2709 | 32.1372 | 40.8422 | 25.2876 | 2.0066 | 32.0616 | 64.3349 | 15.978 | 14.3686 | -0.1105 | 105.1771 | 235.3525 | 181.0271 | 233.0655 | 161.1538
2012-10-30 18:44 | 28.6075 | 249.7025 | 195.3732 | 247.5722 | 175.5 | -1.6014 | 1.2709 | 32.1374 | 40.8424 | 25.2877 | 2.0066 | 32.0617 | 64.3358 | 15.9779 | 14.3781 | -0.1104 | 105.1782 | 235.3244 | 180.9951 | 233.1941 | 161.1219
2012-10-30 18:44 | 28.612 | 249.6757 | 195.3423 | 247.5101 | 175.6684 | -1.6014 | 1.2709 | 32.1372 | 40.8421 | 25.2876 | 2.0066 | 32.0614 | 64.3352 | 15.9778 | 14.3908 | -0.1104 | 105.1772 | 235.2849 | 180.9515 | 233.1193 | 161.2776
2012-10-30 18:44 | 28.6128 | 249.7078 | 195.3392 | 247.4169 | 175.6926 | -1.5997 | 1.2708 | 32.1368 | 40.8409 | 25.2877 | 2.0066 | 32.0611 | 64.3335 | 15.9779 | 14.3997 | -0.1103 | 105.1744 | 235.3081 | 180.9395 | 233.0172 | 161.2929
2012-10-30 18:44 | 28.6098 | 249.6657 | 195.3479 | 247.4883 | 175.7443 | -1.6026 | 1.2708 | 32.1364 | 40.8404 | 25.2874 | 2.0066 | 32.0607 | 64.3316 | 15.978 | 14.3845 | -0.1105 | 105.1721 | 235.2812 | 180.9634 | 233.1038 | 161.3598
2012-10-30 18:45 | 28.6105 | 249.6977 | 195.3801 | 247.5165 | 175.6943 | -1.6084 | 1.2708 | 32.136 | 40.8391 | 25.2875 | 2.0065 | 32.0602 | 64.3298 | 15.978 | 14.4219 | -0.1109 | 105.1689 | 235.2758 | 180.9582 | 233.0946 | 161.2724
2012-10-30 18:45 | 28.6165 | 249.6802 | 195.3704 | 247.5577 | 175.7783 | -1.6014 | 1.2708 | 32.1355 | 40.8377 | 25.2876 | 2.0065 | 32.0596 | 64.3278 | 15.9778 | 14.4116 | -0.1104 | 105.1656 | 235.2686 | 180.9588 | 233.1461 | 161.3667
2012-10-30 18:45 | 28.6128 | 249.6883 | 195.4217 | 247.5109 | 175.8058 | -1.5974 | 1.2708 | 32.1349 | 40.8367 | 25.2874 | 2.0065 | 32.0592 | 64.3258 | 15.9779 | 14.3956 | -0.1101 | 105.1625 | 235.2927 | 181.0261 | 233.1153 | 161.4102
2012-10-30 18:45 | 28.6233 | 249.6988 | 195.4243 | 247.5136 | 175.8046 | -1.6037 | 1.2708 | 32.1348 | 40.8359 | 25.2876 | 2.0065 | 32.0589 | 64.3249 | 15.9779 | 14.4061 | -0.1106 | 105.1608 | 235.2927 | 181.0182 | 233.1075 | 161.3985
2012-10-30 18:45 | 28.6158 | 249.6991 | 195.4207 | 247.4669 | 175.8127 | -1.5997 | 1.2707 | 32.1344 | 40.8348 | 25.2878 | 2.0064 | 32.0587 | 64.3235 | 15.978 | 14.3905 | -0.1103 | 105.1583 | 235.3086 | 181.0302 | 233.0764 | 161.4222
2012-10-30 18:45 | 28.6188 | 249.7138 | 195.4198 | 247.4856 | 175.6596 | -1.6049 | 1.2707 | 32.1338 | 40.8336 | 25.2875 | 2.0064 | 32.058 | 64.3209 | 15.9779 | 14.4098 | -0.1106 | 105.1545 | 235.304 | 181.01 | 233.0758 | 161.2498
2012-10-30 18:45 | 28.6203 | 249.7114 | 195.3977 | 247.4636 | 175.6728 | -1.6014 | 1.268 | 32.0654 | 40.66 | 25.2875 | 2.0021 | 31.9899 | 64.0486 | 15.9778 | 14.395 | -0.1104 | 104.7086 | 235.3164 | 181.0027 | 233.0686 | 161.2778
2012-10-30 18:45 | 28.6158 | 249.7069 | 195.3657 | 247.5335 | 175.7698 | -1.602 | 1.2647 | 31.9802 | 40.4463 | 25.2862 | 1.9968 | 31.9048 | 63.7072 | 15.978 | 14.4068 | -0.1104 | 104.1536 | 235.3001 | 180.9589 | 233.1267 | 161.363
2012-10-30 18:45 | 28.612 | 249.6366 | 195.3501 | 247.4905 | 175.7816 | -1.6003 | 1.2642 | 31.9672 | 40.4141 | 25.2858 | 1.996 | 31.8923 | 63.6569 | 15.9781 | 14.4071 | -0.1103 | 104.0709 | 235.2295 | 180.943 | 233.0834 | 161.3745
2012-10-30 18:45 | 28.6263 | 249.6235 | 195.3802 | 247.4187 | 175.4798 | -1.6049 | 1.2637 | 31.954 | 40.3805 | 25.2859 | 1.9952 | 31.8787 | 63.6031 | 15.978 | 14.4091 | -0.1106 | 103.9836 | 235.2144 | 180.9711 | 233.0096 | 161.0707
2012-10-30 18:45 | 28.624 | 249.6368 | 195.3464 | 247.4712 | 175.5594 | -1.6037 | 1.2635 | 31.9481 | 40.3658 | 25.2858 | 1.9948 | 31.8728 | 63.5792 | 15.9781 | 14.3742 | -0.1106 | 103.9449 | 235.2626 | 180.9722 | 233.097 | 161.1852
2012-10-30 18:45 | 28.6195 | 249.648 | 195.338 | 247.5568 | 175.2427 | -1.6049 | 1.2635 | 31.947 | 40.3636 | 25.2855 | 1.9947 | 31.8717 | 63.5747 | 15.9781 | 14.3819 | -0.1106 | 103.9384 | 235.2661 | 180.9561 | 233.1749 | 160.8608
2012-10-30 18:45 | 28.6203 | 249.5666 | 195.2602 | 247.4479 | 175.0328 | -1.6032 | 1.2634 | 31.9461 | 40.3606 | 25.2859 | 1.9946 | 31.8708 | 63.5702 | 15.9783 | 14.3909 | -0.1105 | 103.9308 | 235.1757 | 180.8693 | 233.057 | 160.6419
2012-10-30 18:45 | 28.6203 | 249.5901 | 195.2288 | 247.4949 | 175.2006 | -1.5991 | 1.2634 | 31.9454 | 40.3585 | 25.286 | 1.9946 | 31.8701 | 63.5675 | 15.9784 | 14.4113 | -0.1102 | 103.926 | 235.1788 | 180.8175 | 233.0836 | 160.7893
2012-10-30 18:45 | 28.618 | 249.5643 | 195.2108 | 247.3869 | 175.4051 | -1.5991 | 1.2635 | 31.95 | 40.3699 | 25.2862 | 1.9949 | 31.8747 | 63.586 | 15.9783 | 14.4335 | -0.1102 | 103.9559 | 235.1308 | 180.7773 | 232.9534 | 160.9716
2012-10-30 18:45 | 28.6173 | 249.5323 | 195.2061 | 247.4097 | 175.5449 | -1.5997 | 1.2636 | 31.952 | 40.376 | 25.2855 | 1.995 | 31.8767 | 63.5933 | 15.9785 | 14.4409 | -0.1103 | 103.9693 | 235.0914 | 180.7652 | 232.9688 | 161.104
2012-10-30 18:45 | 28.615 | 249.4947 | 195.2667 | 247.4387 | 175.7144 | -1.6014 | 1.2675 | 32.052 | 40.6271 | 25.2868 | 2.0013 | 31.9777 | 63.9975 | 15.9783 | 14.4101 | -0.1104 | 104.6247 | 235.0846 | 180.8566 | 233.0286 | 161.3043
2012-10-30 18:45 | 28.609 | 249.5122 | 195.3 | 247.4288 | 175.7396 | -1.5991 | 1.269 | 32.0899 | 40.7224 | 25.2874 | 2.0036 | 32.0145 | 64.1451 | 15.9783 | 14.4367 | -0.1102 | 104.8674 | 235.0755 | 180.8633 | 232.9921 | 161.3029
2012-10-30 18:45 | 28.6031 | 249.5259 | 195.3648 | 247.3798 | 175.9484 | -1.602 | 1.2696 | 32.105 | 40.7605 | 25.2874 | 2.0046 | 32.0293 | 64.2043 | 15.9783 | 14.39 | -0.1104 | 104.9648 | 235.1359 | 180.9748 | 232.9898 | 161.5584
2012-10-30 18:45 | 28.6008 | 249.5001 | 195.3822 | 247.2913 | 175.8836 | -1.6003 | 1.2701 | 32.1168 | 40.7903 | 25.2875 | 2.0053 | 32.041 | 64.2512 | 15.9783 | 14.4122 | -0.1103 | 105.0415 | 235.0879 | 180.97 | 232.8791 | 161.4714
2012-10-30 18:45 | 28.6038 | 249.5462 | 195.4048 | 247.4001 | 175.7773 | -1.5991 | 1.2704 | 32.1249 | 40.8108 | 25.2877 | 2.0058 | 32.0492 | 64.2852 | 15.9781 | 14.407 | -0.1102 | 105.096 | 235.1392 | 180.9978 | 232.9931 | 161.3703
2012-10-30 18:45 | 28.5986 | 249.5136 | 195.4271 | 247.3714 | 175.7838 | -1.5974 | 1.2706 | 32.1309 | 40.8263 | 25.2875 | 2.0062 | 32.0552 | 64.3091 | 15.9781 | 14.3814 | -0.1101 | 105.1353 | 235.1322 | 181.0457 | 232.99 | 161.4024
2012-10-30 18:45 | 28.6098 | 249.5796 | 195.4815 | 247.4022 | 175.6233 | -1.5985 | 1.2707 | 32.1345 | 40.8348 | 25.2879 | 2.0064 | 32.0587 | 64.323 | 15.9781 | 14.4294 | -0.1102 | 105.1578 | 235.1502 | 181.0521 | 232.9728 | 161.1939
2012-10-30 18:46 | 28.5956 | 249.5419 | 195.4123 | 247.3018 | 175.7106 | -1.598 | 1.2709 | 32.1366 | 40.8409 | 25.2875 | 2.0065 | 32.0609 | 64.3316 | 15.9781 | 14.4192 | -0.1102 | 105.1726 | 235.1227 | 180.9931 | 232.8826 | 161.2914
2012-10-30 18:46 | 28.5903 | 249.517 | 195.3835 | 247.3356 | 175.6467 | -1.6037 | 1.2709 | 32.1381 | 40.8438 | 25.288 | 2.0066 | 32.0622 | 64.3373 | 15.9781 | 14.4221 | -0.1106 | 105.1812 | 235.0949 | 180.9614 | 232.9135 | 161.2246
2012-10-30 18:46 | 28.5903 | 249.5209 | 195.4345 | 247.4297 | 175.6819 | -1.6078 | 1.271 | 32.1396 | 40.8481 | 25.2877 | 2.0067 | 32.0638 | 64.344 | 15.978 | 14.4466 | -0.1108 | 105.1921 | 235.0743 | 180.9879 | 232.9831 | 161.2353
2012-10-30 18:46 | 28.5843 | 249.5345 | 195.4207 | 247.3532 | 175.629 | -1.5997 | 1.271 | 32.1405 | 40.8501 | 25.2879 | 2.0068 | 32.0647 | 64.3478 | 15.9779 | 14.4365 | -0.1103 | 105.1979 | 235.098 | 180.9842 | 232.9167 | 161.1925
2012-10-30 18:46 | 28.5896 | 249.5711 | 195.3828 | 247.3271 | 175.7475 | -1.5997 | 1.271 | 32.1413 | 40.8525 | 25.2877 | 2.0069 | 32.0655 | 64.3509 | 15.978 | 14.4132 | -0.1103 | 105.2034 | 235.1579 | 180.9696 | 232.9139 | 161.3343
2012-10-30 18:46 | 28.5896 | 249.4928 | 195.367 | 247.3232 | 175.7046 | -1.598 | 1.271 | 32.1419 | 40.8538 | 25.2878 | 2.0069 | 32.0661 | 64.3534 | 15.9779 | 14.4459 | -0.1102 | 105.2072 | 235.0469 | 180.9211 | 232.8773 | 161.2587
2012-10-30 18:46 | 28.5866 | 249.525 | 195.3798 | 247.422 | 175.6196 | -1.5962 | 1.2711 | 32.1423 | 40.8546 | 25.288 | 2.0069 | 32.0665 | 64.3552 | 15.9779 | 14.4592 | -0.11 | 105.2098 | 235.0658 | 180.9206 | 232.9628 | 161.1604
2012-10-30 18:46 | 28.5918 | 249.542 | 195.3692 | 247.3489 | 175.4804 | -1.6009 | 1.2711 | 32.1425 | 40.8551 | 25.288 | 2.007 | 32.0667 | 64.3563 | 15.9778 | 14.4154 | -0.1104 | 105.2114 | 235.1266 | 180.9538 | 232.9335 | 161.065
2012-10-30 18:46 | 28.5836 | 249.5338 | 195.3532 | 247.2702 | 175.6322 | -1.5991 | 1.2711 | 32.1424 | 40.855 | 25.2878 | 2.0069 | 32.0665 | 64.3552 | 15.9779 | 14.3828 | -0.1102 | 105.2102 | 235.151 | 180.9704 | 232.8874 | 161.2494
2012-10-30 18:46 | 28.5836 | 249.5534 | 195.2746 | 247.376 | 175.6478 | -1.5951 | 1.2711 | 32.1425 | 40.8553 | 25.2878 | 2.007 | 32.0666 | 64.3563 | 15.9777 | 14.4317 | -0.11 | 105.2116 | 235.1217 | 180.8429 | 232.9443 | 161.2161
2012-10-30 18:46 | 28.5881 | 249.5226 | 195.3341 | 247.4079 | 175.7187 | -1.6014 | 1.2711 | 32.1427 | 40.8554 | 25.288 | 2.007 | 32.0669 | 64.3575 | 15.9777 | 14.3913 | -0.1104 | 105.213 | 235.1313 | 180.9428 | 233.0166 | 161.3274
2012-10-30 18:46 | 28.5858 | 249.5164 | 195.32 | 247.339 | 175.7554 | -1.5985 | 1.2711 | 32.1429 | 40.8565 | 25.2877 | 2.007 | 32.0671 | 64.3584 | 15.9777 | 14.3809 | -0.1102 | 105.2149 | 235.1355 | 180.9391 | 232.9581 | 161.3745
2012-10-30 18:46 | 28.5911 | 249.5413 | 195.3607 | 247.3991 | 175.6553 | -1.6049 | 1.2711 | 32.1431 | 40.8565 | 25.2879 | 2.007 | 32.0672 | 64.3591 | 15.9777 | 14.4188 | -0.1106 | 105.2156 | 235.1225 | 180.9419 | 232.9803 | 161.2365
2012-10-30 18:46 | 28.5911 | 249.5413 | 195.3489 | 247.356 | 175.6397 | -1.6014 | 1.2711 | 32.1423 | 40.8547 | 25.2878 | 2.007 | 32.0665 | 64.3562 | 15.9776 | 14.4147 | -0.1104 | 105.211 | 235.1266 | 180.9342 | 232.9413 | 161.225
2012-10-30 18:46 | 28.5933 | 249.5278 | 195.3079 | 247.3739 | 175.6458 | -1.5962 | 1.2711 | 32.1422 | 40.855 | 25.2875 | 2.007 | 32.0663 | 64.3557 | 15.9776 | 14.4047 | -0.11 | 105.2108 | 235.1231 | 180.9032 | 232.9692 | 161.2411
2012-10-30 18:46 | 28.5926 | 249.5545 | 195.3268 | 247.2988 | 175.6334 | -1.5974 | 1.271 | 32.1416 | 40.8529 | 25.2879 | 2.0069 | 32.0659 | 64.354 | 15.9776 | 14.3795 | -0.1101 | 105.2069 | 235.175 | 180.9473 | 232.9193 | 161.2539
2012-10-30 18:46 | 28.5903 | 249.5875 | 195.2852 | 247.4218 | 175.4828 | -1.5985 | 1.271 | 32.1415 | 40.8529 | 25.2877 | 2.0069 | 32.0657 | 64.3532 | 15.9776 | 14.3527 | -0.1102 | 105.2061 | 235.2348 | 180.9325 | 233.0691 | 161.1301
2012-10-30 18:46 | 28.5993 | 249.5495 | 195.3099 | 247.4073 | 175.6245 | -1.6066 | 1.271 | 32.1411 | 40.852 | 25.2876 | 2.0069 | 32.0653 | 64.3525 | 15.9773 | 14.4148 | -0.1108 | 105.2046 | 235.1347 | 180.8951 | 232.9925 | 161.2097
2012-10-30 18:46 | 28.5903 | 249.5561 | 195.3245 | 247.3787 | 175.6311 | -1.6009 | 1.271 | 32.1409 | 40.8516 | 25.2876 | 2.0069 | 32.0651 | 64.3516 | 15.9774 | 14.3854 | -0.1104 | 105.2032 | 235.1707 | 180.9391 | 232.9933 | 161.2457
2012-10-30 18:47 | 28.5903 | 249.5757 | 195.2892 | 247.3278 | 175.6311 | -1.5985 | 1.271 | 32.1406 | 40.8505 | 25.2877 | 2.0069 | 32.0648 | 64.3499 | 15.9775 | 14.4139 | -0.1102 | 105.2004 | 235.1618 | 180.8753 | 232.9139 | 161.2172
2012-10-30 18:47 | 28.6001 | 249.5659 | 195.3186 | 247.3572 | 175.6878 | -1.5991 | 1.271 | 32.1403 | 40.8501 | 25.2875 | 2.0069 | 32.0645 | 64.349 | 15.9775 | 14.3911 | -0.1102 | 105.1991 | 235.1748 | 180.9275 | 232.9661 | 161.2967
2012-10-30 18:47 | 28.5963 | 249.5387 | 195.3069 | 247.4553 | 175.5239 | -1.5997 | 1.271 | 32.1404 | 40.8502 | 25.2876 | 2.0069 | 32.0645 | 64.3491 | 15.9774 | 14.4199 | -0.1103 | 105.1994 | 235.1188 | 180.887 | 233.0354 | 161.104
2012-10-30 18:47 | 28.6053 | 249.5985 | 195.2806 | 247.4016 | 175.6539 | -1.5991 | 1.271 | 32.1397 | 40.8486 | 25.2875 | 2.0068 | 32.0639 | 64.3466 | 15.9774 | 14.4289 | -0.1102 | 105.1952 | 235.1696 | 180.8517 | 232.9727 | 161.225
2012-10-30 18:47 | 28.609 | 249.5709 | 195.3432 | 247.3504 | 175.802 | -1.6032 | 1.271 | 32.1392 | 40.8473 | 25.2875 | 2.0068 | 32.0634 | 64.3447 | 15.9774 | 14.4204 | -0.1105 | 105.192 | 235.1505 | 180.9228 | 232.93 | 161.3816
2012-10-30 18:47 | 28.6053 | 249.5711 | 195.3159 | 247.3467 | 175.6266 | -1.6014 | 1.2709 | 32.1388 | 40.8458 | 25.2878 | 2.0068 | 32.0629 | 64.3429 | 15.9774 | 14.4045 | -0.1104 | 105.1887 | 235.1666 | 180.9114 | 232.9422 | 161.2221
2012-10-30 18:47 | 28.6008 | 249.5588 | 195.28 | 247.4167 | 175.5401 | -1.6009 | 1.2709 | 32.1387 | 40.8457 | 25.2877 | 2.0068 | 32.0628 | 64.3429 | 15.9773 | 14.3959 | -0.1104 | 105.1885 | 235.1629 | 180.8841 | 233.0208 | 161.1442
2012-10-30 18:47 | 28.6135 | 249.5989 | 195.3084 | 247.3863 | 175.7324 | -1.6032 | 1.2709 | 32.1382 | 40.8443 | 25.2878 | 2.0067 | 32.0625 | 64.341 | 15.9774 | 14.4004 | -0.1105 | 105.1854 | 235.1985 | 180.908 | 232.9859 | 161.332
2012-10-30 18:47 | 28.6098 | 249.5678 | 195.3362 | 247.41 | 175.6818 | -1.6037 | 1.2709 | 32.1376 | 40.8429 | 25.2878 | 2.0067 | 32.0619 | 64.3385 | 15.9774 | 14.4171 | -0.1106 | 105.1814 | 235.1507 | 180.9191 | 232.9929 | 161.2647
2012-10-30 18:47 | 28.609 | 249.567 | 195.3589 | 247.37 | 175.5288 | -1.6003 | 1.2709 | 32.1374 | 40.8424 | 25.2878 | 2.0067 | 32.0616 | 64.3374 | 15.9774 | 14.4367 | -0.1103 | 105.1797 | 235.1303 | 180.9222 | 232.9333 | 161.0921
2012-10-30 18:47 | 28.6068 | 249.6275 | 195.3135 | 247.313 | 175.5657 | -1.6003 | 1.2709 | 32.1373 | 40.8422 | 25.2877 | 2.0067 | 32.0616 | 64.338 | 15.9772 | 14.3978 | -0.1103 | 105.1802 | 235.2297 | 180.9157 | 232.9152 | 161.1679
2012-10-30 18:47 | 28.6098 | 249.603 | 195.3087 | 247.3904 | 175.4555 | -1.6009 | 1.2708 | 32.1369 | 40.8406 | 25.288 | 2.0067 | 32.061 | 64.3356 | 15.9773 | 14.4375 | -0.1104 | 105.1763 | 235.1655 | 180.8712 | 232.9529 | 161.018
2012-10-30 18:47 | 28.6023 | 249.5916 | 195.2579 | 247.3986 | 175.3348 | -1.6037 | 1.2708 | 32.1364 | 40.8395 | 25.2879 | 2.0066 | 32.0605 | 64.3332 | 15.9774 | 14.4096 | -0.1106 | 105.1727 | 235.182 | 180.8483 | 232.989 | 160.9252
2012-10-30 18:47 | 28.606 | 249.5444 | 195.2616 | 247.465 | 175.3814 | -1.6003 | 1.2708 | 32.1361 | 40.8386 | 25.2881 | 2.0066 | 32.0602 | 64.3324 | 15.9773 | 14.4419 | -0.1103 | 105.171 | 235.1025 | 180.8197 | 233.0231 | 160.9395
2012-10-30 18:47 | 28.612 | 249.5778 | 195.2362 | 247.42 | 175.4889 | -1.6049 | 1.2708 | 32.1355 | 40.8378 | 25.2876 | 2.0066 | 32.0597 | 64.3306 | 15.9772 | 14.4479 | -0.1106 | 105.1684 | 235.1299 | 180.7883 | 232.9721 | 161.041
2012-10-30 18:47 | 28.6031 | 249.6042 | 195.2587 | 247.4072 | 175.5073 | -1.5985 | 1.2708 | 32.1353 | 40.8373 | 25.2876 | 2.0066 | 32.0595 | 64.3296 | 15.9773 | 14.4308 | -0.1102 | 105.1669 | 235.1734 | 180.8279 | 232.9764 | 161.0765
2012-10-30 18:47 | 28.6031 | 249.612 | 195.2273 | 247.3132 | 175.6556 | -1.5968 | 1.2708 | 32.135 | 40.8363 | 25.2877 | 2.0066 | 32.0592 | 64.329 | 15.9772 | 14.4063 | -0.1101 | 105.1653 | 235.2057 | 180.821 | 232.9069 | 161.2493
2012-10-30 18:47 | 28.606 | 249.6267 | 195.2813 | 247.4845 | 175.639 | -1.6043 | 1.2708 | 32.1348 | 40.8361 | 25.2876 | 2.0066 | 32.059 | 64.3282 | 15.9772 | 14.4215 | -0.1106 | 105.1643 | 235.2052 | 180.8598 | 233.063 | 161.2175
2012-10-30 18:47 | 28.6075 | 249.6203 | 195.2867 | 247.4586 | 175.6288 | -1.6032 | 1.2708 | 32.1344 | 40.8351 | 25.2876 | 2.0065 | 32.0587 | 64.3274 | 15.9771 | 14.4067 | -0.1105 | 105.1624 | 235.2136 | 180.88 | 233.0519 | 161.2221
2012-10-30 18:47 | 28.6038 | 249.6479 | 195.283 | 247.3687 | 175.5939 | -1.6014 | 1.2707 | 32.134 | 40.8339 | 25.2876 | 2.0065 | 32.0582 | 64.3252 | 15.9771 | 14.3703 | -0.1104 | 105.1591 | 235.2776 | 180.9127 | 232.9984 | 161.2236
2012-10-30 18:47 | 28.6023 | 249.5525 | 195.254 | 247.4299 | 175.5416 | -1.6032 | 1.2678 | 32.0596 | 40.6452 | 25.2875 | 2.0016 | 31.9805 | 64.013 | 15.9773 | 14.3933 | -0.1105 | 104.6582 | 235.1592 | 180.8607 | 233.0366 | 161.1483
2012-10-30 18:47 | 28.6038 | 249.601 | 195.2398 | 247.3139 | 175.6251 | -1.5974 | 1.2646 | 31.9764 | 40.4372 | 25.2859 | 1.9967 | 31.901 | 63.6954 | 15.9772 | 14.4478 | -0.1101 | 104.1326 | 235.1532 | 180.792 | 232.8661 | 161.1773
2012-10-30 18:47 | 28.5978 | 249.5676 | 195.2613 | 247.4137 | 175.623 | -1.6003 | 1.2641 | 31.9642 | 40.4062 | 25.286 | 1.9959 | 31.8885 | 63.645 | 15.9773 | 14.4214 | -0.1103 | 104.0512 | 235.1462 | 180.8399 | 232.9923 | 161.2016
2012-10-30 18:48 | 28.6008 | 249.5862 | 195.2761 | 247.4754 | 175.7743 | -1.5997 | 1.2636 | 31.9508 | 40.3725 | 25.2859 | 1.995 | 31.8755 | 63.5932 | 15.9773 | 14.4408 | -0.1103 | 103.9657 | 235.1454 | 180.8353 | 233.0346 | 161.3335
2012-10-30 18:48 | 28.6008 | 249.5236 | 195.2643 | 247.3932 | 175.4465 | -1.6032 | 1.2634 | 31.9452 | 40.3585 | 25.2858 | 1.9947 | 31.8699 | 63.5707 | 15.9774 | 14.4326 | -0.1105 | 103.9292 | 235.091 | 180.8317 | 232.9606 | 161.0139
2012-10-30 18:48 | 28.5911 | 249.4943 | 195.2585 | 247.4109 | 175.3392 | -1.6055 | 1.2633 | 31.9442 | 40.3559 | 25.2858 | 1.9946 | 31.8691 | 63.5666 | 15.9775 | 14.3943 | -0.1107 | 103.9225 | 235.1 | 180.8642 | 233.0166 | 160.9449
2012-10-30 18:48 | 28.5933 | 249.5122 | 195.2411 | 247.3151 | 175.396 | -1.6037 | 1.2633 | 31.9444 | 40.3567 | 25.2856 | 1.9946 | 31.869 | 63.5665 | 15.9775 | 14.4414 | -0.1106 | 103.9231 | 235.0708 | 180.7997 | 232.8737 | 160.9546
2012-10-30 18:48 | 28.5948 | 249.4784 | 195.1876 | 247.4381 | 175.218 | -1.5991 | 1.2634 | 31.9469 | 40.3628 | 25.2858 | 1.9948 | 31.8716 | 63.5769 | 15.9775 | 14.4266 | -0.1102 | 103.9397 | 235.0518 | 180.761 | 233.0115 | 160.7914
2012-10-30 18:48 | 28.5791 | 249.451 | 195.164 | 247.348 | 175.2726 | -1.5997 | 1.2635 | 31.9486 | 40.367 | 25.2858 | 1.9949 | 31.8733 | 63.5835 | 15.9775 | 14.3864 | -0.1103 | 103.9504 | 235.0646 | 180.7776 | 232.9616 | 160.8862
2012-10-30 18:48 | 28.5806 | 249.4368 | 195.1694 | 247.369 | 175.6136 | -1.6009 | 1.2645 | 31.9744 | 40.4314 | 25.2863 | 1.9966 | 31.9008 | 63.6932 | 15.9776 | 14.4532 | -0.1104 | 104.1246 | 234.9836 | 180.7162 | 232.9158 | 161.1604
2012-10-30 18:48 | 28.5956 | 249.4675 | 195.1962 | 247.3057 | 175.7106 | -1.5985 | 1.2685 | 32.0771 | 40.6904 | 25.2871 | 2.0029 | 32.0017 | 64.0968 | 15.9775 | 14.3784 | -0.1102 | 104.7872 | 235.0891 | 180.8178 | 232.9273 | 161.3322
2012-10-30 18:48 | 28.5881 | 249.4561 | 195.2437 | 247.2943 | 175.6562 | -1.6032 | 1.2691 | 32.0927 | 40.7297 | 25.2872 | 2.0039 | 32.0173 | 64.1593 | 15.9775 | 14.3832 | -0.1105 | 104.889 | 235.0729 | 180.8605 | 232.9111 | 161.273
2012-10-30 18:48 | 28.5836 | 249.4398 | 195.2903 | 247.3642 | 175.7064 | -1.6037 | 1.2697 | 32.1075 | 40.7669 | 25.2874 | 2.0048 | 32.032 | 64.2187 | 15.9774 | 14.4113 | -0.1106 | 104.9856 | 235.0285 | 180.879 | 232.9529 | 161.2951
2012-10-30 18:48 | 28.5836 | 249.4398 | 195.3257 | 247.3094 | 175.7259 | -1.5997 | 1.2702 | 32.1202 | 40.7992 | 25.2875 | 2.0056 | 32.0446 | 64.2696 | 15.9773 | 14.444 | -0.1103 | 105.0687 | 234.9958 | 180.8817 | 232.8654 | 161.2819
2012-10-30 18:48 | 28.5926 | 249.4488 | 195.3504 | 247.3066 | 175.8169 | -1.6043 | 1.2705 | 32.128 | 40.8187 | 25.2876 | 2.0061 | 32.0523 | 64.3005 | 15.9773 | 14.4081 | -0.1106 | 105.1192 | 235.0407 | 180.9423 | 232.8985 | 161.4088
2012-10-30 18:48 | 28.5881 | 249.4365 | 195.3498 | 247.2159 | 175.6211 | -1.6003 | 1.2707 | 32.1323 | 40.83 | 25.2873 | 2.0064 | 32.0566 | 64.3174 | 15.9774 | 14.4117 | -0.1103 | 105.1474 | 235.0248 | 180.9381 | 232.8042 | 161.2094
2012-10-30 18:48 | 28.5888 | 249.4998 | 195.268 | 247.3028 | 175.6647 | -1.6043 | 1.2708 | 32.135 | 40.8364 | 25.2877 | 2.0065 | 32.0592 | 64.3285 | 15.9773 | 14.4124 | -0.1106 | 105.1648 | 235.0874 | 180.8556 | 232.8904 | 161.2523
2012-10-30 18:48 | 28.5746 | 249.47 | 195.3285 | 247.3317 | 175.7091 | -1.6072 | 1.2709 | 32.1373 | 40.8422 | 25.2877 | 2.0067 | 32.0614 | 64.3377 | 15.9772 | 14.3942 | -0.1108 | 105.1799 | 235.0758 | 180.9343 | 232.9375 | 161.3149
2012-10-30 18:48 | 28.5896 | 249.4654 | 195.3317 | 247.2331 | 175.6929 | -1.6014 | 1.2709 | 32.1387 | 40.8455 | 25.2879 | 2.0068 | 32.0629 | 64.3436 | 15.9772 | 14.3969 | -0.1104 | 105.1891 | 235.0685 | 180.9348 | 232.8362 | 161.296
2012-10-30 18:48 | 28.5866 | 249.4976 | 195.3523 | 247.1987 | 175.725 | -1.6003 | 1.271 | 32.1396 | 40.8483 | 25.2875 | 2.0069 | 32.0639 | 64.3475 | 15.9772 | 14.3694 | -0.1103 | 105.1958 | 235.1282 | 180.9829 | 232.8293 | 161.3556
2012-10-30 18:48 | 28.5896 | 249.5006 | 195.3553 | 247.2919 | 175.7202 | -1.598 | 1.271 | 32.1404 | 40.8499 | 25.2878 | 2.0069 | 32.0645 | 64.3501 | 15.9772 | 14.3684 | -0.1102 | 105.2 | 235.1322 | 180.9869 | 232.9235 | 161.3518
2012-10-30 18:48 | 28.5911 | 249.4865 | 195.2978 | 247.3678 | 175.6319 | -1.6014 | 1.271 | 32.1411 | 40.8514 | 25.288 | 2.0069 | 32.0653 | 64.3534 | 15.9771 | 14.3862 | -0.1104 | 105.2048 | 235.1003 | 180.9116 | 232.9816 | 161.2457
2012-10-30 18:48 | 28.5896 | 249.4732 | 195.3042 | 247.2723 | 175.6031 | -1.5991 | 1.271 | 32.1415 | 40.853 | 25.2877 | 2.007 | 32.0656 | 64.3549 | 15.9771 | 14.3806 | -0.1102 | 105.2079 | 235.0926 | 180.9236 | 232.8917 | 161.2225
2012-10-30 18:48 | 28.5903 | 249.47 | 195.3049 | 247.2416 | 175.6155 | -1.6009 | 1.271 | 32.1416 | 40.8531 | 25.2877 | 2.007 | 32.0658 | 64.3555 | 15.9771 | 14.3405 | -0.1104 | 105.2087 | 235.1295 | 180.9644 | 232.9011 | 161.275
2012-10-30 18:48 | 28.5971 | 249.5238 | 195.3077 | 247.1779 | 175.6535 | -1.6043 | 1.271 | 32.1416 | 40.8533 | 25.2877 | 2.007 | 32.0657 | 64.3554 | 15.9771 | 14.384 | -0.1106 | 105.2087 | 235.1398 | 180.9237 | 232.7939 | 161.2695
2012-10-30 18:49 | 28.5926 | 249.4567 | 195.2915 | 247.2831 | 175.6881 | -1.6026 | 1.271 | 32.1418 | 40.8537 | 25.2877 | 2.007 | 32.066 | 64.3564 | 15.9771 | 14.3714 | -0.1105 | 105.2101 | 235.0853 | 180.9201 | 232.9117 | 161.3167
2012-10-30 18:49 | 28.6023 | 249.4859 | 195.364 | 247.3163 | 175.6587 | -1.5991 | 1.271 | 32.1418 | 40.8536 | 25.2878 | 2.007 | 32.0659 | 64.3563 | 15.977 | 14.3688 | -0.1102 | 105.2099 | 235.1171 | 180.9952 | 232.9475 | 161.2899
2012-10-30 18:49 | 28.5873 | 249.4983 | 195.2862 | 247.266 | 175.5423 | -1.6009 | 1.271 | 32.1416 | 40.8532 | 25.2877 | 2.007 | 32.0658 | 64.3562 | 15.9769 | 14.3661 | -0.1104 | 105.2094 | 235.1322 | 180.9201 | 232.8999 | 161.1762
2012-10-30 18:49 | 28.5956 | 249.5223 | 195.3573 | 247.2234 | 175.5974 | -1.6014 | 1.271 | 32.1414 | 40.8525 | 25.2878 | 2.007 | 32.0656 | 64.3553 | 15.9769 | 14.3825 | -0.1104 | 105.2078 | 235.1398 | 180.9748 | 232.8409 | 161.2149
2012-10-30 18:49 | 28.6001 | 249.4955 | 195.2793 | 247.2828 | 175.4263 | -1.6037 | 1.271 | 32.1418 | 40.8534 | 25.2878 | 2.007 | 32.066 | 64.357 | 15.9769 | 14.387 | -0.1106 | 105.2104 | 235.1085 | 180.8923 | 232.8958 | 161.0393
2012-10-30 18:49 | 28.5918 | 249.495 | 195.2907 | 247.3293 | 175.5233 | -1.5974 | 1.271 | 32.142 | 40.8539 | 25.2879 | 2.007 | 32.0662 | 64.3578 | 15.9769 | 14.3787 | -0.1101 | 105.2117 | 235.1163 | 180.912 | 232.9506 | 161.1446
2012-10-30 18:49 | 28.6038 | 249.5344 | 195.3105 | 247.2708 | 175.5822 | -1.5974 | 1.271 | 32.1417 | 40.8533 | 25.2878 | 2.007 | 32.066 | 64.3565 | 15.977 | 14.3948 | -0.1101 | 105.2098 | 235.1396 | 180.9157 | 232.876 | 161.1874
2012-10-30 18:49 | 28.5986 | 249.5253 | 195.2582 | 247.2303 | 175.5731 | -1.5991 | 1.271 | 32.1417 | 40.8533 | 25.2878 | 2.007 | 32.0659 | 64.3568 | 15.9769 | 14.4182 | -0.1102 | 105.2101 | 235.1071 | 180.84 | 232.8121 | 161.1549
2012-10-30 18:49 | 28.609 | 249.524 | 195.2686 | 247.3543 | 175.5288 | -1.6032 | 1.271 | 32.1421 | 40.8539 | 25.288 | 2.007 | 32.0662 | 64.3581 | 15.9768 | 14.4204 | -0.1105 | 105.2121 | 235.1036 | 180.8482 | 232.9339 | 161.1084
2012-10-30 18:49 | 28.5896 | 249.4693 | 195.261 | 247.3428 | 175.6382 | -1.5991 | 1.2711 | 32.1422 | 40.8544 | 25.2879 | 2.007 | 32.0662 | 64.3583 | 15.9768 | 14.401 | -0.1102 | 105.2127 | 235.0683 | 180.86 | 232.9418 | 161.2372
2012-10-30 18:49 | 28.5993 | 249.5299 | 195.2785 | 247.2467 | 175.566 | -1.6009 | 1.2711 | 32.142 | 40.8542 | 25.2877 | 2.007 | 32.0663 | 64.3587 | 15.9769 | 14.4066 | -0.1104 | 105.2129 | 235.1233 | 180.8719 | 232.8401 | 161.1594
2012-10-30 18:49 | 28.5918 | 249.5107 | 195.3182 | 247.2862 | 175.496 | -1.6055 | 1.271 | 32.1422 | 40.8539 | 25.2881 | 2.0071 | 32.0662 | 64.3586 | 15.9768 | 14.4073 | -0.1107 | 105.2126 | 235.1034 | 180.9109 | 232.8789 | 161.0887
2012-10-30 18:49 | 28.5768 | 249.4918 | 195.2599 | 247.3809 | 175.52 | -1.5985 | 1.271 | 32.1418 | 40.8536 | 25.2877 | 2.007 | 32.0659 | 64.3574 | 15.9768 | 14.4208 | -0.1102 | 105.211 | 235.071 | 180.8391 | 232.9601 | 161.0992
2012-10-30 18:49 | 28.5843 | 249.5423 | 195.2478 | 247.3179 | 175.4846 | -1.5991 | 1.271 | 32.1416 | 40.8527 | 25.288 | 2.007 | 32.0657 | 64.3565 | 15.9767 | 14.4039 | -0.1102 | 105.2092 | 235.1384 | 180.8439 | 232.914 | 161.0807
2012-10-30 18:49 | 28.5873 | 249.5297 | 195.2508 | 247.27 | 175.3354 | -1.6032 | 1.271 | 32.1414 | 40.8525 | 25.2878 | 2.007 | 32.0656 | 64.3563 | 15.9767 | 14.4069 | -0.1105 | 105.2089 | 235.1228 | 180.8439 | 232.8631 | 160.9285
2012-10-30 18:49 | 28.5866 | 249.5172 | 195.2108 | 247.3359 | 175.4206 | -1.6026 | 1.271 | 32.1406 | 40.8508 | 25.2876 | 2.007 | 32.0648 | 64.3533 | 15.9767 | 14.3654 | -0.1105 | 105.2041 | 235.1518 | 180.8454 | 232.9705 | 161.0552
2012-10-30 18:49 | 28.5761 | 249.495 | 195.2475 | 247.4037 | 175.4881 | -1.602 | 1.271 | 32.1403 | 40.8495 | 25.2879 | 2.007 | 32.0644 | 64.3518 | 15.9766 | 14.4242 | -0.1104 | 105.2014 | 235.0708 | 180.8233 | 232.9795 | 161.0639
2012-10-30 18:49 | 28.5813 | 249.5276 | 195.2094 | 247.3227 | 175.669 | -1.6014 | 1.271 | 32.1401 | 40.8492 | 25.2878 | 2.0069 | 32.0643 | 64.3513 | 15.9767 | 14.3886 | -0.1104 | 105.2006 | 235.139 | 180.8208 | 232.9341 | 161.2804
2012-10-30 18:49 | 28.5903 | 249.5522 | 195.2106 | 247.273 | 175.6077 | -1.6026 | 1.271 | 32.1401 | 40.8496 | 25.2876 | 2.0069 | 32.0643 | 64.3514 | 15.9766 | 14.4017 | -0.1105 | 105.201 | 235.1505 | 180.8089 | 232.8713 | 161.206
2012-10-30 18:49 | 28.5813 | 249.5158 | 195.2251 | 247.3227 | 175.5245 | -1.5962 | 1.271 | 32.1402 | 40.85 | 25.2875 | 2.007 | 32.0644 | 64.3524 | 15.9765 | 14.3723 | -0.11 | 105.2023 | 235.1435 | 180.8528 | 232.9504 | 161.1522
2012-10-30 18:49 | 28.5761 | 249.5106 | 195.2317 | 247.439 | 175.5076 | -1.5997 | 1.271 | 32.1399 | 40.8486 | 25.2878 | 2.0069 | 32.0639 | 64.3505 | 15.9765 | 14.4079 | -0.1103 | 105.199 | 235.1027 | 180.8238 | 233.0311 | 161.0997
2012-10-30 18:49 | 28.5753 | 249.5333 | 195.2624 | 247.3363 | 175.5459 | -1.5968 | 1.271 | 32.1392 | 40.8474 | 25.2875 | 2.0069 | 32.0633 | 64.348 | 15.9765 | 14.4193 | -0.1101 | 105.1955 | 235.114 | 180.8431 | 232.917 | 161.1266
2012-10-30 18:49 | 28.5776 | 249.5395 | 195.2175 | 247.2877 | 175.5091 | -1.6037 | 1.2709 | 32.1384 | 40.8456 | 25.2873 | 2.0069 | 32.0626 | 64.3453 | 15.9764 | 14.4135 | -0.1106 | 105.1908 | 235.126 | 180.804 | 232.8742 | 161.0956
2012-10-30 18:49 | 28.5791 | 249.5097 | 195.2151 | 247.4028 | 175.4521 | -1.5968 | 1.2709 | 32.1384 | 40.8452 | 25.2876 | 2.0069 | 32.0626 | 64.3459 | 15.9763 | 14.4272 | -0.1101 | 105.1911 | 235.0825 | 180.7879 | 232.9756 | 161.0249
2012-10-30 18:50 | 28.5813 | 249.5785 | 195.2212 | 247.4011 | 175.4738 | -1.602 | 1.2709 | 32.1378 | 40.844 | 25.2875 | 2.0068 | 32.062 | 64.3435 | 15.9763 | 14.4294 | -0.1104 | 105.1874 | 235.1491 | 180.7918 | 232.9717 | 161.0444
2012-10-30 18:50 | 28.5723 | 249.5264 | 195.2161 | 247.3137 | 175.4843 | -1.6009 | 1.2709 | 32.1377 | 40.8433 | 25.2877 | 2.0068 | 32.0619 | 64.3429 | 15.9764 | 14.4612 | -0.1104 | 105.1861 | 235.0652 | 180.7549 | 232.8525 | 161.0231
2012-10-30 18:50 | 28.5716 | 249.5335 | 195.2037 | 247.3248 | 175.4836 | -1.6037 | 1.2709 | 32.1374 | 40.8432 | 25.2873 | 2.0068 | 32.0616 | 64.3422 | 15.9763 | 14.3952 | -0.1106 | 105.1853 | 235.1383 | 180.8085 | 232.9296 | 161.0884
2012-10-30 18:50 | 28.5753 | 249.5333 | 195.2113 | 247.3872 | 175.5927 | -1.6014 | 1.2709 | 32.137 | 40.8415 | 25.2878 | 2.0068 | 32.0612 | 64.3404 | 15.9763 | 14.452 | -0.1104 | 105.1819 | 235.0813 | 180.7593 | 232.9352 | 161.1407
2012-10-30 18:50 | 28.5821 | 249.5088 | 195.2338 | 247.394 | 175.5097 | -1.6003 | 1.2709 | 32.137 | 40.8417 | 25.2876 | 2.0068 | 32.0611 | 64.34 | 15.9762 | 14.4098 | -0.1103 | 105.1817 | 235.099 | 180.824 | 232.9842 | 161.0999
2012-10-30 18:50 | 28.5813 | 249.5628 | 195.2016 | 247.2836 | 175.3997 | -1.5957 | 1.2708 | 32.1365 | 40.8402 | 25.2877 | 2.0068 | 32.0606 | 64.3384 | 15.9762 | 14.3886 | -0.11 | 105.1786 | 235.1742 | 180.813 | 232.895 | 161.0111
2012-10-30 18:50 | 28.5768 | 249.574 | 195.1499 | 247.3652 | 175.4576 | -1.6014 | 1.2708 | 32.1361 | 40.8397 | 25.2874 | 2.0067 | 32.0603 | 64.3367 | 15.9763 | 14.4168 | -0.1104 | 105.1764 | 235.1572 | 180.7331 | 232.9484 | 161.0408
2012-10-30 18:50 | 28.5836 | 249.5612 | 195.2117 | 247.423 | 175.4644 | -1.6014 | 1.2708 | 32.1355 | 40.8379 | 25.2876 | 2.0067 | 32.0597 | 64.3353 | 15.976 | 14.444 | -0.1104 | 105.1732 | 235.1172 | 180.7677 | 232.979 | 161.0204
2012-10-30 18:50 | 28.5746 | 249.5796 | 195.1674 | 247.2612 | 175.3695 | -1.6032 | 1.2708 | 32.1358 | 40.8387 | 25.2874 | 2.0067 | 32.0599 | 64.3359 | 15.9761 | 14.3778 | -0.1105 | 105.1746 | 235.2018 | 180.7896 | 232.8834 | 160.9917
2012-10-30 18:50 | 28.5828 | 249.533 | 195.1795 | 247.3791 | 175.3114 | -1.5951 | 1.2708 | 32.1356 | 40.8379 | 25.2878 | 2.0067 | 32.0598 | 64.3353 | 15.9761 | 14.4187 | -0.11 | 105.1732 | 235.1143 | 180.7608 | 232.9604 | 160.8927
2012-10-30 18:50 | 28.5798 | 249.577 | 195.1608 | 247.4388 | 175.2967 | -1.6009 | 1.2708 | 32.135 | 40.8366 | 25.2875 | 2.0067 | 32.0592 | 64.3327 | 15.9762 | 14.4075 | -0.1104 | 105.1693 | 235.1695 | 180.7533 | 233.0313 | 160.8892
2012-10-30 18:50 | 28.5753 | 249.5411 | 195.1759 | 247.3677 | 175.4873 | -1.6043 | 1.2708 | 32.1344 | 40.8349 | 25.2877 | 2.0066 | 32.0586 | 64.3303 | 15.9762 | 14.4275 | -0.1106 | 105.1652 | 235.1136 | 180.7484 | 232.9402 | 161.0598
2012-10-30 18:50 | 28.5806 | 249.5778 | 195.2087 | 247.322 | 175.4068 | -1.6026 | 1.2665 | 32.0266 | 40.563 | 25.2866 | 2 | 31.9514 | 63.9014 | 15.9761 | 14.4083 | -0.1105 | 104.4644 | 235.1695 | 180.8004 | 232.9137 | 160.9985
2012-10-30 18:50 | 28.5828 | 249.5134 | 195.1402 | 247.3712 | 175.4948 | -1.5974 | 1.2647 | 31.9788 | 40.443 | 25.286 | 1.9969 | 31.9035 | 63.7092 | 15.9762 | 14.3901 | -0.1101 | 104.1523 | 235.1233 | 180.7501 | 232.9811 | 161.1047
2012-10-30 18:50 | 28.5828 | 249.5134 | 195.097 | 247.4418 | 175.6861 | -1.5991 | 1.2639 | 31.959 | 40.3939 | 25.2855 | 1.9957 | 31.8839 | 63.63 | 15.9764 | 14.4309 | -0.1102 | 104.0239 | 235.0825 | 180.6661 | 233.0109 | 161.2552
2012-10-30 18:50 | 28.5798 | 249.5339 | 195.1451 | 247.3682 | 175.7416 | -1.6009 | 1.2636 | 31.9495 | 40.3699 | 25.2855 | 1.9951 | 31.8743 | 63.5914 | 15.9766 | 14.3871 | -0.1104 | 103.9613 | 235.1468 | 180.758 | 232.9811 | 161.3545
2012-10-30 18:50 | 28.5873 | 249.5414 | 195.1761 | 247.2739 | 175.5696 | -1.6014 | 1.2634 | 31.9462 | 40.3609 | 25.2858 | 1.9948 | 31.8707 | 63.5773 | 15.9765 | 14.4069 | -0.1104 | 103.9382 | 235.1345 | 180.7692 | 232.867 | 161.1627
2012-10-30 18:50 | 28.5783 | 249.4893 | 195.1475 | 247.3942 | 175.4396 | -1.6009 | 1.2633 | 31.9445 | 40.3567 | 25.2858 | 1.9947 | 31.8692 | 63.5705 | 15.9767 | 14.3815 | -0.1104 | 103.9272 | 235.1078 | 180.766 | 233.0127 | 161.0581
2012-10-30 18:50 | 28.5701 | 249.4537 | 195.1393 | 247.386 | 175.4353 | -1.5985 | 1.2634 | 31.9465 | 40.3617 | 25.2858 | 1.9949 | 31.8711 | 63.5791 | 15.9764 | 14.4223 | -0.1102 | 103.9408 | 235.0314 | 180.717 | 232.9637 | 161.013
2012-10-30 18:50 | 28.5708 | 249.4505 | 195.1636 | 247.2848 | 175.6936 | -1.6009 | 1.2635 | 31.9483 | 40.3665 | 25.2857 | 1.995 | 31.8728 | 63.5854 | 15.9765 | 14.3985 | -0.1104 | 103.952 | 235.052 | 180.7651 | 232.8863 | 161.2951
2012-10-30 18:50 | 28.5701 | 249.4694 | 195.1511 | 247.335 | 175.8022 | -1.5974 | 1.2636 | 31.9509 | 40.3733 | 25.2855 | 1.9951 | 31.8755 | 63.596 | 15.9766 | 14.4141 | -0.1101 | 103.9693 | 235.0553 | 180.737 | 232.9209 | 161.3881
2012-10-30 18:50 | 28.5671 | 249.4312 | 195.1874 | 247.43 | 175.9787 | -1.598 | 1.2636 | 31.9515 | 40.3739 | 25.2861 | 1.9952 | 31.876 | 63.5978 | 15.9767 | 14.3867 | -0.1102 | 103.9717 | 235.0445 | 180.8007 | 233.0433 | 161.592
2012-10-30 18:50 | 28.5649 | 249.4407 | 195.307 | 247.2789 | 176.0858 | -1.6084 | 1.2685 | 32.0752 | 40.686 | 25.2869 | 2.0027 | 31.9969 | 64.0811 | 15.9766 | 14.413 | -0.1109 | 104.7671 | 235.0277 | 180.894 | 232.8659 | 161.6728
2012-10-30 18:50 | 28.5664 | 249.4461 | 195.2888 | 247.2922 | 175.8258 | -1.5991 | 1.2691 | 32.0918 | 40.7271 | 25.2874 | 2.0039 | 32.0159 | 64.1575 | 15.9766 | 14.4023 | -0.1102 | 104.8845 | 235.0438 | 180.8865 | 232.8899 | 161.4235
2012-10-30 18:50 | 28.5551 | 249.4152 | 195.3168 | 247.3984 | 175.7677 | -1.6055 | 1.2697 | 32.1064 | 40.7643 | 25.2874 | 2.0049 | 32.031 | 64.2183 | 15.9765 | 14.342 | -0.1107 | 104.9826 | 235.0732 | 180.9748 | 233.0564 | 161.4257
2012-10-30 18:51 | 28.5664 | 249.407 | 195.3438 | 247.2843 | 175.8531 | -1.5968 | 1.2701 | 32.1181 | 40.794 | 25.2873 | 2.0056 | 32.0423 | 64.2638 | 15.9765 | 14.4104 | -0.1101 | 105.0578 | 234.9966 | 180.9334 | 232.8739 | 161.4427
2012-10-30 18:51 | 28.5611 | 249.4252 | 195.3503 | 247.232 | 175.637 | -1.5991 | 1.2704 | 32.1259 | 40.8135 | 25.2875 | 2.0061 | 32.0499 | 64.2944 | 15.9765 | 14.3847 | -0.1102 | 105.108 | 235.0405 | 180.9656 | 232.8473 | 161.2523
2012-10-30 18:51 | 28.5604 | 249.4636 | 195.3771 | 247.3293 | 175.7261 | -1.6055 | 1.2706 | 32.131 | 40.826 | 25.2878 | 2.0064 | 32.0551 | 64.315 | 15.9765 | 14.3963 | -0.1107 | 105.141 | 235.0673 | 180.9808 | 232.933 | 161.3298
2012-10-30 18:51 | 28.5626 | 249.4893 | 195.3715 | 247.351 | 175.9391 | -1.6026 | 1.2707 | 32.1341 | 40.834 | 25.2878 | 2.0066 | 32.0582 | 64.3282 | 15.9763 | 14.423 | -0.1105 | 105.1622 | 235.0663 | 180.9485 | 232.928 | 161.5161
2012-10-30 18:51 | 28.5634 | 249.4431 | 195.4076 | 247.1481 | 175.9828 | -1.5962 | 1.2708 | 32.1363 | 40.8397 | 25.2877 | 2.0067 | 32.0604 | 64.3366 | 15.9764 | 14.3789 | -0.11 | 105.1764 | 235.0642 | 181.0287 | 232.7692 | 161.6039
2012-10-30 18:51 | 28.5686 | 249.464 | 195.4246 | 247.3492 | 175.9022 | -1.6009 | 1.2709 | 32.1386 | 40.8451 | 25.288 | 2.0069 | 32.0627 | 64.3457 | 15.9765 | 14.3922 | -0.1104 | 105.1908 | 235.0718 | 181.0324 | 232.957 | 161.51
2012-10-30 18:51 | 28.5506 | 249.4538 | 195.4223 | 247.3038 | 175.9857 | -1.5985 | 1.271 | 32.14 | 40.849 | 25.2878 | 2.007 | 32.0641 | 64.352 | 15.9763 | 14.4191 | -0.1102 | 105.201 | 235.0347 | 181.0032 | 232.8847 | 161.5666
2012-10-30 18:51 | 28.5626 | 249.458 | 195.4619 | 247.2649 | 176.107 | -1.602 | 1.271 | 32.1409 | 40.8509 | 25.2879 | 2.007 | 32.0649 | 64.3551 | 15.9763 | 14.4148 | -0.1104 | 105.206 | 235.0432 | 181.0471 | 232.8501 | 161.6922
2012-10-30 18:51 | 28.5634 | 249.5292 | 195.5452 | 247.203 | 176.1 | -1.6049 | 1.271 | 32.1417 | 40.8534 | 25.2878 | 2.0071 | 32.0658 | 64.3586 | 15.9764 | 14.4278 | -0.1106 | 105.212 | 235.1014 | 181.1174 | 232.7752 | 161.6722
2012-10-30 18:51 | 28.5574 | 249.4371 | 195.5235 | 247.3341 | 175.8012 | -1.5951 | 1.2711 | 32.1422 | 40.8547 | 25.2877 | 2.0071 | 32.0662 | 64.3604 | 15.9763 | 14.4463 | -0.11 | 105.2151 | 234.9908 | 181.0772 | 232.8878 | 161.3549
2012-10-30 18:51 | 28.5664 | 249.5205 | 195.5403 | 247.3039 | 175.9117 | -1.5985 | 1.2711 | 32.1426 | 40.8552 | 25.2881 | 2.0071 | 32.0667 | 64.3615 | 15.9765 | 14.4227 | -0.1102 | 105.2167 | 235.0978 | 181.1176 | 232.8812 | 161.489
2012-10-30 18:51 | 28.5559 | 249.5061 | 195.5102 | 247.2033 | 175.866 | -1.6072 | 1.2711 | 32.1428 | 40.8555 | 25.2882 | 2.0071 | 32.0669 | 64.3629 | 15.9763 | 14.3999 | -0.1108 | 105.2184 | 235.1062 | 181.1103 | 232.8034 | 161.4661
2012-10-30 18:51 | 28.5604 | 249.4636 | 195.4636 | 247.3253 | 175.7378 | -1.6003 | 1.2711 | 32.1428 | 40.8557 | 25.2881 | 2.0072 | 32.067 | 64.3636 | 15.9763 | 14.3922 | -0.1103 | 105.2193 | 235.0714 | 181.0714 | 232.9331 | 161.3456
2012-10-30 18:51 | 28.5506 | 249.489 | 195.4695 | 247.3821 | 175.7163 | -1.5997 | 1.2711 | 32.1431 | 40.857 | 25.2877 | 2.0072 | 32.0671 | 64.3645 | 15.9762 | 14.415 | -0.1103 | 105.2215 | 235.074 | 181.0545 | 232.9671 | 161.3013
2012-10-30 18:51 | 28.5551 | 249.4975 | 195.4701 | 247.2103 | 175.8262 | -1.6061 | 1.2711 | 32.1428 | 40.856 | 25.2879 | 2.0072 | 32.0669 | 64.3641 | 15.9761 | 14.3869 | -0.1107 | 105.2201 | 235.1106 | 181.0832 | 232.8234 | 161.4393
2012-10-30 18:51 | 28.5611 | 249.5152 | 195.4132 | 247.3182 | 175.9181 | -1.6043 | 1.2711 | 32.1427 | 40.8557 | 25.2879 | 2.0072 | 32.0669 | 64.3635 | 15.9762 | 14.3684 | -0.1106 | 105.2192 | 235.1468 | 181.0448 | 232.9498 | 161.5497
2012-10-30 18:52 | 28.5664 | 249.5518 | 195.4853 | 247.3274 | 175.9117 | -1.5997 | 1.2711 | 32.1428 | 40.8561 | 25.2878 | 2.0072 | 32.0669 | 64.3642 | 15.976 | 14.3574 | -0.1103 | 105.2203 | 235.1944 | 181.1279 | 232.97 | 161.5543
2012-10-30 18:52 | 28.5776 | 249.5748 | 195.4769 | 247.3112 | 175.8175 | -1.5997 | 1.2711 | 32.1419 | 40.8542 | 25.2876 | 2.0071 | 32.066 | 64.361 | 15.9759 | 14.3604 | -0.1103 | 105.2152 | 235.2144 | 181.1165 | 232.9508 | 161.4571
2012-10-30 18:52 | 28.5716 | 249.6197 | 195.408 | 247.3248 | 175.7529 | -1.6026 | 1.271 | 32.1415 | 40.8526 | 25.2879 | 2.0071 | 32.0656 | 64.3595 | 15.9759 | 14.3177 | -0.1105 | 105.2121 | 235.302 | 181.0903 | 233.0071 | 161.4352
2012-10-30 18:52 | 28.5731 | 249.582 | 195.4566 | 247.4242 | 175.7115 | -1.598 | 1.271 | 32.142 | 40.8538 | 25.2879 | 2.0071 | 32.066 | 64.3612 | 15.9759 | 14.3886 | -0.1102 | 105.215 | 235.1934 | 181.068 | 233.0356 | 161.3229
2012-10-30 18:52 | 28.5716 | 249.6118 | 195.4512 | 247.2934 | 175.6671 | -1.6014 | 1.271 | 32.1418 | 40.8536 | 25.2877 | 2.0071 | 32.0659 | 64.3608 | 15.9759 | 14.3218 | -0.1104 | 105.2144 | 235.29 | 181.1294 | 232.9716 | 161.3453
2012-10-30 18:52 | 28.5723 | 249.6008 | 195.393 | 247.3607 | 175.5585 | -1.6014 | 1.271 | 32.1416 | 40.8527 | 25.288 | 2.0071 | 32.0657 | 64.3606 | 15.9758 | 14.3511 | -0.1104 | 105.2133 | 235.2497 | 181.0419 | 233.0096 | 161.2074
2012-10-30 18:52 | 28.5768 | 249.5857 | 195.4014 | 247.4397 | 175.6293 | -1.6026 | 1.271 | 32.1413 | 40.8524 | 25.2878 | 2.0071 | 32.0655 | 64.3596 | 15.9758 | 14.3352 | -0.1105 | 105.212 | 235.2505 | 181.0662 | 233.1045 | 161.2941
2012-10-30 18:52 | 28.5806 | 249.5934 | 195.3816 | 247.3024 | 175.8595 | -1.6003 | 1.271 | 32.1414 | 40.8525 | 25.2877 | 2.0071 | 32.0655 | 64.3595 | 15.9758 | 14.3226 | -0.1103 | 105.212 | 235.2708 | 181.059 | 232.9798 | 161.5369
2012-10-30 18:52 | 28.5753 | 249.5881 | 195.3881 | 247.3598 | 175.7176 | -1.6032 | 1.271 | 32.141 | 40.8519 | 25.2876 | 2.0071 | 32.0652 | 64.3588 | 15.9757 | 14.3622 | -0.1105 | 105.2107 | 235.2259 | 181.0259 | 232.9976 | 161.3554
2012-10-30 18:52 | 28.5723 | 249.5812 | 195.389 | 247.4469 | 175.7224 | -1.6049 | 1.271 | 32.1411 | 40.8519 | 25.2877 | 2.0071 | 32.0652 | 64.3584 | 15.9758 | 14.3633 | -0.1106 | 105.2103 | 235.2179 | 181.0257 | 233.0836 | 161.3591
2012-10-30 18:52 | 28.5881 | 249.5853 | 195.4284 | 247.4235 | 175.8631 | -1.6043 | 1.271 | 32.1409 | 40.8509 | 25.2881 | 2.0071 | 32.0649 | 64.3578 | 15.9757 | 14.3873 | -0.1106 | 105.2086 | 235.198 | 181.0411 | 233.0362 | 161.4758
2012-10-30 18:52 | 28.5798 | 249.5926 | 195.4083 | 247.3487 | 175.7806 | -1.6014 | 1.271 | 32.1403 | 40.8498 | 25.2877 | 2.0071 | 32.0644 | 64.3562 | 15.9756 | 14.4238 | -0.1104 | 105.2061 | 235.1688 | 180.9845 | 232.9249 | 161.3568
2012-10-30 18:52 | 28.5836 | 249.6551 | 195.4632 | 247.4621 | 175.8664 | -1.6014 | 1.271 | 32.1396 | 40.8481 | 25.2876 | 2.007 | 32.0636 | 64.3527 | 15.9756 | 14.4113 | -0.1104 | 105.2008 | 235.2438 | 181.0519 | 233.0508 | 161.4551
2012-10-30 18:52 | 28.5798 | 249.5887 | 195.4476 | 247.3839 | 175.9056 | -1.5991 | 1.2709 | 32.139 | 40.8464 | 25.2877 | 2.007 | 32.063 | 64.35 | 15.9757 | 14.4483 | -0.1102 | 105.1965 | 235.1404 | 180.9993 | 232.9356 | 161.4573
2012-10-30 18:52 | 28.5873 | 249.6588 | 195.5141 | 247.3366 | 175.796 | -1.6037 | 1.2709 | 32.1387 | 40.8454 | 25.2879 | 2.007 | 32.0628 | 64.3489 | 15.9757 | 14.4354 | -0.1106 | 105.1943 | 235.2234 | 181.0787 | 232.9012 | 161.3606
2012-10-30 18:52 | 28.5791 | 249.6193 | 195.4862 | 247.4185 | 175.8112 | -1.5951 | 1.2709 | 32.1381 | 40.844 | 25.2878 | 2.0069 | 32.0621 | 64.3469 | 15.9756 | 14.4517 | -0.11 | 105.1908 | 235.1676 | 181.0345 | 232.9668 | 161.3595
2012-10-30 18:52 | 28.5821 | 249.5988 | 195.5206 | 247.3274 | 175.8259 | -1.5991 | 1.2709 | 32.1377 | 40.8432 | 25.2877 | 2.0069 | 32.0619 | 64.3457 | 15.9756 | 14.4302 | -0.1102 | 105.1889 | 235.1686 | 181.0904 | 232.8972 | 161.3957
2012-10-30 18:52 | 28.5888 | 249.6369 | 195.4605 | 247.3929 | 175.8014 | -1.6037 | 1.2709 | 32.1376 | 40.8435 | 25.2874 | 2.0069 | 32.0617 | 64.3453 | 15.9756 | 14.4124 | -0.1106 | 105.1888 | 235.2245 | 181.0481 | 232.9805 | 161.389
2012-10-30 18:52 | 28.5881 | 249.6048 | 195.4795 | 247.5528 | 175.547 | -1.6055 | 1.2709 | 32.1374 | 40.8424 | 25.2878 | 2.0069 | 32.0615 | 64.3449 | 15.9755 | 14.3628 | -0.1107 | 105.1873 | 235.242 | 181.1167 | 233.19 | 161.1842
2012-10-30 18:52 | 28.5828 | 249.6426 | 195.5213 | 247.3595 | 175.6236 | -1.6032 | 1.2709 | 32.1372 | 40.8422 | 25.2876 | 2.0069 | 32.0613 | 64.3444 | 15.9754 | 14.4391 | -0.1105 | 105.1866 | 235.2035 | 181.0822 | 232.9204 | 161.1845
2012-10-30 18:52 | 28.5866 | 249.6581 | 195.3994 | 247.3946 | 175.5533 | -1.6037 | 1.2709 | 32.1371 | 40.8419 | 25.2876 | 2.0069 | 32.0611 | 64.3435 | 15.9755 | 14.4306 | -0.1106 | 105.1854 | 235.2275 | 180.9688 | 232.964 | 161.1227
2012-10-30 18:52 | 28.5776 | 249.6257 | 195.4022 | 247.5031 | 175.5521 | -1.5997 | 1.2709 | 32.1369 | 40.8417 | 25.2874 | 2.0069 | 32.061 | 64.3431 | 15.9755 | 14.4176 | -0.1103 | 105.1847 | 235.2081 | 180.9846 | 233.0855 | 161.1345
2012-10-30 18:52 | 28.5873 | 249.6158 | 195.4276 | 247.4071 | 175.5618 | -1.5945 | 1.2708 | 32.1363 | 40.8398 | 25.2876 | 2.0069 | 32.0605 | 64.3406 | 15.9755 | 14.4069 | -0.1099 | 105.1804 | 235.2089 | 181.0207 | 233.0002 | 161.1549
2012-10-30 18:52 | 28.5873 | 249.651 | 195.4001 | 247.3326 | 175.5696 | -1.6003 | 1.2708 | 32.1358 | 40.8385 | 25.2876 | 2.0068 | 32.0599 | 64.3385 | 15.9754 | 14.3742 | -0.1103 | 105.177 | 235.2768 | 181.0259 | 232.9584 | 161.1954
2012-10-30 18:52 | 28.5933 | 249.6688 | 195.3904 | 247.4836 | 175.5522 | -1.6089 | 1.2708 | 32.1354 | 40.837 | 25.2879 | 2.0068 | 32.0596 | 64.3373 | 15.9755 | 14.4496 | -0.1109 | 105.1743 | 235.2192 | 180.9408 | 233.034 | 161.1026
2012-10-30 18:52 | 28.5843 | 249.6245 | 195.3735 | 247.4511 | 175.59 | -1.6055 | 1.2708 | 32.1353 | 40.8373 | 25.2876 | 2.0068 | 32.0595 | 64.3371 | 15.9754 | 14.412 | -0.1107 | 105.1744 | 235.2125 | 180.9615 | 233.0391 | 161.178
2012-10-30 18:53 | 28.5813 | 249.645 | 195.3627 | 247.311 | 175.5363 | -1.5968 | 1.2708 | 32.1347 | 40.8357 | 25.2877 | 2.0068 | 32.0588 | 64.3349 | 15.9753 | 14.3682 | -0.1101 | 105.1706 | 235.2768 | 180.9945 | 232.9428 | 161.1681
2012-10-30 18:53 | 28.5828 | 249.6113 | 195.3406 | 247.4378 | 175.4597 | -1.6055 | 1.2651 | 31.9893 | 40.4687 | 25.2865 | 1.9974 | 31.91 | 63.7381 | 15.9755 | 14.4105 | -0.1107 | 104.2069 | 235.2008 | 180.9301 | 233.0273 | 161.0492
2012-10-30 18:53 | 28.5881 | 249.6127 | 195.3223 | 247.4392 | 175.547 | -1.6078 | 1.2646 | 31.9768 | 40.438 | 25.2859 | 1.9969 | 31.9012 | 63.7033 | 15.9754 | 14.3954 | -0.1108 | 104.1413 | 235.2173 | 180.9269 | 233.0438 | 161.1516
2012-10-30 18:53 | 28.5941 | 249.5913 | 195.3008 | 247.3473 | 175.6974 | -1.6003 | 1.2641 | 31.9631 | 40.4036 | 25.2858 | 1.996 | 31.8878 | 63.6493 | 15.9756 | 14.43 | -0.1103 | 104.0529 | 235.1613 | 180.8708 | 232.9173 | 161.2674
2012-10-30 18:53 | 28.5881 | 249.5344 | 195.2477 | 247.4862 | 175.3713 | -1.6043 | 1.2637 | 31.9528 | 40.3778 | 25.2858 | 1.9954 | 31.8775 | 63.6079 | 15.9756 | 14.4485 | -0.1106 | 103.9857 | 235.0859 | 180.7992 | 233.0377 | 160.9228
2012-10-30 18:53 | 28.5821 | 249.5284 | 195.2377 | 247.4528 | 175.5175 | -1.5991 | 1.2635 | 31.9487 | 40.3671 | 25.286 | 1.9951 | 31.8733 | 63.5911 | 15.9756 | 14.4302 | -0.1102 | 103.9582 | 235.0982 | 180.8075 | 233.0226 | 161.0873
2012-10-30 18:53 | 28.5828 | 249.5095 | 195.2895 | 247.3399 | 175.8071 | -1.598 | 1.2634 | 31.9458 | 40.3609 | 25.2852 | 1.995 | 31.8705 | 63.5801 | 15.9755 | 14.4024 | -0.1102 | 103.9409 | 235.1071 | 180.8871 | 232.9375 | 161.4047
2012-10-30 18:53 | 28.5798 | 249.5143 | 195.314 | 247.3251 | 175.7924 | -1.5985 | 1.2635 | 31.9473 | 40.3638 | 25.2857 | 1.995 | 31.8718 | 63.5854 | 15.9756 | 14.4198 | -0.1102 | 103.9491 | 235.0945 | 180.8942 | 232.9053 | 161.3726
2012-10-30 18:53 | 28.5836 | 249.4946 | 195.306 | 247.4426 | 175.8079 | -1.5997 | 1.2636 | 31.9502 | 40.3711 | 25.2857 | 1.9952 | 31.8748 | 63.5968 | 15.9757 | 14.4276 | -0.1103 | 103.9679 | 235.067 | 180.8784 | 233.015 | 161.3803
2012-10-30 18:53 | 28.5858 | 249.4499 | 195.3318 | 247.3311 | 175.9272 | -1.5991 | 1.2636 | 31.951 | 40.3736 | 25.2855 | 1.9953 | 31.8757 | 63.6005 | 15.9756 | 14.438 | -0.1102 | 103.9741 | 235.0119 | 180.8938 | 232.8931 | 161.4892
2012-10-30 18:53 | 28.5753 | 249.4433 | 195.3763 | 247.3206 | 175.9128 | -1.6009 | 1.2637 | 31.9533 | 40.3784 | 25.2861 | 1.9954 | 31.8778 | 63.6083 | 15.9758 | 14.3989 | -0.1104 | 103.9867 | 235.0444 | 180.9774 | 232.9217 | 161.5139
2012-10-30 18:53 | 28.5813 | 249.4336 | 195.4138 | 247.3658 | 175.747 | -1.602 | 1.2638 | 31.9553 | 40.3841 | 25.2857 | 1.9955 | 31.8799 | 63.6165 | 15.9758 | 14.3927 | -0.1104 | 104.0006 | 235.0409 | 181.0211 | 232.9731 | 161.3543
2012-10-30 18:53 | 28.5761 | 249.401 | 195.385 | 247.3332 | 175.6481 | -1.6049 | 1.2666 | 32.0277 | 40.5661 | 25.2865 | 2 | 31.9524 | 63.906 | 15.9759 | 14.412 | -0.1106 | 104.4721 | 234.989 | 180.973 | 232.9212 | 161.2361
2012-10-30 18:53 | 28.5761 | 249.4323 | 195.3418 | 247.2078 | 175.6208 | -1.5991 | 1.2688 | 32.0838 | 40.7078 | 25.2868 | 2.0035 | 32.008 | 64.1291 | 15.9758 | 14.4242 | -0.1102 | 104.8369 | 235.0081 | 180.9176 | 232.7836 | 161.1966
2012-10-30 18:53 | 28.5768 | 249.433 | 195.3385 | 247.3652 | 175.3912 | -1.6032 | 1.2694 | 32.0999 | 40.7476 | 25.2875 | 2.0045 | 32.0237 | 64.1919 | 15.9758 | 14.4249 | -0.1105 | 104.9395 | 235.0081 | 180.9136 | 232.9403 | 160.9663
2012-10-30 18:53 | 28.5738 | 249.4457 | 195.3198 | 247.2956 | 175.439 | -1.6009 | 1.27 | 32.1158 | 40.788 | 25.2875 | 2.0055 | 32.04 | 64.2577 | 15.9757 | 14.3974 | -0.1104 | 105.0457 | 235.0483 | 180.9224 | 232.8982 | 161.0416
2012-10-30 18:53 | 28.5783 | 249.4189 | 195.34 | 247.3119 | 175.7362 | -1.5957 | 1.2704 | 32.1264 | 40.8149 | 25.2875 | 2.0062 | 32.0505 | 64.3003 | 15.9756 | 14.4223 | -0.11 | 105.1152 | 234.9966 | 180.9177 | 232.8896 | 161.3139
2012-10-30 18:53 | 28.5866 | 249.4937 | 195.368 | 247.1909 | 175.7757 | -1.6032 | 1.2707 | 32.1316 | 40.8281 | 25.2875 | 2.0065 | 32.0558 | 64.321 | 15.9757 | 14.447 | -0.1105 | 105.1492 | 235.0467 | 180.921 | 232.7439 | 161.3287
2012-10-30 18:53 | 28.5821 | 249.4031 | 195.3281 | 247.3627 | 175.7556 | -1.6066 | 1.2708 | 32.1355 | 40.8377 | 25.2877 | 2.0068 | 32.0596 | 64.3377 | 15.9754 | 14.4261 | -0.1108 | 105.1753 | 234.977 | 180.902 | 232.9366 | 161.3295
2012-10-30 18:53 | 28.5723 | 249.4481 | 195.4087 | 247.2315 | 175.8746 | -1.6061 | 1.2709 | 32.1376 | 40.8432 | 25.2876 | 2.0069 | 32.0617 | 64.3453 | 15.9756 | 14.4245 | -0.1107 | 105.1886 | 235.0236 | 180.9842 | 232.807 | 161.4501
2012-10-30 18:53 | 28.5783 | 249.4424 | 195.399 | 247.2375 | 175.5918 | -1.5991 | 1.271 | 32.1396 | 40.8478 | 25.2878 | 2.007 | 32.0636 | 64.3528 | 15.9756 | 14.4019 | -0.1102 | 105.2006 | 235.0405 | 180.9971 | 232.8356 | 161.1899
2012-10-30 18:53 | 28.5753 | 249.4394 | 195.337 | 247.3755 | 175.5693 | -1.5991 | 1.271 | 32.1411 | 40.8519 | 25.2877 | 2.0072 | 32.0652 | 64.3601 | 15.9754 | 14.4275 | -0.1102 | 105.212 | 235.0119 | 180.9095 | 232.948 | 161.1418
2012-10-30 18:53 | 28.5791 | 249.4353 | 195.384 | 247.2069 | 175.6746 | -1.6037 | 1.2711 | 32.142 | 40.8542 | 25.2877 | 2.0072 | 32.066 | 64.3629 | 15.9755 | 14.4191 | -0.1106 | 105.2171 | 235.0162 | 180.9649 | 232.7878 | 161.2555
2012-10-30 18:53 | 28.5693 | 249.4529 | 195.3507 | 247.1579 | 175.6765 | -1.6049 | 1.2711 | 32.1427 | 40.8564 | 25.2875 | 2.0073 | 32.0668 | 64.3661 | 15.9754 | 14.3889 | -0.1106 | 105.2225 | 235.064 | 180.9618 | 232.769 | 161.2876
2012-10-30 18:53 | 28.5723 | 249.4285 | 195.3419 | 247.3255 | 175.8044 | -1.5997 | 1.2711 | 32.1436 | 40.858 | 25.2878 | 2.0073 | 32.0677 | 64.3697 | 15.9754 | 14.4123 | -0.1103 | 105.2277 | 235.0162 | 180.9296 | 232.9132 | 161.3921
2012-10-30 18:54 | 28.5716 | 249.4435 | 195.3923 | 247.2934 | 175.8935 | -1.5991 | 1.2711 | 32.1441 | 40.8592 | 25.288 | 2.0073 | 32.0681 | 64.3717 | 15.9754 | 14.3993 | -0.1102 | 105.2309 | 235.0442 | 180.993 | 232.8941 | 161.4942
2012-10-30 18:54 | 28.5768 | 249.4604 | 195.4093 | 247.1733 | 175.9806 | -1.5985 | 1.2711 | 32.1446 | 40.8603 | 25.2879 | 2.0074 | 32.0685 | 64.373 | 15.9755 | 14.4331 | -0.1102 | 105.2334 | 235.0273 | 180.9762 | 232.7402 | 161.5475
2012-10-30 18:54 | 28.5731 | 249.4959 | 195.4291 | 247.3067 | 175.731 | -1.5997 | 1.2711 | 32.1445 | 40.8603 | 25.2879 | 2.0074 | 32.0686 | 64.3734 | 15.9755 | 14.3927 | -0.1103 | 105.2337 | 235.1032 | 181.0364 | 232.914 | 161.3383
2012-10-30 18:54 | 28.5828 | 249.4821 | 195.4467 | 247.3007 | 175.9671 | -1.6014 | 1.2712 | 32.1448 | 40.8612 | 25.2877 | 2.0074 | 32.0687 | 64.3742 | 15.9754 | 14.4187 | -0.1104 | 105.2354 | 235.0634 | 181.028 | 232.882 | 161.5484
2012-10-30 18:54 | 28.5776 | 249.4808 | 195.4454 | 247.1976 | 176.001 | -1.5951 | 1.2712 | 32.1447 | 40.8611 | 25.2876 | 2.0074 | 32.0687 | 64.374 | 15.9754 | 14.3808 | -0.11 | 105.2351 | 235.1 | 181.0646 | 232.8168 | 161.6202
2012-10-30 18:54 | 28.5753 | 249.4903 | 195.4628 | 247.2619 | 175.9518 | -1.5997 | 1.2711 | 32.1446 | 40.8603 | 25.2879 | 2.0074 | 32.0686 | 64.3743 | 15.9752 | 14.3867 | -0.1103 | 105.2346 | 235.1036 | 181.0761 | 232.8752 | 161.5651
2012-10-30 18:54 | 28.5776 | 249.4573 | 195.4651 | 247.2955 | 175.8331 | -1.5968 | 1.2712 | 32.1449 | 40.8617 | 25.2876 | 2.0074 | 32.0688 | 64.3748 | 15.9754 | 14.4094 | -0.1101 | 105.2365 | 235.0479 | 181.0557 | 232.8861 | 161.4237
2012-10-30 18:54 | 28.5701 | 249.4576 | 195.5047 | 247.3546 | 175.7749 | -1.6009 | 1.2712 | 32.1452 | 40.8622 | 25.2878 | 2.0074 | 32.0693 | 64.377 | 15.9753 | 14.4019 | -0.1104 | 105.2391 | 235.0557 | 181.1028 | 232.9527 | 161.373
2012-10-30 18:54 | 28.5783 | 249.552 | 195.4893 | 247.21 | 175.6386 | -1.6026 | 1.2712 | 32.1453 | 40.8624 | 25.2878 | 2.0074 | 32.0694 | 64.3774 | 15.9753 | 14.3815 | -0.1105 | 105.2398 | 235.1705 | 181.1078 | 232.8285 | 161.2571
2012-10-30 18:54 | 28.5731 | 249.5115 | 195.4527 | 247.2832 | 175.5593 | -1.6014 | 1.2712 | 32.1451 | 40.8622 | 25.2876 | 2.0074 | 32.0691 | 64.3772 | 15.9751 | 14.4171 | -0.1104 | 105.2394 | 235.0944 | 181.0356 | 232.8661 | 161.1422
2012-10-30 18:54 | 28.5738 | 249.5083 | 195.4141 | 247.3348 | 175.5248 | -1.5974 | 1.2712 | 32.145 | 40.8614 | 25.2879 | 2.0074 | 32.069 | 64.3757 | 15.9753 | 14.4342 | -0.1101 | 105.2371 | 235.0741 | 180.9799 | 232.9006 | 161.0906
2012-10-30 18:54 | 28.5753 | 249.5333 | 195.392 | 247.3206 | 175.6083 | -1.6026 | 1.2712 | 32.1449 | 40.8615 | 25.2877 | 2.0074 | 32.0689 | 64.376 | 15.9752 | 14.4724 | -0.1105 | 105.2375 | 235.0609 | 180.9196 | 232.8482 | 161.1359
2012-10-30 18:54 | 28.5783 | 249.4776 | 195.3793 | 247.21 | 175.5489 | -1.6026 | 1.2712 | 32.1449 | 40.8617 | 25.2876 | 2.0074 | 32.069 | 64.3764 | 15.9751 | 14.4142 | -0.1105 | 105.2381 | 235.0634 | 180.9651 | 232.7958 | 161.1347
2012-10-30 18:54 | 28.5761 | 249.5106 | 195.3732 | 247.2784 | 175.5779 | -1.6032 | 1.2712 | 32.1449 | 40.8616 | 25.2876 | 2.0074 | 32.069 | 64.3767 | 15.9751 | 14.4242 | -0.1105 | 105.2383 | 235.0864 | 180.949 | 232.8542 | 161.1537
2012-10-30 18:54 | 28.5753 | 249.502 | 195.2899 | 247.3363 | 175.5303 | -1.6049 | 1.2712 | 32.1447 | 40.861 | 25.2877 | 2.0074 | 32.0688 | 64.3758 | 15.9751 | 14.4316 | -0.1106 | 105.2368 | 235.0704 | 180.8583 | 232.9047 | 161.0987
2012-10-30 18:54 | 28.5708 | 249.5327 | 195.3089 | 247.3201 | 175.4594 | -1.6043 | 1.2712 | 32.1449 | 40.8615 | 25.2877 | 2.0074 | 32.0689 | 64.3758 | 15.9751 | 14.4271 | -0.1106 | 105.2373 | 235.1056 | 180.8818 | 232.893 | 161.0323
2012-10-30 18:54 | 28.5761 | 249.5419 | 195.275 | 247.3214 | 175.5076 | -1.5962 | 1.2712 | 32.1447 | 40.8613 | 25.2876 | 2.0074 | 32.0688 | 64.376 | 15.975 | 14.4446 | -0.11 | 105.2373 | 235.0973 | 180.8304 | 232.8768 | 161.063
2012-10-30 18:54 | 28.5753 | 249.5177 | 195.3213 | 247.2619 | 175.6747 | -1.6078 | 1.2712 | 32.1445 | 40.8605 | 25.2877 | 2.0074 | 32.0686 | 64.3752 | 15.975 | 14.4112 | -0.1108 | 105.2357 | 235.1065 | 180.9101 | 232.8507 | 161.2635
2012-10-30 18:54 | 28.5858 | 249.5242 | 195.3082 | 247.3155 | 175.5915 | -1.5957 | 1.2711 | 32.1442 | 40.8598 | 25.2877 | 2.0074 | 32.0683 | 64.3743 | 15.975 | 14.4217 | -0.11 | 105.2341 | 235.1025 | 180.8865 | 232.8938 | 161.1698
2012-10-30 18:54 | 28.5761 | 249.5341 | 195.326 | 247.2705 | 175.4686 | -1.5968 | 1.2711 | 32.144 | 40.859 | 25.2878 | 2.0074 | 32.0681 | 64.3738 | 15.9748 | 14.3589 | -0.1101 | 105.2328 | 235.1752 | 180.9671 | 232.9116 | 161.1097
2012-10-30 18:54 | 28.5828 | 249.5212 | 195.2817 | 247.3791 | 175.6314 | -1.6026 | 1.2711 | 32.144 | 40.8594 | 25.2876 | 2.0074 | 32.0681 | 64.373 | 15.9751 | 14.3983 | -0.1105 | 105.2324 | 235.1229 | 180.8834 | 232.9808 | 161.2331
2012-10-30 18:54 | 28.5783 | 249.5089 | 195.285 | 247.2923 | 175.7206 | -1.6026 | 1.2711 | 32.1436 | 40.858 | 25.2879 | 2.0074 | 32.0676 | 64.3714 | 15.975 | 14.4183 | -0.1105 | 105.2294 | 235.0906 | 180.8667 | 232.874 | 161.3023
2012-10-30 18:54 | 28.5881 | 249.5774 | 195.3066 | 247.2864 | 175.7148 | -1.6043 | 1.2711 | 32.1436 | 40.8577 | 25.2879 | 2.0074 | 32.0676 | 64.3719 | 15.9748 | 14.3832 | -0.1106 | 105.2296 | 235.1942 | 180.9234 | 232.9032 | 161.3316
2012-10-30 18:54 | 28.5911 | 249.5491 | 195.29 | 247.3521 | 175.5304 | -1.5962 | 1.2711 | 32.1431 | 40.8574 | 25.2875 | 2.0073 | 32.0671 | 64.3696 | 15.9749 | 14.4229 | -0.11 | 105.227 | 235.1262 | 180.8671 | 232.9292 | 161.1075
2012-10-30 18:54 | 28.5881 | 249.4991 | 195.287 | 247.4118 | 175.504 | -1.598 | 1.2711 | 32.1429 | 40.8562 | 25.2879 | 2.0073 | 32.0671 | 64.3697 | 15.9748 | 14.3954 | -0.1102 | 105.226 | 235.1037 | 180.8916 | 233.0164 | 161.1086
2012-10-30 18:54 | 28.5926 | 249.5624 | 195.2875 | 247.3223 | 175.4851 | -1.602 | 1.2711 | 32.1427 | 40.856 | 25.2877 | 2.0073 | 32.0668 | 64.3686 | 15.9749 | 14.3754 | -0.1104 | 105.2246 | 235.187 | 180.9121 | 232.9469 | 161.1097
2012-10-30 18:55 | 28.5903 | 249.5875 | 195.2813 | 247.3239 | 175.4828 | -1.598 | 1.2711 | 32.1426 | 40.8555 | 25.2878 | 2.0073 | 32.0666 | 64.3679 | 15.9749 | 14.3731 | -0.1102 | 105.2234 | 235.2144 | 180.9082 | 232.9508 | 161.1097
2012-10-30 18:55 | 28.5963 | 249.5974 | 195.2205 | 247.3652 | 175.4342 | -1.6014 | 1.2711 | 32.1427 | 40.8563 | 25.2875 | 2.0073 | 32.0667 | 64.3685 | 15.9748 | 14.3302 | -0.1104 | 105.2248 | 235.2672 | 180.8903 | 233.035 | 161.104
2012-10-30 18:55 | 28.5978 | 249.591 | 195.2613 | 247.3627 | 175.5176 | -1.5985 | 1.2711 | 32.1422 | 40.8546 | 25.2878 | 2.0073 | 32.0662 | 64.3668 | 15.9747 | 14.4174 | -0.1102 | 105.2214 | 235.1736 | 180.8439 | 232.9453 | 161.1002
2012-10-30 18:55 | 28.5918 | 249.5811 | 195.2356 | 247.3176 | 175.6326 | -1.6026 | 1.2711 | 32.1425 | 40.855 | 25.288 | 2.0073 | 32.0665 | 64.3681 | 15.9747 | 14.3665 | -0.1105 | 105.2232 | 235.2146 | 180.8691 | 232.9511 | 161.2661
2012-10-30 18:55 | 28.5963 | 249.5739 | 195.2441 | 247.2829 | 175.5864 | -1.6032 | 1.271 | 32.1419 | 40.8539 | 25.2877 | 2.0073 | 32.066 | 64.3661 | 15.9747 | 14.4077 | -0.1105 | 105.22 | 235.1662 | 180.8364 | 232.8752 | 161.1787
2012-10-30 18:55 | 28.5956 | 249.5967 | 195.267 | 247.3409 | 175.5935 | -1.6014 | 1.271 | 32.1415 | 40.8523 | 25.2881 | 2.0073 | 32.0656 | 64.3643 | 15.9747 | 14.3907 | -0.1104 | 105.2166 | 235.206 | 180.8763 | 232.9502 | 161.2028
2012-10-30 18:55 | 28.5956 | 249.581 | 195.2787 | 247.4663 | 175.652 | -1.6043 | 1.271 | 32.1408 | 40.8509 | 25.2879 | 2.0072 | 32.0649 | 64.3618 | 15.9746 | 14.4396 | -0.1106 | 105.2127 | 235.1414 | 180.8391 | 233.0267 | 161.2124
2012-10-30 18:55 | 28.5933 | 249.5787 | 195.2843 | 247.3778 | 175.8215 | -1.5997 | 1.271 | 32.1403 | 40.8499 | 25.2876 | 2.0072 | 32.0644 | 64.3597 | 15.9747 | 14.4373 | -0.1103 | 105.2096 | 235.1414 | 180.847 | 232.9405 | 161.3842
2012-10-30 18:55 | 28.5978 | 249.6341 | 195.3359 | 247.2961 | 175.7947 | -1.5985 | 1.271 | 32.1401 | 40.8494 | 25.2877 | 2.0072 | 32.0641 | 64.3588 | 15.9746 | 14.3725 | -0.1102 | 105.2082 | 235.2616 | 180.9634 | 232.9236 | 161.4222
2012-10-30 18:55 | 28.5956 | 249.6163 | 195.322 | 247.3801 | 175.5193 | -1.6009 | 1.271 | 32.1399 | 40.8489 | 25.2877 | 2.0072 | 32.064 | 64.3579 | 15.9747 | 14.3621 | -0.1104 | 105.2068 | 235.2542 | 180.9599 | 233.018 | 161.1572
2012-10-30 18:55 | 28.5941 | 249.5873 | 195.3008 | 247.4805 | 175.3539 | -1.5985 | 1.271 | 32.1395 | 40.8479 | 25.2876 | 2.0072 | 32.0636 | 64.3573 | 15.9745 | 14.3933 | -0.1102 | 105.2052 | 235.194 | 180.9075 | 233.0872 | 160.9606
2012-10-30 18:55 | 28.5896 | 249.5828 | 195.3042 | 247.3232 | 175.5406 | -1.5968 | 1.2709 | 32.1388 | 40.8466 | 25.2874 | 2.0071 | 32.063 | 64.3547 | 15.9746 | 14.3439 | -0.1101 | 105.2012 | 235.2389 | 180.9603 | 232.9793 | 161.1967
2012-10-30 18:55 | 28.5956 | 249.6163 | 195.2591 | 247.3723 | 175.4842 | -1.6026 | 1.2709 | 32.1383 | 40.8449 | 25.2877 | 2.0071 | 32.0625 | 64.3526 | 15.9745 | 14.4029 | -0.1105 | 105.1976 | 235.2134 | 180.8562 | 232.9694 | 161.0813
2012-10-30 18:55 | 28.5963 | 249.5895 | 195.2559 | 247.4239 | 175.4264 | -1.6037 | 1.2709 | 32.1384 | 40.8452 | 25.2876 | 2.0071 | 32.0624 | 64.3525 | 15.9745 | 14.4281 | -0.1106 | 105.1977 | 235.1614 | 180.8278 | 232.9958 | 160.9983
2012-10-30 18:55 | 28.5888 | 249.5781 | 195.2837 | 247.4125 | 175.4969 | -1.5991 | 1.2709 | 32.138 | 40.8439 | 25.2878 | 2.0071 | 32.0621 | 64.3513 | 15.9744 | 14.4084 | -0.1102 | 105.1952 | 235.1697 | 180.8753 | 233.0041 | 161.0885
2012-10-30 18:55 | 28.5903 | 249.5875 | 195.2656 | 247.3435 | 175.557 | -1.5997 | 1.2701 | 32.1165 | 40.7898 | 25.2875 | 2.0052 | 32.0325 | 64.2317 | 15.9746 | 14.4017 | -0.1103 | 105.0215 | 235.1858 | 180.8639 | 232.9418 | 161.1553
2012-10-30 18:55 | 28.5873 | 249.6001 | 195.2469 | 247.3679 | 175.433 | -1.5997 | 1.2646 | 31.9764 | 40.4371 | 25.286 | 1.997 | 31.9018 | 63.7094 | 15.9745 | 14.4191 | -0.1103 | 104.1465 | 235.181 | 180.8278 | 232.9488 | 161.0139
2012-10-30 18:55 | 28.5911 | 249.5648 | 195.2192 | 247.4305 | 175.4016 | -1.5962 | 1.2644 | 31.9709 | 40.4229 | 25.2862 | 1.9966 | 31.8953 | 63.6831 | 15.9746 | 14.4433 | -0.11 | 104.106 | 235.1215 | 180.7759 | 232.9872 | 160.9583
2012-10-30 18:55 | 28.5866 | 249.5446 | 195.2462 | 247.4456 | 175.4284 | -1.6014 | 1.2639 | 31.9585 | 40.3923 | 25.2857 | 1.9959 | 31.8833 | 63.6345 | 15.9747 | 14.4021 | -0.1104 | 104.0269 | 235.1425 | 180.8441 | 233.0435 | 161.0263
2012-10-30 18:55 | 28.5941 | 249.5639 | 195.2222 | 247.3003 | 175.0183 | -1.6009 | 1.2637 | 31.9536 | 40.3798 | 25.2858 | 1.9955 | 31.8781 | 63.6134 | 15.9748 | 14.4055 | -0.1104 | 103.9931 | 235.1584 | 180.8167 | 232.8948 | 160.6128
2012-10-30 18:55 | 28.6001 | 249.519 | 195.1614 | 247.3611 | 174.7941 | -1.5985 | 1.2636 | 31.9505 | 40.3716 | 25.2859 | 1.9953 | 31.875 | 63.6007 | 15.9749 | 14.4319 | -0.1102 | 103.9723 | 235.0871 | 180.7295 | 232.9292 | 160.3622
2012-10-30 18:55 | 28.5926 | 249.4958 | 195.0871 | 247.432 | 174.7164 | -1.5985 | 1.2635 | 31.9486 | 40.3671 | 25.2857 | 1.9952 | 31.873 | 63.5932 | 15.9748 | 14.3714 | -0.1102 | 103.9602 | 235.1244 | 180.7157 | 233.0606 | 160.345
2012-10-30 18:55 | 28.5911 | 249.463 | 195.0385 | 247.2738 | 174.9295 | -1.598 | 1.2635 | 31.9478 | 40.3655 | 25.2855 | 1.9951 | 31.8723 | 63.5893 | 15.9751 | 14.4147 | -0.1102 | 103.9548 | 235.0483 | 180.6238 | 232.8591 | 160.5148
2012-10-30 18:55 | 28.5858 | 249.5008 | 194.9468 | 247.3272 | 174.8149 | -1.6032 | 1.2635 | 31.949 | 40.3684 | 25.2857 | 1.9952 | 31.8737 | 63.5947 | 15.9751 | 14.3686 | -0.1105 | 103.9631 | 235.1322 | 180.5782 | 232.9586 | 160.4463
2012-10-30 18:55 | 28.5888 | 249.4607 | 194.9498 | 247.3694 | 175.0832 | -1.6014 | 1.2637 | 31.9525 | 40.3768 | 25.2858 | 1.9954 | 31.8771 | 63.6084 | 15.975 | 14.4247 | -0.1104 | 103.9852 | 235.036 | 180.5251 | 232.9447 | 160.6585
2012-10-30 18:55 | 28.5813 | 249.4767 | 194.958 | 247.2326 | 174.9392 | -1.6009 | 1.2637 | 31.9546 | 40.3825 | 25.2856 | 1.9956 | 31.8793 | 63.6172 | 15.9751 | 14.3968 | -0.1104 | 103.9997 | 235.0799 | 180.5612 | 232.8358 | 160.5424
2012-10-30 18:56 | 28.5918 | 249.4559 | 194.9645 | 247.4351 | 174.899 | -1.598 | 1.2656 | 32.002 | 40.5019 | 25.2859 | 1.9985 | 31.9267 | 63.8067 | 15.9751 | 14.4318 | -0.1102 | 104.3086 | 235.0241 | 180.5327 | 233.0033 | 160.4672
2012-10-30 18:56 | 28.5881 | 249.4561 | 194.9451 | 247.353 | 175.0279 | -1.602 | 1.2688 | 32.0838 | 40.7074 | 25.287 | 2.0036 | 32.008 | 64.1321 | 15.975 | 14.4036 | -0.1104 | 104.8396 | 235.0525 | 180.5415 | 232.9494 | 160.6243
2012-10-30 18:56 | 28.5806 | 249.4721 | 195.0241 | 247.2515 | 175.2116 | -1.5974 | 1.2693 | 32.0975 | 40.7416 | 25.2873 | 2.0045 | 32.0216 | 64.1869 | 15.9749 | 14.3879 | -0.1101 | 104.9286 | 235.0842 | 180.6362 | 232.8636 | 160.8237
2012-10-30 18:56 | 28.5873 | 249.4748 | 194.9325 | 247.3405 | 175.1364 | -1.5968 | 1.2698 | 32.1094 | 40.7723 | 25.2871 | 2.0052 | 32.0334 | 64.2347 | 15.9749 | 14.4273 | -0.1101 | 105.0069 | 235.0475 | 180.5052 | 232.9132 | 160.7091
2012-10-30 18:56 | 28.5791 | 249.4158 | 195.0147 | 247.2578 | 175.5145 | -1.598 | 1.2702 | 32.1196 | 40.7979 | 25.2874 | 2.0059 | 32.0437 | 64.2759 | 15.9749 | 14.4027 | -0.1102 | 105.0738 | 235.0131 | 180.612 | 232.8551 | 161.1118
2012-10-30 18:56 | 28.5791 | 249.5175 | 195.0461 | 247.2187 | 175.5965 | -1.598 | 1.2705 | 32.1272 | 40.8171 | 25.2873 | 2.0063 | 32.0514 | 64.3063 | 15.975 | 14.4191 | -0.1102 | 105.1234 | 235.0984 | 180.627 | 232.7996 | 161.1774
2012-10-30 18:56 | 28.5761 | 249.4597 | 195.0746 | 247.3175 | 175.6872 | -1.5991 | 1.2707 | 32.1321 | 40.8295 | 25.2874 | 2.0067 | 32.0561 | 64.3261 | 15.9748 | 14.3916 | -0.1102 | 105.1556 | 235.0681 | 180.683 | 232.9259 | 161.2956
2012-10-30 18:56 | 28.5791 | 249.4745 | 195.1679 | 247.25 | 175.7799 | -1.6003 | 1.2708 | 32.135 | 40.8366 | 25.2875 | 2.0068 | 32.0589 | 64.3373 | 15.9748 | 14.4027 | -0.1103 | 105.1739 | 235.0718 | 180.7652 | 232.8473 | 161.3772
2012-10-30 18:56 | 28.5806 | 249.5269 | 195.2087 | 247.2202 | 175.6644 | -1.6049 | 1.2709 | 32.1375 | 40.843 | 25.2876 | 2.007 | 32.0615 | 64.348 | 15.9747 | 14.4002 | -0.1106 | 105.191 | 235.1267 | 180.8085 | 232.82 | 161.2642
2012-10-30 18:56 | 28.5708 | 249.4583 | 195.1596 | 247.3397 | 175.6155 | -1.6003 | 1.2709 | 32.1393 | 40.8472 | 25.2878 | 2.0071 | 32.0634 | 64.3549 | 15.9748 | 14.4108 | -0.1103 | 105.2021 | 235.0475 | 180.7488 | 232.9289 | 161.2047
2012-10-30 18:56 | 28.5753 | 249.4315 | 195.2663 | 247.211 | 175.8152 | -1.5974 | 1.271 | 32.1407 | 40.8507 | 25.2878 | 2.0072 | 32.0646 | 64.3604 | 15.9747 | 14.35 | -0.1101 | 105.2111 | 235.0815 | 180.9163 | 232.861 | 161.4652
2012-10-30 18:56 | 28.5761 | 249.5028 | 195.2278 | 247.2078 | 175.6559 | -1.6014 | 1.271 | 32.1415 | 40.8527 | 25.2878 | 2.0073 | 32.0655 | 64.3638 | 15.9748 | 14.3957 | -0.1104 | 105.2164 | 235.1071 | 180.8321 | 232.8121 | 161.2602
2012-10-30 18:56 | 28.5649 | 249.472 | 195.2127 | 247.3102 | 175.5823 | -1.5985 | 1.271 | 32.1421 | 40.854 | 25.288 | 2.0073 | 32.0662 | 64.367 | 15.9747 | 14.36 | -0.1102 | 105.221 | 235.112 | 180.8527 | 232.9502 | 161.2223
2012-10-30 18:56 | 28.5701 | 249.4655 | 195.265 | 247.2567 | 175.7514 | -1.6003 | 1.2711 | 32.1427 | 40.8561 | 25.2876 | 2.0073 | 32.0668 | 64.3691 | 15.9747 | 14.4101 | -0.1103 | 105.2252 | 235.0554 | 180.8549 | 232.8466 | 161.3413
2012-10-30 18:56 | 28.5716 | 249.4513 | 195.3215 | 247.1877 | 175.6124 | -1.6032 | 1.2711 | 32.1431 | 40.857 | 25.2877 | 2.0074 | 32.067 | 64.3704 | 15.9746 | 14.3871 | -0.1105 | 105.2273 | 235.0642 | 180.9344 | 232.8006 | 161.2253
2012-10-30 18:56 | 28.5731 | 249.4724 | 195.268 | 247.338 | 175.6256 | -1.6003 | 1.2711 | 32.1433 | 40.8576 | 25.2876 | 2.0074 | 32.0674 | 64.3718 | 15.9746 | 14.4049 | -0.1103 | 105.2295 | 235.0675 | 180.8631 | 232.9331 | 161.2207
2012-10-30 18:56 | 28.5821 | 249.5049 | 195.2888 | 247.2334 | 175.6893 | -1.6014 | 1.2711 | 32.1437 | 40.8584 | 25.2877 | 2.0074 | 32.0677 | 64.3731 | 15.9746 | 14.4425 | -0.1104 | 105.2315 | 235.0624 | 180.8463 | 232.7909 | 161.2468
2012-10-30 18:56 | 28.5708 | 249.4897 | 195.2697 | 247.1751 | 175.5804 | -1.5991 | 1.2711 | 32.144 | 40.8593 | 25.2877 | 2.0074 | 32.068 | 64.3749 | 15.9745 | 14.3659 | -0.1102 | 105.2342 | 235.1238 | 180.9038 | 232.8092 | 161.2145
2012-10-30 18:56 | 28.5731 | 249.4646 | 195.2798 | 247.2989 | 175.4578 | -1.6049 | 1.2711 | 32.1441 | 40.8597 | 25.2876 | 2.0075 | 32.0682 | 64.3757 | 15.9745 | 14.3927 | -0.1106 | 105.2353 | 235.0719 | 180.8871 | 232.9062 | 161.0651
2012-10-30 18:56 | 28.5783 | 249.4815 | 195.3086 | 247.2806 | 175.7284 | -1.5974 | 1.2711 | 32.1444 | 40.8599 | 25.2879 | 2.0075 | 32.0683 | 64.3757 | 15.9746 | 14.3693 | -0.1101 | 105.2356 | 235.1122 | 180.9393 | 232.9113 | 161.3591
2012-10-30 18:56 | 28.5761 | 249.5067 | 195.3103 | 247.2235 | 175.4959 | -1.6003 | 1.2711 | 32.1446 | 40.8603 | 25.2879 | 2.0075 | 32.0686 | 64.377 | 15.9745 | 14.3875 | -0.1103 | 105.2373 | 235.1192 | 180.9228 | 232.836 | 161.1084
2012-10-30 18:56 | 28.5783 | 249.4737 | 195.2889 | 247.3197 | 175.4162 | -1.6009 | 1.2712 | 32.1445 | 40.8606 | 25.2876 | 2.0075 | 32.0685 | 64.3762 | 15.9747 | 14.3938 | -0.1104 | 105.2368 | 235.0799 | 180.8951 | 232.9259 | 161.0224
2012-10-30 18:56 | 28.5723 | 249.5068 | 195.2515 | 247.2001 | 175.418 | -1.6003 | 1.2711 | 32.1447 | 40.8604 | 25.2881 | 2.0075 | 32.0687 | 64.3775 | 15.9745 | 14.3878 | -0.1103 | 105.2379 | 235.119 | 180.8637 | 232.8123 | 161.0302
2012-10-30 18:56 | 28.5731 | 249.5272 | 195.2248 | 247.2401 | 175.3329 | -1.6032 | 1.2711 | 32.1444 | 40.8604 | 25.2877 | 2.0075 | 32.0684 | 64.3769 | 15.9744 | 14.3763 | -0.1105 | 105.2372 | 235.1509 | 180.8485 | 232.8638 | 160.9566
2012-10-30 18:56 | 28.5671 | 249.4781 | 195.207 | 247.3124 | 175.4167 | -1.5997 | 1.2712 | 32.1447 | 40.8607 | 25.2878 | 2.0075 | 32.0687 | 64.3774 | 15.9745 | 14.3785 | -0.1103 | 105.2382 | 235.0996 | 180.8285 | 232.9339 | 161.0382
2012-10-30 18:56 | 28.5783 | 249.5011 | 195.2064 | 247.21 | 175.5489 | -1.5991 | 1.2711 | 32.1445 | 40.8603 | 25.2879 | 2.0075 | 32.0685 | 64.3771 | 15.9745 | 14.4101 | -0.1102 | 105.2373 | 235.091 | 180.7963 | 232.7999 | 161.1388
2012-10-30 18:56 | 28.5806 | 249.5034 | 195.2402 | 247.2985 | 175.5473 | -1.5968 | 1.2711 | 32.1446 | 40.8606 | 25.2878 | 2.0075 | 32.0686 | 64.3774 | 15.9745 | 14.4287 | -0.1101 | 105.2379 | 235.0747 | 180.8115 | 232.8698 | 161.1186
2012-10-30 18:57 | 28.5731 | 249.4724 | 195.2405 | 247.3772 | 175.6022 | -1.5957 | 1.2711 | 32.1443 | 40.8599 | 25.2878 | 2.0075 | 32.0683 | 64.3765 | 15.9744 | 14.3559 | -0.11 | 105.2364 | 235.1165 | 180.8846 | 233.0213 | 161.2463
2012-10-30 18:57 | 28.5821 | 249.501 | 195.2456 | 247.2138 | 175.58 | -1.6014 | 1.2712 | 32.1443 | 40.8603 | 25.2875 | 2.0075 | 32.0683 | 64.3764 | 15.9744 | 14.3935 | -0.1104 | 105.2367 | 235.1075 | 180.8521 | 232.8203 | 161.1865
2012-10-30 18:57 | 28.5783 | 249.5167 | 195.23 | 247.3785 | 175.5645 | -1.6026 | 1.2711 | 32.1441 | 40.8597 | 25.2876 | 2.0075 | 32.0681 | 64.3762 | 15.9743 | 14.3897 | -0.1105 | 105.2359 | 235.127 | 180.8403 | 232.9888 | 161.1748
2012-10-30 18:57 | 28.5821 | 249.544 | 195.2613 | 247.2413 | 175.5371 | -1.6003 | 1.2711 | 32.1437 | 40.8586 | 25.2877 | 2.0074 | 32.0678 | 64.3744 | 15.9744 | 14.4017 | -0.1103 | 105.2329 | 235.1423 | 180.8596 | 232.8396 | 161.1354
2012-10-30 18:57 | 28.5806 | 249.5269 | 195.2598 | 247.2672 | 175.6058 | -1.6009 | 1.2711 | 32.1435 | 40.8576 | 25.2879 | 2.0074 | 32.0675 | 64.3734 | 15.9744 | 14.4002 | -0.1104 | 105.231 | 235.1267 | 180.8596 | 232.867 | 161.2056
2012-10-30 18:57 | 28.5866 | 249.5524 | 195.2658 | 247.3359 | 175.4362 | -1.5968 | 1.2711 | 32.1435 | 40.8581 | 25.2877 | 2.0074 | 32.0674 | 64.3734 | 15.9743 | 14.4225 | -0.1101 | 105.2315 | 235.1299 | 180.8433 | 232.9134 | 161.0137
2012-10-30 18:57 | 28.5843 | 249.4718 | 195.2596 | 247.3257 | 175.5549 | -1.5985 | 1.2711 | 32.1432 | 40.8573 | 25.2877 | 2.0074 | 32.0673 | 64.3726 | 15.9743 | 14.4039 | -0.1102 | 105.2299 | 235.0679 | 180.8557 | 232.9218 | 161.151
2012-10-30 18:57 | 28.5836 | 249.5494 | 195.2392 | 247.2428 | 175.4137 | -1.5974 | 1.2711 | 32.1429 | 40.857 | 25.2875 | 2.0074 | 32.0669 | 64.3716 | 15.9742 | 14.4113 | -0.1101 | 105.2286 | 235.1381 | 180.8279 | 232.8315 | 161.0024
2012-10-30 18:57 | 28.5843 | 249.5306 | 195.2439 | 247.4158 | 175.2934 | -1.5991 | 1.2711 | 32.1426 | 40.8557 | 25.2878 | 2.0074 | 32.0668 | 64.3709 | 15.9743 | 14.4243 | -0.1102 | 105.2266 | 235.1063 | 180.8196 | 232.9915 | 160.8691
2012-10-30 18:57 | 28.5873 | 249.514 | 195.1958 | 247.3013 | 175.3354 | -1.5991 | 1.2711 | 32.1426 | 40.8561 | 25.2875 | 2.0074 | 32.0666 | 64.3703 | 15.9743 | 14.4517 | -0.1102 | 105.2264 | 235.0623 | 180.7441 | 232.8496 | 160.8837
2012-10-30 18:57 | 28.5821 | 249.5166 | 195.1827 | 247.2726 | 175.4044 | -1.6043 | 1.2711 | 32.1425 | 40.8556 | 25.2876 | 2.0074 | 32.0665 | 64.3704 | 15.9742 | 14.418 | -0.1106 | 105.2259 | 235.0986 | 180.7647 | 232.8546 | 160.9864
2012-10-30 18:57 | 28.5888 | 249.5429 | 195.1973 | 247.4164 | 175.4189 | -1.6055 | 1.2711 | 32.1422 | 40.8545 | 25.2879 | 2.0074 | 32.0662 | 64.3693 | 15.9741 | 14.388 | -0.1107 | 105.2238 | 235.1549 | 180.8093 | 233.0284 | 161.0309
2012-10-30 18:57 | 28.5806 | 249.5504 | 195.1734 | 247.1888 | 175.5473 | -1.6043 | 1.271 | 32.1421 | 40.854 | 25.288 | 2.0074 | 32.0661 | 64.3681 | 15.9743 | 14.4002 | -0.1106 | 105.2221 | 235.1502 | 180.7732 | 232.7886 | 161.1471
2012-10-30 18:57 | 28.5926 | 249.5624 | 195.1775 | 247.3066 | 175.4461 | -1.6043 | 1.2711 | 32.1417 | 40.854 | 25.2874 | 2.0073 | 32.0658 | 64.3669 | 15.9743 | 14.4162 | -0.1106 | 105.2209 | 235.1462 | 180.7613 | 232.8904 | 161.0299
2012-10-30 18:57 | 28.5918 | 249.5224 | 195.1924 | 247.3607 | 175.5233 | -1.5985 | 1.271 | 32.1411 | 40.8517 | 25.2878 | 2.0073 | 32.0652 | 64.3649 | 15.9742 | 14.4114 | -0.1102 | 105.2167 | 235.111 | 180.781 | 232.9493 | 161.1119
2012-10-30 18:57 | 28.5933 | 249.52 | 195.2096 | 247.3308 | 175.681 | -1.6037 | 1.271 | 32.141 | 40.852 | 25.2875 | 2.0073 | 32.0651 | 64.3654 | 15.9739 | 14.3557 | -0.1106 | 105.2173 | 235.1643 | 180.8539 | 232.9751 | 161.3253
2012-10-30 18:57 | 28.5926 | 249.5428 | 195.1932 | 247.2949 | 175.4383 | -1.6078 | 1.271 | 32.1408 | 40.8513 | 25.2876 | 2.0073 | 32.0648 | 64.3633 | 15.9742 | 14.3958 | -0.1108 | 105.2146 | 235.147 | 180.7974 | 232.8991 | 161.0425
2012-10-30 18:57 | 28.5798 | 249.4713 | 195.2276 | 247.3761 | 175.5543 | -1.5997 | 1.271 | 32.1408 | 40.8509 | 25.2878 | 2.0073 | 32.0648 | 64.3635 | 15.9741 | 14.3912 | -0.1103 | 105.2143 | 235.0801 | 180.8364 | 232.9849 | 161.1631
2012-10-30 18:57 | 28.5888 | 249.5429 | 195.2169 | 247.2597 | 175.5047 | -1.5968 | 1.271 | 32.1403 | 40.8507 | 25.2872 | 2.0073 | 32.0644 | 64.3623 | 15.974 | 14.3635 | -0.1101 | 105.2131 | 235.1794 | 180.8534 | 232.8962 | 161.1412
2012-10-30 18:57 | 28.5956 | 249.5184 | 195.2041 | 247.3488 | 175.4686 | -1.6026 | 1.271 | 32.14 | 40.8488 | 25.2878 | 2.0073 | 32.064 | 64.3609 | 15.974 | 14.4233 | -0.1105 | 105.2098 | 235.0951 | 180.7808 | 232.9255 | 161.0453
2012-10-30 18:57 | 28.5866 | 249.4467 | 195.2422 | 247.3868 | 175.5884 | -1.5985 | 1.271 | 32.1396 | 40.8486 | 25.2874 | 2.0072 | 32.0636 | 64.359 | 15.974 | 14.3654 | -0.1102 | 105.2076 | 235.0813 | 180.8768 | 233.0214 | 161.223
2012-10-30 18:57 | 28.5896 | 249.5515 | 195.1784 | 247.2488 | 175.5836 | -1.6009 | 1.271 | 32.1395 | 40.8481 | 25.2875 | 2.0072 | 32.0634 | 64.3587 | 15.974 | 14.3724 | -0.1104 | 105.2068 | 235.1791 | 180.806 | 232.8764 | 161.2112
2012-10-30 18:57 | 28.5821 | 249.544 | 195.222 | 247.4489 | 175.5839 | -1.5997 | 1.2709 | 32.1392 | 40.8471 | 25.2876 | 2.0072 | 32.0631 | 64.3576 | 15.9739 | 14.3813 | -0.1103 | 105.2047 | 235.1627 | 180.8407 | 233.0676 | 161.2026
2012-10-30 18:57 | 28.5873 | 249.5297 | 195.2351 | 247.2621 | 175.7101 | -1.5991 | 1.2709 | 32.1385 | 40.8454 | 25.2876 | 2.0072 | 32.0626 | 64.3553 | 15.974 | 14.3538 | -0.1102 | 105.2007 | 235.1759 | 180.8813 | 232.9083 | 161.3563
2012-10-30 18:57 | 28.5843 | 249.5345 | 195.2753 | 247.2826 | 175.7188 | -1.6026 | 1.2709 | 32.1383 | 40.8449 | 25.2876 | 2.0072 | 32.0623 | 64.3544 | 15.9739 | 14.3835 | -0.1105 | 105.1994 | 235.151 | 180.8918 | 232.8991 | 161.3353
2012-10-30 18:57 | 28.5948 | 249.5332 | 195.2701 | 247.4381 | 175.7215 | -1.6003 | 1.2709 | 32.1375 | 40.8427 | 25.2878 | 2.0071 | 32.0616 | 64.3514 | 15.9739 | 14.3858 | -0.1103 | 105.1941 | 235.1474 | 180.8843 | 233.0523 | 161.3357
2012-10-30 18:58 | 28.5948 | 249.5606 | 195.3605 | 247.3206 | 175.9869 | -1.5997 | 1.2709 | 32.137 | 40.8416 | 25.2876 | 2.0071 | 32.0611 | 64.3494 | 15.9739 | 14.394 | -0.1103 | 105.191 | 235.1666 | 180.9665 | 232.9266 | 161.5929
2012-10-30 18:58 | 28.5836 | 249.5299 | 195.31 | 247.3133 | 175.6439 | -1.6049 | 1.2708 | 32.1363 | 40.8401 | 25.2874 | 2.0071 | 32.0604 | 64.3471 | 15.9738 | 14.3664 | -0.1106 | 105.1872 | 235.1635 | 180.9436 | 232.9469 | 161.2775
2012-10-30 18:58 | 28.5918 | 249.542 | 195.326 | 247.4468 | 175.6014 | -1.5985 | 1.2708 | 32.1362 | 40.8398 | 25.2875 | 2.007 | 32.0603 | 64.3462 | 15.9739 | 14.3991 | -0.1102 | 105.186 | 235.1429 | 180.9269 | 233.0477 | 161.2023
2012-10-30 18:58 | 28.5911 | 249.5648 | 195.3292 | 247.262 | 175.6749 | -1.6072 | 1.2708 | 32.1363 | 40.8399 | 25.2875 | 2.007 | 32.0603 | 64.3465 | 15.9739 | 14.3617 | -0.1108 | 105.1864 | 235.2031 | 180.9675 | 232.9003 | 161.3132
2012-10-30 18:58 | 28.5941 | 249.556 | 195.3205 | 247.3551 | 175.6193 | -1.6032 | 1.2708 | 32.1361 | 40.8397 | 25.2873 | 2.007 | 32.0601 | 64.3457 | 15.9739 | 14.3973 | -0.1105 | 105.1854 | 235.1587 | 180.9232 | 232.9578 | 161.222
2012-10-30 18:58 | 28.5896 | 249.5711 | 195.3199 | 247.4094 | 175.646 | -1.6043 | 1.2708 | 32.1363 | 40.8395 | 25.2878 | 2.0071 | 32.0603 | 64.3468 | 15.9738 | 14.401 | -0.1106 | 105.1863 | 235.1701 | 180.9189 | 233.0084 | 161.245
2012-10-30 18:58 | 28.5948 | 249.5606 | 195.329 | 247.25 | 175.6512 | -1.5933 | 1.2708 | 32.136 | 40.8389 | 25.2877 | 2.007 | 32.0601 | 64.3458 | 15.9739 | 14.3899 | -0.1098 | 105.1846 | 235.1707 | 180.9391 | 232.8601 | 161.2613
2012-10-30 18:58 | 28.5888 | 249.5703 | 195.3505 | 247.4713 | 175.3174 | -1.6003 | 1.2708 | 32.1355 | 40.8381 | 25.2875 | 2.007 | 32.0596 | 64.3435 | 15.9739 | 14.4165 | -0.1103 | 105.1817 | 235.1538 | 180.934 | 233.0548 | 160.9009
2012-10-30 18:58 | 28.5948 | 249.5332 | 195.2937 | 247.3558 | 175.5068 | -1.6009 | 1.2647 | 31.9797 | 40.4444 | 25.2865 | 1.9968 | 31.8973 | 63.6935 | 15.9739 | 14.4388 | -0.1104 | 104.1379 | 235.0944 | 180.8549 | 232.917 | 161.068
2012-10-30 18:58 | 28.5956 | 249.5693 | 195.2905 | 247.3253 | 175.4335 | -1.598 | 1.2645 | 31.9739 | 40.4308 | 25.2859 | 1.9969 | 31.8987 | 63.6987 | 15.974 | 14.4274 | -0.1102 | 104.1295 | 235.1419 | 180.8631 | 232.8979 | 161.0061
2012-10-30 18:58 | 28.5933 | 249.4965 | 195.2018 | 247.4327 | 175.4351 | -1.602 | 1.264 | 31.9613 | 40.3989 | 25.2859 | 1.9961 | 31.8858 | 63.6479 | 15.9739 | 14.3639 | -0.1104 | 104.0468 | 235.1326 | 180.8379 | 233.0688 | 161.0712
2012-10-30 18:58 | 28.5978 | 249.4893 | 195.2377 | 247.3353 | 175.6113 | -1.6049 | 1.2637 | 31.9544 | 40.3819 | 25.2856 | 1.9956 | 31.8788 | 63.6188 | 15.9742 | 14.4214 | -0.1106 | 104.0007 | 235.0679 | 180.8163 | 232.9139 | 161.1899
2012-10-30 18:58 | 28.5963 | 249.4604 | 195.2284 | 247.2829 | 175.3522 | -1.6003 | 1.2636 | 31.95 | 40.3708 | 25.2857 | 1.9954 | 31.8746 | 63.6019 | 15.9742 | 14.3832 | -0.1103 | 103.9727 | 235.0772 | 180.8452 | 232.8997 | 160.969
2012-10-30 18:58 | 28.5933 | 249.4143 | 195.1743 | 247.4287 | 175.5014 | -1.5974 | 1.2634 | 31.9471 | 40.3635 | 25.2856 | 1.9952 | 31.8717 | 63.5901 | 15.9742 | 14.3761 | -0.1101 | 103.9536 | 235.0382 | 180.7982 | 233.0526 | 161.1253
2012-10-30 18:58 | 28.5993 | 249.436 | 195.2196 | 247.3329 | 175.4255 | -1.5991 | 1.2635 | 31.9477 | 40.3646 | 25.2858 | 1.9952 | 31.8723 | 63.5924 | 15.9743 | 14.3781 | -0.1102 | 103.957 | 235.0579 | 180.8415 | 232.9548 | 161.0474
2012-10-30 18:58 | 28.5978 | 249.4462 | 195.222 | 247.2687 | 175.0727 | -1.5968 | 1.2636 | 31.9499 | 40.3705 | 25.2856 | 1.9954 | 31.8744 | 63.6014 | 15.9742 | 14.3847 | -0.1101 | 103.972 | 235.0615 | 180.8373 | 232.884 | 160.688
2012-10-30 18:58 | 28.5896 | 249.4145 | 195.1391 | 247.429 | 175.1816 | -1.6026 | 1.2637 | 31.9526 | 40.3772 | 25.2857 | 1.9955 | 31.8771 | 63.6116 | 15.9743 | 14.4051 | -0.1105 | 103.9888 | 235.0094 | 180.734 | 233.0239 | 160.7765
2012-10-30 18:58 | 28.5978 | 249.4305 | 195.1749 | 247.2452 | 175.6816 | -1.602 | 1.2637 | 31.9545 | 40.3821 | 25.2857 | 1.9956 | 31.8791 | 63.6194 | 15.9743 | 14.4378 | -0.1104 | 104.0015 | 234.9927 | 180.7371 | 232.8074 | 161.2438
2012-10-30 18:58 | 28.5903 | 249.3956 | 195.1713 | 247.3474 | 175.2799 | -1.6037 | 1.2645 | 31.9733 | 40.4291 | 25.286 | 1.9968 | 31.8975 | 63.6921 | 15.9745 | 14.4017 | -0.1106 | 104.1212 | 234.9939 | 180.7696 | 232.9457 | 160.8782
2012-10-30 18:58 | 28.5956 | 249.3579 | 195.2198 | 247.3801 | 175.5779 | -1.598 | 1.2686 | 32.0781 | 40.6934 | 25.2868 | 2.0034 | 32.0022 | 64.1117 | 15.9744 | 14.4478 | -0.1102 | 104.8051 | 234.9101 | 180.772 | 232.9323 | 161.1301
2012-10-30 18:58 | 28.5933 | 249.4143 | 195.2371 | 247.1741 | 175.6927 | -1.6049 | 1.2691 | 32.0925 | 40.7294 | 25.2871 | 2.0043 | 32.0167 | 64.1701 | 15.9742 | 14.4129 | -0.1106 | 104.8994 | 235.0014 | 180.8242 | 232.7612 | 161.2798
2012-10-30 18:58 | 28.5926 | 249.4175 | 195.2718 | 247.2988 | 175.6803 | -1.5968 | 1.2696 | 32.1058 | 40.763 | 25.2872 | 2.0051 | 32.0297 | 64.2225 | 15.9741 | 14.457 | -0.1101 | 104.9855 | 234.9605 | 180.8148 | 232.8418 | 161.2233
2012-10-30 18:58 | 28.5836 | 249.4242 | 195.2864 | 247.325 | 175.8118 | -1.6026 | 1.2701 | 32.1167 | 40.7902 | 25.2875 | 2.0058 | 32.0406 | 64.2665 | 15.9741 | 14.4276 | -0.1105 | 105.0567 | 234.9966 | 180.8588 | 232.8974 | 161.3842
2012-10-30 18:58 | 28.5843 | 249.4601 | 195.3303 | 247.1925 | 175.7773 | -1.6078 | 1.2704 | 32.1248 | 40.8112 | 25.2873 | 2.0063 | 32.0487 | 64.2987 | 15.9742 | 14.412 | -0.1108 | 105.1099 | 235.0481 | 180.9183 | 232.7805 | 161.3653
2012-10-30 18:58 | 28.5821 | 249.4109 | 195.2927 | 247.2295 | 175.6034 | -1.5974 | 1.2706 | 32.131 | 40.8265 | 25.2875 | 2.0067 | 32.0549 | 64.3241 | 15.974 | 14.4261 | -0.1101 | 105.1506 | 234.9848 | 180.8666 | 232.8034 | 161.1773
2012-10-30 18:58 | 28.5926 | 249.4018 | 195.3425 | 247.3105 | 175.6256 | -1.6014 | 1.2708 | 32.1345 | 40.8352 | 25.2876 | 2.0069 | 32.0584 | 64.3381 | 15.9741 | 14.4652 | -0.1104 | 105.1733 | 234.9366 | 180.8773 | 232.8453 | 161.1604
2012-10-30 18:58 | 28.5941 | 249.4699 | 195.3401 | 247.1749 | 175.7637 | -1.6037 | 1.2709 | 32.1367 | 40.8411 | 25.2874 | 2.007 | 32.0606 | 64.347 | 15.974 | 14.4014 | -0.1106 | 105.1881 | 235.0685 | 180.9387 | 232.7735 | 161.3623
2012-10-30 18:58 | 28.5888 | 249.3902 | 195.3269 | 247.2205 | 175.575 | -1.6003 | 1.2709 | 32.1384 | 40.845 | 25.2878 | 2.0072 | 32.0625 | 64.3545 | 15.9741 | 14.4655 | -0.1103 | 105.1994 | 234.9247 | 180.8614 | 232.755 | 161.1095
2012-10-30 18:59 | 28.5903 | 249.4348 | 195.3284 | 247.3905 | 175.6311 | -1.6009 | 1.271 | 32.1401 | 40.85 | 25.2874 | 2.0073 | 32.0641 | 64.3612 | 15.974 | 14.4507 | -0.1104 | 105.2112 | 234.9841 | 180.8777 | 232.9398 | 161.1804
2012-10-30 18:59 | 28.5881 | 249.4443 | 195.3302 | 247.2551 | 175.7031 | -1.5997 | 1.271 | 32.1409 | 40.8512 | 25.2878 | 2.0073 | 32.0649 | 64.3648 | 15.9739 | 14.424 | -0.1103 | 105.216 | 235.0203 | 180.9062 | 232.8311 | 161.2791
2012-10-30 18:59 | 28.5911 | 249.4278 | 195.341 | 247.1837 | 175.6631 | -1.5991 | 1.271 | 32.1415 | 40.8527 | 25.2879 | 2.0074 | 32.0654 | 64.3667 | 15.974 | 14.4147 | -0.1102 | 105.2194 | 235.0131 | 180.9263 | 232.769 | 161.2484
2012-10-30 18:59 | 28.5836 | 249.4124 | 195.2825 | 247.376 | 175.5073 | -1.5985 | 1.271 | 32.1418 | 40.8533 | 25.2879 | 2.0074 | 32.0658 | 64.3683 | 15.9739 | 14.3991 | -0.1102 | 105.2216 | 235.0133 | 180.8834 | 232.9769 | 161.1082
2012-10-30 18:59 | 28.5911 | 249.5061 | 195.3057 | 247.2346 | 175.4836 | -1.6026 | 1.2711 | 32.1422 | 40.8545 | 25.2879 | 2.0074 | 32.0663 | 64.3702 | 15.9739 | 14.3984 | -0.1105 | 105.2247 | 235.1077 | 180.9073 | 232.8362 | 161.0852
2012-10-30 18:59 | 28.5903 | 249.4504 | 195.3127 | 247.2142 | 175.4477 | -1.5974 | 1.2711 | 32.1426 | 40.8555 | 25.2878 | 2.0074 | 32.0665 | 64.3714 | 15.9739 | 14.4507 | -0.1101 | 105.2269 | 234.9997 | 180.862 | 232.7635 | 160.997
2012-10-30 18:59 | 28.5903 | 249.4661 | 195.2499 | 247.2651 | 175.4555 | -1.5974 | 1.2711 | 32.1425 | 40.8557 | 25.2876 | 2.0074 | 32.0665 | 64.371 | 15.974 | 14.418 | -0.1101 | 105.2266 | 235.0481 | 180.8319 | 232.8471 | 161.0375
2012-10-30 18:59 | 28.5933 | 249.4691 | 195.2804 | 247.3386 | 175.5404 | -1.6037 | 1.2711 | 32.143 | 40.8568 | 25.2876 | 2.0075 | 32.067 | 64.3736 | 15.9738 | 14.4659 | -0.1106 | 105.2304 | 235.0032 | 180.8145 | 232.8727 | 161.0745
2012-10-30 18:59 | 28.5881 | 249.4561 | 195.2909 | 247.2473 | 175.6914 | -1.602 | 1.2711 | 32.1433 | 40.8573 | 25.2878 | 2.0075 | 32.0672 | 64.3747 | 15.9737 | 14.4077 | -0.1104 | 105.232 | 235.0484 | 180.8832 | 232.8396 | 161.2837
2012-10-30 18:59 | 28.5993 | 249.4751 | 195.2824 | 247.1958 | 175.6323 | -1.5991 | 1.2711 | 32.1434 | 40.8576 | 25.2878 | 2.0075 | 32.0674 | 64.3754 | 15.9738 | 14.3985 | -0.1102 | 105.233 | 235.0766 | 180.8839 | 232.7973 | 161.2338
2012-10-30 18:59 | 28.5918 | 249.4519 | 195.2749 | 247.3646 | 175.5741 | -1.5974 | 1.2711 | 32.1435 | 40.8584 | 25.2875 | 2.0075 | 32.0676 | 64.3763 | 15.9738 | 14.4399 | -0.1101 | 105.2347 | 235.012 | 180.835 | 232.9247 | 161.1342
2012-10-30 18:59 | 28.5843 | 249.464 | 195.2792 | 247.2826 | 175.551 | -1.5985 | 1.2711 | 32.1438 | 40.8587 | 25.2878 | 2.0075 | 32.0677 | 64.3769 | 15.9737 | 14.4243 | -0.1102 | 105.2356 | 235.0397 | 180.8549 | 232.8583 | 161.1267
2012-10-30 18:59 | 28.5948 | 249.4354 | 195.3094 | 247.2344 | 175.6473 | -1.5985 | 1.2711 | 32.1437 | 40.8583 | 25.2878 | 2.0075 | 32.0678 | 64.3766 | 15.9738 | 14.4674 | -0.1102 | 105.2349 | 234.968 | 180.842 | 232.767 | 161.1799
2012-10-30 18:59 | 28.5918 | 249.4872 | 195.271 | 247.2588 | 175.5936 | -1.6026 | 1.2711 | 32.1434 | 40.8577 | 25.2877 | 2.0075 | 32.0674 | 64.3754 | 15.9738 | 14.4318 | -0.1105 | 105.2331 | 235.0554 | 180.8392 | 232.827 | 161.1618
2012-10-30 18:59 | 28.6016 | 249.4735 | 195.2926 | 247.3626 | 175.5136 | -1.6014 | 1.2711 | 32.1435 | 40.858 | 25.2877 | 2.0075 | 32.0674 | 64.3759 | 15.9737 | 14.4089 | -0.1104 | 105.2339 | 235.0646 | 180.8837 | 232.9537 | 161.1047
2012-10-30 18:59 | 28.5963 | 249.4878 | 195.3344 | 247.2711 | 175.563 | -1.6003 | 1.2711 | 32.1429 | 40.8563 | 25.2878 | 2.0075 | 32.0669 | 64.3738 | 15.9737 | 14.4363 | -0.1103 | 105.2301 | 235.0515 | 180.8981 | 232.8348 | 161.1267
2012-10-30 18:59 | 28.6001 | 249.4876 | 195.2636 | 247.2671 | 175.5668 | -1.6009 | 1.2711 | 32.1427 | 40.8561 | 25.2877 | 2.0075 | 32.0669 | 64.3736 | 15.9737 | 14.4686 | -0.1104 | 105.2297 | 235.019 | 180.795 | 232.7985 | 161.0982
2012-10-30 18:59 | 28.6016 | 249.4617 | 195.273 | 247.296 | 175.5214 | -1.6003 | 1.2711 | 32.1427 | 40.8559 | 25.2877 | 2.0075 | 32.0667 | 64.3728 | 15.9737 | 14.4212 | -0.1103 | 105.2288 | 235.0405 | 180.8518 | 232.8748 | 161.1002
2012-10-30 18:59 | 28.6068 | 249.4826 | 195.2821 | 247.4305 | 175.5266 | -1.5939 | 1.2711 | 32.1428 | 40.8559 | 25.2879 | 2.0075 | 32.0668 | 64.3734 | 15.9737 | 14.4223 | -0.1099 | 105.2293 | 235.0603 | 180.8598 | 233.0082 | 161.1043
2012-10-30 18:59 | 28.6068 | 249.4787 | 195.2664 | 247.3012 | 175.6437 | -1.602 | 1.2711 | 32.1428 | 40.8563 | 25.2877 | 2.0075 | 32.0668 | 64.3739 | 15.9736 | 14.4631 | -0.1104 | 105.2302 | 235.0156 | 180.8033 | 232.8381 | 161.1806
2012-10-30 18:59 | 28.6098 | 249.5326 | 195.3322 | 247.2729 | 175.5608 | -1.6003 | 1.2711 | 32.1425 | 40.8557 | 25.2876 | 2.0075 | 32.0666 | 64.3731 | 15.9735 | 14.4538 | -0.1103 | 105.2288 | 235.0788 | 180.8784 | 232.8191 | 161.107
2012-10-30 18:59 | 28.6098 | 249.4778 | 195.3126 | 247.3473 | 175.5023 | -1.6009 | 1.2711 | 32.1425 | 40.8552 | 25.2879 | 2.0075 | 32.0665 | 64.3728 | 15.9736 | 14.4457 | -0.1104 | 105.228 | 235.0321 | 180.8669 | 232.9016 | 161.0566
2012-10-30 18:59 | 28.6008 | 249.4492 | 195.2839 | 247.3383 | 175.5753 | -1.602 | 1.2711 | 32.1422 | 40.8548 | 25.2876 | 2.0075 | 32.0661 | 64.3715 | 15.9735 | 14.4489 | -0.1104 | 105.2263 | 235.0003 | 180.835 | 232.8894 | 161.1264
2012-10-30 18:59 | 28.6053 | 249.5046 | 195.2688 | 247.2645 | 175.6344 | -1.5974 | 1.2711 | 32.142 | 40.8541 | 25.2877 | 2.0075 | 32.0659 | 64.3708 | 15.9734 | 14.4167 | -0.1101 | 105.2249 | 235.0879 | 180.8521 | 232.8478 | 161.2177
2012-10-30 18:59 | 28.6113 | 249.5028 | 195.263 | 247.3566 | 175.6014 | -1.5916 | 1.2711 | 32.142 | 40.8545 | 25.2875 | 2.0074 | 32.0659 | 64.3706 | 15.9735 | 14.4676 | -0.1097 | 105.2251 | 235.0352 | 180.7954 | 232.889 | 161.1338
2012-10-30 18:59 | 28.6023 | 249.4703 | 195.3208 | 247.3947 | 175.7056 | -1.6026 | 1.271 | 32.1417 | 40.8529 | 25.288 | 2.0074 | 32.0657 | 64.3699 | 15.9735 | 14.4749 | -0.1105 | 105.2228 | 234.9954 | 180.8459 | 232.9198 | 161.2307
2012-10-30 19:00 | 28.6105 | 249.5137 | 195.2936 | 247.3284 | 175.6513 | -1.6037 | 1.271 | 32.1418 | 40.8535 | 25.2878 | 2.0074 | 32.0658 | 64.3704 | 15.9735 | 14.479 | -0.1106 | 105.2239 | 235.0347 | 180.8146 | 232.8494 | 161.1723
2012-10-30 19:00 | 28.6046 | 249.5196 | 195.331 | 247.2677 | 175.5283 | -1.6026 | 1.271 | 32.1416 | 40.853 | 25.2878 | 2.0074 | 32.0656 | 64.3694 | 15.9734 | 14.465 | -0.1105 | 105.2224 | 235.0546 | 180.866 | 232.8027 | 161.0633
2012-10-30 19:00 | 28.6008 | 249.5314 | 195.2761 | 247.401 | 175.4738 | -1.6072 | 1.271 | 32.1411 | 40.8519 | 25.2876 | 2.0074 | 32.065 | 64.3674 | 15.9734 | 14.4775 | -0.1108 | 105.2194 | 235.0539 | 180.7986 | 232.9235 | 160.9963
2012-10-30 19:00 | 28.5971 | 249.5355 | 195.3156 | 247.4521 | 175.5521 | -1.6009 | 1.271 | 32.1409 | 40.8516 | 25.2875 | 2.0074 | 32.0649 | 64.3669 | 15.9734 | 14.4779 | -0.1104 | 105.2184 | 235.0576 | 180.8377 | 232.9742 | 161.0742
2012-10-30 19:00 | 28.6016 | 249.5087 | 195.2926 | 247.3274 | 175.4629 | -1.6026 | 1.271 | 32.1407 | 40.8515 | 25.2874 | 2.0074 | 32.0647 | 64.3659 | 15.9734 | 14.4783 | -0.1105 | 105.2174 | 235.0304 | 180.8143 | 232.8491 | 160.9846
2012-10-30 19:00 | 28.6046 | 249.547 | 195.276 | 247.3852 | 175.3332 | -1.5945 | 1.271 | 32.1408 | 40.8515 | 25.2875 | 2.0074 | 32.0648 | 64.3665 | 15.9734 | 14.469 | -0.1099 | 105.218 | 235.078 | 180.807 | 232.9162 | 160.8642
2012-10-30 19:00 | 28.6008 | 249.5275 | 195.2525 | 247.4598 | 175.4465 | -1.6003 | 1.271 | 32.1403 | 40.8494 | 25.2879 | 2.0074 | 32.0643 | 64.3644 | 15.9734 | 14.4571 | -0.1103 | 105.2138 | 235.0704 | 180.7954 | 233.0027 | 160.9894
2012-10-30 19:00 | 28.6008 | 249.4923 | 195.225 | 247.3461 | 175.5792 | -1.5985 | 1.271 | 32.1399 | 40.8486 | 25.2878 | 2.0073 | 32.0639 | 64.3634 | 15.9733 | 14.4612 | -0.1102 | 105.212 | 235.0311 | 180.7638 | 232.8849 | 161.118
2012-10-30 19:00 | 28.5986 | 249.5331 | 195.2621 | 247.3518 | 175.4872 | -1.5968 | 1.271 | 32.1397 | 40.8484 | 25.2877 | 2.0073 | 32.0636 | 64.3619 | 15.9733 | 14.4508 | -0.1101 | 105.2103 | 235.0823 | 180.8113 | 232.901 | 161.0364
2012-10-30 19:00 | 28.5948 | 249.5293 | 195.2819 | 247.3754 | 175.5263 | -1.6084 | 1.2709 | 32.1392 | 40.8467 | 25.2879 | 2.0073 | 32.0633 | 64.3611 | 15.9732 | 14.4837 | -0.1109 | 105.2078 | 235.0456 | 180.7982 | 232.8917 | 161.0426
2012-10-30 19:00 | 28.5986 | 249.5644 | 195.2778 | 247.387 | 175.6511 | -1.591 | 1.2709 | 32.1393 | 40.8474 | 25.2877 | 2.0073 | 32.0633 | 64.3612 | 15.9732 | 14.4753 | -0.1097 | 105.2086 | 235.0891 | 180.8025 | 232.9117 | 161.1758
2012-10-30 19:00 | 28.5993 | 249.5143 | 195.2785 | 247.3172 | 175.5933 | -1.5968 | 1.2709 | 32.1387 | 40.8459 | 25.2876 | 2.0073 | 32.0627 | 64.3583 | 15.9733 | 14.4433 | -0.1101 | 105.2042 | 235.071 | 180.8352 | 232.8739 | 161.15
2012-10-30 19:00 | 28.5948 | 249.5685 | 195.2583 | 247.3754 | 175.5341 | -1.6003 | 1.2709 | 32.1384 | 40.8446 | 25.2879 | 2.0073 | 32.0623 | 64.3575 | 15.9731 | 14.4592 | -0.1103 | 105.2021 | 235.1093 | 180.7991 | 232.9162 | 161.0749
2012-10-30 19:00 | 28.5986 | 249.5527 | 195.2542 | 247.4184 | 175.616 | -1.6003 | 1.2709 | 32.1376 | 40.8431 | 25.2876 | 2.0072 | 32.0617 | 64.3546 | 15.9733 | 14.4957 | -0.1103 | 105.1977 | 235.057 | 180.7585 | 232.9227 | 161.1203
2012-10-30 19:00 | 28.5948 | 249.5293 | 195.2701 | 247.3989 | 175.5537 | -1.6003 | 1.2709 | 32.1376 | 40.8432 | 25.2876 | 2.0072 | 32.0616 | 64.3548 | 15.9731 | 14.4715 | -0.1103 | 105.198 | 235.0578 | 180.7986 | 232.9274 | 161.0822
2012-10-30 19:00 | 28.5896 | 249.528 | 195.312 | 247.2801 | 175.4899 | -1.6003 | 1.2709 | 32.1373 | 40.8426 | 25.2875 | 2.0072 | 32.0612 | 64.353 | 15.9732 | 14.4663 | -0.1103 | 105.1956 | 235.0617 | 180.8457 | 232.8138 | 161.0236
2012-10-30 19:00 | 28.5926 | 249.5467 | 195.2718 | 247.381 | 175.5124 | -1.602 | 1.2709 | 32.1368 | 40.8411 | 25.2876 | 2.0072 | 32.0608 | 64.3514 | 15.9731 | 14.457 | -0.1104 | 105.1925 | 235.0897 | 180.8148 | 232.924 | 161.0554
2012-10-30 19:00 | 28.5956 | 249.5106 | 195.2748 | 247.4389 | 175.4335 | -1.5957 | 1.2709 | 32.1366 | 40.8408 | 25.2875 | 2.0072 | 32.0607 | 64.3512 | 15.9731 | 14.3988 | -0.11 | 105.192 | 235.1118 | 180.876 | 233.0401 | 161.0347
2012-10-30 19:00 | 28.5978 | 249.5793 | 195.2377 | 247.398 | 175.2444 | -1.5985 | 1.2708 | 32.1363 | 40.8402 | 25.2875 | 2.0071 | 32.0604 | 64.35 | 15.9731 | 14.4378 | -0.1102 | 105.1901 | 235.1415 | 180.7999 | 232.9602 | 160.8066
2012-10-30 19:00 | 28.5866 | 249.5642 | 195.2265 | 247.3515 | 175.401 | -1.5991 | 1.2708 | 32.1357 | 40.838 | 25.2879 | 2.0071 | 32.0598 | 64.3475 | 15.9731 | 14.451 | -0.1102 | 105.1855 | 235.1132 | 180.7755 | 232.9005 | 160.95
2012-10-30 19:00 | 28.5926 | 249.535 | 195.2168 | 247.334 | 175.3992 | -1.6009 | 1.2708 | 32.1355 | 40.8375 | 25.2878 | 2.0071 | 32.0595 | 64.3467 | 15.973 | 14.4203 | -0.1104 | 105.1842 | 235.1147 | 180.7965 | 232.9137 | 160.9789
2012-10-30 19:00 | 28.5881 | 249.5383 | 195.1966 | 247.4431 | 175.3557 | -1.6026 | 1.2695 | 32.1024 | 40.754 | 25.2874 | 2.0049 | 32.0243 | 64.2054 | 15.9731 | 14.4444 | -0.1105 | 104.9594 | 235.0939 | 180.7522 | 232.9987 | 160.9113
2012-10-30 19:00 | 28.5843 | 249.5345 | 195.1614 | 247.4589 | 175.387 | -1.5997 | 1.2646 | 31.9766 | 40.437 | 25.2863 | 1.9972 | 31.9018 | 63.7148 | 15.9731 | 14.4406 | -0.1103 | 104.1518 | 235.0939 | 180.7208 | 233.0183 | 160.9464
2012-10-30 19:00 | 28.5933 | 249.52 | 195.1861 | 247.4131 | 175.4429 | -1.5997 | 1.2643 | 31.9689 | 40.4177 | 25.2862 | 1.9967 | 31.8934 | 63.6807 | 15.9733 | 14.4741 | -0.1103 | 104.0984 | 235.0459 | 180.712 | 232.939 | 160.9688
2012-10-30 19:00 | 28.5798 | 249.4987 | 195.1726 | 247.38 | 175.5308 | -1.5962 | 1.2637 | 31.9539 | 40.38 | 25.2861 | 1.9957 | 31.8783 | 63.6207 | 15.9732 | 14.4075 | -0.11 | 104.0007 | 235.0912 | 180.7651 | 232.9725 | 161.1233
2012-10-30 19:01 | 28.5888 | 249.4685 | 195.1776 | 247.3146 | 175.5867 | -1.6003 | 1.2635 | 31.9482 | 40.3661 | 25.2858 | 1.9954 | 31.8726 | 63.5972 | 15.9734 | 14.4614 | -0.1103 | 103.9633 | 235.0071 | 180.7162 | 232.8532 | 161.1253
2012-10-30 19:01 | 28.5851 | 249.4531 | 195.1189 | 247.3775 | 175.3176 | -1.6055 | 1.2634 | 31.9459 | 40.3604 | 25.2857 | 1.9952 | 31.8704 | 63.5886 | 15.9733 | 14.4781 | -0.1107 | 103.949 | 234.975 | 180.6408 | 232.8994 | 160.8395
2012-10-30 19:01 | 28.5806 | 249.4212 | 195.1223 | 247.467 | 175.4458 | -1.5991 | 1.2634 | 31.9468 | 40.3623 | 25.2859 | 1.9953 | 31.8713 | 63.592 | 15.9734 | 14.4369 | -0.1102 | 103.9542 | 234.9843 | 180.6854 | 233.0301 | 161.0089
2012-10-30 19:01 | 28.5873 | 249.424 | 195.1172 | 247.3483 | 175.5891 | -1.6026 | 1.2635 | 31.9486 | 40.367 | 25.2858 | 1.9954 | 31.8733 | 63.5992 | 15.9735 | 14.3987 | -0.1105 | 103.9662 | 235.0253 | 180.7185 | 232.9496 | 161.1904
2012-10-30 19:01 | 28.5806 | 249.3938 | 195.1616 | 247.3573 | 175.6839 | -1.6009 | 1.2636 | 31.9504 | 40.372 | 25.2856 | 1.9955 | 31.8749 | 63.6059 | 15.9735 | 14.4165 | -0.1104 | 103.9779 | 234.9773 | 180.7451 | 232.9408 | 161.2674
2012-10-30 19:01 | 28.5836 | 249.4242 | 195.1489 | 247.2859 | 175.5698 | -1.6014 | 1.2636 | 31.9524 | 40.3765 | 25.2859 | 1.9956 | 31.8768 | 63.6134 | 15.9735 | 14.4113 | -0.1104 | 103.99 | 235.0129 | 180.7376 | 232.8746 | 161.1585
2012-10-30 19:01 | 28.5881 | 249.413 | 195.1691 | 247.404 | 175.4104 | -1.5997 | 1.2674 | 32.0489 | 40.6198 | 25.2865 | 2.0018 | 31.9756 | 64.0079 | 15.9736 | 14.4199 | -0.1103 | 104.6277 | 234.9931 | 180.7492 | 232.9841 | 160.9905
2012-10-30 19:01 | 28.5896 | 249.4458 | 195.1745 | 247.3193 | 175.81 | -1.6043 | 1.2689 | 32.0882 | 40.7182 | 25.2873 | 2.0041 | 32.0127 | 64.157 | 15.9735 | 14.4214 | -0.1106 | 104.8752 | 235.0244 | 180.7531 | 232.8979 | 161.3886
2012-10-30 19:01 | 28.5836 | 249.4007 | 195.1921 | 247.3485 | 175.9367 | -1.5933 | 1.2696 | 32.1056 | 40.7622 | 25.2874 | 2.0052 | 32.0299 | 64.2264 | 15.9734 | 14.3868 | -0.1098 | 104.9886 | 235.0139 | 180.8053 | 232.9617 | 161.5499
2012-10-30 19:01 | 28.5836 | 249.4398 | 195.2628 | 247.2193 | 175.9211 | -1.602 | 1.2701 | 32.1181 | 40.7935 | 25.2876 | 2.006 | 32.0422 | 64.276 | 15.9734 | 14.4072 | -0.1104 | 105.0695 | 235.0326 | 180.8556 | 232.8121 | 161.5139
2012-10-30 19:01 | 28.5866 | 249.4507 | 195.313 | 247.2418 | 175.7484 | -1.6003 | 1.2704 | 32.1261 | 40.8139 | 25.2876 | 2.0065 | 32.0502 | 64.3083 | 15.9733 | 14.4306 | -0.1103 | 105.1222 | 235.0201 | 180.8824 | 232.8112 | 161.3178
2012-10-30 19:01 | 28.5971 | 249.422 | 195.3628 | 247.3111 | 175.6184 | -1.6043 | 1.2706 | 32.1305 | 40.8255 | 25.2874 | 2.0067 | 32.0546 | 64.3254 | 15.9734 | 14.3922 | -0.1106 | 105.1509 | 235.0298 | 180.9706 | 232.9189 | 161.2262
2012-10-30 19:01 | 28.5918 | 249.4519 | 195.3575 | 247.2627 | 175.7458 | -1.5939 | 1.2707 | 32.1341 | 40.834 | 25.2877 | 2.007 | 32.0579 | 64.3394 | 15.9733 | 14.3991 | -0.1099 | 105.1734 | 235.0528 | 180.9584 | 232.8636 | 161.3467
2012-10-30 19:01 | 28.5903 | 249.4191 | 195.3206 | 247.3435 | 175.756 | -1.5962 | 1.2708 | 32.137 | 40.8412 | 25.2879 | 2.0072 | 32.0609 | 64.3512 | 15.9733 | 14.418 | -0.11 | 105.1924 | 235.0011 | 180.9026 | 232.9255 | 161.338
2012-10-30 19:01 | 28.5888 | 249.4529 | 195.3348 | 247.2401 | 175.8131 | -1.6009 | 1.2709 | 32.1389 | 40.846 | 25.2878 | 2.0073 | 32.0627 | 64.3592 | 15.9732 | 14.4043 | -0.1104 | 105.2052 | 235.0486 | 180.9305 | 232.8358 | 161.4088
2012-10-30 19:01 | 28.5918 | 249.4754 | 195.3142 | 247.1844 | 175.6756 | -1.5985 | 1.271 | 32.1399 | 40.8485 | 25.2879 | 2.0074 | 32.0638 | 64.3636 | 15.9731 | 14.444 | -0.1102 | 105.2121 | 235.0314 | 180.8702 | 232.7404 | 161.2316
2012-10-30 19:01 | 28.5963 | 249.476 | 195.3305 | 247.2633 | 175.6488 | -1.6043 | 1.271 | 32.1407 | 40.8509 | 25.2877 | 2.0074 | 32.0647 | 64.3667 | 15.9732 | 14.4036 | -0.1106 | 105.2175 | 235.0724 | 180.9269 | 232.8597 | 161.2452
2012-10-30 19:01 | 28.5956 | 249.4831 | 195.2866 | 247.3488 | 175.6911 | -1.6014 | 1.271 | 32.141 | 40.8513 | 25.288 | 2.0074 | 32.065 | 64.3677 | 15.9733 | 14.407 | -0.1104 | 105.219 | 235.0761 | 180.8796 | 232.9418 | 161.2841
2012-10-30 19:01 | 28.5986 | 249.4587 | 195.3289 | 247.3518 | 175.7253 | -1.602 | 1.271 | 32.1416 | 40.8529 | 25.2879 | 2.0075 | 32.0655 | 64.3705 | 15.9731 | 14.4508 | -0.1104 | 105.2234 | 235.0079 | 180.8781 | 232.901 | 161.2745
2012-10-30 19:01 | 28.5948 | 249.5098 | 195.3644 | 247.2735 | 175.8074 | -1.5997 | 1.271 | 32.1418 | 40.8529 | 25.2881 | 2.0075 | 32.0657 | 64.3712 | 15.9731 | 14.3817 | -0.1103 | 105.2241 | 235.1281 | 180.9827 | 232.8918 | 161.4257
2012-10-30 19:01 | 28.6031 | 249.5181 | 195.3805 | 247.2152 | 175.8039 | -1.602 | 1.271 | 32.1417 | 40.853 | 25.288 | 2.0075 | 32.0657 | 64.3708 | 15.9732 | 14.4227 | -0.1104 | 105.2238 | 235.0954 | 180.9578 | 232.7925 | 161.3812
2012-10-30 19:01 | 28.5948 | 249.498 | 195.3565 | 247.2892 | 175.6473 | -1.5991 | 1.2711 | 32.142 | 40.8544 | 25.2875 | 2.0075 | 32.066 | 64.3722 | 15.9731 | 14.3899 | -0.1102 | 105.2267 | 235.1081 | 180.9666 | 232.8993 | 161.2574
2012-10-30 19:01 | 28.5866 | 249.4624 | 195.313 | 247.2889 | 175.483 | -1.6009 | 1.2711 | 32.1422 | 40.8546 | 25.2877 | 2.0075 | 32.0661 | 64.3736 | 15.973 | 14.4062 | -0.1104 | 105.2282 | 235.0562 | 180.9068 | 232.8827 | 161.0768
2012-10-30 19:01 | 28.5918 | 249.5068 | 195.326 | 247.3763 | 175.5389 | -1.6055 | 1.2711 | 32.1426 | 40.8555 | 25.2878 | 2.0075 | 32.0666 | 64.3752 | 15.9731 | 14.4154 | -0.1107 | 105.2307 | 235.0914 | 180.9106 | 232.9609 | 161.1235
2012-10-30 19:02 | 28.5888 | 249.5038 | 195.3191 | 247.2793 | 175.6491 | -1.5997 | 1.2711 | 32.1428 | 40.8556 | 25.2881 | 2.0076 | 32.0667 | 64.3765 | 15.9728 | 14.3798 | -0.1103 | 105.2321 | 235.124 | 180.9393 | 232.8995 | 161.2693
2012-10-30 19:02 | 28.5963 | 249.5113 | 195.3187 | 247.2672 | 175.7581 | -1.6049 | 1.2711 | 32.143 | 40.856 | 25.2881 | 2.0076 | 32.0669 | 64.377 | 15.9729 | 14.4118 | -0.1106 | 105.233 | 235.0995 | 180.9069 | 232.8554 | 161.3463
2012-10-30 19:02 | 28.6001 | 249.5346 | 195.3304 | 247.2984 | 175.7307 | -1.5951 | 1.2711 | 32.1425 | 40.8554 | 25.2878 | 2.0075 | 32.0665 | 64.3748 | 15.973 | 14.4441 | -0.11 | 105.2302 | 235.0905 | 180.8863 | 232.8543 | 161.2866
2012-10-30 19:02 | 28.6023 | 249.5486 | 195.364 | 247.3476 | 175.6002 | -1.598 | 1.2711 | 32.1425 | 40.8555 | 25.2877 | 2.0075 | 32.0664 | 64.3749 | 15.9729 | 14.43 | -0.1102 | 105.2304 | 235.1186 | 180.934 | 232.9176 | 161.1702
2012-10-30 19:02 | 28.6016 | 249.5205 | 195.3122 | 247.394 | 175.5526 | -1.598 | 1.271 | 32.1424 | 40.8546 | 25.2881 | 2.0075 | 32.0663 | 64.3745 | 15.9729 | 14.4334 | -0.1102 | 105.2291 | 235.0871 | 180.8788 | 232.9606 | 161.1192
2012-10-30 19:02 | 28.6008 | 249.5314 | 195.3272 | 247.3344 | 175.6221 | -1.5957 | 1.2711 | 32.1424 | 40.8553 | 25.2877 | 2.0076 | 32.0664 | 64.3752 | 15.9729 | 14.4326 | -0.11 | 105.2305 | 235.0988 | 180.8946 | 232.9018 | 161.1895
2012-10-30 19:02 | 28.6031 | 249.5181 | 195.3098 | 247.3054 | 175.6517 | -1.5968 | 1.271 | 32.1424 | 40.8544 | 25.2882 | 2.0075 | 32.0663 | 64.3746 | 15.9729 | 14.4349 | -0.1101 | 105.2291 | 235.0832 | 180.8749 | 232.8705 | 161.2168
2012-10-30 19:02 | 28.5948 | 249.5802 | 195.2976 | 247.254 | 175.659 | -1.5997 | 1.2711 | 32.1419 | 40.8541 | 25.2877 | 2.0075 | 32.0659 | 64.3726 | 15.9729 | 14.4184 | -0.1103 | 105.2267 | 235.1618 | 180.8792 | 232.8356 | 161.2406
2012-10-30 19:02 | 28.6038 | 249.5383 | 195.3262 | 247.3491 | 175.6329 | -1.6049 | 1.271 | 32.1417 | 40.8527 | 25.2881 | 2.0075 | 32.0657 | 64.372 | 15.9729 | 14.4234 | -0.1106 | 105.2247 | 235.1149 | 180.9028 | 232.9257 | 161.2095
2012-10-30 19:02 | 28.6038 | 249.5736 | 195.338 | 247.3844 | 175.5392 | -1.6037 | 1.271 | 32.1417 | 40.8535 | 25.2877 | 2.0075 | 32.0657 | 64.3726 | 15.9728 | 14.4356 | -0.1106 | 105.2261 | 235.138 | 180.9024 | 232.9488 | 161.1036
2012-10-30 19:02 | 28.5971 | 249.5825 | 195.3392 | 247.366 | 175.4545 | -1.6014 | 1.271 | 32.1413 | 40.8523 | 25.2877 | 2.0075 | 32.0654 | 64.3706 | 15.973 | 14.4044 | -0.1104 | 105.2229 | 235.1781 | 180.9348 | 232.9616 | 161.0501
2012-10-30 19:02 | 28.6023 | 249.5642 | 195.2972 | 247.3281 | 175.6314 | -1.6026 | 1.271 | 32.1412 | 40.8521 | 25.2878 | 2.0075 | 32.0652 | 64.3705 | 15.9728 | 14.4219 | -0.1105 | 105.2226 | 235.1423 | 180.8753 | 232.9062 | 161.2095
2012-10-30 19:02 | 28.6046 | 249.5509 | 195.2838 | 247.346 | 175.7118 | -1.5974 | 1.271 | 32.1411 | 40.8514 | 25.2879 | 2.0075 | 32.065 | 64.3699 | 15.9728 | 14.3915 | -0.1101 | 105.2214 | 235.1594 | 180.8923 | 232.9545 | 161.3203
2012-10-30 19:02 | 28.6068 | 249.4983 | 195.3253 | 247.3913 | 175.5149 | -1.6003 | 1.271 | 32.1412 | 40.852 | 25.2878 | 2.0075 | 32.0652 | 64.3711 | 15.9726 | 14.4141 | -0.1103 | 105.2231 | 235.0842 | 180.9112 | 232.9772 | 161.1008
2012-10-30 19:02 | 28.5963 | 249.5661 | 195.2755 | 247.4396 | 175.5161 | -1.6003 | 1.271 | 32.141 | 40.8514 | 25.2879 | 2.0075 | 32.065 | 64.3699 | 15.9727 | 14.4036 | -0.1103 | 105.2213 | 235.1625 | 180.8719 | 233.036 | 161.1125
2012-10-30 19:02 | 28.6031 | 249.5846 | 195.2745 | 247.3876 | 175.6322 | -1.6014 | 1.271 | 32.1405 | 40.8498 | 25.2879 | 2.0074 | 32.0644 | 64.3676 | 15.9727 | 14.4512 | -0.1104 | 105.2175 | 235.1334 | 180.8233 | 232.9364 | 161.181
2012-10-30 19:02 | 28.5971 | 249.5786 | 195.3431 | 247.3934 | 175.6535 | -1.598 | 1.271 | 32.1402 | 40.8493 | 25.288 | 2.0074 | 32.0642 | 64.3669 | 15.9727 | 14.433 | -0.1102 | 105.2162 | 235.1456 | 180.9101 | 232.9604 | 161.2205
2012-10-30 19:02 | 28.6075 | 249.6164 | 195.2828 | 247.3372 | 175.3868 | -1.6003 | 1.271 | 32.1402 | 40.8496 | 25.2878 | 2.0074 | 32.0642 | 64.3671 | 15.9726 | 14.4107 | -0.1103 | 105.2166 | 235.2057 | 180.8721 | 232.9265 | 160.9761
2012-10-30 19:02 | 28.612 | 249.5974 | 195.2676 | 247.4122 | 175.4577 | -1.6003 | 1.271 | 32.1398 | 40.8492 | 25.2873 | 2.0074 | 32.0638 | 64.3657 | 15.9726 | 14.3989 | -0.1103 | 105.2149 | 235.1985 | 180.8687 | 233.0133 | 161.0588
2012-10-30 19:02 | 28.6038 | 249.5853 | 195.2005 | 247.4902 | 175.5939 | -1.6055 | 1.2709 | 32.1395 | 40.8476 | 25.2878 | 2.0074 | 32.0635 | 64.364 | 15.9727 | 14.4274 | -0.1107 | 105.2116 | 235.1579 | 180.7731 | 233.0628 | 161.1665
2012-10-30 19:02 | 28.6068 | 249.604 | 195.2546 | 247.3874 | 175.6593 | -1.5991 | 1.2709 | 32.1391 | 40.8465 | 25.2879 | 2.0074 | 32.0631 | 64.3633 | 15.9725 | 14.406 | -0.1102 | 105.2098 | 235.198 | 180.8486 | 232.9814 | 161.2533
2012-10-30 19:02 | 28.609 | 249.6375 | 195.2607 | 247.3034 | 175.6498 | -1.6009 | 1.2709 | 32.1386 | 40.8452 | 25.2878 | 2.0074 | 32.0626 | 64.361 | 15.9726 | 14.4286 | -0.1104 | 105.2062 | 235.2089 | 180.8321 | 232.8748 | 161.2212
2012-10-30 19:02 | 28.609 | 249.6336 | 195.2804 | 247.4092 | 175.4859 | -1.6049 | 1.2709 | 32.1383 | 40.8443 | 25.2879 | 2.0074 | 32.0622 | 64.3603 | 15.9724 | 14.4286 | -0.1106 | 105.2047 | 235.205 | 180.8518 | 232.9806 | 161.0573
2012-10-30 19:02 | 28.6128 | 249.5943 | 195.2449 | 247.4913 | 175.4975 | -1.5928 | 1.2709 | 32.138 | 40.844 | 25.2877 | 2.0073 | 32.062 | 64.3582 | 15.9726 | 14.4201 | -0.1098 | 105.2022 | 235.1742 | 180.8248 | 233.0712 | 161.0774
2012-10-30 19:03 | 28.6233 | 249.6126 | 195.2357 | 247.4196 | 175.6329 | -1.5962 | 1.2709 | 32.1375 | 40.8427 | 25.2877 | 2.0073 | 32.0615 | 64.3568 | 15.9725 | 14.4388 | -0.11 | 105.1995 | 235.1738 | 180.7969 | 232.9808 | 161.1941
2012-10-30 19:03 | 28.6098 | 249.5991 | 195.2812 | 247.4217 | 175.5648 | -1.602 | 1.2709 | 32.1372 | 40.8419 | 25.2878 | 2.0073 | 32.0612 | 64.356 | 15.9724 | 14.4008 | -0.1104 | 105.1979 | 235.1983 | 180.8804 | 233.0209 | 161.164
2012-10-30 19:03 | 28.621 | 249.6338 | 195.257 | 247.3663 | 175.6657 | -1.5974 | 1.2709 | 32.1372 | 40.8417 | 25.2878 | 2.0073 | 32.0611 | 64.3561 | 15.9723 | 14.3916 | -0.1101 | 105.1978 | 235.2422 | 180.8654 | 232.9747 | 161.2741
2012-10-30 19:03 | 28.6173 | 249.6027 | 195.2887 | 247.4567 | 175.6035 | -1.6095 | 1.2708 | 32.1366 | 40.8401 | 25.2879 | 2.0072 | 32.0607 | 64.3537 | 15.9724 | 14.4124 | -0.111 | 105.1939 | 235.1903 | 180.8763 | 233.0443 | 161.1911
2012-10-30 19:03 | 28.6233 | 249.6361 | 195.2789 | 247.4744 | 175.8593 | -1.602 | 1.2708 | 32.1359 | 40.8383 | 25.2878 | 2.0072 | 32.0598 | 64.3499 | 15.9726 | 14.4102 | -0.1104 | 105.1882 | 235.2259 | 180.8687 | 233.0642 | 161.4491
2012-10-30 19:03 | 28.6233 | 249.6322 | 195.2671 | 247.3882 | 175.6524 | -1.5991 | 1.2708 | 32.1352 | 40.8367 | 25.2878 | 2.0071 | 32.0592 | 64.3475 | 15.9725 | 14.4102 | -0.1102 | 105.1842 | 235.222 | 180.8569 | 232.978 | 161.2422
2012-10-30 19:03 | 28.6255 | 249.6148 | 195.3636 | 247.3787 | 175.5297 | -1.602 | 1.2708 | 32.135 | 40.8359 | 25.2879 | 2.0071 | 32.0589 | 64.3469 | 15.9724 | 14.3716 | -0.1104 | 105.1828 | 235.2432 | 180.992 | 233.0071 | 161.1581
2012-10-30 19:03 | 28.6233 | 249.6048 | 195.3104 | 247.4823 | 175.5158 | -1.5962 | 1.2708 | 32.1349 | 40.8358 | 25.2879 | 2.0071 | 32.0587 | 64.3461 | 15.9724 | 14.3939 | -0.11 | 105.1819 | 235.2109 | 180.9165 | 233.0884 | 161.1219
2012-10-30 19:03 | 28.6285 | 249.6335 | 195.3195 | 247.4561 | 175.6147 | -1.6003 | 1.2708 | 32.1349 | 40.836 | 25.2878 | 2.0071 | 32.0589 | 64.3465 | 15.9724 | 14.4113 | -0.1103 | 105.1825 | 235.2222 | 180.9082 | 233.0448 | 161.2034
2012-10-30 19:03 | 28.6255 | 249.6266 | 195.2536 | 247.3317 | 175.4907 | -1.5997 | 1.2693 | 32.098 | 40.7426 | 25.2876 | 2.0047 | 32.0191 | 64.1875 | 15.9724 | 14.4369 | -0.1103 | 104.93 | 235.1897 | 180.8167 | 232.8948 | 161.0538
2012-10-30 19:03 | 28.6315 | 249.6796 | 195.295 | 247.4474 | 175.403 | -1.6009 | 1.2646 | 31.9781 | 40.4404 | 25.2865 | 1.9974 | 31.9031 | 63.723 | 15.9724 | 14.4184 | -0.1104 | 104.1635 | 235.2612 | 180.8766 | 233.029 | 160.9846
2012-10-30 19:03 | 28.6263 | 249.6 | 195.223 | 247.5009 | 175.39 | -1.5985 | 1.2643 | 31.9697 | 40.4196 | 25.2863 | 1.9968 | 31.8937 | 63.6852 | 15.9725 | 14.4091 | -0.1102 | 104.1047 | 235.1909 | 180.8139 | 233.0918 | 160.9809
2012-10-30 19:03 | 28.6293 | 249.6069 | 195.2339 | 247.3746 | 175.5374 | -1.6026 | 1.2637 | 31.9541 | 40.3803 | 25.2862 | 1.9958 | 31.8784 | 63.6241 | 15.9725 | 14.4489 | -0.1105 | 104.0044 | 235.158 | 180.785 | 232.9257 | 161.0885
2012-10-30 19:03 | 28.63 | 249.5998 | 195.1678 | 247.391 | 175.7059 | -1.598 | 1.2635 | 31.9478 | 40.365 | 25.2858 | 1.9954 | 31.8722 | 63.5987 | 15.9727 | 14.4414 | -0.1102 | 103.9636 | 235.1584 | 180.7264 | 232.9496 | 161.2645
2012-10-30 19:03 | 28.6345 | 249.5729 | 195.2076 | 247.4269 | 175.5504 | -1.6014 | 1.2634 | 31.9458 | 40.3603 | 25.2857 | 1.9953 | 31.8703 | 63.5911 | 15.9727 | 14.4704 | -0.1104 | 103.9514 | 235.1025 | 180.7372 | 232.9565 | 161.08
2012-10-30 19:03 | 28.6293 | 249.5482 | 195.226 | 247.453 | 175.4672 | -1.5985 | 1.2634 | 31.9465 | 40.3615 | 25.286 | 1.9953 | 31.871 | 63.5928 | 15.9729 | 14.4448 | -0.1102 | 103.9543 | 235.1034 | 180.7812 | 233.0082 | 161.0224
2012-10-30 19:03 | 28.6308 | 249.5183 | 195.2354 | 247.4427 | 175.7614 | -1.6026 | 1.2635 | 31.9475 | 40.3644 | 25.2857 | 1.9954 | 31.8719 | 63.5972 | 15.9727 | 14.4014 | -0.1105 | 103.9616 | 235.1169 | 180.834 | 233.0413 | 161.36
2012-10-30 19:03 | 28.6368 | 249.5518 | 195.2885 | 247.3155 | 175.342 | -1.5962 | 1.2636 | 31.9506 | 40.3716 | 25.2861 | 1.9956 | 31.875 | 63.6089 | 15.9728 | 14.44 | -0.11 | 103.9805 | 235.1118 | 180.8485 | 232.8755 | 160.902
2012-10-30 19:03 | 28.6203 | 249.5118 | 195.2052 | 247.3774 | 174.8182 | -1.6037 | 1.2646 | 31.976 | 40.4361 | 25.286 | 1.9971 | 31.8993 | 63.7059 | 15.9728 | 14.4317 | -0.1106 | 104.142 | 235.0801 | 180.7735 | 232.9457 | 160.3865
2012-10-30 19:03 | 28.6263 | 249.5138 | 195.0855 | 247.4265 | 174.672 | -1.6072 | 1.2686 | 32.0791 | 40.6955 | 25.2871 | 2.0036 | 32.0032 | 64.1211 | 15.9729 | 14.4663 | -0.1108 | 104.8166 | 235.0475 | 180.6192 | 232.9602 | 160.2057
2012-10-30 19:03 | 28.6233 | 249.5069 | 195.0786 | 247.4352 | 175.1607 | -1.5974 | 1.2692 | 32.0952 | 40.7355 | 25.2876 | 2.0046 | 32.0192 | 64.1861 | 15.9728 | 14.4959 | -0.1101 | 104.9216 | 235.011 | 180.5827 | 232.9393 | 160.6648
2012-10-30 19:03 | 28.6195 | 249.5188 | 195.1023 | 247.3139 | 175.5198 | -1.598 | 1.2698 | 32.1105 | 40.7741 | 25.2877 | 2.0056 | 32.0341 | 64.2461 | 15.9727 | 14.4391 | -0.1102 | 105.0202 | 235.0797 | 180.6632 | 232.8748 | 161.0807
2012-10-30 19:03 | 28.6113 | 249.538 | 195.1569 | 247.3136 | 175.6482 | -1.6078 | 1.2703 | 32.1223 | 40.8038 | 25.2879 | 2.0063 | 32.0461 | 64.2941 | 15.9728 | 14.4635 | -0.1108 | 105.098 | 235.0745 | 180.6934 | 232.8501 | 161.1847
2012-10-30 19:04 | 28.6001 | 249.5268 | 195.1732 | 247.3024 | 175.6214 | -1.6078 | 1.2705 | 32.1293 | 40.8216 | 25.2879 | 2.0068 | 32.0533 | 64.3235 | 15.9727 | 14.436 | -0.1108 | 105.1451 | 235.0908 | 180.7372 | 232.8664 | 161.1854
2012-10-30 19:04 | 28.6075 | 249.5029 | 195.1728 | 247.4273 | 175.7732 | -1.6049 | 1.2707 | 32.1334 | 40.8319 | 25.288 | 2.007 | 32.0573 | 64.3395 | 15.9727 | 14.4515 | -0.1106 | 105.1714 | 235.0514 | 180.7213 | 232.9758 | 161.3217
2012-10-30 19:04 | 28.6113 | 249.5184 | 195.2944 | 247.3253 | 176.019 | -1.6014 | 1.2708 | 32.1356 | 40.838 | 25.2876 | 2.0072 | 32.0595 | 64.3485 | 15.9726 | 14.4961 | -0.1104 | 105.1865 | 235.0223 | 180.7983 | 232.8292 | 161.5229
2012-10-30 19:04 | 28.5941 | 249.5521 | 195.3558 | 247.265 | 176.1268 | -1.5997 | 1.2708 | 32.1372 | 40.8413 | 25.2881 | 2.0073 | 32.0612 | 64.3552 | 15.9726 | 14.4789 | -0.1103 | 105.1965 | 235.0732 | 180.8769 | 232.7861 | 161.6479
2012-10-30 19:04 | 28.5986 | 249.5801 | 195.4193 | 247.3009 | 175.9009 | -1.5991 | 1.2709 | 32.1386 | 40.8454 | 25.2878 | 2.0073 | 32.0626 | 64.3603 | 15.9727 | 14.4426 | -0.1102 | 105.2057 | 235.1375 | 180.9767 | 232.8583 | 161.4583
2012-10-30 19:04 | 28.6023 | 249.5564 | 195.3955 | 247.4221 | 175.8656 | -1.602 | 1.2709 | 32.1389 | 40.8457 | 25.288 | 2.0074 | 32.063 | 64.3624 | 15.9726 | 14.4463 | -0.1104 | 105.2081 | 235.1101 | 180.9492 | 232.9758 | 161.4193
2012-10-30 19:04 | 28.6001 | 249.5659 | 195.4404 | 247.318 | 175.8751 | -1.5985 | 1.2709 | 32.1395 | 40.8477 | 25.2878 | 2.0074 | 32.0635 | 64.3648 | 15.9725 | 14.4972 | -0.1102 | 105.2125 | 235.0687 | 180.9432 | 232.8208 | 161.3779
2012-10-30 19:04 | 28.5926 | 249.5506 | 195.4368 | 247.287 | 175.8481 | -1.5997 | 1.271 | 32.1404 | 40.85 | 25.2878 | 2.0075 | 32.0644 | 64.3685 | 15.9725 | 14.453 | -0.1103 | 105.2185 | 235.0976 | 180.9838 | 232.834 | 161.3951
2012-10-30 19:04 | 28.5843 | 249.5775 | 195.4482 | 247.3414 | 175.953 | -1.5997 | 1.271 | 32.1409 | 40.8509 | 25.288 | 2.0075 | 32.0648 | 64.37 | 15.9726 | 14.461 | -0.1103 | 105.2209 | 235.1165 | 180.9872 | 232.8804 | 161.492
2012-10-30 19:04 | 28.5866 | 249.5877 | 195.4701 | 247.3633 | 176.0568 | -1.5968 | 1.271 | 32.1413 | 40.852 | 25.2879 | 2.0075 | 32.0651 | 64.3718 | 15.9724 | 14.4878 | -0.1101 | 105.2238 | 235.0999 | 180.9823 | 232.8755 | 161.569
2012-10-30 19:04 | 28.5873 | 249.5805 | 195.5337 | 247.4227 | 175.9482 | -1.6003 | 1.271 | 32.1414 | 40.8522 | 25.288 | 2.0075 | 32.0652 | 64.3724 | 15.9724 | 14.4313 | -0.1103 | 105.2246 | 235.1492 | 181.1024 | 232.9914 | 161.5169
2012-10-30 19:04 | 28.5851 | 249.6058 | 195.5354 | 247.2834 | 176.0514 | -1.598 | 1.271 | 32.1414 | 40.8517 | 25.2882 | 2.0075 | 32.0653 | 64.3724 | 15.9724 | 14.4332 | -0.1102 | 105.2241 | 235.1726 | 181.1022 | 232.8502 | 161.6182
2012-10-30 19:04 | 28.5896 | 249.5828 | 195.5596 | 247.3232 | 176.1262 | -1.6003 | 1.271 | 32.1416 | 40.8524 | 25.2881 | 2.0075 | 32.0654 | 64.3727 | 15.9724 | 14.45 | -0.1103 | 105.2251 | 235.1328 | 181.1096 | 232.8732 | 161.6762
2012-10-30 19:04 | 28.5926 | 249.6367 | 195.594 | 247.3967 | 176.0433 | -1.6037 | 1.271 | 32.1414 | 40.8517 | 25.2883 | 2.0075 | 32.0654 | 64.3721 | 15.9726 | 14.4774 | -0.1106 | 105.2238 | 235.1593 | 181.1166 | 232.9193 | 161.5659
2012-10-30 19:04 | 28.5956 | 249.628 | 195.6206 | 247.435 | 175.9643 | -1.602 | 1.271 | 32.1418 | 40.8531 | 25.2881 | 2.0076 | 32.0658 | 64.3748 | 15.9724 | 14.46 | -0.1104 | 105.2279 | 235.168 | 181.1606 | 232.975 | 161.5043
2012-10-30 19:04 | 28.5918 | 249.6007 | 195.5736 | 247.2862 | 175.9995 | -1.6032 | 1.271 | 32.1419 | 40.853 | 25.2882 | 2.0075 | 32.0657 | 64.3734 | 15.9726 | 14.4685 | -0.1105 | 105.2263 | 235.1322 | 181.1051 | 232.8177 | 161.531
2012-10-30 19:04 | 28.5941 | 249.5991 | 195.5444 | 247.4178 | 175.5803 | -1.602 | 1.271 | 32.142 | 40.8532 | 25.2883 | 2.0076 | 32.0658 | 64.3748 | 15.9724 | 14.4463 | -0.1104 | 105.228 | 235.1528 | 181.0981 | 232.9715 | 161.134
2012-10-30 19:04 | 28.5971 | 249.6256 | 195.5514 | 247.3895 | 175.634 | -1.6037 | 1.271 | 32.1423 | 40.8542 | 25.2881 | 2.0076 | 32.0663 | 64.3763 | 15.9724 | 14.4697 | -0.1106 | 105.2305 | 235.1559 | 181.0817 | 232.9198 | 161.1643
2012-10-30 19:04 | 28.5963 | 249.6248 | 195.4956 | 247.4161 | 175.762 | -1.598 | 1.271 | 32.1421 | 40.8537 | 25.2882 | 2.0076 | 32.0661 | 64.3757 | 15.9724 | 14.4567 | -0.1102 | 105.2294 | 235.1681 | 181.0389 | 232.9594 | 161.3053
2012-10-30 19:04 | 28.6135 | 249.6459 | 195.5363 | 247.3236 | 175.7519 | -1.6003 | 1.271 | 32.1419 | 40.8529 | 25.2883 | 2.0076 | 32.0658 | 64.3746 | 15.9724 | 14.4412 | -0.1103 | 105.2275 | 235.2047 | 181.0951 | 232.8824 | 161.3107
2012-10-30 19:04 | 28.6075 | 249.6399 | 195.4753 | 247.4194 | 175.7537 | -1.602 | 1.271 | 32.1418 | 40.8531 | 25.288 | 2.0076 | 32.0657 | 64.3747 | 15.9722 | 14.4923 | -0.1104 | 105.2278 | 235.1476 | 180.983 | 232.9271 | 161.2614
2012-10-30 19:04 | 28.6053 | 249.6377 | 195.422 | 247.4016 | 175.6656 | -1.602 | 1.271 | 32.1415 | 40.8524 | 25.2881 | 2.0076 | 32.0655 | 64.3739 | 15.9722 | 14.4371 | -0.1104 | 105.2263 | 235.2006 | 180.9849 | 232.9645 | 161.2285
2012-10-30 19:04 | 28.6165 | 249.6567 | 195.4097 | 247.4598 | 175.7354 | -1.5997 | 1.271 | 32.1414 | 40.8521 | 25.2881 | 2.0076 | 32.0653 | 64.3734 | 15.9722 | 14.4483 | -0.1103 | 105.2255 | 235.2084 | 180.9614 | 233.0115 | 161.2871
2012-10-30 19:05 | 28.6143 | 249.6232 | 195.4546 | 247.3518 | 175.78 | -1.5962 | 1.271 | 32.1412 | 40.8516 | 25.288 | 2.0076 | 32.0651 | 64.3727 | 15.9722 | 14.4379 | -0.11 | 105.2243 | 235.1853 | 181.0167 | 232.9139 | 161.3421
2012-10-30 19:05 | 28.606 | 249.6697 | 195.4188 | 247.4532 | 175.7093 | -1.6026 | 1.271 | 32.1408 | 40.8508 | 25.2879 | 2.0075 | 32.0647 | 64.3705 | 15.9723 | 14.4541 | -0.1105 | 105.2213 | 235.2156 | 180.9647 | 232.9991 | 161.2552
2012-10-30 19:05 | 28.6053 | 249.6299 | 195.4102 | 247.4172 | 175.6344 | -1.6003 | 1.271 | 32.1403 | 40.8493 | 25.2881 | 2.0075 | 32.0643 | 64.3696 | 15.9721 | 14.4167 | -0.1103 | 105.219 | 235.2132 | 180.9935 | 233.0005 | 161.2177
2012-10-30 19:05 | 28.6173 | 249.6654 | 195.379 | 247.441 | 175.5683 | -1.6037 | 1.271 | 32.1402 | 40.8492 | 25.2879 | 2.0075 | 32.0641 | 64.3695 | 15.972 | 14.4491 | -0.1106 | 105.2188 | 235.2163 | 180.9299 | 232.9919 | 161.1192
2012-10-30 19:05 | 28.6158 | 249.6325 | 195.4129 | 247.3807 | 175.7074 | -1.6009 | 1.271 | 32.1403 | 40.8487 | 25.2884 | 2.0075 | 32.0641 | 64.3692 | 15.972 | 14.3701 | -0.1104 | 105.2179 | 235.2624 | 181.0428 | 233.0106 | 161.3373
2012-10-30 19:05 | 28.6128 | 249.6256 | 195.3706 | 247.4443 | 175.6809 | -1.6043 | 1.2709 | 32.1399 | 40.8479 | 25.2882 | 2.0075 | 32.0637 | 64.3676 | 15.9721 | 14.412 | -0.1106 | 105.2155 | 235.2136 | 180.9586 | 233.0323 | 161.2689
2012-10-30 19:05 | 28.606 | 249.6658 | 195.3717 | 247.5002 | 175.6546 | -1.6032 | 1.2709 | 32.1394 | 40.8474 | 25.2878 | 2.0075 | 32.0634 | 64.3664 | 15.972 | 14.4092 | -0.1105 | 105.2139 | 235.2566 | 180.9625 | 233.091 | 161.2454
2012-10-30 19:05 | 28.6158 | 249.6913 | 195.3657 | 247.4042 | 175.5551 | -1.5991 | 1.2709 | 32.139 | 40.8463 | 25.2879 | 2.0074 | 32.0629 | 64.3646 | 15.972 | 14.4394 | -0.1102 | 105.2109 | 235.2519 | 180.9263 | 232.9648 | 161.1157
2012-10-30 19:05 | 28.615 | 249.6983 | 195.3296 | 247.3682 | 175.6012 | -1.5962 | 1.2709 | 32.1391 | 40.8462 | 25.2881 | 2.0074 | 32.063 | 64.3645 | 15.9721 | 14.406 | -0.11 | 105.2107 | 235.2923 | 180.9236 | 232.9622 | 161.1952
2012-10-30 19:05 | 28.6105 | 249.6507 | 195.3329 | 247.5086 | 175.5381 | -1.5974 | 1.2709 | 32.1385 | 40.8453 | 25.2877 | 2.0074 | 32.0625 | 64.363 | 15.9719 | 14.4545 | -0.1101 | 105.2083 | 235.1962 | 180.8784 | 233.0541 | 161.0836
2012-10-30 19:05 | 28.6128 | 249.6765 | 195.3274 | 247.4796 | 175.6966 | -1.6026 | 1.2709 | 32.138 | 40.8435 | 25.288 | 2.0074 | 32.062 | 64.3608 | 15.972 | 14.3916 | -0.1105 | 105.2043 | 235.2849 | 180.9358 | 233.088 | 161.305
2012-10-30 19:05 | 28.615 | 249.6709 | 195.3217 | 247.3839 | 175.6909 | -1.6003 | 1.2709 | 32.1378 | 40.8432 | 25.288 | 2.0074 | 32.0617 | 64.36 | 15.9719 | 14.4264 | -0.1103 | 105.2032 | 235.2445 | 180.8953 | 232.9575 | 161.2645
2012-10-30 19:05 | 28.6128 | 249.6687 | 195.3195 | 247.4678 | 175.4155 | -1.602 | 1.2709 | 32.1372 | 40.8418 | 25.2879 | 2.0074 | 32.0612 | 64.3586 | 15.9718 | 14.4405 | -0.1104 | 105.2004 | 235.2282 | 180.879 | 233.0273 | 160.975
2012-10-30 19:05 | 28.6165 | 249.7194 | 195.2839 | 247.5342 | 175.3646 | -1.5951 | 1.2708 | 32.1368 | 40.8407 | 25.2878 | 2.0073 | 32.0607 | 64.3564 | 15.9718 | 14.4197 | -0.11 | 105.1971 | 235.2997 | 180.8642 | 233.1145 | 160.9449
2012-10-30 19:05 | 28.612 | 249.6953 | 195.2716 | 247.5258 | 175.5162 | -1.5985 | 1.2708 | 32.1367 | 40.8403 | 25.288 | 2.0073 | 32.0606 | 64.3561 | 15.9718 | 14.4438 | -0.1102 | 105.1964 | 235.2515 | 180.8278 | 233.082 | 161.0724
2012-10-30 19:05 | 28.621 | 249.6965 | 195.2845 | 247.4094 | 175.3964 | -1.6009 | 1.2708 | 32.1367 | 40.8405 | 25.2878 | 2.0073 | 32.0606 | 64.3564 | 15.9717 | 14.4977 | -0.1104 | 105.1968 | 235.1988 | 180.7868 | 232.9117 | 160.8987
2012-10-30 19:05 | 28.6225 | 249.6862 | 195.2428 | 247.5559 | 175.3394 | -1.5991 | 1.2708 | 32.1367 | 40.8402 | 25.288 | 2.0073 | 32.0605 | 64.3561 | 15.9717 | 14.4584 | -0.1102 | 105.1963 | 235.2278 | 180.7844 | 233.0975 | 160.881
2012-10-30 19:05 | 28.6195 | 249.7028 | 195.2476 | 247.5059 | 175.5354 | -1.6078 | 1.2708 | 32.1362 | 40.8391 | 25.288 | 2.0073 | 32.0602 | 64.3552 | 15.9716 | 14.5043 | -0.1108 | 105.1943 | 235.1985 | 180.7433 | 233.0016 | 161.0311
2012-10-30 19:05 | 28.618 | 249.6739 | 195.2501 | 247.4495 | 175.6471 | -1.5997 | 1.2708 | 32.1357 | 40.8379 | 25.2879 | 2.0073 | 32.0597 | 64.3526 | 15.9718 | 14.4947 | -0.1103 | 105.1905 | 235.1792 | 180.7554 | 232.9548 | 161.1524
2012-10-30 19:05 | 28.6248 | 249.7237 | 195.3079 | 247.4211 | 175.5485 | -1.602 | 1.2708 | 32.1353 | 40.8369 | 25.2878 | 2.0072 | 32.0593 | 64.351 | 15.9718 | 14.4933 | -0.1104 | 105.1878 | 235.2304 | 180.8146 | 232.9278 | 161.0552
2012-10-30 19:05 | 28.6188 | 249.7177 | 195.2862 | 247.5639 | 175.3903 | -1.6014 | 1.2708 | 32.1347 | 40.8359 | 25.2876 | 2.0072 | 32.0587 | 64.3486 | 15.9717 | 14.5077 | -0.1104 | 105.1846 | 235.21 | 180.7785 | 233.0562 | 160.8826
2012-10-30 19:05 | 28.621 | 249.6808 | 195.2727 | 247.4604 | 175.4667 | -1.6049 | 1.2708 | 32.1344 | 40.835 | 25.2876 | 2.0072 | 32.0583 | 64.3474 | 15.9717 | 14.465 | -0.1106 | 105.1824 | 235.2158 | 180.8077 | 232.9954 | 161.0017
2012-10-30 19:05 | 28.6113 | 249.7298 | 195.2787 | 247.3762 | 175.5545 | -1.602 | 1.2707 | 32.1342 | 40.8341 | 25.2878 | 2.0072 | 32.0582 | 64.3469 | 15.9717 | 14.4839 | -0.1104 | 105.181 | 235.2459 | 180.7948 | 232.8923 | 161.0706
2012-10-30 19:05 | 28.6225 | 249.6784 | 195.2663 | 247.5089 | 175.5735 | -1.6026 | 1.2707 | 32.1337 | 40.833 | 25.2878 | 2.0072 | 32.0577 | 64.3449 | 15.9717 | 14.4584 | -0.1105 | 105.1779 | 235.22 | 180.8079 | 233.0505 | 161.1151
2012-10-30 19:05 | 28.615 | 249.7179 | 195.3335 | 247.5327 | 175.6051 | -1.602 | 1.2707 | 32.1334 | 40.8325 | 25.2876 | 2.0071 | 32.0574 | 64.3437 | 15.9717 | 14.4835 | -0.1104 | 105.1762 | 235.2344 | 180.85 | 233.0492 | 161.1216
2012-10-30 19:06 | 28.6188 | 249.6825 | 195.3019 | 247.419 | 175.6791 | -1.6037 | 1.2653 | 31.9943 | 40.4811 | 25.2868 | 1.9983 | 31.9161 | 63.7776 | 15.9717 | 14.4914 | -0.1106 | 104.2587 | 235.1911 | 180.8105 | 232.9276 | 161.1877
2012-10-30 19:06 | 28.6218 | 249.6659 | 195.2539 | 247.4416 | 175.6314 | -1.5974 | 1.2644 | 31.9724 | 40.4266 | 25.2861 | 1.997 | 31.8963 | 63.6979 | 15.9719 | 14.4903 | -0.1101 | 104.1245 | 235.1756 | 180.7636 | 232.9513 | 161.1411
2012-10-30 19:06 | 28.6165 | 249.6332 | 195.2486 | 247.5225 | 175.5246 | -1.6014 | 1.2637 | 31.9551 | 40.3832 | 25.2859 | 1.996 | 31.8793 | 63.6304 | 15.9717 | 14.485 | -0.1104 | 104.0136 | 235.1482 | 180.7636 | 233.0375 | 161.0396
2012-10-30 19:06 | 28.6188 | 249.6395 | 195.2509 | 247.4425 | 175.5855 | -1.6014 | 1.2635 | 31.9491 | 40.3682 | 25.2858 | 1.9956 | 31.8734 | 63.6067 | 15.9718 | 14.4955 | -0.1104 | 103.9749 | 235.144 | 180.7554 | 232.947 | 161.09
2012-10-30 19:06 | 28.6128 | 249.5982 | 195.2488 | 247.4365 | 175.2477 | -1.5997 | 1.2634 | 31.9465 | 40.3616 | 25.2859 | 1.9954 | 31.871 | 63.5961 | 15.972 | 14.4691 | -0.1103 | 103.9577 | 235.1291 | 180.7797 | 232.9674 | 160.7786
2012-10-30 19:06 | 28.6225 | 249.5844 | 195.2231 | 247.505 | 175.074 | -1.6014 | 1.2634 | 31.9464 | 40.3614 | 25.2859 | 1.9954 | 31.8708 | 63.5953 | 15.9721 | 14.4951 | -0.1104 | 103.9567 | 235.0893 | 180.728 | 233.0099 | 160.5789
2012-10-30 19:06 | 28.6128 | 249.5552 | 195.0916 | 247.4247 | 174.8614 | -1.6026 | 1.2637 | 31.9529 | 40.3777 | 25.286 | 1.9958 | 31.8771 | 63.6204 | 15.9721 | 14.4976 | -0.1105 | 103.9981 | 235.0576 | 180.594 | 232.9271 | 160.3638
2012-10-30 19:06 | 28.6023 | 249.5877 | 194.9711 | 247.3046 | 175.0226 | -1.5968 | 1.2683 | 32.0722 | 40.678 | 25.287 | 2.0031 | 31.9942 | 64.0886 | 15.9721 | 14.5116 | -0.1101 | 104.7666 | 235.0761 | 180.4595 | 232.793 | 160.511
2012-10-30 19:06 | 28.6068 | 249.4904 | 194.956 | 247.4187 | 175.0427 | -1.6049 | 1.2689 | 32.0872 | 40.7158 | 25.2873 | 2.0042 | 32.0116 | 64.1588 | 15.9719 | 14.5202 | -0.1106 | 104.8745 | 234.9702 | 180.4358 | 232.8985 | 160.5225
2012-10-30 19:06 | 28.6053 | 249.5359 | 195.0213 | 247.3898 | 175.4315 | -1.5985 | 1.2695 | 32.1012 | 40.7513 | 25.2873 | 2.0051 | 32.0255 | 64.2147 | 15.9719 | 14.4983 | -0.1102 | 104.966 | 235.0376 | 180.523 | 232.8915 | 160.9332
2012-10-30 19:06 | 28.6053 | 249.5242 | 195.0802 | 247.2919 | 175.4432 | -1.6014 | 1.27 | 32.1147 | 40.785 | 25.2876 | 2.0059 | 32.0387 | 64.2678 | 15.9719 | 14.5187 | -0.1104 | 105.0528 | 235.0055 | 180.5615 | 232.7732 | 160.9245
2012-10-30 19:06 | 28.6083 | 249.5428 | 195.0518 | 247.4633 | 175.4852 | -1.6032 | 1.2703 | 32.1237 | 40.8074 | 25.2879 | 2.0065 | 32.0474 | 64.3025 | 15.972 | 14.5095 | -0.1105 | 105.11 | 235.0333 | 180.5423 | 232.9538 | 160.9757
2012-10-30 19:06 | 28.5926 | 249.5467 | 195.1814 | 247.3497 | 175.6451 | -1.6014 | 1.2708 | 32.1349 | 40.8356 | 25.288 | 2.0072 | 32.0587 | 64.3482 | 15.9719 | 14.5019 | -0.1104 | 105.1838 | 235.0448 | 180.6795 | 232.8478 | 161.1432
2012-10-30 19:06 | 28.5791 | 249.5371 | 195.1876 | 247.391 | 175.4248 | -1.5945 | 1.2709 | 32.1397 | 40.8479 | 25.288 | 2.0075 | 32.0635 | 64.3675 | 15.9719 | 14.468 | -0.1099 | 105.2154 | 235.0691 | 180.7196 | 232.923 | 160.9568
2012-10-30 19:06 | 28.5783 | 249.505 | 195.2418 | 247.3824 | 175.6972 | -1.602 | 1.271 | 32.1403 | 40.8491 | 25.2881 | 2.0076 | 32.0641 | 64.3708 | 15.9716 | 14.4876 | -0.1104 | 105.2199 | 235.0174 | 180.7542 | 232.8948 | 161.2096
2012-10-30 19:06 | 28.5776 | 249.5865 | 195.1979 | 247.2485 | 175.6457 | -1.6043 | 1.271 | 32.1407 | 40.8507 | 25.2879 | 2.0076 | 32.0646 | 64.3722 | 15.9717 | 14.4624 | -0.1106 | 105.2229 | 235.1241 | 180.7355 | 232.7861 | 161.1833
2012-10-30 19:06 | 28.5798 | 249.5887 | 195.2079 | 247.4348 | 175.4489 | -1.5997 | 1.271 | 32.1412 | 40.8519 | 25.2879 | 2.0076 | 32.0651 | 64.3745 | 15.9717 | 14.4932 | -0.1103 | 105.2264 | 235.0955 | 180.7147 | 232.9416 | 160.9557
2012-10-30 19:06 | 28.5836 | 249.5651 | 195.2589 | 247.3525 | 175.5073 | -1.598 | 1.271 | 32.1416 | 40.8531 | 25.2878 | 2.0076 | 32.0655 | 64.3761 | 15.9717 | 14.444 | -0.1102 | 105.2291 | 235.1211 | 180.8149 | 232.9085 | 161.0633
2012-10-30 19:07 | 28.5701 | 249.5516 | 195.2493 | 247.2214 | 175.6695 | -1.5974 | 1.2711 | 32.1425 | 40.8552 | 25.2879 | 2.0077 | 32.0664 | 64.3795 | 15.9717 | 14.4672 | -0.1101 | 105.2347 | 235.0844 | 180.7821 | 232.7542 | 161.2023
2012-10-30 19:07 | 28.5738 | 249.5788 | 195.2294 | 247.4249 | 175.5522 | -1.6003 | 1.2711 | 32.1428 | 40.8556 | 25.288 | 2.0077 | 32.0665 | 64.3808 | 15.9716 | 14.4423 | -0.1103 | 105.2365 | 235.1365 | 180.7871 | 232.9826 | 161.1099
2012-10-30 19:07 | 28.5753 | 249.5725 | 195.2663 | 247.4343 | 175.5654 | -1.5939 | 1.2711 | 32.143 | 40.8564 | 25.288 | 2.0077 | 32.0669 | 64.382 | 15.9716 | 14.4316 | -0.1099 | 105.2383 | 235.1409 | 180.8347 | 233.0027 | 161.1338
2012-10-30 19:07 | 28.5746 | 249.5835 | 195.2617 | 247.222 | 175.5842 | -1.6049 | 1.2711 | 32.1435 | 40.8571 | 25.2882 | 2.0078 | 32.0673 | 64.3836 | 15.9716 | 14.4431 | -0.1106 | 105.2407 | 235.1404 | 180.8186 | 232.7789 | 161.1411
2012-10-30 19:07 | 28.5783 | 249.5598 | 195.1986 | 247.3511 | 175.6074 | -1.5991 | 1.2711 | 32.1436 | 40.8578 | 25.2879 | 2.0078 | 32.0674 | 64.385 | 15.9714 | 14.4591 | -0.1102 | 105.2429 | 235.1007 | 180.7395 | 232.892 | 161.1483
2012-10-30 19:07 | 28.5693 | 249.5469 | 195.3192 | 247.4243 | 175.6413 | -1.6043 | 1.2711 | 32.1434 | 40.8571 | 25.2881 | 2.0078 | 32.0673 | 64.3842 | 15.9715 | 14.4623 | -0.1106 | 105.2413 | 235.0846 | 180.8569 | 232.962 | 161.179
2012-10-30 19:07 | 28.5708 | 249.5758 | 195.2697 | 247.2652 | 175.6194 | -1.6003 | 1.2711 | 32.1433 | 40.8574 | 25.2878 | 2.0078 | 32.0672 | 64.3837 | 15.9715 | 14.4312 | -0.1103 | 105.2411 | 235.1446 | 180.8385 | 232.834 | 161.1882
2012-10-30 19:07 | 28.5753 | 249.5803 | 195.2074 | 247.4029 | 175.3351 | -1.5997 | 1.2711 | 32.1432 | 40.8572 | 25.2878 | 2.0078 | 32.0671 | 64.3831 | 15.9716 | 14.4846 | -0.1103 | 105.2404 | 235.0957 | 180.7228 | 232.9183 | 160.8505
2012-10-30 19:07 | 28.5656 | 249.551 | 195.2134 | 247.3775 | 175.38 | -1.6003 | 1.2711 | 32.1431 | 40.8564 | 25.2881 | 2.0078 | 32.067 | 64.3834 | 15.9714 | 14.4504 | -0.1103 | 105.2398 | 235.1006 | 180.763 | 232.9271 | 160.9296
2012-10-30 19:07 | 28.5693 | 249.6291 | 195.2053 | 247.3146 | 175.4852 | -1.602 | 1.2711 | 32.1424 | 40.8548 | 25.288 | 2.0077 | 32.0663 | 64.3806 | 15.9714 | 14.4174 | -0.1104 | 105.2354 | 235.2117 | 180.7879 | 232.8972 | 161.0678
2012-10-30 19:07 | 28.5708 | 249.5836 | 195.1675 | 247.3632 | 175.4009 | -1.6003 | 1.2711 | 32.1423 | 40.8547 | 25.2878 | 2.0077 | 32.0662 | 64.38 | 15.9714 | 14.4556 | -0.1103 | 105.2347 | 235.128 | 180.7119 | 232.9076 | 160.9453
2012-10-30 19:07 | 28.5701 | 249.5868 | 195.2179 | 247.4526 | 175.4743 | -1.6037 | 1.271 | 32.1422 | 40.854 | 25.288 | 2.0077 | 32.066 | 64.3794 | 15.9714 | 14.4182 | -0.1106 | 105.2335 | 235.1686 | 180.7997 | 233.0344 | 161.0561
2012-10-30 19:07 | 28.5701 | 249.6221 | 195.2022 | 247.382 | 175.5094 | -1.6055 | 1.271 | 32.142 | 40.8541 | 25.2878 | 2.0077 | 32.0659 | 64.3786 | 15.9715 | 14.4345 | -0.1107 | 105.2327 | 235.1876 | 180.7677 | 232.9475 | 161.0749
2012-10-30 19:07 | 28.5649 | 249.6247 | 195.2087 | 247.3377 | 175.5862 | -1.6032 | 1.271 | 32.1416 | 40.8527 | 25.2879 | 2.0077 | 32.0653 | 64.3767 | 15.9714 | 14.4538 | -0.1105 | 105.2294 | 235.1709 | 180.7549 | 232.8839 | 161.1324
2012-10-30 19:07 | 28.5619 | 249.5669 | 195.2333 | 247.4326 | 175.4388 | -1.598 | 1.271 | 32.1411 | 40.8508 | 25.2883 | 2.0076 | 32.0649 | 64.3748 | 15.9714 | 14.4182 | -0.1102 | 105.2256 | 235.1487 | 180.8151 | 233.0144 | 161.0206
2012-10-30 19:07 | 28.5634 | 249.6075 | 195.2308 | 247.4889 | 175.4208 | -1.6061 | 1.271 | 32.1408 | 40.8509 | 25.2878 | 2.0076 | 32.0645 | 64.3739 | 15.9713 | 14.5013 | -0.1107 | 105.2247 | 235.1062 | 180.7295 | 232.9876 | 160.9195
2012-10-30 19:07 | 28.5664 | 249.6184 | 195.2181 | 247.3274 | 175.4706 | -1.598 | 1.271 | 32.1406 | 40.8504 | 25.2879 | 2.0076 | 32.0645 | 64.3732 | 15.9714 | 14.4757 | -0.1102 | 105.2235 | 235.1427 | 180.7424 | 232.8517 | 160.9949
2012-10-30 19:07 | 28.5581 | 249.6179 | 195.1902 | 247.4092 | 175.3882 | -1.6014 | 1.271 | 32.1406 | 40.8501 | 25.288 | 2.0076 | 32.0644 | 64.3731 | 15.9714 | 14.4674 | -0.1104 | 105.2232 | 235.1505 | 180.7228 | 232.9418 | 160.9208
2012-10-30 19:07 | 28.5574 | 249.5937 | 195.1345 | 247.479 | 175.3836 | -1.6026 | 1.271 | 32.1401 | 40.8491 | 25.2878 | 2.0076 | 32.064 | 64.3715 | 15.9713 | 14.4626 | -0.1105 | 105.2206 | 235.1311 | 180.6719 | 233.0164 | 160.921
2012-10-30 19:07 | 28.5671 | 249.6152 | 195.2267 | 247.3203 | 175.6977 | -1.598 | 1.271 | 32.1399 | 40.8482 | 25.2881 | 2.0076 | 32.0637 | 64.3711 | 15.9711 | 14.4642 | -0.1102 | 105.2193 | 235.151 | 180.7625 | 232.8561 | 161.2335
2012-10-30 19:07 | 28.5671 | 249.6426 | 195.1363 | 247.3908 | 175.5923 | -1.6032 | 1.2709 | 32.1395 | 40.8476 | 25.2878 | 2.0076 | 32.0633 | 64.369 | 15.9713 | 14.4397 | -0.1105 | 105.2166 | 235.2029 | 180.6966 | 232.9511 | 161.1526
2012-10-30 19:07 | 28.5581 | 249.614 | 195.2177 | 247.5189 | 175.5052 | -1.5997 | 1.2709 | 32.139 | 40.8465 | 25.2877 | 2.0075 | 32.0629 | 64.367 | 15.9713 | 14.4511 | -0.1103 | 105.2135 | 235.1629 | 180.7666 | 233.0678 | 161.0541
2012-10-30 19:07 | 28.5581 | 249.6336 | 195.2019 | 247.37 | 175.5248 | -1.6003 | 1.2709 | 32.1389 | 40.8466 | 25.2876 | 2.0075 | 32.0628 | 64.3674 | 15.9712 | 14.3777 | -0.1103 | 105.214 | 235.2559 | 180.8242 | 232.9923 | 161.1471
2012-10-30 19:07 | 28.5634 | 249.6271 | 195.1719 | 247.3283 | 175.5886 | -1.6026 | 1.2709 | 32.1388 | 40.8455 | 25.2881 | 2.0075 | 32.0626 | 64.3666 | 15.9712 | 14.4238 | -0.1105 | 105.2121 | 235.2033 | 180.7481 | 232.9045 | 161.1648
2012-10-30 19:08 | 28.5634 | 249.5801 | 195.1129 | 247.4459 | 175.6589 | -1.602 | 1.2709 | 32.1383 | 40.8448 | 25.2877 | 2.0075 | 32.0622 | 64.365 | 15.9711 | 14.4482 | -0.1104 | 105.2098 | 235.1319 | 180.6647 | 232.9977 | 161.2107
2012-10-30 19:08 | 28.5596 | 249.6233 | 195.2349 | 247.3598 | 175.6668 | -1.5985 | 1.2709 | 32.1381 | 40.8438 | 25.2879 | 2.0075 | 32.062 | 64.3642 | 15.9712 | 14.3996 | -0.1102 | 105.208 | 235.2237 | 180.8353 | 232.9602 | 161.2672
2012-10-30 19:08 | 28.5596 | 249.6233 | 195.1524 | 247.4068 | 175.5419 | -1.6009 | 1.2709 | 32.1376 | 40.8429 | 25.2878 | 2.0075 | 32.0615 | 64.3626 | 15.9711 | 14.4648 | -0.1104 | 105.2055 | 235.1585 | 180.6876 | 232.942 | 161.0771
2012-10-30 19:08 | 28.5649 | 249.6443 | 195.2205 | 247.4474 | 175.512 | -1.6032 | 1.2709 | 32.1374 | 40.842 | 25.288 | 2.0075 | 32.0613 | 64.3621 | 15.9711 | 14.4171 | -0.1105 | 105.2041 | 235.2272 | 180.8034 | 233.0303 | 161.0949
2012-10-30 19:08 | 28.5626 | 249.6733 | 195.2222 | 247.3354 | 175.5995 | -1.5968 | 1.2709 | 32.1377 | 40.8434 | 25.2876 | 2.0075 | 32.0616 | 64.3631 | 15.971 | 14.4434 | -0.1101 | 105.2065 | 235.2299 | 180.7788 | 232.892 | 161.1561
2012-10-30 19:08 | 28.5596 | 249.6116 | 195.172 | 247.4303 | 175.3311 | -1.6055 | 1.2709 | 32.1377 | 40.8429 | 25.2879 | 2.0075 | 32.0615 | 64.3623 | 15.9712 | 14.42 | -0.1107 | 105.2052 | 235.1916 | 180.752 | 233.0103 | 160.9111
2012-10-30 19:08 | 28.5604 | 249.6202 | 195.2318 | 247.4311 | 175.4841 | -1.6043 | 1.2709 | 32.1375 | 40.8422 | 25.288 | 2.0074 | 32.0612 | 64.3612 | 15.9711 | 14.4208 | -0.1106 | 105.2035 | 235.1994 | 180.811 | 233.0103 | 161.0633
2012-10-30 19:08 | 28.5581 | 249.6218 | 195.1234 | 247.323 | 175.4974 | -1.6055 | 1.2709 | 32.1372 | 40.8418 | 25.2877 | 2.0074 | 32.0611 | 64.3606 | 15.9711 | 14.4389 | -0.1107 | 105.2024 | 235.1829 | 180.6845 | 232.8841 | 161.0585
2012-10-30 19:08 | 28.5693 | 249.6369 | 195.1699 | 247.4753 | 175.4813 | -1.6049 | 1.2708 | 32.1367 | 40.8403 | 25.2879 | 2.0074 | 32.0606 | 64.3592 | 15.971 | 14.4297 | -0.1106 | 105.1995 | 235.2072 | 180.7402 | 233.0456 | 161.0516
2012-10-30 19:08 | 28.5731 | 249.6995 | 195.2169 | 247.3694 | 175.5398 | -1.6043 | 1.2708 | 32.1363 | 40.8395 | 25.2878 | 2.0074 | 32.0602 | 64.3578 | 15.971 | 14.4457 | -0.1106 | 105.1972 | 235.2538 | 180.7712 | 232.9237 | 161.0941
2012-10-30 19:08 | 28.5664 | 249.6145 | 195.1513 | 247.3235 | 175.4862 | -1.6009 | 1.2708 | 32.1362 | 40.8393 | 25.2878 | 2.0074 | 32.06 | 64.3566 | 15.9711 | 14.4064 | -0.1104 | 105.1959 | 235.2081 | 180.7449 | 232.9171 | 161.0798
2012-10-30 19:08 | 28.5664 | 249.6027 | 195.167 | 247.5076 | 175.5877 | -1.6009 | 1.2708 | 32.1359 | 40.8386 | 25.2877 | 2.0074 | 32.0598 | 64.356 | 15.971 | 14.4308 | -0.1104 | 105.1947 | 235.1719 | 180.7362 | 233.0768 | 161.1569
2012-10-30 19:08 | 28.5686 | 249.648 | 195.2203 | 247.3296 | 175.5314 | -1.598 | 1.2708 | 32.1356 | 40.8381 | 25.2876 | 2.0074 | 32.0595 | 64.355 | 15.971 | 14.4575 | -0.1102 | 105.1932 | 235.1905 | 180.7628 | 232.8721 | 161.0739
2012-10-30 19:08 | 28.5619 | 249.6648 | 195.1782 | 247.5149 | 175.5364 | -1.6043 | 1.2708 | 32.1354 | 40.8372 | 25.2878 | 2.0074 | 32.0592 | 64.3542 | 15.9709 | 14.4263 | -0.1106 | 105.1914 | 235.2385 | 180.7519 | 233.0886 | 161.1101
2012-10-30 19:08 | 28.5753 | 249.6977 | 195.2624 | 247.3677 | 175.7488 | -1.5997 | 1.2708 | 32.1354 | 40.8376 | 25.2876 | 2.0073 | 32.0593 | 64.3539 | 15.971 | 14.452 | -0.1103 | 105.1915 | 235.2457 | 180.8104 | 232.9157 | 161.2968
2012-10-30 19:08 | 28.5634 | 249.6467 | 195.1915 | 247.3518 | 175.6159 | -1.602 | 1.2708 | 32.1355 | 40.8375 | 25.2878 | 2.0074 | 32.0595 | 64.3549 | 15.971 | 14.4482 | -0.1104 | 105.1923 | 235.1985 | 180.7433 | 232.9036 | 161.1677
2012-10-30 19:08 | 28.5649 | 249.6286 | 195.2913 | 247.4669 | 175.6447 | -1.602 | 1.2708 | 32.1356 | 40.8374 | 25.288 | 2.0074 | 32.0595 | 64.3549 | 15.9709 | 14.4579 | -0.1104 | 105.1923 | 235.1707 | 180.8334 | 233.009 | 161.1868
2012-10-30 19:08 | 28.5783 | 249.6733 | 195.2261 | 247.402 | 175.7167 | -1.6003 | 1.2708 | 32.135 | 40.8368 | 25.2875 | 2.0074 | 32.0589 | 64.3535 | 15.9708 | 14.4427 | -0.1103 | 105.1903 | 235.2306 | 180.7834 | 232.9593 | 161.274
2012-10-30 19:08 | 28.5678 | 249.655 | 195.302 | 247.509 | 175.8038 | -1.5997 | 1.2644 | 31.9727 | 40.427 | 25.2863 | 1.9969 | 31.8924 | 63.6857 | 15.971 | 14.4078 | -0.1103 | 104.1126 | 235.2472 | 180.8942 | 233.1012 | 161.396
2012-10-30 19:08 | 28.5753 | 249.6586 | 195.2859 | 247.3402 | 175.7996 | -1.5974 | 1.2645 | 31.9756 | 40.4346 | 25.2862 | 1.9973 | 31.8995 | 63.7137 | 15.9711 | 14.4642 | -0.1101 | 104.1483 | 235.1944 | 180.8217 | 232.876 | 161.3354
2012-10-30 19:08 | 28.5671 | 249.6895 | 195.2463 | 247.3555 | 175.6704 | -1.6066 | 1.2639 | 31.9599 | 40.3953 | 25.286 | 1.9964 | 31.884 | 63.6521 | 15.971 | 14.4805 | -0.1108 | 104.0474 | 235.209 | 180.7658 | 232.875 | 161.1899
2012-10-30 19:08 | 28.5656 | 249.5667 | 195.2527 | 247.4951 | 175.7274 | -1.6009 | 1.2636 | 31.9515 | 40.3746 | 25.2856 | 1.9958 | 31.8757 | 63.6183 | 15.9712 | 14.4912 | -0.1104 | 103.9929 | 235.0755 | 180.7615 | 233.0039 | 161.2362
2012-10-30 19:08 | 28.5634 | 249.6349 | 195.3291 | 247.2852 | 175.9516 | -1.6009 | 1.2635 | 31.9489 | 40.3675 | 25.286 | 1.9957 | 31.8732 | 63.6087 | 15.9711 | 14.4482 | -0.1104 | 103.9762 | 235.1867 | 180.8809 | 232.837 | 161.5034
2012-10-30 19:08 | 28.5634 | 249.5958 | 195.2858 | 247.4419 | 175.9165 | -1.5985 | 1.2635 | 31.9479 | 40.3658 | 25.2855 | 1.9956 | 31.8724 | 63.605 | 15.9712 | 14.4646 | -0.1102 | 103.9708 | 235.1312 | 180.8212 | 232.9773 | 161.4519
2012-10-30 19:08 | 28.5626 | 249.5911 | 195.3283 | 247.4686 | 175.9899 | -1.6049 | 1.2634 | 31.9474 | 40.3638 | 25.286 | 1.9956 | 31.8717 | 63.6026 | 15.9712 | 14.4719 | -0.1106 | 103.9664 | 235.1192 | 180.8564 | 232.9967 | 161.518
2012-10-30 19:08 | 28.5506 | 249.5713 | 195.3398 | 247.3038 | 175.974 | -1.6078 | 1.2636 | 31.9502 | 40.3708 | 25.286 | 1.9958 | 31.8745 | 63.6139 | 15.9711 | 14.4885 | -0.1108 | 103.9847 | 235.0828 | 180.8513 | 232.8153 | 161.4855
2012-10-30 19:08 | 28.5634 | 249.5292 | 195.3408 | 247.5634 | 175.9243 | -1.598 | 1.2635 | 31.949 | 40.3677 | 25.286 | 1.9957 | 31.8733 | 63.6085 | 15.9713 | 14.5053 | -0.1102 | 103.9762 | 235.0239 | 180.8355 | 233.0581 | 161.419
2012-10-30 19:08 | 28.5566 | 249.5616 | 195.3733 | 247.2589 | 176.0034 | -1.6049 | 1.2635 | 31.9488 | 40.367 | 25.2861 | 1.9956 | 31.8731 | 63.6074 | 15.9713 | 14.4618 | -0.1106 | 103.9744 | 235.0998 | 180.9115 | 232.7971 | 161.5416
2012-10-30 19:08 | 28.5619 | 249.5238 | 195.3668 | 247.3934 | 175.6652 | -1.5945 | 1.2636 | 31.9507 | 40.3723 | 25.2859 | 1.9958 | 31.875 | 63.6148 | 15.9714 | 14.4671 | -0.1099 | 103.9871 | 235.0567 | 180.8997 | 232.9263 | 161.1981
2012-10-30 19:09 | 28.5596 | 249.5176 | 195.3292 | 247.4812 | 175.6434 | -1.6014 | 1.2657 | 32.0045 | 40.508 | 25.2861 | 1.9991 | 31.9277 | 63.8252 | 15.9714 | 14.4934 | -0.1104 | 104.3332 | 235.0242 | 180.8358 | 232.9878 | 161.15
2012-10-30 19:09 | 28.5529 | 249.5305 | 195.3539 | 247.3257 | 175.4649 | -1.5962 | 1.2686 | 32.0793 | 40.6954 | 25.2874 | 2.0038 | 32.0036 | 64.1292 | 15.9713 | 14.4581 | -0.11 | 104.8246 | 235.0724 | 180.8958 | 232.8676 | 161.0068
2012-10-30 19:09 | 28.5589 | 249.5169 | 195.3403 | 247.3708 | 174.8778 | -1.5951 | 1.2691 | 32.0921 | 40.7282 | 25.2872 | 2.0046 | 32.0164 | 64.1805 | 15.9713 | 14.5294 | -0.11 | 104.9087 | 234.9875 | 180.8109 | 232.8414 | 160.3484
2012-10-30 19:09 | 28.5446 | 249.5104 | 195.322 | 247.337 | 175.4137 | -1.6003 | 1.2696 | 32.105 | 40.7606 | 25.2874 | 2.0054 | 32.0293 | 64.2323 | 15.9713 | 14.4213 | -0.1103 | 104.9929 | 235.0891 | 180.9007 | 232.9157 | 160.9924
2012-10-30 19:09 | 28.5596 | 249.592 | 195.3135 | 247.254 | 175.8151 | -1.5962 | 1.2701 | 32.1169 | 40.7909 | 25.2875 | 2.0062 | 32.0409 | 64.2789 | 15.9713 | 14.473 | -0.11 | 105.0697 | 235.119 | 180.8405 | 232.781 | 161.3421
2012-10-30 19:09 | 28.5454 | 249.5426 | 195.315 | 247.3573 | 175.6135 | -1.6026 | 1.2704 | 32.1257 | 40.8128 | 25.2877 | 2.0067 | 32.0493 | 64.3129 | 15.9713 | 14.4996 | -0.1105 | 105.1257 | 235.043 | 180.8154 | 232.8577 | 161.1139
2012-10-30 19:09 | 28.5559 | 249.5139 | 195.353 | 247.3326 | 175.8114 | -1.5985 | 1.2706 | 32.1305 | 40.8247 | 25.2878 | 2.007 | 32.0543 | 64.3332 | 15.9712 | 14.4897 | -0.1102 | 105.1579 | 235.0242 | 180.8633 | 232.8429 | 161.3217
2012-10-30 19:09 | 28.5514 | 249.5446 | 195.376 | 247.2106 | 176.0489 | -1.6014 | 1.2707 | 32.1336 | 40.8321 | 25.2882 | 2.0072 | 32.0575 | 64.3462 | 15.9711 | 14.4893 | -0.1104 | 105.1783 | 235.0553 | 180.8867 | 232.7213 | 161.5596
2012-10-30 19:09 | 28.5499 | 249.5314 | 195.4099 | 247.2874 | 175.7351 | -1.6026 | 1.2708 | 32.1356 | 40.8373 | 25.288 | 2.0073 | 32.0593 | 64.3537 | 15.9711 | 14.4551 | -0.1105 | 105.191 | 235.0763 | 180.9548 | 232.8323 | 161.28
2012-10-30 19:09 | 28.5559 | 249.5687 | 195.4552 | 247.3835 | 175.6201 | -1.6003 | 1.2708 | 32.1371 | 40.841 | 25.2881 | 2.0074 | 32.0608 | 64.3597 | 15.9711 | 14.4611 | -0.1103 | 105.2007 | 235.1076 | 180.9941 | 232.9224 | 161.159
2012-10-30 19:09 | 28.5529 | 249.5853 | 195.4011 | 247.1455 | 175.7537 | -1.6061 | 1.2709 | 32.1383 | 40.8443 | 25.288 | 2.0075 | 32.0622 | 64.3655 | 15.971 | 14.4908 | -0.1107 | 105.2097 | 235.0945 | 180.9103 | 232.6547 | 161.2629
2012-10-30 19:09 | 28.5604 | 249.5615 | 195.3929 | 247.388 | 175.6402 | -1.5985 | 1.2709 | 32.1393 | 40.8468 | 25.288 | 2.0076 | 32.0631 | 64.3688 | 15.9711 | 14.4534 | -0.1102 | 105.2156 | 235.1081 | 180.9395 | 232.9346 | 161.1868
2012-10-30 19:09 | 28.5559 | 249.5844 | 195.4394 | 247.3326 | 175.8582 | -1.5997 | 1.271 | 32.1401 | 40.8493 | 25.2878 | 2.0076 | 32.064 | 64.3726 | 15.9711 | 14.4693 | -0.1103 | 105.2219 | 235.1151 | 180.9701 | 232.8633 | 161.3889
2012-10-30 19:09 | 28.5649 | 249.5621 | 195.3738 | 247.228 | 175.7931 | -1.5985 | 1.271 | 32.1408 | 40.8507 | 25.288 | 2.0077 | 32.0647 | 64.3755 | 15.971 | 14.4661 | -0.1102 | 105.2262 | 235.096 | 180.9077 | 232.7619 | 161.327
2012-10-30 19:09 | 28.5596 | 249.5254 | 195.4392 | 247.4225 | 175.819 | -1.5974 | 1.271 | 32.141 | 40.8517 | 25.2877 | 2.0077 | 32.0649 | 64.3763 | 15.971 | 14.4485 | -0.1101 | 105.228 | 235.0769 | 180.9907 | 232.974 | 161.3705
2012-10-30 19:09 | 28.5506 | 249.6026 | 195.442 | 247.2333 | 176.0247 | -1.6014 | 1.271 | 32.1413 | 40.852 | 25.288 | 2.0077 | 32.0652 | 64.3781 | 15.9709 | 14.4273 | -0.1104 | 105.2301 | 235.1753 | 181.0147 | 232.806 | 161.5974
2012-10-30 19:09 | 28.5506 | 249.5869 | 195.4341 | 247.3077 | 175.7436 | -1.6014 | 1.271 | 32.1421 | 40.8541 | 25.2879 | 2.0078 | 32.0658 | 64.381 | 15.9708 | 14.4395 | -0.1104 | 105.2351 | 235.1474 | 180.9946 | 232.8682 | 161.3041
2012-10-30 19:09 | 28.5641 | 249.5613 | 195.4358 | 247.4074 | 175.6674 | -1.602 | 1.2711 | 32.1423 | 40.8546 | 25.2879 | 2.0078 | 32.0661 | 64.3813 | 15.971 | 14.4979 | -0.1104 | 105.2359 | 235.0634 | 180.9379 | 232.9095 | 161.1695
2012-10-30 19:09 | 28.5656 | 249.6411 | 195.4609 | 247.2365 | 175.8289 | -1.598 | 1.2711 | 32.1425 | 40.8551 | 25.2879 | 2.0078 | 32.0664 | 64.3825 | 15.971 | 14.4178 | -0.1102 | 105.2376 | 235.2233 | 181.0431 | 232.8187 | 161.4111
2012-10-30 19:09 | 28.5581 | 249.6101 | 195.363 | 247.3622 | 175.6536 | -1.6014 | 1.2711 | 32.1422 | 40.8546 | 25.2877 | 2.0078 | 32.066 | 64.3808 | 15.971 | 14.4225 | -0.1104 | 105.2353 | 235.1876 | 180.9405 | 232.9397 | 161.2311
2012-10-30 19:09 | 28.5678 | 249.5845 | 195.4553 | 247.368 | 175.9014 | -1.602 | 1.271 | 32.1418 | 40.8534 | 25.2879 | 2.0078 | 32.0657 | 64.3801 | 15.9709 | 14.4322 | -0.1104 | 105.2335 | 235.1523 | 181.0231 | 232.9358 | 161.4692
2012-10-30 19:09 | 28.5716 | 249.6275 | 195.4394 | 247.2112 | 175.9754 | -1.5985 | 1.271 | 32.1418 | 40.8533 | 25.2879 | 2.0078 | 32.0657 | 64.3805 | 15.9708 | 14.4156 | -0.1102 | 105.2337 | 235.2119 | 181.0238 | 232.7956 | 161.5598
2012-10-30 19:09 | 28.5716 | 249.5962 | 195.4237 | 247.3796 | 175.8622 | -1.5985 | 1.2711 | 32.1421 | 40.8543 | 25.2878 | 2.0078 | 32.0659 | 64.382 | 15.9706 | 14.436 | -0.1102 | 105.2363 | 235.1602 | 180.9877 | 232.9436 | 161.4262
2012-10-30 19:09 | 28.5783 | 249.6224 | 195.45 | 247.4412 | 175.9041 | -1.6009 | 1.271 | 32.1421 | 40.8541 | 25.288 | 2.0078 | 32.0661 | 64.3825 | 15.9707 | 14.4264 | -0.1104 | 105.2366 | 235.196 | 181.0236 | 233.0148 | 161.4777
2012-10-30 19:09 | 28.5686 | 249.6049 | 195.4325 | 247.2513 | 175.9295 | -1.5939 | 1.2711 | 32.1426 | 40.8552 | 25.288 | 2.0078 | 32.0664 | 64.3832 | 15.9708 | 14.3882 | -0.1099 | 105.2383 | 235.2167 | 181.0443 | 232.8631 | 161.5413
2012-10-30 19:09 | 28.5806 | 249.5856 | 195.4484 | 247.4004 | 175.8556 | -1.5957 | 1.2711 | 32.1423 | 40.8549 | 25.2877 | 2.0078 | 32.0661 | 64.3818 | 15.9708 | 14.4206 | -0.11 | 105.2367 | 235.165 | 181.0278 | 232.9798 | 161.435
2012-10-30 19:09 | 28.5843 | 249.5893 | 195.456 | 247.4002 | 175.9881 | -1.6049 | 1.2711 | 32.1425 | 40.8552 | 25.2879 | 2.0078 | 32.0664 | 64.3829 | 15.9709 | 14.4324 | -0.1106 | 105.2381 | 235.1569 | 181.0236 | 232.9678 | 161.5557
2012-10-30 19:09 | 28.5776 | 249.6413 | 195.4572 | 247.3229 | 175.9229 | -1.6049 | 1.2711 | 32.1427 | 40.8558 | 25.2879 | 2.0078 | 32.0664 | 64.384 | 15.9707 | 14.4788 | -0.1106 | 105.2398 | 235.1625 | 180.9784 | 232.8441 | 161.4441
2012-10-30 19:09 | 28.5836 | 249.616 | 195.3885 | 247.3799 | 175.761 | -1.6009 | 1.2711 | 32.1427 | 40.8557 | 25.2878 | 2.0078 | 32.0664 | 64.3838 | 15.9707 | 14.4562 | -0.1104 | 105.2395 | 235.1598 | 180.9323 | 232.9237 | 161.3048
2012-10-30 19:10 | 28.5963 | 249.5582 | 195.4602 | 247.4592 | 175.9923 | -1.5985 | 1.2711 | 32.1425 | 40.8554 | 25.2878 | 2.0078 | 32.0663 | 64.3835 | 15.9707 | 14.4607 | -0.1102 | 105.2389 | 235.0975 | 180.9995 | 232.9985 | 161.5316
2012-10-30 19:10 | 28.5821 | 249.6341 | 195.4774 | 247.2648 | 176.0913 | -1.5991 | 1.2711 | 32.1424 | 40.8555 | 25.2876 | 2.0078 | 32.0663 | 64.3839 | 15.9706 | 14.4425 | -0.1102 | 105.2394 | 235.1916 | 181.0349 | 232.8223 | 161.6488
2012-10-30 19:10 | 28.5881 | 249.6322 | 195.4481 | 247.4118 | 175.8904 | -1.6003 | 1.2711 | 32.1424 | 40.8554 | 25.2876 | 2.0078 | 32.0661 | 64.3828 | 15.9706 | 14.4689 | -0.1103 | 105.2382 | 235.1633 | 180.9792 | 232.9429 | 161.4215
2012-10-30 19:10 | 28.5888 | 249.6251 | 195.5038 | 247.4047 | 176.0551 | -1.6037 | 1.271 | 32.1423 | 40.8544 | 25.288 | 2.0078 | 32.0661 | 64.3827 | 15.9707 | 14.4859 | -0.1106 | 105.2371 | 235.1392 | 181.0179 | 232.9188 | 161.5692
2012-10-30 19:10 | 28.5888 | 249.629 | 195.5549 | 247.2754 | 176.1137 | -1.598 | 1.271 | 32.1422 | 40.8543 | 25.288 | 2.0078 | 32.066 | 64.3821 | 15.9707 | 14.4777 | -0.1102 | 105.2364 | 235.1513 | 181.0772 | 232.7977 | 161.636
2012-10-30 19:10 | 28.5873 | 249.6628 | 195.5612 | 247.4071 | 175.7491 | -1.6032 | 1.271 | 32.1419 | 40.8529 | 25.2883 | 2.0078 | 32.0657 | 64.3815 | 15.9706 | 14.464 | -0.1105 | 105.2344 | 235.1988 | 181.0972 | 232.9431 | 161.2851
2012-10-30 19:10 | 28.6023 | 249.6347 | 195.5801 | 247.4848 | 175.7836 | -1.598 | 1.271 | 32.1415 | 40.8528 | 25.2878 | 2.0078 | 32.0654 | 64.3802 | 15.9706 | 14.479 | -0.1102 | 105.233 | 235.1557 | 181.1011 | 233.0058 | 161.3046
2012-10-30 19:10 | 28.5911 | 249.647 | 195.5296 | 247.3051 | 175.7646 | -1.6043 | 1.271 | 32.1415 | 40.8526 | 25.2879 | 2.0078 | 32.0653 | 64.3797 | 15.9706 | 14.4882 | -0.1106 | 105.2323 | 235.1588 | 181.0414 | 232.8169 | 161.2764
2012-10-30 19:10 | 28.6023 | 249.6308 | 195.478 | 247.4025 | 175.8344 | -1.6003 | 1.271 | 32.1413 | 40.8521 | 25.2879 | 2.0078 | 32.0651 | 64.379 | 15.9705 | 14.4586 | -0.1103 | 105.2311 | 235.1722 | 181.0194 | 232.9439 | 161.3758
2012-10-30 19:10 | 28.5993 | 249.6552 | 195.475 | 247.5013 | 175.8275 | -1.6026 | 1.271 | 32.1411 | 40.8518 | 25.2878 | 2.0077 | 32.0649 | 64.3781 | 15.9706 | 14.4801 | -0.1105 | 105.2299 | 235.1751 | 180.9949 | 233.0212 | 161.3474
2012-10-30 19:10 | 28.6001 | 249.656 | 195.5347 | 247.369 | 175.9805 | -1.6026 | 1.271 | 32.1407 | 40.8508 | 25.2877 | 2.0077 | 32.0644 | 64.376 | 15.9706 | 14.4564 | -0.1105 | 105.2268 | 235.1996 | 181.0783 | 232.9126 | 161.5241
2012-10-30 19:10 | 28.6053 | 249.6416 | 195.4613 | 247.3742 | 175.7047 | -1.6055 | 1.271 | 32.1403 | 40.8495 | 25.2879 | 2.0077 | 32.064 | 64.3747 | 15.9705 | 14.4453 | -0.1107 | 105.2242 | 235.1963 | 181.016 | 232.9289 | 161.2594
2012-10-30 19:10 | 28.6158 | 249.6599 | 195.5111 | 247.5453 | 175.7074 | -1.6003 | 1.2709 | 32.14 | 40.848 | 25.2883 | 2.0077 | 32.0639 | 64.374 | 15.9706 | 14.4517 | -0.1103 | 105.2221 | 235.2082 | 181.0594 | 233.0936 | 161.2557
2012-10-30 19:10 | 28.6188 | 249.6629 | 195.5063 | 247.3171 | 175.9094 | -1.5962 | 1.271 | 32.1401 | 40.8491 | 25.2879 | 2.0077 | 32.064 | 64.3752 | 15.9705 | 14.4588 | -0.11 | 105.2243 | 235.2041 | 181.0475 | 232.8583 | 161.4506
2012-10-30 19:10 | 28.6128 | 249.6452 | 195.4335 | 247.4365 | 175.759 | -1.5997 | 1.271 | 32.14 | 40.8486 | 25.2879 | 2.0077 | 32.0638 | 64.3733 | 15.9707 | 14.4405 | -0.1103 | 105.2219 | 235.2047 | 180.993 | 232.996 | 161.3185
2012-10-30 19:10 | 28.6135 | 249.595 | 195.4735 | 247.4725 | 175.7012 | -1.6049 | 1.271 | 32.1396 | 40.8482 | 25.2877 | 2.0077 | 32.0636 | 64.3734 | 15.9705 | 14.4616 | -0.1106 | 105.2216 | 235.1334 | 181.0119 | 233.0109 | 161.2396
2012-10-30 19:10 | 28.6143 | 249.7054 | 195.4585 | 247.3949 | 175.8932 | -1.6113 | 1.271 | 32.1396 | 40.848 | 25.2878 | 2.0077 | 32.0636 | 64.3735 | 15.9704 | 14.4665 | -0.1111 | 105.2215 | 235.2389 | 180.992 | 232.9284 | 161.4267
2012-10-30 19:10 | 28.6248 | 249.6924 | 195.4376 | 247.3662 | 175.7944 | -1.6009 | 1.2709 | 32.139 | 40.846 | 25.2881 | 2.0076 | 32.0629 | 64.3707 | 15.9705 | 14.4648 | -0.1104 | 105.2167 | 235.2276 | 180.9728 | 232.9014 | 161.3296
2012-10-30 19:10 | 28.624 | 249.629 | 195.5075 | 247.577 | 175.7975 | -1.5985 | 1.2709 | 32.1387 | 40.8457 | 25.2878 | 2.0076 | 32.0626 | 64.3699 | 15.9703 | 14.4272 | -0.1102 | 105.2156 | 235.2018 | 181.0803 | 233.1498 | 161.3703
2012-10-30 19:10 | 28.6203 | 249.6918 | 195.4606 | 247.4087 | 175.7509 | -1.6026 | 1.2709 | 32.1383 | 40.8443 | 25.288 | 2.0076 | 32.0621 | 64.368 | 15.9704 | 14.4358 | -0.1105 | 105.2123 | 235.256 | 181.0248 | 232.9729 | 161.3151
2012-10-30 19:10 | 28.6263 | 249.6665 | 195.4705 | 247.3129 | 175.8076 | -1.6043 | 1.2709 | 32.1377 | 40.8431 | 25.2878 | 2.0076 | 32.0615 | 64.3655 | 15.9703 | 14.4867 | -0.1106 | 105.2087 | 235.1798 | 180.9838 | 232.8262 | 161.3209
2012-10-30 19:10 | 28.6293 | 249.6539 | 195.4185 | 247.5666 | 175.7209 | -1.6043 | 1.2709 | 32.1374 | 40.8424 | 25.2878 | 2.0076 | 32.0613 | 64.3648 | 15.9703 | 14.4244 | -0.1106 | 105.2072 | 235.2295 | 180.9941 | 233.1422 | 161.2965
2012-10-30 19:10 | 28.6338 | 249.6819 | 195.478 | 247.4222 | 175.983 | -1.6026 | 1.2709 | 32.1371 | 40.8416 | 25.2878 | 2.0075 | 32.0609 | 64.3635 | 15.9702 | 14.4166 | -0.1105 | 105.2051 | 235.2653 | 181.0614 | 233.0056 | 161.5664
2012-10-30 19:10 | 28.6338 | 249.6466 | 195.4387 | 247.387 | 175.7293 | -1.6032 | 1.2708 | 32.1366 | 40.84 | 25.2879 | 2.0075 | 32.0604 | 64.3611 | 15.9704 | 14.437 | -0.1105 | 105.2012 | 235.2096 | 181.0017 | 232.95 | 161.2923
2012-10-30 19:10 | 28.6278 | 249.6798 | 195.4445 | 247.4672 | 175.5788 | -1.6032 | 1.2708 | 32.1365 | 40.84 | 25.2878 | 2.0075 | 32.0605 | 64.3612 | 15.9704 | 14.4229 | -0.1105 | 105.2012 | 235.2569 | 181.0216 | 233.0443 | 161.1559
2012-10-30 19:10 | 28.642 | 249.7292 | 195.4391 | 247.4931 | 175.675 | -1.6032 | 1.2708 | 32.1365 | 40.84 | 25.2878 | 2.0075 | 32.0603 | 64.3615 | 15.9702 | 14.4534 | -0.1105 | 105.2016 | 235.2758 | 180.9857 | 233.0397 | 161.2216
2012-10-30 19:10 | 28.6465 | 249.7376 | 195.4397 | 247.3879 | 175.781 | -1.5985 | 1.2708 | 32.1362 | 40.8396 | 25.2876 | 2.0075 | 32.06 | 64.36 | 15.9702 | 14.462 | -0.1102 | 105.1996 | 235.2756 | 180.9777 | 232.9259 | 161.319
2012-10-30 19:10 | 28.6375 | 249.6699 | 195.3756 | 247.5278 | 175.5105 | -1.6072 | 1.2708 | 32.136 | 40.8388 | 25.2877 | 2.0075 | 32.0598 | 64.3592 | 15.9702 | 14.4244 | -0.1108 | 105.198 | 235.2455 | 180.9512 | 233.1034 | 161.0861
2012-10-30 19:11 | 28.6495 | 249.6897 | 195.4152 | 247.5163 | 175.5264 | -1.5928 | 1.2708 | 32.1354 | 40.8372 | 25.2878 | 2.0074 | 32.0592 | 64.3563 | 15.9703 | 14.4201 | -0.1098 | 105.1935 | 235.2696 | 180.9951 | 233.0962 | 161.1063
2012-10-30 19:11 | 28.642 | 249.6861 | 195.3409 | 247.403 | 175.5852 | -1.6026 | 1.2708 | 32.1347 | 40.8355 | 25.2878 | 2.0074 | 32.0585 | 64.3538 | 15.9703 | 14.4371 | -0.1105 | 105.1892 | 235.249 | 180.9038 | 232.9659 | 161.1481
2012-10-30 19:11 | 28.6532 | 249.7052 | 195.2892 | 247.4495 | 175.6628 | -1.6014 | 1.2708 | 32.1346 | 40.8351 | 25.2878 | 2.0074 | 32.0584 | 64.3535 | 15.9702 | 14.4524 | -0.1104 | 105.1886 | 235.2528 | 180.8368 | 232.9971 | 161.2104
2012-10-30 19:11 | 28.648 | 249.6687 | 195.3233 | 247.5226 | 175.7005 | -1.5991 | 1.2707 | 32.1344 | 40.8348 | 25.2878 | 2.0074 | 32.0581 | 64.3524 | 15.9702 | 14.4023 | -0.1102 | 105.1872 | 235.2664 | 180.921 | 233.1203 | 161.2982
2012-10-30 19:11 | 28.6555 | 249.7075 | 195.3622 | 247.3891 | 175.7041 | -1.602 | 1.2702 | 32.1198 | 40.7977 | 25.2877 | 2.0067 | 32.0478 | 64.311 | 15.9702 | 14.4302 | -0.1104 | 105.1086 | 235.2773 | 180.932 | 232.9589 | 161.2739
2012-10-30 19:11 | 28.6592 | 249.6838 | 195.3266 | 247.4516 | 175.5829 | -1.5985 | 1.2644 | 31.9722 | 40.4258 | 25.2863 | 1.9971 | 31.8952 | 63.6995 | 15.9703 | 14.3849 | -0.1102 | 104.1252 | 235.2989 | 180.9417 | 233.0667 | 161.198
2012-10-30 19:11 | 28.6585 | 249.6557 | 195.2945 | 247.5175 | 175.5822 | -1.6049 | 1.2642 | 31.9678 | 40.415 | 25.2862 | 1.997 | 31.8924 | 63.6879 | 15.9705 | 14.425 | -0.1106 | 104.1029 | 235.2307 | 180.8695 | 233.0925 | 161.1572
2012-10-30 19:11 | 28.6645 | 249.6734 | 195.3044 | 247.5274 | 175.678 | -1.5951 | 1.2638 | 31.9569 | 40.3872 | 25.2863 | 1.9963 | 31.8812 | 63.6432 | 15.9704 | 14.4106 | -0.11 | 104.0305 | 235.2628 | 180.8938 | 233.1168 | 161.2674
2012-10-30 19:11 | 28.6682 | 249.7045 | 195.316 | 247.4096 | 175.7676 | -1.5962 | 1.2636 | 31.9504 | 40.3711 | 25.2861 | 1.9959 | 31.8746 | 63.6172 | 15.9704 | 14.398 | -0.11 | 103.9882 | 235.3065 | 180.918 | 233.0116 | 161.3696
2012-10-30 19:11 | 28.6742 | 249.6596 | 195.263 | 247.494 | 175.4262 | -1.5951 | 1.2635 | 31.948 | 40.3651 | 25.286 | 1.9957 | 31.8724 | 63.6072 | 15.9706 | 14.4448 | -0.11 | 103.9723 | 235.2148 | 180.8182 | 233.0492 | 160.9814
2012-10-30 19:11 | 28.6705 | 249.6011 | 195.2947 | 247.4864 | 175.282 | -1.5991 | 1.2634 | 31.9465 | 40.3614 | 25.286 | 1.9956 | 31.8708 | 63.6016 | 15.9705 | 14.4493 | -0.1102 | 103.9631 | 235.1518 | 180.8454 | 233.0371 | 160.8327
2012-10-30 19:11 | 28.678 | 249.6086 | 195.2472 | 247.4351 | 175.4066 | -1.602 | 1.2635 | 31.9479 | 40.365 | 25.2859 | 1.9957 | 31.8722 | 63.6069 | 15.9705 | 14.3956 | -0.1104 | 103.9719 | 235.213 | 180.8516 | 233.0395 | 161.011
2012-10-30 19:11 | 28.6697 | 249.5886 | 195.2585 | 247.3837 | 175.371 | -1.6037 | 1.2635 | 31.9499 | 40.3695 | 25.2862 | 1.9958 | 31.8742 | 63.6146 | 15.9706 | 14.424 | -0.1106 | 103.9841 | 235.1646 | 180.8345 | 232.9597 | 160.947
2012-10-30 19:11 | 28.6697 | 249.6081 | 195.2271 | 247.5012 | 175.1251 | -1.6032 | 1.2636 | 31.9513 | 40.3736 | 25.286 | 1.9959 | 31.8755 | 63.6197 | 15.9707 | 14.4077 | -0.1105 | 103.9932 | 235.2004 | 180.8194 | 233.0935 | 160.7174
2012-10-30 19:11 | 28.6817 | 249.5614 | 195.2116 | 247.4584 | 174.9498 | -1.6003 | 1.2636 | 31.9524 | 40.3762 | 25.2861 | 1.9959 | 31.8767 | 63.6236 | 15.9708 | 14.4523 | -0.1103 | 103.9998 | 235.1091 | 180.7593 | 233.0061 | 160.4975
2012-10-30 19:11 | 28.6712 | 249.5587 | 195.1382 | 247.4205 | 174.7013 | -1.6009 | 1.267 | 32.0381 | 40.5921 | 25.2868 | 2.0012 | 31.96 | 63.9573 | 15.9707 | 14.4296 | -0.1104 | 104.5494 | 235.1291 | 180.7086 | 232.9909 | 160.2717
2012-10-30 19:11 | 28.669 | 249.5565 | 195.0496 | 247.3203 | 174.8474 | -1.6026 | 1.2689 | 32.0874 | 40.7158 | 25.2875 | 2.0044 | 32.0113 | 64.1624 | 15.9708 | 14.4355 | -0.1105 | 104.8782 | 235.121 | 180.6141 | 232.8848 | 160.4119
2012-10-30 19:11 | 28.6772 | 249.553 | 195.0382 | 247.4656 | 175.0234 | -1.6026 | 1.2694 | 32.101 | 40.7507 | 25.2874 | 2.0052 | 32.025 | 64.2179 | 15.9706 | 14.456 | -0.1105 | 104.9685 | 235.097 | 180.5822 | 233.0096 | 160.5674
2012-10-30 19:11 | 28.6697 | 249.569 | 195.0817 | 247.3602 | 175.3827 | -1.5939 | 1.2699 | 32.1135 | 40.782 | 25.2876 | 2.006 | 32.0373 | 64.2675 | 15.9706 | 14.4321 | -0.1099 | 105.0495 | 235.1369 | 180.6496 | 232.9281 | 160.9506
2012-10-30 19:11 | 28.6667 | 249.5582 | 195.0316 | 247.3141 | 175.2704 | -1.6003 | 1.2703 | 32.1228 | 40.8051 | 25.2879 | 2.0066 | 32.0465 | 64.3045 | 15.9706 | 14.4577 | -0.1103 | 105.1095 | 235.1005 | 180.5739 | 232.8564 | 160.8127
2012-10-30 19:12 | 28.6847 | 249.584 | 195.1832 | 247.3595 | 175.5421 | -1.6003 | 1.2709 | 32.1383 | 40.8445 | 25.2879 | 2.0076 | 32.0621 | 64.3676 | 15.9704 | 14.439 | -0.1103 | 105.2121 | 235.145 | 180.7442 | 232.9205 | 161.1031
2012-10-30 19:12 | 28.6892 | 249.6042 | 195.2663 | 247.2426 | 175.5271 | -1.6043 | 1.271 | 32.1422 | 40.8541 | 25.288 | 2.0079 | 32.066 | 64.3841 | 15.9702 | 14.4272 | -0.1106 | 105.2382 | 235.177 | 180.8391 | 232.8154 | 161.0999
2012-10-30 19:12 | 28.6869 | 249.594 | 195.2679 | 247.3696 | 175.5521 | -1.5991 | 1.2711 | 32.1428 | 40.8554 | 25.2882 | 2.0079 | 32.0666 | 64.3862 | 15.9703 | 14.3881 | -0.1102 | 105.2415 | 235.2059 | 180.8798 | 232.9815 | 161.164
2012-10-30 19:12 | 28.6937 | 249.5812 | 195.2747 | 247.4038 | 175.555 | -1.6009 | 1.2711 | 32.1426 | 40.8551 | 25.2882 | 2.0079 | 32.0665 | 64.3858 | 15.9703 | 14.4357 | -0.1104 | 105.2409 | 235.1455 | 180.839 | 232.9681 | 161.1193
2012-10-30 19:12 | 28.6899 | 249.6283 | 195.3062 | 247.4078 | 175.6683 | -1.5985 | 1.2711 | 32.143 | 40.8566 | 25.2877 | 2.0079 | 32.0667 | 64.3864 | 15.9703 | 14.4401 | -0.1102 | 105.243 | 235.1882 | 180.8661 | 232.9677 | 161.2282
2012-10-30 19:12 | 28.6877 | 249.5713 | 195.3355 | 247.3508 | 175.6114 | -1.6003 | 1.2711 | 32.143 | 40.8559 | 25.2882 | 2.0079 | 32.0667 | 64.3871 | 15.9702 | 14.442 | -0.1103 | 105.2431 | 235.1293 | 180.8935 | 232.9088 | 161.1694
2012-10-30 19:12 | 28.6974 | 249.6163 | 195.2745 | 247.3291 | 175.6289 | -1.6037 | 1.2711 | 32.1431 | 40.8566 | 25.288 | 2.0079 | 32.0669 | 64.3875 | 15.9703 | 14.4802 | -0.1106 | 105.244 | 235.1361 | 180.7943 | 232.8489 | 161.1487
2012-10-30 19:12 | 28.6997 | 249.5912 | 195.2728 | 247.3863 | 175.6195 | -1.5991 | 1.2711 | 32.1428 | 40.8556 | 25.2881 | 2.0079 | 32.0666 | 64.3861 | 15.9703 | 14.4336 | -0.1102 | 105.2417 | 235.1576 | 180.8392 | 232.9527 | 161.1859
2012-10-30 19:12 | 28.7034 | 249.6262 | 195.2608 | 247.4644 | 175.553 | -1.5962 | 1.271 | 32.1425 | 40.8546 | 25.2882 | 2.0079 | 32.0663 | 64.3852 | 15.9703 | 14.4699 | -0.11 | 105.2397 | 235.1563 | 180.7909 | 232.9945 | 161.0831
2012-10-30 19:12 | 28.6982 | 249.5779 | 195.2674 | 247.3574 | 175.54 | -1.5985 | 1.2711 | 32.1427 | 40.8556 | 25.2879 | 2.0079 | 32.0665 | 64.3865 | 15.9701 | 14.4525 | -0.1102 | 105.2422 | 235.1254 | 180.8149 | 232.9049 | 161.0875
2012-10-30 19:12 | 28.7034 | 249.5988 | 195.2765 | 247.3273 | 175.3344 | -1.598 | 1.2711 | 32.1427 | 40.8558 | 25.2879 | 2.0079 | 32.0664 | 64.3859 | 15.9702 | 14.4577 | -0.1102 | 105.2417 | 235.1411 | 180.8188 | 232.8696 | 160.8767
2012-10-30 19:12 | 28.6929 | 249.5765 | 195.2189 | 247.4304 | 175.2927 | -1.5968 | 1.2711 | 32.143 | 40.856 | 25.2882 | 2.0079 | 32.0669 | 64.3878 | 15.9702 | 14.4553 | -0.1101 | 105.2438 | 235.1212 | 180.7636 | 232.9751 | 160.8374
2012-10-30 19:12 | 28.7057 | 249.6324 | 195.2317 | 247.404 | 175.5045 | -1.5933 | 1.2711 | 32.1428 | 40.8552 | 25.2883 | 2.0079 | 32.0665 | 64.3866 | 15.9701 | 14.4437 | -0.1098 | 105.2418 | 235.1887 | 180.788 | 232.9603 | 161.0608
2012-10-30 19:12 | 28.7079 | 249.6307 | 195.2299 | 247.3827 | 175.6004 | -1.5997 | 1.2711 | 32.143 | 40.8562 | 25.288 | 2.0079 | 32.0668 | 64.3884 | 15.97 | 14.4826 | -0.1103 | 105.2446 | 235.1481 | 180.7473 | 232.9001 | 161.1178
2012-10-30 19:12 | 28.7109 | 249.575 | 195.2801 | 247.3387 | 175.4902 | -1.6055 | 1.2711 | 32.1425 | 40.8551 | 25.288 | 2.0079 | 32.0664 | 64.3863 | 15.9701 | 14.4489 | -0.1107 | 105.2414 | 235.1261 | 180.8312 | 232.8898 | 161.0413
2012-10-30 19:12 | 28.7072 | 249.5869 | 195.2175 | 247.4761 | 175.4748 | -1.598 | 1.2711 | 32.1427 | 40.8555 | 25.288 | 2.0079 | 32.0664 | 64.3869 | 15.9699 | 14.4411 | -0.1102 | 105.2424 | 235.1458 | 180.7764 | 233.035 | 161.0337
2012-10-30 19:12 | 28.7049 | 249.5611 | 195.2269 | 247.4228 | 175.4569 | -1.6014 | 1.2711 | 32.1425 | 40.8554 | 25.2878 | 2.0079 | 32.0664 | 64.3864 | 15.9701 | 14.4551 | -0.1104 | 105.2417 | 235.106 | 180.7718 | 232.9677 | 161.0018
2012-10-30 19:12 | 28.7064 | 249.6057 | 195.2717 | 247.3812 | 175.5404 | -1.5997 | 1.271 | 32.1422 | 40.8536 | 25.2883 | 2.0079 | 32.066 | 64.3847 | 15.97 | 14.4892 | -0.1103 | 105.2383 | 235.1165 | 180.7825 | 232.892 | 161.0512
2012-10-30 19:12 | 28.7132 | 249.6438 | 195.2549 | 247.341 | 175.6721 | -1.5968 | 1.271 | 32.1422 | 40.8542 | 25.2879 | 2.0079 | 32.0659 | 64.3844 | 15.97 | 14.4756 | -0.1101 | 105.2386 | 235.1682 | 180.7793 | 232.8654 | 161.1965
2012-10-30 19:12 | 28.7087 | 249.5884 | 195.2543 | 247.4736 | 175.5973 | -1.602 | 1.271 | 32.1419 | 40.8536 | 25.2879 | 2.0079 | 32.0658 | 64.3842 | 15.97 | 14.5119 | -0.1104 | 105.2378 | 235.0765 | 180.7424 | 232.9617 | 161.0854
2012-10-30 19:12 | 28.7117 | 249.6384 | 195.277 | 247.4453 | 175.694 | -1.6032 | 1.271 | 32.1418 | 40.8533 | 25.2879 | 2.0079 | 32.0656 | 64.3838 | 15.9699 | 14.4945 | -0.1105 | 105.2371 | 235.1439 | 180.7825 | 232.9508 | 161.1995
2012-10-30 19:12 | 28.7132 | 249.6203 | 195.2667 | 247.3841 | 175.6486 | -1.5939 | 1.271 | 32.1413 | 40.8522 | 25.2879 | 2.0078 | 32.0651 | 64.3818 | 15.9699 | 14.4552 | -0.1099 | 105.234 | 235.1651 | 180.8115 | 232.9289 | 161.1934
2012-10-30 19:13 | 28.7154 | 249.6656 | 195.2925 | 247.3628 | 175.604 | -1.5985 | 1.271 | 32.1408 | 40.8508 | 25.288 | 2.0078 | 32.0646 | 64.3804 | 15.9698 | 14.4697 | -0.1102 | 105.2311 | 235.1959 | 180.8228 | 232.8931 | 161.1343
2012-10-30 19:13 | 28.7147 | 249.6062 | 195.3193 | 247.4522 | 175.654 | -1.5991 | 1.271 | 32.1407 | 40.8509 | 25.2877 | 2.0078 | 32.0645 | 64.3798 | 15.9698 | 14.5139 | -0.1102 | 105.2307 | 235.0923 | 180.8054 | 232.9383 | 161.1401
2012-10-30 19:13 | 28.7117 | 249.5836 | 195.3084 | 247.5041 | 175.7057 | -1.5997 | 1.271 | 32.1406 | 40.8505 | 25.2877 | 2.0078 | 32.0644 | 64.3794 | 15.9698 | 14.4864 | -0.1103 | 105.2299 | 235.0972 | 180.822 | 233.0177 | 161.2193
2012-10-30 19:13 | 28.7177 | 249.6287 | 195.2908 | 247.3455 | 175.5907 | -1.5991 | 1.271 | 32.1401 | 40.8489 | 25.288 | 2.0078 | 32.064 | 64.378 | 15.9698 | 14.4761 | -0.1102 | 105.2269 | 235.1526 | 180.8147 | 232.8694 | 161.1146
2012-10-30 19:13 | 28.7162 | 249.6116 | 195.3129 | 247.4185 | 175.4799 | -1.6055 | 1.2709 | 32.1398 | 40.848 | 25.2881 | 2.0078 | 32.0636 | 64.3764 | 15.9697 | 14.4868 | -0.1107 | 105.2244 | 235.1248 | 180.8261 | 232.9317 | 160.9931
2012-10-30 19:13 | 28.7147 | 249.6375 | 195.2682 | 247.4953 | 175.4979 | -1.5962 | 1.271 | 32.1395 | 40.8477 | 25.2877 | 2.0077 | 32.0633 | 64.3749 | 15.9698 | 14.4323 | -0.11 | 105.2226 | 235.2052 | 180.8359 | 233.063 | 161.0656
2012-10-30 19:13 | 28.7132 | 249.6438 | 195.2863 | 247.5056 | 175.6096 | -1.6089 | 1.2709 | 32.1393 | 40.8468 | 25.288 | 2.0077 | 32.0631 | 64.3746 | 15.9697 | 14.4593 | -0.1109 | 105.2213 | 235.1845 | 180.827 | 233.0463 | 161.1503
2012-10-30 19:13 | 28.7109 | 249.6376 | 195.2958 | 247.3583 | 175.7322 | -1.598 | 1.2709 | 32.139 | 40.8465 | 25.2878 | 2.0077 | 32.0628 | 64.3733 | 15.9697 | 14.4366 | -0.1102 | 105.2198 | 235.201 | 180.8592 | 232.9217 | 161.2956
2012-10-30 19:13 | 28.7117 | 249.6658 | 195.2888 | 247.4414 | 175.6706 | -1.5991 | 1.2709 | 32.1388 | 40.8461 | 25.2876 | 2.0077 | 32.0627 | 64.373 | 15.9697 | 14.4497 | -0.1102 | 105.2191 | 235.2161 | 180.8391 | 232.9917 | 161.2209
2012-10-30 19:13 | 28.7132 | 249.6556 | 195.2863 | 247.486 | 175.5823 | -1.6066 | 1.2709 | 32.1381 | 40.8441 | 25.2878 | 2.0077 | 32.0619 | 64.3695 | 15.9697 | 14.4675 | -0.1108 | 105.2136 | 235.1881 | 180.8188 | 233.0185 | 161.1148
2012-10-30 19:13 | 28.7132 | 249.6556 | 195.3649 | 247.5095 | 175.6603 | -1.6026 | 1.2709 | 32.1381 | 40.8439 | 25.2879 | 2.0077 | 32.0618 | 64.3692 | 15.9697 | 14.4756 | -0.1105 | 105.2131 | 235.18 | 180.8893 | 233.0339 | 161.1847
2012-10-30 19:13 | 28.7132 | 249.6869 | 195.3531 | 247.3802 | 175.7618 | -1.6003 | 1.2709 | 32.1379 | 40.8435 | 25.2879 | 2.0077 | 32.0616 | 64.3691 | 15.9696 | 14.4471 | -0.1103 | 105.2126 | 235.2398 | 180.906 | 232.9331 | 161.3147
2012-10-30 19:13 | 28.7162 | 249.682 | 195.3286 | 247.5242 | 175.5189 | -1.6055 | 1.2709 | 32.1373 | 40.8417 | 25.288 | 2.0076 | 32.061 | 64.366 | 15.9698 | 14.4746 | -0.1107 | 105.2076 | 235.2074 | 180.854 | 233.0496 | 161.0443
2012-10-30 19:13 | 28.7079 | 249.6581 | 195.3321 | 247.5277 | 175.6082 | -1.6032 | 1.2708 | 32.1363 | 40.8391 | 25.288 | 2.0076 | 32.0601 | 64.3625 | 15.9697 | 14.4785 | -0.1105 | 105.2017 | 235.1796 | 180.8536 | 233.0492 | 161.1297
2012-10-30 19:13 | 28.7154 | 249.6852 | 195.3435 | 247.449 | 175.6586 | -1.5968 | 1.2708 | 32.1357 | 40.8379 | 25.2879 | 2.0075 | 32.0596 | 64.3611 | 15.9695 | 14.5023 | -0.1101 | 105.1989 | 235.1829 | 180.8412 | 232.9467 | 161.1563
2012-10-30 19:13 | 28.7094 | 249.6831 | 195.3218 | 247.3568 | 175.6097 | -1.5991 | 1.2708 | 32.1355 | 40.8371 | 25.288 | 2.0075 | 32.0594 | 64.3596 | 15.9697 | 14.4841 | -0.1102 | 105.1967 | 235.199 | 180.8377 | 232.8727 | 161.1256
2012-10-30 19:13 | 28.7117 | 249.658 | 195.3713 | 247.5589 | 175.5769 | -1.6026 | 1.2708 | 32.1352 | 40.8364 | 25.2879 | 2.0075 | 32.059 | 64.3579 | 15.9697 | 14.4905 | -0.1105 | 105.1944 | 235.1675 | 180.8808 | 233.0684 | 161.0864
2012-10-30 19:13 | 28.7199 | 249.6897 | 195.3323 | 247.5044 | 175.6007 | -1.5962 | 1.2665 | 32.0253 | 40.5594 | 25.2869 | 2.0006 | 31.949 | 63.9174 | 15.9697 | 14.515 | -0.11 | 104.4767 | 235.1747 | 180.8173 | 232.9894 | 161.0857
2012-10-30 19:13 | 28.7139 | 249.6484 | 195.2988 | 247.5023 | 175.5518 | -1.6049 | 1.2645 | 31.9738 | 40.4295 | 25.2866 | 1.9974 | 31.898 | 63.7132 | 15.9698 | 14.4682 | -0.1106 | 104.1427 | 235.1802 | 180.8306 | 233.0341 | 161.0836
2012-10-30 19:13 | 28.7147 | 249.6571 | 195.2957 | 247.5031 | 175.7243 | -1.6014 | 1.2638 | 31.9572 | 40.388 | 25.2862 | 1.9964 | 31.8816 | 63.6473 | 15.9698 | 14.4567 | -0.1104 | 104.0353 | 235.2004 | 180.839 | 233.0464 | 161.2676
2012-10-30 19:13 | 28.7094 | 249.6126 | 195.2511 | 247.3921 | 175.5317 | -1.6014 | 1.2636 | 31.9502 | 40.3707 | 25.286 | 1.9959 | 31.8745 | 63.6188 | 15.9698 | 14.4474 | -0.1104 | 103.9895 | 235.1652 | 180.8037 | 232.9447 | 161.0843
2012-10-30 19:13 | 28.7049 | 249.565 | 195.2348 | 247.4424 | 175.4374 | -1.5997 | 1.2635 | 31.9484 | 40.3659 | 25.2861 | 1.9958 | 31.8727 | 63.6114 | 15.9699 | 14.4877 | -0.1103 | 103.9773 | 235.0773 | 180.7471 | 232.9547 | 160.9497
2012-10-30 19:13 | 28.7162 | 249.592 | 195.2265 | 247.5477 | 175.1404 | -1.6009 | 1.2635 | 31.9491 | 40.3679 | 25.286 | 1.9958 | 31.8734 | 63.6137 | 15.97 | 14.4705 | -0.1104 | 103.9816 | 235.1215 | 180.756 | 233.0772 | 160.6699
2012-10-30 19:14 | 28.7154 | 249.5716 | 195.2296 | 247.449 | 175.4128 | -1.5997 | 1.2635 | 31.9492 | 40.3679 | 25.2862 | 1.9958 | 31.8735 | 63.6141 | 15.97 | 14.5227 | -0.1103 | 103.982 | 235.0489 | 180.7069 | 232.9263 | 160.8901
2012-10-30 19:14 | 28.7154 | 249.6147 | 195.2335 | 247.3471 | 175.604 | -1.5974 | 1.2636 | 31.9507 | 40.3715 | 25.2864 | 1.9959 | 31.8748 | 63.6195 | 15.97 | 14.4819 | -0.1101 | 103.991 | 235.1328 | 180.7516 | 232.8652 | 161.1221
2012-10-30 19:14 | 28.7064 | 249.5744 | 195.1813 | 247.3969 | 175.3843 | -1.6043 | 1.2636 | 31.9529 | 40.3771 | 25.2864 | 1.996 | 31.8771 | 63.6282 | 15.9701 | 14.5137 | -0.1106 | 104.0053 | 235.0607 | 180.6676 | 232.8832 | 160.8706
2012-10-30 19:14 | 28.7109 | 249.5515 | 195.1465 | 247.4562 | 175.0727 | -1.602 | 1.266 | 32.0128 | 40.5284 | 25.2865 | 1.9997 | 31.9359 | 63.8633 | 15.9701 | 14.4978 | -0.1104 | 104.3917 | 235.0537 | 180.6487 | 232.9584 | 160.5749
2012-10-30 19:14 | 28.7079 | 249.5406 | 195.116 | 247.4611 | 175.4638 | -1.6037 | 1.2687 | 32.0818 | 40.7013 | 25.2877 | 2.0041 | 32.0054 | 64.1422 | 15.97 | 14.4948 | -0.1106 | 104.8435 | 235.0458 | 180.6212 | 232.9663 | 160.969
2012-10-30 19:14 | 28.7072 | 249.5595 | 195.2135 | 247.3703 | 175.7675 | -1.6003 | 1.2693 | 32.0986 | 40.7441 | 25.2877 | 2.0051 | 32.0221 | 64.209 | 15.97 | 14.5308 | -0.1103 | 104.9531 | 235.0287 | 180.6827 | 232.8395 | 161.2367
2012-10-30 19:14 | 28.7057 | 249.5698 | 195.2552 | 247.3061 | 175.6724 | -1.5991 | 1.2699 | 32.1132 | 40.7807 | 25.2878 | 2.0061 | 32.0369 | 64.2683 | 15.97 | 14.5171 | -0.1102 | 105.049 | 235.0527 | 180.7381 | 232.789 | 161.1553
2012-10-30 19:14 | 28.7049 | 249.6003 | 195.2859 | 247.3758 | 175.7457 | -1.6026 | 1.2703 | 32.1242 | 40.8084 | 25.2881 | 2.0068 | 32.048 | 64.3128 | 15.97 | 14.5163 | -0.1105 | 105.1212 | 235.084 | 180.7696 | 232.8595 | 161.2294
2012-10-30 19:14 | 28.7079 | 249.5837 | 195.2928 | 247.4062 | 175.819 | -1.5997 | 1.2706 | 32.1313 | 40.8266 | 25.2879 | 2.0072 | 32.055 | 64.341 | 15.9699 | 14.4826 | -0.1103 | 105.1676 | 235.1011 | 180.8102 | 232.9236 | 161.3364
2012-10-30 19:14 | 28.7049 | 249.5729 | 195.3605 | 247.3993 | 175.8901 | -1.602 | 1.2708 | 32.1354 | 40.8365 | 25.2882 | 2.0075 | 32.0591 | 64.3577 | 15.9699 | 14.5041 | -0.1104 | 105.1942 | 235.0688 | 180.8564 | 232.8952 | 161.386
2012-10-30 19:14 | 28.7042 | 249.6192 | 195.3952 | 247.3869 | 175.8855 | -1.6026 | 1.2708 | 32.1375 | 40.8418 | 25.2883 | 2.0076 | 32.0612 | 64.3665 | 15.9698 | 14.4748 | -0.1105 | 105.2083 | 235.1444 | 180.9204 | 232.9121 | 161.4107
2012-10-30 19:14 | 28.7034 | 249.634 | 195.4455 | 247.3586 | 175.8769 | -1.5991 | 1.2709 | 32.1392 | 40.8464 | 25.2881 | 2.0077 | 32.0631 | 64.374 | 15.9698 | 14.5352 | -0.1102 | 105.2204 | 235.0988 | 180.9103 | 232.8234 | 161.3417
2012-10-30 19:14 | 28.7019 | 249.6169 | 195.4283 | 247.271 | 175.891 | -1.6043 | 1.271 | 32.1406 | 40.8504 | 25.2879 | 2.0078 | 32.0645 | 64.3798 | 15.9698 | 14.5133 | -0.1106 | 105.2302 | 235.1036 | 180.915 | 232.7577 | 161.3777
2012-10-30 19:14 | 28.7034 | 249.6301 | 195.4101 | 247.3195 | 175.7442 | -1.6003 | 1.271 | 32.1419 | 40.8531 | 25.2882 | 2.0079 | 32.0656 | 64.3849 | 15.9696 | 14.5026 | -0.1103 | 105.238 | 235.1275 | 180.9075 | 232.8169 | 161.2416
2012-10-30 19:14 | 28.6982 | 249.6014 | 195.4049 | 247.3926 | 175.4736 | -1.5933 | 1.271 | 32.1422 | 40.8537 | 25.2883 | 2.0079 | 32.0659 | 64.3859 | 15.9697 | 14.5055 | -0.1098 | 105.2396 | 235.0959 | 180.8994 | 232.8871 | 160.9681
2012-10-30 19:14 | 28.7079 | 249.662 | 195.4343 | 247.3318 | 175.6668 | -1.6066 | 1.2711 | 32.1427 | 40.8551 | 25.2882 | 2.0079 | 32.0663 | 64.3871 | 15.9698 | 14.4744 | -0.1108 | 105.2422 | 235.1876 | 180.9599 | 232.8574 | 161.1924
2012-10-30 19:14 | 28.7132 | 249.6438 | 195.4042 | 247.3293 | 175.7111 | -1.6032 | 1.2711 | 32.1432 | 40.8567 | 25.288 | 2.008 | 32.0669 | 64.3899 | 15.9697 | 14.492 | -0.1105 | 105.2466 | 235.1518 | 180.9122 | 232.8373 | 161.2191
2012-10-30 19:14 | 28.7102 | 249.6486 | 195.3933 | 247.338 | 175.712 | -1.6003 | 1.2711 | 32.1434 | 40.8569 | 25.2882 | 2.008 | 32.0672 | 64.3908 | 15.9697 | 14.489 | -0.1103 | 105.2477 | 235.1596 | 180.9043 | 232.849 | 161.223
2012-10-30 19:14 | 28.7124 | 249.6156 | 195.4388 | 247.446 | 175.6908 | -1.5962 | 1.2711 | 32.1437 | 40.858 | 25.288 | 2.008 | 32.0674 | 64.392 | 15.9696 | 14.4912 | -0.11 | 105.25 | 235.1244 | 180.9476 | 232.9548 | 161.1996
2012-10-30 19:14 | 28.7064 | 249.6996 | 195.3974 | 247.3812 | 175.9385 | -1.6061 | 1.2711 | 32.1437 | 40.8578 | 25.2881 | 2.008 | 32.0674 | 64.3926 | 15.9695 | 14.4852 | -0.1107 | 105.2504 | 235.2144 | 180.9122 | 232.896 | 161.4533
2012-10-30 19:14 | 28.7139 | 249.6758 | 195.4442 | 247.4122 | 176.0006 | -1.598 | 1.2711 | 32.1443 | 40.8591 | 25.2882 | 2.0081 | 32.068 | 64.3947 | 15.9696 | 14.4682 | -0.1102 | 105.2539 | 235.2076 | 180.976 | 232.944 | 161.5324
2012-10-30 19:14 | 28.7124 | 249.6822 | 195.4427 | 247.3598 | 175.9562 | -1.6026 | 1.2711 | 32.1436 | 40.8574 | 25.2882 | 2.008 | 32.0674 | 64.3928 | 15.9695 | 14.4259 | -0.1105 | 105.2501 | 235.2563 | 181.0168 | 232.9339 | 161.5303
2012-10-30 19:14 | 28.7057 | 249.6598 | 195.5264 | 247.408 | 175.9768 | -1.6003 | 1.2711 | 32.1437 | 40.8578 | 25.2881 | 2.008 | 32.0674 | 64.3924 | 15.9696 | 14.4763 | -0.1103 | 105.2502 | 235.1835 | 181.0501 | 232.9317 | 161.5005
2012-10-30 19:15 | 28.7057 | 249.6676 | 195.5421 | 247.4354 | 175.7855 | -1.6026 | 1.2711 | 32.1439 | 40.8587 | 25.2879 | 2.0081 | 32.0676 | 64.3935 | 15.9695 | 14.5049 | -0.1105 | 105.2522 | 235.1627 | 181.0372 | 232.9305 | 161.2806
2012-10-30 19:15 | 28.7064 | 249.6683 | 195.5231 | 247.4008 | 175.7862 | -1.598 | 1.2711 | 32.144 | 40.8585 | 25.2882 | 2.0081 | 32.0677 | 64.3941 | 15.9694 | 14.5015 | -0.1102 | 105.2526 | 235.1668 | 181.0216 | 232.8993 | 161.2847
2012-10-30 19:15 | 28.7049 | 249.7177 | 195.5334 | 247.4385 | 175.7769 | -1.598 | 1.2711 | 32.1435 | 40.8573 | 25.2882 | 2.008 | 32.0673 | 64.3921 | 15.9695 | 14.5122 | -0.1102 | 105.2494 | 235.2055 | 181.0212 | 232.9263 | 161.2647
2012-10-30 19:15 | 28.7057 | 249.7029 | 195.5224 | 247.4197 | 175.887 | -1.5985 | 1.2711 | 32.1437 | 40.8574 | 25.2884 | 2.0081 | 32.0674 | 64.3935 | 15.9693 | 14.5008 | -0.1102 | 105.2508 | 235.2021 | 181.0216 | 232.9189 | 161.3862
2012-10-30 19:15 | 28.7132 | 249.7025 | 195.5299 | 247.4155 | 175.8243 | -1.5991 | 1.2711 | 32.1434 | 40.8569 | 25.2882 | 2.008 | 32.0671 | 64.392 | 15.9694 | 14.5124 | -0.1102 | 105.2488 | 235.1901 | 181.0175 | 232.9031 | 161.3119
2012-10-30 19:15 | 28.7019 | 249.6443 | 195.4676 | 247.4825 | 175.6608 | -1.5985 | 1.2711 | 32.1435 | 40.8572 | 25.2883 | 2.008 | 32.0672 | 64.3923 | 15.9694 | 14.5133 | -0.1102 | 105.2495 | 235.131 | 180.9543 | 232.9692 | 161.1475
2012-10-30 19:15 | 28.7019 | 249.6795 | 195.5029 | 247.4708 | 175.649 | -1.6055 | 1.2711 | 32.1434 | 40.8571 | 25.2882 | 2.008 | 32.0672 | 64.392 | 15.9694 | 14.5133 | -0.1107 | 105.2491 | 235.1662 | 180.9896 | 232.9575 | 161.1357
2012-10-30 19:15 | 28.7064 | 249.7075 | 195.4917 | 247.4713 | 175.8097 | -1.6003 | 1.2711 | 32.143 | 40.8561 | 25.288 | 2.008 | 32.0668 | 64.3903 | 15.9694 | 14.477 | -0.1103 | 105.2464 | 235.2305 | 181.0147 | 232.9943 | 161.3327
2012-10-30 19:15 | 28.7072 | 249.7396 | 195.4689 | 247.4369 | 175.8612 | -1.5939 | 1.2711 | 32.1426 | 40.8552 | 25.2881 | 2.008 | 32.0664 | 64.3892 | 15.9693 | 14.5308 | -0.1099 | 105.2444 | 235.2088 | 180.9381 | 232.9061 | 161.3304
2012-10-30 19:15 | 28.7049 | 249.6747 | 195.4666 | 247.4933 | 175.6559 | -1.6037 | 1.2711 | 32.1425 | 40.8549 | 25.2881 | 2.008 | 32.0663 | 64.3892 | 15.9693 | 14.5245 | -0.1106 | 105.2441 | 235.1502 | 180.9421 | 232.9688 | 161.1314
2012-10-30 19:15 | 28.7027 | 249.6999 | 195.433 | 247.4911 | 175.5211 | -1.598 | 1.271 | 32.1423 | 40.8539 | 25.2883 | 2.008 | 32.0661 | 64.3881 | 15.9693 | 14.5059 | -0.1102 | 105.242 | 235.194 | 180.9271 | 232.9852 | 161.0152
2012-10-30 19:15 | 28.7027 | 249.739 | 195.4408 | 247.5538 | 175.5562 | -1.5974 | 1.271 | 32.1417 | 40.8524 | 25.2883 | 2.0079 | 32.0654 | 64.3851 | 15.9694 | 14.5223 | -0.1101 | 105.2376 | 235.2167 | 180.9185 | 233.0315 | 161.0339
2012-10-30 19:15 | 28.6989 | 249.747 | 195.4213 | 247.4952 | 175.6031 | -1.6061 | 1.271 | 32.1416 | 40.852 | 25.2884 | 2.0079 | 32.0653 | 64.3847 | 15.9694 | 14.4981 | -0.1107 | 105.2367 | 235.2489 | 180.9232 | 232.9971 | 161.105
2012-10-30 19:15 | 28.7004 | 249.725 | 195.4268 | 247.4536 | 175.5578 | -1.5968 | 1.271 | 32.1412 | 40.8517 | 25.2881 | 2.0079 | 32.0651 | 64.3839 | 15.9693 | 14.4832 | -0.1101 | 105.2355 | 235.2418 | 180.9436 | 232.9704 | 161.0746
2012-10-30 19:15 | 28.6974 | 249.7415 | 195.3963 | 247.5172 | 175.5197 | -1.6049 | 1.271 | 32.141 | 40.8503 | 25.2885 | 2.0079 | 32.0647 | 64.3822 | 15.9694 | 14.4843 | -0.1106 | 105.2325 | 235.2572 | 180.912 | 233.0329 | 161.0354
2012-10-30 19:15 | 28.7064 | 249.7231 | 195.3778 | 247.5693 | 175.4233 | -1.602 | 1.271 | 32.1406 | 40.85 | 25.2882 | 2.0079 | 32.0644 | 64.3815 | 15.9692 | 14.5137 | -0.1104 | 105.2315 | 235.2094 | 180.8641 | 233.0556 | 160.9096
2012-10-30 19:15 | 28.7042 | 249.7483 | 195.3991 | 247.5553 | 175.5382 | -1.5997 | 1.271 | 32.1404 | 40.849 | 25.2883 | 2.0079 | 32.0641 | 64.3803 | 15.9693 | 14.4911 | -0.1103 | 105.2293 | 235.2572 | 180.908 | 233.0642 | 161.0471
2012-10-30 19:15 | 28.7004 | 249.768 | 195.3678 | 247.4849 | 175.4992 | -1.6037 | 1.2709 | 32.1399 | 40.8478 | 25.2883 | 2.0078 | 32.0637 | 64.3789 | 15.9692 | 14.4832 | -0.1106 | 105.2267 | 235.2848 | 180.8846 | 233.0017 | 161.016
2012-10-30 19:15 | 28.6952 | 249.7119 | 195.3312 | 247.5032 | 175.4589 | -1.5997 | 1.2709 | 32.1397 | 40.8472 | 25.2884 | 2.0078 | 32.0635 | 64.3785 | 15.9691 | 14.4821 | -0.1103 | 105.2257 | 235.2298 | 180.8491 | 233.0211 | 160.9768
2012-10-30 19:15 | 28.6997 | 249.7438 | 195.3396 | 247.5116 | 175.4751 | -1.598 | 1.2709 | 32.1396 | 40.8478 | 25.2879 | 2.0078 | 32.0634 | 64.3782 | 15.9691 | 14.5233 | -0.1102 | 105.2261 | 235.2205 | 180.8163 | 232.9883 | 160.9518
2012-10-30 19:15 | 28.7012 | 249.7062 | 195.3097 | 247.568 | 175.461 | -1.5991 | 1.2709 | 32.1392 | 40.8464 | 25.2881 | 2.0078 | 32.063 | 64.3764 | 15.9692 | 14.484 | -0.1102 | 105.2228 | 235.2222 | 180.8257 | 233.084 | 160.977
2012-10-30 19:15 | 28.6982 | 249.6993 | 195.342 | 247.5611 | 175.54 | -1.602 | 1.2709 | 32.1387 | 40.8453 | 25.288 | 2.0078 | 32.0625 | 64.3745 | 15.9691 | 14.5218 | -0.1104 | 105.2198 | 235.1775 | 180.8202 | 233.0393 | 161.0182
2012-10-30 19:15 | 28.6989 | 249.7352 | 195.3074 | 247.4521 | 175.4548 | -1.6043 | 1.2709 | 32.1386 | 40.8447 | 25.2882 | 2.0078 | 32.0623 | 64.3741 | 15.969 | 14.4858 | -0.1106 | 105.2188 | 235.2494 | 180.8216 | 232.9663 | 160.969
2012-10-30 19:16 | 28.7004 | 249.7289 | 195.2696 | 247.5202 | 175.3822 | -1.6014 | 1.2709 | 32.1383 | 40.8444 | 25.288 | 2.0078 | 32.062 | 64.373 | 15.969 | 14.4465 | -0.1104 | 105.2173 | 235.2824 | 180.8231 | 233.0737 | 160.9357
2012-10-30 19:16 | 28.7094 | 249.7379 | 195.255 | 247.5409 | 175.5356 | -1.602 | 1.2709 | 32.1379 | 40.8436 | 25.2877 | 2.0077 | 32.0616 | 64.3715 | 15.9689 | 14.4474 | -0.1104 | 105.2151 | 235.2905 | 180.8076 | 233.0935 | 161.0882
2012-10-30 19:16 | 28.7057 | 249.695 | 195.322 | 247.5764 | 175.4694 | -1.6037 | 1.2709 | 32.1374 | 40.8418 | 25.2881 | 2.0077 | 32.0612 | 64.3697 | 15.969 | 14.4967 | -0.1106 | 105.2115 | 235.1983 | 180.8253 | 233.0797 | 160.9727
2012-10-30 19:16 | 28.7109 | 249.7002 | 195.2762 | 247.5463 | 175.4278 | -1.6026 | 1.2709 | 32.1372 | 40.8417 | 25.2879 | 2.0077 | 32.061 | 64.3696 | 15.9689 | 14.4611 | -0.1105 | 105.2113 | 235.2391 | 180.8151 | 233.0852 | 160.9667
2012-10-30 19:16 | 28.7072 | 249.72 | 195.2646 | 247.5348 | 175.4124 | -1.5997 | 1.2708 | 32.1367 | 40.8405 | 25.2879 | 2.0077 | 32.0606 | 64.3673 | 15.969 | 14.4656 | -0.1103 | 105.2078 | 235.2544 | 180.799 | 233.0692 | 160.9468
2012-10-30 19:16 | 28.7034 | 249.7906 | 195.2412 | 247.4409 | 175.5335 | -1.5939 | 1.2708 | 32.1362 | 40.8393 | 25.2878 | 2.0076 | 32.06 | 64.3647 | 15.969 | 14.5026 | -0.1099 | 105.204 | 235.288 | 180.7386 | 232.9383 | 161.0309
2012-10-30 19:16 | 28.7049 | 249.7177 | 195.223 | 247.5168 | 175.4881 | -1.598 | 1.2684 | 32.0747 | 40.6831 | 25.2878 | 2.0039 | 31.9998 | 64.1235 | 15.9689 | 14.4918 | -0.1102 | 104.8066 | 235.2259 | 180.7312 | 233.025 | 160.9963
2012-10-30 19:16 | 28.7079 | 249.7011 | 195.2653 | 247.5512 | 175.3545 | -1.5997 | 1.2647 | 31.9799 | 40.4459 | 25.286 | 1.9979 | 31.9039 | 63.7397 | 15.969 | 14.4989 | -0.1103 | 104.1856 | 235.2022 | 180.7664 | 233.0523 | 160.8556
2012-10-30 19:16 | 28.6952 | 249.6806 | 195.1858 | 247.5267 | 175.5175 | -1.6084 | 1.2641 | 31.9651 | 40.4079 | 25.2862 | 1.9969 | 31.889 | 63.6798 | 15.9691 | 14.4821 | -0.1109 | 104.0877 | 235.1985 | 180.7037 | 233.0446 | 161.0354
2012-10-30 19:16 | 28.7019 | 249.6912 | 195.2239 | 247.5413 | 175.7622 | -1.5985 | 1.2637 | 31.9534 | 40.379 | 25.2859 | 1.9962 | 31.8776 | 63.6341 | 15.9692 | 14.4684 | -0.1102 | 104.0131 | 235.2228 | 180.7555 | 233.0729 | 161.2938
2012-10-30 19:16 | 28.6967 | 249.6939 | 195.2541 | 247.4656 | 175.6712 | -1.5997 | 1.2635 | 31.9477 | 40.3644 | 25.286 | 1.9959 | 31.8719 | 63.6116 | 15.9691 | 14.4918 | -0.1103 | 103.9761 | 235.2021 | 180.7623 | 232.9738 | 161.1794
2012-10-30 19:16 | 28.6899 | 249.6518 | 195.2394 | 247.4901 | 175.6097 | -1.602 | 1.2634 | 31.947 | 40.363 | 25.2859 | 1.9958 | 31.8713 | 63.6083 | 15.9693 | 14.4768 | -0.1104 | 103.9713 | 235.175 | 180.7626 | 233.0133 | 161.1329
2012-10-30 19:16 | 28.6907 | 249.59 | 195.2678 | 247.5849 | 175.6886 | -1.602 | 1.2635 | 31.9501 | 40.3698 | 25.2864 | 1.996 | 31.8743 | 63.6204 | 15.9692 | 14.5021 | -0.1104 | 103.9902 | 235.0879 | 180.7657 | 233.0828 | 161.1865
2012-10-30 19:16 | 28.6802 | 249.6186 | 195.2926 | 247.5235 | 175.7796 | -1.6101 | 1.2636 | 31.9508 | 40.3722 | 25.2861 | 1.996 | 31.8751 | 63.6232 | 15.9694 | 14.4753 | -0.111 | 103.9955 | 235.1433 | 180.8173 | 233.0482 | 161.3043
2012-10-30 19:16 | 28.6862 | 249.5816 | 195.33 | 247.5021 | 175.9807 | -1.6026 | 1.2637 | 31.9535 | 40.3791 | 25.2861 | 1.9962 | 31.8778 | 63.6343 | 15.9693 | 14.5343 | -0.1105 | 104.0133 | 235.0473 | 180.7957 | 232.9678 | 161.4464
2012-10-30 19:16 | 28.6772 | 249.5765 | 195.3525 | 247.4304 | 175.9483 | -1.6037 | 1.2638 | 31.9563 | 40.3862 | 25.286 | 1.9964 | 31.8806 | 63.6452 | 15.9693 | 14.5335 | -0.1106 | 104.0314 | 235.043 | 180.819 | 232.8969 | 161.4148
2012-10-30 19:16 | 28.6727 | 249.5368 | 195.3716 | 247.5003 | 175.7955 | -1.6014 | 1.268 | 32.063 | 40.655 | 25.2869 | 2.0029 | 31.9844 | 64.0601 | 15.9693 | 14.5739 | -0.1104 | 104.7151 | 234.9629 | 180.7977 | 232.9264 | 161.2216
2012-10-30 19:16 | 28.6854 | 249.5808 | 195.3843 | 247.5013 | 175.8043 | -1.5974 | 1.2691 | 32.0925 | 40.7286 | 25.2876 | 2.0049 | 32.0163 | 64.1884 | 15.9692 | 14.5539 | -0.1101 | 104.917 | 235.0269 | 180.8304 | 232.9474 | 161.2504
2012-10-30 19:16 | 28.6869 | 249.6293 | 195.4486 | 247.4479 | 175.8643 | -1.6003 | 1.2697 | 32.1086 | 40.7697 | 25.2874 | 2.0059 | 32.0326 | 64.2544 | 15.9692 | 14.5187 | -0.1103 | 105.0241 | 235.1106 | 180.9299 | 232.9292 | 161.3456
2012-10-30 19:16 | 28.6795 | 249.567 | 195.4687 | 247.2916 | 175.8843 | -1.6032 | 1.2702 | 32.1202 | 40.7983 | 25.288 | 2.0066 | 32.0441 | 64.3008 | 15.9691 | 14.5235 | -0.1105 | 105.0991 | 235.0435 | 180.9452 | 232.7681 | 161.3608
2012-10-30 19:16 | 28.6787 | 249.6054 | 195.4679 | 247.4319 | 175.8405 | -1.6014 | 1.2705 | 32.1284 | 40.8187 | 25.2883 | 2.0071 | 32.0523 | 64.3329 | 15.9693 | 14.5676 | -0.1104 | 105.1515 | 235.0378 | 180.9003 | 232.8643 | 161.2729
2012-10-30 19:16 | 28.6735 | 249.6119 | 195.4667 | 247.4424 | 175.9641 | -1.5991 | 1.2707 | 32.1343 | 40.834 | 25.2881 | 2.0075 | 32.0582 | 64.3569 | 15.9692 | 14.5461 | -0.1102 | 105.1909 | 235.0658 | 180.9206 | 232.8963 | 161.418
2012-10-30 19:16 | 28.675 | 249.6174 | 195.4996 | 247.3851 | 176.0047 | -1.6009 | 1.2708 | 32.1376 | 40.842 | 25.2884 | 2.0077 | 32.0614 | 64.3696 | 15.9693 | 14.515 | -0.1104 | 105.2116 | 235.1024 | 180.9846 | 232.8701 | 161.4897
2012-10-30 19:17 | 28.6757 | 249.6063 | 195.571 | 247.378 | 175.8961 | -1.6009 | 1.2709 | 32.1392 | 40.8458 | 25.2884 | 2.0078 | 32.0629 | 64.3762 | 15.9691 | 14.5157 | -0.1104 | 105.2219 | 235.0906 | 181.0553 | 232.8623 | 161.3804
2012-10-30 19:17 | 28.6772 | 249.6313 | 195.49 | 247.3129 | 175.8585 | -1.602 | 1.271 | 32.1406 | 40.8493 | 25.2885 | 2.0079 | 32.0643 | 64.3815 | 15.9692 | 14.4968 | -0.1104 | 105.2307 | 235.1345 | 180.9932 | 232.8161 | 161.3617
2012-10-30 19:17 | 28.6772 | 249.6235 | 195.5136 | 247.4225 | 175.8273 | -1.6014 | 1.271 | 32.1416 | 40.8524 | 25.2881 | 2.008 | 32.0653 | 64.3863 | 15.969 | 14.4641 | -0.1104 | 105.2387 | 235.1594 | 181.0495 | 232.9584 | 161.3632
2012-10-30 19:17 | 28.6839 | 249.6106 | 195.5203 | 247.4253 | 175.8379 | -1.6003 | 1.271 | 32.1425 | 40.8545 | 25.2882 | 2.008 | 32.0662 | 64.3896 | 15.969 | 14.532 | -0.1103 | 105.2441 | 235.0786 | 180.9883 | 232.8933 | 161.3059
2012-10-30 19:17 | 28.6795 | 249.6219 | 195.5355 | 247.4327 | 175.7906 | -1.6026 | 1.2711 | 32.1431 | 40.8563 | 25.2881 | 2.0081 | 32.0669 | 64.3927 | 15.9689 | 14.4583 | -0.1105 | 105.2491 | 235.1636 | 181.0772 | 232.9744 | 161.3323
2012-10-30 19:17 | 28.6802 | 249.6656 | 195.5205 | 247.3472 | 175.6703 | -1.5974 | 1.2711 | 32.1435 | 40.8571 | 25.2882 | 2.0081 | 32.0672 | 64.3936 | 15.9691 | 14.4794 | -0.1101 | 105.2507 | 235.1862 | 181.0411 | 232.8678 | 161.1909
2012-10-30 19:17 | 28.6847 | 249.6231 | 195.4779 | 247.3595 | 175.6709 | -1.6049 | 1.2711 | 32.1443 | 40.8589 | 25.2883 | 2.0081 | 32.068 | 64.397 | 15.969 | 14.4961 | -0.1106 | 105.256 | 235.127 | 180.9818 | 232.8634 | 161.1748
2012-10-30 19:17 | 28.6824 | 249.6404 | 195.4834 | 247.4238 | 175.6686 | -1.5985 | 1.2711 | 32.1438 | 40.8577 | 25.2883 | 2.0081 | 32.0674 | 64.3951 | 15.9689 | 14.4612 | -0.1102 | 105.2528 | 235.1792 | 181.0222 | 232.9626 | 161.2074
2012-10-30 19:17 | 28.6847 | 249.6662 | 195.4621 | 247.4849 | 175.6982 | -1.6014 | 1.2711 | 32.1439 | 40.8586 | 25.288 | 2.0081 | 32.0676 | 64.3961 | 15.9689 | 14.4594 | -0.1104 | 105.2547 | 235.2068 | 181.0027 | 233.0255 | 161.2388
2012-10-30 19:17 | 28.6922 | 249.6854 | 195.4657 | 247.3945 | 175.6979 | -1.5991 | 1.2711 | 32.1443 | 40.8594 | 25.288 | 2.0082 | 32.068 | 64.3975 | 15.9688 | 14.4383 | -0.1102 | 105.257 | 235.2471 | 181.0274 | 232.9562 | 161.2596
2012-10-30 19:17 | 28.6907 | 249.7114 | 195.4564 | 247.4125 | 175.5364 | -1.6049 | 1.2711 | 32.1444 | 40.8595 | 25.2882 | 2.0082 | 32.0681 | 64.3978 | 15.9689 | 14.4817 | -0.1106 | 105.2573 | 235.2297 | 180.9747 | 232.9308 | 161.0547
2012-10-30 19:17 | 28.6847 | 249.6505 | 195.3875 | 247.4849 | 175.5889 | -1.602 | 1.2711 | 32.1447 | 40.86 | 25.2884 | 2.0082 | 32.0684 | 64.3993 | 15.9689 | 14.4635 | -0.1104 | 105.2593 | 235.187 | 180.924 | 233.0214 | 161.1254
2012-10-30 19:17 | 28.6892 | 249.6276 | 195.4156 | 247.4541 | 175.691 | -1.6014 | 1.2711 | 32.1441 | 40.859 | 25.2881 | 2.0082 | 32.0678 | 64.3969 | 15.9688 | 14.4516 | -0.1104 | 105.2559 | 235.176 | 180.964 | 233.0025 | 161.2394
2012-10-30 19:17 | 28.6862 | 249.6442 | 195.3851 | 247.4198 | 175.6099 | -1.5985 | 1.2711 | 32.1437 | 40.8576 | 25.2883 | 2.0081 | 32.0673 | 64.3955 | 15.9687 | 14.4772 | -0.1102 | 105.2531 | 235.167 | 180.9079 | 232.9426 | 161.1327
2012-10-30 19:17 | 28.6914 | 249.7042 | 195.3549 | 247.3584 | 175.5917 | -1.6009 | 1.2711 | 32.1439 | 40.8585 | 25.288 | 2.0082 | 32.0676 | 64.3966 | 15.9687 | 14.4498 | -0.1104 | 105.255 | 235.2544 | 180.9051 | 232.9086 | 161.1419
2012-10-30 19:17 | 28.6937 | 249.6869 | 195.3729 | 247.4547 | 175.5238 | -1.6014 | 1.2711 | 32.144 | 40.8582 | 25.2884 | 2.0082 | 32.0678 | 64.397 | 15.9688 | 14.4929 | -0.1104 | 105.2553 | 235.194 | 180.88 | 232.9618 | 161.0309
2012-10-30 19:17 | 28.6997 | 249.6734 | 195.3593 | 247.5156 | 175.522 | -1.6055 | 1.2711 | 32.1436 | 40.8573 | 25.2882 | 2.0081 | 32.0673 | 64.3954 | 15.9687 | 14.4948 | -0.1107 | 105.2527 | 235.1786 | 180.8645 | 233.0208 | 161.0272
2012-10-30 19:17 | 28.6937 | 249.64 | 195.3493 | 247.4704 | 175.633 | -1.6043 | 1.2711 | 32.1436 | 40.8572 | 25.2884 | 2.0081 | 32.0674 | 64.3957 | 15.9687 | 14.5133 | -0.1106 | 105.253 | 235.1267 | 180.836 | 232.9571 | 161.1197
2012-10-30 19:17 | 28.7064 | 249.6644 | 195.366 | 247.4087 | 175.6145 | -1.5974 | 1.2711 | 32.1436 | 40.8573 | 25.2883 | 2.0081 | 32.0673 | 64.3957 | 15.9687 | 14.5015 | -0.1101 | 105.253 | 235.1629 | 180.8645 | 232.9072 | 161.113
2012-10-30 19:17 | 28.7019 | 249.6364 | 195.33 | 247.553 | 175.4968 | -1.6037 | 1.2711 | 32.1434 | 40.8567 | 25.2883 | 2.0081 | 32.067 | 64.3938 | 15.9688 | 14.4643 | -0.1106 | 105.2505 | 235.1721 | 180.8657 | 233.0887 | 161.0325
2012-10-30 19:17 | 28.7004 | 249.6349 | 195.3167 | 247.4693 | 175.5734 | -1.6009 | 1.2711 | 32.143 | 40.8558 | 25.2882 | 2.0081 | 32.0666 | 64.393 | 15.9686 | 14.4506 | -0.1104 | 105.2488 | 235.1843 | 180.8661 | 233.0187 | 161.1228
2012-10-30 19:17 | 28.7087 | 249.6902 | 195.2897 | 247.4893 | 175.6168 | -1.5985 | 1.2711 | 32.1428 | 40.8551 | 25.2883 | 2.0081 | 32.0664 | 64.3922 | 15.9686 | 14.463 | -0.1102 | 105.2473 | 235.2272 | 180.8267 | 233.0263 | 161.1538
2012-10-30 19:17 | 28.7132 | 249.6986 | 195.2903 | 247.3528 | 175.6096 | -1.5945 | 1.271 | 32.1424 | 40.8544 | 25.2883 | 2.0081 | 32.066 | 64.3905 | 15.9687 | 14.4797 | -0.1099 | 105.2449 | 235.2189 | 180.8106 | 232.8731 | 161.1299
2012-10-30 19:17 | 28.7124 | 249.6313 | 195.3209 | 247.5126 | 175.7181 | -1.6014 | 1.271 | 32.1422 | 40.8537 | 25.2884 | 2.0081 | 32.066 | 64.3904 | 15.9687 | 14.4667 | -0.1104 | 105.2441 | 235.1646 | 180.8542 | 233.0459 | 161.2514
2012-10-30 19:18 | 28.7147 | 249.6884 | 195.3193 | 247.4757 | 175.7126 | -1.5997 | 1.271 | 32.1417 | 40.8524 | 25.2883 | 2.008 | 32.0654 | 64.3879 | 15.9687 | 14.4731 | -0.1103 | 105.2404 | 235.2153 | 180.8462 | 233.0026 | 161.2395
2012-10-30 19:18 | 28.7184 | 249.6803 | 195.3898 | 247.4363 | 175.7553 | -1.6003 | 1.271 | 32.1414 | 40.852 | 25.2882 | 2.008 | 32.0651 | 64.387 | 15.9686 | 14.4972 | -0.1103 | 105.239 | 235.1831 | 180.8926 | 232.9391 | 161.2581
2012-10-30 19:18 | 28.7206 | 249.7295 | 195.3841 | 247.5678 | 175.6834 | -1.6026 | 1.271 | 32.1413 | 40.8516 | 25.2882 | 2.008 | 32.0651 | 64.387 | 15.9686 | 14.4504 | -0.1105 | 105.2387 | 235.2791 | 180.9337 | 233.1174 | 161.233
2012-10-30 19:18 | 28.7229 | 249.6613 | 195.3864 | 247.5192 | 175.752 | -1.6066 | 1.271 | 32.1414 | 40.8521 | 25.288 | 2.008 | 32.0651 | 64.3871 | 15.9686 | 14.4853 | -0.1108 | 105.2392 | 235.176 | 180.9011 | 233.0339 | 161.2667
2012-10-30 19:18 | 28.7289 | 249.6908 | 195.416 | 247.4703 | 175.7463 | -1.5997 | 1.271 | 32.1412 | 40.851 | 25.2884 | 2.008 | 32.0649 | 64.3861 | 15.9686 | 14.475 | -0.1103 | 105.2371 | 235.2158 | 180.941 | 232.9953 | 161.2713
2012-10-30 19:18 | 28.7192 | 249.685 | 195.3748 | 247.4136 | 175.7756 | -1.6043 | 1.271 | 32.1413 | 40.8516 | 25.2883 | 2.008 | 32.0651 | 64.3876 | 15.9685 | 14.5061 | -0.1106 | 105.2392 | 235.1789 | 180.8687 | 232.9075 | 161.2695
2012-10-30 19:18 | 28.7221 | 249.731 | 195.3502 | 247.5027 | 175.7902 | -1.6009 | 1.271 | 32.1409 | 40.8507 | 25.2881 | 2.008 | 32.0646 | 64.3855 | 15.9685 | 14.4764 | -0.1104 | 105.2362 | 235.2546 | 180.8738 | 233.0263 | 161.3138
2012-10-30 19:18 | 28.7296 | 249.7072 | 195.4127 | 247.565 | 175.747 | -1.6026 | 1.271 | 32.1407 | 40.8503 | 25.2881 | 2.008 | 32.0644 | 64.3856 | 15.9683 | 14.5124 | -0.1105 | 105.2358 | 235.1948 | 180.9003 | 233.0526 | 161.2346
2012-10-30 19:18 | 28.7251 | 249.7301 | 195.4318 | 247.5841 | 175.7347 | -1.598 | 1.271 | 32.1399 | 40.8483 | 25.2881 | 2.0079 | 32.0636 | 64.3816 | 15.9684 | 14.4998 | -0.1102 | 105.2299 | 235.2303 | 180.932 | 233.0843 | 161.2349
2012-10-30 19:18 | 28.7266 | 249.7551 | 195.453 | 247.515 | 175.7557 | -1.6009 | 1.271 | 32.1398 | 40.8482 | 25.288 | 2.0079 | 32.0636 | 64.3817 | 15.9684 | 14.5176 | -0.1104 | 105.23 | 235.2375 | 180.9354 | 232.9974 | 161.2381
2012-10-30 19:18 | 28.7296 | 249.7385 | 195.4206 | 247.5063 | 175.786 | -1.5991 | 1.2709 | 32.1395 | 40.8472 | 25.288 | 2.0079 | 32.0632 | 64.3805 | 15.9683 | 14.5002 | -0.1102 | 105.2276 | 235.2383 | 180.9204 | 233.0061 | 161.2858
2012-10-30 19:18 | 28.7221 | 249.7271 | 195.4131 | 247.5497 | 175.7863 | -1.6009 | 1.2709 | 32.1388 | 40.846 | 25.2878 | 2.0079 | 32.0626 | 64.3785 | 15.9682 | 14.509 | -0.1104 | 105.2245 | 235.2181 | 180.9041 | 233.0407 | 161.2773
2012-10-30 19:18 | 28.7192 | 249.7516 | 195.4298 | 247.5938 | 175.7639 | -1.6032 | 1.2709 | 32.1384 | 40.8445 | 25.288 | 2.0079 | 32.0622 | 64.377 | 15.9682 | 14.5224 | -0.1105 | 105.2215 | 235.2292 | 180.9074 | 233.0714 | 161.2415
2012-10-30 19:18 | 28.7192 | 249.7516 | 195.4573 | 247.5116 | 175.6781 | -1.6026 | 1.2709 | 32.1379 | 40.8435 | 25.2879 | 2.0078 | 32.0617 | 64.3746 | 15.9683 | 14.4939 | -0.1105 | 105.2181 | 235.2577 | 180.9634 | 233.0177 | 161.1842
2012-10-30 19:18 | 28.7281 | 249.7409 | 195.4623 | 247.497 | 175.7026 | -1.6014 | 1.2709 | 32.1374 | 40.842 | 25.288 | 2.0078 | 32.0612 | 64.3728 | 15.9683 | 14.5436 | -0.1104 | 105.2148 | 235.1973 | 180.9187 | 232.9534 | 161.159
2012-10-30 19:18 | 28.7296 | 249.7307 | 195.4206 | 247.5886 | 175.6026 | -1.6026 | 1.2708 | 32.1372 | 40.8412 | 25.2882 | 2.0078 | 32.0609 | 64.372 | 15.9682 | 14.5655 | -0.1105 | 105.2132 | 235.1652 | 180.8551 | 233.0231 | 161.0371
2012-10-30 19:18 | 28.7251 | 249.6949 | 195.4397 | 247.6036 | 175.5786 | -1.5962 | 1.2708 | 32.1371 | 40.8413 | 25.2879 | 2.0078 | 32.0608 | 64.3711 | 15.9682 | 14.4957 | -0.11 | 105.2124 | 235.1992 | 180.944 | 233.1079 | 161.0829
2012-10-30 19:18 | 28.7192 | 249.7633 | 195.4023 | 247.4802 | 175.6702 | -1.598 | 1.2708 | 32.1364 | 40.8395 | 25.2881 | 2.0078 | 32.0602 | 64.3692 | 15.9682 | 14.5102 | -0.1102 | 105.2087 | 235.2531 | 180.8921 | 232.97 | 161.16
2012-10-30 19:18 | 28.7192 | 249.7242 | 195.3788 | 247.4841 | 175.6507 | -1.6026 | 1.2676 | 32.0542 | 40.6322 | 25.2872 | 2.0026 | 31.9771 | 64.0359 | 15.9682 | 14.4857 | -0.1105 | 104.6681 | 235.2385 | 180.8931 | 232.9984 | 161.165
2012-10-30 19:18 | 28.7184 | 249.7391 | 195.378 | 247.6009 | 175.6031 | -1.5997 | 1.2646 | 31.9775 | 40.4392 | 25.2864 | 1.9978 | 31.9017 | 63.7339 | 15.9683 | 14.4972 | -0.1103 | 104.173 | 235.2419 | 180.8808 | 233.1037 | 161.1059
2012-10-30 19:18 | 28.7154 | 249.6969 | 195.3828 | 247.5744 | 175.6821 | -1.5991 | 1.264 | 31.9613 | 40.3989 | 25.286 | 1.9968 | 31.8856 | 63.6692 | 15.9683 | 14.5146 | -0.1102 | 104.0681 | 235.1823 | 180.8682 | 233.0598 | 161.1675
2012-10-30 19:18 | 28.7139 | 249.6797 | 195.3892 | 247.5023 | 175.9343 | -1.602 | 1.2637 | 31.9535 | 40.3792 | 25.286 | 1.9963 | 31.8777 | 63.6365 | 15.9686 | 14.4845 | -0.1104 | 104.0157 | 235.1952 | 180.9047 | 233.0178 | 161.4498
2012-10-30 19:19 | 28.7162 | 249.6664 | 195.4072 | 247.5164 | 175.9131 | -1.6032 | 1.2636 | 31.9515 | 40.3739 | 25.2861 | 1.9961 | 31.8757 | 63.6287 | 15.9686 | 14.4909 | -0.1105 | 104.0026 | 235.1755 | 180.9163 | 233.0255 | 161.4222
2012-10-30 19:19 | 28.7072 | 249.6496 | 195.41 | 247.5348 | 176.0056 | -1.602 | 1.2635 | 31.9494 | 40.3687 | 25.2861 | 1.996 | 31.8737 | 63.6201 | 15.9687 | 14.5023 | -0.1104 | 103.9887 | 235.1473 | 180.9077 | 233.0325 | 161.5033
2012-10-30 19:19 | 28.7214 | 249.5894 | 195.4281 | 247.643 | 175.9027 | -1.6032 | 1.2635 | 31.9502 | 40.3707 | 25.2861 | 1.9961 | 31.8744 | 63.6233 | 15.9686 | 14.5246 | -0.1105 | 103.994 | 235.0648 | 180.9035 | 233.1184 | 161.3781
2012-10-30 19:19 | 28.7162 | 249.6116 | 195.4583 | 247.4929 | 175.679 | -1.598 | 1.2636 | 31.9525 | 40.3766 | 25.286 | 1.9962 | 31.8768 | 63.6325 | 15.9687 | 14.5439 | -0.1102 | 104.0091 | 235.0677 | 180.9144 | 232.949 | 161.1351
2012-10-30 19:19 | 28.7109 | 249.6102 | 195.4373 | 247.4053 | 175.8181 | -1.6003 | 1.2637 | 31.9547 | 40.3817 | 25.2863 | 1.9963 | 31.8789 | 63.641 | 15.9687 | 14.5345 | -0.1103 | 104.0227 | 235.0757 | 180.9028 | 232.8708 | 161.2836
2012-10-30 19:19 | 28.7109 | 249.618 | 195.4333 | 247.4993 | 175.8961 | -1.6032 | 1.2638 | 31.9558 | 40.3844 | 25.2863 | 1.9964 | 31.88 | 63.6449 | 15.9688 | 14.5386 | -0.1105 | 104.0293 | 235.0794 | 180.8947 | 232.9607 | 161.3575
2012-10-30 19:19 | 28.7079 | 249.6033 | 195.4618 | 247.5904 | 175.8619 | -1.6032 | 1.268 | 32.0628 | 40.6541 | 25.287 | 2.0031 | 31.9864 | 64.0703 | 15.9688 | 14.5234 | -0.1105 | 104.7245 | 235.0799 | 180.9384 | 233.067 | 161.3385
2012-10-30 19:19 | 28.7034 | 249.6027 | 195.5083 | 247.4174 | 175.9667 | -1.6026 | 1.2691 | 32.0915 | 40.7258 | 25.2877 | 2.0049 | 32.0155 | 64.1877 | 15.9687 | 14.5148 | -0.1105 | 104.9135 | 235.0879 | 180.9935 | 232.9026 | 161.4519
2012-10-30 19:19 | 28.7034 | 249.6223 | 195.5201 | 247.3508 | 175.9589 | -1.6049 | 1.2697 | 32.1082 | 40.7681 | 25.2878 | 2.0059 | 32.0318 | 64.2534 | 15.9686 | 14.5189 | -0.1106 | 105.0215 | 235.1034 | 181.0012 | 232.8319 | 161.44
2012-10-30 19:19 | 28.6914 | 249.6338 | 195.5435 | 247.4172 | 175.9274 | -1.5985 | 1.2702 | 32.1219 | 40.8025 | 25.288 | 2.0068 | 32.0456 | 64.3087 | 15.9686 | 14.5191 | -0.1102 | 105.1112 | 235.1147 | 181.0244 | 232.8981 | 161.4083
2012-10-30 19:19 | 28.6937 | 249.6047 | 195.5654 | 247.5409 | 175.9063 | -1.5974 | 1.2706 | 32.1298 | 40.8228 | 25.288 | 2.0073 | 32.0536 | 64.341 | 15.9686 | 14.5051 | -0.1101 | 105.1638 | 235.0996 | 181.0603 | 233.0358 | 161.4012
2012-10-30 19:19 | 28.6959 | 249.6657 | 195.5991 | 247.3864 | 176.0802 | -1.602 | 1.2707 | 32.1341 | 40.8333 | 25.2882 | 2.0076 | 32.058 | 64.3586 | 15.9685 | 14.5277 | -0.1104 | 105.1919 | 235.138 | 181.0714 | 232.8587 | 161.5525
2012-10-30 19:19 | 28.6892 | 249.6629 | 195.5845 | 247.3562 | 176.0072 | -1.598 | 1.2708 | 32.1374 | 40.8414 | 25.2884 | 2.0078 | 32.061 | 64.3709 | 15.9685 | 14.4884 | -0.1102 | 105.2123 | 235.1745 | 181.0961 | 232.8678 | 161.5188
2012-10-30 19:19 | 28.6862 | 249.6207 | 195.6051 | 247.4355 | 176.1213 | -1.5957 | 1.2709 | 32.1396 | 40.8471 | 25.2883 | 2.0079 | 32.0634 | 64.3809 | 15.9684 | 14.4772 | -0.11 | 105.2279 | 235.1435 | 181.1279 | 232.9583 | 161.6441
2012-10-30 19:19 | 28.6869 | 249.6293 | 195.6372 | 247.4401 | 175.9346 | -1.598 | 1.271 | 32.1413 | 40.8516 | 25.2883 | 2.008 | 32.065 | 64.3875 | 15.9684 | 14.5269 | -0.1102 | 105.239 | 235.1024 | 181.1103 | 232.9132 | 161.4077
2012-10-30 19:19 | 28.6847 | 249.6779 | 195.6429 | 247.3752 | 175.9285 | -1.6003 | 1.2711 | 32.1425 | 40.855 | 25.288 | 2.0081 | 32.0661 | 64.3916 | 15.9685 | 14.5287 | -0.1103 | 105.2466 | 235.1492 | 181.1142 | 232.8465 | 161.3998
2012-10-30 19:19 | 28.6847 | 249.674 | 195.5957 | 247.3478 | 175.9402 | -1.6003 | 1.2711 | 32.1435 | 40.8574 | 25.288 | 2.0082 | 32.0672 | 64.3964 | 15.9684 | 14.4635 | -0.1103 | 105.2539 | 235.2105 | 181.1322 | 232.8843 | 161.4767
2012-10-30 19:19 | 28.6824 | 249.6365 | 195.5974 | 247.4513 | 175.8286 | -1.5968 | 1.2711 | 32.1442 | 40.8589 | 25.2882 | 2.0082 | 32.0678 | 64.3988 | 15.9684 | 14.4816 | -0.1101 | 105.2578 | 235.1549 | 181.1158 | 232.9697 | 161.347
2012-10-30 19:19 | 28.6839 | 249.6928 | 195.591 | 247.4097 | 175.752 | -1.602 | 1.2711 | 32.1449 | 40.8606 | 25.2882 | 2.0083 | 32.0685 | 64.4018 | 15.9683 | 14.4912 | -0.1104 | 105.2624 | 235.2016 | 181.0998 | 232.9185 | 161.2608
2012-10-30 19:19 | 28.6802 | 249.6656 | 195.548 | 247.3903 | 175.7405 | -1.598 | 1.2711 | 32.1452 | 40.8613 | 25.2883 | 2.0083 | 32.0688 | 64.4036 | 15.9682 | 14.4875 | -0.1102 | 105.2649 | 235.1781 | 181.0605 | 232.9028 | 161.253
2012-10-30 19:19 | 28.6899 | 249.691 | 195.5224 | 247.353 | 175.719 | -1.598 | 1.2712 | 32.1453 | 40.8617 | 25.2883 | 2.0083 | 32.0689 | 64.4038 | 15.9683 | 14.4972 | -0.1102 | 105.2654 | 235.1938 | 181.0252 | 232.8558 | 161.2218
2012-10-30 19:19 | 28.6765 | 249.6697 | 195.4893 | 247.4062 | 175.6315 | -1.5997 | 1.2712 | 32.1455 | 40.8623 | 25.2882 | 2.0083 | 32.0692 | 64.4045 | 15.9684 | 14.5165 | -0.1103 | 105.2668 | 235.1532 | 180.9728 | 232.8897 | 161.115
2012-10-30 19:19 | 28.6839 | 249.7006 | 195.4888 | 247.4449 | 175.6271 | -1.6003 | 1.2712 | 32.1456 | 40.8623 | 25.2884 | 2.0083 | 32.0694 | 64.4055 | 15.9683 | 14.532 | -0.1103 | 105.2678 | 235.1686 | 180.9568 | 232.9129 | 161.0951
2012-10-30 19:20 | 28.6854 | 249.6826 | 195.4668 | 247.466 | 175.7106 | -1.5991 | 1.2712 | 32.1457 | 40.8624 | 25.2883 | 2.0083 | 32.0692 | 64.405 | 15.9682 | 14.5172 | -0.1102 | 105.2674 | 235.1654 | 180.9496 | 232.9488 | 161.1934
2012-10-30 19:20 | 28.6862 | 249.7029 | 195.4558 | 247.3806 | 175.7504 | -1.5997 | 1.2712 | 32.1458 | 40.8627 | 25.2884 | 2.0083 | 32.0694 | 64.4061 | 15.9682 | 14.5017 | -0.1103 | 105.2689 | 235.2012 | 180.9541 | 232.8789 | 161.2487
2012-10-30 19:20 | 28.6884 | 249.6582 | 195.4973 | 247.4024 | 175.7448 | -1.5968 | 1.2712 | 32.1461 | 40.8638 | 25.2882 | 2.0084 | 32.0698 | 64.4075 | 15.9682 | 14.5039 | -0.1101 | 105.2712 | 235.1543 | 180.9934 | 232.8985 | 161.2409
2012-10-30 19:20 | 28.6817 | 249.6945 | 195.4474 | 247.4114 | 175.6484 | -1.6026 | 1.2712 | 32.1462 | 40.8638 | 25.2883 | 2.0084 | 32.0699 | 64.4076 | 15.9683 | 14.4849 | -0.1105 | 105.2714 | 235.2096 | 180.9625 | 232.9265 | 161.1635
2012-10-30 19:20 | 28.6757 | 249.6924 | 195.4256 | 247.4641 | 175.6111 | -1.5991 | 1.2712 | 32.1458 | 40.8634 | 25.288 | 2.0083 | 32.0694 | 64.4064 | 15.9681 | 14.4912 | -0.1102 | 105.2698 | 235.2012 | 180.9344 | 232.9729 | 161.1199
2012-10-30 19:20 | 28.6832 | 249.653 | 195.4331 | 247.4638 | 175.7591 | -1.5997 | 1.2712 | 32.1462 | 40.8638 | 25.2883 | 2.0084 | 32.0698 | 64.408 | 15.9681 | 14.5191 | -0.1103 | 105.2719 | 235.1339 | 180.914 | 232.9447 | 161.24
2012-10-30 19:20 | 28.678 | 249.6556 | 195.4554 | 247.4233 | 175.793 | -1.6043 | 1.2712 | 32.1466 | 40.8646 | 25.2884 | 2.0084 | 32.0702 | 64.4098 | 15.9681 | 14.4853 | -0.1106 | 105.2744 | 235.1703 | 180.9701 | 232.938 | 161.3077
2012-10-30 19:20 | 28.6832 | 249.7234 | 195.4292 | 247.3972 | 175.7552 | -1.5997 | 1.2712 | 32.1461 | 40.8639 | 25.2881 | 2.0084 | 32.0698 | 64.4083 | 15.968 | 14.5068 | -0.1103 | 105.2722 | 235.2166 | 180.9224 | 232.8904 | 161.2484
2012-10-30 19:20 | 28.6795 | 249.661 | 195.3784 | 247.464 | 175.654 | -1.6037 | 1.2712 | 32.1461 | 40.8636 | 25.2883 | 2.0084 | 32.0697 | 64.4077 | 15.968 | 14.4868 | -0.1106 | 105.2713 | 235.1742 | 180.8916 | 232.9772 | 161.1672
2012-10-30 19:20 | 28.6802 | 249.6852 | 195.4262 | 247.5196 | 175.7327 | -1.6032 | 1.2712 | 32.1459 | 40.8636 | 25.2879 | 2.0084 | 32.0695 | 64.4073 | 15.968 | 14.4712 | -0.1105 | 105.2709 | 235.214 | 180.955 | 233.0484 | 161.2615
2012-10-30 19:20 | 28.681 | 249.7095 | 195.4191 | 247.4224 | 175.8545 | -1.6014 | 1.2712 | 32.1456 | 40.8624 | 25.2883 | 2.0084 | 32.0693 | 64.4063 | 15.968 | 14.423 | -0.1104 | 105.2687 | 235.2865 | 180.9961 | 232.9994 | 161.4315
2012-10-30 19:20 | 28.6899 | 249.7066 | 195.4163 | 247.4 | 175.7463 | -1.602 | 1.2712 | 32.1454 | 40.8619 | 25.2883 | 2.0084 | 32.069 | 64.4058 | 15.9678 | 14.4931 | -0.1104 | 105.2677 | 235.2135 | 180.9232 | 232.9069 | 161.2532
2012-10-30 19:20 | 28.6877 | 249.7397 | 195.418 | 247.4291 | 175.7285 | -1.6003 | 1.2711 | 32.145 | 40.8607 | 25.2883 | 2.0083 | 32.0688 | 64.4041 | 15.968 | 14.4542 | -0.1103 | 105.2648 | 235.2855 | 180.9638 | 232.9749 | 161.2743
2012-10-30 19:20 | 28.6907 | 249.7192 | 195.3778 | 247.5026 | 175.7198 | -1.6061 | 1.2711 | 32.1448 | 40.8608 | 25.2881 | 2.0083 | 32.0685 | 64.4035 | 15.9679 | 14.4654 | -0.1107 | 105.2643 | 235.2538 | 180.9124 | 233.0372 | 161.2544
2012-10-30 19:20 | 28.6795 | 249.7197 | 195.3941 | 247.5189 | 175.6423 | -1.5991 | 1.2711 | 32.1445 | 40.8602 | 25.2878 | 2.0083 | 32.0681 | 64.4022 | 15.9679 | 14.4542 | -0.1102 | 105.2624 | 235.2655 | 180.9399 | 233.0647 | 161.1881
2012-10-30 19:20 | 28.6847 | 249.6936 | 195.3953 | 247.4575 | 175.7606 | -1.6003 | 1.2711 | 32.1442 | 40.8592 | 25.2881 | 2.0083 | 32.0679 | 64.4012 | 15.9679 | 14.5083 | -0.1103 | 105.2604 | 235.1853 | 180.887 | 232.9492 | 161.2523
2012-10-30 19:20 | 28.6832 | 249.6882 | 195.3781 | 247.4325 | 175.6303 | -1.602 | 1.2711 | 32.1436 | 40.8577 | 25.288 | 2.0082 | 32.0673 | 64.3991 | 15.9678 | 14.4783 | -0.1104 | 105.2568 | 235.2099 | 180.8998 | 232.9542 | 161.152
2012-10-30 19:20 | 28.6832 | 249.6843 | 195.3781 | 247.4168 | 175.6655 | -1.6009 | 1.2711 | 32.1433 | 40.857 | 25.288 | 2.0082 | 32.0669 | 64.397 | 15.9679 | 14.5068 | -0.1104 | 105.254 | 235.1775 | 180.8713 | 232.91 | 161.1587
2012-10-30 19:20 | 28.6862 | 249.7069 | 195.3654 | 247.4002 | 175.7504 | -1.6009 | 1.2711 | 32.1428 | 40.856 | 25.2879 | 2.0082 | 32.0665 | 64.396 | 15.9678 | 14.4894 | -0.1104 | 105.252 | 235.2175 | 180.876 | 232.9108 | 161.261
2012-10-30 19:20 | 28.675 | 249.6839 | 195.366 | 247.4948 | 175.6456 | -1.6032 | 1.2711 | 32.1426 | 40.8552 | 25.2881 | 2.0082 | 32.0663 | 64.395 | 15.9678 | 14.4782 | -0.1105 | 105.2502 | 235.2057 | 180.8878 | 233.0166 | 161.1674
2012-10-30 19:20 | 28.6765 | 249.7402 | 195.3439 | 247.4884 | 175.6158 | -1.6014 | 1.271 | 32.1422 | 40.8539 | 25.2882 | 2.0082 | 32.0658 | 64.3934 | 15.9678 | 14.4838 | -0.1104 | 105.2473 | 235.2564 | 180.8601 | 233.0046 | 161.132
2012-10-30 19:20 | 28.6862 | 249.7108 | 195.3654 | 247.4864 | 175.5163 | -1.598 | 1.271 | 32.1415 | 40.8522 | 25.2881 | 2.0081 | 32.0652 | 64.3909 | 15.9677 | 14.4854 | -0.1102 | 105.2431 | 235.2254 | 180.88 | 233.001 | 161.0309
2012-10-30 19:21 | 28.6854 | 249.7178 | 195.3371 | 247.4072 | 175.5974 | -1.598 | 1.271 | 32.1412 | 40.8521 | 25.2877 | 2.0081 | 32.0648 | 64.3896 | 15.9677 | 14.4805 | -0.1102 | 105.2417 | 235.2373 | 180.8566 | 232.9267 | 161.1169
2012-10-30 19:21 | 28.6929 | 249.7488 | 195.3289 | 247.4461 | 175.5698 | -1.6055 | 1.271 | 32.1412 | 40.852 | 25.2879 | 2.0081 | 32.0649 | 64.3898 | 15.9677 | 14.4676 | -0.1107 | 105.2417 | 235.2812 | 180.8613 | 232.9785 | 161.1022
2012-10-30 19:21 | 28.678 | 249.6712 | 195.3218 | 247.4939 | 175.5315 | -1.6014 | 1.271 | 32.1408 | 40.8506 | 25.288 | 2.0081 | 32.0645 | 64.3886 | 15.9676 | 14.4935 | -0.1104 | 105.2392 | 235.1777 | 180.8283 | 233.0004 | 161.038
2012-10-30 19:21 | 28.6824 | 249.7539 | 195.3302 | 247.4826 | 175.5203 | -1.6095 | 1.271 | 32.1407 | 40.8503 | 25.288 | 2.0081 | 32.0643 | 64.3875 | 15.9676 | 14.4612 | -0.111 | 105.2378 | 235.2927 | 180.869 | 233.0214 | 161.0591
2012-10-30 19:21 | 28.6862 | 249.7225 | 195.3379 | 247.4159 | 175.5241 | -1.6009 | 1.271 | 32.1402 | 40.8492 | 25.288 | 2.008 | 32.064 | 64.3859 | 15.9678 | 14.4772 | -0.1104 | 105.2351 | 235.2453 | 180.8607 | 232.9387 | 161.0469
2012-10-30 19:21 | 28.6772 | 249.7644 | 195.2975 | 247.4304 | 175.5346 | -1.598 | 1.271 | 32.1397 | 40.8482 | 25.2878 | 2.008 | 32.0634 | 64.3841 | 15.9676 | 14.4927 | -0.1102 | 105.2323 | 235.2717 | 180.8048 | 232.9377 | 161.0419
2012-10-30 19:21 | 28.6757 | 249.7355 | 195.2803 | 247.4837 | 175.5253 | -1.6055 | 1.2709 | 32.1393 | 40.8473 | 25.2878 | 2.008 | 32.0631 | 64.3826 | 15.9677 | 14.4912 | -0.1107 | 105.2299 | 235.2443 | 180.7891 | 232.9925 | 161.0341
2012-10-30 19:21 | 28.6832 | 249.7391 | 195.2995 | 247.4951 | 175.5055 | -1.602 | 1.2709 | 32.1391 | 40.8461 | 25.2881 | 2.008 | 32.0628 | 64.3821 | 15.9675 | 14.462 | -0.1104 | 105.2283 | 235.2771 | 180.8375 | 233.0331 | 161.0435
2012-10-30 19:21 | 28.6787 | 249.762 | 195.2793 | 247.5768 | 175.5829 | -1.602 | 1.2709 | 32.1382 | 40.8444 | 25.2878 | 2.0079 | 32.0619 | 64.3785 | 15.9675 | 14.4534 | -0.1104 | 105.2229 | 235.3086 | 180.8259 | 233.1234 | 161.1295
2012-10-30 19:21 | 28.6832 | 249.7978 | 195.3153 | 247.4246 | 175.6342 | -1.6009 | 1.2709 | 32.1378 | 40.843 | 25.288 | 2.0079 | 32.0615 | 64.377 | 15.9675 | 14.5191 | -0.1104 | 105.22 | 235.2787 | 180.7962 | 232.9055 | 161.1151
2012-10-30 19:21 | 28.6765 | 249.748 | 195.2889 | 247.4454 | 175.6471 | -1.602 | 1.2673 | 32.0459 | 40.6108 | 25.2874 | 2.002 | 31.9668 | 63.9966 | 15.9676 | 14.4879 | -0.1104 | 104.6074 | 235.2601 | 180.801 | 232.9575 | 161.1592
2012-10-30 19:21 | 28.6772 | 249.737 | 195.2582 | 247.5244 | 175.5658 | -1.6009 | 1.2648 | 31.9813 | 40.4489 | 25.2863 | 1.9981 | 31.9052 | 63.7507 | 15.9676 | 14.4723 | -0.1104 | 104.1996 | 235.2647 | 180.7859 | 233.0521 | 161.0935
2012-10-30 19:21 | 28.6802 | 249.7048 | 195.2612 | 247.4961 | 175.6547 | -1.6026 | 1.2641 | 31.9631 | 40.4031 | 25.2862 | 1.997 | 31.8873 | 63.6784 | 15.9677 | 14.512 | -0.1105 | 104.0815 | 235.1928 | 180.7492 | 232.9841 | 161.1427
2012-10-30 19:21 | 28.6795 | 249.6767 | 195.2683 | 247.4523 | 175.6969 | -1.6049 | 1.2636 | 31.9523 | 40.3761 | 25.2861 | 1.9963 | 31.8768 | 63.6365 | 15.9677 | 14.4583 | -0.1106 | 104.0126 | 235.2184 | 180.81 | 232.994 | 161.2386
2012-10-30 19:21 | 28.6877 | 249.677 | 195.249 | 247.4448 | 175.2836 | -1.5957 | 1.2634 | 31.9476 | 40.3641 | 25.286 | 1.996 | 31.8718 | 63.6158 | 15.9679 | 14.4787 | -0.11 | 103.9799 | 235.1983 | 180.7703 | 232.9661 | 160.8049
2012-10-30 19:21 | 28.6802 | 249.6578 | 195.2101 | 247.4373 | 175.4478 | -1.6043 | 1.2635 | 31.9486 | 40.3666 | 25.2861 | 1.996 | 31.8728 | 63.6196 | 15.9679 | 14.4671 | -0.1106 | 103.9862 | 235.1907 | 180.743 | 232.9702 | 160.9807
2012-10-30 19:21 | 28.6772 | 249.6078 | 195.2228 | 247.4696 | 175.4565 | -1.5991 | 1.2636 | 31.9506 | 40.3714 | 25.2862 | 1.9962 | 31.8747 | 63.6269 | 15.968 | 14.4478 | -0.1102 | 103.9983 | 235.16 | 180.775 | 233.0218 | 161.0087
2012-10-30 19:21 | 28.6705 | 249.605 | 195.2004 | 247.5334 | 175.5981 | -1.6043 | 1.2636 | 31.9516 | 40.3738 | 25.2863 | 1.9962 | 31.8758 | 63.6311 | 15.9681 | 14.4697 | -0.1106 | 104.0049 | 235.1353 | 180.7307 | 233.0637 | 161.1284
2012-10-30 19:21 | 28.6742 | 249.6283 | 195.208 | 247.4195 | 175.836 | -1.5985 | 1.2637 | 31.9553 | 40.3833 | 25.2862 | 1.9965 | 31.8794 | 63.6456 | 15.968 | 14.4774 | -0.1102 | 104.0289 | 235.1509 | 180.7306 | 232.9421 | 161.3586
2012-10-30 19:21 | 28.6757 | 249.6024 | 195.2449 | 247.421 | 175.7595 | -1.5985 | 1.2655 | 31.9995 | 40.4948 | 25.2864 | 1.9992 | 31.9236 | 63.8226 | 15.968 | 14.4545 | -0.1102 | 104.3174 | 235.1479 | 180.7904 | 232.9665 | 161.305
2012-10-30 19:21 | 28.6712 | 249.5822 | 195.2483 | 247.3382 | 175.8564 | -1.6009 | 1.2688 | 32.0838 | 40.7065 | 25.2875 | 2.0045 | 32.0076 | 64.1583 | 15.9681 | 14.4663 | -0.1104 | 104.8648 | 235.1159 | 180.782 | 232.8719 | 161.3901
2012-10-30 19:21 | 28.6645 | 249.552 | 195.3044 | 247.3981 | 175.8341 | -1.6037 | 1.2695 | 32.1019 | 40.752 | 25.2878 | 2.0056 | 32.0256 | 64.231 | 15.968 | 14.4473 | -0.1106 | 104.9831 | 235.1047 | 180.8571 | 232.9508 | 161.3868
2012-10-30 19:21 | 28.6712 | 249.5901 | 195.319 | 247.4675 | 175.8213 | -1.5991 | 1.27 | 32.116 | 40.7881 | 25.2878 | 2.0065 | 32.0399 | 64.2888 | 15.9679 | 14.4704 | -0.1102 | 105.0768 | 235.1197 | 180.8486 | 232.9971 | 161.3509
2012-10-30 19:22 | 28.6675 | 249.6216 | 195.3978 | 247.3893 | 175.9152 | -1.6014 | 1.2704 | 32.1262 | 40.8137 | 25.288 | 2.0071 | 32.0499 | 64.3286 | 15.9679 | 14.4911 | -0.1104 | 105.1422 | 235.1305 | 180.9067 | 232.8982 | 161.4241
2012-10-30 19:22 | 28.6667 | 249.6091 | 195.3813 | 247.3454 | 175.8675 | -1.602 | 1.2707 | 32.133 | 40.831 | 25.2879 | 2.0076 | 32.0567 | 64.3557 | 15.968 | 14.4903 | -0.1104 | 105.1867 | 235.1188 | 180.891 | 232.8551 | 161.3772
2012-10-30 19:22 | 28.6675 | 249.602 | 195.3703 | 247.3423 | 175.884 | -1.6043 | 1.2708 | 32.137 | 40.8406 | 25.2883 | 2.0078 | 32.0606 | 64.3719 | 15.9679 | 14.4952 | -0.1106 | 105.2124 | 235.1068 | 180.8751 | 232.8471 | 161.3888
2012-10-30 19:22 | 28.6645 | 249.6147 | 195.3791 | 247.3707 | 175.799 | -1.6026 | 1.2709 | 32.1393 | 40.8465 | 25.2882 | 2.008 | 32.0629 | 64.3814 | 15.9678 | 14.4881 | -0.1105 | 105.2279 | 235.1266 | 180.891 | 232.8826 | 161.3109
2012-10-30 19:22 | 28.6652 | 249.5841 | 195.4112 | 247.4771 | 175.7138 | -1.6014 | 1.271 | 32.1412 | 40.8513 | 25.2881 | 2.0081 | 32.0648 | 64.3887 | 15.9679 | 14.5011 | -0.1104 | 105.24 | 235.083 | 180.9101 | 232.976 | 161.2127
2012-10-30 19:22 | 28.6585 | 249.6439 | 195.4124 | 247.3647 | 175.5276 | -1.6003 | 1.271 | 32.1425 | 40.8547 | 25.2882 | 2.0082 | 32.0661 | 64.3944 | 15.9678 | 14.4862 | -0.1103 | 105.249 | 235.1577 | 180.9262 | 232.8785 | 161.0414
2012-10-30 19:22 | 28.6585 | 249.6517 | 195.3691 | 247.3568 | 175.6291 | -1.6037 | 1.2711 | 32.1437 | 40.858 | 25.2881 | 2.0083 | 32.0673 | 64.3993 | 15.9678 | 14.4781 | -0.1106 | 105.2573 | 235.1736 | 180.891 | 232.8787 | 161.151
2012-10-30 19:22 | 28.6592 | 249.6446 | 195.3463 | 247.2753 | 175.5595 | -1.6032 | 1.2711 | 32.1443 | 40.8592 | 25.2882 | 2.0083 | 32.0679 | 64.4016 | 15.9678 | 14.5032 | -0.1105 | 105.2609 | 235.1414 | 180.8431 | 232.7721 | 161.0563
2012-10-30 19:22 | 28.6652 | 249.6232 | 195.2894 | 247.387 | 175.4601 | -1.5991 | 1.2711 | 32.1449 | 40.8607 | 25.2883 | 2.0083 | 32.0685 | 64.4046 | 15.9676 | 14.4562 | -0.1102 | 105.2653 | 235.167 | 180.8332 | 232.9308 | 161.0039
2012-10-30 19:22 | 28.6652 | 249.6271 | 195.3012 | 247.4184 | 175.4796 | -1.6009 | 1.2712 | 32.1452 | 40.8618 | 25.288 | 2.0084 | 32.0689 | 64.4058 | 15.9677 | 14.4766 | -0.1104 | 105.2676 | 235.1505 | 180.8246 | 232.9418 | 161.003
2012-10-30 19:22 | 28.6682 | 249.6262 | 195.2806 | 247.3587 | 175.5334 | -1.6032 | 1.2712 | 32.1455 | 40.862 | 25.2883 | 2.0084 | 32.0691 | 64.407 | 15.9676 | 14.4878 | -0.1105 | 105.269 | 235.1384 | 180.7928 | 232.8709 | 161.0456
2012-10-30 19:22 | 28.6637 | 249.6687 | 195.284 | 247.3503 | 175.603 | -1.6032 | 1.2712 | 32.1457 | 40.8623 | 25.2885 | 2.0084 | 32.0693 | 64.408 | 15.9676 | 14.4833 | -0.1105 | 105.2703 | 235.1854 | 180.8007 | 232.867 | 161.1197
2012-10-30 19:22 | 28.663 | 249.7071 | 195.2558 | 247.33 | 175.4618 | -1.6014 | 1.2712 | 32.1461 | 40.8635 | 25.2883 | 2.0084 | 32.0696 | 64.4093 | 15.9676 | 14.4826 | -0.1104 | 105.2728 | 235.2245 | 180.7732 | 232.8474 | 160.9792
2012-10-30 19:22 | 28.6637 | 249.6883 | 195.2604 | 247.3895 | 175.4313 | -1.602 | 1.2712 | 32.1461 | 40.864 | 25.2882 | 2.0084 | 32.0698 | 64.4105 | 15.9675 | 14.4914 | -0.1104 | 105.2745 | 235.1969 | 180.769 | 232.8981 | 160.9399
2012-10-30 19:22 | 28.6607 | 249.6579 | 195.2849 | 247.4491 | 175.3581 | -1.6003 | 1.2712 | 32.1458 | 40.8631 | 25.2881 | 2.0084 | 32.0695 | 64.4085 | 15.9676 | 14.468 | -0.1103 | 105.2716 | 235.1899 | 180.8169 | 232.9811 | 160.8901
2012-10-30 19:22 | 28.6615 | 249.6782 | 195.2307 | 247.4147 | 175.5072 | -1.6026 | 1.2712 | 32.1456 | 40.8621 | 25.2885 | 2.0084 | 32.0692 | 64.4079 | 15.9675 | 14.4892 | -0.1105 | 105.27 | 235.189 | 180.7415 | 232.9255 | 161.018
2012-10-30 19:22 | 28.6645 | 249.7008 | 195.2376 | 247.2962 | 175.6038 | -1.6055 | 1.2712 | 32.1458 | 40.8628 | 25.2883 | 2.0084 | 32.0694 | 64.4086 | 15.9675 | 14.4555 | -0.1107 | 105.2714 | 235.2453 | 180.7821 | 232.8407 | 161.1483
2012-10-30 19:22 | 28.6592 | 249.6681 | 195.2402 | 247.3536 | 175.4619 | -1.5968 | 1.2712 | 32.1456 | 40.8624 | 25.2882 | 2.0084 | 32.0691 | 64.4077 | 15.9675 | 14.491 | -0.1101 | 105.2701 | 235.1771 | 180.7492 | 232.8626 | 160.9709
2012-10-30 19:22 | 28.666 | 249.6279 | 195.188 | 247.427 | 175.4726 | -1.6003 | 1.2711 | 32.1448 | 40.8605 | 25.2882 | 2.0084 | 32.0684 | 64.4049 | 15.9675 | 14.457 | -0.1103 | 105.2654 | 235.1709 | 180.731 | 232.97 | 161.0156
2012-10-30 19:22 | 28.6622 | 249.6907 | 195.2275 | 247.388 | 175.4806 | -1.602 | 1.2711 | 32.1447 | 40.8599 | 25.2884 | 2.0084 | 32.0684 | 64.4045 | 15.9675 | 14.4654 | -0.1104 | 105.2644 | 235.2253 | 180.7621 | 232.9226 | 161.0152
2012-10-30 19:22 | 28.6607 | 249.6814 | 195.2338 | 247.4256 | 175.5376 | -1.6043 | 1.2711 | 32.1449 | 40.8602 | 25.2885 | 2.0084 | 32.0685 | 64.4053 | 15.9674 | 14.4639 | -0.1106 | 105.2655 | 235.2175 | 180.7699 | 232.9617 | 161.0737
2012-10-30 19:22 | 28.657 | 249.7051 | 195.2105 | 247.3671 | 175.4675 | -1.5997 | 1.2711 | 32.1447 | 40.8599 | 25.2884 | 2.0084 | 32.0683 | 64.4045 | 15.9675 | 14.448 | -0.1103 | 105.2644 | 235.2571 | 180.7625 | 232.9191 | 161.0195
2012-10-30 19:23 | 28.66 | 249.7041 | 195.1978 | 247.3544 | 175.5096 | -1.6003 | 1.2711 | 32.1442 | 40.8585 | 25.2884 | 2.0083 | 32.0679 | 64.403 | 15.9674 | 14.4673 | -0.1103 | 105.2615 | 235.2368 | 180.7305 | 232.8871 | 161.0423
2012-10-30 19:23 | 28.6562 | 249.6925 | 195.2018 | 247.4564 | 175.4238 | -1.598 | 1.2711 | 32.1441 | 40.8589 | 25.2882 | 2.0083 | 32.0678 | 64.4029 | 15.9673 | 14.4717 | -0.1102 | 105.2618 | 235.2208 | 180.7301 | 232.9847 | 160.9521
2012-10-30 19:23 | 28.6562 | 249.7082 | 195.2254 | 247.4368 | 175.5487 | -1.6014 | 1.2711 | 32.144 | 40.8583 | 25.2883 | 2.0083 | 32.0677 | 64.4023 | 15.9674 | 14.4758 | -0.1104 | 105.2607 | 235.2324 | 180.7496 | 232.961 | 161.0729
2012-10-30 19:23 | 28.6547 | 249.6754 | 195.2475 | 247.357 | 175.7228 | -1.6032 | 1.2711 | 32.1436 | 40.8577 | 25.288 | 2.0083 | 32.0672 | 64.4004 | 15.9673 | 14.4783 | -0.1105 | 105.2581 | 235.1971 | 180.7692 | 232.8787 | 161.2445
2012-10-30 19:23 | 28.66 | 249.7041 | 195.2567 | 247.3583 | 175.615 | -1.6026 | 1.2711 | 32.1433 | 40.8564 | 25.2884 | 2.0083 | 32.067 | 64.4003 | 15.9672 | 14.4388 | -0.1105 | 105.2567 | 235.2653 | 180.8179 | 232.9195 | 161.1762
2012-10-30 19:23 | 28.654 | 249.7216 | 195.2586 | 247.4464 | 175.6909 | -1.5974 | 1.2711 | 32.1431 | 40.8557 | 25.2884 | 2.0083 | 32.0666 | 64.3985 | 15.9673 | 14.4776 | -0.1101 | 105.2542 | 235.244 | 180.781 | 232.9688 | 161.2133
2012-10-30 19:23 | 28.6607 | 249.6735 | 195.281 | 247.4648 | 175.6664 | -1.6014 | 1.2711 | 32.1428 | 40.8558 | 25.288 | 2.0083 | 32.0665 | 64.3985 | 15.9671 | 14.5047 | -0.1104 | 105.2543 | 235.1688 | 180.7763 | 232.9601 | 161.1617
2012-10-30 19:23 | 28.6577 | 249.6901 | 195.3369 | 247.4383 | 175.6361 | -1.5985 | 1.2711 | 32.1427 | 40.8555 | 25.288 | 2.0082 | 32.0663 | 64.3971 | 15.9673 | 14.4487 | -0.1102 | 105.2527 | 235.2414 | 180.8882 | 232.9896 | 161.1874
2012-10-30 19:23 | 28.654 | 249.7099 | 195.2703 | 247.3641 | 175.5855 | -1.5997 | 1.2711 | 32.1427 | 40.8554 | 25.2881 | 2.0083 | 32.0663 | 64.3979 | 15.9671 | 14.4817 | -0.1103 | 105.2533 | 235.2282 | 180.7886 | 232.8824 | 161.1038
2012-10-30 19:23 | 28.66 | 249.7081 | 195.2488 | 247.4719 | 175.5954 | -1.5962 | 1.271 | 32.1423 | 40.8538 | 25.2884 | 2.0082 | 32.0658 | 64.3957 | 15.9672 | 14.4306 | -0.11 | 105.2495 | 235.2775 | 180.8182 | 233.0413 | 161.1648
2012-10-30 19:23 | 28.6697 | 249.7099 | 195.2546 | 247.4621 | 175.609 | -1.5974 | 1.271 | 32.1414 | 40.8518 | 25.2883 | 2.0082 | 32.0651 | 64.3933 | 15.9671 | 14.4729 | -0.1101 | 105.2451 | 235.237 | 180.7817 | 232.9892 | 161.1361
2012-10-30 19:23 | 28.6645 | 249.6969 | 195.3044 | 247.3981 | 175.5804 | -1.5957 | 1.271 | 32.141 | 40.8513 | 25.288 | 2.0082 | 32.0646 | 64.3909 | 15.9671 | 14.4963 | -0.11 | 105.2422 | 235.2006 | 180.8081 | 232.9018 | 161.0841
2012-10-30 19:23 | 28.6675 | 249.7273 | 195.3074 | 247.3737 | 175.6654 | -1.5985 | 1.271 | 32.1413 | 40.8512 | 25.2884 | 2.0082 | 32.0649 | 64.3924 | 15.9671 | 14.5238 | -0.1102 | 105.2436 | 235.2035 | 180.7836 | 232.8499 | 161.1416
2012-10-30 19:23 | 28.6585 | 249.7066 | 195.2748 | 247.4665 | 175.5471 | -1.6078 | 1.271 | 32.1408 | 40.85 | 25.2884 | 2.0082 | 32.0644 | 64.3903 | 15.9671 | 14.4699 | -0.1108 | 105.2403 | 235.2367 | 180.8049 | 232.9966 | 161.0772
2012-10-30 19:23 | 28.6615 | 249.7056 | 195.3132 | 247.4578 | 175.5111 | -1.5985 | 1.271 | 32.14 | 40.8484 | 25.2881 | 2.0081 | 32.0637 | 64.3873 | 15.9671 | 14.5137 | -0.1102 | 105.2357 | 235.1919 | 180.7995 | 232.9441 | 160.9974
2012-10-30 19:23 | 28.6615 | 249.7213 | 195.2857 | 247.5205 | 175.4018 | -1.6014 | 1.2709 | 32.1398 | 40.8479 | 25.2881 | 2.0081 | 32.0636 | 64.3871 | 15.967 | 14.4892 | -0.1104 | 105.235 | 235.2321 | 180.7965 | 233.0313 | 160.9126
2012-10-30 19:23 | 28.6615 | 249.7096 | 195.2386 | 247.3559 | 175.4369 | -1.6014 | 1.2709 | 32.1397 | 40.8473 | 25.2884 | 2.0081 | 32.0633 | 64.3862 | 15.967 | 14.477 | -0.1104 | 105.2335 | 235.2326 | 180.7616 | 232.8789 | 160.9599
2012-10-30 19:23 | 28.6622 | 249.6985 | 195.255 | 247.4898 | 175.461 | -1.6009 | 1.2709 | 32.1392 | 40.8464 | 25.2881 | 2.0081 | 32.0627 | 64.3839 | 15.967 | 14.4777 | -0.1104 | 105.2303 | 235.2208 | 180.7773 | 233.0121 | 160.9833
2012-10-30 19:23 | 28.666 | 249.7493 | 195.2706 | 247.4936 | 175.4258 | -1.6043 | 1.2709 | 32.1389 | 40.8457 | 25.2881 | 2.008 | 32.0625 | 64.383 | 15.967 | 14.5345 | -0.1106 | 105.2287 | 235.2148 | 180.7361 | 232.9591 | 160.8913
2012-10-30 19:23 | 28.6682 | 249.7515 | 195.257 | 247.4527 | 175.428 | -1.602 | 1.2709 | 32.1385 | 40.8445 | 25.2882 | 2.008 | 32.0622 | 64.3819 | 15.9669 | 14.5286 | -0.1104 | 105.2264 | 235.2229 | 180.7284 | 232.9241 | 160.8994
2012-10-30 19:23 | 28.672 | 249.7631 | 195.253 | 247.4252 | 175.4552 | -1.6003 | 1.2709 | 32.1383 | 40.8441 | 25.288 | 2.008 | 32.0619 | 64.3809 | 15.967 | 14.5201 | -0.1103 | 105.225 | 235.243 | 180.7329 | 232.9051 | 160.9351
2012-10-30 19:23 | 28.6742 | 249.7183 | 195.267 | 247.5332 | 175.4613 | -1.5968 | 1.2709 | 32.1378 | 40.8429 | 25.2881 | 2.008 | 32.0614 | 64.3791 | 15.9669 | 14.4938 | -0.1101 | 105.2221 | 235.2245 | 180.7732 | 233.0394 | 160.9675
2012-10-30 19:23 | 28.6645 | 249.7047 | 195.2337 | 247.4764 | 175.4321 | -1.5974 | 1.2709 | 32.1373 | 40.8419 | 25.288 | 2.008 | 32.061 | 64.3775 | 15.9669 | 14.4514 | -0.1101 | 105.2195 | 235.2533 | 180.7823 | 233.025 | 160.9807
2012-10-30 19:24 | 28.6742 | 249.7223 | 195.1884 | 247.4901 | 175.5238 | -1.6037 | 1.2701 | 32.1169 | 40.7901 | 25.2879 | 2.0071 | 32.0469 | 64.3206 | 15.9669 | 14.5264 | -0.1106 | 105.1107 | 235.1959 | 180.662 | 232.9637 | 160.9974
2012-10-30 19:24 | 28.6757 | 249.7003 | 195.2292 | 247.4485 | 175.4823 | -1.5997 | 1.2645 | 31.9745 | 40.4314 | 25.2865 | 1.9977 | 31.897 | 63.7201 | 15.967 | 14.5157 | -0.1103 | 104.1515 | 235.1846 | 180.7135 | 232.9328 | 160.9666
2012-10-30 19:24 | 28.6667 | 249.6913 | 195.2123 | 247.6001 | 175.4733 | -1.6014 | 1.2646 | 31.9768 | 40.4372 | 25.2865 | 1.9979 | 31.9009 | 63.7356 | 15.967 | 14.523 | -0.1104 | 104.1728 | 235.1683 | 180.6893 | 233.0771 | 160.9503
2012-10-30 19:24 | 28.666 | 249.6475 | 195.1959 | 247.4113 | 175.7302 | -1.598 | 1.2639 | 31.9601 | 40.3954 | 25.2862 | 1.9969 | 31.8842 | 63.6687 | 15.9671 | 14.4978 | -0.1102 | 104.0642 | 235.1497 | 180.6981 | 232.9135 | 161.2324
2012-10-30 19:24 | 28.6697 | 249.6708 | 195.1957 | 247.4346 | 175.4959 | -1.5968 | 1.2636 | 31.9519 | 40.3749 | 25.2861 | 1.9964 | 31.876 | 63.6362 | 15.9671 | 14.526 | -0.1101 | 104.0111 | 235.1448 | 180.6697 | 232.9086 | 160.9699
2012-10-30 19:24 | 28.6615 | 249.5921 | 195.1796 | 247.4539 | 175.3784 | -1.5997 | 1.2636 | 31.9516 | 40.3745 | 25.2859 | 1.9963 | 31.8758 | 63.6344 | 15.9673 | 14.5423 | -0.1103 | 104.0089 | 235.0498 | 180.6373 | 232.9116 | 160.8361
2012-10-30 19:24 | 28.6525 | 249.5909 | 195.1745 | 247.5271 | 175.5333 | -1.6037 | 1.2635 | 31.9497 | 40.3697 | 25.286 | 1.9962 | 31.8738 | 63.6268 | 15.9672 | 14.5006 | -0.1106 | 103.9965 | 235.0903 | 180.6739 | 233.0265 | 161.0327
2012-10-30 19:24 | 28.6532 | 249.5721 | 195.1556 | 247.3241 | 175.7018 | -1.6032 | 1.2636 | 31.9505 | 40.3712 | 25.2862 | 1.9962 | 31.8746 | 63.6297 | 15.9673 | 14.485 | -0.1105 | 104.0009 | 235.0871 | 180.6706 | 232.8391 | 161.2168
2012-10-30 19:24 | 28.6465 | 249.538 | 195.1843 | 247.4741 | 175.6288 | -1.6066 | 1.2637 | 31.9534 | 40.3786 | 25.2862 | 1.9964 | 31.8775 | 63.6411 | 15.9672 | 14.5232 | -0.1108 | 104.0197 | 235.0148 | 180.6611 | 232.9509 | 161.1056
2012-10-30 19:24 | 28.6495 | 249.5723 | 195.2226 | 247.4379 | 175.7801 | -1.602 | 1.2638 | 31.9564 | 40.3867 | 25.2859 | 1.9966 | 31.8805 | 63.6528 | 15.9673 | 14.5547 | -0.1104 | 104.0395 | 235.0176 | 180.6679 | 232.8832 | 161.2254
2012-10-30 19:24 | 28.6442 | 249.5748 | 195.237 | 247.37 | 175.7397 | -1.5974 | 1.2639 | 31.9591 | 40.3927 | 25.2863 | 1.9968 | 31.8832 | 63.6634 | 15.9674 | 14.5658 | -0.1101 | 104.0562 | 235.009 | 180.6712 | 232.8042 | 161.1739
2012-10-30 19:24 | 28.6353 | 249.5503 | 195.2399 | 247.3297 | 175.641 | -1.6014 | 1.2666 | 32.0285 | 40.5679 | 25.2866 | 2.0011 | 31.9519 | 63.938 | 15.9674 | 14.5487 | -0.1104 | 104.5059 | 235.0016 | 180.6912 | 232.781 | 161.0923
2012-10-30 19:24 | 28.645 | 249.5482 | 195.2613 | 247.5 | 175.6429 | -1.5991 | 1.2689 | 32.0875 | 40.7163 | 25.2873 | 2.0048 | 32.0112 | 64.1759 | 15.9673 | 14.5298 | -0.1102 | 104.8922 | 235.0184 | 180.7315 | 232.9702 | 161.1131
2012-10-30 19:24 | 28.6368 | 249.5752 | 195.3042 | 247.3782 | 175.8845 | -1.6032 | 1.2695 | 32.1038 | 40.757 | 25.2878 | 2.0058 | 32.0272 | 64.2404 | 15.9673 | 14.5135 | -0.1105 | 104.9974 | 235.0617 | 180.7907 | 232.8647 | 161.371
2012-10-30 19:24 | 28.6308 | 249.5653 | 195.3375 | 247.1959 | 175.7731 | -1.5997 | 1.2701 | 32.1178 | 40.7921 | 25.2881 | 2.0067 | 32.0415 | 64.2979 | 15.9671 | 14.5442 | -0.1103 | 105.09 | 235.0211 | 180.7933 | 232.6517 | 161.2289
2012-10-30 19:24 | 28.63 | 249.5763 | 195.3524 | 247.4654 | 175.8621 | -1.602 | 1.2704 | 32.1268 | 40.8149 | 25.2881 | 2.0073 | 32.0506 | 64.3347 | 15.9671 | 14.5597 | -0.1104 | 105.1496 | 235.0166 | 180.7927 | 232.9057 | 161.3024
2012-10-30 19:24 | 28.6263 | 249.5921 | 195.3527 | 247.2972 | 175.9716 | -1.6009 | 1.2707 | 32.1327 | 40.8301 | 25.288 | 2.0076 | 32.0563 | 64.3572 | 15.9672 | 14.5438 | -0.1104 | 105.1873 | 235.0483 | 180.8089 | 232.7534 | 161.4278
2012-10-30 19:24 | 28.6203 | 249.6018 | 195.386 | 247.2873 | 175.8094 | -1.6043 | 1.2708 | 32.1365 | 40.8397 | 25.2879 | 2.0079 | 32.0602 | 64.3728 | 15.9672 | 14.5337 | -0.1106 | 105.2125 | 235.0681 | 180.8523 | 232.7536 | 161.2757
2012-10-30 19:24 | 28.6195 | 249.5971 | 195.4323 | 247.3844 | 175.8828 | -1.6014 | 1.2709 | 32.14 | 40.8482 | 25.2883 | 2.0081 | 32.0636 | 64.3869 | 15.9671 | 14.5166 | -0.1104 | 105.2351 | 235.0805 | 180.9157 | 232.8678 | 161.3662
2012-10-30 19:24 | 28.6293 | 249.5756 | 195.4067 | 247.4412 | 175.7794 | -1.598 | 1.271 | 32.142 | 40.8532 | 25.2883 | 2.0082 | 32.0656 | 64.395 | 15.9671 | 14.5264 | -0.1102 | 105.2482 | 235.0492 | 180.8803 | 232.9148 | 161.253
2012-10-30 19:24 | 28.6255 | 249.6501 | 195.4462 | 247.2533 | 175.8615 | -1.602 | 1.2711 | 32.143 | 40.856 | 25.2882 | 2.0083 | 32.0666 | 64.3986 | 15.9672 | 14.543 | -0.1104 | 105.2547 | 235.1071 | 180.9032 | 232.7103 | 161.3185
2012-10-30 19:24 | 28.624 | 249.6447 | 195.3857 | 247.3223 | 175.7507 | -1.6026 | 1.2711 | 32.144 | 40.8586 | 25.2881 | 2.0084 | 32.0676 | 64.4036 | 15.967 | 14.5211 | -0.1105 | 105.2622 | 235.1236 | 180.8646 | 232.8012 | 161.2296
2012-10-30 19:24 | 28.6248 | 249.618 | 195.4219 | 247.4289 | 175.8491 | -1.6014 | 1.2711 | 32.1448 | 40.8603 | 25.2884 | 2.0084 | 32.0685 | 64.4069 | 15.9671 | 14.5178 | -0.1104 | 105.2672 | 235.1002 | 180.9041 | 232.9111 | 161.3313
2012-10-30 19:24 | 28.6255 | 249.6657 | 195.4501 | 247.3082 | 175.7171 | -1.5985 | 1.2712 | 32.1458 | 40.8626 | 25.2884 | 2.0085 | 32.0693 | 64.4104 | 15.967 | 14.4981 | -0.1102 | 105.2729 | 235.1676 | 180.952 | 232.8101 | 161.219
2012-10-30 19:25 | 28.6323 | 249.6843 | 195.4412 | 247.264 | 175.8097 | -1.5974 | 1.2712 | 32.1464 | 40.8639 | 25.2886 | 2.0085 | 32.07 | 64.4133 | 15.9669 | 14.4967 | -0.1101 | 105.2772 | 235.1876 | 180.9445 | 232.7673 | 161.313
2012-10-30 19:25 | 28.6285 | 249.6139 | 195.4295 | 247.4091 | 175.7474 | -1.5991 | 1.2712 | 32.1467 | 40.8651 | 25.2883 | 2.0085 | 32.0703 | 64.4143 | 15.967 | 14.5297 | -0.1102 | 105.2795 | 235.0842 | 180.8998 | 232.8794 | 161.2177
2012-10-30 19:25 | 28.6293 | 249.646 | 195.4382 | 247.3864 | 176.0058 | -1.6014 | 1.2712 | 32.1467 | 40.865 | 25.2884 | 2.0085 | 32.0702 | 64.414 | 15.967 | 14.5305 | -0.1104 | 105.279 | 235.1155 | 180.9077 | 232.8559 | 161.4753
2012-10-30 19:25 | 28.633 | 249.6732 | 195.4655 | 247.3353 | 175.8963 | -1.6003 | 1.2712 | 32.147 | 40.8657 | 25.2884 | 2.0086 | 32.0705 | 64.4153 | 15.9669 | 14.5464 | -0.1103 | 105.281 | 235.1268 | 180.9191 | 232.7889 | 161.3499
2012-10-30 19:25 | 28.627 | 249.6711 | 195.4555 | 247.3214 | 175.8435 | -1.5997 | 1.2712 | 32.1466 | 40.8651 | 25.2882 | 2.0085 | 32.0701 | 64.4142 | 15.9669 | 14.5037 | -0.1103 | 105.2793 | 235.1674 | 180.9518 | 232.8177 | 161.3398
2012-10-30 19:25 | 28.63 | 249.6389 | 195.4664 | 247.4811 | 175.7801 | -1.6026 | 1.2712 | 32.1465 | 40.8645 | 25.2884 | 2.0085 | 32.07 | 64.4139 | 15.9669 | 14.5189 | -0.1105 | 105.2784 | 235.12 | 180.9475 | 232.9622 | 161.2612
2012-10-30 19:25 | 28.636 | 249.6958 | 195.5117 | 247.3422 | 175.7939 | -1.5997 | 1.2712 | 32.1465 | 40.8646 | 25.2884 | 2.0085 | 32.0701 | 64.4139 | 15.9669 | 14.5576 | -0.1103 | 105.2785 | 235.1382 | 180.9541 | 232.7846 | 161.2363
2012-10-30 19:25 | 28.6398 | 249.6957 | 195.4644 | 247.3538 | 175.7704 | -1.5962 | 1.2712 | 32.1468 | 40.8657 | 25.2882 | 2.0086 | 32.0704 | 64.4156 | 15.9668 | 14.5369 | -0.11 | 105.2813 | 235.1588 | 180.9275 | 232.8169 | 161.2335
2012-10-30 19:25 | 28.63 | 249.6507 | 195.4428 | 247.489 | 175.6708 | -1.6055 | 1.2712 | 32.1471 | 40.8665 | 25.2881 | 2.0086 | 32.0707 | 64.417 | 15.9668 | 14.5108 | -0.1107 | 105.2835 | 235.1399 | 180.932 | 232.9782 | 161.16
2012-10-30 19:25 | 28.639 | 249.6831 | 195.4518 | 247.3804 | 175.8086 | -1.5997 | 1.2712 | 32.1471 | 40.866 | 25.2884 | 2.0086 | 32.0708 | 64.4168 | 15.9669 | 14.5238 | -0.1103 | 105.2828 | 235.1593 | 180.928 | 232.8566 | 161.2848
2012-10-30 19:25 | 28.6368 | 249.7005 | 195.4142 | 247.3547 | 175.6893 | -1.5945 | 1.2712 | 32.1468 | 40.8656 | 25.2882 | 2.0086 | 32.0704 | 64.4156 | 15.9668 | 14.4727 | -0.1099 | 105.2812 | 235.2278 | 180.9415 | 232.882 | 161.2166
2012-10-30 19:25 | 28.6353 | 249.6912 | 195.456 | 247.4394 | 175.559 | -1.6014 | 1.2712 | 32.1465 | 40.8646 | 25.2884 | 2.0085 | 32.0701 | 64.4143 | 15.9668 | 14.5528 | -0.1104 | 105.2789 | 235.1384 | 180.9032 | 232.8866 | 161.0062
2012-10-30 19:25 | 28.642 | 249.7057 | 195.4391 | 247.4187 | 175.7258 | -1.5991 | 1.2712 | 32.1464 | 40.8641 | 25.2885 | 2.0085 | 32.0699 | 64.4136 | 15.9668 | 14.5105 | -0.1102 | 105.2777 | 235.1952 | 180.9286 | 232.9082 | 161.2153
2012-10-30 19:25 | 28.6413 | 249.7011 | 195.458 | 247.3475 | 175.6392 | -1.5997 | 1.2712 | 32.1462 | 40.8636 | 25.2884 | 2.0085 | 32.0698 | 64.4133 | 15.9667 | 14.5343 | -0.1103 | 105.2769 | 235.1668 | 180.9237 | 232.8132 | 161.1049
2012-10-30 19:25 | 28.6525 | 249.681 | 195.4103 | 247.4997 | 175.6699 | -1.6061 | 1.2712 | 32.1457 | 40.8627 | 25.2882 | 2.0085 | 32.0693 | 64.4114 | 15.9668 | 14.521 | -0.1107 | 105.274 | 235.16 | 180.8893 | 232.9787 | 161.1489
2012-10-30 19:25 | 28.642 | 249.7057 | 195.4548 | 247.4265 | 175.8272 | -1.5991 | 1.2711 | 32.1454 | 40.8616 | 25.2885 | 2.0085 | 32.069 | 64.4105 | 15.9667 | 14.5105 | -0.1102 | 105.2721 | 235.1952 | 180.9443 | 232.916 | 161.3167
2012-10-30 19:25 | 28.6442 | 249.6805 | 195.3666 | 247.3543 | 175.6694 | -1.6026 | 1.2712 | 32.1454 | 40.862 | 25.2882 | 2.0085 | 32.069 | 64.4102 | 15.9667 | 14.5331 | -0.1105 | 105.2721 | 235.1474 | 180.8335 | 232.8212 | 161.1363
2012-10-30 19:25 | 28.6517 | 249.688 | 195.3898 | 247.5224 | 175.6691 | -1.6043 | 1.2711 | 32.1454 | 40.8614 | 25.2886 | 2.0085 | 32.0689 | 64.4103 | 15.9666 | 14.4917 | -0.1106 | 105.2717 | 235.1963 | 180.8981 | 233.0307 | 161.1774
2012-10-30 19:25 | 28.6487 | 249.6811 | 195.4301 | 247.394 | 175.8691 | -1.6026 | 1.2711 | 32.1452 | 40.8609 | 25.2886 | 2.0085 | 32.0687 | 64.4095 | 15.9666 | 14.5335 | -0.1105 | 105.2704 | 235.1476 | 180.8966 | 232.8605 | 161.3356
2012-10-30 19:25 | 28.6502 | 249.6865 | 195.4237 | 247.4073 | 175.6832 | -1.5997 | 1.2711 | 32.1449 | 40.8605 | 25.2883 | 2.0085 | 32.0685 | 64.4087 | 15.9666 | 14.4902 | -0.1103 | 105.2692 | 235.1963 | 180.9335 | 232.9171 | 161.193
2012-10-30 19:25 | 28.6465 | 249.6867 | 195.4475 | 247.5015 | 175.7263 | -1.602 | 1.2711 | 32.1448 | 40.86 | 25.2885 | 2.0085 | 32.0684 | 64.4079 | 15.9667 | 14.4946 | -0.1104 | 105.268 | 235.1921 | 180.9529 | 233.0069 | 161.2317
2012-10-30 19:25 | 28.6577 | 249.694 | 195.4194 | 247.4148 | 175.6868 | -1.6003 | 1.2711 | 32.1446 | 40.8595 | 25.2884 | 2.0085 | 32.0682 | 64.4076 | 15.9666 | 14.5221 | -0.1103 | 105.2672 | 235.1719 | 180.8973 | 232.8927 | 161.1647
2012-10-30 19:25 | 28.6615 | 249.6978 | 195.4114 | 247.5244 | 175.5423 | -1.5968 | 1.2711 | 32.1443 | 40.8592 | 25.2882 | 2.0084 | 32.0679 | 64.4063 | 15.9666 | 14.5178 | -0.1101 | 105.2655 | 235.18 | 180.8936 | 233.0066 | 161.0245
2012-10-30 19:25 | 28.654 | 249.7295 | 195.3882 | 247.4503 | 175.7495 | -1.602 | 1.2711 | 32.1439 | 40.8582 | 25.2883 | 2.0084 | 32.0675 | 64.4048 | 15.9666 | 14.4654 | -0.1104 | 105.2631 | 235.2641 | 180.9228 | 232.9849 | 161.2841
2012-10-30 19:25 | 28.6622 | 249.7533 | 195.3768 | 247.4193 | 175.5235 | -1.6003 | 1.2711 | 32.1437 | 40.8573 | 25.2884 | 2.0084 | 32.0673 | 64.404 | 15.9666 | 14.5389 | -0.1103 | 105.2612 | 235.2144 | 180.8379 | 232.8804 | 160.9846
2012-10-30 19:26 | 28.6607 | 249.6539 | 195.391 | 247.6098 | 175.6859 | -1.6009 | 1.2711 | 32.1435 | 40.8568 | 25.2884 | 2.0084 | 32.0671 | 64.4035 | 15.9665 | 14.5374 | -0.1104 | 105.2603 | 235.1165 | 180.8536 | 233.0724 | 161.1485
2012-10-30 19:26 | 28.6645 | 249.7987 | 195.3791 | 247.3824 | 175.7092 | -1.6049 | 1.271 | 32.1431 | 40.8554 | 25.2887 | 2.0084 | 32.0668 | 64.4022 | 15.9665 | 14.5412 | -0.1106 | 105.2577 | 235.2575 | 180.8379 | 232.8412 | 161.168
2012-10-30 19:26 | 28.6585 | 249.6948 | 195.3731 | 247.5253 | 175.6876 | -1.6037 | 1.2711 | 32.1427 | 40.8551 | 25.2883 | 2.0083 | 32.0663 | 64.4002 | 15.9666 | 14.5229 | -0.1106 | 105.2553 | 235.1719 | 180.8502 | 233.0024 | 161.1647
2012-10-30 19:26 | 28.6637 | 249.7196 | 195.4176 | 247.5266 | 175.9348 | -1.5957 | 1.271 | 32.1421 | 40.8538 | 25.2881 | 2.0083 | 32.0658 | 64.3983 | 15.9665 | 14.5159 | -0.11 | 105.2521 | 235.2037 | 180.9017 | 233.0107 | 161.4189
2012-10-30 19:26 | 28.6555 | 249.7466 | 195.3976 | 247.4048 | 175.751 | -1.6049 | 1.271 | 32.1421 | 40.8527 | 25.2887 | 2.0083 | 32.0656 | 64.3975 | 15.9665 | 14.5199 | -0.1106 | 105.2502 | 235.2267 | 180.8777 | 232.8849 | 161.2311
2012-10-30 19:26 | 28.6547 | 249.7223 | 195.4636 | 247.6312 | 175.5784 | -1.602 | 1.271 | 32.1419 | 40.8525 | 25.2885 | 2.0083 | 32.0654 | 64.397 | 15.9665 | 14.5273 | -0.1104 | 105.2496 | 235.195 | 180.9363 | 233.1039 | 161.0511
2012-10-30 19:26 | 28.6555 | 249.7466 | 195.4369 | 247.4048 | 175.5363 | -1.6032 | 1.271 | 32.1415 | 40.8517 | 25.2884 | 2.0083 | 32.0651 | 64.3958 | 15.9664 | 14.4995 | -0.1105 | 105.2475 | 235.2471 | 180.9374 | 232.9053 | 161.0368
2012-10-30 19:26 | 28.651 | 249.7265 | 195.3616 | 247.5726 | 175.4069 | -1.6014 | 1.271 | 32.1414 | 40.8518 | 25.2883 | 2.0083 | 32.0649 | 64.3951 | 15.9664 | 14.5195 | -0.1104 | 105.2469 | 235.207 | 180.8421 | 233.0531 | 160.8874
2012-10-30 19:26 | 28.6547 | 249.7615 | 195.3614 | 247.4823 | 175.4887 | -1.5991 | 1.271 | 32.141 | 40.8507 | 25.2883 | 2.0083 | 32.0647 | 64.394 | 15.9664 | 14.5885 | -0.1102 | 105.2446 | 235.173 | 180.7729 | 232.8938 | 160.9002
2012-10-30 19:26 | 28.6592 | 249.7895 | 195.3659 | 247.5221 | 175.4307 | -1.5985 | 1.271 | 32.1407 | 40.8501 | 25.2881 | 2.0082 | 32.0642 | 64.3918 | 15.9666 | 14.5359 | -0.1102 | 105.2419 | 235.2536 | 180.83 | 232.9862 | 160.8948
2012-10-30 19:26 | 28.657 | 249.7442 | 195.3284 | 247.6061 | 175.7056 | -1.602 | 1.271 | 32.1403 | 40.8493 | 25.2881 | 2.0082 | 32.064 | 64.3912 | 15.9665 | 14.5704 | -0.1104 | 105.2405 | 235.1738 | 180.758 | 233.0357 | 161.1352
2012-10-30 19:26 | 28.657 | 249.8186 | 195.403 | 247.4102 | 175.7681 | -1.5991 | 1.2709 | 32.14 | 40.8482 | 25.2882 | 2.0082 | 32.0636 | 64.3905 | 15.9662 | 14.5622 | -0.1102 | 105.2387 | 235.2564 | 180.8408 | 232.848 | 161.2059
2012-10-30 19:26 | 28.6525 | 249.7632 | 195.3631 | 247.5545 | 175.6387 | -1.5991 | 1.2709 | 32.1398 | 40.8479 | 25.2881 | 2.0082 | 32.0635 | 64.3896 | 15.9664 | 14.5373 | -0.1102 | 105.2375 | 235.2259 | 180.8258 | 233.0172 | 161.1014
2012-10-30 19:26 | 28.6525 | 249.724 | 195.3828 | 247.4723 | 175.7363 | -1.5997 | 1.2709 | 32.1392 | 40.8461 | 25.2883 | 2.0082 | 32.0628 | 64.3873 | 15.9663 | 14.5129 | -0.1103 | 105.2334 | 235.2111 | 180.8699 | 232.9594 | 161.2234
2012-10-30 19:26 | 28.6622 | 249.769 | 195.3689 | 247.4937 | 175.5664 | -1.6003 | 1.2709 | 32.1389 | 40.8455 | 25.2882 | 2.0081 | 32.0625 | 64.3856 | 15.9663 | 14.547 | -0.1103 | 105.2311 | 235.222 | 180.8219 | 232.9467 | 161.0194
2012-10-30 19:26 | 28.6622 | 249.7651 | 195.3414 | 247.623 | 175.5664 | -1.6014 | 1.2709 | 32.1386 | 40.8451 | 25.2879 | 2.0081 | 32.0624 | 64.3853 | 15.9663 | 14.5226 | -0.1104 | 105.2304 | 235.2425 | 180.8188 | 233.1004 | 161.0438
2012-10-30 19:26 | 28.6585 | 249.7653 | 195.3849 | 247.4313 | 175.6447 | -1.6014 | 1.2709 | 32.1381 | 40.8437 | 25.288 | 2.0081 | 32.0618 | 64.3831 | 15.9663 | 14.5189 | -0.1104 | 105.2268 | 235.2464 | 180.866 | 232.9124 | 161.1258
2012-10-30 19:26 | 28.6555 | 249.7231 | 195.3347 | 247.5458 | 175.7432 | -1.6032 | 1.2654 | 31.9977 | 40.4896 | 25.2868 | 1.9993 | 31.9215 | 63.8206 | 15.9664 | 14.5403 | -0.1105 | 104.3102 | 235.1828 | 180.7944 | 233.0055 | 161.2029
2012-10-30 19:26 | 28.6577 | 249.7018 | 195.3408 | 247.4187 | 175.9483 | -1.602 | 1.2635 | 31.9503 | 40.3703 | 25.2864 | 1.9963 | 31.8744 | 63.6318 | 15.9665 | 14.5466 | -0.1104 | 104.0021 | 235.1552 | 180.7942 | 232.8721 | 161.4017
2012-10-30 19:26 | 28.6577 | 249.6392 | 195.2976 | 247.5167 | 175.9444 | -1.5991 | 1.2636 | 31.9507 | 40.3719 | 25.2861 | 1.9963 | 31.8748 | 63.633 | 15.9666 | 14.5466 | -0.1102 | 104.0049 | 235.0926 | 180.751 | 232.9701 | 161.3978
2012-10-30 19:26 | 28.6555 | 249.6448 | 195.3779 | 247.538 | 176.071 | -1.6003 | 1.2636 | 31.9525 | 40.3757 | 25.2866 | 1.9965 | 31.8766 | 63.6401 | 15.9666 | 14.5485 | -0.1103 | 104.0158 | 235.0963 | 180.8294 | 232.9895 | 161.5225
2012-10-30 19:26 | 28.666 | 249.6788 | 195.3963 | 247.38 | 175.9879 | -1.6061 | 1.2637 | 31.9537 | 40.3789 | 25.2865 | 1.9965 | 31.8779 | 63.645 | 15.9667 | 14.5631 | -0.1107 | 104.0239 | 235.1157 | 180.8332 | 232.8169 | 161.4248
2012-10-30 19:27 | 28.6592 | 249.672 | 195.4249 | 247.5221 | 175.7351 | -1.5997 | 1.2638 | 31.9559 | 40.3845 | 25.2863 | 1.9967 | 31.88 | 63.6532 | 15.9667 | 14.5563 | -0.1103 | 104.0377 | 235.1157 | 180.8686 | 232.9658 | 161.1788
2012-10-30 19:27 | 28.666 | 249.6358 | 195.4356 | 247.4388 | 175.7654 | -1.5997 | 1.2639 | 31.9589 | 40.3925 | 25.2861 | 1.9968 | 31.883 | 63.6651 | 15.9667 | 14.5631 | -0.1103 | 104.0576 | 235.0727 | 180.8725 | 232.8757 | 161.2023
2012-10-30 19:27 | 28.6585 | 249.6596 | 195.3691 | 247.4665 | 175.7501 | -1.6014 | 1.2663 | 32.0212 | 40.5491 | 25.2868 | 2.0008 | 31.9466 | 63.92 | 15.9666 | 14.5393 | -0.1104 | 104.4691 | 235.1203 | 180.8298 | 232.9272 | 161.2108
2012-10-30 19:27 | 28.6495 | 249.584 | 195.3916 | 247.4967 | 175.866 | -1.5968 | 1.2689 | 32.0861 | 40.7127 | 25.2874 | 2.0048 | 32.0096 | 64.172 | 15.9667 | 14.4976 | -0.1101 | 104.8847 | 235.0864 | 180.894 | 232.9991 | 161.3684
2012-10-30 19:27 | 28.6413 | 249.6306 | 195.3991 | 247.277 | 175.9788 | -1.5985 | 1.2694 | 32.0991 | 40.7452 | 25.2877 | 2.0056 | 32.0224 | 64.2237 | 15.9666 | 14.5098 | -0.1102 | 104.9689 | 235.1208 | 180.8893 | 232.7672 | 161.469
2012-10-30 19:27 | 28.6465 | 249.6241 | 195.475 | 247.5329 | 175.863 | -1.6009 | 1.2699 | 32.113 | 40.7804 | 25.2877 | 2.0065 | 32.0363 | 64.2792 | 15.9666 | 14.4905 | -0.1104 | 105.0595 | 235.1336 | 180.9845 | 233.0424 | 161.3725
2012-10-30 19:27 | 28.6398 | 249.637 | 195.4408 | 247.3695 | 175.9734 | -1.5997 | 1.2703 | 32.1234 | 40.8066 | 25.2878 | 2.0071 | 32.0468 | 64.3222 | 15.9665 | 14.5165 | -0.1103 | 105.1288 | 235.1205 | 180.9243 | 232.853 | 161.4569
2012-10-30 19:27 | 28.6405 | 249.6181 | 195.414 | 247.4015 | 175.896 | -1.6026 | 1.2705 | 32.1294 | 40.8214 | 25.2882 | 2.0075 | 32.053 | 64.3464 | 15.9666 | 14.5417 | -0.1105 | 105.1678 | 235.0764 | 180.8723 | 232.8598 | 161.3543
2012-10-30 19:27 | 28.6398 | 249.6565 | 195.5233 | 247.4282 | 176.1686 | -1.5997 | 1.2707 | 32.1337 | 40.8327 | 25.2879 | 2.0078 | 32.0573 | 64.364 | 15.9666 | 14.4798 | -0.1103 | 105.1967 | 235.1767 | 181.0435 | 232.9484 | 161.6888
2012-10-30 19:27 | 28.6368 | 249.7083 | 195.54 | 247.3469 | 176.0094 | -1.6003 | 1.2708 | 32.1373 | 40.8416 | 25.2881 | 2.008 | 32.0607 | 64.3777 | 15.9665 | 14.5012 | -0.1103 | 105.2193 | 235.2071 | 181.0388 | 232.8457 | 161.5082
2012-10-30 19:27 | 28.6308 | 249.6201 | 195.5222 | 247.5054 | 175.898 | -1.5957 | 1.2709 | 32.1399 | 40.848 | 25.2882 | 2.0082 | 32.0634 | 64.3889 | 15.9665 | 14.532 | -0.11 | 105.2369 | 235.0881 | 180.9902 | 232.9734 | 161.366
2012-10-30 19:27 | 28.6285 | 249.7314 | 195.5513 | 247.413 | 175.9035 | -1.5991 | 1.271 | 32.1415 | 40.8518 | 25.2884 | 2.0083 | 32.0651 | 64.3961 | 15.9664 | 14.5337 | -0.1102 | 105.2478 | 235.1977 | 181.0176 | 232.8793 | 161.3698
2012-10-30 19:27 | 28.627 | 249.6711 | 195.483 | 247.3606 | 175.8122 | -1.5945 | 1.271 | 32.1423 | 40.8542 | 25.2882 | 2.0083 | 32.0659 | 64.3993 | 15.9664 | 14.4874 | -0.1099 | 105.2535 | 235.1837 | 180.9956 | 232.8732 | 161.3248
2012-10-30 19:27 | 28.6293 | 249.6852 | 195.4971 | 247.4883 | 175.8653 | -1.6055 | 1.2711 | 32.143 | 40.8556 | 25.2884 | 2.0084 | 32.0665 | 64.4013 | 15.9664 | 14.5101 | -0.1107 | 105.2569 | 235.1751 | 180.987 | 232.9782 | 161.3552
2012-10-30 19:27 | 28.6285 | 249.7118 | 195.5474 | 247.3974 | 175.8801 | -1.598 | 1.2711 | 32.1438 | 40.8575 | 25.2885 | 2.0084 | 32.0674 | 64.405 | 15.9664 | 14.5093 | -0.1102 | 105.2625 | 235.2025 | 181.0381 | 232.8881 | 161.3708
2012-10-30 19:27 | 28.6195 | 249.695 | 195.5188 | 247.3178 | 175.8321 | -1.5997 | 1.2711 | 32.1445 | 40.8595 | 25.2884 | 2.0085 | 32.068 | 64.4075 | 15.9664 | 14.5288 | -0.1103 | 105.267 | 235.1662 | 180.99 | 232.789 | 161.3033
2012-10-30 19:27 | 28.633 | 249.7202 | 195.5126 | 247.535 | 175.748 | -1.602 | 1.2711 | 32.1454 | 40.8616 | 25.2885 | 2.0085 | 32.069 | 64.4114 | 15.9664 | 14.5342 | -0.1104 | 105.273 | 235.186 | 180.9784 | 233.0008 | 161.2138
2012-10-30 19:27 | 28.633 | 249.7633 | 195.5519 | 247.3117 | 175.9431 | -1.5997 | 1.2712 | 32.1455 | 40.862 | 25.2884 | 2.0085 | 32.069 | 64.4115 | 15.9664 | 14.5627 | -0.1103 | 105.2735 | 235.2006 | 180.9892 | 232.749 | 161.3804
2012-10-30 19:27 | 28.6345 | 249.7491 | 195.4473 | 247.513 | 175.7183 | -1.6037 | 1.2712 | 32.1457 | 40.8628 | 25.2882 | 2.0086 | 32.0693 | 64.4128 | 15.9664 | 14.5193 | -0.1106 | 105.2756 | 235.2298 | 180.928 | 232.9937 | 161.199
2012-10-30 19:27 | 28.6225 | 249.7449 | 195.4864 | 247.3757 | 175.9639 | -1.6003 | 1.2712 | 32.1459 | 40.8626 | 25.2886 | 2.0086 | 32.0694 | 64.4133 | 15.9663 | 14.4706 | -0.1103 | 105.2759 | 235.2743 | 181.0158 | 232.9051 | 161.4933
2012-10-30 19:27 | 28.627 | 249.6946 | 195.4909 | 247.4037 | 175.9645 | -1.5991 | 1.2712 | 32.1461 | 40.8632 | 25.2886 | 2.0086 | 32.0697 | 64.4144 | 15.9664 | 14.5486 | -0.1102 | 105.2776 | 235.146 | 180.9423 | 232.8551 | 161.4159
2012-10-30 19:27 | 28.6278 | 249.6563 | 195.5153 | 247.5259 | 175.9614 | -1.5968 | 1.2712 | 32.1463 | 40.8643 | 25.2883 | 2.0086 | 32.0699 | 64.4152 | 15.9664 | 14.5045 | -0.1101 | 105.2795 | 235.1518 | 181.0108 | 233.0214 | 161.4569
2012-10-30 19:27 | 28.624 | 249.7543 | 195.535 | 247.2871 | 176.0747 | -1.6014 | 1.2712 | 32.1465 | 40.8649 | 25.2882 | 2.0086 | 32.0701 | 64.4169 | 15.9661 | 14.5252 | -0.1104 | 105.2818 | 235.2291 | 181.0098 | 232.7619 | 161.5495
2012-10-30 19:27 | 28.6278 | 249.7502 | 195.5074 | 247.4828 | 175.7116 | -1.6066 | 1.2712 | 32.1466 | 40.8652 | 25.2882 | 2.0086 | 32.0702 | 64.4169 | 15.9662 | 14.5208 | -0.1108 | 105.2821 | 235.2294 | 180.9866 | 232.962 | 161.1908
2012-10-30 19:27 | 28.6315 | 249.6874 | 195.5897 | 247.4474 | 175.926 | -1.5997 | 1.2712 | 32.1463 | 40.8645 | 25.2881 | 2.0086 | 32.0699 | 64.416 | 15.9662 | 14.5408 | -0.1103 | 105.2805 | 235.1466 | 181.0489 | 232.9066 | 161.3852
2012-10-30 19:27 | 28.6308 | 249.7415 | 195.5183 | 247.4192 | 175.8395 | -1.6026 | 1.2712 | 32.1462 | 40.8641 | 25.2881 | 2.0086 | 32.0698 | 64.4155 | 15.9662 | 14.5564 | -0.1105 | 105.2796 | 235.1851 | 180.9619 | 232.8628 | 161.2831
2012-10-30 19:27 | 28.636 | 249.6762 | 195.5431 | 247.6086 | 175.7549 | -1.602 | 1.2712 | 32.1464 | 40.8645 | 25.2881 | 2.0086 | 32.0699 | 64.4161 | 15.9662 | 14.5249 | -0.1104 | 105.2807 | 235.1513 | 181.0182 | 233.0837 | 161.23
2012-10-30 19:28 | 28.6285 | 249.7235 | 195.5435 | 247.4052 | 175.8606 | -1.6009 | 1.2712 | 32.1463 | 40.8635 | 25.2886 | 2.0086 | 32.0698 | 64.4157 | 15.9662 | 14.5541 | -0.1104 | 105.2792 | 235.1694 | 180.9894 | 232.8511 | 161.3065
2012-10-30 19:28 | 28.6315 | 249.7422 | 195.519 | 247.4513 | 175.8011 | -1.6009 | 1.2712 | 32.1458 | 40.8627 | 25.2883 | 2.0086 | 32.0693 | 64.4144 | 15.966 | 14.5408 | -0.1104 | 105.2772 | 235.2014 | 180.9782 | 232.9105 | 161.2603
2012-10-30 19:28 | 28.633 | 249.7006 | 195.5401 | 247.5429 | 175.9197 | -1.6037 | 1.2712 | 32.1459 | 40.8624 | 25.2887 | 2.0086 | 32.0695 | 64.4148 | 15.9661 | 14.5423 | -0.1106 | 105.2772 | 235.1583 | 180.9978 | 233.0006 | 161.3774
2012-10-30 19:28 | 28.6375 | 249.7599 | 195.521 | 247.4064 | 175.7408 | -1.5985 | 1.2712 | 32.1456 | 40.8621 | 25.2884 | 2.0086 | 32.0692 | 64.4134 | 15.9662 | 14.5591 | -0.1102 | 105.2756 | 235.2008 | 180.9619 | 232.8473 | 161.1817
2012-10-30 19:28 | 28.6278 | 249.7307 | 195.4642 | 247.5063 | 175.6569 | -1.6037 | 1.2712 | 32.1458 | 40.8635 | 25.2879 | 2.0086 | 32.0692 | 64.4135 | 15.9662 | 14.5167 | -0.1106 | 105.2769 | 235.214 | 180.9475 | 232.9896 | 161.1402
2012-10-30 19:28 | 28.63 | 249.7524 | 195.5568 | 247.4654 | 175.8894 | -1.6061 | 1.2712 | 32.1453 | 40.8618 | 25.2881 | 2.0086 | 32.0688 | 64.4126 | 15.966 | 14.5434 | -0.1107 | 105.2744 | 235.209 | 181.0134 | 232.922 | 161.346
2012-10-30 19:28 | 28.6375 | 249.756 | 195.4935 | 247.4299 | 175.7173 | -1.6055 | 1.2711 | 32.145 | 40.8606 | 25.2884 | 2.0086 | 32.0686 | 64.4116 | 15.966 | 14.4938 | -0.1107 | 105.2723 | 235.2622 | 180.9997 | 232.9361 | 161.2235
2012-10-30 19:28 | 28.6405 | 249.7434 | 195.4926 | 247.57 | 175.8687 | -1.5957 | 1.2711 | 32.1448 | 40.8604 | 25.2883 | 2.0085 | 32.0685 | 64.4108 | 15.9661 | 14.5539 | -0.11 | 105.2711 | 235.1895 | 180.9387 | 233.0161 | 161.3148
2012-10-30 19:28 | 28.627 | 249.7847 | 195.538 | 247.3684 | 175.945 | -1.5997 | 1.2711 | 32.1446 | 40.8598 | 25.2882 | 2.0085 | 32.0681 | 64.4091 | 15.9661 | 14.5118 | -0.1103 | 105.2689 | 235.2729 | 181.0262 | 232.8566 | 161.4332
2012-10-30 19:28 | 28.63 | 249.7524 | 195.5096 | 247.5046 | 175.7255 | -1.6009 | 1.2711 | 32.1444 | 40.8598 | 25.288 | 2.0085 | 32.068 | 64.4093 | 15.9659 | 14.5393 | -0.1104 | 105.2691 | 235.2131 | 180.9703 | 232.9653 | 161.1862
2012-10-30 19:28 | 28.6323 | 249.743 | 195.5158 | 247.5343 | 175.798 | -1.6066 | 1.2711 | 32.1442 | 40.8585 | 25.2884 | 2.0085 | 32.0677 | 64.4081 | 15.9659 | 14.5579 | -0.1108 | 105.2666 | 235.1851 | 180.9579 | 232.9764 | 161.2401
2012-10-30 19:28 | 28.63 | 249.7603 | 195.49 | 247.4654 | 175.8035 | -1.5985 | 1.2711 | 32.1438 | 40.8579 | 25.2883 | 2.0085 | 32.0675 | 64.407 | 15.966 | 14.5108 | -0.1102 | 105.2649 | 235.2495 | 180.9792 | 232.9546 | 161.2927
2012-10-30 19:28 | 28.6368 | 249.7592 | 195.5203 | 247.6681 | 175.7206 | -1.6061 | 1.2711 | 32.1437 | 40.8578 | 25.2881 | 2.0085 | 32.0673 | 64.4064 | 15.966 | 14.5584 | -0.1107 | 105.2642 | 235.2008 | 180.9619 | 233.1097 | 161.1622
2012-10-30 19:28 | 28.6345 | 249.7491 | 195.5102 | 247.3603 | 175.8471 | -1.6089 | 1.2711 | 32.1432 | 40.8563 | 25.2882 | 2.0085 | 32.0668 | 64.4046 | 15.9659 | 14.5642 | -0.1109 | 105.261 | 235.1849 | 180.946 | 232.7961 | 161.2829
2012-10-30 19:28 | 28.6345 | 249.7491 | 195.4827 | 247.5405 | 175.7026 | -1.5997 | 1.2711 | 32.143 | 40.8562 | 25.2881 | 2.0084 | 32.0665 | 64.4036 | 15.9659 | 14.5234 | -0.1103 | 105.2597 | 235.2257 | 180.9593 | 233.0171 | 161.1792
2012-10-30 19:28 | 28.6428 | 249.777 | 195.5263 | 247.5762 | 175.7226 | -1.6032 | 1.2711 | 32.1428 | 40.8552 | 25.2883 | 2.0084 | 32.0664 | 64.4032 | 15.9659 | 14.5521 | -0.1105 | 105.2585 | 235.2249 | 180.9742 | 233.0241 | 161.1705
2012-10-30 19:28 | 28.6383 | 249.7803 | 195.4432 | 247.4111 | 175.6557 | -1.6084 | 1.271 | 32.1426 | 40.8546 | 25.2883 | 2.0084 | 32.0662 | 64.4025 | 15.9659 | 14.4783 | -0.1109 | 105.2571 | 235.302 | 180.9649 | 232.9328 | 161.1774
2012-10-30 19:28 | 28.6345 | 249.7491 | 195.463 | 247.6462 | 175.6285 | -1.5997 | 1.2711 | 32.1427 | 40.8555 | 25.288 | 2.0084 | 32.0664 | 64.4034 | 15.9658 | 14.5234 | -0.1103 | 105.259 | 235.2257 | 180.9396 | 233.1228 | 161.1051
2012-10-30 19:28 | 28.645 | 249.7792 | 195.4971 | 247.504 | 175.7483 | -1.6014 | 1.2711 | 32.1424 | 40.8549 | 25.2879 | 2.0084 | 32.066 | 64.4017 | 15.9659 | 14.5421 | -0.1104 | 105.2566 | 235.2371 | 180.955 | 232.9619 | 161.2062
2012-10-30 19:28 | 28.639 | 249.8162 | 195.4243 | 247.4 | 175.7852 | -1.5962 | 1.271 | 32.1419 | 40.8532 | 25.2882 | 2.0084 | 32.0655 | 64.3998 | 15.9658 | 14.5279 | -0.11 | 105.253 | 235.2883 | 180.8964 | 232.8721 | 161.2573
2012-10-30 19:28 | 28.6405 | 249.7903 | 195.469 | 247.6405 | 175.7555 | -1.5957 | 1.271 | 32.1414 | 40.8521 | 25.288 | 2.0084 | 32.065 | 64.3979 | 15.9658 | 14.5498 | -0.11 | 105.2501 | 235.2405 | 180.9192 | 233.0907 | 161.2057
2012-10-30 19:28 | 28.6487 | 249.7907 | 195.4929 | 247.4411 | 175.8027 | -1.6014 | 1.271 | 32.1412 | 40.8519 | 25.2878 | 2.0083 | 32.0648 | 64.397 | 15.9658 | 14.558 | -0.1104 | 105.2489 | 235.2327 | 180.9349 | 232.8831 | 161.2447
2012-10-30 19:28 | 28.6435 | 249.7503 | 195.417 | 247.6591 | 175.4853 | -1.6026 | 1.271 | 32.1409 | 40.8512 | 25.2879 | 2.0083 | 32.0644 | 64.3961 | 15.9657 | 14.5283 | -0.1105 | 105.2473 | 235.222 | 180.8887 | 233.1308 | 160.957
2012-10-30 19:28 | 28.645 | 249.7988 | 195.4657 | 247.5549 | 175.6741 | -1.6037 | 1.271 | 32.1404 | 40.8498 | 25.2879 | 2.0083 | 32.064 | 64.3935 | 15.9658 | 14.5543 | -0.1106 | 105.2433 | 235.2445 | 180.9114 | 233.0006 | 161.1198
2012-10-30 19:28 | 28.6442 | 249.7431 | 195.4059 | 247.4796 | 175.6109 | -1.602 | 1.271 | 32.1403 | 40.8493 | 25.288 | 2.0083 | 32.0638 | 64.3935 | 15.9657 | 14.5413 | -0.1104 | 105.2428 | 235.2018 | 180.8646 | 232.9383 | 161.0696
2012-10-30 19:28 | 28.6428 | 249.7574 | 195.4045 | 247.6154 | 175.4846 | -1.6014 | 1.2709 | 32.1398 | 40.8479 | 25.2881 | 2.0083 | 32.0634 | 64.3918 | 15.9658 | 14.5276 | -0.1104 | 105.2397 | 235.2298 | 180.8769 | 233.0878 | 160.957
2012-10-30 19:28 | 28.6428 | 249.7731 | 195.432 | 247.4469 | 175.6524 | -1.5985 | 1.2709 | 32.1395 | 40.8472 | 25.288 | 2.0082 | 32.063 | 64.3904 | 15.9657 | 14.5113 | -0.1102 | 105.2375 | 235.2618 | 180.9207 | 232.9356 | 161.1411
2012-10-30 19:29 | 28.6495 | 249.7837 | 195.3876 | 247.5829 | 175.6201 | -1.5928 | 1.2709 | 32.1394 | 40.8467 | 25.2881 | 2.0082 | 32.0629 | 64.3904 | 15.9656 | 14.5466 | -0.1098 | 105.2371 | 235.2371 | 180.841 | 233.0363 | 161.0735
2012-10-30 19:29 | 28.642 | 249.7918 | 195.3919 | 247.6067 | 175.6321 | -1.598 | 1.2709 | 32.1391 | 40.8461 | 25.2881 | 2.0082 | 32.0628 | 64.3894 | 15.9657 | 14.5391 | -0.1102 | 105.2355 | 235.2527 | 180.8528 | 233.0676 | 161.093
2012-10-30 19:29 | 28.6487 | 249.7868 | 195.3947 | 247.4998 | 175.7012 | -1.6014 | 1.2709 | 32.1391 | 40.8466 | 25.2879 | 2.0082 | 32.0626 | 64.3898 | 15.9655 | 14.505 | -0.1104 | 105.2364 | 235.2818 | 180.8897 | 232.9948 | 161.1962
2012-10-30 19:29 | 28.6547 | 249.7576 | 195.3418 | 247.6468 | 175.5472 | -1.5997 | 1.2709 | 32.1388 | 40.8458 | 25.2879 | 2.0082 | 32.0625 | 64.3885 | 15.9656 | 14.5191 | -0.1103 | 105.2343 | 235.2385 | 180.8227 | 233.1277 | 161.0281
2012-10-30 19:29 | 28.6517 | 249.7898 | 195.3781 | 247.5655 | 175.6574 | -1.5957 | 1.2709 | 32.1385 | 40.8448 | 25.288 | 2.0082 | 32.0621 | 64.387 | 15.9656 | 14.5365 | -0.11 | 105.2318 | 235.2533 | 180.8416 | 233.029 | 161.1209
2012-10-30 19:29 | 28.6562 | 249.8021 | 195.3668 | 247.4995 | 175.6853 | -1.5985 | 1.2709 | 32.1383 | 40.8443 | 25.2879 | 2.0082 | 32.0619 | 64.3866 | 15.9655 | 14.5084 | -0.1102 | 105.2309 | 235.2937 | 180.8584 | 232.9911 | 161.1769
2012-10-30 19:29 | 28.6525 | 249.7828 | 195.371 | 247.578 | 175.5528 | -1.6032 | 1.2698 | 32.1096 | 40.7716 | 25.2879 | 2.0066 | 32.0368 | 64.2859 | 15.9655 | 14.4884 | -0.1105 | 105.0576 | 235.2944 | 180.8826 | 233.0896 | 161.0644
2012-10-30 19:29 | 28.654 | 249.7686 | 195.4197 | 247.5913 | 175.5972 | -1.6014 | 1.2647 | 31.9793 | 40.4437 | 25.2865 | 1.9982 | 31.9024 | 63.7471 | 15.9656 | 14.5307 | -0.1104 | 104.1908 | 235.2379 | 180.889 | 233.0606 | 161.0665
2012-10-30 19:29 | 28.648 | 249.7509 | 195.3311 | 247.4521 | 175.6186 | -1.6003 | 1.2646 | 31.9759 | 40.4354 | 25.2863 | 1.998 | 31.9004 | 63.7384 | 15.9658 | 14.5369 | -0.1103 | 104.1738 | 235.214 | 180.7942 | 232.9152 | 161.0817
2012-10-30 19:29 | 28.6495 | 249.7132 | 195.3444 | 247.7043 | 175.7762 | -1.6026 | 1.264 | 31.9624 | 40.4011 | 25.2863 | 1.9972 | 31.8867 | 63.6834 | 15.9659 | 14.5262 | -0.1105 | 104.0845 | 235.187 | 180.8182 | 233.1781 | 161.25
2012-10-30 19:29 | 28.66 | 249.7629 | 195.3785 | 247.4249 | 175.7281 | -1.5951 | 1.2637 | 31.9538 | 40.3796 | 25.2862 | 1.9966 | 31.8781 | 63.6492 | 15.9659 | 14.5448 | -0.11 | 104.0289 | 235.2181 | 180.8337 | 232.8801 | 161.1833
2012-10-30 19:29 | 28.6592 | 249.6877 | 195.3659 | 247.6122 | 175.5907 | -1.6003 | 1.2636 | 31.9507 | 40.3719 | 25.286 | 1.9964 | 31.8748 | 63.6353 | 15.966 | 14.5114 | -0.1103 | 104.0072 | 235.1763 | 180.8545 | 233.1008 | 161.0793
2012-10-30 19:29 | 28.6585 | 249.6596 | 195.3416 | 247.5919 | 175.2739 | -1.6026 | 1.2635 | 31.9499 | 40.3697 | 25.2862 | 1.9964 | 31.8739 | 63.6317 | 15.966 | 14.5311 | -0.1105 | 104.0014 | 235.1285 | 180.8105 | 233.0608 | 160.7428
2012-10-30 19:29 | 28.654 | 249.6981 | 195.2978 | 247.4699 | 175.207 | -1.6066 | 1.2635 | 31.9495 | 40.3687 | 25.2862 | 1.9963 | 31.8736 | 63.6303 | 15.966 | 14.4858 | -0.1108 | 103.999 | 235.2123 | 180.812 | 232.9841 | 160.7212
2012-10-30 19:29 | 28.657 | 249.6424 | 195.2223 | 247.6061 | 175.3388 | -1.5997 | 1.2636 | 31.9504 | 40.3709 | 25.2861 | 1.9964 | 31.8745 | 63.6337 | 15.9661 | 14.4929 | -0.1103 | 104.0046 | 235.1495 | 180.7294 | 233.1132 | 160.8459
2012-10-30 19:29 | 28.6532 | 249.6347 | 195.246 | 247.4847 | 175.5301 | -1.6003 | 1.2636 | 31.9523 | 40.3756 | 25.2863 | 1.9965 | 31.8764 | 63.6413 | 15.9661 | 14.5176 | -0.1103 | 104.0169 | 235.1171 | 180.7284 | 232.9671 | 161.0125
2012-10-30 19:29 | 28.6637 | 249.61 | 195.2329 | 247.4835 | 175.4313 | -1.6061 | 1.2637 | 31.9542 | 40.3804 | 25.2863 | 1.9966 | 31.8783 | 63.6488 | 15.9662 | 14.5485 | -0.1107 | 104.0292 | 235.0615 | 180.6844 | 232.935 | 160.8828
2012-10-30 19:29 | 28.6532 | 249.5603 | 195.2067 | 247.6414 | 175.2257 | -1.5974 | 1.2663 | 32.0212 | 40.549 | 25.2868 | 2.0008 | 31.9447 | 63.914 | 15.9662 | 14.5176 | -0.1101 | 104.4629 | 235.0427 | 180.6891 | 233.1238 | 160.7081
2012-10-30 19:29 | 28.6562 | 249.6142 | 195.2568 | 247.3976 | 175.4004 | -1.5991 | 1.2687 | 32.0837 | 40.706 | 25.2878 | 2.0047 | 32.0076 | 64.1658 | 15.9662 | 14.5002 | -0.1102 | 104.8719 | 235.114 | 180.7566 | 232.8974 | 160.9002
2012-10-30 19:29 | 28.6562 | 249.622 | 195.2372 | 247.5112 | 175.5253 | -1.5997 | 1.2693 | 32.0969 | 40.7391 | 25.2881 | 2.0055 | 32.0206 | 64.2186 | 15.966 | 14.541 | -0.1103 | 104.9577 | 235.081 | 180.6962 | 232.9702 | 160.9843
2012-10-30 19:29 | 28.654 | 249.6629 | 195.2389 | 247.462 | 175.8314 | -1.6003 | 1.2698 | 32.1099 | 40.7718 | 25.2882 | 2.0063 | 32.0332 | 64.2689 | 15.9661 | 14.5429 | -0.1103 | 105.0408 | 235.12 | 180.696 | 232.9191 | 161.2885
2012-10-30 19:29 | 28.6435 | 249.6328 | 195.2756 | 247.4045 | 175.86 | -1.6043 | 1.2702 | 32.1212 | 40.801 | 25.2879 | 2.007 | 32.0447 | 64.3153 | 15.9661 | 14.5283 | -0.1106 | 105.1162 | 235.1045 | 180.7473 | 232.8762 | 161.3317
2012-10-30 19:29 | 28.6487 | 249.6302 | 195.2965 | 247.5507 | 175.7754 | -1.6043 | 1.2705 | 32.1286 | 40.8196 | 25.2881 | 2.0076 | 32.0522 | 64.3465 | 15.9658 | 14.5091 | -0.1106 | 105.166 | 235.1211 | 180.7874 | 233.0416 | 161.2663
2012-10-30 19:29 | 28.6577 | 249.6509 | 195.3683 | 247.4266 | 175.9327 | -1.6003 | 1.2707 | 32.1332 | 40.8314 | 25.288 | 2.0078 | 32.0569 | 64.3646 | 15.9659 | 14.5466 | -0.1103 | 105.196 | 235.1043 | 180.8217 | 232.88 | 161.3861
2012-10-30 19:29 | 28.6457 | 249.6194 | 195.3721 | 247.3871 | 175.8895 | -1.6078 | 1.2708 | 32.1364 | 40.8393 | 25.2882 | 2.008 | 32.0599 | 64.3773 | 15.9659 | 14.5061 | -0.1108 | 105.2166 | 235.1133 | 180.866 | 232.881 | 161.3834
2012-10-30 19:29 | 28.6577 | 249.6235 | 195.3841 | 247.5127 | 175.6985 | -1.6037 | 1.2709 | 32.1383 | 40.8443 | 25.288 | 2.0082 | 32.0619 | 64.3855 | 15.9658 | 14.5344 | -0.1106 | 105.2298 | 235.0891 | 180.8497 | 232.9783 | 161.1641
2012-10-30 19:30 | 28.6547 | 249.6401 | 195.4125 | 247.451 | 175.8399 | -1.6009 | 1.271 | 32.1403 | 40.849 | 25.2883 | 2.0083 | 32.0639 | 64.3925 | 15.966 | 14.5559 | -0.1104 | 105.2415 | 235.0842 | 180.8566 | 232.8951 | 161.284
2012-10-30 19:30 | 28.6517 | 249.6919 | 195.3938 | 247.3892 | 175.9189 | -1.6009 | 1.271 | 32.1415 | 40.8521 | 25.2882 | 2.0084 | 32.0651 | 64.3982 | 15.9658 | 14.5161 | -0.1104 | 105.2503 | 235.1758 | 180.8777 | 232.8731 | 161.4028
2012-10-30 19:30 | 28.645 | 249.6226 | 195.391 | 247.504 | 175.7873 | -1.5997 | 1.2711 | 32.1427 | 40.8551 | 25.2883 | 2.0084 | 32.0663 | 64.4034 | 15.9657 | 14.5298 | -0.1103 | 105.2585 | 235.0928 | 180.8612 | 232.9742 | 161.2575
2012-10-30 19:30 | 28.651 | 249.6599 | 195.4245 | 247.4003 | 175.8753 | -1.602 | 1.2711 | 32.1434 | 40.8565 | 25.2885 | 2.0085 | 32.067 | 64.4061 | 15.9657 | 14.5318 | -0.1104 | 105.2626 | 235.1281 | 180.8927 | 232.8685 | 161.3435
2012-10-30 19:30 | 28.6442 | 249.6453 | 195.4216 | 247.3856 | 175.767 | -1.598 | 1.2711 | 32.144 | 40.8586 | 25.2881 | 2.0085 | 32.0676 | 64.4085 | 15.9657 | 14.5617 | -0.1102 | 105.2671 | 235.0836 | 180.8599 | 232.8239 | 161.2053
2012-10-30 19:30 | 28.651 | 249.6403 | 195.4599 | 247.459 | 175.766 | -1.6014 | 1.2711 | 32.1446 | 40.8598 | 25.2883 | 2.0086 | 32.0682 | 64.4106 | 15.9658 | 14.5073 | -0.1104 | 105.2704 | 235.133 | 180.9526 | 232.9517 | 161.2587
2012-10-30 19:30 | 28.6502 | 249.6356 | 195.463 | 247.4543 | 175.7183 | -1.602 | 1.2711 | 32.1448 | 40.8604 | 25.2883 | 2.0086 | 32.0684 | 64.4114 | 15.9658 | 14.5962 | -0.1104 | 105.2718 | 235.0394 | 180.8668 | 232.8581 | 161.1221
2012-10-30 19:30 | 28.6442 | 249.6688 | 195.4295 | 247.37 | 175.7826 | -1.6014 | 1.2711 | 32.1448 | 40.8604 | 25.2883 | 2.0086 | 32.0684 | 64.412 | 15.9657 | 14.525 | -0.1104 | 105.2725 | 235.1438 | 180.9045 | 232.845 | 161.2576
2012-10-30 19:30 | 28.6442 | 249.657 | 195.4374 | 247.5658 | 175.7201 | -1.6003 | 1.2711 | 32.1452 | 40.8609 | 25.2886 | 2.0086 | 32.0686 | 64.4127 | 15.9657 | 14.5209 | -0.1103 | 105.2736 | 235.1361 | 180.9165 | 233.0449 | 161.1992
2012-10-30 19:30 | 28.6368 | 249.6692 | 195.4614 | 247.3861 | 175.8103 | -1.5928 | 1.2711 | 32.1449 | 40.861 | 25.2881 | 2.0086 | 32.0685 | 64.4127 | 15.9657 | 14.5257 | -0.1098 | 105.2738 | 235.1435 | 180.9357 | 232.8604 | 161.2846
2012-10-30 19:30 | 28.6465 | 249.6906 | 195.4593 | 247.4154 | 175.6834 | -1.6055 | 1.2711 | 32.1452 | 40.8613 | 25.2883 | 2.0086 | 32.0687 | 64.4136 | 15.9656 | 14.5313 | -0.1107 | 105.2749 | 235.1593 | 180.928 | 232.8841 | 161.1521
2012-10-30 19:30 | 28.6465 | 249.6358 | 195.4082 | 247.5368 | 175.6561 | -1.5974 | 1.2711 | 32.1451 | 40.8608 | 25.2885 | 2.0086 | 32.0687 | 64.4136 | 15.9656 | 14.5069 | -0.1101 | 105.2745 | 235.1289 | 180.9013 | 233.0299 | 161.1492
2012-10-30 19:30 | 28.6487 | 249.7202 | 195.3986 | 247.394 | 175.791 | -1.5962 | 1.2712 | 32.1453 | 40.862 | 25.288 | 2.0086 | 32.0688 | 64.4139 | 15.9656 | 14.5376 | -0.11 | 105.2759 | 235.1826 | 180.861 | 232.8564 | 161.2534
2012-10-30 19:30 | 28.642 | 249.6744 | 195.3723 | 247.4735 | 175.6165 | -1.598 | 1.2712 | 32.1455 | 40.8622 | 25.2882 | 2.0086 | 32.069 | 64.4143 | 15.9657 | 14.5268 | -0.1102 | 105.2765 | 235.1476 | 180.8455 | 232.9467 | 161.0897
2012-10-30 19:30 | 28.6413 | 249.6502 | 195.3677 | 247.5394 | 175.7836 | -1.5997 | 1.2712 | 32.1454 | 40.8618 | 25.2883 | 2.0086 | 32.0689 | 64.4146 | 15.9655 | 14.5017 | -0.1103 | 105.2763 | 235.1485 | 180.866 | 233.0377 | 161.2819
2012-10-30 19:30 | 28.6517 | 249.6489 | 195.4252 | 247.4558 | 175.8096 | -1.598 | 1.2711 | 32.1452 | 40.861 | 25.2886 | 2.0086 | 32.0688 | 64.4136 | 15.9657 | 14.5529 | -0.1102 | 105.2746 | 235.096 | 180.8723 | 232.9029 | 161.2567
2012-10-30 19:30 | 28.6472 | 249.7305 | 195.4168 | 247.42 | 175.6139 | -1.5997 | 1.2712 | 32.1454 | 40.8618 | 25.2884 | 2.0086 | 32.069 | 64.4147 | 15.9656 | 14.5484 | -0.1103 | 105.2765 | 235.1821 | 180.8684 | 232.8716 | 161.0655
2012-10-30 19:30 | 28.645 | 249.6696 | 195.3871 | 247.5666 | 175.6312 | -1.6003 | 1.2711 | 32.145 | 40.8606 | 25.2884 | 2.0086 | 32.0686 | 64.4135 | 15.9655 | 14.5298 | -0.1103 | 105.2741 | 235.1398 | 180.8573 | 233.0368 | 161.1014
2012-10-30 19:30 | 28.642 | 249.7331 | 195.3998 | 247.4265 | 175.5774 | -1.6049 | 1.2712 | 32.1452 | 40.8614 | 25.2882 | 2.0086 | 32.0687 | 64.414 | 15.9655 | 14.5228 | -0.1106 | 105.2754 | 235.2103 | 180.877 | 232.9037 | 161.0546
2012-10-30 19:30 | 28.6442 | 249.657 | 195.3195 | 247.5267 | 175.5796 | -1.602 | 1.2711 | 32.1449 | 40.8607 | 25.2882 | 2.0086 | 32.0684 | 64.4131 | 15.9654 | 14.5413 | -0.1104 | 105.2738 | 235.1157 | 180.7782 | 232.9854 | 161.0383
2012-10-30 19:30 | 28.648 | 249.7195 | 195.3233 | 247.5148 | 175.6459 | -1.5985 | 1.2711 | 32.1445 | 40.8596 | 25.2882 | 2.0086 | 32.068 | 64.4109 | 15.9656 | 14.5247 | -0.1102 | 105.2705 | 235.1948 | 180.7986 | 232.9901 | 161.1212
2012-10-30 19:30 | 28.6435 | 249.6759 | 195.3816 | 247.4006 | 175.8248 | -1.6014 | 1.2711 | 32.1443 | 40.8594 | 25.2881 | 2.0086 | 32.0679 | 64.4108 | 15.9655 | 14.4957 | -0.1104 | 105.2702 | 235.1802 | 180.8859 | 232.9049 | 161.3291
2012-10-30 19:30 | 28.6547 | 249.6988 | 195.3968 | 247.4862 | 175.6409 | -1.5997 | 1.2711 | 32.1442 | 40.8585 | 25.2885 | 2.0086 | 32.0677 | 64.4106 | 15.9654 | 14.5191 | -0.1103 | 105.269 | 235.1797 | 180.8777 | 232.9671 | 161.1218
2012-10-30 19:30 | 28.6517 | 249.6958 | 195.3388 | 247.5929 | 175.6184 | -1.6061 | 1.2711 | 32.1439 | 40.8577 | 25.2885 | 2.0086 | 32.0675 | 64.4094 | 15.9655 | 14.5529 | -0.1107 | 105.2671 | 235.1429 | 180.7859 | 233.04 | 161.0655
2012-10-30 19:30 | 28.651 | 249.6873 | 195.3616 | 247.4277 | 175.6567 | -1.5997 | 1.2711 | 32.1438 | 40.8577 | 25.2883 | 2.0086 | 32.0674 | 64.4092 | 15.9654 | 14.4991 | -0.1103 | 105.2669 | 235.1882 | 180.8625 | 232.9286 | 161.1576
2012-10-30 19:30 | 28.6562 | 249.7317 | 195.3276 | 247.476 | 175.5799 | -1.6032 | 1.2711 | 32.1435 | 40.857 | 25.2884 | 2.0086 | 32.0671 | 64.4086 | 15.9653 | 14.5084 | -0.1105 | 105.2656 | 235.2233 | 180.8192 | 232.9676 | 161.0715
2012-10-30 19:31 | 28.6585 | 249.7066 | 195.3141 | 247.5292 | 175.4925 | -1.5997 | 1.2711 | 32.1436 | 40.8566 | 25.2887 | 2.0085 | 32.0671 | 64.408 | 15.9654 | 14.5311 | -0.1103 | 105.2646 | 235.1755 | 180.783 | 232.9981 | 160.9614
2012-10-30 19:31 | 28.6517 | 249.6958 | 195.3309 | 247.5655 | 175.6106 | -1.6066 | 1.2711 | 32.143 | 40.856 | 25.2881 | 2.0085 | 32.0666 | 64.4066 | 15.9653 | 14.5202 | -0.1108 | 105.2625 | 235.1756 | 180.8107 | 233.0453 | 161.0904
2012-10-30 19:31 | 28.663 | 249.7345 | 195.3501 | 247.4044 | 175.7077 | -1.5985 | 1.271 | 32.1428 | 40.8548 | 25.2885 | 2.0085 | 32.0663 | 64.405 | 15.9653 | 14.5111 | -0.1102 | 105.2598 | 235.2234 | 180.839 | 232.8933 | 161.1966
2012-10-30 19:31 | 28.6622 | 249.7103 | 195.3139 | 247.4859 | 175.5976 | -1.6003 | 1.2711 | 32.1426 | 40.855 | 25.2881 | 2.0085 | 32.0662 | 64.4047 | 15.9653 | 14.5552 | -0.1103 | 105.2597 | 235.1551 | 180.7587 | 232.9307 | 161.0424
2012-10-30 19:31 | 28.66 | 249.6846 | 195.3314 | 247.566 | 175.6735 | -1.5985 | 1.271 | 32.1423 | 40.8543 | 25.2882 | 2.0085 | 32.0659 | 64.4033 | 15.9654 | 14.5326 | -0.1102 | 105.2576 | 235.152 | 180.7988 | 233.0334 | 161.1409
2012-10-30 19:31 | 28.663 | 249.7306 | 195.3658 | 247.4749 | 175.817 | -1.598 | 1.271 | 32.1421 | 40.8533 | 25.2883 | 2.0085 | 32.0656 | 64.4025 | 15.9652 | 14.5111 | -0.1102 | 105.2558 | 235.2195 | 180.8547 | 232.9638 | 161.3059
2012-10-30 19:31 | 28.6577 | 249.7292 | 195.3762 | 247.4579 | 175.7493 | -1.6009 | 1.271 | 32.142 | 40.8531 | 25.2884 | 2.0085 | 32.0655 | 64.4021 | 15.9653 | 14.5385 | -0.1104 | 105.2552 | 235.1907 | 180.8377 | 232.9194 | 161.2108
2012-10-30 19:31 | 28.6645 | 249.7126 | 195.3948 | 247.5861 | 175.6195 | -1.6026 | 1.271 | 32.1414 | 40.8523 | 25.288 | 2.0084 | 32.065 | 64.4009 | 15.9651 | 14.5289 | -0.1105 | 105.2532 | 235.1837 | 180.8659 | 233.0572 | 161.0906
2012-10-30 19:31 | 28.6712 | 249.7388 | 195.374 | 247.5341 | 175.7354 | -1.5968 | 1.271 | 32.1409 | 40.8507 | 25.2881 | 2.0084 | 32.0645 | 64.3984 | 15.9651 | 14.5112 | -0.1101 | 105.2491 | 235.2276 | 180.8628 | 233.0229 | 161.2242
2012-10-30 19:31 | 28.675 | 249.7465 | 195.3896 | 247.5104 | 175.7978 | -1.5962 | 1.271 | 32.1405 | 40.85 | 25.2879 | 2.0084 | 32.0641 | 64.3964 | 15.9653 | 14.5517 | -0.11 | 105.2464 | 235.1948 | 180.8379 | 232.9587 | 161.2461
2012-10-30 19:31 | 28.6675 | 249.739 | 195.3978 | 247.499 | 175.7942 | -1.5962 | 1.271 | 32.1401 | 40.8485 | 25.2882 | 2.0084 | 32.0637 | 64.3954 | 15.9652 | 14.5442 | -0.11 | 105.2438 | 235.1948 | 180.8536 | 232.9548 | 161.25
2012-10-30 19:31 | 28.6682 | 249.6928 | 195.3985 | 247.6173 | 175.7754 | -1.6003 | 1.2709 | 32.14 | 40.8479 | 25.2884 | 2.0084 | 32.0635 | 64.3951 | 15.965 | 14.4959 | -0.1103 | 105.243 | 235.1969 | 180.9026 | 233.1214 | 161.2795
2012-10-30 19:31 | 28.6742 | 249.7575 | 195.4477 | 247.4391 | 175.7658 | -1.6072 | 1.2709 | 32.1397 | 40.8475 | 25.2881 | 2.0083 | 32.0632 | 64.3935 | 15.9651 | 14.5142 | -0.1108 | 105.241 | 235.2433 | 180.9335 | 232.9249 | 161.2516
2012-10-30 19:31 | 28.6697 | 249.7295 | 195.3921 | 247.5443 | 175.7105 | -1.5957 | 1.2709 | 32.1394 | 40.847 | 25.2881 | 2.0083 | 32.063 | 64.3925 | 15.9651 | 14.5301 | -0.11 | 105.2395 | 235.1994 | 180.862 | 233.0142 | 161.1804
2012-10-30 19:31 | 28.6667 | 249.7187 | 195.3695 | 247.5139 | 175.6568 | -1.602 | 1.2709 | 32.1388 | 40.8455 | 25.288 | 2.0083 | 32.0624 | 64.3902 | 15.9652 | 14.5107 | -0.1104 | 105.2357 | 235.208 | 180.8588 | 233.0032 | 161.1461
2012-10-30 19:31 | 28.666 | 249.761 | 195.3609 | 247.6151 | 175.578 | -1.6055 | 1.2709 | 32.1387 | 40.8456 | 25.2878 | 2.0083 | 32.0623 | 64.3901 | 15.9651 | 14.4978 | -0.1107 | 105.2356 | 235.2632 | 180.8631 | 233.1173 | 161.0802
2012-10-30 19:31 | 28.672 | 249.7592 | 195.3787 | 247.4487 | 175.6074 | -1.6026 | 1.2709 | 32.1384 | 40.8445 | 25.288 | 2.0083 | 32.062 | 64.389 | 15.965 | 14.5201 | -0.1105 | 105.2335 | 235.2391 | 180.8586 | 232.9286 | 161.0873
2012-10-30 19:31 | 28.672 | 249.7553 | 195.3237 | 247.5231 | 175.4825 | -1.5957 | 1.2709 | 32.1379 | 40.8431 | 25.2881 | 2.0082 | 32.0616 | 64.3872 | 15.9651 | 14.4997 | -0.11 | 105.2303 | 235.2556 | 180.824 | 233.0234 | 160.9828
2012-10-30 19:31 | 28.6727 | 249.7521 | 195.2773 | 247.5826 | 175.4481 | -1.6014 | 1.2709 | 32.1375 | 40.8426 | 25.2878 | 2.0082 | 32.0611 | 64.3852 | 15.9651 | 14.5127 | -0.1104 | 105.2278 | 235.2394 | 180.7646 | 233.0699 | 160.9354
2012-10-30 19:31 | 28.669 | 249.7327 | 195.2775 | 247.5436 | 175.6083 | -1.5933 | 1.2709 | 32.1375 | 40.8425 | 25.2878 | 2.0082 | 32.0611 | 64.3859 | 15.9649 | 14.4967 | -0.1098 | 105.2284 | 235.236 | 180.7808 | 233.0469 | 161.1116
2012-10-30 19:31 | 28.6735 | 249.745 | 195.3056 | 247.4463 | 175.8353 | -1.598 | 1.2709 | 32.1374 | 40.8418 | 25.2881 | 2.0082 | 32.061 | 64.3852 | 15.965 | 14.4767 | -0.1102 | 105.227 | 235.2683 | 180.8289 | 232.9696 | 161.3586
2012-10-30 19:31 | 28.6667 | 249.7069 | 195.3184 | 247.6079 | 175.5241 | -1.6014 | 1.2708 | 32.1369 | 40.8409 | 25.2879 | 2.0082 | 32.0604 | 64.383 | 15.9649 | 14.4536 | -0.1104 | 105.2239 | 235.2533 | 180.8648 | 233.1543 | 161.0705
2012-10-30 19:31 | 28.6765 | 249.795 | 195.3086 | 247.5237 | 175.6002 | -1.5985 | 1.2648 | 31.9815 | 40.449 | 25.2865 | 1.9987 | 31.9086 | 63.7742 | 15.965 | 14.4961 | -0.1102 | 104.2232 | 235.2989 | 180.8125 | 233.0276 | 161.1041
2012-10-30 19:32 | 28.6712 | 249.6801 | 195.2954 | 247.4988 | 175.7237 | -1.5997 | 1.264 | 31.9609 | 40.3974 | 25.2862 | 1.9972 | 31.8852 | 63.6797 | 15.9653 | 14.5275 | -0.1103 | 104.0771 | 235.1526 | 180.7679 | 232.9713 | 161.1962
2012-10-30 19:32 | 28.666 | 249.6162 | 195.3531 | 247.3722 | 175.7537 | -1.5997 | 1.2688 | 32.0834 | 40.7059 | 25.2874 | 2.0048 | 32.0075 | 64.1683 | 15.9655 | 14.5264 | -0.1103 | 104.8741 | 235.0898 | 180.8267 | 232.8458 | 161.2273
2012-10-30 19:32 | 28.6645 | 249.6303 | 195.3673 | 247.4686 | 175.7326 | -1.6037 | 1.2693 | 32.0987 | 40.7437 | 25.2879 | 2.0057 | 32.0222 | 64.2282 | 15.9653 | 14.5208 | -0.1106 | 104.9718 | 235.1095 | 180.8465 | 232.9478 | 161.2118
2012-10-30 19:32 | 28.6622 | 249.5811 | 195.3453 | 247.5368 | 175.746 | -1.6009 | 1.2699 | 32.1123 | 40.7785 | 25.2878 | 2.0066 | 32.0358 | 64.283 | 15.9652 | 14.5103 | -0.1104 | 105.0615 | 235.0708 | 180.835 | 233.0265 | 161.2357
2012-10-30 19:32 | 28.6765 | 249.6502 | 195.3793 | 247.3944 | 175.7524 | -1.5968 | 1.2703 | 32.1226 | 40.8043 | 25.288 | 2.0072 | 32.0461 | 64.324 | 15.9653 | 14.5165 | -0.1101 | 105.1283 | 235.1337 | 180.8628 | 232.8779 | 161.2359
2012-10-30 19:32 | 28.6645 | 249.6382 | 195.3751 | 247.3746 | 175.596 | -1.6026 | 1.2706 | 32.1303 | 40.8238 | 25.2881 | 2.0077 | 32.0538 | 64.3555 | 15.9652 | 14.5412 | -0.1105 | 105.1793 | 235.097 | 180.8339 | 232.8334 | 161.0548
2012-10-30 19:32 | 28.6615 | 249.6547 | 195.3407 | 247.3951 | 175.5969 | -1.5974 | 1.2708 | 32.1357 | 40.8376 | 25.288 | 2.0081 | 32.0593 | 64.377 | 15.9653 | 14.5096 | -0.1101 | 105.2147 | 235.1451 | 180.8311 | 232.8855 | 161.0873
2012-10-30 19:32 | 28.657 | 249.6385 | 195.3559 | 247.4924 | 175.7603 | -1.6009 | 1.2709 | 32.1387 | 40.8452 | 25.2881 | 2.0083 | 32.0623 | 64.3894 | 15.9652 | 14.5133 | -0.1104 | 105.2346 | 235.1252 | 180.8426 | 232.9791 | 161.247
2012-10-30 19:32 | 28.6682 | 249.6536 | 195.3749 | 247.3391 | 175.7754 | -1.5968 | 1.271 | 32.1406 | 40.8496 | 25.2883 | 2.0084 | 32.0641 | 64.3969 | 15.9652 | 14.5816 | -0.1101 | 105.2465 | 235.072 | 180.7933 | 232.7575 | 161.1938
2012-10-30 19:32 | 28.666 | 249.7062 | 195.3963 | 247.4466 | 175.6795 | -1.5991 | 1.271 | 32.1422 | 40.8539 | 25.2882 | 2.0085 | 32.0656 | 64.403 | 15.9651 | 14.5345 | -0.1102 | 105.2568 | 235.1717 | 180.8618 | 232.9121 | 161.145
2012-10-30 19:32 | 28.6667 | 249.6599 | 195.4206 | 247.4238 | 175.7036 | -1.6032 | 1.2711 | 32.1434 | 40.8564 | 25.2885 | 2.0086 | 32.0668 | 64.4083 | 15.965 | 14.5679 | -0.1105 | 105.2647 | 235.092 | 180.8527 | 232.8559 | 161.1357
2012-10-30 19:32 | 28.6577 | 249.6744 | 195.4037 | 247.3991 | 175.8898 | -1.6026 | 1.2711 | 32.1442 | 40.8589 | 25.2882 | 2.0086 | 32.0677 | 64.4117 | 15.965 | 14.5548 | -0.1105 | 105.2706 | 235.1196 | 180.8489 | 232.8443 | 161.335
2012-10-30 19:32 | 28.6547 | 249.7106 | 195.4204 | 247.3021 | 175.918 | -1.6014 | 1.2711 | 32.1445 | 40.8597 | 25.2883 | 2.0086 | 32.068 | 64.4131 | 15.965 | 14.5395 | -0.1104 | 105.2728 | 235.1711 | 180.8809 | 232.7626 | 161.3785
2012-10-30 19:32 | 28.6622 | 249.6829 | 195.4554 | 247.4741 | 175.824 | -1.6037 | 1.2711 | 32.1449 | 40.8604 | 25.2884 | 2.0087 | 32.0683 | 64.4149 | 15.9649 | 14.5593 | -0.1106 | 105.2752 | 235.1236 | 180.8961 | 232.9148 | 161.2647
2012-10-30 19:32 | 28.6562 | 249.6847 | 195.4729 | 247.4486 | 175.697 | -1.6043 | 1.2712 | 32.1451 | 40.8617 | 25.288 | 2.0087 | 32.0686 | 64.4154 | 15.965 | 14.5288 | -0.1106 | 105.277 | 235.1559 | 180.9441 | 232.9198 | 161.1682
2012-10-30 19:32 | 28.6517 | 249.6919 | 195.4291 | 247.3735 | 175.8018 | -1.6003 | 1.2712 | 32.1454 | 40.8618 | 25.2883 | 2.0087 | 32.0688 | 64.4164 | 15.965 | 14.561 | -0.1103 | 105.2782 | 235.1309 | 180.8681 | 232.8125 | 161.2408
2012-10-30 19:32 | 28.6502 | 249.6904 | 195.3883 | 247.3995 | 175.7652 | -1.6009 | 1.2712 | 32.1459 | 40.8632 | 25.2882 | 2.0087 | 32.0694 | 64.419 | 15.9649 | 14.5432 | -0.1104 | 105.2822 | 235.1472 | 180.8451 | 232.8563 | 161.222
2012-10-30 19:32 | 28.6547 | 249.6636 | 195.4125 | 247.5097 | 175.7775 | -1.6003 | 1.2712 | 32.1461 | 40.8633 | 25.2884 | 2.0087 | 32.0696 | 64.4196 | 15.965 | 14.564 | -0.1103 | 105.2829 | 235.0996 | 180.8485 | 232.9457 | 161.2135
2012-10-30 19:32 | 28.6562 | 249.6729 | 195.4219 | 247.4838 | 175.8414 | -1.5991 | 1.2712 | 32.1459 | 40.8628 | 25.2885 | 2.0087 | 32.0694 | 64.4184 | 15.9651 | 14.5288 | -0.1102 | 105.2812 | 235.1441 | 180.8931 | 232.955 | 161.3126
2012-10-30 19:33 | 28.6555 | 249.7114 | 195.4172 | 247.4479 | 175.8251 | -1.6032 | 1.2712 | 32.146 | 40.8635 | 25.2882 | 2.0087 | 32.0695 | 64.4196 | 15.9649 | 14.5118 | -0.1105 | 105.2831 | 235.1996 | 180.9054 | 232.9361 | 161.3133
2012-10-30 19:33 | 28.6555 | 249.731 | 195.4801 | 247.346 | 175.8798 | -1.5991 | 1.2712 | 32.146 | 40.863 | 25.2886 | 2.0087 | 32.0694 | 64.4193 | 15.9649 | 14.5199 | -0.1102 | 105.2823 | 235.2111 | 180.9602 | 232.8261 | 161.3599
2012-10-30 19:33 | 28.6547 | 249.6675 | 195.4282 | 247.498 | 175.6994 | -1.6014 | 1.2712 | 32.146 | 40.8638 | 25.2881 | 2.0088 | 32.0695 | 64.4199 | 15.9648 | 14.5355 | -0.1104 | 105.2837 | 235.132 | 180.8927 | 232.9625 | 161.1639
2012-10-30 19:33 | 28.66 | 249.712 | 195.4453 | 247.519 | 175.8648 | -1.5997 | 1.2712 | 32.1459 | 40.8634 | 25.2881 | 2.0087 | 32.0694 | 64.419 | 15.9649 | 14.5163 | -0.1103 | 105.2824 | 235.1957 | 180.929 | 233.0027 | 161.3485
2012-10-30 19:33 | 28.6592 | 249.6799 | 195.4445 | 247.432 | 175.9537 | -1.6014 | 1.2712 | 32.1455 | 40.8623 | 25.2882 | 2.0087 | 32.069 | 64.4176 | 15.9649 | 14.5277 | -0.1104 | 105.2799 | 235.1522 | 180.9168 | 232.9043 | 161.426
2012-10-30 19:33 | 28.6585 | 249.6792 | 195.4163 | 247.4234 | 175.8359 | -1.6037 | 1.2711 | 32.1452 | 40.8614 | 25.2883 | 2.0087 | 32.0686 | 64.4167 | 15.9647 | 14.4617 | -0.1106 | 105.2781 | 235.2175 | 180.9546 | 232.9617 | 161.3742
2012-10-30 19:33 | 28.672 | 249.724 | 195.4769 | 247.5153 | 175.9197 | -1.6003 | 1.2712 | 32.1452 | 40.8615 | 25.2882 | 2.0087 | 32.0686 | 64.4167 | 15.9647 | 14.5364 | -0.1103 | 105.2782 | 235.1876 | 180.9405 | 232.9789 | 161.3833
2012-10-30 19:33 | 28.6675 | 249.7273 | 195.496 | 247.3815 | 175.9659 | -1.6014 | 1.2711 | 32.1449 | 40.8607 | 25.2883 | 2.0087 | 32.0685 | 64.4157 | 15.9649 | 14.5279 | -0.1104 | 105.2765 | 235.1994 | 180.9681 | 232.8536 | 161.438
2012-10-30 19:33 | 28.6697 | 249.7178 | 195.4668 | 247.4816 | 175.8432 | -1.5997 | 1.2711 | 32.145 | 40.861 | 25.2882 | 2.0087 | 32.0685 | 64.416 | 15.9648 | 14.5056 | -0.1103 | 105.277 | 235.2122 | 180.9612 | 232.976 | 161.3376
2012-10-30 19:33 | 28.6682 | 249.7084 | 195.4731 | 247.5467 | 175.8261 | -1.5962 | 1.2711 | 32.1448 | 40.8602 | 25.2884 | 2.0087 | 32.0683 | 64.4153 | 15.9647 | 14.5775 | -0.11 | 105.2755 | 235.1309 | 180.8956 | 232.9692 | 161.2486
2012-10-30 19:33 | 28.6705 | 249.7538 | 195.5147 | 247.4668 | 176.0002 | -1.5962 | 1.2711 | 32.1444 | 40.8596 | 25.2881 | 2.0086 | 32.0678 | 64.413 | 15.9649 | 14.486 | -0.11 | 105.2726 | 235.2678 | 181.0287 | 232.9808 | 161.5142
2012-10-30 19:33 | 28.6757 | 249.7668 | 195.461 | 247.4485 | 175.9195 | -1.6032 | 1.2711 | 32.1439 | 40.8585 | 25.288 | 2.0086 | 32.0674 | 64.4117 | 15.9648 | 14.5279 | -0.1105 | 105.2702 | 235.2389 | 180.9331 | 232.9206 | 161.3916
2012-10-30 19:33 | 28.6735 | 249.745 | 195.5177 | 247.5716 | 176.0305 | -1.5957 | 1.2711 | 32.1441 | 40.8584 | 25.2885 | 2.0087 | 32.0677 | 64.4129 | 15.9648 | 14.5298 | -0.11 | 105.2713 | 235.2152 | 180.9879 | 233.0418 | 161.5007
2012-10-30 19:33 | 28.6712 | 249.7858 | 195.5154 | 247.4009 | 175.9306 | -1.6043 | 1.2711 | 32.1441 | 40.8594 | 25.2878 | 2.0086 | 32.0676 | 64.4122 | 15.9648 | 14.5234 | -0.1106 | 105.2716 | 235.2624 | 180.992 | 232.8775 | 161.4072
2012-10-30 19:33 | 28.666 | 249.7689 | 195.5024 | 247.5093 | 175.7771 | -1.5985 | 1.2711 | 32.1442 | 40.8588 | 25.2884 | 2.0087 | 32.0677 | 64.413 | 15.9648 | 14.559 | -0.1102 | 105.2718 | 235.2099 | 180.9434 | 232.9503 | 161.2181
2012-10-30 19:33 | 28.6712 | 249.7623 | 195.4919 | 247.5223 | 175.7276 | -1.6009 | 1.2711 | 32.1437 | 40.858 | 25.2881 | 2.0086 | 32.0674 | 64.4115 | 15.9648 | 14.5234 | -0.1104 | 105.2695 | 235.2389 | 180.9685 | 232.9989 | 161.2042
2012-10-30 19:33 | 28.6675 | 249.7351 | 195.4842 | 247.5813 | 175.7825 | -1.5991 | 1.2711 | 32.1435 | 40.8573 | 25.2881 | 2.0086 | 32.067 | 64.4104 | 15.9647 | 14.536 | -0.1102 | 105.2677 | 235.1991 | 180.9482 | 233.0453 | 161.2465
2012-10-30 19:33 | 28.672 | 249.8023 | 195.528 | 247.4604 | 175.8065 | -1.6049 | 1.2711 | 32.1433 | 40.8566 | 25.2882 | 2.0086 | 32.0668 | 64.4094 | 15.9647 | 14.5405 | -0.1106 | 105.2661 | 235.2618 | 180.9875 | 232.9199 | 161.266
2012-10-30 19:33 | 28.6637 | 249.747 | 195.5079 | 247.5462 | 175.806 | -1.6009 | 1.2711 | 32.1431 | 40.8563 | 25.2881 | 2.0086 | 32.0667 | 64.4094 | 15.9646 | 14.5608 | -0.1104 | 105.2658 | 235.1862 | 180.9471 | 232.9854 | 161.2452
2012-10-30 19:33 | 28.6667 | 249.7656 | 195.5227 | 247.5139 | 175.8675 | -1.5957 | 1.2711 | 32.1431 | 40.8563 | 25.2882 | 2.0086 | 32.0666 | 64.4089 | 15.9647 | 14.5311 | -0.11 | 105.2652 | 235.2345 | 180.9916 | 232.9828 | 161.3364
2012-10-30 19:33 | 28.6652 | 249.815 | 195.5527 | 247.5281 | 175.8426 | -1.5991 | 1.271 | 32.1426 | 40.8543 | 25.2886 | 2.0086 | 32.0661 | 64.4072 | 15.9646 | 14.5582 | -0.1102 | 105.2615 | 235.2568 | 180.9945 | 232.9699 | 161.2844
2012-10-30 19:33 | 28.6765 | 249.8381 | 195.5168 | 247.5237 | 175.772 | -1.6003 | 1.271 | 32.1423 | 40.8543 | 25.2881 | 2.0086 | 32.0658 | 64.4063 | 15.9645 | 14.5491 | -0.1103 | 105.2606 | 235.289 | 180.9677 | 232.9746 | 161.2229
2012-10-30 19:33 | 28.666 | 249.7454 | 195.4749 | 247.5641 | 175.539 | -1.6026 | 1.271 | 32.1421 | 40.8537 | 25.2881 | 2.0085 | 32.0656 | 64.4053 | 15.9646 | 14.5549 | -0.1105 | 105.259 | 235.1905 | 180.92 | 233.0092 | 160.9841
2012-10-30 19:33 | 28.6802 | 249.8105 | 195.4851 | 247.5979 | 175.6312 | -1.5974 | 1.271 | 32.142 | 40.8532 | 25.2883 | 2.0085 | 32.0655 | 64.405 | 15.9645 | 14.5487 | -0.1101 | 105.2582 | 235.2618 | 180.9364 | 233.0492 | 161.0825
2012-10-30 19:34 | 28.6742 | 249.7497 | 195.4359 | 247.5527 | 175.6252 | -1.6003 | 1.271 | 32.1419 | 40.8528 | 25.2884 | 2.0085 | 32.0654 | 64.4046 | 15.9646 | 14.5386 | -0.1103 | 105.2573 | 235.2111 | 180.8973 | 233.0141 | 161.0866
2012-10-30 19:34 | 28.672 | 249.8062 | 195.4062 | 247.5388 | 175.6933 | -1.6003 | 1.271 | 32.1419 | 40.8532 | 25.2882 | 2.0085 | 32.0655 | 64.4047 | 15.9646 | 14.516 | -0.1103 | 105.2579 | 235.2902 | 180.8902 | 233.0228 | 161.1773
2012-10-30 19:34 | 28.6712 | 249.8054 | 195.3858 | 247.4675 | 175.6301 | -1.6003 | 1.271 | 32.1416 | 40.8521 | 25.2883 | 2.0085 | 32.0651 | 64.4034 | 15.9646 | 14.5275 | -0.1103 | 105.2555 | 235.2779 | 180.8583 | 232.94 | 161.1026
2012-10-30 19:34 | 28.6697 | 249.7765 | 195.3725 | 247.6383 | 175.531 | -1.6026 | 1.271 | 32.1411 | 40.8507 | 25.2884 | 2.0085 | 32.0645 | 64.4007 | 15.9646 | 14.4974 | -0.1105 | 105.2514 | 235.2791 | 180.8751 | 233.1409 | 161.0336
2012-10-30 19:34 | 28.6817 | 249.7806 | 195.4081 | 247.5485 | 175.6484 | -1.5985 | 1.271 | 32.1402 | 40.849 | 25.288 | 2.0084 | 32.0637 | 64.3976 | 15.9646 | 14.538 | -0.1102 | 105.2466 | 235.2426 | 180.8701 | 233.0105 | 161.1104
2012-10-30 19:34 | 28.6832 | 249.8056 | 195.3663 | 247.5265 | 175.6421 | -1.6009 | 1.2709 | 32.1396 | 40.8472 | 25.2883 | 2.0084 | 32.0632 | 64.3957 | 15.9645 | 14.5191 | -0.1104 | 105.2429 | 235.2865 | 180.8472 | 233.0074 | 161.123
2012-10-30 19:34 | 28.6854 | 249.8 | 195.3607 | 247.607 | 175.7106 | -1.6026 | 1.2709 | 32.1397 | 40.8474 | 25.2883 | 2.0084 | 32.0632 | 64.3966 | 15.9644 | 14.5458 | -0.1105 | 105.244 | 235.2542 | 180.8149 | 233.0612 | 161.1648
2012-10-30 19:34 | 28.6824 | 249.7696 | 195.4048 | 247.5649 | 175.8832 | -1.598 | 1.2709 | 32.1396 | 40.8471 | 25.2882 | 2.0084 | 32.0632 | 64.3955 | 15.9646 | 14.5346 | -0.1102 | 105.2426 | 235.235 | 180.8702 | 233.0303 | 161.3486
2012-10-30 19:34 | 28.6884 | 249.8343 | 195.399 | 247.5904 | 175.9322 | -1.6014 | 1.2709 | 32.1391 | 40.8461 | 25.2881 | 2.0084 | 32.0627 | 64.3944 | 15.9644 | 14.5528 | -0.1104 | 105.2405 | 235.2815 | 180.8462 | 233.0376 | 161.3794
2012-10-30 19:34 | 28.6899 | 249.8202 | 195.487 | 247.5371 | 176.0352 | -1.6014 | 1.2709 | 32.1387 | 40.8447 | 25.2884 | 2.0084 | 32.0621 | 64.3922 | 15.9643 | 14.5625 | -0.1104 | 105.2369 | 235.2577 | 180.9245 | 232.9746 | 161.4727
2012-10-30 19:34 | 28.6884 | 249.7991 | 195.4816 | 247.5826 | 175.8814 | -1.5962 | 1.2641 | 31.9653 | 40.4076 | 25.2868 | 1.9976 | 31.8909 | 63.7058 | 15.9645 | 14.5692 | -0.11 | 104.1134 | 235.2299 | 180.9124 | 233.0134 | 161.3122
2012-10-30 19:34 | 28.6899 | 249.7419 | 195.5302 | 247.6468 | 176.043 | -1.602 | 1.2647 | 31.9791 | 40.4431 | 25.2865 | 1.9984 | 31.9028 | 63.753 | 15.9645 | 14.5339 | -0.1104 | 104.1961 | 235.208 | 180.9963 | 233.1129 | 161.5091
2012-10-30 19:34 | 28.6839 | 249.775 | 195.5006 | 247.5272 | 176.0838 | -1.5991 | 1.264 | 31.9629 | 40.4019 | 25.2865 | 1.9973 | 31.8865 | 63.6877 | 15.9646 | 14.5606 | -0.1102 | 104.0896 | 235.2144 | 180.94 | 232.9666 | 161.5232
2012-10-30 19:34 | 28.6899 | 249.7536 | 195.4477 | 247.5998 | 175.9103 | -1.5997 | 1.2637 | 31.9549 | 40.382 | 25.2864 | 1.9968 | 31.8789 | 63.6572 | 15.9646 | 14.5135 | -0.1103 | 104.0392 | 235.2401 | 180.9342 | 233.0863 | 161.3968
2012-10-30 19:34 | 28.6884 | 249.7482 | 195.4894 | 247.5944 | 176.0532 | -1.6014 | 1.2636 | 31.9521 | 40.3751 | 25.2862 | 1.9967 | 31.8761 | 63.6455 | 15.9647 | 14.5406 | -0.1104 | 104.0206 | 235.2076 | 180.9488 | 233.0538 | 161.5126
2012-10-30 19:34 | 28.6795 | 249.7197 | 195.5237 | 247.5541 | 175.9389 | -1.6026 | 1.2636 | 31.9519 | 40.3746 | 25.2863 | 1.9966 | 31.876 | 63.6451 | 15.9648 | 14.5562 | -0.1105 | 104.0197 | 235.1635 | 180.9675 | 232.9979 | 161.3827
2012-10-30 19:34 | 28.6832 | 249.6804 | 195.5353 | 247.4756 | 176.0207 | -1.6014 | 1.2636 | 31.9522 | 40.3752 | 25.2864 | 1.9967 | 31.8763 | 63.6461 | 15.9648 | 14.5558 | -0.1104 | 104.0213 | 235.1246 | 180.9795 | 232.9198 | 161.4649
2012-10-30 19:34 | 28.675 | 249.6957 | 195.5075 | 247.6515 | 176.0515 | -1.5968 | 1.2637 | 31.9534 | 40.3781 | 25.2865 | 1.9967 | 31.8773 | 63.6495 | 15.965 | 14.5517 | -0.1101 | 104.0277 | 235.144 | 180.9558 | 233.0998 | 161.4998
2012-10-30 19:34 | 28.6787 | 249.6837 | 195.5347 | 247.5768 | 175.7625 | -1.6009 | 1.2637 | 31.954 | 40.3794 | 25.2866 | 1.9967 | 31.878 | 63.6521 | 15.965 | 14.5635 | -0.1104 | 104.0316 | 235.1202 | 180.9712 | 233.0133 | 161.199
2012-10-30 19:34 | 28.6795 | 249.7041 | 195.5473 | 247.5189 | 175.5603 | -1.6032 | 1.2637 | 31.9555 | 40.3836 | 25.2863 | 1.9968 | 31.8794 | 63.6583 | 15.9649 | 14.6092 | -0.1105 | 104.042 | 235.0949 | 180.9381 | 232.9097 | 160.9511
2012-10-30 19:34 | 28.6712 | 249.6919 | 195.4683 | 247.4871 | 175.3256 | -1.602 | 1.2666 | 32.0275 | 40.5649 | 25.2869 | 2.0014 | 31.9525 | 63.9501 | 15.965 | 14.605 | -0.1104 | 104.515 | 235.0869 | 180.8633 | 232.8821 | 160.7206
2012-10-30 19:34 | 28.6705 | 249.6872 | 195.4362 | 247.5373 | 175.8011 | -1.6026 | 1.2689 | 32.0873 | 40.7152 | 25.2877 | 2.0051 | 32.0109 | 64.1841 | 15.9649 | 14.5553 | -0.1105 | 104.8994 | 235.1319 | 180.8809 | 232.982 | 161.2458
2012-10-30 19:34 | 28.6675 | 249.6451 | 195.4842 | 247.5578 | 176.1299 | -1.6078 | 1.2695 | 32.1034 | 40.7558 | 25.2879 | 2.0061 | 32.0268 | 64.2488 | 15.9647 | 14.5605 | -0.1108 | 105.0047 | 235.0846 | 180.9237 | 232.9973 | 161.5694
2012-10-30 19:34 | 28.6637 | 249.6765 | 195.5001 | 247.5031 | 176.1651 | -1.5985 | 1.2701 | 32.1189 | 40.7948 | 25.288 | 2.0071 | 32.0422 | 64.3103 | 15.9648 | 14.573 | -0.1102 | 105.1051 | 235.1035 | 180.9271 | 232.9301 | 161.5921
2012-10-30 19:35 | 28.669 | 249.7092 | 195.5683 | 247.477 | 176.2954 | -1.5997 | 1.2705 | 32.1288 | 40.8199 | 25.2882 | 2.0077 | 32.0522 | 64.3511 | 15.9647 | 14.5865 | -0.1103 | 105.171 | 235.1227 | 180.9818 | 232.8905 | 161.7089
2012-10-30 19:35 | 28.6637 | 249.6844 | 195.6769 | 247.5697 | 176.212 | -1.6084 | 1.2707 | 32.1338 | 40.8327 | 25.2881 | 2.008 | 32.0572 | 64.3707 | 15.9648 | 14.5934 | -0.1109 | 105.2034 | 235.091 | 181.0835 | 232.9763 | 161.6186
2012-10-30 19:35 | 28.663 | 249.715 | 195.6801 | 247.4789 | 176.4104 | -1.6066 | 1.2708 | 32.1373 | 40.8408 | 25.2885 | 2.0082 | 32.0607 | 64.3847 | 15.9648 | 14.5723 | -0.1108 | 105.2255 | 235.1427 | 181.1078 | 232.9066 | 161.8381
2012-10-30 19:35 | 28.6607 | 249.744 | 195.6857 | 247.406 | 176.33 | -1.5933 | 1.2709 | 32.1395 | 40.8468 | 25.2883 | 2.0084 | 32.0629 | 64.3945 | 15.9646 | 14.57 | -0.1098 | 105.2412 | 235.174 | 181.1157 | 232.836 | 161.76
2012-10-30 19:35 | 28.672 | 249.7279 | 195.7363 | 247.48 | 176.1695 | -1.6037 | 1.271 | 32.1412 | 40.8512 | 25.2883 | 2.0085 | 32.0646 | 64.4013 | 15.9646 | 14.5772 | -0.1106 | 105.2525 | 235.1507 | 181.1591 | 232.9028 | 161.5923
2012-10-30 19:35 | 28.6622 | 249.722 | 195.7186 | 247.529 | 176.0934 | -1.6043 | 1.271 | 32.1427 | 40.8548 | 25.2884 | 2.0086 | 32.0661 | 64.4074 | 15.9646 | 14.5307 | -0.1106 | 105.2622 | 235.1913 | 181.1879 | 232.9983 | 161.5627
2012-10-30 19:35 | 28.6675 | 249.7312 | 195.7043 | 247.5578 | 176.3602 | -1.5985 | 1.2711 | 32.1435 | 40.8573 | 25.2882 | 2.0086 | 32.067 | 64.4103 | 15.9647 | 14.5523 | -0.1102 | 105.2676 | 235.1789 | 181.152 | 233.0055 | 161.8079
2012-10-30 19:35 | 28.66 | 249.712 | 195.7557 | 247.4328 | 176.4386 | -1.6043 | 1.2711 | 32.1442 | 40.859 | 25.2882 | 2.0087 | 32.0678 | 64.4137 | 15.9647 | 14.5489 | -0.1106 | 105.2727 | 235.1631 | 181.2068 | 232.8839 | 161.8897
2012-10-30 19:35 | 28.6652 | 249.7289 | 195.7531 | 247.4536 | 176.1471 | -1.5985 | 1.2711 | 32.145 | 40.8608 | 25.2883 | 2.0087 | 32.0684 | 64.4168 | 15.9646 | 14.5011 | -0.1102 | 105.2775 | 235.2278 | 181.252 | 232.9525 | 161.646
2012-10-30 19:35 | 28.6645 | 249.7243 | 195.7524 | 247.5078 | 175.9786 | -1.602 | 1.2712 | 32.1452 | 40.8615 | 25.2883 | 2.0088 | 32.0687 | 64.4183 | 15.9645 | 14.5616 | -0.1104 | 105.2798 | 235.1627 | 181.1908 | 232.9462 | 161.417
2012-10-30 19:35 | 28.663 | 249.7071 | 195.7391 | 247.5611 | 175.8638 | -1.6055 | 1.2711 | 32.1455 | 40.8617 | 25.2886 | 2.0088 | 32.0691 | 64.4199 | 15.9644 | 14.5846 | -0.1107 | 105.2817 | 235.1225 | 181.1545 | 232.9765 | 161.2792
2012-10-30 19:35 | 28.6607 | 249.7636 | 195.7328 | 247.4805 | 176.0333 | -1.5997 | 1.2712 | 32.1461 | 40.8633 | 25.2885 | 2.0088 | 32.0695 | 64.4218 | 15.9644 | 14.5496 | -0.1103 | 105.2851 | 235.214 | 181.1832 | 232.9309 | 161.4837
2012-10-30 19:35 | 28.6652 | 249.725 | 195.6863 | 247.4497 | 175.827 | -1.6014 | 1.2712 | 32.1462 | 40.8641 | 25.2881 | 2.0088 | 32.0696 | 64.4218 | 15.9645 | 14.6031 | -0.1104 | 105.2859 | 235.1219 | 181.0832 | 232.8466 | 161.2239
2012-10-30 19:35 | 28.6637 | 249.7352 | 195.669 | 247.5501 | 175.8138 | -1.6026 | 1.2712 | 32.146 | 40.8631 | 25.2884 | 2.0088 | 32.0693 | 64.4211 | 15.9644 | 14.5404 | -0.1105 | 105.2842 | 235.1948 | 181.1286 | 233.0097 | 161.2734
2012-10-30 19:35 | 28.6682 | 249.7084 | 195.6225 | 247.6094 | 175.912 | -1.6026 | 1.2712 | 32.1462 | 40.8639 | 25.2883 | 2.0088 | 32.0696 | 64.4224 | 15.9643 | 14.5816 | -0.1105 | 105.2863 | 235.1268 | 181.0409 | 233.0278 | 161.3304
2012-10-30 19:35 | 28.6592 | 249.7307 | 195.6253 | 247.4829 | 175.985 | -1.5997 | 1.2712 | 32.1461 | 40.8636 | 25.2883 | 2.0088 | 32.0696 | 64.4226 | 15.9643 | 14.6093 | -0.1103 | 105.2862 | 235.1214 | 181.016 | 232.8736 | 161.3757
2012-10-30 19:35 | 28.663 | 249.7737 | 195.6133 | 247.4593 | 175.9771 | -1.5968 | 1.2712 | 32.1463 | 40.864 | 25.2884 | 2.0088 | 32.0697 | 64.4224 | 15.9644 | 14.5723 | -0.1101 | 105.2864 | 235.2014 | 181.041 | 232.887 | 161.4048
2012-10-30 19:35 | 28.675 | 249.7505 | 195.6253 | 247.6201 | 175.9656 | -1.5933 | 1.2712 | 32.1462 | 40.8638 | 25.2883 | 2.0088 | 32.0697 | 64.4226 | 15.9643 | 14.6251 | -0.1098 | 105.2864 | 235.1254 | 181.0002 | 232.995 | 161.3405
2012-10-30 19:35 | 28.6705 | 249.7186 | 195.6523 | 247.5765 | 176.0275 | -1.6003 | 1.2712 | 32.1462 | 40.8635 | 25.2885 | 2.0088 | 32.0696 | 64.4225 | 15.9643 | 14.5961 | -0.1103 | 105.2859 | 235.1225 | 181.0562 | 232.9804 | 161.4314
2012-10-30 19:35 | 28.6727 | 249.7442 | 195.6545 | 247.4729 | 176.1429 | -1.6043 | 1.2712 | 32.1461 | 40.8631 | 25.2885 | 2.0088 | 32.0695 | 64.4222 | 15.9642 | 14.5494 | -0.1106 | 105.2853 | 235.1948 | 181.1051 | 232.9235 | 161.5935
2012-10-30 19:35 | 28.6705 | 249.7225 | 195.6523 | 247.5608 | 176.0158 | -1.6003 | 1.2712 | 32.1459 | 40.8632 | 25.2882 | 2.0088 | 32.0693 | 64.4217 | 15.9642 | 14.5513 | -0.1103 | 105.2849 | 235.1712 | 181.101 | 233.0095 | 161.4645
2012-10-30 19:35 | 28.6727 | 249.7168 | 195.6348 | 247.5826 | 175.8501 | -1.6095 | 1.2711 | 32.1454 | 40.8617 | 25.2885 | 2.0088 | 32.0689 | 64.4202 | 15.9642 | 14.5331 | -0.111 | 105.2818 | 235.1837 | 181.1017 | 233.0495 | 161.317
2012-10-30 19:35 | 28.6817 | 249.7728 | 195.5967 | 247.5524 | 175.9723 | -1.6055 | 1.2711 | 32.1452 | 40.8612 | 25.2884 | 2.0088 | 32.0687 | 64.4193 | 15.9642 | 14.538 | -0.1107 | 105.2805 | 235.2348 | 181.0587 | 233.0144 | 161.4343
2012-10-30 19:35 | 28.6772 | 249.7566 | 195.5647 | 247.5636 | 175.761 | -1.5951 | 1.2712 | 32.1453 | 40.8616 | 25.2882 | 2.0088 | 32.0688 | 64.4199 | 15.9641 | 14.5661 | -0.11 | 105.2816 | 235.1905 | 180.9986 | 232.9975 | 161.1949
2012-10-30 19:35 | 28.6772 | 249.7644 | 195.5568 | 247.5557 | 175.8702 | -1.598 | 1.2711 | 32.1452 | 40.8612 | 25.2884 | 2.0088 | 32.0686 | 64.4193 | 15.9641 | 14.5049 | -0.1102 | 105.2805 | 235.2595 | 181.0519 | 233.0508 | 161.3653
2012-10-30 19:36 | 28.6847 | 249.7719 | 195.5564 | 247.5789 | 175.8387 | -1.6084 | 1.2711 | 32.1451 | 40.8609 | 25.2884 | 2.0088 | 32.0685 | 64.4184 | 15.9642 | 14.5491 | -0.1109 | 105.2793 | 235.2228 | 181.0073 | 233.0298 | 161.2896
2012-10-30 19:36 | 28.6929 | 249.741 | 195.5529 | 247.6615 | 175.7337 | -1.6003 | 1.2711 | 32.1446 | 40.8595 | 25.2884 | 2.0088 | 32.068 | 64.4166 | 15.9641 | 14.5696 | -0.1103 | 105.2761 | 235.1714 | 180.9833 | 233.0919 | 161.1641
2012-10-30 19:36 | 28.6854 | 249.7765 | 195.5296 | 247.5639 | 175.6208 | -1.5991 | 1.2711 | 32.1442 | 40.8588 | 25.2883 | 2.0087 | 32.0677 | 64.4157 | 15.9641 | 14.5417 | -0.1102 | 105.2745 | 235.2348 | 180.9879 | 233.0222 | 161.0791
2012-10-30 19:36 | 28.6877 | 249.7906 | 195.5005 | 247.5819 | 175.6153 | -1.5968 | 1.2711 | 32.1436 | 40.8572 | 25.2884 | 2.0087 | 32.0672 | 64.4136 | 15.9641 | 14.5032 | -0.1101 | 105.2708 | 235.2874 | 180.9973 | 233.0787 | 161.1121
2012-10-30 19:36 | 28.6922 | 249.799 | 195.4814 | 247.5472 | 175.7838 | -1.6026 | 1.2711 | 32.1431 | 40.8563 | 25.2881 | 2.0087 | 32.0666 | 64.4117 | 15.964 | 14.475 | -0.1105 | 105.2681 | 235.324 | 181.0064 | 233.0722 | 161.3088
2012-10-30 19:36 | 28.7042 | 249.811 | 195.4699 | 247.5827 | 175.7645 | -1.6003 | 1.2711 | 32.1428 | 40.8554 | 25.2882 | 2.0087 | 32.0662 | 64.41 | 15.964 | 14.5278 | -0.1103 | 105.2655 | 235.2832 | 180.9421 | 233.0549 | 161.2367
2012-10-30 19:36 | 28.6952 | 249.7511 | 195.4766 | 247.6678 | 175.8804 | -1.6066 | 1.2711 | 32.1429 | 40.8559 | 25.2881 | 2.0087 | 32.0664 | 64.4106 | 15.964 | 14.5515 | -0.1108 | 105.2665 | 235.1996 | 180.9251 | 233.1163 | 161.3289
2012-10-30 19:36 | 28.7049 | 249.7843 | 195.5098 | 247.6187 | 175.9175 | -1.598 | 1.2711 | 32.1428 | 40.8556 | 25.2881 | 2.0087 | 32.0664 | 64.4106 | 15.964 | 14.5245 | -0.1102 | 105.2662 | 235.2598 | 180.9853 | 233.0942 | 161.393
2012-10-30 19:36 | 28.7034 | 249.8102 | 195.5005 | 247.5741 | 175.7637 | -1.5991 | 1.271 | 32.1424 | 40.8542 | 25.2884 | 2.0086 | 32.0659 | 64.4087 | 15.964 | 14.527 | -0.1102 | 105.2628 | 235.2832 | 180.9735 | 233.0471 | 161.2367
2012-10-30 19:36 | 28.7042 | 249.7875 | 195.5249 | 247.6376 | 175.667 | -1.5985 | 1.271 | 32.1422 | 40.8536 | 25.2883 | 2.0086 | 32.0657 | 64.4078 | 15.964 | 14.5686 | -0.1102 | 105.2614 | 235.2189 | 180.9563 | 233.069 | 161.0984
2012-10-30 19:36 | 28.7079 | 249.7677 | 195.5011 | 247.6765 | 175.7175 | -1.6009 | 1.271 | 32.1421 | 40.8535 | 25.2882 | 2.0086 | 32.0655 | 64.4076 | 15.9639 | 14.5193 | -0.1104 | 105.2612 | 235.2484 | 180.9818 | 233.1572 | 161.1982
2012-10-30 19:36 | 28.7102 | 249.7817 | 195.4955 | 247.6397 | 175.673 | -1.598 | 1.271 | 32.1418 | 40.8527 | 25.2883 | 2.0086 | 32.0652 | 64.4066 | 15.9639 | 14.5461 | -0.1102 | 105.2593 | 235.2356 | 180.9494 | 233.0936 | 161.1269
2012-10-30 19:36 | 28.7124 | 249.8153 | 195.4938 | 247.5635 | 175.6674 | -1.6009 | 1.271 | 32.1413 | 40.8516 | 25.2883 | 2.0086 | 32.0647 | 64.4042 | 15.9639 | 14.536 | -0.1104 | 105.2558 | 235.2793 | 180.9578 | 233.0275 | 161.1314
2012-10-30 19:36 | 28.7192 | 249.8221 | 195.4259 | 247.6017 | 175.6702 | -1.5997 | 1.271 | 32.1406 | 40.8501 | 25.288 | 2.0085 | 32.0641 | 64.402 | 15.9639 | 14.5592 | -0.1103 | 105.2521 | 235.2629 | 180.8667 | 233.0425 | 161.111
2012-10-30 19:36 | 28.7087 | 249.8272 | 195.4979 | 247.7126 | 175.7573 | -1.6014 | 1.271 | 32.1403 | 40.8493 | 25.2881 | 2.0085 | 32.0639 | 64.4008 | 15.9639 | 14.4875 | -0.1104 | 105.2501 | 235.3397 | 181.0104 | 233.2251 | 161.2698
2012-10-30 19:36 | 28.7154 | 249.8261 | 195.4418 | 247.5626 | 175.6704 | -1.5997 | 1.2709 | 32.1395 | 40.8472 | 25.2881 | 2.0085 | 32.0631 | 64.3981 | 15.9639 | 14.4901 | -0.1103 | 105.2453 | 235.336 | 180.9517 | 233.0725 | 161.1803
2012-10-30 19:36 | 28.7199 | 249.7993 | 195.4502 | 247.6298 | 175.6631 | -1.598 | 1.2709 | 32.1391 | 40.8459 | 25.2883 | 2.0084 | 32.0627 | 64.396 | 15.964 | 14.5517 | -0.1102 | 105.242 | 235.2476 | 180.8985 | 233.0781 | 161.1114
2012-10-30 19:36 | 28.7266 | 249.7629 | 195.3783 | 247.7266 | 175.4903 | -1.598 | 1.2709 | 32.1391 | 40.8458 | 25.2883 | 2.0085 | 32.0626 | 64.3966 | 15.9638 | 14.538 | -0.1102 | 105.2424 | 235.2249 | 180.8403 | 233.1886 | 160.9523
2012-10-30 19:36 | 28.7311 | 249.8144 | 195.4457 | 247.6371 | 175.7017 | -1.6037 | 1.2709 | 32.139 | 40.8456 | 25.2883 | 2.0084 | 32.0624 | 64.3958 | 15.9638 | 14.5343 | -0.1106 | 105.2413 | 235.2801 | 180.9114 | 233.1028 | 161.1674
2012-10-30 19:36 | 28.7177 | 249.8245 | 195.3891 | 247.5727 | 175.6297 | -1.6014 | 1.2709 | 32.1389 | 40.8455 | 25.2883 | 2.0084 | 32.0624 | 64.3956 | 15.9637 | 14.5128 | -0.1104 | 105.2411 | 235.3117 | 180.8763 | 233.0599 | 161.1169
2012-10-30 19:36 | 28.7281 | 249.7957 | 195.4152 | 247.732 | 175.7845 | -1.602 | 1.2709 | 32.1384 | 40.8442 | 25.2882 | 2.0084 | 32.0619 | 64.3932 | 15.9639 | 14.5109 | -0.1104 | 105.2374 | 235.2848 | 180.9043 | 233.2211 | 161.2736
2012-10-30 19:36 | 28.7281 | 249.7996 | 195.427 | 247.6262 | 175.9133 | -1.6009 | 1.2709 | 32.138 | 40.8435 | 25.288 | 2.0084 | 32.0615 | 64.3919 | 15.9638 | 14.5354 | -0.1104 | 105.2354 | 235.2642 | 180.8916 | 233.0908 | 161.3779
2012-10-30 19:36 | 28.7296 | 249.809 | 195.4442 | 247.6003 | 175.8368 | -1.6014 | 1.2709 | 32.1378 | 40.8432 | 25.2879 | 2.0084 | 32.0613 | 64.3913 | 15.9637 | 14.5614 | -0.1104 | 105.2345 | 235.2476 | 180.8828 | 233.0389 | 161.2754
2012-10-30 19:36 | 28.7177 | 249.7814 | 195.4166 | 247.6472 | 175.8444 | -1.598 | 1.2687 | 32.0821 | 40.7017 | 25.2879 | 2.0048 | 32.0041 | 64.1617 | 15.9638 | 14.5046 | -0.1102 | 104.8634 | 235.2768 | 180.912 | 233.1426 | 161.3398
2012-10-30 19:37 | 28.7311 | 249.7987 | 195.4968 | 247.7154 | 176.0295 | -1.5957 | 1.2648 | 31.9816 | 40.4493 | 25.2865 | 1.9986 | 31.9052 | 63.7655 | 15.9638 | 14.5507 | -0.11 | 104.2147 | 235.248 | 180.9461 | 233.1647 | 161.4788
2012-10-30 19:37 | 28.7184 | 249.8173 | 195.4683 | 247.6165 | 176.0325 | -1.6003 | 1.2643 | 31.9702 | 40.4204 | 25.2865 | 1.9979 | 31.8943 | 63.7216 | 15.9639 | 14.5135 | -0.1103 | 104.142 | 235.3038 | 180.9548 | 233.103 | 161.519
2012-10-30 19:37 | 28.7192 | 249.7633 | 195.4652 | 247.6056 | 176.0567 | -1.5974 | 1.2639 | 31.9583 | 40.3909 | 25.2862 | 1.9971 | 31.8824 | 63.674 | 15.964 | 14.5551 | -0.1101 | 104.065 | 235.2082 | 180.9101 | 233.0505 | 161.5016
2012-10-30 19:37 | 28.7251 | 249.7418 | 195.5025 | 247.7447 | 176.0821 | -1.6032 | 1.2636 | 31.953 | 40.3774 | 25.2863 | 1.9968 | 31.877 | 63.6525 | 15.964 | 14.5243 | -0.1105 | 104.0299 | 235.2175 | 180.9782 | 233.2204 | 161.5578
2012-10-30 19:37 | 28.7199 | 249.7484 | 195.5131 | 247.5749 | 176.1784 | -1.598 | 1.2635 | 31.95 | 40.3701 | 25.2861 | 1.9966 | 31.8741 | 63.6405 | 15.9641 | 14.5884 | -0.1102 | 104.0106 | 235.16 | 180.9247 | 232.9865 | 161.59
2012-10-30 19:37 | 28.7199 | 249.721 | 195.5091 | 247.6337 | 176.0808 | -1.598 | 1.2636 | 31.9523 | 40.3757 | 25.2862 | 1.9967 | 31.8763 | 63.6488 | 15.9641 | 14.5191 | -0.1102 | 104.0245 | 235.2019 | 180.99 | 233.1146 | 161.5617
2012-10-30 19:37 | 28.7244 | 249.7176 | 195.5765 | 247.6656 | 176.1985 | -1.5974 | 1.2637 | 31.954 | 40.3798 | 25.2863 | 1.9968 | 31.8778 | 63.6548 | 15.9642 | 14.5113 | -0.1101 | 104.0346 | 235.2063 | 181.0652 | 233.1543 | 161.6872
2012-10-30 19:37 | 28.7244 | 249.7294 | 195.6119 | 247.6225 | 176.2961 | -1.602 | 1.2637 | 31.9547 | 40.3814 | 25.2865 | 1.9969 | 31.8787 | 63.6581 | 15.9642 | 14.5195 | -0.1104 | 104.0395 | 235.2099 | 181.0924 | 233.103 | 161.7766
2012-10-30 19:37 | 28.7206 | 249.7099 | 195.6159 | 247.5012 | 176.1869 | -1.6014 | 1.2637 | 31.9559 | 40.3843 | 25.2866 | 1.997 | 31.8799 | 63.663 | 15.9642 | 14.5157 | -0.1104 | 104.0472 | 235.1942 | 181.1002 | 232.9855 | 161.6712
2012-10-30 19:37 | 28.7251 | 249.6518 | 195.6204 | 247.6859 | 176.1524 | -1.602 | 1.2638 | 31.9582 | 40.3902 | 25.2865 | 1.9971 | 31.8821 | 63.672 | 15.9642 | 14.512 | -0.1104 | 104.0622 | 235.1398 | 181.1084 | 233.1739 | 161.6404
2012-10-30 19:37 | 28.7274 | 249.6854 | 195.6227 | 247.5864 | 176.221 | -1.6014 | 1.2684 | 32.0746 | 40.6834 | 25.2874 | 2.0041 | 31.9947 | 64.1219 | 15.9643 | 14.5266 | -0.1104 | 104.8054 | 235.1588 | 181.0961 | 233.0598 | 161.6944
2012-10-30 19:37 | 28.7206 | 249.6708 | 195.7142 | 247.4895 | 176.2611 | -1.5985 | 1.2694 | 32.0997 | 40.7465 | 25.2878 | 2.0059 | 32.0228 | 64.2355 | 15.9641 | 14.4953 | -0.1102 | 104.9819 | 235.1755 | 181.2189 | 232.9942 | 161.7658
2012-10-30 19:37 | 28.7154 | 249.7008 | 195.6854 | 247.5273 | 176.2051 | -1.598 | 1.2699 | 32.1145 | 40.7835 | 25.2882 | 2.0069 | 32.038 | 64.2961 | 15.9642 | 14.5105 | -0.1102 | 105.0796 | 235.1903 | 181.1749 | 233.0168 | 161.6946
2012-10-30 19:37 | 28.7124 | 249.6704 | 195.7138 | 247.6184 | 176.206 | -1.6009 | 1.2704 | 32.1253 | 40.8112 | 25.288 | 2.0076 | 32.0488 | 64.34 | 15.9641 | 14.5034 | -0.1104 | 105.1513 | 235.167 | 181.2104 | 233.115 | 161.7026
2012-10-30 19:37 | 28.7124 | 249.7174 | 195.7531 | 247.4852 | 176.3388 | -1.5991 | 1.2706 | 32.1324 | 40.8286 | 25.2884 | 2.008 | 32.0559 | 64.3682 | 15.9641 | 14.4952 | -0.1102 | 105.1968 | 235.2222 | 181.2579 | 232.99 | 161.8436
2012-10-30 19:37 | 28.7079 | 249.7599 | 195.7525 | 247.5198 | 176.1781 | -1.602 | 1.2709 | 32.1377 | 40.8423 | 25.2883 | 2.0083 | 32.0612 | 64.3896 | 15.9641 | 14.4948 | -0.1104 | 105.2319 | 235.2651 | 181.2577 | 233.025 | 161.6833
2012-10-30 19:37 | 28.7117 | 249.7519 | 195.7288 | 247.5432 | 176.1507 | -1.6003 | 1.271 | 32.141 | 40.8506 | 25.2883 | 2.0086 | 32.0645 | 64.4038 | 15.9638 | 14.5313 | -0.1103 | 105.2545 | 235.2206 | 181.1975 | 233.0119 | 161.6194
2012-10-30 19:37 | 28.7072 | 249.7279 | 195.7597 | 247.6367 | 176.1579 | -1.5991 | 1.2711 | 32.1431 | 40.8557 | 25.2884 | 2.0087 | 32.0665 | 64.4112 | 15.9641 | 14.5186 | -0.1102 | 105.2669 | 235.2093 | 181.2411 | 233.1181 | 161.6393
2012-10-30 19:37 | 28.7012 | 249.7453 | 195.7262 | 247.4426 | 176.105 | -1.6037 | 1.2711 | 32.1442 | 40.859 | 25.2882 | 2.0087 | 32.0677 | 64.4158 | 15.9641 | 14.5534 | -0.1106 | 105.2748 | 235.1919 | 181.1728 | 232.8892 | 161.5516
2012-10-30 19:37 | 28.7019 | 249.7813 | 195.7269 | 247.506 | 175.9613 | -1.6003 | 1.2711 | 32.1452 | 40.8612 | 25.2885 | 2.0088 | 32.0686 | 64.4199 | 15.964 | 14.5255 | -0.1103 | 105.2811 | 235.2558 | 181.2014 | 232.9805 | 161.4358
2012-10-30 19:37 | 28.6974 | 249.8042 | 195.6595 | 247.5054 | 175.9841 | -1.5974 | 1.2712 | 32.146 | 40.8629 | 25.2886 | 2.0089 | 32.0694 | 64.4229 | 15.964 | 14.521 | -0.1101 | 105.2858 | 235.2832 | 181.1385 | 232.9844 | 161.4631
2012-10-30 19:37 | 28.6997 | 249.7673 | 195.6854 | 247.6292 | 176.0059 | -1.602 | 1.2712 | 32.1461 | 40.8637 | 25.2883 | 2.0089 | 32.0695 | 64.4242 | 15.9638 | 14.5152 | -0.1104 | 105.2879 | 235.2521 | 181.1702 | 233.114 | 161.4907
2012-10-30 19:37 | 28.7004 | 249.8228 | 195.6979 | 247.528 | 176.0808 | -1.6014 | 1.2712 | 32.147 | 40.8658 | 25.2884 | 2.0089 | 32.0704 | 64.4273 | 15.9639 | 14.5036 | -0.1104 | 105.2931 | 235.3192 | 181.1943 | 233.0244 | 161.5772
2012-10-30 19:37 | 28.6982 | 249.8011 | 195.6957 | 247.4318 | 176.1762 | -1.598 | 1.2712 | 32.1475 | 40.8671 | 25.2885 | 2.009 | 32.0709 | 64.4298 | 15.9638 | 14.481 | -0.1102 | 105.2968 | 235.3201 | 181.2147 | 232.9508 | 161.6952
2012-10-30 19:38 | 28.7019 | 249.8048 | 195.723 | 247.5217 | 176.2033 | -1.6003 | 1.2712 | 32.1476 | 40.8674 | 25.2884 | 2.009 | 32.0711 | 64.4305 | 15.9637 | 14.5011 | -0.1103 | 105.2979 | 235.3037 | 181.2219 | 233.0206 | 161.7022
2012-10-30 19:38 | 28.6907 | 249.821 | 195.7078 | 247.5457 | 176.0906 | -1.6014 | 1.2712 | 32.1479 | 40.868 | 25.2884 | 2.009 | 32.0713 | 64.4319 | 15.9637 | 14.4776 | -0.1104 | 105.2999 | 235.3434 | 181.2302 | 233.0681 | 161.613
2012-10-30 19:38 | 28.6952 | 249.8098 | 195.7084 | 247.6599 | 176.0561 | -1.6032 | 1.2713 | 32.148 | 40.8687 | 25.2882 | 2.009 | 32.0714 | 64.4318 | 15.9638 | 14.5188 | -0.1105 | 105.3005 | 235.291 | 181.1896 | 233.1411 | 161.5373
2012-10-30 19:38 | 28.7027 | 249.8056 | 195.7356 | 247.5186 | 176.1417 | -1.5974 | 1.2713 | 32.1478 | 40.868 | 25.2883 | 2.009 | 32.0712 | 64.4309 | 15.9638 | 14.4733 | -0.1101 | 105.2988 | 235.3323 | 181.2623 | 233.0453 | 161.6684
2012-10-30 19:38 | 28.6937 | 249.8553 | 195.6637 | 247.5527 | 175.9414 | -1.5985 | 1.2712 | 32.1478 | 40.8679 | 25.2884 | 2.009 | 32.0713 | 64.4317 | 15.9637 | 14.4929 | -0.1102 | 105.2996 | 235.3624 | 181.1708 | 233.0598 | 161.4485
2012-10-30 19:38 | 28.6944 | 249.7894 | 195.6762 | 247.5298 | 176.0241 | -1.5985 | 1.2712 | 32.1477 | 40.8678 | 25.2883 | 2.009 | 32.0711 | 64.4311 | 15.9636 | 14.5058 | -0.1102 | 105.2989 | 235.2836 | 181.1704 | 233.024 | 161.5183
2012-10-30 19:38 | 28.6982 | 249.8011 | 195.676 | 247.6198 | 176.0474 | -1.602 | 1.2713 | 32.1483 | 40.869 | 25.2884 | 2.009 | 32.0717 | 64.4329 | 15.9638 | 14.5096 | -0.1104 | 105.3019 | 235.2915 | 181.1664 | 233.1102 | 161.5378
2012-10-30 19:38 | 28.6944 | 249.7894 | 195.7076 | 247.6356 | 176.0631 | -1.6003 | 1.2713 | 32.1481 | 40.8684 | 25.2884 | 2.009 | 32.0714 | 64.4329 | 15.9635 | 14.5221 | -0.1103 | 105.3014 | 235.2673 | 181.1855 | 233.1135 | 161.541
2012-10-30 19:38 | 28.7004 | 249.8307 | 195.694 | 247.5711 | 176.1277 | -1.6003 | 1.2712 | 32.1481 | 40.8681 | 25.2887 | 2.009 | 32.0715 | 64.4332 | 15.9636 | 14.4955 | -0.1103 | 105.3013 | 235.3352 | 181.1985 | 233.0756 | 161.6322
2012-10-30 19:38 | 28.6914 | 249.7825 | 195.6535 | 247.5621 | 176.0055 | -1.6037 | 1.2713 | 32.148 | 40.8689 | 25.2881 | 2.009 | 32.0713 | 64.4325 | 15.9635 | 14.4906 | -0.1106 | 105.3014 | 235.2919 | 181.1629 | 233.0715 | 161.5149
2012-10-30 19:38 | 28.6944 | 249.8012 | 195.6605 | 247.5925 | 175.825 | -1.6032 | 1.2712 | 32.1477 | 40.8677 | 25.2882 | 2.009 | 32.071 | 64.4308 | 15.9636 | 14.5344 | -0.1105 | 105.2985 | 235.2668 | 181.1261 | 233.0581 | 161.2906
2012-10-30 19:38 | 28.7012 | 249.8393 | 195.6555 | 247.6385 | 175.7537 | -1.6026 | 1.2712 | 32.1474 | 40.8671 | 25.2882 | 2.009 | 32.0707 | 64.4302 | 15.9635 | 14.5412 | -0.1105 | 105.2972 | 235.2981 | 181.1143 | 233.0973 | 161.2125
2012-10-30 19:38 | 28.6952 | 249.8529 | 195.6613 | 247.609 | 175.8258 | -1.6026 | 1.2712 | 32.1471 | 40.8661 | 25.2884 | 2.009 | 32.0706 | 64.4297 | 15.9635 | 14.5352 | -0.1105 | 105.2958 | 235.3177 | 181.1261 | 233.0738 | 161.2906
2012-10-30 19:38 | 28.7019 | 249.8361 | 195.6208 | 247.5648 | 175.9769 | -1.6026 | 1.2712 | 32.1469 | 40.8661 | 25.2881 | 2.009 | 32.0703 | 64.4287 | 15.9635 | 14.5419 | -0.1105 | 105.2948 | 235.2942 | 181.0789 | 233.0229 | 161.435
2012-10-30 19:38 | 28.6952 | 249.8294 | 195.6416 | 247.6012 | 175.7985 | -1.5997 | 1.2712 | 32.1468 | 40.8654 | 25.2883 | 2.009 | 32.0702 | 64.428 | 15.9635 | 14.4984 | -0.1103 | 105.2935 | 235.331 | 181.1432 | 233.1028 | 161.3001
2012-10-30 19:38 | 28.6982 | 249.8206 | 195.5896 | 247.6159 | 175.7507 | -1.6043 | 1.2712 | 32.1465 | 40.8642 | 25.2885 | 2.009 | 32.0699 | 64.4269 | 15.9635 | 14.5137 | -0.1106 | 105.2911 | 235.3069 | 181.0759 | 233.1022 | 161.237
2012-10-30 19:38 | 28.6982 | 249.8402 | 195.6053 | 247.6903 | 175.8834 | -1.5997 | 1.2712 | 32.1461 | 40.8635 | 25.2884 | 2.0089 | 32.0696 | 64.4256 | 15.9635 | 14.5055 | -0.1103 | 105.2892 | 235.3347 | 181.0998 | 233.1848 | 161.3779
2012-10-30 19:38 | 28.6989 | 249.8644 | 195.5706 | 247.597 | 175.81 | -1.6043 | 1.2712 | 32.1457 | 40.8626 | 25.2883 | 2.0089 | 32.0691 | 64.4243 | 15.9634 | 14.5307 | -0.1106 | 105.2869 | 235.3337 | 181.0399 | 233.0663 | 161.2793
2012-10-30 19:38 | 28.6967 | 249.8035 | 195.537 | 247.6301 | 175.7102 | -1.6003 | 1.2712 | 32.1457 | 40.8621 | 25.2886 | 2.0089 | 32.069 | 64.4235 | 15.9635 | 14.4795 | -0.1103 | 105.2857 | 235.324 | 181.0575 | 233.1506 | 161.2307
2012-10-30 19:38 | 28.7049 | 249.843 | 195.5413 | 247.5913 | 175.8004 | -1.5957 | 1.2711 | 32.1453 | 40.8614 | 25.2884 | 2.0089 | 32.0687 | 64.4223 | 15.9635 | 14.5449 | -0.11 | 105.2837 | 235.2981 | 180.9964 | 233.0464 | 161.2555
2012-10-30 19:38 | 28.6944 | 249.8403 | 195.5151 | 247.6513 | 175.7313 | -1.5985 | 1.2712 | 32.1451 | 40.8613 | 25.2882 | 2.0089 | 32.0686 | 64.4222 | 15.9634 | 14.5344 | -0.1102 | 105.2836 | 235.3059 | 180.9807 | 233.1169 | 161.1969
2012-10-30 19:38 | 28.7064 | 249.8641 | 195.4996 | 247.6633 | 175.8019 | -1.5991 | 1.2711 | 32.1448 | 40.8606 | 25.2882 | 2.0089 | 32.0683 | 64.421 | 15.9633 | 14.5137 | -0.1102 | 105.2816 | 235.3504 | 180.9859 | 233.1496 | 161.2882
2012-10-30 19:38 | 28.7072 | 249.8257 | 195.5082 | 247.6053 | 175.7675 | -1.6009 | 1.2711 | 32.1445 | 40.8599 | 25.2881 | 2.0089 | 32.068 | 64.42 | 15.9633 | 14.4941 | -0.1104 | 105.2799 | 235.3316 | 181.0141 | 233.1112 | 161.2734
2012-10-30 19:39 | 28.7034 | 249.7984 | 195.4926 | 247.6603 | 175.9199 | -1.6032 | 1.2711 | 32.144 | 40.8583 | 25.2883 | 2.0088 | 32.0673 | 64.4171 | 15.9633 | 14.4781 | -0.1105 | 105.2754 | 235.3203 | 181.0145 | 233.1822 | 161.4418
2012-10-30 19:39 | 28.7042 | 249.8462 | 195.4777 | 247.5475 | 175.9011 | -1.6037 | 1.2711 | 32.1435 | 40.857 | 25.2882 | 2.0088 | 32.067 | 64.4162 | 15.9632 | 14.4666 | -0.1106 | 105.2732 | 235.3796 | 181.0111 | 233.0809 | 161.4345
2012-10-30 19:39 | 28.7132 | 249.8591 | 195.5024 | 247.6623 | 176.0077 | -1.5997 | 1.2711 | 32.143 | 40.8559 | 25.2882 | 2.0088 | 32.0664 | 64.4139 | 15.9632 | 14.4756 | -0.1103 | 105.2698 | 235.3835 | 181.0268 | 233.1867 | 161.5321
2012-10-30 19:39 | 28.7102 | 249.8522 | 195.5702 | 247.6788 | 175.9813 | -1.5974 | 1.271 | 32.1425 | 40.8545 | 25.2883 | 2.0087 | 32.0659 | 64.4121 | 15.9632 | 14.5053 | -0.1101 | 105.2666 | 235.3469 | 181.0649 | 233.1735 | 161.476
2012-10-30 19:39 | 28.7154 | 249.8652 | 195.6068 | 247.6645 | 176.0412 | -1.5997 | 1.2711 | 32.1426 | 40.8549 | 25.2882 | 2.0087 | 32.066 | 64.4121 | 15.9632 | 14.486 | -0.1103 | 105.267 | 235.3792 | 181.1208 | 233.1785 | 161.5552
2012-10-30 19:39 | 28.7124 | 249.8427 | 195.592 | 247.587 | 176.0694 | -1.5962 | 1.271 | 32.1424 | 40.8543 | 25.2883 | 2.0087 | 32.0659 | 64.4121 | 15.9632 | 14.5279 | -0.11 | 105.2664 | 235.3148 | 181.0641 | 233.0591 | 161.5415
2012-10-30 19:39 | 28.7079 | 249.8851 | 195.6307 | 247.5747 | 176.1703 | -1.6037 | 1.271 | 32.1422 | 40.8538 | 25.2883 | 2.0087 | 32.0656 | 64.4111 | 15.9631 | 14.4703 | -0.1106 | 105.2649 | 235.4148 | 181.1604 | 233.1044 | 161.7
2012-10-30 19:39 | 28.7139 | 249.8207 | 195.6092 | 247.71 | 176.1255 | -1.5974 | 1.271 | 32.1421 | 40.8538 | 25.2881 | 2.0087 | 32.0655 | 64.4106 | 15.9632 | 14.509 | -0.1101 | 105.2644 | 235.3117 | 181.1002 | 233.201 | 161.6165
2012-10-30 19:39 | 28.7162 | 249.8934 | 195.7058 | 247.6417 | 176.2411 | -1.6037 | 1.271 | 32.1418 | 40.8529 | 25.2881 | 2.0087 | 32.0652 | 64.4093 | 15.9632 | 14.5113 | -0.1106 | 105.2622 | 235.3821 | 181.1945 | 233.1304 | 161.7298
2012-10-30 19:39 | 28.7154 | 249.8261 | 195.6657 | 247.7193 | 176.3183 | -1.6009 | 1.271 | 32.1413 | 40.8519 | 25.288 | 2.0087 | 32.0648 | 64.4072 | 15.9633 | 14.4901 | -0.1104 | 105.2592 | 235.336 | 181.1756 | 233.2292 | 161.8282
2012-10-30 19:39 | 28.7079 | 249.846 | 195.7525 | 247.5864 | 176.3343 | -1.5985 | 1.271 | 32.1407 | 40.8502 | 25.2882 | 2.0086 | 32.0642 | 64.4051 | 15.9632 | 14.4622 | -0.1102 | 105.2552 | 235.3838 | 181.2903 | 233.1242 | 161.8721
2012-10-30 19:39 | 28.7162 | 249.9208 | 195.7569 | 247.6339 | 176.1435 | -1.5928 | 1.271 | 32.1404 | 40.8492 | 25.2883 | 2.0086 | 32.0639 | 64.4036 | 15.9633 | 14.499 | -0.1098 | 105.2528 | 235.4218 | 181.2579 | 233.1349 | 161.6445
2012-10-30 19:39 | 28.7177 | 249.8636 | 195.7427 | 247.6824 | 176.2386 | -1.5985 | 1.2686 | 32.0803 | 40.6971 | 25.288 | 2.005 | 32.0063 | 64.1728 | 15.9632 | 14.4842 | -0.1102 | 104.8699 | 235.3794 | 181.2585 | 233.1982 | 161.7544
2012-10-30 19:39 | 28.7169 | 249.8746 | 195.7419 | 247.6581 | 176.1871 | -1.6014 | 1.2649 | 31.9848 | 40.4573 | 25.2866 | 1.9989 | 31.9086 | 63.7816 | 15.9632 | 14.4997 | -0.1104 | 104.239 | 235.3749 | 181.2422 | 233.1584 | 161.6874
2012-10-30 19:39 | 28.7199 | 249.8384 | 195.737 | 247.6337 | 175.991 | -1.6009 | 1.2641 | 31.9658 | 40.4095 | 25.2865 | 1.9977 | 31.8898 | 63.7059 | 15.9633 | 14.4905 | -0.1104 | 104.1154 | 235.3479 | 181.2465 | 233.1432 | 161.5005
2012-10-30 19:39 | 28.7177 | 249.8088 | 195.6798 | 247.6002 | 175.8093 | -1.5991 | 1.2637 | 31.9552 | 40.383 | 25.2863 | 1.997 | 31.879 | 63.6627 | 15.9634 | 14.5128 | -0.1102 | 104.0458 | 235.296 | 181.167 | 233.0874 | 161.2965
2012-10-30 19:39 | 28.7199 | 249.8306 | 195.5995 | 247.618 | 175.788 | -1.5991 | 1.2636 | 31.9523 | 40.3758 | 25.2862 | 1.9968 | 31.8763 | 63.6517 | 15.9634 | 14.4252 | -0.1102 | 104.0275 | 235.4054 | 181.1743 | 233.1928 | 161.3628
2012-10-30 19:39 | 28.7289 | 249.773 | 195.6046 | 247.6427 | 175.8673 | -1.6003 | 1.2636 | 31.9513 | 40.3729 | 25.2864 | 1.9967 | 31.8751 | 63.6464 | 15.9636 | 14.5281 | -0.1103 | 104.0193 | 235.2449 | 181.0765 | 233.1146 | 161.3392
2012-10-30 19:39 | 28.7296 | 249.7816 | 195.5856 | 247.7022 | 175.626 | -1.6061 | 1.2636 | 31.9515 | 40.3737 | 25.2862 | 1.9968 | 31.8754 | 63.6473 | 15.9636 | 14.5165 | -0.1107 | 104.021 | 235.2651 | 181.0691 | 233.1857 | 161.1095
2012-10-30 19:39 | 28.7266 | 249.7512 | 195.6101 | 247.5856 | 175.7674 | -1.5991 | 1.2636 | 31.9521 | 40.3749 | 25.2864 | 1.9968 | 31.876 | 63.6498 | 15.9636 | 14.4931 | -0.1102 | 104.0248 | 235.2581 | 181.117 | 233.0925 | 161.2743
2012-10-30 19:39 | 28.7244 | 249.7529 | 195.5765 | 247.5638 | 175.976 | -1.598 | 1.2637 | 31.9541 | 40.3798 | 25.2864 | 1.9969 | 31.8779 | 63.6569 | 15.9637 | 14.4828 | -0.1102 | 104.0367 | 235.2701 | 181.0937 | 233.081 | 161.4932
2012-10-30 19:39 | 28.7251 | 249.7223 | 195.5811 | 247.5958 | 175.9181 | -1.598 | 1.265 | 31.9882 | 40.4661 | 25.2865 | 1.9988 | 31.909 | 63.7807 | 15.9638 | 14.5161 | -0.1102 | 104.2468 | 235.2062 | 181.065 | 233.0797 | 161.402
2012-10-30 19:39 | 28.7311 | 249.7283 | 195.5596 | 247.5901 | 175.6821 | -1.6003 | 1.2687 | 32.0823 | 40.7023 | 25.2879 | 2.0049 | 32.0057 | 64.1683 | 15.9637 | 14.5139 | -0.1103 | 104.8706 | 235.2144 | 181.0457 | 233.0762 | 161.1682
2012-10-30 19:39 | 28.7296 | 249.7111 | 195.5503 | 247.6434 | 175.4933 | -1.602 | 1.2694 | 32.0997 | 40.7464 | 25.2879 | 2.006 | 32.0233 | 64.2391 | 15.9636 | 14.5043 | -0.1104 | 104.9856 | 235.2068 | 181.046 | 233.1391 | 160.989
2012-10-30 19:40 | 28.7296 | 249.7385 | 195.5424 | 247.5415 | 175.7938 | -1.5997 | 1.2699 | 32.1144 | 40.7835 | 25.288 | 2.0069 | 32.0382 | 64.2987 | 15.9637 | 14.4961 | -0.1103 | 105.0822 | 235.2424 | 181.0463 | 233.0454 | 161.2977
2012-10-30 19:40 | 28.7274 | 249.7481 | 195.5284 | 247.5198 | 175.9048 | -1.6014 | 1.2704 | 32.1249 | 40.8103 | 25.288 | 2.0076 | 32.0484 | 64.3402 | 15.9635 | 14.498 | -0.1104 | 105.1505 | 235.2501 | 181.0304 | 233.0218 | 161.4068
2012-10-30 19:40 | 28.7266 | 249.712 | 195.5394 | 247.5307 | 175.7908 | -1.5997 | 1.2706 | 32.132 | 40.8279 | 25.2882 | 2.008 | 32.0555 | 64.3688 | 15.9636 | 14.5258 | -0.1103 | 105.1967 | 235.1862 | 181.0136 | 233.0049 | 161.265
2012-10-30 19:40 | 28.7221 | 249.7075 | 195.5192 | 247.5497 | 175.9425 | -1.6032 | 1.2708 | 32.1364 | 40.839 | 25.2883 | 2.0083 | 32.0599 | 64.3864 | 15.9636 | 14.5049 | -0.1105 | 105.2253 | 235.2026 | 181.0143 | 233.0448 | 161.4376
2012-10-30 19:40 | 28.7281 | 249.7018 | 195.5134 | 247.5949 | 176.2373 | -1.5997 | 1.2709 | 32.139 | 40.8457 | 25.2883 | 2.0085 | 32.0625 | 64.3973 | 15.9635 | 14.4865 | -0.1103 | 105.2431 | 235.2153 | 181.0269 | 233.1084 | 161.7508
2012-10-30 19:40 | 28.7266 | 249.7159 | 195.6455 | 247.5699 | 176.3022 | -1.5997 | 1.271 | 32.1407 | 40.8498 | 25.2883 | 2.0086 | 32.0641 | 64.403 | 15.9636 | 14.4809 | -0.1103 | 105.2529 | 235.235 | 181.1646 | 233.089 | 161.8213
2012-10-30 19:40 | 28.7341 | 249.7704 | 195.6373 | 247.5225 | 176.1575 | -1.6049 | 1.271 | 32.1413 | 40.8515 | 25.2883 | 2.0086 | 32.0648 | 64.4065 | 15.9634 | 14.5129 | -0.1106 | 105.2579 | 235.2575 | 181.1244 | 233.0096 | 161.6446
2012-10-30 19:40 | 28.7349 | 249.7477 | 195.7049 | 247.4685 | 175.9631 | -1.6014 | 1.271 | 32.1421 | 40.853 | 25.2885 | 2.0087 | 32.0654 | 64.4095 | 15.9634 | 14.4933 | -0.1104 | 105.2625 | 235.2544 | 181.2116 | 232.9752 | 161.4698
2012-10-30 19:40 | 28.7281 | 249.7527 | 195.6156 | 247.54 | 175.8119 | -1.6009 | 1.2711 | 32.1435 | 40.8573 | 25.2882 | 2.0088 | 32.0669 | 64.415 | 15.9635 | 14.462 | -0.1104 | 105.2723 | 235.2907 | 181.1536 | 233.078 | 161.3499
2012-10-30 19:40 | 28.7281 | 249.7213 | 195.6156 | 247.6458 | 175.8626 | -1.6032 | 1.2711 | 32.1436 | 40.857 | 25.2885 | 2.0088 | 32.067 | 64.4154 | 15.9634 | 14.5028 | -0.1105 | 105.2724 | 235.2185 | 181.1128 | 233.143 | 161.3598
2012-10-30 19:40 | 28.7311 | 249.7713 | 195.5871 | 247.3981 | 175.9202 | -1.6003 | 1.2711 | 32.1438 | 40.8575 | 25.2885 | 2.0088 | 32.0672 | 64.4166 | 15.9634 | 14.5588 | -0.1103 | 105.2741 | 235.2125 | 181.0283 | 232.8393 | 161.3614
2012-10-30 19:40 | 28.7349 | 249.7125 | 195.5949 | 247.5273 | 175.9357 | -1.6049 | 1.2711 | 32.1445 | 40.8596 | 25.2882 | 2.0088 | 32.0679 | 64.4193 | 15.9634 | 14.5137 | -0.1106 | 105.2789 | 235.1988 | 181.0812 | 233.0136 | 161.422
2012-10-30 19:40 | 28.7371 | 249.7538 | 195.5696 | 247.5412 | 175.7974 | -1.6014 | 1.2711 | 32.1445 | 40.8594 | 25.2885 | 2.0088 | 32.0679 | 64.4196 | 15.9633 | 14.5363 | -0.1104 | 105.279 | 235.2175 | 181.0333 | 233.0049 | 161.2611
2012-10-30 19:40 | 28.7416 | 249.6918 | 195.5819 | 247.6084 | 175.88 | -1.6037 | 1.2711 | 32.1443 | 40.8591 | 25.2883 | 2.0088 | 32.0677 | 64.4191 | 15.9633 | 14.5244 | -0.1106 | 105.2782 | 235.1674 | 181.0575 | 233.084 | 161.3556
2012-10-30 19:40 | 28.7386 | 249.7279 | 195.6143 | 247.4565 | 175.8809 | -1.5985 | 1.2711 | 32.1443 | 40.8585 | 25.2886 | 2.0088 | 32.0676 | 64.4184 | 15.9633 | 14.5092 | -0.1102 | 105.2769 | 235.2187 | 181.1051 | 232.9473 | 161.3717
2012-10-30 19:40 | 28.7461 | 249.7433 | 195.6021 | 247.5855 | 175.8767 | -1.6037 | 1.2711 | 32.1445 | 40.859 | 25.2886 | 2.0088 | 32.0678 | 64.4193 | 15.9633 | 14.5126 | -0.1106 | 105.2783 | 235.2307 | 181.0895 | 233.0729 | 161.3641
2012-10-30 19:40 | 28.7446 | 249.7104 | 195.581 | 247.5526 | 175.8284 | -1.5997 | 1.2711 | 32.1447 | 40.8597 | 25.2884 | 2.0089 | 32.0681 | 64.4207 | 15.9633 | 14.5111 | -0.1103 | 105.2804 | 235.1993 | 181.0699 | 233.0415 | 161.3173
2012-10-30 19:40 | 28.7491 | 249.7463 | 195.6051 | 247.6276 | 175.9265 | -1.6037 | 1.2711 | 32.1444 | 40.8592 | 25.2884 | 2.0089 | 32.0678 | 64.4194 | 15.9632 | 14.5115 | -0.1106 | 105.2786 | 235.2348 | 181.0936 | 233.1161 | 161.415
2012-10-30 19:40 | 28.7446 | 249.777 | 195.6006 | 247.5683 | 176.0508 | -1.5985 | 1.2711 | 32.1445 | 40.8596 | 25.2883 | 2.0089 | 32.0679 | 64.4202 | 15.9632 | 14.4866 | -0.1102 | 105.2798 | 235.2904 | 181.114 | 233.0817 | 161.5642
2012-10-30 19:40 | 28.7409 | 249.7146 | 195.6048 | 247.451 | 176.1057 | -1.5991 | 1.2711 | 32.1443 | 40.859 | 25.2884 | 2.0088 | 32.0677 | 64.419 | 15.9633 | 14.5074 | -0.1102 | 105.278 | 235.2072 | 181.0974 | 232.9436 | 161.5983
2012-10-30 19:40 | 28.7544 | 249.7437 | 195.6261 | 247.5781 | 176.1075 | -1.6049 | 1.2711 | 32.1439 | 40.8582 | 25.2883 | 2.0088 | 32.0674 | 64.4185 | 15.9631 | 14.5209 | -0.1106 | 105.2767 | 235.2228 | 181.1052 | 233.0572 | 161.5866
2012-10-30 19:40 | 28.7581 | 249.767 | 195.6731 | 247.6523 | 176.1112 | -1.5997 | 1.2711 | 32.1441 | 40.858 | 25.2886 | 2.0089 | 32.0674 | 64.4187 | 15.9631 | 14.5124 | -0.1103 | 105.2767 | 235.2546 | 181.1607 | 233.1399 | 161.5988
2012-10-30 19:41 | 28.7499 | 249.7392 | 195.6845 | 247.5893 | 176.1342 | -1.6055 | 1.2711 | 32.1438 | 40.8576 | 25.2884 | 2.0088 | 32.0672 | 64.4173 | 15.9632 | 14.4879 | -0.1107 | 105.2749 | 235.2513 | 181.1966 | 233.1014 | 161.6463
2012-10-30 19:41 | 28.7536 | 249.7273 | 195.6764 | 247.5616 | 176.0481 | -1.6009 | 1.2711 | 32.1439 | 40.8578 | 25.2884 | 2.0088 | 32.0673 | 64.4183 | 15.963 | 14.4916 | -0.1104 | 105.2761 | 235.2357 | 181.1848 | 233.07 | 161.5565
2012-10-30 19:41 | 28.7559 | 249.7491 | 195.663 | 247.5287 | 175.9411 | -1.5997 | 1.2711 | 32.1437 | 40.8575 | 25.2883 | 2.0088 | 32.0671 | 64.4174 | 15.9631 | 14.4898 | -0.1103 | 105.2749 | 235.2593 | 181.1732 | 233.0389 | 161.4513
2012-10-30 19:41 | 28.7544 | 249.7437 | 195.6065 | 247.6055 | 175.7601 | -1.5985 | 1.2711 | 32.144 | 40.8581 | 25.2884 | 2.0089 | 32.0673 | 64.4185 | 15.9629 | 14.5046 | -0.1102 | 105.2766 | 235.2391 | 181.1019 | 233.1009 | 161.2555
2012-10-30 19:41 | 28.7514 | 249.7603 | 195.6153 | 247.6652 | 175.921 | -1.5991 | 1.2711 | 32.1437 | 40.8574 | 25.2884 | 2.0088 | 32.0672 | 64.4177 | 15.9631 | 14.5057 | -0.1102 | 105.2751 | 235.2546 | 181.1096 | 233.1595 | 161.4153
2012-10-30 19:41 | 28.7506 | 249.7752 | 195.6145 | 247.59 | 175.9436 | -1.598 | 1.2711 | 32.1436 | 40.8575 | 25.2882 | 2.0088 | 32.0671 | 64.4174 | 15.963 | 14.4763 | -0.1102 | 105.2749 | 235.2989 | 181.1382 | 233.1137 | 161.4673
2012-10-30 19:41 | 28.7574 | 249.7663 | 195.6331 | 247.5302 | 175.806 | -1.6032 | 1.2711 | 32.1435 | 40.8567 | 25.2886 | 2.0088 | 32.0669 | 64.417 | 15.963 | 14.5076 | -0.1105 | 105.2737 | 235.2587 | 181.1255 | 233.0226 | 161.2984
2012-10-30 19:41 | 28.7603 | 249.7379 | 195.5849 | 247.6271 | 175.8714 | -1.6009 | 1.2711 | 32.1432 | 40.8562 | 25.2884 | 2.0088 | 32.0666 | 64.4159 | 15.9629 | 14.4983 | -0.1104 | 105.2721 | 235.2396 | 181.0866 | 233.1288 | 161.3731
2012-10-30 19:41 | 28.7514 | 249.7525 | 195.576 | 247.6221 | 175.8156 | -1.6026 | 1.2711 | 32.143 | 40.8557 | 25.2883 | 2.0088 | 32.0664 | 64.4147 | 15.963 | 14.4934 | -0.1105 | 105.2704 | 235.2591 | 181.0826 | 233.1287 | 161.3222
2012-10-30 19:41 | 28.7514 | 249.7564 | 195.5996 | 247.5986 | 175.7454 | -1.6061 | 1.2711 | 32.1431 | 40.8563 | 25.2882 | 2.0088 | 32.0665 | 64.4159 | 15.9629 | 14.5016 | -0.1107 | 105.2722 | 235.2548 | 181.098 | 233.097 | 161.2438
2012-10-30 19:41 | 28.7574 | 249.7585 | 195.5781 | 247.6359 | 175.8372 | -1.6037 | 1.271 | 32.1426 | 40.8547 | 25.2883 | 2.0088 | 32.066 | 64.4135 | 15.9629 | 14.4954 | -0.1106 | 105.2682 | 235.2631 | 181.0827 | 233.1405 | 161.3418
2012-10-30 19:41 | 28.7551 | 249.7758 | 195.5561 | 247.5631 | 175.9481 | -1.6009 | 1.271 | 32.1413 | 40.8516 | 25.2883 | 2.0087 | 32.0648 | 64.4089 | 15.9629 | 14.5012 | -0.1104 | 105.2605 | 235.2746 | 181.0549 | 233.0619 | 161.4469
2012-10-30 19:41 | 28.7603 | 249.7927 | 195.5692 | 247.5801 | 175.887 | -1.5997 | 1.271 | 32.1409 | 40.8505 | 25.2881 | 2.0087 | 32.0643 | 64.4068 | 15.9629 | 14.4901 | -0.1103 | 105.2574 | 235.3026 | 181.0791 | 233.09 | 161.3969
2012-10-30 19:41 | 28.7581 | 249.7396 | 195.5591 | 247.6053 | 175.8731 | -1.602 | 1.271 | 32.1406 | 40.8498 | 25.2882 | 2.0087 | 32.064 | 64.4057 | 15.9629 | 14.4838 | -0.1104 | 105.2554 | 235.2558 | 181.0753 | 233.1215 | 161.3893
2012-10-30 19:41 | 28.7551 | 249.7483 | 195.5836 | 247.712 | 175.9481 | -1.6003 | 1.271 | 32.1403 | 40.8493 | 25.288 | 2.0087 | 32.0638 | 64.405 | 15.9628 | 14.4482 | -0.1103 | 105.2543 | 235.3001 | 181.1354 | 233.2638 | 161.4999
2012-10-30 19:41 | 28.7491 | 249.7737 | 195.5973 | 247.6316 | 175.9421 | -1.5985 | 1.271 | 32.1398 | 40.8483 | 25.288 | 2.0086 | 32.0632 | 64.4029 | 15.9628 | 14.4993 | -0.1102 | 105.2512 | 235.2744 | 181.098 | 233.1323 | 161.4428
2012-10-30 19:41 | 28.7596 | 249.7959 | 195.5763 | 247.6068 | 175.9487 | -1.6049 | 1.2709 | 32.1397 | 40.8475 | 25.2882 | 2.0086 | 32.0631 | 64.4025 | 15.9628 | 14.5302 | -0.1106 | 105.25 | 235.2657 | 181.0461 | 233.0766 | 161.4185
2012-10-30 19:41 | 28.7588 | 249.799 | 195.5834 | 247.5903 | 176.0182 | -1.6032 | 1.2709 | 32.1397 | 40.8475 | 25.2883 | 2.0086 | 32.0631 | 64.4023 | 15.9629 | 14.4886 | -0.1105 | 105.2498 | 235.3104 | 181.0948 | 233.1017 | 161.5296
2012-10-30 19:41 | 28.7603 | 249.7575 | 195.5967 | 247.5957 | 175.8011 | -1.598 | 1.2709 | 32.1395 | 40.8473 | 25.288 | 2.0086 | 32.0628 | 64.4016 | 15.9627 | 14.4983 | -0.1102 | 105.2489 | 235.2592 | 181.0984 | 233.0974 | 161.3028
2012-10-30 19:41 | 28.7551 | 249.7836 | 195.5325 | 247.6219 | 175.8662 | -1.6032 | 1.2709 | 32.1391 | 40.8459 | 25.2883 | 2.0086 | 32.0626 | 64.4009 | 15.9627 | 14.4727 | -0.1105 | 105.2467 | 235.3109 | 181.0598 | 233.1492 | 161.3935
2012-10-30 19:41 | 28.7544 | 249.7829 | 195.5829 | 247.6212 | 175.9826 | -1.598 | 1.2709 | 32.1389 | 40.8452 | 25.2883 | 2.0086 | 32.0623 | 64.3994 | 15.9627 | 14.4638 | -0.1102 | 105.2446 | 235.3191 | 181.1191 | 233.1574 | 161.5188
2012-10-30 19:41 | 28.7559 | 249.7844 | 195.5766 | 247.6697 | 175.9177 | -1.5991 | 1.2709 | 32.1384 | 40.8438 | 25.2884 | 2.0085 | 32.0618 | 64.3971 | 15.9628 | 14.5061 | -0.1102 | 105.2409 | 235.2783 | 181.0705 | 233.1636 | 161.4116
2012-10-30 19:41 | 28.7544 | 249.779 | 195.6301 | 247.6212 | 176.0528 | -1.6014 | 1.2709 | 32.1379 | 40.8428 | 25.2882 | 2.0085 | 32.0614 | 64.3955 | 15.9628 | 14.4801 | -0.1104 | 105.2384 | 235.2989 | 181.15 | 233.1411 | 161.5727
2012-10-30 19:41 | 28.7566 | 249.8242 | 195.648 | 247.5881 | 176.0121 | -1.5945 | 1.2709 | 32.1376 | 40.8421 | 25.2883 | 2.0085 | 32.061 | 64.3941 | 15.9628 | 14.4905 | -0.1099 | 105.2362 | 235.3337 | 181.1575 | 233.0976 | 161.5216
2012-10-30 19:42 | 28.7559 | 249.7883 | 195.6119 | 247.5874 | 176.0895 | -1.6009 | 1.2708 | 32.1374 | 40.8415 | 25.2883 | 2.0085 | 32.0608 | 64.3931 | 15.9628 | 14.5183 | -0.1104 | 105.2346 | 235.27 | 181.0936 | 233.0691 | 161.5712
2012-10-30 19:42 | 28.7559 | 249.7844 | 195.6473 | 247.6227 | 175.9255 | -1.6043 | 1.2708 | 32.1363 | 40.8389 | 25.2883 | 2.0084 | 32.0597 | 64.3892 | 15.9627 | 14.5143 | -0.1106 | 105.2281 | 235.2701 | 181.133 | 233.1084 | 161.4112
2012-10-30 19:42 | 28.7574 | 249.7781 | 195.6095 | 247.6908 | 175.8646 | -1.6014 | 1.2653 | 31.997 | 40.4873 | 25.2872 | 2 | 31.9255 | 63.8507 | 15.9629 | 14.5117 | -0.1104 | 104.338 | 235.2664 | 181.0978 | 233.1791 | 161.3529
2012-10-30 19:42 | 28.7618 | 249.7472 | 195.6296 | 247.6678 | 175.7792 | -1.5997 | 1.2646 | 31.9771 | 40.4375 | 25.2867 | 1.9985 | 31.9012 | 63.7538 | 15.9628 | 14.512 | -0.1103 | 104.1913 | 235.2352 | 181.1176 | 233.1558 | 161.2672
2012-10-30 19:42 | 28.7529 | 249.7501 | 195.5343 | 247.6197 | 175.864 | -1.6026 | 1.2639 | 31.9607 | 40.3965 | 25.2865 | 1.9974 | 31.8845 | 63.6871 | 15.9628 | 14.4745 | -0.1105 | 104.0836 | 235.2756 | 181.0598 | 233.1452 | 161.3895
2012-10-30 19:42 | 28.7641 | 249.7378 | 195.5258 | 247.6348 | 175.9806 | -1.6049 | 1.2635 | 31.9502 | 40.3701 | 25.2865 | 1.9968 | 31.8741 | 63.6449 | 15.963 | 14.5388 | -0.1106 | 104.015 | 235.199 | 180.987 | 233.096 | 161.4418
2012-10-30 19:42 | 28.7551 | 249.7249 | 195.5168 | 247.5905 | 175.9872 | -1.6014 | 1.2634 | 31.9477 | 40.3638 | 25.2863 | 1.9966 | 31.8716 | 63.6351 | 15.9629 | 14.5175 | -0.1104 | 103.9989 | 235.2074 | 180.9993 | 233.073 | 161.4697
2012-10-30 19:42 | 28.7536 | 249.7038 | 195.5743 | 247.6517 | 176.2121 | -1.6009 | 1.2634 | 31.9482 | 40.3649 | 25.2865 | 1.9966 | 31.8721 | 63.6366 | 15.963 | 14.4997 | -0.1104 | 104.0014 | 235.2041 | 181.0746 | 233.152 | 161.7124
2012-10-30 19:42 | 28.7506 | 249.7321 | 195.6106 | 247.6605 | 176.3067 | -1.6032 | 1.2636 | 31.951 | 40.3724 | 25.2862 | 1.9968 | 31.8749 | 63.6468 | 15.9632 | 14.5171 | -0.1105 | 104.0192 | 235.215 | 181.0935 | 233.1434 | 161.7896
2012-10-30 19:42 | 28.7379 | 249.6724 | 195.6293 | 247.589 | 175.8724 | -1.6026 | 1.2636 | 31.9532 | 40.3775 | 25.2865 | 1.9969 | 31.8771 | 63.6557 | 15.9632 | 14.4963 | -0.1105 | 104.0332 | 235.1761 | 181.133 | 233.0927 | 161.3761
2012-10-30 19:42 | 28.7251 | 249.6518 | 195.5576 | 247.6232 | 175.8947 | -1.6043 | 1.2644 | 31.9732 | 40.4279 | 25.2867 | 1.9982 | 31.8979 | 63.7391 | 15.9632 | 14.4916 | -0.1106 | 104.167 | 235.1602 | 181.066 | 233.1316 | 161.4031
2012-10-30 19:42 | 28.7334 | 249.6249 | 195.6012 | 247.5375 | 175.5283 | -1.5974 | 1.2686 | 32.0801 | 40.6971 | 25.2876 | 2.0049 | 32.0039 | 64.1633 | 15.9631 | 14.4999 | -0.1101 | 104.8603 | 235.125 | 181.1013 | 233.0376 | 161.0284
2012-10-30 19:42 | 28.7326 | 249.6515 | 195.5179 | 247.5328 | 175.6251 | -1.6032 | 1.2692 | 32.0963 | 40.7382 | 25.2877 | 2.0059 | 32.0198 | 64.2273 | 15.9631 | 14.5032 | -0.1105 | 104.9655 | 235.1483 | 181.0147 | 233.0296 | 161.1219
2012-10-30 19:42 | 28.7169 | 249.6084 | 195.4708 | 247.6542 | 175.8631 | -1.6003 | 1.2699 | 32.1131 | 40.7801 | 25.2881 | 2.0069 | 32.0364 | 64.2943 | 15.9631 | 14.463 | -0.1103 | 105.0744 | 235.1454 | 181.0078 | 233.1912 | 161.4001
2012-10-30 19:42 | 28.7221 | 249.6645 | 195.5428 | 247.5262 | 175.9815 | -1.5997 | 1.2703 | 32.1238 | 40.8075 | 25.288 | 2.0076 | 32.0473 | 64.3377 | 15.9631 | 14.4682 | -0.1103 | 105.1452 | 235.1963 | 181.0746 | 233.058 | 161.5133
2012-10-30 19:42 | 28.7184 | 249.6803 | 195.5548 | 247.4951 | 176.0559 | -1.6055 | 1.2706 | 32.1302 | 40.8234 | 25.2882 | 2.008 | 32.0536 | 64.3637 | 15.9629 | 14.4768 | -0.1107 | 105.1871 | 235.2035 | 181.078 | 233.0183 | 161.5791
2012-10-30 19:42 | 28.7124 | 249.69 | 195.592 | 247.4813 | 176.0889 | -1.6026 | 1.2707 | 32.1345 | 40.8338 | 25.2884 | 2.0083 | 32.0579 | 64.3809 | 15.9629 | 14.4871 | -0.1105 | 105.2147 | 235.2029 | 181.1049 | 232.9942 | 161.6018
2012-10-30 19:42 | 28.7184 | 249.649 | 195.653 | 247.5264 | 176.0012 | -1.6009 | 1.2708 | 32.1376 | 40.8417 | 25.2884 | 2.0085 | 32.061 | 64.3932 | 15.963 | 14.4849 | -0.1104 | 105.235 | 235.1641 | 181.1681 | 233.0415 | 161.5163
2012-10-30 19:42 | 28.7169 | 249.6671 | 195.6279 | 247.5014 | 175.9607 | -1.5985 | 1.2709 | 32.1397 | 40.8477 | 25.2881 | 2.0086 | 32.0631 | 64.4023 | 15.9628 | 14.4753 | -0.1102 | 105.25 | 235.1918 | 181.1526 | 233.0261 | 161.4854
2012-10-30 19:42 | 28.7132 | 249.6477 | 195.6203 | 247.5447 | 175.8126 | -1.6026 | 1.271 | 32.1413 | 40.8513 | 25.2884 | 2.0087 | 32.0646 | 64.4084 | 15.9628 | 14.4389 | -0.1105 | 105.2596 | 235.2088 | 181.1814 | 233.1058 | 161.3737
2012-10-30 19:42 | 28.7117 | 249.6462 | 195.607 | 247.5119 | 175.9984 | -1.5985 | 1.271 | 32.1428 | 40.8549 | 25.2885 | 2.0088 | 32.0661 | 64.4148 | 15.9628 | 14.4497 | -0.1102 | 105.2697 | 235.1965 | 181.1573 | 233.0622 | 161.5487
2012-10-30 19:43 | 28.7109 | 249.7041 | 195.6102 | 247.4837 | 175.9859 | -1.6066 | 1.2711 | 32.1437 | 40.857 | 25.2886 | 2.0089 | 32.067 | 64.4178 | 15.9628 | 14.4937 | -0.1108 | 105.2749 | 235.2104 | 181.1165 | 232.99 | 161.4922
2012-10-30 19:43 | 28.7132 | 249.679 | 195.6007 | 247.4938 | 175.957 | -1.6014 | 1.2711 | 32.1438 | 40.8579 | 25.2883 | 2.0089 | 32.0672 | 64.4187 | 15.9628 | 14.4512 | -0.1104 | 105.2765 | 235.2278 | 181.1495 | 233.0426 | 161.5058
2012-10-30 19:43 | 28.7057 | 249.6637 | 195.6246 | 247.5137 | 175.7777 | -1.6037 | 1.2711 | 32.1441 | 40.8582 | 25.2885 | 2.0089 | 32.0675 | 64.4198 | 15.9628 | 14.4804 | -0.1106 | 105.278 | 235.1833 | 181.1442 | 233.0333 | 161.2973
2012-10-30 19:43 | 28.7079 | 249.6268 | 195.5796 | 247.5551 | 175.8346 | -1.5991 | 1.2711 | 32.1444 | 40.8592 | 25.2884 | 2.0089 | 32.068 | 64.4221 | 15.9627 | 14.4785 | -0.1102 | 105.2814 | 235.1483 | 181.1011 | 233.0766 | 161.3561
2012-10-30 19:43 | 28.7132 | 249.6947 | 195.5732 | 247.5251 | 175.8165 | -1.6037 | 1.2711 | 32.1447 | 40.8597 | 25.2885 | 2.0089 | 32.068 | 64.4223 | 15.9628 | 14.496 | -0.1106 | 105.2819 | 235.1987 | 181.0772 | 233.0291 | 161.3205
2012-10-30 19:43 | 28.7192 | 249.6929 | 195.5281 | 247.4254 | 175.8303 | -1.6003 | 1.2711 | 32.145 | 40.8609 | 25.2883 | 2.009 | 32.0685 | 64.424 | 15.9628 | 14.4694 | -0.1103 | 105.2849 | 235.2235 | 181.0587 | 232.956 | 161.3609
2012-10-30 19:43 | 28.7154 | 249.6734 | 195.5243 | 247.5234 | 175.7211 | -1.6009 | 1.2711 | 32.1452 | 40.8609 | 25.2886 | 2.009 | 32.0686 | 64.425 | 15.9627 | 14.4697 | -0.1104 | 105.2859 | 235.2037 | 181.0546 | 233.0537 | 161.2514
2012-10-30 19:43 | 28.7221 | 249.6879 | 195.4642 | 247.5889 | 175.638 | -1.6014 | 1.2711 | 32.1453 | 40.8613 | 25.2885 | 2.009 | 32.0687 | 64.4253 | 15.9627 | 14.5009 | -0.1104 | 105.2866 | 235.187 | 180.9633 | 233.088 | 161.1371
2012-10-30 19:43 | 28.7199 | 249.674 | 195.5131 | 247.5044 | 175.8349 | -1.602 | 1.2711 | 32.1449 | 40.8604 | 25.2884 | 2.009 | 32.0685 | 64.4244 | 15.9627 | 14.4946 | -0.1104 | 105.2848 | 235.1794 | 181.0185 | 233.0098 | 161.3403
2012-10-30 19:43 | 28.7214 | 249.699 | 195.4988 | 247.4667 | 175.9652 | -1.5991 | 1.2711 | 32.1451 | 40.8612 | 25.2882 | 2.009 | 32.0685 | 64.4249 | 15.9626 | 14.5002 | -0.1102 | 105.2861 | 235.1988 | 180.9986 | 232.9665 | 161.465
2012-10-30 19:43 | 28.7236 | 249.6816 | 195.4735 | 247.5277 | 175.7488 | -1.6009 | 1.2711 | 32.1448 | 40.8602 | 25.2884 | 2.009 | 32.0682 | 64.4235 | 15.9627 | 14.486 | -0.1104 | 105.2838 | 235.1956 | 180.9875 | 233.0417 | 161.2628
2012-10-30 19:43 | 28.7244 | 249.6198 | 195.4586 | 247.6069 | 175.7925 | -1.6014 | 1.2712 | 32.1451 | 40.8615 | 25.2881 | 2.009 | 32.0685 | 64.4254 | 15.9625 | 14.4256 | -0.1104 | 105.2869 | 235.1942 | 181.033 | 233.1813 | 161.3669
2012-10-30 19:43 | 28.7274 | 249.6854 | 195.5323 | 247.5315 | 175.8072 | -1.6026 | 1.2712 | 32.1454 | 40.8617 | 25.2884 | 2.009 | 32.0687 | 64.4257 | 15.9626 | 14.4694 | -0.1105 | 105.2874 | 235.216 | 181.0629 | 233.0621 | 161.3378
2012-10-30 19:43 | 28.7177 | 249.6522 | 195.4362 | 247.4395 | 175.7351 | -1.6009 | 1.2711 | 32.1453 | 40.8615 | 25.2884 | 2.009 | 32.0687 | 64.426 | 15.9625 | 14.4924 | -0.1104 | 105.2876 | 235.1598 | 180.9438 | 232.9471 | 161.2427
2012-10-30 19:43 | 28.7139 | 249.7111 | 195.4717 | 247.5807 | 175.7703 | -1.5991 | 1.2711 | 32.1452 | 40.8611 | 25.2884 | 2.009 | 32.0686 | 64.4253 | 15.9626 | 14.4763 | -0.1102 | 105.2864 | 235.2348 | 180.9954 | 233.1044 | 161.294
2012-10-30 19:43 | 28.7154 | 249.693 | 195.485 | 247.5234 | 175.7679 | -1.5985 | 1.2711 | 32.1448 | 40.86 | 25.2885 | 2.009 | 32.068 | 64.4239 | 15.9624 | 14.4656 | -0.1102 | 105.284 | 235.2274 | 181.0194 | 233.0578 | 161.3023
2012-10-30 19:43 | 28.7117 | 249.6775 | 195.4695 | 247.4531 | 175.7057 | -1.5985 | 1.2711 | 32.1444 | 40.8593 | 25.2883 | 2.0089 | 32.0678 | 64.4223 | 15.9626 | 14.4129 | -0.1102 | 105.2816 | 235.2646 | 181.0566 | 233.0402 | 161.2928
2012-10-30 19:43 | 28.7087 | 249.7019 | 195.4665 | 247.6068 | 175.7964 | -1.5968 | 1.2711 | 32.1442 | 40.8588 | 25.2882 | 2.0089 | 32.0675 | 64.4212 | 15.9625 | 14.4875 | -0.1101 | 105.28 | 235.2144 | 180.979 | 233.1193 | 161.3089
2012-10-30 19:43 | 28.7192 | 249.6772 | 195.4966 | 247.5821 | 175.9552 | -1.6009 | 1.2711 | 32.1438 | 40.8581 | 25.2881 | 2.0089 | 32.0672 | 64.4205 | 15.9624 | 14.4776 | -0.1104 | 105.2786 | 235.1996 | 181.019 | 233.1045 | 161.4776
2012-10-30 19:43 | 28.7229 | 249.6927 | 195.4689 | 247.4878 | 175.834 | -1.602 | 1.2711 | 32.1437 | 40.8572 | 25.2885 | 2.0089 | 32.067 | 64.4194 | 15.9625 | 14.4935 | -0.1104 | 105.2766 | 235.1992 | 180.9754 | 232.9943 | 161.3405
2012-10-30 19:43 | 28.7177 | 249.6992 | 195.4991 | 247.608 | 175.7663 | -1.6043 | 1.2711 | 32.1435 | 40.8572 | 25.2881 | 2.0089 | 32.0668 | 64.4191 | 15.9624 | 14.4557 | -0.1106 | 105.2763 | 235.2435 | 181.0434 | 233.1523 | 161.3106
2012-10-30 19:43 | 28.7139 | 249.7111 | 195.4992 | 247.6434 | 175.7742 | -1.6014 | 1.2711 | 32.1432 | 40.8562 | 25.2883 | 2.0089 | 32.0665 | 64.418 | 15.9623 | 14.4559 | -0.1104 | 105.2742 | 235.2552 | 181.0433 | 233.1875 | 161.3183
2012-10-30 19:43 | 28.7154 | 249.7713 | 195.4535 | 247.5117 | 175.7796 | -1.5945 | 1.2711 | 32.1428 | 40.8551 | 25.2883 | 2.0089 | 32.0662 | 64.4169 | 15.9623 | 14.4534 | -0.1099 | 105.272 | 235.3179 | 181.0001 | 233.0583 | 161.3262
2012-10-30 19:43 | 28.7192 | 249.7085 | 195.4141 | 247.6134 | 175.7561 | -1.6003 | 1.2711 | 32.1427 | 40.8552 | 25.2882 | 2.0089 | 32.0661 | 64.4165 | 15.9623 | 14.4776 | -0.1103 | 105.2717 | 235.2309 | 180.9365 | 233.1358 | 161.2785
2012-10-30 19:43 | 28.7049 | 249.6903 | 195.4548 | 247.6109 | 175.8277 | -1.5945 | 1.271 | 32.1424 | 40.8545 | 25.2882 | 2.0088 | 32.0659 | 64.4155 | 15.9624 | 14.451 | -0.1099 | 105.27 | 235.2393 | 181.0038 | 233.1599 | 161.3767
2012-10-30 19:44 | 28.7162 | 249.7447 | 195.4936 | 247.5321 | 175.6009 | -1.5968 | 1.271 | 32.1424 | 40.854 | 25.2884 | 2.0088 | 32.0657 | 64.4151 | 15.9622 | 14.5072 | -0.1101 | 105.2691 | 235.2375 | 180.9864 | 233.0249 | 161.0937
2012-10-30 19:44 | 28.7192 | 249.6929 | 195.4141 | 247.5507 | 175.4868 | -1.5974 | 1.271 | 32.142 | 40.8538 | 25.288 | 2.0088 | 32.0654 | 64.4139 | 15.9623 | 14.4653 | -0.1101 | 105.2676 | 235.2276 | 180.9488 | 233.0854 | 161.0215
2012-10-30 19:44 | 28.7199 | 249.768 | 195.3952 | 247.618 | 175.5578 | -1.6032 | 1.271 | 32.1417 | 40.8524 | 25.2883 | 2.0088 | 32.0651 | 64.4122 | 15.9623 | 14.4701 | -0.1105 | 105.2647 | 235.2979 | 180.9251 | 233.1479 | 161.0877
2012-10-30 19:44 | 28.7102 | 249.7426 | 195.3619 | 247.5927 | 175.6808 | -1.5945 | 1.271 | 32.1415 | 40.8521 | 25.2882 | 2.0088 | 32.0648 | 64.4115 | 15.9622 | 14.4522 | -0.1099 | 105.2636 | 235.2904 | 180.9097 | 233.1405 | 161.2286
2012-10-30 19:44 | 28.7244 | 249.7372 | 195.38 | 247.552 | 175.6833 | -1.6014 | 1.271 | 32.1412 | 40.8516 | 25.288 | 2.0088 | 32.0646 | 64.4104 | 15.9623 | 14.4379 | -0.1104 | 105.262 | 235.2993 | 180.9421 | 233.1141 | 161.2454
2012-10-30 19:44 | 28.7236 | 249.7443 | 195.4028 | 247.6335 | 175.6981 | -1.6032 | 1.271 | 32.1407 | 40.8498 | 25.2883 | 2.0087 | 32.0641 | 64.4084 | 15.9623 | 14.4575 | -0.1105 | 105.2582 | 235.2868 | 180.9453 | 233.176 | 161.2406
2012-10-30 19:44 | 28.7169 | 249.765 | 195.4197 | 247.521 | 175.8241 | -1.6009 | 1.271 | 32.1406 | 40.8504 | 25.2878 | 2.0087 | 32.064 | 64.4083 | 15.9623 | 14.5242 | -0.1104 | 105.2588 | 235.2408 | 180.8955 | 232.9968 | 161.2999
2012-10-30 19:44 | 28.7192 | 249.7633 | 195.3906 | 247.5899 | 175.7366 | -1.6014 | 1.271 | 32.1406 | 40.8497 | 25.2883 | 2.0087 | 32.0639 | 64.4081 | 15.9621 | 14.4816 | -0.1104 | 105.2577 | 235.2817 | 180.909 | 233.1083 | 161.255
2012-10-30 19:44 | 28.7117 | 249.7676 | 195.4616 | 247.6686 | 175.7837 | -1.6003 | 1.271 | 32.1404 | 40.8497 | 25.288 | 2.0087 | 32.0639 | 64.4081 | 15.9621 | 14.4456 | -0.1103 | 105.2578 | 235.322 | 181.016 | 233.223 | 161.3381
2012-10-30 19:44 | 28.7139 | 249.7541 | 195.4953 | 247.5376 | 175.9967 | -1.6061 | 1.271 | 32.1404 | 40.8492 | 25.2882 | 2.0087 | 32.0637 | 64.4078 | 15.962 | 14.4845 | -0.1107 | 105.2571 | 235.2696 | 181.0108 | 233.0531 | 161.5122
2012-10-30 19:44 | 28.7154 | 249.7674 | 195.4968 | 247.5744 | 175.9241 | -1.6026 | 1.271 | 32.1402 | 40.8489 | 25.2881 | 2.0087 | 32.0635 | 64.4073 | 15.962 | 14.5472 | -0.1105 | 105.2562 | 235.2202 | 180.9496 | 233.0272 | 161.3769
2012-10-30 19:44 | 28.7147 | 249.7471 | 195.5079 | 247.7108 | 175.8882 | -1.6032 | 1.271 | 32.14 | 40.8485 | 25.2881 | 2.0087 | 32.0635 | 64.4068 | 15.9621 | 14.5261 | -0.1105 | 105.2553 | 235.221 | 180.9818 | 233.1847 | 161.3621
2012-10-30 19:44 | 28.7124 | 249.7448 | 195.5527 | 247.6379 | 176.0538 | -1.5951 | 1.2709 | 32.1397 | 40.8477 | 25.288 | 2.0087 | 32.0631 | 64.4056 | 15.962 | 14.5238 | -0.11 | 105.2533 | 235.221 | 181.0289 | 233.1141 | 161.53
2012-10-30 19:44 | 28.7064 | 249.7858 | 195.5192 | 247.534 | 176.0322 | -1.6037 | 1.2709 | 32.1393 | 40.8462 | 25.2883 | 2.0087 | 32.0626 | 64.404 | 15.9619 | 14.5953 | -0.1106 | 105.2502 | 235.1905 | 180.9239 | 232.9387 | 161.4369
2012-10-30 19:44 | 28.7049 | 249.7569 | 195.5374 | 247.6618 | 175.9097 | -1.598 | 1.2709 | 32.1391 | 40.8459 | 25.2882 | 2.0087 | 32.0625 | 64.4029 | 15.9621 | 14.5408 | -0.1102 | 105.2487 | 235.2161 | 180.9966 | 233.121 | 161.3689
2012-10-30 19:44 | 28.7049 | 249.7647 | 195.6277 | 247.6657 | 176.0229 | -1.6072 | 1.2709 | 32.1393 | 40.8466 | 25.2881 | 2.0087 | 32.0626 | 64.4033 | 15.9621 | 14.5408 | -0.1108 | 105.2499 | 235.2239 | 181.0869 | 233.1249 | 161.4821
2012-10-30 19:44 | 28.6989 | 249.7861 | 195.6335 | 247.5657 | 176.0481 | -1.5991 | 1.2653 | 31.9962 | 40.4848 | 25.2874 | 2.0001 | 31.9251 | 63.8519 | 15.9621 | 14.5552 | -0.1102 | 104.3367 | 235.2309 | 181.0783 | 233.0105 | 161.4929
2012-10-30 19:44 | 28.7034 | 249.7671 | 195.5594 | 247.6133 | 175.7481 | -1.5991 | 1.2648 | 31.9823 | 40.4511 | 25.2866 | 1.9989 | 31.9061 | 63.7758 | 15.9621 | 14.5393 | -0.1102 | 104.2268 | 235.2278 | 181.0201 | 233.074 | 161.2088
2012-10-30 19:44 | 28.6824 | 249.7187 | 195.5109 | 247.655 | 175.7818 | -1.6032 | 1.2642 | 31.9672 | 40.4129 | 25.2865 | 1.9979 | 31.891 | 63.7152 | 15.9622 | 14.5509 | -0.1105 | 104.1281 | 235.1678 | 180.96 | 233.1041 | 161.2309
2012-10-30 19:44 | 28.6967 | 249.6939 | 195.5645 | 247.634 | 176.0068 | -1.598 | 1.2637 | 31.9555 | 40.3834 | 25.2864 | 1.9972 | 31.8795 | 63.669 | 15.9623 | 14.5203 | -0.1102 | 104.0524 | 235.1736 | 181.0442 | 233.1137 | 161.4865
2012-10-30 19:44 | 28.6974 | 249.7181 | 195.5574 | 247.5446 | 176.0349 | -1.602 | 1.2637 | 31.9532 | 40.3783 | 25.2861 | 1.997 | 31.8771 | 63.6592 | 15.9623 | 14.5088 | -0.1104 | 104.0375 | 235.2093 | 181.0486 | 233.0358 | 161.5261
2012-10-30 19:44 | 28.681 | 249.639 | 195.537 | 247.6065 | 175.8311 | -1.6049 | 1.2637 | 31.9532 | 40.3779 | 25.2863 | 1.997 | 31.8771 | 63.6592 | 15.9624 | 14.525 | -0.1106 | 104.0371 | 235.114 | 181.012 | 233.0815 | 161.3061
2012-10-30 19:44 | 28.6869 | 249.6253 | 195.5429 | 247.683 | 175.9268 | -1.6055 | 1.2636 | 31.9523 | 40.3755 | 25.2864 | 1.9969 | 31.8761 | 63.6548 | 15.9624 | 14.5228 | -0.1107 | 104.0304 | 235.1025 | 181.0201 | 233.1602 | 161.404
2012-10-30 19:44 | 28.6884 | 249.6973 | 195.5641 | 247.5395 | 176.1196 | -1.6032 | 1.2637 | 31.9536 | 40.3793 | 25.2861 | 1.997 | 31.8774 | 63.6597 | 15.9625 | 14.5202 | -0.1105 | 104.039 | 235.1771 | 181.0439 | 233.0193 | 161.5994
2012-10-30 19:44 | 28.6974 | 249.675 | 195.5849 | 247.5838 | 176.1325 | -1.6032 | 1.2637 | 31.9551 | 40.3825 | 25.2864 | 1.9971 | 31.8789 | 63.666 | 15.9625 | 14.5496 | -0.1105 | 104.0484 | 235.1254 | 181.0353 | 233.0342 | 161.5829
2012-10-30 19:44 | 28.6854 | 249.6317 | 195.6161 | 247.5874 | 176.1439 | -1.5997 | 1.2638 | 31.9559 | 40.3845 | 25.2865 | 1.9972 | 31.8798 | 63.6692 | 15.9626 | 14.5784 | -0.1103 | 104.0537 | 235.0533 | 181.0377 | 233.009 | 161.5655
2012-10-30 19:45 | 28.6854 | 249.6395 | 195.6672 | 247.7402 | 176.2649 | -1.6014 | 1.2667 | 32.0302 | 40.5717 | 25.287 | 2.0019 | 31.9562 | 63.9744 | 15.9626 | 14.5131 | -0.1104 | 104.5461 | 235.1264 | 181.1541 | 233.2271 | 161.7518
2012-10-30 19:45 | 28.678 | 249.6556 | 195.6676 | 247.5017 | 176.1989 | -1.5968 | 1.2689 | 32.089 | 40.7193 | 25.288 | 2.0055 | 32.0124 | 64.2006 | 15.9624 | 14.522 | -0.1101 | 104.9199 | 235.1336 | 181.1456 | 232.9797 | 161.6769
2012-10-30 19:45 | 28.6787 | 249.6954 | 195.6605 | 247.5847 | 176.1255 | -1.6049 | 1.2695 | 32.1037 | 40.7571 | 25.2876 | 2.0064 | 32.0269 | 64.2585 | 15.9624 | 14.5064 | -0.1106 | 105.0155 | 235.189 | 181.1541 | 233.0783 | 161.6191
2012-10-30 19:45 | 28.6839 | 249.6458 | 195.7089 | 247.6212 | 176.2127 | -1.5985 | 1.2701 | 32.1171 | 40.7904 | 25.2881 | 2.0072 | 32.0403 | 64.3128 | 15.9623 | 14.5402 | -0.1102 | 105.1032 | 235.1056 | 181.1687 | 233.081 | 161.6725
2012-10-30 19:45 | 28.6765 | 249.7011 | 195.7447 | 247.461 | 176.3458 | -1.5974 | 1.2704 | 32.1269 | 40.8147 | 25.2883 | 2.0079 | 32.0502 | 64.3525 | 15.9623 | 14.5083 | -0.1101 | 105.1671 | 235.1928 | 181.2364 | 232.9527 | 161.8375
2012-10-30 19:45 | 28.6712 | 249.6684 | 195.7237 | 247.4518 | 176.2429 | -1.6009 | 1.2707 | 32.1328 | 40.8296 | 25.2884 | 2.0082 | 32.056 | 64.3761 | 15.9623 | 14.5071 | -0.1104 | 105.2057 | 235.1613 | 181.2166 | 232.9447 | 161.7358
2012-10-30 19:45 | 28.6742 | 249.687 | 195.7542 | 247.6311 | 176.1249 | -1.5968 | 1.2708 | 32.1372 | 40.8409 | 25.2883 | 2.0085 | 32.0603 | 64.393 | 15.9623 | 14.5182 | -0.1101 | 105.2339 | 235.1688 | 181.236 | 233.1129 | 161.6067
2012-10-30 19:45 | 28.6637 | 249.6609 | 195.7516 | 247.4286 | 176.2901 | -1.6009 | 1.2709 | 32.14 | 40.8479 | 25.2884 | 2.0087 | 32.0633 | 64.4053 | 15.9623 | 14.5241 | -0.1104 | 105.2532 | 235.1368 | 181.2275 | 232.9045 | 161.766
2012-10-30 19:45 | 28.6645 | 249.7126 | 195.7602 | 247.5195 | 176.1581 | -1.5985 | 1.271 | 32.1424 | 40.8543 | 25.2883 | 2.0088 | 32.0658 | 64.4154 | 15.9623 | 14.5412 | -0.1102 | 105.2697 | 235.1714 | 181.219 | 232.9783 | 161.6169
2012-10-30 19:45 | 28.657 | 249.6855 | 195.7174 | 247.5669 | 176.0882 | -1.6043 | 1.2711 | 32.1442 | 40.8581 | 25.2888 | 2.009 | 32.0676 | 64.4222 | 15.9623 | 14.5337 | -0.1106 | 105.2803 | 235.1518 | 181.1837 | 233.0332 | 161.5545
2012-10-30 19:45 | 28.654 | 249.6668 | 195.7654 | 247.5795 | 176.1749 | -1.5962 | 1.2711 | 32.1453 | 40.8615 | 25.2884 | 2.009 | 32.0687 | 64.427 | 15.9622 | 14.5388 | -0.11 | 105.2885 | 235.128 | 181.2266 | 233.0407 | 161.6361
2012-10-30 19:45 | 28.651 | 249.703 | 195.7349 | 247.4512 | 176.1251 | -1.602 | 1.2712 | 32.146 | 40.8632 | 25.2884 | 2.0091 | 32.0693 | 64.4301 | 15.9621 | 14.4828 | -0.1104 | 105.2933 | 235.2202 | 181.2521 | 232.9684 | 161.6423
2012-10-30 19:45 | 28.6607 | 249.7244 | 195.7604 | 247.5392 | 176.0372 | -1.6032 | 1.2712 | 32.1469 | 40.8653 | 25.2885 | 2.0091 | 32.0702 | 64.4336 | 15.9621 | 14.5211 | -0.1105 | 105.2989 | 235.2033 | 181.2393 | 233.0181 | 161.5161
2012-10-30 19:45 | 28.6562 | 249.7512 | 195.7244 | 247.617 | 176.0366 | -1.5997 | 1.2712 | 32.1471 | 40.8655 | 25.2887 | 2.0092 | 32.0704 | 64.4344 | 15.9622 | 14.4717 | -0.1103 | 105.2999 | 235.2795 | 181.2527 | 233.1453 | 161.5649
2012-10-30 19:45 | 28.669 | 249.7484 | 195.749 | 247.524 | 176.2329 | -1.6049 | 1.2712 | 32.1472 | 40.866 | 25.2885 | 2.0092 | 32.0706 | 64.4349 | 15.9622 | 14.4886 | -0.1106 | 105.3008 | 235.2598 | 181.2604 | 233.0354 | 161.7443
2012-10-30 19:45 | 28.6562 | 249.7199 | 195.7205 | 247.4603 | 176.142 | -1.5985 | 1.2712 | 32.1474 | 40.8668 | 25.2883 | 2.0092 | 32.0707 | 64.4355 | 15.9621 | 14.5002 | -0.1102 | 105.3024 | 235.2197 | 181.2203 | 232.9601 | 161.6418
2012-10-30 19:45 | 28.66 | 249.7198 | 195.7361 | 247.6326 | 176.1029 | -1.598 | 1.2712 | 32.1476 | 40.8671 | 25.2885 | 2.0092 | 32.071 | 64.4369 | 15.9621 | 14.4632 | -0.1102 | 105.304 | 235.2566 | 181.2729 | 233.1694 | 161.6397
2012-10-30 19:45 | 28.6652 | 249.7407 | 195.7177 | 247.5947 | 176.1432 | -1.6003 | 1.2712 | 32.148 | 40.8681 | 25.2886 | 2.0092 | 32.0714 | 64.4386 | 15.9621 | 14.5011 | -0.1103 | 105.3067 | 235.2396 | 181.2166 | 233.0936 | 161.6421
2012-10-30 19:45 | 28.6637 | 249.7157 | 195.728 | 247.4443 | 176.1573 | -1.5962 | 1.2713 | 32.1486 | 40.8695 | 25.2886 | 2.0093 | 32.0718 | 64.4407 | 15.962 | 14.4506 | -0.11 | 105.3102 | 235.2651 | 181.2774 | 232.9937 | 161.7067
2012-10-30 19:45 | 28.6607 | 249.7205 | 195.6582 | 247.6098 | 175.8498 | -1.5997 | 1.2713 | 32.1484 | 40.8693 | 25.2885 | 2.0093 | 32.0718 | 64.4402 | 15.962 | 14.4558 | -0.1103 | 105.3095 | 235.2647 | 181.2024 | 233.154 | 161.394
2012-10-30 19:45 | 28.6652 | 249.7367 | 195.6745 | 247.6182 | 175.8817 | -1.6009 | 1.2713 | 32.1484 | 40.8695 | 25.2883 | 2.0093 | 32.0716 | 64.4402 | 15.9619 | 14.4684 | -0.1104 | 105.3097 | 235.2683 | 181.2061 | 233.1498 | 161.4133
2012-10-30 19:45 | 28.6682 | 249.7554 | 195.7246 | 247.5428 | 175.9315 | -1.6061 | 1.2713 | 32.1486 | 40.8699 | 25.2883 | 2.0093 | 32.0719 | 64.4408 | 15.962 | 14.4633 | -0.1107 | 105.3107 | 235.2921 | 181.2613 | 233.0795 | 161.4682
2012-10-30 19:45 | 28.6645 | 249.7634 | 195.6581 | 247.5587 | 175.9044 | -1.6014 | 1.2713 | 32.1484 | 40.8693 | 25.2885 | 2.0093 | 32.0718 | 64.4409 | 15.9619 | 14.48 | -0.1104 | 105.3102 | 235.2834 | 181.1781 | 233.0787 | 161.4244
2012-10-30 19:45 | 28.6705 | 249.7734 | 195.609 | 247.6196 | 175.6645 | -1.6003 | 1.2713 | 32.1483 | 40.869 | 25.2884 | 2.0092 | 32.0716 | 64.4395 | 15.9621 | 14.5023 | -0.1103 | 105.3085 | 235.2711 | 181.1067 | 233.1173 | 161.1622
2012-10-30 19:45 | 28.6705 | 249.7734 | 195.6208 | 247.6431 | 175.9026 | -1.5997 | 1.2712 | 32.1479 | 40.8674 | 25.2888 | 2.0092 | 32.0712 | 64.4382 | 15.962 | 14.4778 | -0.1103 | 105.3055 | 235.2956 | 181.143 | 233.1653 | 161.4248
2012-10-30 19:46 | 28.669 | 249.8149 | 195.5997 | 247.4966 | 175.9752 | -1.6009 | 1.2712 | 32.1475 | 40.8672 | 25.2883 | 2.0092 | 32.0709 | 64.437 | 15.9619 | 14.4478 | -0.1104 | 105.3041 | 235.3671 | 181.1519 | 233.0488 | 161.5274
2012-10-30 19:46 | 28.6682 | 249.7515 | 195.5871 | 247.6251 | 175.7871 | -1.6049 | 1.2712 | 32.1474 | 40.8669 | 25.2882 | 2.0092 | 32.0707 | 64.4368 | 15.9618 | 14.4755 | -0.1106 | 105.3037 | 235.276 | 181.1116 | 233.1496 | 161.3116
2012-10-30 19:46 | 28.6757 | 249.7864 | 195.5396 | 247.6248 | 175.7477 | -1.598 | 1.2712 | 32.1471 | 40.8663 | 25.2882 | 2.0092 | 32.0704 | 64.4351 | 15.962 | 14.4667 | -0.1102 | 105.3014 | 235.3197 | 181.0729 | 233.1581 | 161.281
2012-10-30 19:46 | 28.6795 | 249.7706 | 195.6141 | 247.5306 | 175.7867 | -1.602 | 1.2712 | 32.1474 | 40.8667 | 25.2885 | 2.0092 | 32.0708 | 64.437 | 15.9619 | 14.4419 | -0.1104 | 105.3037 | 235.3287 | 181.1722 | 233.0887 | 161.3448
2012-10-30 19:46 | 28.6682 | 249.7671 | 195.5124 | 247.6133 | 175.6895 | -1.602 | 1.2712 | 32.1474 | 40.867 | 25.2882 | 2.0092 | 32.0707 | 64.4371 | 15.9618 | 14.4388 | -0.1104 | 105.3041 | 235.3283 | 181.0736 | 233.1745 | 161.2507
2012-10-30 19:46 | 28.6795 | 249.7706 | 195.5355 | 247.6795 | 175.8335 | -1.6049 | 1.2712 | 32.1472 | 40.8661 | 25.2885 | 2.0092 | 32.0706 | 64.4367 | 15.9617 | 14.4787 | -0.1106 | 105.3028 | 235.2919 | 181.0568 | 233.2008 | 161.3548
2012-10-30 19:46 | 28.6802 | 249.7987 | 195.548 | 247.6097 | 175.7561 | -1.5974 | 1.2712 | 32.1466 | 40.8649 | 25.2884 | 2.0092 | 32.07 | 64.4344 | 15.9618 | 14.4834 | -0.1101 | 105.2993 | 235.3153 | 181.0646 | 233.1263 | 161.2727
2012-10-30 19:46 | 28.678 | 249.8004 | 195.5458 | 247.5487 | 175.8632 | -1.6043 | 1.2712 | 32.1463 | 40.8637 | 25.2885 | 2.0092 | 32.0695 | 64.433 | 15.9616 | 14.4853 | -0.1106 | 105.2967 | 235.3151 | 181.0605 | 233.0634 | 161.3779
2012-10-30 19:46 | 28.6817 | 249.7924 | 195.522 | 247.6072 | 175.8435 | -1.598 | 1.2712 | 32.1457 | 40.8627 | 25.2883 | 2.0091 | 32.0691 | 64.4312 | 15.9616 | 14.5502 | -0.1102 | 105.2939 | 235.2422 | 180.9718 | 233.057 | 161.2933
2012-10-30 19:46 | 28.6847 | 249.7915 | 195.5329 | 247.6808 | 175.8699 | -1.6003 | 1.2712 | 32.1456 | 40.8621 | 25.2884 | 2.0091 | 32.069 | 64.4307 | 15.9617 | 14.4879 | -0.1103 | 105.2928 | 235.3036 | 181.045 | 233.1929 | 161.382
2012-10-30 19:46 | 28.6854 | 249.7922 | 195.5611 | 247.6384 | 175.8316 | -1.6014 | 1.2711 | 32.1452 | 40.8609 | 25.2885 | 2.0091 | 32.0686 | 64.4287 | 15.9617 | 14.4886 | -0.1104 | 105.2895 | 235.3036 | 181.0725 | 233.1498 | 161.343
2012-10-30 19:46 | 28.6877 | 249.7945 | 195.5319 | 247.5897 | 175.6348 | -1.6032 | 1.2711 | 32.1446 | 40.8599 | 25.2882 | 2.0091 | 32.068 | 64.4265 | 15.9617 | 14.4869 | -0.1105 | 105.2864 | 235.3076 | 181.045 | 233.1028 | 161.1479
2012-10-30 19:46 | 28.6862 | 249.7969 | 195.5383 | 247.6392 | 175.4967 | -1.602 | 1.2711 | 32.1447 | 40.86 | 25.2883 | 2.009 | 32.068 | 64.4261 | 15.9618 | 14.4854 | -0.1104 | 105.2861 | 235.3115 | 181.0529 | 233.1538 | 161.0113
2012-10-30 19:46 | 28.6884 | 249.7717 | 195.4383 | 247.7197 | 175.7878 | -1.598 | 1.2711 | 32.1439 | 40.858 | 25.2883 | 2.009 | 32.0673 | 64.4235 | 15.9618 | 14.4876 | -0.1102 | 105.2815 | 235.2841 | 180.9507 | 233.2321 | 161.3002
2012-10-30 19:46 | 28.6869 | 249.7741 | 195.5076 | 247.7025 | 175.8917 | -1.6037 | 1.2711 | 32.1437 | 40.8573 | 25.2885 | 2.009 | 32.0671 | 64.4229 | 15.9617 | 14.4779 | -0.1106 | 105.2801 | 235.2962 | 181.0297 | 233.2246 | 161.4138
2012-10-30 19:46 | 28.6847 | 249.815 | 195.5132 | 247.6494 | 175.6982 | -1.598 | 1.2711 | 32.1438 | 40.8573 | 25.2886 | 2.009 | 32.067 | 64.4227 | 15.9617 | 14.4879 | -0.1102 | 105.28 | 235.3271 | 181.0253 | 233.1615 | 161.2103
2012-10-30 19:46 | 28.6959 | 249.8066 | 195.5048 | 247.692 | 175.807 | -1.6003 | 1.271 | 32.1425 | 40.8544 | 25.2884 | 2.0089 | 32.0659 | 64.4184 | 15.9616 | 14.491 | -0.1103 | 105.2728 | 235.3156 | 181.0138 | 233.201 | 161.316
2012-10-30 19:46 | 28.6914 | 249.7825 | 195.5199 | 247.6209 | 175.8376 | -1.5939 | 1.271 | 32.1413 | 40.8521 | 25.2879 | 2.0089 | 32.0647 | 64.4139 | 15.9616 | 14.4824 | -0.1099 | 105.266 | 235.3001 | 181.0375 | 233.1385 | 161.3552
2012-10-30 19:46 | 28.7027 | 249.7782 | 195.5509 | 247.7144 | 175.8372 | -1.6009 | 1.271 | 32.1412 | 40.8512 | 25.2882 | 2.0089 | 32.0646 | 64.4136 | 15.9615 | 14.4488 | -0.1104 | 105.2648 | 235.3294 | 181.1021 | 233.2656 | 161.3884
2012-10-30 19:46 | 28.6974 | 249.7885 | 195.577 | 247.7326 | 176.0232 | -1.6014 | 1.2709 | 32.1403 | 40.8487 | 25.2885 | 2.0088 | 32.0637 | 64.41 | 15.9615 | 14.4598 | -0.1104 | 105.2587 | 235.3287 | 181.1172 | 233.2728 | 161.5634
2012-10-30 19:47 | 28.6952 | 249.7902 | 195.5787 | 247.5894 | 176.0951 | -1.6003 | 1.2709 | 32.14 | 40.8481 | 25.2883 | 2.0088 | 32.0633 | 64.4087 | 15.9614 | 14.478 | -0.1103 | 105.2567 | 235.3122 | 181.1007 | 233.1114 | 161.6171
2012-10-30 19:47 | 28.6952 | 249.7941 | 195.5709 | 247.7383 | 175.818 | -1.5991 | 1.2709 | 32.1399 | 40.8481 | 25.2881 | 2.0088 | 32.0633 | 64.4089 | 15.9613 | 14.4617 | -0.1102 | 105.257 | 235.3324 | 181.1092 | 233.2766 | 161.3563
2012-10-30 19:47 | 28.7012 | 249.8315 | 195.5769 | 247.6463 | 175.6718 | -1.6037 | 1.2693 | 32.0973 | 40.74 | 25.2881 | 2.0063 | 32.0234 | 64.2485 | 15.9614 | 14.4636 | -0.1106 | 104.9885 | 235.3679 | 181.1133 | 233.1827 | 161.2082
2012-10-30 19:47 | 28.6967 | 249.8113 | 195.4938 | 247.7202 | 175.7336 | -1.5997 | 1.2648 | 31.9836 | 40.4537 | 25.2869 | 1.999 | 31.9066 | 63.7802 | 15.9615 | 14.4755 | -0.1103 | 104.2339 | 235.3358 | 181.0183 | 233.2447 | 161.2581
2012-10-30 19:47 | 28.6982 | 249.7893 | 195.5149 | 247.7099 | 175.8717 | -1.5962 | 1.2645 | 31.9753 | 40.4334 | 25.2865 | 1.9985 | 31.8991 | 63.7502 | 15.9615 | 14.4688 | -0.11 | 104.1836 | 235.3205 | 181.0461 | 233.2411 | 161.4029
2012-10-30 19:47 | 28.6997 | 249.7947 | 195.4811 | 247.6723 | 175.9123 | -1.602 | 1.264 | 31.9611 | 40.3974 | 25.2865 | 1.9976 | 31.8848 | 63.6925 | 15.9617 | 14.4825 | -0.1104 | 104.09 | 235.3122 | 180.9986 | 233.1898 | 161.4298
2012-10-30 19:47 | 28.7057 | 249.7577 | 195.491 | 247.643 | 175.9261 | -1.602 | 1.2637 | 31.9549 | 40.3821 | 25.2863 | 1.9972 | 31.8787 | 63.6676 | 15.9618 | 14.4804 | -0.1104 | 104.0497 | 235.2773 | 181.0106 | 233.1626 | 161.4457
2012-10-30 19:47 | 28.7049 | 249.7216 | 195.4666 | 247.6304 | 175.7965 | -1.5985 | 1.2636 | 31.9525 | 40.376 | 25.2864 | 1.997 | 31.8764 | 63.6587 | 15.9618 | 14.4714 | -0.1102 | 104.0347 | 235.2502 | 180.9952 | 233.159 | 161.3251
2012-10-30 19:47 | 28.6989 | 249.7156 | 195.4449 | 247.7185 | 175.8412 | -1.5962 | 1.2637 | 31.954 | 40.3796 | 25.2864 | 1.9971 | 31.8779 | 63.6647 | 15.9618 | 14.4654 | -0.11 | 104.0443 | 235.2502 | 180.9795 | 233.2531 | 161.3758
2012-10-30 19:47 | 28.6952 | 249.6924 | 195.4766 | 247.7226 | 175.9819 | -1.6043 | 1.2637 | 31.955 | 40.3821 | 25.2865 | 1.9972 | 31.8789 | 63.6678 | 15.9619 | 14.4454 | -0.1106 | 104.0499 | 235.247 | 181.0312 | 233.2772 | 161.5365
2012-10-30 19:47 | 28.7049 | 249.7177 | 195.5177 | 247.6226 | 176.0424 | -1.6049 | 1.2638 | 31.9574 | 40.3881 | 25.2866 | 1.9973 | 31.8813 | 63.6773 | 15.962 | 14.5081 | -0.1106 | 104.0654 | 235.2096 | 181.0096 | 233.1145 | 161.5343
2012-10-30 19:47 | 28.7019 | 249.7265 | 195.5461 | 247.5687 | 176.3165 | -1.5997 | 1.2639 | 31.9596 | 40.3942 | 25.2862 | 1.9975 | 31.8836 | 63.6866 | 15.9619 | 14.4399 | -0.1103 | 104.0809 | 235.2866 | 181.1062 | 233.1288 | 161.8766
2012-10-30 19:47 | 28.6929 | 249.6705 | 195.5764 | 247.6263 | 176.2334 | -1.5985 | 1.2683 | 32.073 | 40.679 | 25.2877 | 2.0046 | 31.9968 | 64.1398 | 15.962 | 14.439 | -0.1102 | 104.8189 | 235.2315 | 181.1374 | 233.1873 | 161.7944
2012-10-30 19:47 | 28.6854 | 249.7452 | 195.6593 | 247.5639 | 176.2142 | -1.5985 | 1.2705 | 32.1287 | 40.8196 | 25.2882 | 2.008 | 32.0519 | 64.3617 | 15.9618 | 14.4642 | -0.1102 | 105.1813 | 235.281 | 181.1951 | 233.0997 | 161.75
2012-10-30 19:47 | 28.6839 | 249.775 | 195.6814 | 247.5899 | 176.0916 | -1.6014 | 1.2709 | 32.1401 | 40.8481 | 25.2884 | 2.0088 | 32.0633 | 64.4078 | 15.9617 | 14.4749 | -0.1104 | 105.2559 | 235.3001 | 181.2065 | 233.115 | 161.6167
2012-10-30 19:47 | 28.6824 | 249.7539 | 195.676 | 247.5727 | 175.8286 | -1.5974 | 1.271 | 32.1427 | 40.8549 | 25.2883 | 2.0089 | 32.066 | 64.4185 | 15.9617 | 14.4571 | -0.1101 | 105.2735 | 235.2968 | 181.2189 | 233.1156 | 161.3715
2012-10-30 19:47 | 28.6854 | 249.7765 | 195.6161 | 247.5796 | 175.8902 | -1.6003 | 1.2711 | 32.1439 | 40.8577 | 25.2885 | 2.009 | 32.0673 | 64.4239 | 15.9616 | 14.4478 | -0.1103 | 105.2816 | 235.3287 | 181.1683 | 233.1318 | 161.4424
2012-10-30 19:47 | 28.6839 | 249.7515 | 195.6067 | 247.6134 | 175.9121 | -1.6003 | 1.2711 | 32.1449 | 40.8598 | 25.2888 | 2.0091 | 32.0682 | 64.4277 | 15.9616 | 14.4341 | -0.1103 | 105.2875 | 235.3174 | 181.1726 | 233.1793 | 161.478
2012-10-30 19:47 | 28.6892 | 249.796 | 195.6042 | 247.5011 | 175.9447 | -1.6003 | 1.2712 | 32.1456 | 40.8623 | 25.2884 | 2.0091 | 32.0689 | 64.4305 | 15.9616 | 14.4394 | -0.1103 | 105.2928 | 235.3566 | 181.1648 | 233.0617 | 161.5053
2012-10-30 19:48 | 28.6847 | 249.7758 | 195.5682 | 247.5084 | 176.0105 | -1.6014 | 1.2712 | 32.1461 | 40.8633 | 25.2885 | 2.0092 | 32.0693 | 64.4324 | 15.9616 | 14.4267 | -0.1104 | 105.2957 | 235.3491 | 181.1415 | 233.0817 | 161.5838
2012-10-30 19:48 | 28.6847 | 249.8032 | 195.5604 | 247.5397 | 175.9948 | -1.6003 | 1.2712 | 32.1467 | 40.8646 | 25.2886 | 2.0092 | 32.07 | 64.4348 | 15.9616 | 14.4757 | -0.1103 | 105.2994 | 235.3275 | 181.0847 | 233.064 | 161.5191
2012-10-30 19:48 | 28.6899 | 249.8045 | 195.6127 | 247.588 | 176.0118 | -1.6037 | 1.2712 | 32.1471 | 40.8653 | 25.2888 | 2.0092 | 32.0704 | 64.4362 | 15.9617 | 14.4523 | -0.1106 | 105.3015 | 235.3522 | 181.1604 | 233.1357 | 161.5595
2012-10-30 19:48 | 28.6937 | 249.8005 | 195.6322 | 247.5527 | 175.9726 | -1.6026 | 1.2712 | 32.1473 | 40.8664 | 25.2884 | 2.0093 | 32.0706 | 64.4381 | 15.9614 | 14.4602 | -0.1105 | 105.3045 | 235.3403 | 181.172 | 233.0925 | 161.5124
2012-10-30 19:48 | 28.6847 | 249.7993 | 195.6468 | 247.575 | 175.987 | -1.6003 | 1.2712 | 32.1477 | 40.8674 | 25.2885 | 2.0093 | 32.071 | 64.4391 | 15.9615 | 14.4757 | -0.1103 | 105.3065 | 235.3236 | 181.1711 | 233.0993 | 161.5113
2012-10-30 19:48 | 28.6944 | 249.7973 | 195.6212 | 247.4789 | 175.8913 | -1.6026 | 1.2712 | 32.1482 | 40.8681 | 25.2889 | 2.0093 | 32.0715 | 64.4415 | 15.9614 | 14.5099 | -0.1105 | 105.3096 | 235.2874 | 181.1113 | 232.969 | 161.3814
2012-10-30 19:48 | 28.6907 | 249.7818 | 195.5782 | 247.5418 | 175.9072 | -1.5974 | 1.2713 | 32.1482 | 40.8685 | 25.2886 | 2.0093 | 32.0715 | 64.4414 | 15.9615 | 14.4776 | -0.1101 | 105.3099 | 235.3042 | 181.1006 | 233.0642 | 161.4296
2012-10-30 19:48 | 28.6982 | 249.8285 | 195.5975 | 247.6237 | 175.8054 | -1.6043 | 1.2712 | 32.1481 | 40.8682 | 25.2885 | 2.0093 | 32.0714 | 64.4414 | 15.9614 | 14.4974 | -0.1106 | 105.3096 | 235.3311 | 181.1001 | 233.1263 | 161.308
2012-10-30 19:48 | 28.7012 | 249.8158 | 195.6122 | 247.5719 | 176.0074 | -1.5962 | 1.2713 | 32.1484 | 40.8693 | 25.2884 | 2.0093 | 32.0716 | 64.4425 | 15.9614 | 14.4963 | -0.11 | 105.3118 | 235.3195 | 181.1159 | 233.0756 | 161.5111
2012-10-30 19:48 | 28.6997 | 249.8221 | 195.6265 | 247.5665 | 175.9942 | -1.6061 | 1.2713 | 32.1482 | 40.8691 | 25.2882 | 2.0093 | 32.0715 | 64.4417 | 15.9614 | 14.4744 | -0.1107 | 105.3109 | 235.3477 | 181.1521 | 233.0921 | 161.5198
2012-10-30 19:48 | 28.7057 | 249.8203 | 195.6639 | 247.5176 | 175.9651 | -1.6014 | 1.2712 | 32.1476 | 40.8673 | 25.2883 | 2.0093 | 32.0709 | 64.4392 | 15.9614 | 14.4845 | -0.1104 | 105.3065 | 235.3358 | 181.1794 | 233.0331 | 161.4806
2012-10-30 19:48 | 28.6944 | 249.8286 | 195.6015 | 247.5573 | 175.743 | -1.5945 | 1.2713 | 32.1476 | 40.8677 | 25.2881 | 2.0093 | 32.0709 | 64.4398 | 15.9612 | 14.4732 | -0.1099 | 105.3075 | 235.3554 | 181.1283 | 233.0841 | 161.2698
2012-10-30 19:48 | 28.7004 | 249.8072 | 195.5564 | 247.6886 | 175.8895 | -1.6049 | 1.2712 | 32.1474 | 40.8663 | 25.2887 | 2.0093 | 32.0707 | 64.4389 | 15.9614 | 14.4628 | -0.1106 | 105.3051 | 235.3444 | 181.0936 | 233.2258 | 161.4267
2012-10-30 19:48 | 28.6997 | 249.8417 | 195.5832 | 247.5743 | 175.9474 | -1.5968 | 1.2712 | 32.1471 | 40.8656 | 25.2886 | 2.0093 | 32.0704 | 64.4376 | 15.9613 | 14.4621 | -0.1101 | 105.3032 | 235.3796 | 181.1211 | 233.1122 | 161.4853
2012-10-30 19:48 | 28.7034 | 249.8258 | 195.5751 | 247.5232 | 175.8886 | -1.5962 | 1.2712 | 32.1472 | 40.8665 | 25.2883 | 2.0093 | 32.0706 | 64.4389 | 15.9612 | 14.4495 | -0.11 | 105.3054 | 235.3763 | 181.1256 | 233.0737 | 161.4391
2012-10-30 19:48 | 28.7012 | 249.7845 | 195.5651 | 247.5562 | 175.7732 | -1.602 | 1.2712 | 32.1471 | 40.866 | 25.2884 | 2.0093 | 32.0704 | 64.4375 | 15.9614 | 14.48 | -0.1104 | 105.3035 | 235.3045 | 181.0851 | 233.0762 | 161.2932
2012-10-30 19:48 | 28.7019 | 249.8165 | 195.5815 | 247.6627 | 175.7193 | -1.5945 | 1.2712 | 32.1471 | 40.8658 | 25.2885 | 2.0093 | 32.0703 | 64.4374 | 15.9613 | 14.4276 | -0.1099 | 105.3032 | 235.3889 | 181.1539 | 233.2351 | 161.2917
2012-10-30 19:48 | 28.7019 | 249.8048 | 195.5501 | 247.5844 | 175.8871 | -1.6032 | 1.2712 | 32.1466 | 40.8648 | 25.2883 | 2.0092 | 32.07 | 64.4361 | 15.9613 | 14.497 | -0.1105 | 105.3009 | 235.3078 | 181.0531 | 233.0874 | 161.3901
2012-10-30 19:48 | 28.7019 | 249.8322 | 195.5619 | 247.6235 | 175.8871 | -1.6009 | 1.2712 | 32.1465 | 40.8642 | 25.2886 | 2.0092 | 32.0698 | 64.4354 | 15.9613 | 14.4929 | -0.1104 | 105.2995 | 235.3393 | 181.069 | 233.1306 | 161.3942
2012-10-30 19:48 | 28.7064 | 249.8602 | 195.5781 | 247.5614 | 176.1219 | -1.6049 | 1.2712 | 32.1466 | 40.8645 | 25.2885 | 2.0093 | 32.0699 | 64.4365 | 15.9611 | 14.4403 | -0.1106 | 105.301 | 235.4199 | 181.1378 | 233.1211 | 161.6816
2012-10-30 19:48 | 28.7064 | 249.8093 | 195.586 | 247.6084 | 176.1844 | -1.5985 | 1.2712 | 32.1464 | 40.8643 | 25.2884 | 2.0092 | 32.0697 | 64.4357 | 15.9611 | 14.4648 | -0.1102 | 105.3001 | 235.3445 | 181.1212 | 233.1436 | 161.7196
2012-10-30 19:48 | 28.7064 | 249.868 | 195.6096 | 247.6633 | 176.079 | -1.6009 | 1.2712 | 32.1462 | 40.8639 | 25.2883 | 2.0092 | 32.0695 | 64.4349 | 15.9611 | 14.4729 | -0.1104 | 105.2989 | 235.3951 | 181.1367 | 233.1904 | 161.6061
2012-10-30 19:48 | 28.7079 | 249.8343 | 195.6464 | 247.6139 | 176.0532 | -1.6049 | 1.2712 | 32.1457 | 40.8629 | 25.2882 | 2.0092 | 32.0691 | 64.4335 | 15.9611 | 14.454 | -0.1106 | 105.2964 | 235.3803 | 181.1924 | 233.1599 | 161.5992
2012-10-30 19:49 | 28.7064 | 249.8719 | 195.6489 | 247.6633 | 175.8409 | -1.6009 | 1.2711 | 32.1453 | 40.8615 | 25.2884 | 2.0092 | 32.0686 | 64.4318 | 15.961 | 14.4199 | -0.1104 | 105.2933 | 235.452 | 181.229 | 233.2434 | 161.421
2012-10-30 19:49 | 28.7042 | 249.858 | 195.6742 | 247.6258 | 175.9948 | -1.598 | 1.2712 | 32.1451 | 40.8615 | 25.288 | 2.0092 | 32.0685 | 64.4309 | 15.9611 | 14.4462 | -0.1102 | 105.2924 | 235.4118 | 181.228 | 233.1796 | 161.5486
2012-10-30 19:49 | 28.7064 | 249.821 | 195.6567 | 247.5536 | 175.9385 | -1.5962 | 1.2711 | 32.1442 | 40.859 | 25.2882 | 2.0091 | 32.0675 | 64.427 | 15.9611 | 14.4811 | -0.11 | 105.286 | 235.3399 | 181.1756 | 233.0725 | 161.4574
2012-10-30 19:49 | 28.7012 | 249.8119 | 195.5965 | 247.6503 | 175.8513 | -1.6049 | 1.2711 | 32.1436 | 40.8574 | 25.2882 | 2.0091 | 32.0669 | 64.4249 | 15.961 | 14.4555 | -0.1106 | 105.2823 | 235.3564 | 181.141 | 233.1948 | 161.3958
2012-10-30 19:49 | 28.7094 | 249.8201 | 195.6322 | 247.6467 | 175.6604 | -1.6043 | 1.2711 | 32.143 | 40.8559 | 25.2882 | 2.009 | 32.0664 | 64.4222 | 15.9612 | 14.4637 | -0.1106 | 105.2781 | 235.3564 | 181.1685 | 233.183 | 161.1967
2012-10-30 19:49 | 28.7019 | 249.8478 | 195.554 | 247.698 | 175.6764 | -1.5997 | 1.2711 | 32.1428 | 40.8551 | 25.2883 | 2.009 | 32.0661 | 64.4218 | 15.961 | 14.4807 | -0.1103 | 105.2769 | 235.3671 | 181.0733 | 233.2173 | 161.1957
2012-10-30 19:49 | 28.7154 | 249.8222 | 195.6029 | 247.6566 | 175.9982 | -1.6049 | 1.271 | 32.1424 | 40.8544 | 25.2883 | 2.009 | 32.0658 | 64.4204 | 15.9611 | 14.4656 | -0.1106 | 105.2748 | 235.3566 | 181.1373 | 233.191 | 161.5326
2012-10-30 19:49 | 28.7072 | 249.8414 | 195.5868 | 247.6406 | 175.6934 | -1.5997 | 1.271 | 32.142 | 40.8529 | 25.2884 | 2.009 | 32.0653 | 64.4188 | 15.9609 | 14.4778 | -0.1103 | 105.2717 | 235.3636 | 181.109 | 233.1628 | 161.2156
2012-10-30 19:49 | 28.7087 | 249.8781 | 195.5451 | 247.6656 | 175.7964 | -1.6026 | 1.271 | 32.1413 | 40.8516 | 25.2882 | 2.0089 | 32.0646 | 64.4157 | 15.961 | 14.4834 | -0.1105 | 105.2673 | 235.3947 | 181.0617 | 233.1822 | 161.313
2012-10-30 19:49 | 28.7079 | 249.8421 | 195.5718 | 247.6139 | 175.737 | -1.5997 | 1.271 | 32.1408 | 40.8496 | 25.2886 | 2.0089 | 32.064 | 64.4138 | 15.9609 | 14.4907 | -0.1103 | 105.2634 | 235.3514 | 181.0811 | 233.1232 | 161.2463
2012-10-30 19:49 | 28.7087 | 249.839 | 195.4979 | 247.6773 | 175.6832 | -1.5968 | 1.271 | 32.1403 | 40.8491 | 25.2881 | 2.0089 | 32.0635 | 64.4118 | 15.9609 | 14.4956 | -0.1101 | 105.2609 | 235.3434 | 181.0023 | 233.1817 | 161.1876
2012-10-30 19:49 | 28.7087 | 249.8507 | 195.5412 | 247.6891 | 175.6558 | -1.602 | 1.2709 | 32.1399 | 40.848 | 25.2883 | 2.0089 | 32.0633 | 64.4112 | 15.9609 | 14.4875 | -0.1104 | 105.2591 | 235.3632 | 181.0537 | 233.2016 | 161.1683
2012-10-30 19:49 | 28.7206 | 249.843 | 195.5138 | 247.6422 | 175.6912 | -1.6061 | 1.2709 | 32.1399 | 40.8479 | 25.2882 | 2.0089 | 32.0632 | 64.4103 | 15.961 | 14.4953 | -0.1107 | 105.2582 | 235.3477 | 181.0185 | 233.1469 | 161.1959
2012-10-30 19:49 | 28.7162 | 249.7956 | 195.5133 | 247.7083 | 175.7414 | -1.6003 | 1.2669 | 32.0378 | 40.59 | 25.2875 | 2.0023 | 31.9588 | 63.9918 | 15.9609 | 14.4542 | -0.1103 | 104.5818 | 235.3414 | 181.0591 | 233.2541 | 161.2872
2012-10-30 19:49 | 28.7079 | 249.8303 | 195.4618 | 247.653 | 175.7448 | -1.6026 | 1.2648 | 31.9835 | 40.4543 | 25.2864 | 1.9991 | 31.9073 | 63.7856 | 15.9609 | 14.4663 | -0.1105 | 104.2399 | 235.364 | 180.9955 | 233.1867 | 161.2785
2012-10-30 19:49 | 28.7162 | 249.8347 | 195.4543 | 247.5987 | 175.6868 | -1.6009 | 1.2642 | 31.966 | 40.4102 | 25.2863 | 1.998 | 31.89 | 63.7161 | 15.9609 | 14.4868 | -0.1104 | 104.1263 | 235.3479 | 180.9675 | 233.1119 | 161.2
2012-10-30 19:49 | 28.7154 | 249.8065 | 195.4064 | 247.6527 | 175.7679 | -1.6003 | 1.2638 | 31.9559 | 40.3852 | 25.286 | 1.9973 | 31.8797 | 63.6749 | 15.961 | 14.4697 | -0.1103 | 104.0601 | 235.3368 | 180.9367 | 233.183 | 161.2982
2012-10-30 19:49 | 28.7109 | 249.755 | 195.4373 | 247.6835 | 175.8025 | -1.6043 | 1.2637 | 31.9536 | 40.3787 | 25.2864 | 1.9972 | 31.8774 | 63.665 | 15.9612 | 14.4897 | -0.1106 | 104.0437 | 235.2653 | 180.9476 | 233.1938 | 161.3128
2012-10-30 19:49 | 28.7109 | 249.7316 | 195.4608 | 247.656 | 175.4824 | -1.6009 | 1.2636 | 31.9527 | 40.3765 | 25.2864 | 1.9971 | 31.8765 | 63.6613 | 15.9612 | 14.4897 | -0.1104 | 104.0378 | 235.2419 | 180.9711 | 233.1663 | 160.9927
2012-10-30 19:49 | 28.7072 | 249.7396 | 195.4729 | 247.668 | 175.627 | -1.602 | 1.2637 | 31.9536 | 40.3787 | 25.2864 | 1.9972 | 31.8775 | 63.6652 | 15.9612 | 14.4982 | -0.1104 | 104.0439 | 235.2414 | 180.9747 | 233.1698 | 161.1288
2012-10-30 19:49 | 28.7087 | 249.7059 | 195.4076 | 247.6382 | 175.7495 | -1.6003 | 1.2637 | 31.955 | 40.3824 | 25.2863 | 1.9973 | 31.8788 | 63.6701 | 15.9613 | 14.516 | -0.1103 | 104.0524 | 235.1899 | 180.8916 | 233.1222 | 161.2335
2012-10-30 19:49 | 28.6952 | 249.7276 | 195.4412 | 247.5581 | 175.9975 | -1.6014 | 1.2638 | 31.9564 | 40.3862 | 25.2861 | 1.9973 | 31.8803 | 63.6756 | 15.9614 | 14.4984 | -0.1104 | 104.0618 | 235.2292 | 180.9428 | 233.0597 | 161.4991
2012-10-30 19:50 | 28.7012 | 249.6984 | 195.4511 | 247.6032 | 175.9762 | -1.5974 | 1.267 | 32.0393 | 40.5945 | 25.2872 | 2.0026 | 31.9649 | 64.0143 | 15.9614 | 14.4881 | -0.1101 | 104.6088 | 235.2103 | 180.963 | 233.1151 | 161.4881
2012-10-30 19:50 | 28.6937 | 249.7144 | 195.5301 | 247.6584 | 175.9687 | -1.6003 | 1.269 | 32.0909 | 40.7244 | 25.2876 | 2.0057 | 32.0139 | 64.2109 | 15.9613 | 14.5092 | -0.1103 | 104.9354 | 235.2052 | 181.0209 | 233.1492 | 161.4595
2012-10-30 19:50 | 28.6974 | 249.6711 | 195.5809 | 247.6151 | 176.0115 | -1.6003 | 1.2698 | 32.1096 | 40.7715 | 25.288 | 2.0069 | 32.0329 | 64.2875 | 15.9612 | 14.517 | -0.1103 | 105.059 | 235.1541 | 181.0639 | 233.0981 | 161.4945
2012-10-30 19:50 | 28.6959 | 249.7009 | 195.5794 | 247.5705 | 176.0178 | -1.6043 | 1.2703 | 32.124 | 40.8079 | 25.288 | 2.0078 | 32.0476 | 64.3459 | 15.9613 | 14.5032 | -0.1106 | 105.1538 | 235.1977 | 181.0762 | 233.0673 | 161.5146
2012-10-30 19:50 | 28.6869 | 249.7115 | 195.5586 | 247.5459 | 176.0439 | -1.6037 | 1.2706 | 32.1316 | 40.8273 | 25.2879 | 2.0083 | 32.055 | 64.376 | 15.9612 | 14.4861 | -0.1106 | 105.2033 | 235.2254 | 181.0725 | 233.0598 | 161.5578
2012-10-30 19:50 | 28.6944 | 249.7151 | 195.5976 | 247.5338 | 176.0241 | -1.6043 | 1.2708 | 32.1366 | 40.8396 | 25.2882 | 2.0086 | 32.06 | 64.3964 | 15.9612 | 14.5507 | -0.1106 | 105.236 | 235.1644 | 181.0469 | 232.9831 | 161.4734
2012-10-30 19:50 | 28.6817 | 249.6945 | 195.6202 | 247.6308 | 175.9997 | -1.5951 | 1.271 | 32.1401 | 40.8486 | 25.2882 | 2.0088 | 32.0633 | 64.41 | 15.9611 | 14.4645 | -0.11 | 105.2586 | 235.23 | 181.1557 | 233.1663 | 161.5352
2012-10-30 19:50 | 28.6854 | 249.7139 | 195.5964 | 247.5639 | 176.0229 | -1.6043 | 1.271 | 32.1425 | 40.8546 | 25.2882 | 2.009 | 32.0657 | 64.4195 | 15.9612 | 14.4927 | -0.1106 | 105.2741 | 235.2212 | 181.1037 | 233.0712 | 161.5302
2012-10-30 19:50 | 28.6832 | 249.7234 | 195.6296 | 247.5735 | 176.1066 | -1.6003 | 1.2711 | 32.144 | 40.8578 | 25.2886 | 2.0091 | 32.0673 | 64.426 | 15.9611 | 14.515 | -0.1103 | 105.2838 | 235.2084 | 181.1146 | 233.0585 | 161.5916
2012-10-30 19:50 | 28.6817 | 249.7219 | 195.6556 | 247.4936 | 176.0153 | -1.6003 | 1.2711 | 32.1447 | 40.8605 | 25.2881 | 2.0091 | 32.068 | 64.429 | 15.9611 | 14.4931 | -0.1103 | 105.2894 | 235.2288 | 181.1625 | 233.0005 | 161.5222
2012-10-30 19:50 | 28.6735 | 249.7255 | 195.6317 | 247.5755 | 175.8509 | -1.6032 | 1.2711 | 32.1453 | 40.8614 | 25.2885 | 2.0092 | 32.0686 | 64.4318 | 15.961 | 14.44 | -0.1105 | 105.2932 | 235.2855 | 181.1917 | 233.1355 | 161.4109
2012-10-30 19:50 | 28.6787 | 249.6954 | 195.6605 | 247.6082 | 175.7937 | -1.5974 | 1.2712 | 32.1459 | 40.8629 | 25.2885 | 2.0092 | 32.0692 | 64.4337 | 15.9611 | 14.4819 | -0.1101 | 105.2966 | 235.2135 | 181.1786 | 233.1263 | 161.3118
2012-10-30 19:50 | 28.6787 | 249.7111 | 195.5897 | 247.5572 | 175.9225 | -1.6009 | 1.2712 | 32.1466 | 40.8643 | 25.2886 | 2.0093 | 32.0699 | 64.4365 | 15.9611 | 14.5187 | -0.1104 | 105.3008 | 235.1924 | 181.071 | 233.0385 | 161.4038
2012-10-30 19:50 | 28.6772 | 249.7487 | 195.604 | 247.5362 | 176.003 | -1.6009 | 1.2712 | 32.1471 | 40.8659 | 25.2884 | 2.0093 | 32.0704 | 64.4386 | 15.9611 | 14.4845 | -0.1104 | 105.3045 | 235.2642 | 181.1195 | 233.0517 | 161.5185
2012-10-30 19:50 | 28.6824 | 249.7618 | 195.562 | 247.51 | 175.7935 | -1.6003 | 1.2712 | 32.1475 | 40.8673 | 25.2882 | 2.0093 | 32.0707 | 64.4404 | 15.961 | 14.5428 | -0.1103 | 105.3076 | 235.219 | 181.0192 | 232.9672 | 161.2507
2012-10-30 19:50 | 28.6735 | 249.7294 | 195.5924 | 247.6343 | 175.7651 | -1.6009 | 1.2713 | 32.1481 | 40.8685 | 25.2885 | 2.0094 | 32.0714 | 64.4432 | 15.961 | 14.489 | -0.1104 | 105.3116 | 235.2404 | 181.1034 | 233.1453 | 161.2761
2012-10-30 19:50 | 28.6795 | 249.7471 | 195.5552 | 247.5894 | 175.9936 | -1.5991 | 1.2713 | 32.1484 | 40.8693 | 25.2885 | 2.0094 | 32.0718 | 64.4447 | 15.9609 | 14.4746 | -0.1102 | 105.314 | 235.2725 | 181.0806 | 233.1148 | 161.519
2012-10-30 19:50 | 28.6839 | 249.732 | 195.5556 | 247.5507 | 175.9394 | -1.6084 | 1.2713 | 32.1486 | 40.8693 | 25.2887 | 2.0094 | 32.0718 | 64.4447 | 15.961 | 14.5075 | -0.1109 | 105.3139 | 235.2245 | 181.0481 | 233.0432 | 161.4319
2012-10-30 19:50 | 28.6869 | 249.735 | 195.5311 | 247.5459 | 175.9385 | -1.5991 | 1.2713 | 32.1482 | 40.8687 | 25.2885 | 2.0094 | 32.0715 | 64.444 | 15.9609 | 14.4901 | -0.1102 | 105.3127 | 235.2449 | 181.041 | 233.0558 | 161.4484
2012-10-30 19:50 | 28.6817 | 249.7337 | 195.581 | 247.6151 | 175.8591 | -1.6014 | 1.2713 | 32.148 | 40.8685 | 25.2883 | 2.0094 | 32.0714 | 64.443 | 15.961 | 14.4645 | -0.1104 | 105.3115 | 235.2692 | 181.1165 | 233.1506 | 161.3946
2012-10-30 19:50 | 28.6929 | 249.7762 | 195.5529 | 247.6106 | 175.8079 | -1.6003 | 1.2713 | 32.148 | 40.8683 | 25.2883 | 2.0094 | 32.0712 | 64.4429 | 15.9608 | 14.5084 | -0.1103 | 105.3112 | 235.2678 | 181.0445 | 233.1022 | 161.2995
2012-10-30 19:50 | 28.6929 | 249.7566 | 195.5764 | 247.6655 | 175.8625 | -1.5997 | 1.2713 | 32.1479 | 40.8682 | 25.2884 | 2.0094 | 32.0712 | 64.4435 | 15.9607 | 14.439 | -0.1103 | 105.3116 | 235.3176 | 181.1374 | 233.2265 | 161.4235
2012-10-30 19:50 | 28.6824 | 249.7344 | 195.5659 | 247.5844 | 175.9262 | -1.5997 | 1.2713 | 32.1479 | 40.8682 | 25.2884 | 2.0094 | 32.0712 | 64.4427 | 15.9609 | 14.4367 | -0.1103 | 105.3109 | 235.2977 | 181.1292 | 233.1477 | 161.4895
2012-10-30 19:51 | 28.6982 | 249.7463 | 195.5817 | 247.5493 | 175.8795 | -1.6026 | 1.2713 | 32.148 | 40.8682 | 25.2884 | 2.0094 | 32.0712 | 64.4434 | 15.9607 | 14.4525 | -0.1105 | 105.3116 | 235.2938 | 181.1292 | 233.0968 | 161.427
2012-10-30 19:51 | 28.6952 | 249.8098 | 195.5277 | 247.6012 | 175.7516 | -1.6026 | 1.2712 | 32.1478 | 40.8678 | 25.2884 | 2.0094 | 32.071 | 64.4423 | 15.9608 | 14.478 | -0.1105 | 105.31 | 235.3318 | 181.0497 | 233.1232 | 161.2736
2012-10-30 19:51 | 28.6914 | 249.7473 | 195.5435 | 247.6953 | 175.7634 | -1.6003 | 1.2712 | 32.1473 | 40.8666 | 25.2883 | 2.0093 | 32.0707 | 64.4409 | 15.9608 | 14.4538 | -0.1103 | 105.3076 | 235.2935 | 181.0897 | 233.2415 | 161.3096
2012-10-30 19:51 | 28.6937 | 249.7887 | 195.5458 | 247.6506 | 175.797 | -1.6026 | 1.2712 | 32.147 | 40.8659 | 25.2883 | 2.0093 | 32.0703 | 64.4401 | 15.9606 | 14.4643 | -0.1105 | 105.3061 | 235.3244 | 181.0815 | 233.1863 | 161.3327
2012-10-30 19:51 | 28.6937 | 249.7926 | 195.5104 | 247.6075 | 175.8711 | -1.602 | 1.2712 | 32.1472 | 40.8663 | 25.2884 | 2.0093 | 32.0704 | 64.4401 | 15.9608 | 14.4153 | -0.1104 | 105.3064 | 235.3773 | 181.0951 | 233.1922 | 161.4558
2012-10-30 19:51 | 28.6967 | 249.7917 | 195.4899 | 247.5478 | 175.7531 | -1.6026 | 1.2712 | 32.1469 | 40.8655 | 25.2884 | 2.0093 | 32.0702 | 64.4393 | 15.9607 | 14.4347 | -0.1105 | 105.3048 | 235.357 | 181.0552 | 233.1131 | 161.3184
2012-10-30 19:51 | 28.7012 | 249.8001 | 195.4944 | 247.6463 | 175.7069 | -1.5974 | 1.2712 | 32.1465 | 40.8645 | 25.2884 | 2.0093 | 32.0698 | 64.4383 | 15.9606 | 14.4392 | -0.1101 | 105.3028 | 235.3609 | 181.0552 | 233.2071 | 161.2677
2012-10-30 19:51 | 28.7004 | 249.8033 | 195.5093 | 247.6651 | 175.7607 | -1.6003 | 1.2712 | 32.146 | 40.8638 | 25.2881 | 2.0093 | 32.0694 | 64.4364 | 15.9606 | 14.4873 | -0.1103 | 105.3002 | 235.316 | 181.022 | 233.1778 | 161.2734
2012-10-30 19:51 | 28.6989 | 249.7822 | 195.5235 | 247.6323 | 175.9388 | -1.6003 | 1.2712 | 32.1455 | 40.862 | 25.2883 | 2.0093 | 32.0688 | 64.4343 | 15.9605 | 14.4573 | -0.1103 | 105.2964 | 235.3249 | 181.0662 | 233.175 | 161.4815
2012-10-30 19:51 | 28.6877 | 249.8062 | 195.5162 | 247.5702 | 175.8144 | -1.5957 | 1.2712 | 32.1454 | 40.8622 | 25.2881 | 2.0092 | 32.0687 | 64.4338 | 15.9606 | 14.442 | -0.11 | 105.296 | 235.3642 | 181.0742 | 233.1282 | 161.3724
2012-10-30 19:51 | 28.6877 | 249.8219 | 195.5241 | 247.6563 | 175.7363 | -1.6003 | 1.2712 | 32.1452 | 40.8614 | 25.2882 | 2.0092 | 32.0685 | 64.4329 | 15.9606 | 14.5032 | -0.1103 | 105.2943 | 235.3187 | 181.0209 | 233.1531 | 161.2331
2012-10-30 19:51 | 28.6952 | 249.7589 | 195.4609 | 247.6913 | 175.7516 | -1.6014 | 1.2711 | 32.1449 | 40.8606 | 25.2883 | 2.0092 | 32.0682 | 64.4324 | 15.9605 | 14.4944 | -0.1104 | 105.2929 | 235.2645 | 180.9665 | 233.1969 | 161.2572
2012-10-30 19:51 | 28.6839 | 249.8337 | 195.4771 | 247.68 | 175.8067 | -1.6026 | 1.2711 | 32.1448 | 40.8603 | 25.2883 | 2.0092 | 32.0681 | 64.4318 | 15.9605 | 14.479 | -0.1105 | 105.2921 | 235.3547 | 180.9981 | 233.201 | 161.3277
2012-10-30 19:51 | 28.6974 | 249.8316 | 195.4788 | 247.6269 | 175.8202 | -1.5997 | 1.2711 | 32.1446 | 40.8603 | 25.288 | 2.0092 | 32.0679 | 64.4311 | 15.9605 | 14.5129 | -0.1103 | 105.2914 | 235.3187 | 180.9659 | 233.114 | 161.3073
2012-10-30 19:51 | 28.6884 | 249.7952 | 195.5326 | 247.657 | 175.8229 | -1.6026 | 1.2711 | 32.1442 | 40.8592 | 25.2881 | 2.0092 | 32.0676 | 64.4295 | 15.9605 | 14.4794 | -0.1105 | 105.2887 | 235.3158 | 181.0532 | 233.1776 | 161.3435
2012-10-30 19:51 | 28.6892 | 249.8077 | 195.5099 | 247.65 | 175.7964 | -1.6003 | 1.2711 | 32.1442 | 40.8593 | 25.288 | 2.0092 | 32.0675 | 64.4297 | 15.9604 | 14.4761 | -0.1103 | 105.289 | 235.3316 | 181.0338 | 233.1739 | 161.3203
2012-10-30 19:51 | 28.6922 | 249.8107 | 195.4854 | 247.6569 | 175.8501 | -1.5974 | 1.2711 | 32.144 | 40.8581 | 25.2884 | 2.0092 | 32.0673 | 64.4293 | 15.9603 | 14.5281 | -0.1101 | 105.2874 | 235.2826 | 180.9573 | 233.1288 | 161.322
2012-10-30 19:51 | 28.6824 | 249.844 | 195.4795 | 247.7607 | 175.8169 | -1.6037 | 1.2711 | 32.1432 | 40.8562 | 25.2884 | 2.0091 | 32.0664 | 64.4254 | 15.9604 | 14.502 | -0.1106 | 105.2815 | 235.342 | 180.9775 | 233.2587 | 161.3149
2012-10-30 19:51 | 28.6824 | 249.8401 | 195.5581 | 247.6667 | 175.7857 | -1.6009 | 1.2711 | 32.1426 | 40.855 | 25.2881 | 2.0091 | 32.0658 | 64.4229 | 15.9605 | 14.5305 | -0.1104 | 105.2779 | 235.3096 | 181.0276 | 233.1362 | 161.2552
2012-10-30 19:51 | 28.6854 | 249.7843 | 195.5061 | 247.6893 | 175.7262 | -1.6014 | 1.2711 | 32.1425 | 40.8548 | 25.288 | 2.0091 | 32.0656 | 64.4224 | 15.9603 | 14.5254 | -0.1104 | 105.2772 | 235.2589 | 180.9807 | 233.1639 | 161.2008
2012-10-30 19:51 | 28.6929 | 249.8349 | 195.4939 | 247.5793 | 175.8313 | -1.5997 | 1.271 | 32.1421 | 40.8539 | 25.2881 | 2.0091 | 32.0655 | 64.4218 | 15.9603 | 14.5002 | -0.1103 | 105.2757 | 235.3347 | 180.9937 | 233.0791 | 161.3311
2012-10-30 19:51 | 28.6884 | 249.8382 | 195.5444 | 247.6805 | 175.737 | -1.6026 | 1.271 | 32.1419 | 40.8529 | 25.2883 | 2.009 | 32.0652 | 64.4204 | 15.9604 | 14.4998 | -0.1105 | 105.2733 | 235.3384 | 181.0446 | 233.1807 | 161.2372
2012-10-30 19:52 | 28.6854 | 249.8392 | 195.4864 | 247.6854 | 175.6794 | -1.6003 | 1.271 | 32.1414 | 40.8522 | 25.288 | 2.009 | 32.0648 | 64.4189 | 15.9604 | 14.5213 | -0.1103 | 105.2711 | 235.3179 | 180.9651 | 233.1641 | 161.1581
2012-10-30 19:52 | 28.6892 | 249.7842 | 195.447 | 247.65 | 175.7066 | -1.6009 | 1.2656 | 32.0029 | 40.5029 | 25.2867 | 2.0003 | 31.9263 | 63.863 | 15.9606 | 14.5047 | -0.1104 | 104.3659 | 235.2795 | 180.9423 | 233.1453 | 161.2019
2012-10-30 19:52 | 28.675 | 249.6996 | 195.4485 | 247.6201 | 175.7392 | -1.5991 | 1.2639 | 31.9582 | 40.3905 | 25.2863 | 1.9975 | 31.8819 | 63.6841 | 15.9609 | 14.5027 | -0.1102 | 104.0746 | 235.1969 | 180.9458 | 233.1174 | 161.2365
2012-10-30 19:52 | 28.666 | 249.6436 | 195.4592 | 247.6816 | 176.0035 | -1.6032 | 1.2639 | 31.9598 | 40.3943 | 25.2865 | 1.9976 | 31.8835 | 63.69 | 15.961 | 14.4937 | -0.1105 | 104.0843 | 235.1499 | 180.9655 | 233.1879 | 161.5098
2012-10-30 19:52 | 28.6697 | 249.6512 | 195.4904 | 247.5678 | 176.0774 | -1.6014 | 1.2679 | 32.0615 | 40.6509 | 25.2871 | 2.0038 | 31.9832 | 64.0895 | 15.9609 | 14.5219 | -0.1104 | 104.7404 | 235.1293 | 180.9685 | 233.0459 | 161.5555
2012-10-30 19:52 | 28.6637 | 249.6922 | 195.5276 | 247.5422 | 175.7631 | -1.5962 | 1.2692 | 32.0961 | 40.7375 | 25.2877 | 2.0061 | 32.0197 | 64.2363 | 15.9608 | 14.4914 | -0.11 | 104.9737 | 235.2008 | 181.0362 | 233.0508 | 161.2717
2012-10-30 19:52 | 28.6577 | 249.6705 | 195.5177 | 247.6068 | 175.9327 | -1.6003 | 1.2699 | 32.1136 | 40.7816 | 25.288 | 2.0073 | 32.0372 | 64.3069 | 15.9607 | 14.465 | -0.1103 | 105.0885 | 235.2055 | 181.0527 | 233.1418 | 161.4677
2012-10-30 19:52 | 28.657 | 249.6972 | 195.5484 | 247.6217 | 176.0686 | -1.5997 | 1.2704 | 32.1258 | 40.8124 | 25.288 | 2.008 | 32.0492 | 64.3551 | 15.9607 | 14.4684 | -0.1103 | 105.1675 | 235.2288 | 181.08 | 233.1533 | 161.6002
2012-10-30 19:52 | 28.654 | 249.6981 | 195.5847 | 247.5639 | 176.0774 | -1.6014 | 1.2707 | 32.1339 | 40.8328 | 25.2882 | 2.0085 | 32.0571 | 64.3871 | 15.9606 | 14.4736 | -0.1104 | 105.2199 | 235.2245 | 181.1111 | 233.0903 | 161.6038
2012-10-30 19:52 | 28.6532 | 249.7521 | 195.6153 | 247.4886 | 176.1312 | -1.598 | 1.2709 | 32.1386 | 40.8444 | 25.2884 | 2.0088 | 32.0619 | 64.4059 | 15.9607 | 14.4401 | -0.1102 | 105.2502 | 235.312 | 181.1752 | 233.0485 | 161.6911
2012-10-30 19:52 | 28.648 | 249.7274 | 195.6258 | 247.5696 | 175.9542 | -1.6003 | 1.271 | 32.1426 | 40.8547 | 25.2883 | 2.0091 | 32.0659 | 64.4221 | 15.9607 | 14.4757 | -0.1103 | 105.2768 | 235.2517 | 181.1501 | 233.0939 | 161.4785
2012-10-30 19:52 | 28.6495 | 249.7093 | 195.6352 | 247.5594 | 175.9596 | -1.5968 | 1.2711 | 32.1451 | 40.8608 | 25.2886 | 2.0092 | 32.0685 | 64.4332 | 15.9605 | 14.4813 | -0.1101 | 105.294 | 235.228 | 181.1539 | 233.0781 | 161.4783
2012-10-30 19:52 | 28.651 | 249.7186 | 195.6642 | 247.5139 | 176.25 | -1.6026 | 1.2712 | 32.1467 | 40.8653 | 25.2882 | 2.0093 | 32.0699 | 64.439 | 15.9605 | 14.4583 | -0.1105 | 105.3043 | 235.2603 | 181.2059 | 233.0556 | 161.7917
2012-10-30 19:52 | 28.6315 | 249.7579 | 195.6683 | 247.4474 | 176.2305 | -1.6043 | 1.2712 | 32.1471 | 40.8657 | 25.2886 | 2.0093 | 32.0704 | 64.4405 | 15.9607 | 14.4674 | -0.1106 | 105.3062 | 235.2905 | 181.2009 | 232.98 | 161.7631
2012-10-30 19:52 | 28.6413 | 249.7324 | 195.7449 | 247.5394 | 176.3418 | -1.5991 | 1.2713 | 32.1478 | 40.8681 | 25.2882 | 2.0094 | 32.0711 | 64.4437 | 15.9605 | 14.4772 | -0.1102 | 105.3118 | 235.2552 | 181.2677 | 233.0622 | 161.8646
2012-10-30 19:53 | 28.6323 | 249.7273 | 195.7673 | 247.6323 | 176.0986 | -1.5951 | 1.2713 | 32.1482 | 40.8687 | 25.2885 | 2.0094 | 32.0714 | 64.4453 | 15.9605 | 14.4641 | -0.11 | 105.314 | 235.2632 | 181.3032 | 233.1682 | 161.6345
2012-10-30 19:53 | 28.648 | 249.7783 | 195.7516 | 247.5422 | 176.126 | -1.6037 | 1.2713 | 32.1483 | 40.8688 | 25.2886 | 2.0094 | 32.0715 | 64.4455 | 15.9605 | 14.4757 | -0.1106 | 105.3142 | 235.3026 | 181.2759 | 233.0665 | 161.6503
2012-10-30 19:53 | 28.6428 | 249.7652 | 195.7739 | 247.4743 | 176.0349 | -1.598 | 1.2713 | 32.1487 | 40.8697 | 25.2886 | 2.0095 | 32.0719 | 64.447 | 15.9605 | 14.4991 | -0.1102 | 105.3167 | 235.2661 | 181.2748 | 232.9752 | 161.5358
2012-10-30 19:53 | 28.6338 | 249.7132 | 195.6981 | 247.6416 | 175.9244 | -1.6003 | 1.2713 | 32.1487 | 40.8705 | 25.2882 | 2.0095 | 32.072 | 64.4476 | 15.9605 | 14.4534 | -0.1103 | 105.3181 | 235.2598 | 181.2447 | 233.1882 | 161.471
2012-10-30 19:53 | 28.6323 | 249.7469 | 195.7162 | 247.55 | 176.1142 | -1.5991 | 1.2713 | 32.149 | 40.8709 | 25.2884 | 2.0095 | 32.0722 | 64.4489 | 15.9604 | 14.4886 | -0.1102 | 105.3197 | 235.2583 | 181.2276 | 233.0614 | 161.6256
2012-10-30 19:53 | 28.6368 | 249.8062 | 195.7168 | 247.4722 | 176.1343 | -1.5968 | 1.2713 | 32.1491 | 40.8709 | 25.2885 | 2.0095 | 32.0724 | 64.4499 | 15.9603 | 14.4727 | -0.1101 | 105.3207 | 235.3335 | 181.2441 | 232.9995 | 161.6616
2012-10-30 19:53 | 28.63 | 249.7759 | 195.7061 | 247.5869 | 175.8699 | -1.6037 | 1.2713 | 32.1495 | 40.8721 | 25.2884 | 2.0095 | 32.0727 | 64.4508 | 15.9604 | 14.4455 | -0.1106 | 105.3229 | 235.3304 | 181.2606 | 233.1414 | 161.4244
2012-10-30 19:53 | 28.6323 | 249.7743 | 195.6769 | 247.597 | 175.7785 | -1.5968 | 1.2713 | 32.1496 | 40.8718 | 25.2888 | 2.0095 | 32.0728 | 64.4515 | 15.9603 | 14.4886 | -0.1101 | 105.3233 | 235.2857 | 181.1883 | 233.1084 | 161.2899
2012-10-30 19:53 | 28.6398 | 249.7662 | 195.6491 | 247.5693 | 175.8289 | -1.6032 | 1.2713 | 32.1493 | 40.8713 | 25.2886 | 2.0095 | 32.0726 | 64.451 | 15.9602 | 14.4757 | -0.1105 | 105.3223 | 235.2905 | 181.1734 | 233.0936 | 161.3532
2012-10-30 19:53 | 28.636 | 249.7937 | 195.6846 | 247.495 | 175.9657 | -1.5933 | 1.2713 | 32.1488 | 40.8704 | 25.2884 | 2.0095 | 32.072 | 64.4481 | 15.9604 | 14.4515 | -0.1098 | 105.3185 | 235.3422 | 181.2331 | 233.0435 | 161.5142
2012-10-30 19:53 | 28.6293 | 249.787 | 195.5954 | 247.594 | 175.7989 | -1.6003 | 1.2713 | 32.1488 | 40.8703 | 25.2883 | 2.0095 | 32.072 | 64.4481 | 15.9604 | 14.4774 | -0.1103 | 105.3184 | 235.3096 | 181.118 | 233.1166 | 161.3215
2012-10-30 19:53 | 28.6353 | 249.8282 | 195.6406 | 247.6627 | 175.8908 | -1.602 | 1.2713 | 32.1485 | 40.8692 | 25.2887 | 2.0095 | 32.0719 | 64.4471 | 15.9605 | 14.4793 | -0.1104 | 105.3163 | 235.3489 | 181.1613 | 233.1834 | 161.4115
2012-10-30 19:53 | 28.6338 | 249.8071 | 195.5959 | 247.5045 | 175.8698 | -1.5985 | 1.2713 | 32.1489 | 40.8707 | 25.2882 | 2.0095 | 32.0722 | 64.4493 | 15.9602 | 14.4778 | -0.1102 | 105.3201 | 235.3293 | 181.1181 | 233.0267 | 161.392
2012-10-30 19:53 | 28.6315 | 249.8205 | 195.5858 | 247.5766 | 175.6801 | -1.6078 | 1.2713 | 32.1487 | 40.8699 | 25.2885 | 2.0095 | 32.072 | 64.4482 | 15.9603 | 14.5 | -0.1108 | 105.3181 | 235.3205 | 181.0858 | 233.0766 | 161.1801
2012-10-30 19:53 | 28.6308 | 249.718 | 195.5929 | 247.7365 | 175.7146 | -1.6003 | 1.2713 | 32.148 | 40.8684 | 25.2883 | 2.0095 | 32.0713 | 64.4457 | 15.9602 | 14.4544 | -0.1103 | 105.3142 | 235.2636 | 181.1385 | 233.2821 | 161.2602
2012-10-30 19:53 | 28.633 | 249.8063 | 195.5755 | 247.5311 | 175.8182 | -1.6009 | 1.2712 | 32.1478 | 40.8676 | 25.2885 | 2.0095 | 32.0711 | 64.4453 | 15.9601 | 14.473 | -0.1104 | 105.3129 | 235.3333 | 181.1025 | 233.0581 | 161.3452
2012-10-30 19:53 | 28.63 | 249.7564 | 195.5253 | 247.5791 | 175.6552 | -1.6026 | 1.2713 | 32.1479 | 40.8681 | 25.2884 | 2.0095 | 32.0712 | 64.4456 | 15.9602 | 14.4536 | -0.1105 | 105.3137 | 235.3028 | 181.0717 | 233.1255 | 161.2016
2012-10-30 19:53 | 28.6383 | 249.8116 | 195.5376 | 247.7636 | 175.7416 | -1.5939 | 1.2712 | 32.1476 | 40.8671 | 25.2886 | 2.0094 | 32.071 | 64.4449 | 15.9601 | 14.466 | -0.1099 | 105.312 | 235.3456 | 181.0716 | 233.2976 | 161.2756
2012-10-30 19:53 | 28.633 | 249.7946 | 195.5323 | 247.6251 | 175.826 | -1.6032 | 1.2712 | 32.1475 | 40.8669 | 25.2884 | 2.0094 | 32.0707 | 64.4436 | 15.9601 | 14.473 | -0.1105 | 105.3105 | 235.3216 | 181.0593 | 233.1521 | 161.353
2012-10-30 19:53 | 28.6315 | 249.8244 | 195.5347 | 247.5375 | 175.9963 | -1.6037 | 1.2712 | 32.1472 | 40.866 | 25.2885 | 2.0094 | 32.0704 | 64.4428 | 15.9601 | 14.4878 | -0.1106 | 105.3088 | 235.3366 | 181.0469 | 233.0497 | 161.5085
2012-10-30 19:53 | 28.6338 | 249.7915 | 195.5409 | 247.6651 | 175.8893 | -1.6061 | 1.2712 | 32.1465 | 40.8645 | 25.2884 | 2.0094 | 32.0698 | 64.4406 | 15.96 | 14.4289 | -0.1107 | 105.3051 | 235.3626 | 181.112 | 233.2362 | 161.4604
2012-10-30 19:53 | 28.624 | 249.813 | 195.5115 | 247.6357 | 175.8405 | -1.6037 | 1.2712 | 32.1467 | 40.8652 | 25.2882 | 2.0094 | 32.0699 | 64.4408 | 15.96 | 14.4884 | -0.1106 | 105.3061 | 235.3246 | 181.0231 | 233.1473 | 161.3521
2012-10-30 19:53 | 28.6338 | 249.7993 | 195.5684 | 247.6612 | 175.8386 | -1.6032 | 1.2712 | 32.1465 | 40.8648 | 25.2882 | 2.0094 | 32.0698 | 64.4402 | 15.96 | 14.4738 | -0.1105 | 105.305 | 235.3255 | 181.0946 | 233.1874 | 161.3648
2012-10-30 19:53 | 28.6383 | 249.7764 | 195.5729 | 247.6265 | 175.968 | -1.6037 | 1.2712 | 32.1456 | 40.862 | 25.2885 | 2.0093 | 32.0689 | 64.4364 | 15.9601 | 14.5191 | -0.1106 | 105.2984 | 235.2573 | 181.0538 | 233.1074 | 161.4489
2012-10-30 19:54 | 28.636 | 249.8367 | 195.547 | 247.589 | 176.0633 | -1.6026 | 1.2711 | 32.1451 | 40.8611 | 25.2884 | 2.0093 | 32.0685 | 64.435 | 15.9601 | 14.4515 | -0.1105 | 105.296 | 235.3852 | 181.0955 | 233.1375 | 161.6118
2012-10-30 19:54 | 28.6383 | 249.749 | 195.5572 | 247.6892 | 176.0773 | -1.602 | 1.2711 | 32.145 | 40.8605 | 25.2885 | 2.0093 | 32.0682 | 64.4335 | 15.9602 | 14.4619 | -0.1104 | 105.294 | 235.2871 | 181.0953 | 233.2273 | 161.6154
2012-10-30 19:54 | 28.6413 | 249.7794 | 195.6624 | 247.6648 | 175.889 | -1.5968 | 1.2711 | 32.1447 | 40.8601 | 25.2883 | 2.0093 | 32.068 | 64.433 | 15.96 | 14.5057 | -0.1101 | 105.2931 | 235.2737 | 181.1567 | 233.1591 | 161.3833
2012-10-30 19:54 | 28.6375 | 249.8147 | 195.5839 | 247.6688 | 175.8149 | -1.598 | 1.2711 | 32.1444 | 40.8596 | 25.2882 | 2.0093 | 32.0677 | 64.4323 | 15.96 | 14.5101 | -0.1102 | 105.2919 | 235.3046 | 181.0738 | 233.1587 | 161.3048
2012-10-30 19:54 | 28.6428 | 249.82 | 195.5892 | 247.6036 | 175.8827 | -1.5974 | 1.2711 | 32.1443 | 40.8589 | 25.2883 | 2.0092 | 32.0674 | 64.4311 | 15.96 | 14.4909 | -0.1101 | 105.2901 | 235.3291 | 181.0983 | 233.1127 | 161.3918
2012-10-30 19:54 | 28.6442 | 249.798 | 195.5906 | 247.7147 | 175.97 | -1.6032 | 1.2711 | 32.144 | 40.8585 | 25.2881 | 2.0092 | 32.0673 | 64.4307 | 15.96 | 14.5086 | -0.1105 | 105.2893 | 235.2894 | 181.082 | 233.2061 | 161.4614
2012-10-30 19:54 | 28.6465 | 249.8042 | 195.6008 | 247.7248 | 175.9957 | -1.5957 | 1.2711 | 32.1441 | 40.8587 | 25.2882 | 2.0092 | 32.0673 | 64.431 | 15.9599 | 14.5028 | -0.11 | 105.2898 | 235.3014 | 181.098 | 233.222 | 161.4929
2012-10-30 19:54 | 28.639 | 249.7888 | 195.5775 | 247.6625 | 175.8828 | -1.6032 | 1.2711 | 32.1439 | 40.858 | 25.2883 | 2.0092 | 32.0671 | 64.4303 | 15.9599 | 14.4382 | -0.1105 | 105.2883 | 235.3506 | 181.1393 | 233.2243 | 161.4446
2012-10-30 19:54 | 28.6442 | 249.8254 | 195.6338 | 247.5815 | 175.888 | -1.6061 | 1.2711 | 32.1435 | 40.8566 | 25.2886 | 2.0092 | 32.0668 | 64.4289 | 15.9599 | 14.4597 | -0.1107 | 105.2855 | 235.3657 | 181.1741 | 233.1218 | 161.4283
2012-10-30 19:54 | 28.6413 | 249.842 | 195.6113 | 247.6491 | 175.9475 | -1.5968 | 1.2711 | 32.143 | 40.8558 | 25.2882 | 2.0092 | 32.0663 | 64.4269 | 15.9599 | 14.4894 | -0.1101 | 105.2826 | 235.3526 | 181.1219 | 233.1597 | 161.4581
2012-10-30 19:54 | 28.6383 | 249.7725 | 195.6201 | 247.7832 | 176.0109 | -1.6061 | 1.271 | 32.1423 | 40.8537 | 25.2885 | 2.0091 | 32.0655 | 64.4239 | 15.9599 | 14.4619 | -0.1107 | 105.2776 | 235.3106 | 181.1582 | 233.3213 | 161.549
2012-10-30 19:54 | 28.633 | 249.8259 | 195.6659 | 247.6134 | 176.0954 | -1.6026 | 1.271 | 32.1417 | 40.8526 | 25.2881 | 2.0091 | 32.065 | 64.4218 | 15.9599 | 14.4566 | -0.1105 | 105.2744 | 235.3693 | 181.2093 | 233.1568 | 161.6388
2012-10-30 19:54 | 28.6375 | 249.8226 | 195.6153 | 247.7158 | 176.1155 | -1.598 | 1.271 | 32.1414 | 40.8518 | 25.2883 | 2.0091 | 32.0647 | 64.4212 | 15.9597 | 14.4571 | -0.1102 | 105.2731 | 235.3655 | 181.1582 | 233.2587 | 161.6584
2012-10-30 19:54 | 28.642 | 249.8349 | 195.6473 | 247.6341 | 176.0497 | -1.6049 | 1.271 | 32.1407 | 40.8497 | 25.2885 | 2.009 | 32.064 | 64.418 | 15.9598 | 14.4656 | -0.1106 | 105.2676 | 235.3693 | 181.1817 | 233.1685 | 161.5841
2012-10-30 19:54 | 28.6495 | 249.8385 | 195.6745 | 247.7396 | 175.8933 | -1.6026 | 1.2655 | 32.0012 | 40.4979 | 25.2871 | 2 | 31.9201 | 63.8404 | 15.96 | 14.4895 | -0.1105 | 104.3383 | 235.349 | 181.185 | 233.2501 | 161.4038
2012-10-30 19:54 | 28.651 | 249.7969 | 195.6013 | 247.745 | 175.9299 | -1.6009 | 1.2648 | 31.9827 | 40.452 | 25.2866 | 1.9992 | 31.9065 | 63.7863 | 15.96 | 14.4624 | -0.1104 | 104.2382 | 235.3345 | 181.1389 | 233.2826 | 161.4675
2012-10-30 19:54 | 28.6472 | 249.7814 | 195.6211 | 247.6785 | 176.0198 | -1.5974 | 1.2642 | 31.9669 | 40.412 | 25.2866 | 1.9982 | 31.8905 | 63.7221 | 15.96 | 14.4912 | -0.1101 | 104.1341 | 235.2902 | 181.1299 | 233.1873 | 161.5286
2012-10-30 19:54 | 28.648 | 249.7783 | 195.5905 | 247.6166 | 175.9894 | -1.5933 | 1.2639 | 31.9584 | 40.3908 | 25.2864 | 1.9976 | 31.8821 | 63.6882 | 15.96 | 14.4961 | -0.1098 | 104.079 | 235.2822 | 181.0944 | 233.1205 | 161.4933
2012-10-30 19:54 | 28.6435 | 249.7268 | 195.5742 | 247.6356 | 176.1332 | -1.6014 | 1.2637 | 31.9537 | 40.3792 | 25.2863 | 1.9973 | 31.8776 | 63.6701 | 15.9601 | 14.4467 | -0.1104 | 104.0493 | 235.2801 | 181.1275 | 233.1889 | 161.6865
2012-10-30 19:54 | 28.6413 | 249.7089 | 195.5916 | 247.8097 | 176.1076 | -1.5985 | 1.2637 | 31.9539 | 40.3795 | 25.2864 | 1.9973 | 31.8777 | 63.6703 | 15.9601 | 14.4609 | -0.1102 | 104.0499 | 235.248 | 181.1307 | 233.3488 | 161.6467
2012-10-30 19:54 | 28.6413 | 249.7089 | 195.6309 | 247.6413 | 176.252 | -1.5997 | 1.2637 | 31.9546 | 40.3811 | 25.2865 | 1.9974 | 31.8783 | 63.6726 | 15.9602 | 14.4527 | -0.1103 | 104.0537 | 235.2562 | 181.1782 | 233.1886 | 161.7993
2012-10-30 19:54 | 28.6457 | 249.7368 | 195.6235 | 247.63 | 176.2447 | -1.5985 | 1.2638 | 31.956 | 40.3848 | 25.2864 | 1.9974 | 31.8797 | 63.6778 | 15.9603 | 14.4449 | -0.1102 | 104.0626 | 235.2919 | 181.1786 | 233.1851 | 161.7998
2012-10-30 19:54 | 28.639 | 249.6988 | 195.6365 | 247.592 | 176.1717 | -1.5933 | 1.2638 | 31.9575 | 40.3884 | 25.2865 | 1.9975 | 31.8813 | 63.6839 | 15.9603 | 14.4096 | -0.1098 | 104.0723 | 235.2892 | 181.2269 | 233.1824 | 161.7621
2012-10-30 19:55 | 28.642 | 249.6861 | 195.6473 | 247.6968 | 175.7531 | -1.6032 | 1.2639 | 31.9589 | 40.3916 | 25.2867 | 1.9976 | 31.8826 | 63.6892 | 15.9603 | 14.4412 | -0.1105 | 104.0808 | 235.2449 | 181.2061 | 233.2556 | 161.3119
2012-10-30 19:55 | 28.639 | 249.6792 | 195.6522 | 247.6351 | 176.0819 | -1.5957 | 1.2682 | 32.0691 | 40.6698 | 25.2872 | 2.0043 | 31.9902 | 64.1193 | 15.9605 | 14.4708 | -0.11 | 104.7891 | 235.2084 | 181.1814 | 233.1643 | 161.6111
2012-10-30 19:55 | 28.6398 | 249.6839 | 195.6962 | 247.6084 | 176.1725 | -1.602 | 1.2693 | 32.0971 | 40.7397 | 25.2881 | 2.0063 | 32.0205 | 64.2412 | 15.9604 | 14.443 | -0.1104 | 104.9808 | 235.2409 | 181.2532 | 233.1654 | 161.7295
2012-10-30 19:55 | 28.6375 | 249.6777 | 195.6782 | 247.5591 | 176.2678 | -1.5997 | 1.2699 | 32.1134 | 40.7804 | 25.2884 | 2.0073 | 32.0367 | 64.3064 | 15.9603 | 14.4448 | -0.1103 | 105.0868 | 235.2329 | 181.2334 | 233.1143 | 161.823
2012-10-30 19:55 | 28.6338 | 249.7406 | 195.7059 | 247.6769 | 176.2289 | -1.5962 | 1.2704 | 32.1254 | 40.8105 | 25.2885 | 2.008 | 32.0486 | 64.3538 | 15.9604 | 14.4493 | -0.11 | 105.1643 | 235.2913 | 181.2566 | 233.2276 | 161.7796
2012-10-30 19:55 | 28.6353 | 249.6912 | 195.7978 | 247.5687 | 176.371 | -1.6003 | 1.2707 | 32.1327 | 40.8295 | 25.2883 | 2.0085 | 32.056 | 64.3841 | 15.9603 | 14.3896 | -0.1103 | 105.2136 | 235.3016 | 181.4082 | 233.1791 | 161.9814
2012-10-30 19:55 | 28.6225 | 249.7684 | 195.7536 | 247.4579 | 176.3582 | -1.6061 | 1.2708 | 32.1376 | 40.8418 | 25.2885 | 2.0088 | 32.0609 | 64.404 | 15.9602 | 14.3849 | -0.1107 | 105.2458 | 235.3835 | 181.3687 | 233.073 | 161.9733
2012-10-30 19:55 | 28.63 | 249.7407 | 195.7572 | 247.5673 | 176.5453 | -1.5991 | 1.271 | 32.1409 | 40.8504 | 25.2883 | 2.009 | 32.0642 | 64.4178 | 15.9601 | 14.4169 | -0.1102 | 105.2683 | 235.3238 | 181.3403 | 233.1504 | 162.1284
2012-10-30 19:55 | 28.6315 | 249.7383 | 195.8255 | 247.6824 | 176.3984 | -1.6026 | 1.2711 | 32.1431 | 40.856 | 25.2882 | 2.0091 | 32.0663 | 64.4256 | 15.9602 | 14.4103 | -0.1105 | 105.2816 | 235.328 | 181.4152 | 233.2721 | 161.9881
2012-10-30 19:55 | 28.6308 | 249.7532 | 195.8798 | 247.5877 | 176.4094 | -1.6003 | 1.2711 | 32.1445 | 40.8588 | 25.2887 | 2.0092 | 32.0676 | 64.431 | 15.9602 | 14.4136 | -0.1103 | 105.2898 | 235.3396 | 181.4662 | 233.1741 | 161.9958
2012-10-30 19:55 | 28.63 | 249.7759 | 195.8436 | 247.5046 | 176.2173 | -1.5991 | 1.2712 | 32.1459 | 40.8627 | 25.2885 | 2.0093 | 32.0691 | 64.4373 | 15.9601 | 14.4618 | -0.1102 | 105.3 | 235.3141 | 181.3818 | 233.0428 | 161.7555
2012-10-30 19:55 | 28.6255 | 249.7127 | 195.8234 | 247.6607 | 176.1074 | -1.5997 | 1.2712 | 32.1463 | 40.8639 | 25.2885 | 2.0094 | 32.0696 | 64.4392 | 15.9602 | 14.4451 | -0.1103 | 105.3031 | 235.2676 | 181.3783 | 233.2156 | 161.6623
2012-10-30 19:55 | 28.6278 | 249.7894 | 195.8296 | 247.6356 | 176.1058 | -1.6009 | 1.2712 | 32.1473 | 40.8667 | 25.2883 | 2.0094 | 32.0704 | 64.4433 | 15.96 | 14.4433 | -0.1104 | 105.3101 | 235.3461 | 181.3863 | 233.1923 | 161.6625
2012-10-30 19:55 | 28.6353 | 249.8204 | 195.845 | 247.5844 | 176.1484 | -1.6043 | 1.2712 | 32.1478 | 40.8676 | 25.2885 | 2.0095 | 32.0709 | 64.4449 | 15.96 | 14.4957 | -0.1106 | 105.3125 | 235.3247 | 181.3493 | 233.0887 | 161.6527
2012-10-30 19:55 | 28.6293 | 249.787 | 195.7918 | 247.5313 | 176.1776 | -1.6055 | 1.2712 | 32.148 | 40.8682 | 25.2885 | 2.0095 | 32.0711 | 64.4455 | 15.9601 | 14.4203 | -0.1107 | 105.3137 | 235.3667 | 181.3715 | 233.111 | 161.7573
2012-10-30 19:55 | 28.6135 | 249.779 | 195.7917 | 247.6683 | 175.99 | -1.6037 | 1.2713 | 32.1483 | 40.8687 | 25.2887 | 2.0095 | 32.0714 | 64.4469 | 15.9601 | 14.4779 | -0.1106 | 105.3156 | 235.3011 | 181.3138 | 233.1904 | 161.5121
2012-10-30 19:55 | 28.6353 | 249.7617 | 195.7467 | 247.5687 | 175.9611 | -1.6049 | 1.2713 | 32.1488 | 40.87 | 25.2886 | 2.0095 | 32.072 | 64.4504 | 15.9598 | 14.4549 | -0.1106 | 105.3204 | 235.3068 | 181.2918 | 233.1138 | 161.5062
2012-10-30 19:55 | 28.6345 | 249.7961 | 195.7577 | 247.6071 | 175.972 | -1.6003 | 1.2713 | 32.1488 | 40.8701 | 25.2886 | 2.0095 | 32.072 | 64.4498 | 15.96 | 14.45 | -0.1103 | 105.3198 | 235.3461 | 181.3077 | 233.1571 | 161.522
2012-10-30 19:55 | 28.639 | 249.7849 | 195.7347 | 247.5371 | 176.0389 | -1.5991 | 1.2713 | 32.1487 | 40.8701 | 25.2885 | 2.0095 | 32.0719 | 64.449 | 15.96 | 14.4545 | -0.1102 | 105.319 | 235.3304 | 181.2802 | 233.0826 | 161.5844
2012-10-30 19:55 | 28.63 | 249.819 | 195.6982 | 247.5712 | 175.9597 | -1.5974 | 1.2713 | 32.1487 | 40.8694 | 25.2888 | 2.0095 | 32.0719 | 64.4495 | 15.9599 | 14.4414 | -0.1101 | 105.3189 | 235.3776 | 181.2568 | 233.1298 | 161.5183
2012-10-30 19:55 | 28.6345 | 249.7726 | 195.7066 | 247.6501 | 176.0227 | -1.5997 | 1.2713 | 32.1487 | 40.8699 | 25.2885 | 2.0095 | 32.0719 | 64.4495 | 15.9599 | 14.4704 | -0.1103 | 105.3194 | 235.3022 | 181.2362 | 233.1797 | 161.5523
2012-10-30 19:55 | 28.639 | 249.7693 | 195.7111 | 247.6429 | 176.078 | -1.598 | 1.2713 | 32.1489 | 40.87 | 25.2887 | 2.0095 | 32.072 | 64.45 | 15.9599 | 14.479 | -0.1102 | 105.32 | 235.2903 | 181.2321 | 233.1639 | 161.599
2012-10-30 19:55 | 28.6308 | 249.7728 | 195.6794 | 247.6034 | 176.1322 | -1.5974 | 1.2713 | 32.1488 | 40.87 | 25.2886 | 2.0095 | 32.072 | 64.4503 | 15.9598 | 14.4871 | -0.1101 | 105.3203 | 235.2857 | 181.1923 | 233.1163 | 161.6451
2012-10-30 19:55 | 28.6338 | 249.8189 | 195.7335 | 247.5868 | 176.1157 | -1.6014 | 1.2713 | 32.1484 | 40.8691 | 25.2886 | 2.0095 | 32.0717 | 64.4494 | 15.9597 | 14.5023 | -0.1104 | 105.3185 | 235.3166 | 181.2312 | 233.0845 | 161.6134
2012-10-30 19:55 | 28.6285 | 249.8175 | 195.736 | 247.6011 | 176.1221 | -1.602 | 1.2713 | 32.1486 | 40.8696 | 25.2885 | 2.0095 | 32.0718 | 64.4495 | 15.9598 | 14.4889 | -0.1104 | 105.3191 | 235.3286 | 181.2471 | 233.1122 | 161.6332
2012-10-30 19:56 | 28.6345 | 249.8 | 195.742 | 247.5914 | 176.1203 | -1.6014 | 1.2713 | 32.1485 | 40.8692 | 25.2887 | 2.0095 | 32.0717 | 64.4495 | 15.9597 | 14.4704 | -0.1104 | 105.3188 | 235.3296 | 181.2716 | 233.121 | 161.6499
2012-10-30 19:56 | 28.6323 | 249.8174 | 195.728 | 247.6244 | 175.9346 | -1.6037 | 1.2713 | 32.1484 | 40.8695 | 25.2882 | 2.0095 | 32.0716 | 64.4487 | 15.9597 | 14.4804 | -0.1106 | 105.3182 | 235.337 | 181.2476 | 233.144 | 161.4542
2012-10-30 19:56 | 28.6248 | 249.7786 | 195.693 | 247.7266 | 176.0599 | -1.6009 | 1.2712 | 32.1475 | 40.8665 | 25.2886 | 2.0095 | 32.0706 | 64.4451 | 15.9597 | 14.4484 | -0.1104 | 105.3116 | 235.3302 | 181.2446 | 233.2782 | 161.6115
2012-10-30 19:56 | 28.6285 | 249.8018 | 195.6928 | 247.5619 | 175.8996 | -1.6014 | 1.2712 | 32.1474 | 40.8672 | 25.2882 | 2.0095 | 32.0707 | 64.4453 | 15.9597 | 14.5174 | -0.1104 | 105.3125 | 235.2844 | 181.1754 | 233.0445 | 161.3822
2012-10-30 19:56 | 28.6315 | 249.844 | 195.6447 | 247.5766 | 175.848 | -1.5991 | 1.2712 | 32.1471 | 40.8662 | 25.2883 | 2.0095 | 32.0704 | 64.4443 | 15.9597 | 14.4715 | -0.1102 | 105.3105 | 235.3725 | 181.1732 | 233.1051 | 161.3765
2012-10-30 19:56 | 28.63 | 249.8307 | 195.6707 | 247.63 | 175.7918 | -1.6026 | 1.2712 | 32.1468 | 40.8654 | 25.2883 | 2.0094 | 32.0701 | 64.4433 | 15.9597 | 14.4822 | -0.1105 | 105.3087 | 235.3485 | 181.1885 | 233.1478 | 161.3096
2012-10-30 19:56 | 28.6383 | 249.8273 | 195.6633 | 247.6539 | 175.8821 | -1.5991 | 1.2712 | 32.1465 | 40.8643 | 25.2885 | 2.0094 | 32.0697 | 64.4416 | 15.9597 | 14.4252 | -0.1102 | 105.3058 | 235.4021 | 181.2381 | 233.2287 | 161.4569
2012-10-30 19:56 | 28.636 | 249.8367 | 195.6531 | 247.7222 | 175.9696 | -1.5985 | 1.2712 | 32.1465 | 40.8645 | 25.2884 | 2.0094 | 32.0697 | 64.4416 | 15.9596 | 14.48 | -0.1102 | 105.3061 | 235.3567 | 181.1731 | 233.2422 | 161.4896
2012-10-30 19:56 | 28.6315 | 249.8479 | 195.6211 | 247.6354 | 175.8597 | -1.6003 | 1.2712 | 32.1465 | 40.8643 | 25.2885 | 2.0094 | 32.0697 | 64.4417 | 15.9596 | 14.4674 | -0.1103 | 105.3059 | 235.3805 | 181.1537 | 233.168 | 161.3923
2012-10-30 19:56 | 28.636 | 249.8054 | 195.5903 | 247.6477 | 175.9071 | -1.5991 | 1.2712 | 32.1466 | 40.8652 | 25.2881 | 2.0094 | 32.0698 | 64.4426 | 15.9595 | 14.4882 | -0.1102 | 105.3078 | 235.3172 | 181.1021 | 233.1595 | 161.4189
2012-10-30 19:56 | 28.6383 | 249.8038 | 195.5454 | 247.6422 | 175.886 | -1.6014 | 1.2712 | 32.1463 | 40.8642 | 25.2883 | 2.0094 | 32.0696 | 64.4417 | 15.9596 | 14.4701 | -0.1104 | 105.3059 | 235.3337 | 181.0753 | 233.1721 | 161.4159
2012-10-30 19:56 | 28.6413 | 249.7716 | 195.5838 | 247.7 | 175.9944 | -1.6014 | 1.2712 | 32.1461 | 40.8636 | 25.2883 | 2.0094 | 32.0693 | 64.441 | 15.9594 | 14.4323 | -0.1104 | 105.3047 | 235.3393 | 181.1515 | 233.2677 | 161.5621
2012-10-30 19:56 | 28.6398 | 249.8288 | 195.6019 | 247.6006 | 175.9539 | -1.6032 | 1.2712 | 32.1455 | 40.8621 | 25.2883 | 2.0094 | 32.0688 | 64.4387 | 15.9595 | 14.4308 | -0.1105 | 105.3008 | 235.398 | 181.1711 | 233.1698 | 161.5231
2012-10-30 19:56 | 28.6525 | 249.8102 | 195.6264 | 247.6407 | 175.9509 | -1.6043 | 1.2712 | 32.145 | 40.8613 | 25.288 | 2.0094 | 32.0683 | 64.4366 | 15.9595 | 14.4476 | -0.1106 | 105.2979 | 235.3626 | 181.1788 | 233.1931 | 161.5033
2012-10-30 19:56 | 28.657 | 249.846 | 195.6073 | 247.5825 | 175.7993 | -1.5991 | 1.2711 | 32.145 | 40.8607 | 25.2884 | 2.0094 | 32.0683 | 64.4367 | 15.9595 | 14.4398 | -0.1102 | 105.2974 | 235.4062 | 181.1675 | 233.1427 | 161.3595
2012-10-30 19:56 | 28.6517 | 249.8133 | 195.5745 | 247.7339 | 175.7862 | -1.6049 | 1.2711 | 32.1448 | 40.8606 | 25.2881 | 2.0093 | 32.068 | 64.4351 | 15.9596 | 14.4427 | -0.1106 | 105.2958 | 235.3706 | 181.1318 | 233.2912 | 161.3435
2012-10-30 19:56 | 28.6562 | 249.8021 | 195.5555 | 247.6718 | 175.8727 | -1.6003 | 1.2711 | 32.1443 | 40.8587 | 25.2885 | 2.0093 | 32.0675 | 64.4339 | 15.9594 | 14.4309 | -0.1103 | 105.2926 | 235.3712 | 181.1246 | 233.2409 | 161.4418
2012-10-30 19:56 | 28.654 | 249.8469 | 195.5808 | 247.6031 | 175.8275 | -1.602 | 1.2711 | 32.1439 | 40.8581 | 25.2883 | 2.0093 | 32.0672 | 64.4321 | 15.9595 | 14.4491 | -0.1104 | 105.2902 | 235.3978 | 181.1317 | 233.154 | 161.3784
2012-10-30 19:56 | 28.6592 | 249.8364 | 195.5702 | 247.6435 | 175.8561 | -1.5968 | 1.2711 | 32.1434 | 40.8568 | 25.2883 | 2.0093 | 32.0668 | 64.4309 | 15.9594 | 14.4053 | -0.1101 | 105.2877 | 235.4311 | 181.1649 | 233.2382 | 161.4508
2012-10-30 19:56 | 28.66 | 249.8216 | 195.5868 | 247.6521 | 175.9819 | -1.598 | 1.2711 | 32.1427 | 40.8551 | 25.2882 | 2.0092 | 32.0659 | 64.4273 | 15.9594 | 14.4714 | -0.1102 | 105.2824 | 235.3502 | 181.1154 | 233.1807 | 161.5105
2012-10-30 19:56 | 28.66 | 249.8177 | 195.5475 | 247.7266 | 175.8296 | -1.6032 | 1.271 | 32.1426 | 40.8547 | 25.2884 | 2.0092 | 32.0658 | 64.4271 | 15.9594 | 14.4551 | -0.1105 | 105.2818 | 235.3626 | 181.0924 | 233.2715 | 161.3745
2012-10-30 19:56 | 28.66 | 249.8372 | 195.5671 | 247.7148 | 175.7984 | -1.5991 | 1.271 | 32.1423 | 40.8542 | 25.2882 | 2.0092 | 32.0656 | 64.4259 | 15.9594 | 14.4469 | -0.1102 | 105.2801 | 235.3903 | 181.1202 | 233.2679 | 161.3515
2012-10-30 19:57 | 28.6675 | 249.8291 | 195.606 | 247.6401 | 175.7903 | -1.602 | 1.271 | 32.1421 | 40.8535 | 25.2882 | 2.0092 | 32.0653 | 64.4248 | 15.9594 | 14.4544 | -0.1104 | 105.2784 | 235.3747 | 181.1516 | 233.1857 | 161.3359
2012-10-30 19:57 | 28.666 | 249.8667 | 195.5613 | 247.6072 | 175.8044 | -1.598 | 1.271 | 32.1419 | 40.8531 | 25.2882 | 2.0092 | 32.0651 | 64.4244 | 15.9593 | 14.4815 | -0.1102 | 105.2775 | 235.3852 | 181.0798 | 233.1257 | 161.3229
2012-10-30 19:57 | 28.66 | 249.8138 | 195.4807 | 247.7109 | 175.8569 | -1.602 | 1.271 | 32.1421 | 40.8538 | 25.2881 | 2.0092 | 32.0653 | 64.426 | 15.9592 | 14.4347 | -0.1104 | 105.2798 | 235.3791 | 181.046 | 233.2762 | 161.4222
2012-10-30 19:57 | 28.666 | 249.8511 | 195.581 | 247.7051 | 175.9957 | -1.6014 | 1.2704 | 32.1272 | 40.8158 | 25.2883 | 2.0085 | 32.0545 | 64.382 | 15.9593 | 14.4856 | -0.1104 | 105.1977 | 235.3655 | 181.0954 | 233.2195 | 161.5101
2012-10-30 19:57 | 28.6795 | 249.8411 | 195.5591 | 247.6011 | 175.9389 | -1.5939 | 1.2646 | 31.9786 | 40.4415 | 25.2867 | 1.9989 | 31.9014 | 63.7685 | 15.9593 | 14.4746 | -0.1099 | 104.2101 | 235.3665 | 181.0845 | 233.1265 | 161.4643
2012-10-30 19:57 | 28.6772 | 249.8192 | 195.5765 | 247.6968 | 175.9288 | -1.5985 | 1.2643 | 31.971 | 40.4224 | 25.2865 | 1.9985 | 31.8945 | 63.7406 | 15.9594 | 14.4682 | -0.1102 | 104.1631 | 235.351 | 181.1083 | 233.2286 | 161.4606
2012-10-30 19:57 | 28.6862 | 249.7538 | 195.5304 | 247.7567 | 175.7621 | -1.5974 | 1.2639 | 31.9591 | 40.3928 | 25.2863 | 1.9977 | 31.8829 | 63.6936 | 15.9595 | 14.4731 | -0.1101 | 104.0864 | 235.2807 | 181.0573 | 233.2836 | 161.289
2012-10-30 19:57 | 28.6824 | 249.7657 | 195.5384 | 247.7255 | 176.0901 | -1.6049 | 1.2637 | 31.9537 | 40.3792 | 25.2862 | 1.9974 | 31.8775 | 63.6718 | 15.9595 | 14.502 | -0.1106 | 104.051 | 235.2637 | 181.0364 | 233.2235 | 161.5881
2012-10-30 19:57 | 28.6795 | 249.7667 | 195.5473 | 247.5855 | 175.8062 | -1.5991 | 1.2636 | 31.9516 | 40.3736 | 25.2864 | 1.9973 | 31.8754 | 63.6633 | 15.9596 | 14.4868 | -0.1102 | 104.0369 | 235.2799 | 181.0605 | 233.0987 | 161.3194
2012-10-30 19:57 | 28.6884 | 249.7286 | 195.5287 | 247.708 | 175.7253 | -1.6003 | 1.2636 | 31.9526 | 40.3764 | 25.2863 | 1.9973 | 31.8763 | 63.6671 | 15.9596 | 14.4835 | -0.1103 | 104.0434 | 235.2451 | 181.0452 | 233.2245 | 161.2418
2012-10-30 19:57 | 28.6772 | 249.7135 | 195.5136 | 247.7359 | 175.8039 | -1.5985 | 1.2637 | 31.9538 | 40.379 | 25.2864 | 1.9974 | 31.8775 | 63.6711 | 15.9598 | 14.4641 | -0.1102 | 104.0501 | 235.2494 | 181.0495 | 233.2718 | 161.3398
2012-10-30 19:57 | 28.6854 | 249.7726 | 195.5768 | 247.5404 | 175.9253 | -1.6014 | 1.2638 | 31.9563 | 40.3852 | 25.2866 | 1.9975 | 31.88 | 63.681 | 15.9598 | 14.4927 | -0.1104 | 104.0662 | 235.2799 | 181.0841 | 233.0477 | 161.4326
2012-10-30 19:57 | 28.6899 | 249.7106 | 195.487 | 247.6429 | 175.8322 | -1.6049 | 1.2638 | 31.958 | 40.3898 | 25.2864 | 1.9976 | 31.8818 | 63.6878 | 15.9598 | 14.4931 | -0.1106 | 104.0776 | 235.2175 | 180.9939 | 233.1498 | 161.3391
2012-10-30 19:57 | 28.6862 | 249.6677 | 195.5344 | 247.694 | 176.1213 | -1.6043 | 1.2665 | 32.0254 | 40.5595 | 25.2869 | 2.0019 | 31.9508 | 63.9639 | 15.9598 | 14.4609 | -0.1106 | 104.5233 | 235.2068 | 181.0735 | 233.2331 | 161.6604
2012-10-30 19:57 | 28.6959 | 249.6813 | 195.5834 | 247.6058 | 176.0607 | -1.5991 | 1.269 | 32.0907 | 40.724 | 25.2876 | 2.0059 | 32.014 | 64.2176 | 15.9598 | 14.4991 | -0.1102 | 104.9416 | 235.1822 | 181.0843 | 233.1067 | 161.5616
2012-10-30 19:57 | 28.6862 | 249.7029 | 195.6287 | 247.5608 | 176.2032 | -1.6055 | 1.2698 | 32.1103 | 40.7731 | 25.288 | 2.0072 | 32.0336 | 64.2963 | 15.9597 | 14.4935 | -0.1107 | 105.0694 | 235.2094 | 181.1352 | 233.0673 | 161.7097
2012-10-30 19:57 | 28.6884 | 249.6621 | 195.6309 | 247.5904 | 176.1703 | -1.6037 | 1.2703 | 32.1231 | 40.8052 | 25.2883 | 2.0079 | 32.0463 | 64.3472 | 15.9597 | 14.4549 | -0.1106 | 105.1524 | 235.2072 | 181.176 | 233.1355 | 161.7154
2012-10-30 19:57 | 28.6847 | 249.6623 | 195.6586 | 247.6533 | 176.112 | -1.6026 | 1.2706 | 32.131 | 40.8253 | 25.2883 | 2.0084 | 32.0541 | 64.3789 | 15.9597 | 14.4675 | -0.1105 | 105.2042 | 235.1948 | 181.1911 | 233.1858 | 161.6445
2012-10-30 19:57 | 28.6907 | 249.6839 | 195.6607 | 247.5967 | 176.0438 | -1.6003 | 1.2708 | 32.1365 | 40.8394 | 25.2883 | 2.0088 | 32.0598 | 64.4022 | 15.9596 | 14.498 | -0.1103 | 105.2416 | 235.1859 | 181.1627 | 233.0987 | 161.5458
2012-10-30 19:57 | 28.6832 | 249.7078 | 195.6493 | 247.4873 | 176.1807 | -1.6078 | 1.2709 | 32.1398 | 40.8474 | 25.2884 | 2.009 | 32.0631 | 64.4149 | 15.9597 | 14.4783 | -0.1108 | 105.2623 | 235.2295 | 181.171 | 233.009 | 161.7024
2012-10-30 19:57 | 28.6914 | 249.716 | 195.6614 | 247.5112 | 176.228 | -1.6049 | 1.271 | 32.1423 | 40.8536 | 25.2885 | 2.0092 | 32.0655 | 64.425 | 15.9596 | 14.4783 | -0.1106 | 105.2786 | 235.2377 | 181.1831 | 233.0329 | 161.7497
2012-10-30 19:57 | 28.6862 | 249.6638 | 195.6758 | 247.647 | 176.1291 | -1.6032 | 1.2711 | 32.144 | 40.8578 | 25.2886 | 2.0093 | 32.0673 | 64.4319 | 15.9596 | 14.4486 | -0.1105 | 105.2897 | 235.2152 | 181.2272 | 233.1984 | 161.6805
2012-10-30 19:57 | 28.6832 | 249.6843 | 195.6689 | 247.5617 | 176.2198 | -1.6009 | 1.2711 | 32.1451 | 40.8608 | 25.2885 | 2.0094 | 32.0684 | 64.437 | 15.9595 | 14.4497 | -0.1104 | 105.2978 | 235.2346 | 181.2192 | 233.112 | 161.7701
2012-10-30 19:57 | 28.681 | 249.7212 | 195.706 | 247.5635 | 176.1668 | -1.6037 | 1.2712 | 32.1459 | 40.8628 | 25.2885 | 2.0094 | 32.0692 | 64.4399 | 15.9595 | 14.4638 | -0.1106 | 105.3027 | 235.2574 | 181.2422 | 233.0997 | 161.703
2012-10-30 19:58 | 28.6787 | 249.7228 | 195.7037 | 247.5416 | 176.2973 | -1.6049 | 1.2712 | 32.1462 | 40.8637 | 25.2885 | 2.0094 | 32.0694 | 64.4404 | 15.9596 | 14.5064 | -0.1106 | 105.3041 | 235.2164 | 181.1973 | 233.0352 | 161.7909
2012-10-30 19:58 | 28.6802 | 249.7204 | 195.6973 | 247.5548 | 176.209 | -1.602 | 1.2712 | 32.1468 | 40.8647 | 25.2887 | 2.0095 | 32.0699 | 64.4431 | 15.9595 | 14.463 | -0.1104 | 105.3078 | 235.2574 | 181.2343 | 233.0918 | 161.746
2012-10-30 19:58 | 28.6772 | 249.7135 | 195.6747 | 247.6537 | 176.1357 | -1.5968 | 1.2712 | 32.1473 | 40.8659 | 25.2887 | 2.0095 | 32.0704 | 64.4452 | 15.9595 | 14.4764 | -0.1101 | 105.3111 | 235.2371 | 181.1983 | 233.1773 | 161.6593
2012-10-30 19:58 | 28.6802 | 249.7204 | 195.7209 | 247.5862 | 176.1543 | -1.5985 | 1.2712 | 32.1477 | 40.8675 | 25.2885 | 2.0095 | 32.0709 | 64.4475 | 15.9594 | 14.4875 | -0.1102 | 105.3149 | 235.2329 | 181.2334 | 233.0987 | 161.6668
2012-10-30 19:58 | 28.6757 | 249.7238 | 195.7125 | 247.6052 | 175.9468 | -1.6009 | 1.2712 | 32.148 | 40.8681 | 25.2885 | 2.0095 | 32.0711 | 64.4483 | 15.9594 | 14.483 | -0.1104 | 105.3164 | 235.2408 | 181.2295 | 233.1222 | 161.4638
2012-10-30 19:58 | 28.6772 | 249.7566 | 195.7533 | 247.544 | 175.9522 | -1.6072 | 1.2712 | 32.1477 | 40.8672 | 25.2885 | 2.0095 | 32.0709 | 64.4473 | 15.9595 | 14.5172 | -0.1108 | 105.3145 | 235.2394 | 181.2361 | 233.0268 | 161.435
2012-10-30 19:58 | 28.672 | 249.7122 | 195.6773 | 247.5114 | 175.9743 | -1.6003 | 1.2712 | 32.1473 | 40.8665 | 25.2884 | 2.0095 | 32.0706 | 64.4458 | 15.9595 | 14.4548 | -0.1103 | 105.3122 | 235.2574 | 181.2225 | 233.0566 | 161.5195
2012-10-30 19:58 | 28.6839 | 249.7672 | 195.6814 | 247.6055 | 175.8691 | -1.5997 | 1.2712 | 32.1473 | 40.8665 | 25.2885 | 2.0095 | 32.0704 | 64.4457 | 15.9594 | 14.532 | -0.1103 | 105.3122 | 235.2352 | 181.1494 | 233.0735 | 161.3371
2012-10-30 19:58 | 28.6802 | 249.7517 | 195.6502 | 247.641 | 175.7718 | -1.602 | 1.2712 | 32.1473 | 40.8665 | 25.2885 | 2.0095 | 32.0705 | 64.4458 | 15.9594 | 14.4386 | -0.1104 | 105.3123 | 235.3131 | 181.2116 | 233.2024 | 161.3332
2012-10-30 19:58 | 28.672 | 249.6887 | 195.6105 | 247.6093 | 175.9782 | -1.6009 | 1.2712 | 32.1472 | 40.8666 | 25.2883 | 2.0095 | 32.0704 | 64.4462 | 15.9593 | 14.4997 | -0.1104 | 105.3128 | 235.189 | 181.1108 | 233.1096 | 161.4785
2012-10-30 19:58 | 28.6735 | 249.7098 | 195.6238 | 247.6304 | 176.132 | -1.6049 | 1.2712 | 32.1472 | 40.8659 | 25.2885 | 2.0095 | 32.0703 | 64.446 | 15.9592 | 14.4482 | -0.1106 | 105.312 | 235.2616 | 181.1756 | 233.1822 | 161.6838
2012-10-30 19:58 | 28.6705 | 249.7146 | 195.6483 | 247.5765 | 176.0158 | -1.6009 | 1.2712 | 32.1473 | 40.8666 | 25.2884 | 2.0095 | 32.0706 | 64.4472 | 15.9591 | 14.5023 | -0.1104 | 105.3138 | 235.2123 | 181.146 | 233.0742 | 161.5135
2012-10-30 19:58 | 28.6765 | 249.7246 | 195.6386 | 247.6216 | 175.9359 | -1.6078 | 1.2712 | 32.1478 | 40.8673 | 25.2886 | 2.0095 | 32.0709 | 64.4481 | 15.9592 | 14.4634 | -0.1108 | 105.3154 | 235.2612 | 181.1752 | 233.1582 | 161.4725
2012-10-30 19:58 | 28.6675 | 249.8017 | 195.6139 | 247.593 | 175.9425 | -1.6014 | 1.2712 | 32.1472 | 40.8658 | 25.2886 | 2.0095 | 32.0704 | 64.4464 | 15.9591 | 14.4952 | -0.1104 | 105.3122 | 235.3065 | 181.1187 | 233.0978 | 161.4473
2012-10-30 19:58 | 28.6682 | 249.7671 | 195.5989 | 247.6564 | 175.9198 | -1.6043 | 1.2712 | 32.1477 | 40.8673 | 25.2886 | 2.0096 | 32.0709 | 64.4484 | 15.9592 | 14.4837 | -0.1106 | 105.3157 | 235.2834 | 181.1152 | 233.1727 | 161.4361
2012-10-30 19:58 | 28.6697 | 249.753 | 195.624 | 247.701 | 175.7495 | -1.6061 | 1.2712 | 32.1475 | 40.8672 | 25.2883 | 2.0095 | 32.0707 | 64.4476 | 15.9592 | 14.4811 | -0.1107 | 105.3147 | 235.2719 | 181.1429 | 233.2199 | 161.2684
2012-10-30 19:58 | 28.6652 | 249.7602 | 195.5762 | 247.673 | 175.8465 | -1.5968 | 1.2712 | 32.1464 | 40.8641 | 25.2885 | 2.0095 | 32.0696 | 64.4429 | 15.9593 | 14.4684 | -0.1101 | 105.3069 | 235.2918 | 181.1078 | 233.2046 | 161.3781
2012-10-30 19:58 | 28.6667 | 249.7343 | 195.5699 | 247.6079 | 175.8636 | -1.6055 | 1.2712 | 32.1456 | 40.8619 | 25.2885 | 2.0094 | 32.0688 | 64.4394 | 15.9593 | 14.4495 | -0.1107 | 105.3013 | 235.2848 | 181.1204 | 233.1584 | 161.4141
2012-10-30 19:58 | 28.6712 | 249.7584 | 195.5626 | 247.6242 | 175.9033 | -1.5997 | 1.2712 | 32.1458 | 40.8622 | 25.2887 | 2.0094 | 32.069 | 64.4407 | 15.9591 | 14.4622 | -0.1103 | 105.3029 | 235.2962 | 181.1004 | 233.162 | 161.4411
2012-10-30 19:58 | 28.6622 | 249.7925 | 195.5457 | 247.6543 | 176.0504 | -1.6003 | 1.2712 | 32.1457 | 40.8622 | 25.2885 | 2.0094 | 32.0688 | 64.4402 | 15.9591 | 14.4491 | -0.1103 | 105.3024 | 235.3434 | 181.0966 | 233.2052 | 161.6013
2012-10-30 19:58 | 28.6682 | 249.7711 | 195.595 | 247.7348 | 176.1306 | -1.602 | 1.2711 | 32.1454 | 40.8613 | 25.2887 | 2.0094 | 32.0687 | 64.4399 | 15.9591 | 14.4102 | -0.1104 | 105.3013 | 235.3609 | 181.1848 | 233.3246 | 161.7204
2012-10-30 19:58 | 28.6645 | 249.7909 | 195.6266 | 247.6488 | 176.0098 | -1.6066 | 1.2711 | 32.1452 | 40.8609 | 25.2885 | 2.0094 | 32.0685 | 64.4389 | 15.9591 | 14.4433 | -0.1108 | 105.2998 | 235.3476 | 181.1833 | 233.2055 | 161.5665
2012-10-30 19:59 | 28.6645 | 249.7987 | 195.5952 | 247.7154 | 175.8302 | -1.6014 | 1.2711 | 32.145 | 40.8606 | 25.2885 | 2.0094 | 32.0682 | 64.4381 | 15.9591 | 14.4473 | -0.1104 | 105.2986 | 235.3514 | 181.1479 | 233.2681 | 161.3829
2012-10-30 19:59 | 28.6705 | 249.8203 | 195.5815 | 247.647 | 175.9182 | -1.6009 | 1.2711 | 32.1444 | 40.8594 | 25.2883 | 2.0094 | 32.0676 | 64.4356 | 15.9591 | 14.4452 | -0.1104 | 105.2949 | 235.3751 | 181.1363 | 233.2018 | 161.473
2012-10-30 19:59 | 28.6757 | 249.8373 | 195.5907 | 247.6483 | 175.7634 | -1.5997 | 1.2711 | 32.1444 | 40.8588 | 25.2885 | 2.0094 | 32.0675 | 64.4355 | 15.959 | 14.4463 | -0.1103 | 105.2943 | 235.391 | 181.1444 | 233.202 | 161.3171
2012-10-30 19:59 | 28.6712 | 249.8171 | 195.539 | 247.6986 | 175.8525 | -1.6009 | 1.2711 | 32.1439 | 40.8578 | 25.2884 | 2.0093 | 32.067 | 64.4332 | 15.959 | 14.4051 | -0.1104 | 105.291 | 235.412 | 181.1339 | 233.2935 | 161.4474
2012-10-30 19:59 | 28.678 | 249.8592 | 195.5419 | 247.7407 | 175.8203 | -1.5974 | 1.271 | 32.1428 | 40.855 | 25.2885 | 2.0093 | 32.066 | 64.4293 | 15.959 | 14.42 | -0.1101 | 105.2843 | 235.4392 | 181.1219 | 233.3207 | 161.4003
2012-10-30 19:59 | 28.6839 | 249.822 | 195.5439 | 247.723 | 175.7794 | -1.5957 | 1.2711 | 32.1427 | 40.8552 | 25.2882 | 2.0093 | 32.066 | 64.4286 | 15.9592 | 14.4137 | -0.11 | 105.2838 | 235.4083 | 181.1302 | 233.3093 | 161.3657
2012-10-30 19:59 | 28.6712 | 249.821 | 195.5194 | 247.7064 | 175.8018 | -1.6037 | 1.271 | 32.1423 | 40.8534 | 25.2886 | 2.0092 | 32.0656 | 64.4275 | 15.9591 | 14.4092 | -0.1106 | 105.2809 | 235.4118 | 181.1102 | 233.2972 | 161.3926
2012-10-30 19:59 | 28.6824 | 249.8362 | 195.5109 | 247.6432 | 175.731 | -1.6014 | 1.271 | 32.1421 | 40.853 | 25.2885 | 2.0092 | 32.0653 | 64.4264 | 15.959 | 14.4571 | -0.1104 | 105.2794 | 235.3791 | 181.0538 | 233.1861 | 161.2739
2012-10-30 19:59 | 28.6824 | 249.8401 | 195.4834 | 247.7333 | 175.8091 | -1.5997 | 1.271 | 32.1417 | 40.8522 | 25.2884 | 2.0092 | 32.0649 | 64.4251 | 15.959 | 14.4163 | -0.1103 | 105.2773 | 235.4238 | 181.0671 | 233.317 | 161.3928
2012-10-30 19:59 | 28.6824 | 249.8166 | 195.5031 | 247.6471 | 175.813 | -1.6009 | 1.271 | 32.1411 | 40.8506 | 25.2886 | 2.0092 | 32.0644 | 64.4228 | 15.959 | 14.3755 | -0.1104 | 105.2734 | 235.4411 | 181.1276 | 233.2716 | 161.4375
2012-10-30 19:59 | 28.6862 | 249.8674 | 195.4636 | 247.7371 | 175.727 | -1.6014 | 1.271 | 32.141 | 40.8499 | 25.2887 | 2.0092 | 32.0643 | 64.4225 | 15.959 | 14.4446 | -0.1104 | 105.2724 | 235.4228 | 181.019 | 233.2925 | 161.2824
2012-10-30 19:59 | 28.6862 | 249.8321 | 195.4911 | 247.7293 | 175.93 | -1.6003 | 1.271 | 32.141 | 40.8505 | 25.2884 | 2.0092 | 32.0642 | 64.4228 | 15.9589 | 14.4486 | -0.1103 | 105.2733 | 235.3835 | 181.0425 | 233.2807 | 161.4814
2012-10-30 19:59 | 28.6802 | 249.83 | 195.5048 | 247.7311 | 175.9318 | -1.5991 | 1.271 | 32.1409 | 40.8504 | 25.2883 | 2.0092 | 32.0641 | 64.4219 | 15.959 | 14.4508 | -0.1102 | 105.2723 | 235.3792 | 181.054 | 233.2803 | 161.481
2012-10-30 19:59 | 28.6869 | 249.8759 | 195.5351 | 247.7417 | 175.9892 | -1.5991 | 1.271 | 32.1407 | 40.8498 | 25.2884 | 2.0092 | 32.064 | 64.4219 | 15.9588 | 14.4575 | -0.1102 | 105.2718 | 235.4184 | 181.0776 | 233.2842 | 161.5317
2012-10-30 19:59 | 28.6839 | 249.8807 | 195.6028 | 247.7074 | 175.9511 | -1.5997 | 1.271 | 32.1406 | 40.8499 | 25.2881 | 2.0092 | 32.0639 | 64.4212 | 15.9589 | 14.4627 | -0.1103 | 105.2712 | 235.418 | 181.1401 | 233.2447 | 161.4884
2012-10-30 19:59 | 28.6839 | 249.8533 | 195.5635 | 247.6878 | 175.8574 | -1.6009 | 1.271 | 32.1403 | 40.8488 | 25.2883 | 2.0091 | 32.0634 | 64.4202 | 15.9587 | 14.479 | -0.1104 | 105.269 | 235.3743 | 181.0845 | 233.2088 | 161.3784
2012-10-30 19:59 | 28.6832 | 249.88 | 195.5549 | 247.6832 | 175.9036 | -1.5985 | 1.2679 | 32.0625 | 40.6522 | 25.2878 | 2.0044 | 31.9881 | 64.1176 | 15.9587 | 14.4905 | -0.1102 | 104.7697 | 235.3895 | 181.0644 | 233.1927 | 161.4131
2012-10-30 19:59 | 28.6892 | 249.8743 | 195.5609 | 247.7754 | 175.9993 | -1.6014 | 1.2649 | 31.9848 | 40.457 | 25.2868 | 1.9994 | 31.9081 | 63.7972 | 15.9588 | 14.4843 | -0.1104 | 104.2542 | 235.39 | 181.0766 | 233.2911 | 161.515
2012-10-30 19:59 | 28.6802 | 249.8731 | 195.5166 | 247.7429 | 175.9552 | -1.5997 | 1.2643 | 31.9687 | 40.4166 | 25.2867 | 1.9984 | 31.8924 | 63.7337 | 15.959 | 14.4712 | -0.1103 | 104.1503 | 235.4019 | 181.0454 | 233.2717 | 161.484
2012-10-30 19:59 | 28.6877 | 249.8532 | 195.5437 | 247.7934 | 175.4124 | -1.5968 | 1.2638 | 31.9563 | 40.3852 | 25.2866 | 1.9976 | 31.8801 | 63.6842 | 15.9591 | 14.4501 | -0.1101 | 104.0694 | 235.4031 | 181.0936 | 233.3433 | 160.9623
2012-10-30 19:59 | 28.6839 | 249.8103 | 195.4888 | 247.7544 | 175.2291 | -1.5991 | 1.2634 | 31.9472 | 40.3625 | 25.2865 | 1.997 | 31.8712 | 63.648 | 15.9592 | 14.479 | -0.1102 | 104.0104 | 235.3313 | 181.0098 | 233.2754 | 160.7501
2012-10-30 19:59 | 28.6832 | 249.8056 | 195.4292 | 247.7223 | 175.5445 | -1.6026 | 1.2633 | 31.9442 | 40.3548 | 25.2866 | 1.9968 | 31.8679 | 63.6347 | 15.9593 | 14.4742 | -0.1105 | 103.9895 | 235.3314 | 180.955 | 233.2481 | 161.0703
2012-10-30 19:59 | 28.6787 | 249.7855 | 195.3579 | 247.7139 | 175.8561 | -1.6055 | 1.2633 | 31.9454 | 40.3576 | 25.2866 | 1.9969 | 31.8691 | 63.6389 | 15.9594 | 14.4371 | -0.1107 | 103.9965 | 235.3484 | 180.9208 | 233.2768 | 161.419
2012-10-30 19:59 | 28.6832 | 249.8017 | 195.4096 | 247.6753 | 175.7865 | -1.5985 | 1.2701 | 32.1165 | 40.7897 | 25.2875 | 2.0083 | 32.0513 | 64.369 | 15.9594 | 14.4334 | -0.1102 | 105.1588 | 235.3683 | 180.9762 | 233.2419 | 161.3531
2012-10-30 20:00 | 28.6899 | 249.7849 | 195.4634 | 247.6664 | 176.1015 | -1.5962 | 1.2934 | 32.7144 | 42.3117 | 25.294 | 2.0449 | 32.636 | 66.7378 | 15.9595 | 14.4156 | -0.11 | 109.0496 | 235.3693 | 181.0478 | 233.2508 | 161.6859
2012-10-30 20:00 | 28.6892 | 249.9017 | 195.5649 | 247.6069 | 176.2101 | -1.6026 | 1.2938 | 32.727 | 42.3419 | 25.2955 | 2.0455 | 32.6449 | 66.7739 | 15.9596 | 14.4149 | -0.1105 | 109.1158 | 235.4868 | 181.15 | 233.192 | 161.7952
2012-10-30 20:00 | 28.6877 | 249.9197 | 195.638 | 247.7112 | 176.072 | -1.6049 | 1.298 | 32.8342 | 42.6174 | 25.2968 | 2.0524 | 32.7549 | 67.2249 | 15.9596 | 14.3971 | -0.1106 | 109.8424 | 235.5226 | 181.2409 | 233.3141 | 161.6749
2012-10-30 20:00 | 28.6854 | 250.0271 | 195.6632 | 247.6815 | 176.2844 | -1.5985 | 1.2985 | 32.848 | 42.6522 | 25.2974 | 2.0532 | 32.7686 | 67.2806 | 15.9597 | 14.4152 | -0.1102 | 109.9328 | 235.6119 | 181.248 | 233.2663 | 161.8692
2012-10-30 20:00 | 28.6892 | 250.0896 | 195.7653 | 247.7244 | 176.3117 | -1.6043 | 1.299 | 32.8606 | 42.6843 | 25.2978 | 2.054 | 32.7814 | 67.3337 | 15.9596 | 14.4027 | -0.1106 | 110.0181 | 235.6869 | 181.3626 | 233.3217 | 161.909
2012-10-30 20:00 | 28.6907 | 250.1928 | 195.7707 | 247.7103 | 176.1726 | -1.5991 | 1.2993 | 32.8711 | 42.7107 | 25.2983 | 2.0547 | 32.792 | 67.3777 | 15.9596 | 14.4164 | -0.1102 | 110.0884 | 235.7764 | 181.3543 | 233.2939 | 161.7562
2012-10-30 20:00 | 28.6937 | 250.3211 | 195.8091 | 247.7564 | 176.2303 | -1.6032 | 1.2995 | 32.8762 | 42.7234 | 25.2986 | 2.055 | 32.7969 | 67.3966 | 15.9598 | 14.3949 | -0.1105 | 110.12 | 235.9262 | 181.4142 | 233.3615 | 161.8354
2012-10-30 20:00 | 28.6899 | 250.3682 | 195.9074 | 247.8191 | 176.285 | -1.6055 | 1.2994 | 32.8721 | 42.7124 | 25.2988 | 2.0548 | 32.793 | 67.3817 | 15.9596 | 14.3952 | -0.1107 | 110.0941 | 235.973 | 181.5122 | 233.4239 | 161.8898
2012-10-30 20:00 | 28.6892 | 250.434 | 195.9303 | 247.8968 | 176.2101 | -1.6014 | 1.2989 | 32.8599 | 42.6808 | 25.2988 | 2.0543 | 32.7863 | 67.3532 | 15.9597 | 14.4108 | -0.1104 | 110.034 | 236.0232 | 181.5195 | 233.486 | 161.7993
2012-10-30 20:00 | 28.6869 | 250.4943 | 195.9398 | 247.9141 | 176.2859 | -1.5985 | 1.299 | 32.8621 | 42.6871 | 25.2985 | 2.0543 | 32.7862 | 67.3538 | 15.9595 | 14.3881 | -0.1102 | 110.0409 | 236.1062 | 181.5517 | 233.526 | 161.8978
2012-10-30 20:00 | 28.6862 | 250.5915 | 196.0177 | 248.0035 | 176.4102 | -1.5991 | 1.2994 | 32.8739 | 42.7168 | 25.2991 | 2.0547 | 32.7927 | 67.3803 | 15.9596 | 14.3589 | -0.1102 | 110.0971 | 236.2326 | 181.6588 | 233.6446 | 162.0513
2012-10-30 20:00 | 28.6914 | 250.628 | 196.019 | 248.0478 | 176.3529 | -1.6014 | 1.2995 | 32.8764 | 42.723 | 25.2992 | 2.0548 | 32.7948 | 67.3882 | 15.9597 | 14.4416 | -0.1104 | 110.1112 | 236.1864 | 181.5774 | 233.6062 | 161.9113
2012-10-30 20:00 | 28.6899 | 250.7126 | 196.045 | 248.0777 | 176.449 | -1.6037 | 1.2995 | 32.8768 | 42.7235 | 25.2995 | 2.0549 | 32.7951 | 67.3898 | 15.9596 | 14.4687 | -0.1106 | 110.1133 | 236.2439 | 181.5763 | 233.609 | 161.9803
2012-10-30 20:00 | 28.6899 | 250.7791 | 196.1393 | 248.1482 | 176.5037 | -1.6003 | 1.2995 | 32.8779 | 42.7263 | 25.2995 | 2.0549 | 32.7959 | 67.3935 | 15.9596 | 14.485 | -0.1103 | 110.1198 | 236.2941 | 181.6543 | 233.6632 | 162.0187
2012-10-30 20:00 | 28.6847 | 250.8209 | 196.2048 | 248.2409 | 176.5844 | -1.5939 | 1.2995 | 32.8784 | 42.7269 | 25.3 | 2.055 | 32.7966 | 67.3957 | 15.9597 | 14.5206 | -0.1099 | 110.1226 | 236.3003 | 181.6842 | 233.7203 | 162.0638
2012-10-30 20:00 | 28.6839 | 250.9297 | 196.2473 | 248.3067 | 176.5836 | -1.6032 | 1.2995 | 32.8782 | 42.7254 | 25.3006 | 2.055 | 32.7963 | 67.3949 | 15.9597 | 14.4912 | -0.1105 | 110.1204 | 236.4385 | 181.7561 | 233.8155 | 162.0924
2012-10-30 20:00 | 28.6862 | 250.9867 | 196.3124 | 248.4147 | 176.6913 | -1.5991 | 1.2995 | 32.877 | 42.7229 | 25.3002 | 2.0549 | 32.7954 | 67.3916 | 15.9596 | 14.4813 | -0.1102 | 110.1144 | 236.5054 | 181.8311 | 233.9334 | 162.21
2012-10-30 20:00 | 28.6922 | 251.0475 | 196.4049 | 248.4638 | 176.7832 | -1.5991 | 1.2993 | 32.8732 | 42.713 | 25.3002 | 2.0549 | 32.7948 | 67.3887 | 15.9596 | 14.475 | -0.1102 | 110.1018 | 236.5725 | 181.9299 | 233.9888 | 162.3082
2012-10-30 20:00 | 28.6967 | 251.1264 | 196.4369 | 248.4918 | 176.6979 | -1.6026 | 1.2994 | 32.8748 | 42.7165 | 25.3006 | 2.055 | 32.7963 | 67.3949 | 15.9596 | 14.4877 | -0.1105 | 110.1114 | 236.6387 | 181.9492 | 234.0041 | 162.2102
2012-10-30 20:00 | 28.7132 | 251.1546 | 196.4731 | 248.5866 | 176.5738 | -1.598 | 1.2994 | 32.876 | 42.7194 | 25.3007 | 2.055 | 32.7974 | 67.3993 | 15.9596 | 14.4756 | -0.1102 | 110.1187 | 236.679 | 181.9975 | 234.111 | 162.0982
2012-10-30 20:00 | 28.7319 | 251.279 | 196.5271 | 248.711 | 176.6277 | -1.6043 | 1.2994 | 32.8764 | 42.7205 | 25.3006 | 2.055 | 32.7976 | 67.4004 | 15.9596 | 14.4699 | -0.1106 | 110.1209 | 236.8091 | 182.0572 | 234.2411 | 162.1578
2012-10-30 20:00 | 28.7529 | 251.3548 | 196.5364 | 248.7516 | 176.6213 | -1.602 | 1.2994 | 32.8768 | 42.7209 | 25.3011 | 2.0551 | 32.7978 | 67.4021 | 15.9594 | 14.4949 | -0.1104 | 110.123 | 236.8599 | 182.0415 | 234.2567 | 162.1264
2012-10-30 20:01 | 28.7753 | 251.3811 | 196.5745 | 248.8132 | 176.7452 | -1.5997 | 1.2995 | 32.8774 | 42.7227 | 25.3009 | 2.0551 | 32.7987 | 67.4049 | 15.9596 | 14.5337 | -0.1103 | 110.1276 | 236.8474 | 182.0408 | 234.2795 | 162.2115
2012-10-30 20:01 | 28.7948 | 251.4827 | 196.6254 | 248.8797 | 176.7296 | -1.6003 | 1.2995 | 32.8777 | 42.7231 | 25.3012 | 2.0551 | 32.7988 | 67.4053 | 15.9596 | 14.545 | -0.1103 | 110.1284 | 236.9377 | 182.0804 | 234.3347 | 162.1846
2012-10-30 20:01 | 28.8233 | 251.5425 | 196.6932 | 248.9865 | 176.7347 | -1.6043 | 1.2995 | 32.8782 | 42.7238 | 25.3015 | 2.0552 | 32.7993 | 67.4077 | 15.9595 | 14.5694 | -0.1106 | 110.1315 | 236.9731 | 182.1238 | 234.4171 | 162.1653
2012-10-30 20:01 | 28.8465 | 251.5579 | 196.6889 | 249.041 | 176.8164 | -1.6003 | 1.2995 | 32.8783 | 42.7242 | 25.3015 | 2.0552 | 32.7994 | 67.4082 | 15.9595 | 14.5273 | -0.1103 | 110.1324 | 237.0306 | 182.1616 | 234.5137 | 162.2891
2012-10-30 20:01 | 28.8682 | 251.5992 | 196.6792 | 249.0549 | 176.8538 | -1.6066 | 1.2995 | 32.8788 | 42.7251 | 25.3017 | 2.0552 | 32.7999 | 67.4105 | 15.9594 | 14.5368 | -0.1108 | 110.1356 | 237.0624 | 182.1424 | 234.5181 | 162.317
2012-10-30 20:01 | 28.8952 | 251.6614 | 196.7258 | 249.1132 | 176.8534 | -1.6026 | 1.2995 | 32.8787 | 42.7244 | 25.3018 | 2.0552 | 32.7997 | 67.4097 | 15.9595 | 14.5148 | -0.1105 | 110.1341 | 237.1466 | 182.211 | 234.5984 | 162.3386
2012-10-30 20:01 | 28.9229 | 251.7204 | 196.7732 | 249.1879 | 176.8343 | -1.602 | 1.2995 | 32.8788 | 42.7244 | 25.302 | 2.0552 | 32.7998 | 67.4095 | 15.9595 | 14.4976 | -0.1104 | 110.1339 | 237.2228 | 182.2756 | 234.6903 | 162.3367
2012-10-30 20:01 | 28.9454 | 251.7507 | 196.8232 | 249.2339 | 176.8724 | -1.6049 | 1.2995 | 32.879 | 42.725 | 25.3019 | 2.0552 | 32.7999 | 67.4106 | 15.9594 | 14.512 | -0.1106 | 110.1356 | 237.2387 | 182.3112 | 234.7219 | 162.3604
2012-10-30 20:01 | 28.9589 | 251.7799 | 196.8367 | 249.314 | 176.9523 | -1.6003 | 1.2995 | 32.8794 | 42.7256 | 25.3023 | 2.0552 | 32.8002 | 67.4116 | 15.9595 | 14.5418 | -0.1103 | 110.1371 | 237.2381 | 182.2949 | 234.7722 | 162.4105
2012-10-30 20:01 | 28.9836 | 251.828 | 196.8653 | 249.3387 | 177.0355 | -1.6009 | 1.2995 | 32.8791 | 42.7252 | 25.3021 | 2.0552 | 32.8001 | 67.4113 | 15.9594 | 14.6563 | -0.1104 | 110.1365 | 237.1717 | 182.209 | 234.6824 | 162.3792
2012-10-30 20:01 | 29.0135 | 251.8892 | 196.9384 | 249.4508 | 177.1435 | -1.6043 | 1.2995 | 32.8804 | 42.7283 | 25.3023 | 2.0553 | 32.8014 | 67.4167 | 15.9594 | 14.6535 | -0.1106 | 110.145 | 237.2357 | 182.2849 | 234.7973 | 162.49
2012-10-30 20:01 | 29.039 | 251.9695 | 196.9993 | 249.488 | 177.2901 | -1.5968 | 1.2996 | 32.8826 | 42.7337 | 25.3024 | 2.0554 | 32.8034 | 67.4251 | 15.9594 | 14.6505 | -0.1101 | 110.1588 | 237.319 | 182.3488 | 234.8375 | 162.6396
2012-10-30 20:01 | 29.06 | 251.9631 | 197.0321 | 249.5678 | 177.3228 | -1.5997 | 1.2996 | 32.8843 | 42.7378 | 25.3026 | 2.0555 | 32.8051 | 67.4321 | 15.9594 | 14.7163 | -0.1103 | 110.1699 | 237.2468 | 182.3158 | 234.8515 | 162.6065
2012-10-30 20:01 | 29.0952 | 252.0531 | 197.1302 | 249.6108 | 177.2838 | -1.6026 | 1.2997 | 32.8859 | 42.7414 | 25.3029 | 2.0556 | 32.8065 | 67.438 | 15.9593 | 14.7189 | -0.1105 | 110.1795 | 237.3342 | 182.4113 | 234.8919 | 162.5649
2012-10-30 20:01 | 29.1207 | 252.0864 | 197.1871 | 249.6794 | 177.2859 | -1.6032 | 1.2997 | 32.886 | 42.7418 | 25.3028 | 2.0556 | 32.8065 | 67.438 | 15.9594 | 14.6464 | -0.1105 | 110.1798 | 237.44 | 182.5407 | 235.033 | 162.6395
2012-10-30 20:01 | 29.1544 | 252.1475 | 197.1894 | 249.717 | 177.343 | -1.5974 | 1.2996 | 32.8852 | 42.7392 | 25.3031 | 2.0556 | 32.8058 | 67.4346 | 15.9595 | 14.6883 | -0.1101 | 110.1739 | 237.4592 | 182.5011 | 235.0287 | 162.6547
2012-10-30 20:01 | 29.1836 | 252.1415 | 197.1832 | 249.7814 | 177.1653 | -1.6043 | 1.2996 | 32.8846 | 42.738 | 25.3029 | 2.0555 | 32.8051 | 67.4319 | 15.9595 | 14.6236 | -0.1106 | 110.1699 | 237.5179 | 182.5596 | 235.1578 | 162.5417
2012-10-30 20:01 | 29.1978 | 252.2222 | 197.2289 | 249.8113 | 177.2615 | -1.6066 | 1.2996 | 32.8846 | 42.7372 | 25.3034 | 2.0555 | 32.805 | 67.4321 | 15.9593 | 14.6052 | -0.1108 | 110.1693 | 237.617 | 182.6237 | 235.2061 | 162.6563
2012-10-30 20:01 | 29.2136 | 252.2693 | 197.2368 | 249.9015 | 177.3085 | -1.5991 | 1.2996 | 32.8843 | 42.7363 | 25.3035 | 2.0555 | 32.8047 | 67.4307 | 15.9594 | 14.5271 | -0.1102 | 110.167 | 237.7422 | 182.7097 | 235.3744 | 162.7814
2012-10-30 20:01 | 29.233 | 252.2887 | 197.2719 | 249.8856 | 177.4216 | -1.6003 | 1.2996 | 32.8836 | 42.7342 | 25.3036 | 2.0555 | 32.804 | 67.4277 | 15.9594 | 14.5138 | -0.1103 | 110.1619 | 237.7749 | 182.7581 | 235.3718 | 162.9078
2012-10-30 20:01 | 29.2488 | 252.3554 | 197.3034 | 249.9641 | 177.414 | -1.6061 | 1.2996 | 32.8836 | 42.7345 | 25.3035 | 2.0555 | 32.8041 | 67.4281 | 15.9593 | 14.5745 | -0.1107 | 110.1626 | 237.7809 | 182.7289 | 235.3896 | 162.8395
2012-10-30 20:01 | 29.26 | 252.394 | 197.3657 | 249.9988 | 177.4838 | -1.6009 | 1.2996 | 32.8836 | 42.7352 | 25.3031 | 2.0555 | 32.8039 | 67.4273 | 15.9593 | 14.5694 | -0.1104 | 110.1625 | 237.8246 | 182.7963 | 235.4294 | 162.9144
2012-10-30 20:01 | 29.2832 | 252.4093 | 197.4125 | 250.0181 | 177.5499 | -1.6026 | 1.2996 | 32.8836 | 42.7342 | 25.3037 | 2.0555 | 32.8039 | 67.4278 | 15.9592 | 14.5354 | -0.1105 | 110.162 | 237.8739 | 182.8771 | 235.4827 | 163.0145
2012-10-30 20:01 | 29.2982 | 252.4322 | 197.455 | 250.0644 | 177.5923 | -1.5985 | 1.2996 | 32.8833 | 42.7339 | 25.3033 | 2.0554 | 32.8035 | 67.4256 | 15.9594 | 14.53 | -0.1102 | 110.1595 | 237.9022 | 182.925 | 235.5344 | 163.0623
2012-10-30 20:02 | 29.3162 | 252.5049 | 197.4809 | 250.1568 | 177.54 | -930789.6067 | 1.2994 | 32.8809 | 42.7268 | 25.3039 | 2.0553 | 32.8011 | 67.4167 | 15.9592 | 14.5888 | -64168.6355 | 110.1435 | 237.9161 | 182.8921 | 235.568 | 162.9512
2012-10-30 20:02 | 29.3312 | 252.5473 | 197.5038 | 250.1835 | 177.637 | 73.0569 | 1.2994 | 32.8803 | 42.7257 | 25.3036 | 2.0553 | 32.8008 | 67.4151 | 15.9592 | 14.6324 | 5.0365 | 110.1409 | 237.9149 | 182.8714 | 235.5511 | 163.0046
2012-10-30 20:02 | 29.3469 | 252.5787 | 197.547 | 250.2462 | 177.6331 | 73.0413 | 1.2994 | 32.8807 | 42.726 | 25.304 | 2.0553 | 32.8013 | 67.4165 | 15.9593 | 14.6195 | 5.0355 | 110.1425 | 237.9592 | 182.9275 | 235.6267 | 163.0136
2012-10-30 20:02 | 29.3619 | 252.6719 | 197.6288 | 250.2651 | 177.3982 | 73.0528 | 1.2994 | 32.8807 | 42.7258 | 25.3042 | 2.0553 | 32.8012 | 67.4166 | 15.9593 | 14.6468 | 5.0363 | 110.1424 | 238.0251 | 182.982 | 235.6183 | 162.7514
2012-10-30 20:02 | 29.3746 | 252.6807 | 197.5943 | 250.313 | 177.6101 | 73.0441 | 1.2994 | 32.8804 | 42.7253 | 25.3041 | 2.0553 | 32.8009 | 67.4151 | 15.9593 | 14.6432 | 5.0357 | 110.1404 | 238.0375 | 182.9511 | 235.6698 | 162.9669
2012-10-30 20:02 | 29.3746 | 252.7276 | 197.61 | 250.3444 | 177.5671 | 73.0424 | 1.2994 | 32.8803 | 42.7241 | 25.3046 | 2.0553 | 32.8007 | 67.4142 | 15.9593 | 14.6187 | 5.0355 | 110.1383 | 238.1089 | 182.9913 | 235.7257 | 162.9484
2012-10-30 20:02 | 29.3911 | 252.7128 | 197.6187 | 250.4 | 177.6734 | 73.0499 | 1.2993 | 32.8779 | 42.7181 | 25.3044 | 2.0551 | 32.7982 | 67.4043 | 15.9592 | 14.5902 | 5.0361 | 110.1224 | 238.1226 | 183.0285 | 235.8098 | 163.0832
2012-10-30 20:02 | 29.4008 | 252.7656 | 197.652 | 250.4449 | 177.6831 | 73.0678 | 1.2993 | 32.8778 | 42.7174 | 25.3047 | 2.0551 | 32.7985 | 67.4053 | 15.9593 | 14.6204 | 5.0373 | 110.1227 | 238.1452 | 183.0316 | 235.8245 | 163.0627
2012-10-30 20:02 | 29.4083 | 252.8044 | 197.6752 | 250.472 | 177.6243 | 73.0615 | 1.2998 | 32.8905 | 42.7503 | 25.3047 | 2.0561 | 32.814 | 67.4687 | 15.9594 | 14.5952 | 5.0369 | 110.219 | 238.2092 | 183.08 | 235.8768 | 163.0291
2012-10-30 20:02 | 29.4098 | 252.8372 | 197.716 | 250.4853 | 177.614 | 73.0678 | 1.2998 | 32.8917 | 42.7532 | 25.3049 | 2.0562 | 32.8153 | 67.4745 | 15.9593 | 14.5763 | 5.0373 | 110.2276 | 238.2609 | 183.1397 | 235.909 | 163.0377
2012-10-30 20:02 | 29.4173 | 252.8525 | 197.692 | 250.5123 | 177.6528 | 73.0771 | 1.2991 | 32.8733 | 42.7049 | 25.3052 | 2.0553 | 32.8017 | 67.4186 | 15.9593 | 14.5838 | 5.0379 | 110.1235 | 238.2687 | 183.1082 | 235.9285 | 163.069
2012-10-30 20:02 | 29.4211 | 252.8563 | 197.7391 | 250.5983 | 177.8479 | 73.0678 | 1.2934 | 32.7288 | 42.333 | 25.3036 | 2.0459 | 32.6512 | 66.8015 | 15.9592 | 14.5671 | 5.0373 | 109.1345 | 238.2892 | 183.172 | 236.0312 | 163.2808
2012-10-30 20:02 | 29.4203 | 252.9259 | 197.8326 | 250.5467 | 177.8823 | 73.0678 | 1.2934 | 32.7276 | 42.3299 | 25.3036 | 2.0459 | 32.6506 | 66.7983 | 15.9594 | 14.6276 | 5.0373 | 109.1281 | 238.2983 | 183.205 | 235.9191 | 163.2547
2012-10-30 20:02 | 29.4271 | 252.964 | 197.8394 | 250.62 | 178.0961 | 73.0806 | 1.2928 | 32.7131 | 42.2925 | 25.3034 | 2.0449 | 32.6362 | 66.7393 | 15.9595 | 14.6181 | 5.0382 | 109.0318 | 238.3459 | 183.2213 | 236.0019 | 163.478
2012-10-30 20:02 | 29.4211 | 252.9854 | 197.8688 | 250.6845 | 177.8792 | 73.0678 | 1.2925 | 32.7036 | 42.2681 | 25.3034 | 2.0443 | 32.6267 | 66.7 | 15.9595 | 14.6202 | 5.0373 | 108.9681 | 238.3652 | 183.2486 | 236.0643 | 163.259
2012-10-30 20:02 | 29.4233 | 252.9915 | 197.9103 | 250.7101 | 177.6353 | 73.0638 | 1.2924 | 32.703 | 42.2668 | 25.3033 | 2.0443 | 32.6259 | 66.6964 | 15.9596 | 14.6143 | 5.037 | 108.9632 | 238.3772 | 183.296 | 236.0958 | 163.021
2012-10-30 20:02 | 29.4226 | 253.0182 | 197.8664 | 250.7173 | 177.5644 | 73.0493 | 1.2924 | 32.7028 | 42.2659 | 25.3034 | 2.0442 | 32.6256 | 66.6945 | 15.9598 | 14.5686 | 5.036 | 108.9605 | 238.4496 | 183.2978 | 236.1487 | 162.9958
2012-10-30 20:02 | 29.4256 | 253.0446 | 197.8536 | 250.7516 | 177.7197 | 73.0609 | 1.2925 | 32.7039 | 42.2688 | 25.3035 | 2.0443 | 32.6266 | 66.6984 | 15.9598 | 14.6166 | 5.0368 | 108.9672 | 238.428 | 183.237 | 236.135 | 163.1031
2012-10-30 20:02 | 29.4331 | 253.0913 | 197.8887 | 250.7982 | 177.9029 | 73.054 | 1.2926 | 32.7061 | 42.2742 | 25.3035 | 2.0444 | 32.6288 | 66.7076 | 15.9598 | 14.5628 | 5.0363 | 108.9819 | 238.5285 | 183.3259 | 236.2354 | 163.3401
2012-10-30 20:02 | 29.4338 | 253.0802 | 197.8776 | 250.8029 | 177.6341 | 73.0765 | 1.2927 | 32.7098 | 42.2833 | 25.3038 | 2.0447 | 32.6325 | 66.7221 | 15.9599 | 14.5431 | 5.0379 | 109.0055 | 238.5371 | 183.3345 | 236.2598 | 163.091
2012-10-30 20:02 | 29.4278 | 253.1212 | 197.8873 | 250.8203 | 177.4368 | 73.0782 | 0.92002 | 23.3293 | 21.4635 | 25.3572 | 1.4389 | 22.996 | 33.0886 | 15.9819 | 14.5371 | 5.038 | 54.552 | 238.5841 | 183.3502 | 236.2832 | 162.8997
2012-10-30 20:02 | 29.4398 | 252.7694 | 197.0739 | 250.8676 | 177.1989 | 71.7762 | 0.006744 | 0.33527 | 0.002261 | 49.7138 | 0.006955 | 0.21682 | 0.001508 | 31.1753 | 14.6104 | 4.9483 | 0.003769 | 238.159 | 182.4635 | 236.2572 | 162.5885
2012-10-30 20:02 | 29.4301 | 250.8422 | 195.2646 | 250.6034 | 177.5953 | 70.8867 | 0.006744 | 0.33514 | 0.00226 | 49.6945 | 0.006955 | 0.2169 | 0.001509 | 31.1856 | 14.5598 | 4.8869 | 0.003769 | 236.2824 | 180.7048 | 236.0436 | 163.0355
2012-10-30 20:02 | 29.4383 | 248.5049 | 194.1651 | 250.0988 | 177.3223 | 70.6469 | 0.006747 | 0.3351 | 0.002261 | 49.6659 | 0.006958 | 0.21694 | 0.001509 | 31.1784 | 14.5517 | 4.8704 | 0.00377 | 233.9532 | 179.6134 | 235.5471 | 162.7706
2012-10-30 20:03 | 29.4458 | 246.6159 | 193.3794 | 249.4328 | 176.9004 | 70.4637 | 0.00676 | 0.33513 | 0.002265 | 49.5747 | 0.006968 | 0.21697 | 0.001512 | 31.1378 | 14.5837 | 4.8578 | 0.003777 | 232.0322 | 178.7957 | 234.8491 | 162.3167
2012-10-30 20:03 | 29.4376 | 245.0279 | 192.688 | 248.7431 | 176.2794 | 70.3256 | 0.006766 | 0.33521 | 0.002268 | 49.5436 | 0.006958 | 0.21705 | 0.00151 | 31.195 | 14.5959 | 4.8482 | 0.003778 | 230.432 | 178.0921 | 234.1472 | 161.6835
2012-10-30 20:03 | 29.4391 | 243.3508 | 191.8456 | 247.922 | 176.0116 | 70.1851 | 0.006773 | 0.3352 | 0.00227 | 49.4902 | 0.006961 | 0.21703 | 0.001511 | 31.1774 | 14.5566 | 4.8386 | 0.003781 | 228.7942 | 177.289 | 233.3654 | 161.455
2012-10-30 20:03 | 29.4368 | 241.4101 | 191.0113 | 246.8342 | 175.4747 | 70.0222 | 0.006779 | 0.33518 | 0.002272 | 49.4443 | 0.006974 | 0.21703 | 0.001514 | 31.1193 | 14.5583 | 4.8273 | 0.003786 | 226.8518 | 176.453 | 232.2759 | 160.9164
2012-10-30 20:03 | 29.4346 | 239.7867 | 190.2322 | 245.6091 | 174.7938 | 69.8557 | 0.006766 | 0.33513 | 0.002267 | 49.5307 | 0.006965 | 0.21705 | 0.001512 | 31.1637 | 14.5521 | 4.8159 | 0.003779 | 225.2346 | 175.6801 | 231.057 | 160.2417
2012-10-30 20:03 | 29.4346 | 238.3534 | 189.432 | 244.4092 | 174.2011 | 69.7089 | 0.006776 | 0.33514 | 0.002271 | 49.4599 | 0.006974 | 0.21698 | 0.001513 | 31.1131 | 14.5929 | 4.8057 | 0.003784 | 223.7605 | 174.8391 | 229.8163 | 159.6082
2012-10-30 20:03 | 29.4331 | 236.6548 | 188.5403 | 243.0937 | 173.4745 | 69.5523 | 0.006766 | 0.33507 | 0.002267 | 49.5222 | 0.006974 | 0.217 | 0.001513 | 31.1151 | 14.6445 | 4.7949 | 0.00378 | 222.0103 | 173.8958 | 228.4492 | 158.83
2012-10-30 20:03 | 29.4331 | 234.8432 | 187.521 | 241.4535 | 172.4613 | 69.372 | 0.006776 | 0.33499 | 0.00227 | 49.4385 | 0.006974 | 0.21698 | 0.001513 | 31.1131 | 14.6077 | 4.7825 | 0.003783 | 220.2355 | 172.9133 | 226.8458 | 157.8536
2012-10-30 20:03 | 29.4248 | 233.1052 | 186.5956 | 239.9457 | 171.6155 | 69.2316 | 0.006763 | 0.33498 | 0.002265 | 49.5314 | 0.006983 | 0.21695 | 0.001515 | 31.0689 | 14.6566 | 4.7728 | 0.00378 | 218.4486 | 171.939 | 225.2891 | 156.9589
2012-10-30 20:03 | 29.4368 | 231.8196 | 185.7809 | 238.8308 | 170.9578 | 69.0917 | 0.006785 | 0.33495 | 0.002273 | 49.3665 | 0.006987 | 0.2169 | 0.001515 | 31.0428 | 14.7094 | 4.7632 | 0.003788 | 217.1102 | 171.0715 | 224.1214 | 156.2484
2012-10-30 20:03 | 29.4316 | 230.4772 | 184.9179 | 237.5216 | 170.3609 | 68.9703 | 0.006782 | 0.33495 | 0.002272 | 49.3884 | 0.006987 | 0.21694 | 0.001516 | 31.0489 | 14.7818 | 4.7548 | 0.003788 | 215.6954 | 170.1361 | 222.7398 | 155.5791
2012-10-30 20:03 | 29.4301 | 229.1854 | 184.145 | 236.086 | 169.7835 | 68.838 | 0.006785 | 0.33485 | 0.002272 | 49.3517 | 0.006977 | 0.21698 | 0.001514 | 31.0998 | 14.8375 | 4.7457 | 0.003786 | 214.3479 | 169.3075 | 221.2485 | 154.946
2012-10-30 20:03 | 29.4241 | 227.6409 | 183.2386 | 234.4136 | 169.0034 | 68.682 | 0.006776 | 0.33488 | 0.002269 | 49.4215 | 0.00698 | 0.21691 | 0.001514 | 31.0759 | 14.8315 | 4.7349 | 0.003783 | 212.8094 | 168.4071 | 219.5821 | 154.1719
2012-10-30 20:03 | 29.4256 | 226.1824 | 182.3205 | 232.878 | 168.2544 | 68.5253 | 0.006782 | 0.33485 | 0.002271 | 49.3735 | 0.006974 | 0.21687 | 0.001512 | 31.0965 | 14.7881 | 4.7241 | 0.003783 | 211.3943 | 167.5324 | 218.0899 | 153.4663
2012-10-30 20:03 | 29.4271 | 224.8457 | 181.5005 | 231.5779 | 167.5057 | 68.4242 | 0.006773 | 0.33479 | 0.002268 | 49.4305 | 0.006974 | 0.21688 | 0.001513 | 31.0985 | 14.7896 | 4.7172 | 0.003781 | 210.0561 | 166.7109 | 216.7883 | 152.7161
2012-10-30 20:03 | 29.4151 | 223.4717 | 180.6438 | 230.1814 | 166.837 | 68.2947 | 0.006773 | 0.33469 | 0.002267 | 49.4156 | 0.00698 | 0.21694 | 0.001514 | 31.0801 | 14.7735 | 4.7082 | 0.003781 | 208.6982 | 165.8703 | 215.4079 | 152.0635
2012-10-30 20:03 | 29.4106 | 222.2782 | 179.9043 | 228.796 | 166.1295 | 68.1734 | 0.006782 | 0.33469 | 0.00227 | 49.35 | 0.006987 | 0.21687 | 0.001515 | 31.0386 | 14.7567 | 4.6999 | 0.003785 | 207.5215 | 165.1476 | 214.0393 | 151.3728
2012-10-30 20:03 | 29.4211 | 220.8831 | 179.0394 | 227.2759 | 165.3362 | 68.0491 | 0.006782 | 0.33471 | 0.00227 | 49.3521 | 0.006996 | 0.21687 | 0.001517 | 30.9987 | 14.7468 | 4.6913 | 0.003787 | 206.1363 | 164.2926 | 212.5291 | 150.5894
2012-10-30 20:03 | 29.4091 | 219.3984 | 178.0195 | 225.6661 | 164.5014 | 67.881 | 0.006798 | 0.33459 | 0.002275 | 49.219 | 0.006983 | 0.21693 | 0.001515 | 31.0647 | 14.7226 | 4.6797 | 0.00379 | 204.6758 | 163.2969 | 210.9435 | 149.7788
2012-10-30 20:03 | 29.4083 | 218.0548 | 177.144 | 224.1577 | 163.7286 | 67.7405 | 0.006766 | 0.33456 | 0.002264 | 49.4475 | 0.006996 | 0.21684 | 0.001517 | 30.9946 | 14.7095 | 4.67 | 0.003781 | 203.3453 | 162.4345 | 209.4482 | 149.0191
2012-10-30 20:03 | 29.4023 | 216.7848 | 176.3963 | 222.888 | 163.0634 | 67.6232 | 0.006776 | 0.33458 | 0.002267 | 49.3766 | 0.006987 | 0.21675 | 0.001514 | 31.022 | 14.7199 | 4.6619 | 0.003781 | 202.0649 | 161.6764 | 208.1681 | 148.3435
2012-10-30 20:03 | 29.4031 | 215.612 | 175.5464 | 221.4006 | 162.2424 | 67.4989 | 0.006773 | 0.33448 | 0.002265 | 49.3836 | 0.006993 | 0.21679 | 0.001516 | 31.0017 | 14.7125 | 4.6534 | 0.003781 | 200.8995 | 160.8339 | 206.6881 | 147.5299
2012-10-30 20:03 | 29.4001 | 214.2938 | 174.5604 | 219.9289 | 161.3018 | 67.3741 | 0.006779 | 0.33439 | 0.002267 | 49.327 | 0.006987 | 0.21675 | 0.001514 | 31.022 | 14.7258 | 4.6448 | 0.003781 | 199.568 | 159.8346 | 205.2031 | 146.576
2012-10-30 20:04 | 29.3956 | 213.0842 | 173.7371 | 218.6092 | 160.6389 | 67.2418 | 0.006788 | 0.33439 | 0.00227 | 49.2616 | 0.006996 | 0.21674 | 0.001516 | 30.9801 | 14.7581 | 4.6356 | 0.003786 | 198.3261 | 158.979 | 203.8511 | 145.8808
2012-10-30 20:04 | 29.3949 | 211.9062 | 172.8984 | 217.1948 | 159.8098 | 67.1314 | 0.006779 | 0.33432 | 0.002266 | 49.3164 | 0.006987 | 0.21672 | 0.001514 | 31.0179 | 14.7125 | 4.628 | 0.00378 | 197.1937 | 158.1859 | 202.4823 | 145.0973
2012-10-30 20:04 | 29.3881 | 210.5488 | 172.015 | 215.7821 | 159.0523 | 67.0065 | 0.006779 | 0.33433 | 0.002266 | 49.3186 | 0.006996 | 0.21671 | 0.001516 | 30.976 | 14.722 | 4.6194 | 0.003782 | 195.8268 | 157.293 | 201.0601 | 144.3303
2012-10-30 20:04 | 29.3806 | 209.2774 | 171.1196 | 214.3806 | 158.3331 | 66.8615 | 0.006795 | 0.33424 | 0.002271 | 49.1897 | 0.00699 | 0.21672 | 0.001515 | 31.0046 | 14.7104 | 4.6094 | 0.003786 | 194.567 | 156.4092 | 199.6702 | 143.6227
2012-10-30 20:04 | 29.3844 | 208.108 | 170.2632 | 213.1361 | 157.6642 | 66.7471 | 0.006788 | 0.33413 | 0.002268 | 49.2233 | 0.006999 | 0.21672 | 0.001517 | 30.9647 | 14.706 | 4.6015 | 0.003785 | 193.402 | 155.5572 | 198.4301 | 142.9582
2012-10-30 20:04 | 29.3919 | 206.8991 | 169.4496 | 211.9071 | 157.0455 | 66.6032 | 0.006785 | 0.3342 | 0.002268 | 49.2559 | 0.006996 | 0.21662 | 0.001515 | 30.9636 | 14.7217 | 4.5916 | 0.003783 | 192.1774 | 154.7279 | 197.1854 | 142.3238
2012-10-30 20:04 | 29.3866 | 205.7052 | 168.6042 | 210.4764 | 156.4723 | 66.4569 | 0.006788 | 0.33416 | 0.002268 | 49.2276 | 0.00699 | 0.21664 | 0.001514 | 30.9923 | 14.6756 | 4.5815 | 0.003782 | 191.0296 | 153.9286 | 195.8008 | 141.7967
2012-10-30 20:04 | 29.3806 | 204.4792 | 167.774 | 209.0254 | 155.7054 | 66.3165 | 0.006791 | 0.33414 | 0.002269 | 49.2038 | 0.006987 | 0.21671 | 0.001514 | 31.0159 | 14.6818 | 4.5719 | 0.003783 | 189.7974 | 153.0922 | 194.3436 | 141.0236
2012-10-30 20:04 | 29.3821 | 203.3829 | 167.0177 | 207.8026 | 154.9503 | 66.2102 | 0.006766 | 0.33401 | 0.00226 | 49.3664 | 0.006993 | 0.21662 | 0.001515 | 30.9768 | 14.6792 | 4.5645 | 0.003775 | 188.7037 | 152.3385 | 193.1234 | 140.2711
2012-10-30 20:04 | 29.3799 | 202.3145 | 166.2619 | 206.5841 | 154.2959 | 66.1015 | 0.006791 | 0.33401 | 0.002268 | 49.1847 | 0.00699 | 0.21661 | 0.001514 | 30.9881 | 14.677 | 4.557 | 0.003782 | 187.6375 | 151.5849 | 191.9071 | 139.6189
2012-10-30 20:04 | 29.3784 | 201.1486 | 165.4721 | 205.3861 | 153.658 | 65.9744 | 0.006791 | 0.33394 | 0.002268 | 49.1739 | 0.00699 | 0.21664 | 0.001514 | 30.9923 | 14.6674 | 4.5483 | 0.003782 | 186.4812 | 150.8047 | 190.7187 | 138.9906
2012-10-30 20:04 | 29.3731 | 199.8927 | 164.5547 | 204.09 | 152.8391 | 65.8466 | 0.006776 | 0.33396 | 0.002263 | 49.285 | 0.007002 | 0.21658 | 0.001516 | 30.9309 | 14.6621 | 4.5395 | 0.003779 | 185.2306 | 149.8926 | 189.4279 | 138.177
2012-10-30 20:04 | 29.3761 | 198.6257 | 163.6964 | 202.7198 | 151.9672 | 65.7068 | 0.006776 | 0.3339 | 0.002262 | 49.2764 | 0.006993 | 0.21655 | 0.001514 | 30.9665 | 14.661 | 4.5298 | 0.003776 | 183.9647 | 149.0354 | 188.0588 | 137.3062
2012-10-30 20:04 | 29.3679 | 197.5601 | 162.8894 | 201.4567 | 151.1424 | 65.5912 | 0.006788 | 0.33391 | 0.002267 | 49.1915 | 0.006983 | 0.21652 | 0.001512 | 31.0067 | 14.7263 | 4.5219 | 0.003779 | 182.8338 | 148.1631 | 186.7304 | 136.4161
2012-10-30 20:04 | 29.3761 | 196.5229 | 162.1069 | 200.3047 | 150.5423 | 65.4866 | 0.006801 | 0.33382 | 0.00227 | 49.0847 | 0.006987 | 0.21655 | 0.001513 | 30.9931 | 14.7181 | 4.5146 | 0.003783 | 181.8048 | 147.3888 | 185.5866 | 135.8242
2012-10-30 20:04 | 29.3716 | 195.3122 | 161.3081 | 199.1481 | 149.7641 | 65.3727 | 0.006782 | 0.33375 | 0.002264 | 49.2116 | 0.006999 | 0.21656 | 0.001516 | 30.9421 | 14.7667 | 4.5068 | 0.00378 | 180.5455 | 146.5414 | 184.3814 | 134.9974
2012-10-30 20:04 | 29.3694 | 194.3555 | 160.5546 | 197.9625 | 149.0464 | 65.2381 | 0.006776 | 0.33372 | 0.002261 | 49.2509 | 0.006983 | 0.21652 | 0.001512 | 31.0067 | 14.7931 | 4.4975 | 0.003773 | 179.5624 | 145.7615 | 183.1694 | 134.2533
2012-10-30 20:04 | 29.3664 | 193.4099 | 159.8432 | 196.9533 | 148.5165 | 65.1433 | 0.006795 | 0.3337 | 0.002267 | 49.1089 | 0.006977 | 0.21649 | 0.00151 | 31.0292 | 14.7248 | 4.491 | 0.003777 | 178.6851 | 145.1184 | 182.2285 | 133.7917
2012-10-30 20:04 | 29.3641 | 192.379 | 159.1211 | 195.9883 | 147.8645 | 65.0456 | 0.006782 | 0.33371 | 0.002263 | 49.2051 | 0.006996 | 0.21656 | 0.001515 | 30.9554 | 14.7674 | 4.4842 | 0.003778 | 177.6116 | 144.3537 | 181.2209 | 133.0971
2012-10-30 20:04 | 29.3686 | 191.5239 | 158.4642 | 194.9635 | 147.1311 | 64.9381 | 0.006776 | 0.33365 | 0.002261 | 49.2401 | 0.006996 | 0.21649 | 0.001515 | 30.945 | 14.7596 | 4.4768 | 0.003776 | 176.7643 | 143.7046 | 180.2039 | 132.3715
2012-10-30 20:04 | 29.3611 | 190.4373 | 157.6486 | 193.8837 | 146.5167 | 64.8162 | 0.006769 | 0.33364 | 0.002258 | 49.289 | 0.00699 | 0.21652 | 0.001513 | 30.9757 | 14.7766 | 4.4684 | 0.003771 | 175.6607 | 142.872 | 179.1071 | 131.7401
2012-10-30 20:04 | 29.3724 | 189.4483 | 156.8212 | 192.6778 | 145.8637 | 64.7035 | 0.006779 | 0.33357 | 0.002261 | 49.2056 | 0.00698 | 0.21651 | 0.001511 | 31.0181 | 14.7471 | 4.4607 | 0.003772 | 174.7012 | 142.0741 | 177.9307 | 131.1166
2012-10-30 20:04 | 29.3679 | 188.5535 | 156.218 | 191.6802 | 145.1797 | 64.5995 | 0.006776 | 0.33351 | 0.00226 | 49.2189 | 0.006987 | 0.21638 | 0.001512 | 30.9684 | 14.7875 | 4.4535 | 0.003772 | 173.766 | 141.4305 | 176.8927 | 130.3922
2012-10-30 20:04 | 29.3649 | 187.5665 | 155.4928 | 190.8021 | 144.4401 | 64.5093 | 0.006773 | 0.33349 | 0.002259 | 49.2386 | 0.006996 | 0.2164 | 0.001514 | 30.9327 | 14.8008 | 4.4473 | 0.003773 | 172.7657 | 140.692 | 176.0013 | 129.6393
2012-10-30 20:05 | 29.3634 | 186.4166 | 154.5611 | 189.6079 | 143.6868 | 64.3972 | 0.006776 | 0.33349 | 0.00226 | 49.2168 | 0.00699 | 0.2164 | 0.001513 | 30.9592 | 14.7667 | 4.4395 | 0.003773 | 171.6499 | 139.7944 | 174.8412 | 128.9201
2012-10-30 20:05 | 29.3649 | 185.3133 | 153.7293 | 188.3937 | 142.9101 | 64.2706 | 0.006785 | 0.33343 | 0.002262 | 49.143 | 0.006983 | 0.21639 | 0.001511 | 30.9881 | 14.7682 | 4.4308 | 0.003773 | 170.5451 | 138.9611 | 173.6255 | 128.1419
2012-10-30 20:05 | 29.3641 | 184.3374 | 153.0421 | 187.3422 | 142.2234 | 64.162 | 0.006785 | 0.33339 | 0.002262 | 49.1365 | 0.006993 | 0.21643 | 0.001514 | 30.9501 | 14.7674 | 4.4233 | 0.003776 | 169.57 | 138.2747 | 172.5748 | 127.456
2012-10-30 20:05 | 29.3739 | 183.4115 | 152.3311 | 186.3801 | 141.5858 | 64.0678 | 0.006785 | 0.33332 | 0.002262 | 49.1259 | 0.006974 | 0.21635 | 0.001509 | 31.0219 | 14.7527 | 4.4168 | 0.003771 | 168.6588 | 137.5784 | 171.6274 | 126.8331
2012-10-30 20:05 | 29.3724 | 182.4825 | 151.586 | 185.3718 | 141.0904 | 63.9568 | 0.006769 | 0.33328 | 0.002256 | 49.2356 | 0.00699 | 0.21632 | 0.001512 | 30.9468 | 14.7553 | 4.4092 | 0.003768 | 167.7272 | 136.8307 | 170.6165 | 126.3351
2012-10-30 20:05 | 29.3679 | 181.5391 | 150.8151 | 184.2981 | 140.4772 | 63.8516 | 0.006776 | 0.3332 | 0.002258 | 49.1741 | 0.006993 | 0.21629 | 0.001513 | 30.9294 | 14.7508 | 4.4019 | 0.003771 | 166.7883 | 136.0643 | 169.5473 | 125.7264
2012-10-30 20:05 | 29.3716 | 180.6433 | 150.1682 | 183.3505 | 139.7691 | 63.7442 | 0.006785 | 0.3332 | 0.002261 | 49.1089 | 0.00698 | 0.21635 | 0.00151 | 30.9953 | 14.7504 | 4.3945 | 0.003771 | 165.8929 | 135.4178 | 168.6001 | 125.0187
2012-10-30 20:05 | 29.3656 | 179.7147 | 149.5619 | 182.4405 | 139.1932 | 63.6627 | 0.006779 | 0.33322 | 0.002259 | 49.1544 | 0.006971 | 0.21632 | 0.001508 | 31.0311 | 14.7689 | 4.3889 | 0.003767 | 164.9458 | 134.793 | 167.6716 | 124.4243
2012-10-30 20:05 | 29.3634 | 178.8723 | 148.8096 | 181.4994 | 138.5944 | 63.5615 | 0.006773 | 0.33316 | 0.002256 | 49.1894 | 0.006971 | 0.21629 | 0.001508 | 31.027 | 14.783 | 4.3819 | 0.003764 | 164.0893 | 134.0266 | 166.7164 | 123.8114
2012-10-30 20:05 | 29.3761 | 177.9279 | 148.0687 | 180.464 | 137.9496 | 63.4379 | 0.006776 | 0.33317 | 0.002258 | 49.1699 | 0.006983 | 0.21629 | 0.00151 | 30.9736 | 14.7794 | 4.3734 | 0.003768 | 163.1485 | 133.2893 | 165.6846 | 123.1702
2012-10-30 20:05 | 29.3731 | 176.9877 | 147.3124 | 179.5385 | 137.3275 | 63.335 | 0.006776 | 0.3331 | 0.002257 | 49.1591 | 0.00699 | 0.21632 | 0.001512 | 30.9468 | 14.7846 | 4.3663 | 0.003769 | 162.2031 | 132.5278 | 164.7539 | 122.5429
2012-10-30 20:05 | 29.3709 | 175.9902 | 146.6379 | 178.5047 | 136.7829 | 63.2431 | 0.006785 | 0.33303 | 0.00226 | 49.0833 | 0.006974 | 0.21623 | 0.001508 | 31.0053 | 14.7456 | 4.36 | 0.003768 | 161.2446 | 131.8923 | 163.7591 | 122.0373
2012-10-30 20:05 | 29.3701 | 175.1584 | 146.0572 | 177.5626 | 136.171 | 63.1292 | 0.006795 | 0.333 | 0.002263 | 49.0068 | 0.00699 | 0.21622 | 0.001511 | 30.9323 | 14.8387 | 4.3521 | 0.003774 | 160.3197 | 131.2185 | 162.7239 | 121.3323
2012-10-30 20:05 | 29.3754 | 174.1926 | 145.3216 | 176.5606 | 135.4472 | 63.0408 | 0.006773 | 0.333 | 0.002255 | 49.166 | 0.006983 | 0.21622 | 0.00151 | 30.9633 | 14.893 | 4.346 | 0.003765 | 159.2996 | 130.4286 | 161.6676 | 120.5542
2012-10-30 20:05 | 29.3716 | 173.2882 | 144.512 | 175.4876 | 134.7259 | 62.9431 | 0.006766 | 0.33286 | 0.002252 | 49.1955 | 0.006993 | 0.2162 | 0.001512 | 30.9169 | 14.832 | 4.3393 | 0.003764 | 158.4562 | 129.68 | 160.6556 | 119.8939
2012-10-30 20:05 | 29.3754 | 172.4075 | 143.8944 | 174.6566 | 134.1041 | 62.8345 | 0.006769 | 0.33286 | 0.002253 | 49.1737 | 0.006987 | 0.21626 | 0.001511 | 30.9518 | 14.8113 | 4.3318 | 0.003764 | 157.5962 | 129.0831 | 159.8453 | 119.2928
2012-10-30 20:05 | 29.3821 | 171.639 | 143.2148 | 173.8328 | 133.7104 | 62.7426 | 0.006769 | 0.33283 | 0.002253 | 49.1694 | 0.006977 | 0.21622 | 0.001509 | 30.99 | 14.8017 | 4.3255 | 0.003762 | 156.8373 | 128.4131 | 159.0311 | 118.9087
2012-10-30 20:05 | 29.3701 | 170.6225 | 142.4708 | 172.7996 | 132.8139 | 62.642 | 0.006779 | 0.3328 | 0.002256 | 49.0926 | 0.006971 | 0.2162 | 0.001507 | 31.0145 | 14.7897 | 4.3185 | 0.003763 | 155.8328 | 127.6811 | 158.0099 | 118.0242
2012-10-30 20:05 | 29.3851 | 169.7346 | 141.7349 | 171.9575 | 132.3182 | 62.5311 | 0.006785 | 0.33281 | 0.002258 | 49.0514 | 0.006961 | 0.21613 | 0.001504 | 31.0487 | 14.8129 | 4.3109 | 0.003762 | 154.9217 | 126.922 | 157.1446 | 117.5053
2012-10-30 20:05 | 29.3709 | 168.9658 | 141.1846 | 171.1099 | 131.702 | 62.449 | 0.006769 | 0.3328 | 0.002253 | 49.1652 | 0.006996 | 0.21617 | 0.001512 | 30.8995 | 14.7905 | 4.3052 | 0.003765 | 154.1753 | 126.3941 | 156.3194 | 116.9115
2012-10-30 20:05 | 29.3649 | 168.287 | 140.612 | 170.2631 | 131.1817 | 62.345 | 0.006754 | 0.33274 | 0.002247 | 49.2658 | 0.00698 | 0.21616 | 0.001509 | 30.9683 | 14.7723 | 4.2981 | 0.003756 | 153.5147 | 125.8397 | 155.4908 | 116.4094
2012-10-30 20:05 | 29.3724 | 167.4704 | 139.9919 | 169.5819 | 130.6942 | 62.267 | 0.006766 | 0.33278 | 0.002252 | 49.1849 | 0.006977 | 0.21612 | 0.001508 | 30.9755 | 14.7757 | 4.2927 | 0.00376 | 152.6947 | 125.2162 | 154.8062 | 115.9185
2012-10-30 20:06 | 29.3671 | 166.5482 | 139.2597 | 168.5536 | 130.1521 | 62.178 | 0.006766 | 0.33276 | 0.002251 | 49.1806 | 0.006993 | 0.21612 | 0.001511 | 30.9046 | 14.8194 | 4.2866 | 0.003762 | 151.7288 | 124.4403 | 153.7342 | 115.3327
2012-10-30 20:06 | 29.3754 | 165.6205 | 138.6179 | 167.5628 | 129.5019 | 62.0664 | 0.006766 | 0.33263 | 0.002251 | 49.1614 | 0.006983 | 0.21612 | 0.001509 | 30.9489 | 14.7705 | 4.2789 | 0.00376 | 150.85 | 123.8474 | 152.7923 | 114.7314
2012-10-30 20:06 | 29.3664 | 164.8041 | 137.9972 | 166.7961 | 128.7624 | 61.9601 | 0.006769 | 0.33261 | 0.002251 | 49.1374 | 0.006996 | 0.21609 | 0.001512 | 30.8872 | 14.7615 | 4.2715 | 0.003763 | 150.0426 | 123.2357 | 152.0346 | 114.0009
2012-10-30 20:06 | 29.3604 | 164.1036 | 137.3531 | 165.9822 | 128.1706 | 61.8653 | 0.006779 | 0.33258 | 0.002255 | 49.0606 | 0.006971 | 0.21604 | 0.001506 | 30.9917 | 14.7269 | 4.265 | 0.003761 | 149.3767 | 122.6262 | 151.2553 | 113.4437
2012-10-30 20:06 | 29.3566 | 163.2232 | 136.6807 | 165.1436 | 127.4938 | 61.7838 | 0.006763 | 0.33263 | 0.00225 | 49.1832 | 0.006977 | 0.21606 | 0.001507 | 30.9672 | 14.7068 | 4.2594 | 0.003757 | 148.5164 | 121.9739 | 150.4368 | 112.787
2012-10-30 20:06 | 29.3611 | 162.3748 | 136.12 | 164.3215 | 126.9319 | 61.6798 | 0.006773 | 0.33248 | 0.002252 | 49.0892 | 0.006968 | 0.21604 | 0.001505 | 31.005 | 14.7358 | 4.2522 | 0.003757 | 147.639 | 121.3842 | 149.5857 | 112.1961
2012-10-30 20:06 | 29.3686 | 161.7274 | 135.5053 | 163.5881 | 126.4795 | 61.6029 | 0.006779 | 0.33244 | 0.002254 | 49.0394 | 0.00699 | 0.21607 | 0.00151 | 30.9116 | 14.776 | 4.2469 | 0.003764 | 146.9514 | 120.7293 | 148.8121 | 111.7035
2012-10-30 20:06 | 29.3619 | 160.9924 | 134.9032 | 162.833 | 125.8382 | 61.5295 | 0.006766 | 0.3325 | 0.00225 | 49.1422 | 0.006977 | 0.21607 | 0.001508 | 30.9692 | 14.7407 | 4.2418 | 0.003758 | 146.2517 | 120.1625 | 148.0923 | 111.0975
2012-10-30 20:06 | 29.3604 | 160.2088 | 134.2531 | 161.9788 | 125.2593 | 61.4324 | 0.006769 | 0.33244 | 0.00225 | 49.1118 | 0.006987 | 0.21602 | 0.001509 | 30.9167 | 14.7759 | 4.2351 | 0.003759 | 145.4329 | 119.4772 | 147.2029 | 110.4834
2012-10-30 20:06 | 29.3514 | 159.3909 | 133.6072 | 161.1524 | 124.6275 | 61.3232 | 0.006766 | 0.33248 | 0.00225 | 49.14 | 0.00699 | 0.21597 | 0.00151 | 30.8971 | 14.771 | 4.2276 | 0.00376 | 144.6199 | 118.8362 | 146.3814 | 109.8565
2012-10-30 20:06 | 29.3559 | 158.7222 | 133.0322 | 160.4753 | 124.1839 | 61.2423 | 0.006776 | 0.33235 | 0.002252 | 49.0483 | 0.00699 | 0.21599 | 0.00151 | 30.8993 | 14.8 | 4.222 | 0.003762 | 143.9222 | 118.2322 | 145.6753 | 109.3839
2012-10-30 20:06 | 29.3559 | 158.0571 | 132.51 | 159.8637 | 123.7625 | 61.1805 | 0.00676 | 0.33235 | 0.002247 | 49.1643 | 0.006968 | 0.21599 | 0.001505 | 30.9968 | 14.8449 | 4.2178 | 0.003752 | 143.2122 | 117.6651 | 145.0188 | 108.9176
2012-10-30 20:06 | 29.3484 | 157.4001 | 131.9804 | 159.0667 | 123.0453 | 61.088 | 0.006763 | 0.33234 | 0.002248 | 49.1405 | 0.006983 | 0.21596 | 0.001508 | 30.9261 | 14.7884 | 4.2114 | 0.003756 | 142.6117 | 117.192 | 144.2783 | 108.2569
2012-10-30 20:06 | 29.3514 | 156.5955 | 131.3549 | 158.2691 | 122.5891 | 60.9973 | 0.006776 | 0.33224 | 0.002251 | 49.0313 | 0.006983 | 0.21594 | 0.001508 | 30.924 | 14.8037 | 4.2052 | 0.003759 | 141.7918 | 116.5512 | 143.4654 | 107.7854
2012-10-30 20:06 | 29.3462 | 155.8062 | 130.7403 | 157.4752 | 122.1438 | 60.9129 | 0.006763 | 0.33225 | 0.002247 | 49.1276 | 0.006987 | 0.21593 | 0.001509 | 30.9043 | 14.823 | 4.1993 | 0.003756 | 140.9832 | 115.9173 | 142.6522 | 107.3208
2012-10-30 20:06 | 29.3372 | 155.0945 | 130.1488 | 156.6856 | 121.6038 | 60.8412 | 0.006766 | 0.33222 | 0.002248 | 49.1015 | 0.006983 | 0.21593 | 0.001508 | 30.922 | 14.8017 | 4.1944 | 0.003756 | 140.2928 | 115.3471 | 141.8839 | 106.8021
2012-10-30 20:06 | 29.3259 | 152.5259 | 128.0449 | 153.9758 | 119.7156 | 60.528 | 0.006769 | 0.33213 | 0.002248 | 49.0669 | 0.00698 | 0.2159 | 0.001507 | 30.9311 | 14.77 | 4.1728 | 0.003755 | 137.7559 | 113.2749 | 139.2058 | 104.9456
2012-10-30 20:07 | 29.3139 | 150.0143 | 125.9954 | 151.4158 | 117.6044 | 60.2298 | 0.006766 | 0.33193 | 0.002246 | 49.0588 | 0.006977 | 0.21583 | 0.001506 | 30.9341 | 14.7907 | 4.1522 | 0.003752 | 135.2236 | 111.2047 | 136.6251 | 102.8137
2012-10-30 20:07 | 29.3117 | 149.408 | 125.5335 | 150.7743 | 117.2765 | 60.1749 | 0.006773 | 0.33203 | 0.002249 | 49.023 | 0.006971 | 0.21587 | 0.001505 | 30.9669 | 14.7885 | 4.1485 | 0.003754 | 134.6195 | 110.745 | 135.9858 | 102.488
2012-10-30 20:07 | 29.3102 | 148.7948 | 125.0078 | 150.153 | 116.9568 | 60.0974 | 0.006763 | 0.33198 | 0.002245 | 49.0871 | 0.006977 | 0.21581 | 0.001506 | 30.9319 | 14.7829 | 4.1431 | 0.003751 | 134.0119 | 110.2249 | 135.3701 | 102.1739
2012-10-30 20:07 | 29.3094 | 148.0673 | 124.4526 | 149.4441 | 116.4864 | 60.0159 | 0.006769 | 0.33189 | 0.002247 | 49.0307 | 0.006965 | 0.21583 | 0.001503 | 30.9874 | 14.7331 | 4.1375 | 0.00375 | 133.3342 | 109.7195 | 134.711 | 101.7533
2012-10-30 20:07 | 29.3094 | 147.4139 | 123.9476 | 148.7364 | 115.9374 | 59.9345 | 0.006773 | 0.3319 | 0.002248 | 49.0038 | 0.006974 | 0.21586 | 0.001505 | 30.9515 | 14.778 | 4.1319 | 0.003753 | 132.6359 | 109.1696 | 133.9584 | 101.1594
2012-10-30 20:07 | 29.3065 | 146.8193 | 123.4854 | 148.1144 | 115.5445 | 59.8605 | 0.006763 | 0.33189 | 0.002245 | 49.0742 | 0.006965 | 0.21583 | 0.001503 | 30.9874 | 14.7629 | 4.1268 | 0.003748 | 132.0564 | 108.7225 | 133.3515 | 100.7816
2012-10-30 20:07 | 29.3087 | 146.1686 | 122.9145 | 147.4324 | 114.9259 | 59.779 | 0.00676 | 0.33187 | 0.002243 | 49.0938 | 0.006993 | 0.21583 | 0.001509 | 30.8633 | 14.7406 | 4.1212 | 0.003752 | 131.428 | 108.1739 | 132.6918 | 100.1853
2012-10-30 20:07 | 29.3005 | 145.4156 | 122.3675 | 146.7211 | 114.384 | 59.7004 | 0.006763 | 0.33187 | 0.002244 | 49.072 | 0.006971 | 0.21575 | 0.001504 | 30.9502 | 14.716 | 4.1157 | 0.003748 | 130.6996 | 107.6515 | 132.0051 | 99.668
2012-10-30 20:07 | 29.3117 | 144.8166 | 121.8248 | 146.1485 | 113.9827 | 59.627 | 0.006757 | 0.33177 | 0.002242 | 49.1006 | 0.006974 | 0.21581 | 0.001505 | 30.9452 | 14.7558 | 4.1107 | 0.003747 | 130.0608 | 107.069 | 131.3927 | 99.2269
2012-10-30 20:07 | 29.305 | 144.2152 | 121.3251 | 145.589 | 113.624 | 59.5675 | 0.006766 | 0.33179 | 0.002245 | 49.0375 | 0.006993 | 0.21575 | 0.001509 | 30.8529 | 14.7532 | 4.1066 | 0.003754 | 129.462 | 106.5719 | 130.8358 | 98.8708
2012-10-30 20:07 | 29.3012 | 143.5633 | 120.8207 | 144.925 | 113.0715 | 59.4923 | 0.006766 | 0.33176 | 0.002245 | 49.0333 | 0.006971 | 0.21578 | 0.001504 | 30.9544 | 14.7249 | 4.1014 | 0.003749 | 128.8384 | 106.0958 | 130.2001 | 98.3466
2012-10-30 20:07 | 29.3035 | 142.8717 | 120.258 | 144.2099 | 112.6198 | 59.4051 | 0.006769 | 0.3318 | 0.002246 | 49.0179 | 0.006983 | 0.21583 | 0.001507 | 30.9075 | 14.7435 | 4.0954 | 0.003753 | 128.1282 | 105.5145 | 129.4664 | 97.8763
2012-10-30 20:07 | 29.2975 | 142.333 | 119.7706 | 143.6631 | 112.0728 | 59.3508 | 0.006763 | 0.33169 | 0.002243 | 49.0444 | 0.006987 | 0.21571 | 0.001507 | 30.8732 | 14.7457 | 4.0916 | 0.00375 | 127.5873 | 105.0249 | 128.9174 | 97.3271
2012-10-30 20:07 | 29.2975 | 141.7469 | 119.3196 | 143.0344 | 111.6378 | 59.275 | 0.006766 | 0.33169 | 0.002244 | 49.0226 | 0.006987 | 0.21573 | 0.001507 | 30.8753 | 14.762 | 4.0864 | 0.003751 | 126.9849 | 104.5576 | 128.2724 | 96.8758
2012-10-30 20:07 | 29.3005 | 141.2291 | 118.8564 | 142.4203 | 111.2512 | 59.2161 | 0.006785 | 0.33166 | 0.00225 | 48.8811 | 0.00698 | 0.21567 | 0.001505 | 30.898 | 14.7854 | 4.0824 | 0.003755 | 126.4437 | 104.071 | 127.6349 | 96.4658
2012-10-30 20:07 | 29.3094 | 140.5488 | 118.1606 | 141.7895 | 110.8327 | 50.8246 | 0.006769 | 0.33159 | 0.002244 | 48.9858 | 0.006971 | 0.21571 | 0.001504 | 30.9441 | 14.7984 | 3.5038 | 0.003748 | 125.7504 | 103.3622 | 126.9911 | 96.0343
2012-10-30 20:07 | 29.302 | 139.7493 | 117.4978 | 141.1502 | 110.3601 | 39.2984 | 0.006779 | 0.33154 | 0.002248 | 48.9072 | 0.006993 | 0.21561 | 0.001508 | 30.8323 | 14.8155 | 2.7092 | 0.003756 | 124.9338 | 102.6823 | 126.3347 | 95.5446
2012-10-30 20:07 | 29.2982 | 139.1411 | 116.9903 | 140.5568 | 109.8003 | 32.2073 | 0.006763 | 0.33154 | 0.002242 | 49.0229 | 0.006983 | 0.21567 | 0.001506 | 30.8847 | 14.7872 | 2.2204 | 0.003748 | 124.3539 | 102.2031 | 125.7696 | 95.0131
2012-10-30 20:07 | 29.293 | 138.547 | 116.4587 | 139.9086 | 109.209 | 26.4708 | 0.006773 | 0.3316 | 0.002246 | 48.9591 | 0.006983 | 0.2156 | 0.001506 | 30.8744 | 14.8147 | 1.8249 | 0.003752 | 123.7323 | 101.644 | 125.0939 | 94.3943
2012-10-30 20:07 | 29.293 | 137.8742 | 115.9248 | 139.1779 | 108.7401 | 21.8514 | 0.006763 | 0.33157 | 0.002242 | 49.0272 | 0.006983 | 0.21565 | 0.001506 | 30.8826 | 14.8228 | 1.5064 | 0.003748 | 123.0514 | 101.102 | 124.3551 | 93.9173
2012-10-30 20:07 | 29.2915 | 136.7952 | 115.0298 | 138.1288 | 107.9256 | 16.8206 | 0.006766 | 0.33151 | 0.002243 | 48.9969 | 0.006977 | 0.2156 | 0.001504 | 30.901 | 14.8213 | 1.1596 | 0.003747 | 121.9739 | 100.2085 | 123.3075 | 93.1043
2012-10-30 20:07 | 29.2855 | 135.8497 | 114.21 | 137.2056 | 107.1635 | 13.3761 | 0.006757 | 0.3315 | 0.00224 | 49.0601 | 0.00698 | 0.21567 | 0.001505 | 30.898 | 14.7949 | 0.9221 | 0.003745 | 121.0548 | 99.4151 | 122.4107 | 92.3686
2012-10-30 20:07 | 29.2877 | 135.245 | 113.7316 | 136.65 | 106.6742 | 11.896 | 0.006757 | 0.3315 | 0.00224 | 49.0601 | 0.00699 | 0.21561 | 0.001507 | 30.8455 | 14.8216 | 0.8201 | 0.003747 | 120.4234 | 98.91 | 121.8284 | 91.8526
2012-10-30 20:08 | 29.29 | 134.7245 | 113.2117 | 136.0489 | 106.2379 | 10.4876 | 0.006782 | 0.33137 | 0.002247 | 48.8601 | 0.006983 | 0.21557 | 0.001505 | 30.8703 | 14.7831 | 0.723 | 0.003752 | 119.9414 | 98.4286 | 121.2658 | 91.4548
2012-10-30 20:08 | 29.2862 | 134.0569 | 112.7349 | 135.4229 | 105.8334 | 9.3722 | 0.006776 | 0.33144 | 0.002246 | 48.914 | 0.006993 | 0.21561 | 0.001508 | 30.8323 | 14.7834 | 0.6461 | 0.003754 | 119.2735 | 97.9515 | 120.6395 | 91.05
2012-10-30 20:08 | 29.2847 | 133.5253 | 112.321 | 134.9481 | 105.4009 | 8.5741 | 0.006779 | 0.33132 | 0.002246 | 48.8752 | 0.006999 | 0.2156 | 0.001509 | 30.8038 | 14.7655 | 0.5911 | 0.003755 | 118.7598 | 97.5555 | 120.1826 | 90.6354
2012-10-30 20:08 | 29.2817 | 133.0075 | 111.8452 | 134.3728 | 104.8535 | 7.8101 | 0.006776 | 0.33137 | 0.002245 | 48.9033 | 0.006983 | 0.21558 | 0.001505 | 30.8723 | 14.7625 | 0.5384 | 0.00375 | 118.245 | 97.0827 | 119.6103 | 90.091
2012-10-30 20:08 | 29.278 | 132.3864 | 111.2514 | 133.7245 | 104.4604 | 6.9889 | 0.006776 | 0.33134 | 0.002245 | 48.8991 | 0.00699 | 0.21564 | 0.001507 | 30.8496 | 14.7793 | 0.4818 | 0.003752 | 117.6071 | 96.4721 | 118.9452 | 89.6811
2012-10-30 20:08 | 29.2757 | 131.7934 | 110.7574 | 133.1464 | 103.8948 | 6.3479 | 0.006763 | 0.33127 | 0.00224 | 48.9824 | 0.00699 | 0.21557 | 0.001507 | 30.8393 | 14.7361 | 0.4376 | 0.003747 | 117.0573 | 96.0213 | 118.4103 | 89.1587
2012-10-30 20:08 | 29.2757 | 131.283 | 110.2997 | 132.6737 | 103.4828 | 5.8798 | 0.006773 | 0.33125 | 0.002244 | 48.9079 | 0.006983 | 0.21558 | 0.001505 | 30.8723 | 14.7157 | 0.4054 | 0.003749 | 116.5673 | 95.584 | 117.958 | 88.7671
2012-10-30 20:08 | 29.2727 | 130.7392 | 109.824 | 132.1257 | 103.0185 | 5.3655 | 0.00676 | 0.3312 | 0.002239 | 48.9933 | 0.00698 | 0.21557 | 0.001505 | 30.8835 | 14.725 | 0.3699 | 0.003744 | 116.0142 | 95.099 | 117.4007 | 88.2935
2012-10-30 20:08 | 29.2682 | 130.057 | 109.2485 | 131.4812 | 102.5226 | 4.8546 | 0.006754 | 0.33121 | 0.002237 | 49.0391 | 0.006987 | 0.21552 | 0.001506 | 30.8463 | 14.7531 | 0.3347 | 0.003743 | 115.3039 | 94.4954 | 116.7281 | 87.7695
2012-10-30 20:08 | 29.2667 | 129.4199 | 108.7592 | 130.8665 | 102.0598 | 4.4587 | 0.006754 | 0.33125 | 0.002237 | 49.0455 | 0.006993 | 0.21554 | 0.001507 | 30.822 | 14.7435 | 0.3074 | 0.003744 | 114.6764 | 94.0157 | 116.123 | 87.3163
2012-10-30 20:08 | 29.2615 | 128.939 | 108.3721 | 130.351 | 101.6349 | 4.153 | 0.006766 | 0.33115 | 0.002241 | 48.9435 | 0.006983 | 0.21552 | 0.001505 | 30.864 | 14.7138 | 0.2863 | 0.003746 | 114.2252 | 93.6583 | 115.6372 | 86.9211
2012-10-30 20:08 | 29.2518 | 128.431 | 107.8822 | 129.8693 | 101.2547 | 3.871 | 0.006769 | 0.33111 | 0.002241 | 48.9155 | 0.006977 | 0.21544 | 0.001503 | 30.8782 | 14.6918 | 0.2669 | 0.003744 | 113.7392 | 93.1904 | 115.1775 | 86.5629
2012-10-30 20:08 | 29.2615 | 127.9462 | 107.4722 | 129.369 | 100.9771 | 3.6074 | 0.006766 | 0.33106 | 0.00224 | 48.9307 | 0.006977 | 0.21551 | 0.001504 | 30.8885 | 14.7342 | 0.2487 | 0.003744 | 113.212 | 92.738 | 114.6348 | 86.2429
2012-10-30 20:08 | 29.2585 | 127.3729 | 107.0004 | 128.8409 | 100.5203 | 3.3508 | 0.006773 | 0.33108 | 0.002242 | 48.8823 | 0.006977 | 0.21544 | 0.001503 | 30.8782 | 14.7761 | 0.231 | 0.003745 | 112.5968 | 92.2243 | 114.0648 | 85.7442
2012-10-30 20:08 | 29.2533 | 126.7784 | 106.5075 | 128.2195 | 100.1559 | 3.1012 | 0.00676 | 0.33106 | 0.002238 | 48.9741 | 0.006974 | 0.21541 | 0.001502 | 30.8873 | 14.7586 | 0.2138 | 0.00374 | 112.0198 | 91.7489 | 113.4609 | 85.3973
2012-10-30 20:08 | 29.2548 | 126.3657 | 106.112 | 127.7456 | 99.783 | 2.8798 | 0.00676 | 0.33096 | 0.002237 | 48.9592 | 0.006974 | 0.21545 | 0.001503 | 30.8935 | 14.7561 | 0.1985 | 0.00374 | 111.6096 | 91.3559 | 112.9895 | 85.0269
2012-10-30 20:08 | 29.2473 | 125.8376 | 105.7378 | 127.24 | 99.3331 | 2.7284 | 0.006763 | 0.33101 | 0.002239 | 48.9438 | 0.006971 | 0.21546 | 0.001502 | 30.9088 | 14.7445 | 0.1881 | 0.003741 | 111.0931 | 90.9933 | 112.4955 | 84.5886
2012-10-30 20:08 | 29.2518 | 125.371 | 105.3718 | 126.8416 | 98.9594 | 2.5585 | 0.006766 | 0.33092 | 0.002239 | 48.9094 | 0.006987 | 0.21539 | 0.001505 | 30.8277 | 14.8224 | 0.1764 | 0.003744 | 110.5486 | 90.5494 | 112.0192 | 84.137
2012-10-30 20:08 | 29.2443 | 124.8811 | 104.9295 | 126.321 | 98.5888 | 2.3886 | 0.006757 | 0.33091 | 0.002236 | 48.9723 | 0.006955 | 0.21538 | 0.001498 | 30.9675 | 14.8639 | 0.1647 | 0.003734 | 110.0172 | 90.0656 | 111.4571 | 83.7249
2012-10-30 20:08 | 29.2563 | 124.3729 | 104.4652 | 125.7478 | 98.2528 | 2.2296 | 0.006757 | 0.33091 | 0.002236 | 48.9723 | 0.00699 | 0.21538 | 0.001505 | 30.8124 | 14.9004 | 0.1537 | 0.003741 | 109.4725 | 89.5648 | 110.8474 | 83.3524
2012-10-30 20:08 | 29.2675 | 123.9057 | 104.1097 | 125.307 | 97.9311 | 2.0857 | 0.006741 | 0.33092 | 0.002231 | 49.0908 | 0.006971 | 0.21538 | 0.001501 | 30.8964 | 14.8586 | 0.1438 | 0.003732 | 109.0471 | 89.2511 | 110.4484 | 83.0725
2012-10-30 20:08 | 29.2682 | 123.4471 | 103.7286 | 124.8481 | 97.6141 | 1.958 | 0.006766 | 0.33093 | 0.002239 | 48.9115 | 0.006977 | 0.21541 | 0.001503 | 30.874 | 14.8552 | 0.135 | 0.003742 | 108.5919 | 88.8734 | 109.9929 | 82.7589
2012-10-30 20:08 | 29.26 | 122.9076 | 103.3083 | 124.3348 | 97.2162 | 1.8291 | 0.006769 | 0.33088 | 0.00224 | 48.8812 | 0.006977 | 0.21538 | 0.001503 | 30.8699 | 14.8388 | 0.1261 | 0.003743 | 108.0688 | 88.4695 | 109.496 | 82.3774
2012-10-30 20:08 | 29.2563 | 122.3461 | 102.9001 | 123.8603 | 96.7434 | 1.7089 | 0.006754 | 0.33085 | 0.002235 | 48.9856 | 0.006971 | 0.21536 | 0.001501 | 30.8943 | 14.8269 | 0.1178 | 0.003736 | 107.5192 | 88.0732 | 109.0334 | 81.9165
2012-10-30 20:09 | 29.2645 | 121.9104 | 102.5113 | 123.3713 | 96.3468 | 1.6211 | 0.006766 | 0.33079 | 0.002238 | 48.8902 | 0.00699 | 0.21538 | 0.001505 | 30.8124 | 14.8066 | 0.1118 | 0.003743 | 107.1038 | 87.7047 | 108.5647 | 81.5402
2012-10-30 20:09 | 29.2667 | 121.5029 | 102.1165 | 122.8877 | 95.9176 | 1.5015 | 0.006769 | 0.33075 | 0.002239 | 48.862 | 0.006983 | 0.21531 | 0.001503 | 30.833 | 14.8455 | 0.1035 | 0.003742 | 106.6574 | 87.271 | 108.0422 | 81.0721
2012-10-30 20:09 | 29.2682 | 121.0417 | 101.7248 | 122.4452 | 95.5028 | 1.4119 | 0.006757 | 0.3307 | 0.002235 | 48.9424 | 0.006974 | 0.21533 | 0.001502 | 30.8768 | 14.8348 | 0.0973 | 0.003737 | 106.2069 | 86.89 | 107.6104 | 80.668
2012-10-30 20:09 | 29.2622 | 120.5163 | 101.2651 | 121.9422 | 94.9745 | 1.3183 | 0.006754 | 0.33069 | 0.002233 | 48.9621 | 0.006974 | 0.21526 | 0.001501 | 30.8665 | 14.8206 | 0.0909 | 0.003734 | 105.6957 | 86.4445 | 107.1216 | 80.1539
2012-10-30 20:09 | 29.2682 | 120.0256 | 100.7985 | 121.4627 | 94.761 | 0.9831 | 0.006766 | 0.3307 | 0.002238 | 48.8773 | 0.00698 | 0.21536 | 0.001503 | 30.8544 | 14.8552 | 0.0678 | 0.003741 | 105.1704 | 85.9433 | 106.6075 | 79.9058
2012-10-30 20:09 | 29.2705 | 119.5579 | 100.4642 | 121.0403 | 94.4339 | -0.6375 | 0.006757 | 0.33073 | 0.002235 | 48.9467 | 0.006971 | 0.21531 | 0.001501 | 30.8861 | 14.8534 | -0.0439 | 0.003736 | 104.7045 | 85.6108 | 106.1869 | 79.5805
2012-10-30 20:09 | 29.2578 | 119.06 | 100.1831 | 120.6029 | 94.16 | -0.9911 | 0.006779 | 0.33069 | 0.002242 | 48.7815 | 0.006977 | 0.21535 | 0.001502 | 30.8657 | 14.8203 | -0.0683 | 0.003744 | 104.2397 | 85.3628 | 105.7826 | 79.3397
2012-10-30 20:09 | 29.263 | 118.5309 | 99.8366 | 120.1645 | 93.8207 | -1.0865 | 0.006763 | 0.33066 | 0.002236 | 48.8927 | 0.00698 | 0.21532 | 0.001503 | 30.8483 | 14.8214 | -0.0749 | 0.003739 | 103.7095 | 85.0152 | 105.3431 | 78.9993
2012-10-30 20:09 | 29.263 | 118.0422 | 99.4433 | 119.6982 | 93.4534 | -1.1934 | 0.006769 | 0.33065 | 0.002238 | 48.8471 | 0.006974 | 0.21539 | 0.001502 | 30.8851 | 14.8418 | -0.0823 | 0.00374 | 103.2004 | 84.6015 | 104.8564 | 78.6116
2012-10-30 20:09 | 29.2652 | 117.6769 | 99.1505 | 119.3062 | 93.0542 | -1.2333 | 0.006769 | 0.33067 | 0.002238 | 48.8514 | 0.006971 | 0.21533 | 0.001501 | 30.8901 | 14.8767 | -0.085 | 0.003739 | 102.8002 | 84.2738 | 104.4295 | 78.1775
2012-10-30 20:09 | 29.263 | 117.2543 | 98.8571 | 118.8909 | 92.6429 | -1.2824 | 0.006763 | 0.33069 | 0.002236 | 48.8969 | 0.006971 | 0.21533 | 0.001501 | 30.8901 | 14.8541 | -0.0884 | 0.003737 | 102.4002 | 84.003 | 104.0368 | 77.7888
2012-10-30 20:09 | 29.2578 | 116.9044 | 98.4964 | 118.4803 | 92.2362 | -1.3102 | 0.006766 | 0.33069 | 0.002237 | 48.8753 | 0.006983 | 0.21532 | 0.001504 | 30.835 | 14.8611 | -0.0903 | 0.003741 | 102.0433 | 83.6353 | 103.6192 | 77.3751
2012-10-30 20:09 | 29.2555 | 116.4704 | 98.0478 | 118.1029 | 91.8058 | -1.3362 | 0.006779 | 0.33062 | 0.002241 | 48.7708 | 0.006974 | 0.21535 | 0.001502 | 30.879 | 14.8302 | -0.0921 | 0.003743 | 101.6402 | 83.2176 | 103.2727 | 76.9756
2012-10-30 20:09 | 29.248 | 116.0804 | 97.696 | 117.6749 | 91.1428 | -1.3523 | 0.00676 | 0.33062 | 0.002235 | 48.9078 | 0.006977 | 0.21528 | 0.001502 | 30.8554 | 14.8023 | -0.0932 | 0.003737 | 101.2781 | 82.8937 | 102.8726 | 76.3405
2012-10-30 20:09 | 29.2525 | 115.7555 | 97.3563 | 117.365 | 90.5068 | -1.3876 | 0.006773 | 0.33056 | 0.002239 | 48.8056 | 0.006968 | 0.21533 | 0.0015 | 30.9034 | 14.8476 | -0.0957 | 0.003739 | 100.9079 | 82.5087 | 102.5174 | 75.6592
2012-10-30 20:09 | 29.2495 | 115.4383 | 97.1036 | 117.0665 | 90.3446 | -1.3974 | 0.00676 | 0.33057 | 0.002235 | 48.9015 | 0.006974 | 0.21525 | 0.001501 | 30.8645 | 14.861 | -0.0963 | 0.003736 | 100.5773 | 82.2426 | 102.2055 | 75.4836
2012-10-30 20:09 | 29.2525 | 115.1536 | 96.785 | 116.7552 | 90.0709 | -1.4165 | 0.006769 | 0.33062 | 0.002238 | 48.8428 | 0.006958 | 0.21526 | 0.001498 | 30.9375 | 14.8109 | -0.0977 | 0.003736 | 100.3427 | 81.9741 | 101.9443 | 75.26
2012-10-30 20:09 | 29.2578 | 114.875 | 96.4687 | 116.4235 | 89.7501 | -1.4269 | 0.006773 | 0.33053 | 0.002239 | 48.8013 | 0.006977 | 0.21536 | 0.001503 | 30.8677 | 14.8162 | -0.0984 | 0.003742 | 100.0588 | 81.6525 | 101.6073 | 74.9339
2012-10-30 20:09 | 29.2465 | 114.5646 | 96.1698 | 116.1092 | 89.3596 | -1.4512 | 0.006754 | 0.33059 | 0.002233 | 48.947 | 0.006974 | 0.21531 | 0.001502 | 30.8728 | 14.8253 | -0.1 | 0.003735 | 99.7393 | 81.3445 | 101.2839 | 74.5343
2012-10-30 20:09 | 29.2495 | 114.3633 | 95.9269 | 115.8926 | 89.0175 | -1.4518 | 0.006754 | 0.3305 | 0.002232 | 48.9343 | 0.006987 | 0.21531 | 0.001504 | 30.8154 | 14.7997 | -0.1001 | 0.003736 | 99.5636 | 81.1272 | 101.0929 | 74.2178
2012-10-30 20:09 | 29.2443 | 114.1385 | 95.6303 | 115.5883 | 88.633 | -1.4593 | 0.006766 | 0.33056 | 0.002237 | 48.856 | 0.00698 | 0.21536 | 0.001503 | 30.8544 | 14.8313 | -0.1006 | 0.00374 | 99.3072 | 80.799 | 100.757 | 73.8017
2012-10-30 20:09 | 29.2443 | 113.9115 | 95.4335 | 115.3612 | 88.1891 | -1.4587 | 0.006763 | 0.33052 | 0.002235 | 48.8712 | 0.00699 | 0.21535 | 0.001505 | 30.8083 | 14.766 | -0.1006 | 0.00374 | 99.1455 | 80.6675 | 100.5952 | 73.4231
2012-10-30 20:09 | 29.2443 | 113.7071 | 95.1723 | 115.1151 | 87.6503 | -1.4598 | 0.006751 | 0.33052 | 0.002231 | 48.9581 | 0.006968 | 0.21523 | 0.0015 | 30.8889 | 14.7986 | -0.1006 | 0.003731 | 98.9085 | 80.3737 | 100.3165 | 72.8517
2012-10-30 20:09 | 29.2488 | 113.5111 | 94.8513 | 114.8963 | 87.2677 | -1.4702 | 0.00676 | 0.3305 | 0.002234 | 48.8908 | 0.006974 | 0.21533 | 0.001502 | 30.8768 | 14.799 | -0.1014 | 0.003736 | 98.7121 | 80.0523 | 100.0973 | 72.4687
2012-10-30 20:10 | 29.2428 | 113.3575 | 94.6031 | 114.7199 | 86.8934 | -1.4645 | 0.006769 | 0.3305 | 0.002237 | 48.8258 | 0.006983 | 0.21531 | 0.001503 | 30.833 | 14.8012 | -0.101 | 0.00374 | 98.5563 | 79.8019 | 99.9187 | 72.0922
2012-10-30 20:10 | 29.2458 | 113.1561 | 94.3032 | 114.4996 | 86.5851 | -1.4691 | 0.006757 | 0.33054 | 0.002233 | 48.9189 | 0.00698 | 0.21531 | 0.001503 | 30.8463 | 14.796 | -0.1013 | 0.003736 | 98.3601 | 79.5072 | 99.7036 | 71.7891
2012-10-30 20:10 | 29.2398 | 112.9383 | 94.0019 | 114.293 | 86.0969 | -1.4795 | 0.006766 | 0.33057 | 0.002237 | 48.8581 | 0.00699 | 0.21532 | 0.001505 | 30.8041 | 14.7778 | -0.102 | 0.003742 | 98.1605 | 79.2241 | 99.5152 | 71.3191
2012-10-30 20:10 | 29.242 | 112.8156 | 93.6899 | 114.0265 | 85.6243 | -1.4575 | 0.00676 | 0.33059 | 0.002235 | 48.9036 | 0.006974 | 0.21535 | 0.001502 | 30.879 | 14.7637 | -0.1005 | 0.003737 | 98.0519 | 78.9262 | 99.2628 | 70.8606
2012-10-30 20:10 | 29.2383 | 112.6151 | 93.406 | 113.8752 | 85.2141 | -1.4598 | 0.00676 | 0.33049 | 0.002234 | 48.8886 | 0.006977 | 0.21531 | 0.001502 | 30.8595 | 14.7355 | -0.1006 | 0.003736 | 97.8796 | 78.6705 | 99.1397 | 70.4786
2012-10-30 20:10 | 29.2285 | 112.3935 | 93.0818 | 113.7216 | 84.8927 | -1.4708 | 0.00676 | 0.33052 | 0.002234 | 48.8929 | 0.006977 | 0.21535 | 0.001502 | 30.8657 | 14.7338 | -0.1014 | 0.003736 | 97.6597 | 78.348 | 98.9878 | 70.1589
2012-10-30 20:10 | 29.236 | 112.2383 | 92.8015 | 113.5323 | 84.5315 | -1.4674 | 0.00676 | 0.33046 | 0.002234 | 48.8843 | 0.006974 | 0.21533 | 0.001502 | 30.8768 | 14.7373 | -0.1012 | 0.003736 | 97.501 | 78.0642 | 98.795 | 69.7942
2012-10-30 20:10 | 29.2353 | 112.0484 | 92.3842 | 113.297 | 84.0784 | -1.4737 | 0.006766 | 0.33044 | 0.002236 | 48.8389 | 0.006974 | 0.21532 | 0.001502 | 30.8748 | 14.7692 | -0.1016 | 0.003738 | 97.2792 | 77.615 | 98.5278 | 69.3092
2012-10-30 20:10 | 29.2323 | 111.8298 | 92.0327 | 113.0821 | 83.6648 | -1.4731 | 0.006754 | 0.3305 | 0.002232 | 48.9343 | 0.006987 | 0.21533 | 0.001505 | 30.8194 | 14.7499 | -0.1016 | 0.003737 | 97.0799 | 77.2828 | 98.3322 | 68.9149
2012-10-30 20:10 | 29.2315 | 111.6361 | 91.6833 | 112.9224 | 83.3217 | -1.4818 | 0.006766 | 0.33044 | 0.002236 | 48.8389 | 0.006971 | 0.21525 | 0.001501 | 30.8778 | 14.7654 | -0.1022 | 0.003737 | 96.8707 | 76.9179 | 98.157 | 68.5563
2012-10-30 20:10 | 29.2323 | 111.4553 | 91.4378 | 112.7605 | 82.9992 | -1.4847 | 0.00676 | 0.3304 | 0.002234 | 48.8759 | 0.006958 | 0.21532 | 0.001498 | 30.9458 | 14.7295 | -0.1024 | 0.003732 | 96.7258 | 76.7083 | 98.031 | 68.2697
2012-10-30 20:10 | 29.2383 | 111.2419 | 91.0156 | 112.5395 | 82.5752 | -1.483 | 0.006769 | 0.3304 | 0.002236 | 48.8109 | 0.00698 | 0.21535 | 0.001503 | 30.8524 | 14.711 | -0.1022 | 0.003739 | 96.5309 | 76.3046 | 97.8285 | 67.8642
2012-10-30 20:10 | 29.2375 | 111.0407 | 90.651 | 112.3344 | 82.1063 | -1.483 | 0.006766 | 0.3304 | 0.002235 | 48.8325 | 0.00698 | 0.21532 | 0.001503 | 30.8483 | 14.6857 | -0.1022 | 0.003738 | 96.355 | 75.9653 | 97.6487 | 67.4206
2012-10-30 20:10 | 29.2375 | 110.8591 | 90.3591 | 112.1906 | 81.7637 | -1.4789 | 0.006757 | 0.33036 | 0.002232 | 48.8911 | 0.00698 | 0.21525 | 0.001502 | 30.838 | 14.6816 | -0.102 | 0.003734 | 96.1775 | 75.6775 | 97.509 | 67.0821
2012-10-30 20:10 | 29.2263 | 110.6815 | 90.0598 | 112.0092 | 81.4784 | -1.4731 | 0.006766 | 0.33039 | 0.002235 | 48.8303 | 0.006983 | 0.21528 | 0.001503 | 30.8289 | 14.699 | -0.1016 | 0.003738 | 95.9825 | 75.3608 | 97.3102 | 66.7794
2012-10-30 20:10 | 29.236 | 110.4643 | 89.751 | 111.8486 | 81.1187 | -1.4743 | 0.00676 | 0.33037 | 0.002233 | 48.8716 | 0.00698 | 0.21531 | 0.001503 | 30.8463 | 14.7005 | -0.1016 | 0.003736 | 95.7638 | 75.0505 | 97.1481 | 66.4182
2012-10-30 20:10 | 29.2383 | 110.2623 | 89.4462 | 111.6769 | 80.7744 | -1.476 | 0.006773 | 0.33034 | 0.002237 | 48.7735 | 0.006983 | 0.21523 | 0.001503 | 30.8226 | 14.6783 | -0.1018 | 0.00374 | 95.584 | 74.7679 | 96.9986 | 66.0961
2012-10-30 20:10 | 29.2353 | 110.1232 | 89.0754 | 111.4508 | 80.459 | -1.4864 | 0.00676 | 0.33042 | 0.002234 | 48.878 | 0.006968 | 0.21529 | 0.0015 | 30.8972 | 14.7039 | -0.1025 | 0.003734 | 95.4193 | 74.3715 | 96.7469 | 65.7551
2012-10-30 20:10 | 29.239 | 109.934 | 88.7567 | 111.2653 | 80.0892 | -1.4708 | 0.006779 | 0.33037 | 0.00224 | 48.7346 | 0.00698 | 0.21532 | 0.001503 | 30.8483 | 14.7198 | -0.1014 | 0.003743 | 95.2142 | 74.0369 | 96.5455 | 65.3694
2012-10-30 20:10 | 29.2353 | 109.7261 | 88.495 | 111.0838 | 79.7692 | -1.4812 | 0.006744 | 0.33026 | 0.002227 | 48.9703 | 0.006965 | 0.21531 | 0.0015 | 30.9127 | 14.6917 | -0.1021 | 0.003727 | 95.0344 | 73.8033 | 96.3921 | 65.0775
2012-10-30 20:10 | 29.242 | 109.5702 | 88.183 | 110.9468 | 79.429 | -1.4726 | 0.006751 | 0.33036 | 0.00223 | 48.9345 | 0.00698 | 0.21523 | 0.001502 | 30.8358 | 14.7228 | -0.1015 | 0.003732 | 94.8474 | 73.4602 | 96.224 | 64.7062
2012-10-30 20:10 | 29.2338 | 109.3729 | 87.8219 | 110.7608 | 79.0966 | -1.4783 | 0.00676 | 0.33034 | 0.002233 | 48.8673 | 0.006974 | 0.21532 | 0.001502 | 30.8748 | 14.7146 | -0.1019 | 0.003735 | 94.6583 | 73.1073 | 96.0462 | 64.382
2012-10-30 20:10 | 29.2338 | 109.0931 | 87.4918 | 110.6096 | 78.7228 | -1.4743 | 0.006763 | 0.33036 | 0.002234 | 48.8477 | 0.00698 | 0.21526 | 0.001503 | 30.84 | 14.6697 | -0.1016 | 0.003737 | 94.4234 | 72.8221 | 95.9399 | 64.0531
2012-10-30 20:10 | 29.2353 | 108.9131 | 87.1403 | 110.403 | 78.3963 | -1.4835 | 0.006763 | 0.33036 | 0.002234 | 48.8477 | 0.006968 | 0.21532 | 0.0015 | 30.9014 | 14.6998 | -0.1023 | 0.003734 | 94.2133 | 72.4405 | 95.7032 | 63.6965
2012-10-30 20:10 | 29.2338 | 108.7338 | 86.8199 | 110.2351 | 78.0857 | -1.4795 | 0.006763 | 0.33033 | 0.002234 | 48.8434 | 0.006955 | 0.21529 | 0.001497 | 30.955 | 14.7024 | -0.102 | 0.003731 | 94.0314 | 72.1175 | 95.5327 | 63.3833
2012-10-30 20:10 | 29.233 | 108.5856 | 86.5647 | 110.049 | 77.7833 | -1.4806 | 0.00676 | 0.33036 | 0.002233 | 48.8694 | 0.006983 | 0.21532 | 0.001504 | 30.835 | 14.6894 | -0.1021 | 0.003737 | 93.8962 | 71.8753 | 95.3596 | 63.0939
2012-10-30 20:11 | 29.2315 | 108.3723 | 86.2481 | 109.8584 | 77.3924 | -1.4754 | 0.006757 | 0.3303 | 0.002232 | 48.8826 | 0.006968 | 0.21523 | 0.0015 | 30.8889 | 14.7246 | -0.1017 | 0.003732 | 93.6477 | 71.5235 | 95.1338 | 62.6678
2012-10-30 20:11 | 29.2338 | 108.1175 | 85.9085 | 109.6489 | 77.0549 | -1.4714 | 0.006757 | 0.33031 | 0.002232 | 48.8847 | 0.006974 | 0.21535 | 0.001502 | 30.879 | 14.7269 | -0.1014 | 0.003734 | 93.3906 | 71.1816 | 94.922 | 62.328
2012-10-30 20:11 | 29.2315 | 107.9413 | 85.61 | 109.4387 | 76.7967 | -1.4772 | 0.006766 | 0.33028 | 0.002235 | 48.8154 | 0.006952 | 0.21529 | 0.001497 | 30.9684 | 14.7083 | -0.1018 | 0.003732 | 93.233 | 70.9017 | 94.7304 | 62.0884
2012-10-30 20:11 | 29.227 | 107.7288 | 85.3091 | 109.298 | 76.5286 | -1.4864 | 0.006763 | 0.33027 | 0.002234 | 48.835 | 0.006965 | 0.21531 | 0.0015 | 30.9127 | 14.7283 | -0.1025 | 0.003734 | 93.0005 | 70.5808 | 94.5697 | 61.8003
2012-10-30 20:11 | 29.2218 | 107.527 | 84.9886 | 109.1075 | 76.1986 | -1.487 | 0.00676 | 0.33027 | 0.002233 | 48.8567 | 0.00698 | 0.21526 | 0.001503 | 30.84 | 14.7067 | -0.1025 | 0.003736 | 92.8203 | 70.2819 | 94.4008 | 61.4919
2012-10-30 20:11 | 29.23 | 107.3386 | 84.6927 | 108.9493 | 75.9202 | -1.4876 | 0.006744 | 0.33024 | 0.002227 | 48.9683 | 0.006971 | 0.21528 | 0.001501 | 30.8819 | 14.7557 | -0.1026 | 0.003728 | 92.5829 | 69.937 | 94.1936 | 61.1645
2012-10-30 20:11 | 29.2293 | 107.1753 | 84.4754 | 108.805 | 75.5983 | -1.4818 | 0.006754 | 0.3302 | 0.00223 | 48.8893 | 0.006971 | 0.21531 | 0.001501 | 30.8861 | 14.7632 | -0.1022 | 0.003731 | 92.4121 | 69.7122 | 94.0418 | 60.8351
2012-10-30 20:11 | 29.2166 | 106.9093 | 84.2308 | 108.6108 | 75.3868 | -1.4772 | 0.006766 | 0.33024 | 0.002234 | 48.809 | 0.006983 | 0.21526 | 0.001503 | 30.8267 | 14.7219 | -0.1018 | 0.003737 | 92.1874 | 69.5089 | 93.8889 | 60.6649
2012-10-30 20:11 | 29.2263 | 106.7564 | 84.001 | 108.4768 | 75.1977 | -1.4812 | 0.006769 | 0.33024 | 0.002235 | 48.7874 | 0.006971 | 0.21526 | 0.001501 | 30.8798 | 14.7398 | -0.1021 | 0.003736 | 92.0166 | 69.2612 | 93.737 | 60.4579
2012-10-30 20:11 | 29.2225 | 106.575 | 83.6854 | 108.2612 | 74.819 | -1.4858 | 0.006779 | 0.3302 | 0.002238 | 48.709 | 0.006977 | 0.21533 | 0.001502 | 30.8636 | 14.736 | -0.1024 | 0.00374 | 91.839 | 68.9494 | 93.5252 | 60.083
2012-10-30 20:11 | 29.2278 | 106.3496 | 83.3636 | 108.0586 | 74.4685 | -1.4882 | 0.006766 | 0.33024 | 0.002234 | 48.809 | 0.00698 | 0.21532 | 0.001503 | 30.8483 | 14.7331 | -0.1026 | 0.003737 | 91.6165 | 68.6305 | 93.3255 | 59.7354
2012-10-30 20:11 | 29.2285 | 106.1689 | 83.1285 | 107.9005 | 74.2549 | -1.4731 | 0.00676 | 0.3302 | 0.002232 | 48.8459 | 0.006965 | 0.21525 | 0.001499 | 30.9044 | 14.7951 | -0.1016 | 0.003731 | 91.3738 | 68.3334 | 93.1054 | 59.4598
2012-10-30 20:11 | 29.2315 | 105.9677 | 82.8652 | 107.7371 | 73.9134 | -1.4968 | 0.006766 | 0.33015 | 0.002234 | 48.7962 | 0.006983 | 0.21532 | 0.001504 | 30.835 | 14.7246 | -0.1032 | 0.003738 | 91.2431 | 68.1406 | 93.0125 | 59.1888
2012-10-30 20:11 | 29.224 | 105.7561 | 82.618 | 107.5784 | 73.5422 | -1.4812 | 0.006751 | 0.33017 | 0.002229 | 48.9067 | 0.00698 | 0.21528 | 0.001503 | 30.8421 | 14.7334 | -0.1021 | 0.003732 | 91.0227 | 67.8846 | 92.845 | 58.8088
2012-10-30 20:11 | 29.2293 | 105.5648 | 82.3379 | 107.3682 | 73.1262 | -1.4737 | 0.006754 | 0.33018 | 0.00223 | 48.8872 | 0.006971 | 0.21523 | 0.0015 | 30.8756 | 14.7306 | -0.1016 | 0.00373 | 90.8342 | 67.6073 | 92.6376 | 58.3956
2012-10-30 20:11 | 29.23 | 105.3349 | 82.0113 | 107.1798 | 72.9315 | -1.487 | 0.006741 | 0.33014 | 0.002225 | 48.9751 | 0.006971 | 0.21531 | 0.001501 | 30.8861 | 14.7476 | -0.1025 | 0.003726 | 90.5873 | 67.2637 | 92.4322 | 58.1839
2012-10-30 20:11 | 29.23 | 105.1458 | 81.7296 | 107.0172 | 72.7437 | -1.4824 | 0.00676 | 0.33013 | 0.002232 | 48.8352 | 0.006983 | 0.21531 | 0.001503 | 30.833 | 14.7925 | -0.1022 | 0.003735 | 90.3533 | 66.9371 | 92.2247 | 57.9512
2012-10-30 20:11 | 29.233 | 104.9182 | 81.4737 | 106.8274 | 72.5475 | -1.4928 | 0.006763 | 0.3302 | 0.002233 | 48.8242 | 0.006968 | 0.21531 | 0.0015 | 30.8994 | 14.7914 | -0.1029 | 0.003733 | 90.1268 | 66.6823 | 92.036 | 57.7561
2012-10-30 20:11 | 29.2248 | 104.7437 | 81.1723 | 106.6566 | 72.2901 | -1.5061 | 0.006751 | 0.33015 | 0.002229 | 48.9046 | 0.006974 | 0.21538 | 0.001502 | 30.8831 | 14.7669 | -0.1038 | 0.003731 | 89.9768 | 66.4054 | 91.8897 | 57.5232
2012-10-30 20:11 | 29.2248 | 104.5093 | 80.9133 | 106.4638 | 71.9451 | -1.4905 | 0.006754 | 0.33017 | 0.00223 | 48.885 | 0.006977 | 0.21528 | 0.001502 | 30.8554 | 14.7832 | -0.1028 | 0.003732 | 89.7261 | 66.1301 | 91.6806 | 57.1619
2012-10-30 20:11 | 29.2225 | 104.2613 | 80.6481 | 106.2271 | 71.6283 | -1.4835 | 0.00676 | 0.33007 | 0.002231 | 48.8266 | 0.006987 | 0.21525 | 0.001504 | 30.8071 | 14.7523 | -0.1023 | 0.003735 | 89.509 | 65.8958 | 91.4748 | 56.876
2012-10-30 20:11 | 29.2233 | 104.0806 | 80.3746 | 106.02 | 71.3798 | -1.4893 | 0.006773 | 0.33011 | 0.002236 | 48.7394 | 0.006971 | 0.21535 | 0.001501 | 30.8923 | 14.8021 | -0.1027 | 0.003737 | 89.2785 | 65.5725 | 91.2179 | 56.5777
2012-10-30 20:11 | 29.2166 | 103.8925 | 80.1202 | 105.8659 | 71.1506 | -1.4893 | 0.006763 | 0.33001 | 0.002232 | 48.7965 | 0.006993 | 0.21535 | 0.001506 | 30.7951 | 14.7464 | -0.1027 | 0.003738 | 89.1461 | 65.3738 | 91.1195 | 56.4042
2012-10-30 20:11 | 29.2121 | 103.6422 | 79.8375 | 105.695 | 70.9772 | -1.4957 | 0.006747 | 0.33008 | 0.002227 | 48.9229 | 0.006977 | 0.21529 | 0.001502 | 30.8574 | 14.7419 | -0.1031 | 0.003729 | 88.9003 | 65.0956 | 90.9531 | 56.2353
2012-10-30 20:11 | 29.2091 | 103.4237 | 79.6057 | 105.4954 | 70.6288 | -1.4783 | 0.006757 | 0.33004 | 0.00223 | 48.844 | 0.00699 | 0.21526 | 0.001505 | 30.7959 | 14.743 | -0.1019 | 0.003735 | 88.6807 | 64.8627 | 90.7524 | 55.8858
2012-10-30 20:12 | 29.2166 | 103.3103 | 79.354 | 105.2988 | 70.3638 | -1.4824 | 0.00676 | 0.33005 | 0.002231 | 48.8246 | 0.006983 | 0.21528 | 0.001503 | 30.8289 | 14.6974 | -0.1022 | 0.003734 | 88.6129 | 64.6566 | 90.6014 | 55.6664
2012-10-30 20:12 | 29.2188 | 103.0516 | 79.1541 | 105.1762 | 70.1356 | -1.4749 | 0.006738 | 0.33005 | 0.002224 | 48.984 | 0.006958 | 0.21529 | 0.001498 | 30.9416 | 14.7323 | -0.1017 | 0.003722 | 88.3193 | 64.4218 | 90.4439 | 55.4033
2012-10-30 20:12 | 29.2076 | 102.9005 | 78.8988 | 105.0176 | 69.9247 | -1.4916 | 0.006763 | 0.33 | 0.002232 | 48.7943 | 0.006968 | 0.21526 | 0.0015 | 30.8931 | 14.7252 | -0.1028 | 0.003732 | 88.1753 | 64.1736 | 90.2924 | 55.1995
2012-10-30 20:12 | 29.2038 | 101.7852 | 77.6129 | 103.9666 | 68.6527 | -1.4882 | 0.006754 | 0.33 | 0.002229 | 48.8593 | 0.006974 | 0.21528 | 0.001501 | 30.8687 | 14.701 | -0.1026 | 0.00373 | 87.0842 | 62.9119 | 89.2656 | 53.9517
2012-10-30 20:12 | 29.1986 | 100.6381 | 76.29 | 102.9293 | 67.2123 | -1.4934 | 0.006751 | 0.32994 | 0.002227 | 48.8725 | 0.006968 | 0.21522 | 0.0015 | 30.8869 | 14.7937 | -0.103 | 0.003727 | 85.8444 | 61.4963 | 88.1356 | 52.4186
2012-10-30 20:12 | 29.2016 | 100.4294 | 76.0905 | 102.7546 | 67.0421 | -1.4835 | 0.00676 | 0.32995 | 0.00223 | 48.8096 | 0.006968 | 0.21529 | 0.0015 | 30.8972 | 14.8702 | -0.1023 | 0.00373 | 85.5592 | 61.2203 | 87.8844 | 52.1719
2012-10-30 20:12 | 29.1948 | 100.1692 | 75.8582 | 102.5664 | 66.8735 | -1.4858 | 0.006763 | 0.32995 | 0.002231 | 48.788 | 0.006971 | 0.21519 | 0.0015 | 30.8695 | 14.8593 | -0.1024 | 0.003731 | 85.3099 | 60.9989 | 87.7071 | 52.0142
2012-10-30 20:12 | 29.2053 | 99.9755 | 75.6279 | 102.35 | 66.6761 | -1.4766 | 0.006757 | 0.32987 | 0.002229 | 48.8184 | 0.006974 | 0.21523 | 0.001501 | 30.8623 | 14.9392 | -0.1018 | 0.00373 | 85.0363 | 60.6887 | 87.4108 | 51.7369
2012-10-30 20:12 | 29.1911 | 99.7912 | 75.4531 | 102.1165 | 66.4615 | -1.4899 | 0.006757 | 0.33001 | 0.00223 | 48.8399 | 0.006968 | 0.21521 | 0.0015 | 30.8848 | 14.9209 | -0.1027 | 0.00373 | 84.8703 | 60.5322 | 87.1956 | 51.5406
2012-10-30 20:12 | 29.1888 | 99.5809 | 75.2175 | 101.9857 | 66.2743 | -1.4778 | 0.006766 | 0.32991 | 0.002232 | 48.7598 | 0.006977 | 0.21525 | 0.001502 | 30.8512 | 14.9349 | -0.1019 | 0.003734 | 84.646 | 60.2826 | 87.0508 | 51.3394
2012-10-30 20:12 | 29.1881 | 99.3987 | 74.9644 | 101.7581 | 66.0154 | -1.4841 | 0.006754 | 0.32994 | 0.002228 | 48.8508 | 0.006974 | 0.21523 | 0.001501 | 30.8623 | 14.9179 | -0.1023 | 0.003729 | 84.4808 | 60.0465 | 86.8402 | 51.0975
2012-10-30 20:12 | 29.1911 | 99.1634 | 74.7302 | 101.5872 | 65.7987 | -1.4893 | 0.006757 | 0.32984 | 0.002229 | 48.8141 | 0.006974 | 0.21523 | 0.001501 | 30.8623 | 14.8719 | -0.1027 | 0.00373 | 84.2915 | 59.8583 | 86.7153 | 50.9268
2012-10-30 20:12 | 29.1956 | 98.9372 | 74.5358 | 101.3838 | 65.595 | -1.487 | 0.006757 | 0.32989 | 0.002229 | 48.8227 | 0.00698 | 0.21528 | 0.001503 | 30.8421 | 14.8683 | -0.1025 | 0.003732 | 84.0689 | 59.6675 | 86.5155 | 50.7267
2012-10-30 20:12 | 29.1888 | 98.7413 | 74.3759 | 101.2371 | 65.3954 | -1.483 | 0.006773 | 0.32979 | 0.002234 | 48.6925 | 0.006965 | 0.21525 | 0.001499 | 30.9044 | 14.788 | -0.1022 | 0.003733 | 83.9533 | 59.5879 | 86.4491 | 50.6074
2012-10-30 20:12 | 29.1798 | 98.547 | 74.1105 | 101.0315 | 65.155 | -1.476 | 0.006763 | 0.32981 | 0.00223 | 48.7665 | 0.006983 | 0.21528 | 0.001503 | 30.8289 | 14.7913 | -0.1018 | 0.003733 | 83.7557 | 59.3192 | 86.2402 | 50.3637
2012-10-30 20:12 | 29.1896 | 98.3601 | 73.9059 | 100.8447 | 64.8408 | -1.4812 | 0.006754 | 0.32982 | 0.002228 | 48.8337 | 0.006974 | 0.21525 | 0.001501 | 30.8645 | 14.8419 | -0.1021 | 0.003729 | 83.5182 | 59.064 | 86.0028 | 49.9989
2012-10-30 20:13 | 29.1903 | 98.1225 | 73.6846 | 100.645 | 64.6215 | -1.4841 | 0.006763 | 0.32981 | 0.00223 | 48.7665 | 0.006965 | 0.21525 | 0.001499 | 30.9044 | 14.8711 | -0.1023 | 0.003729 | 83.2514 | 58.8135 | 85.7739 | 49.7504
2012-10-30 20:13 | 29.1821 | 97.8495 | 73.462 | 100.425 | 64.2737 | -1.4945 | 0.006763 | 0.32974 | 0.00223 | 48.7559 | 0.006971 | 0.21515 | 0.0015 | 30.8633 | 14.8425 | -0.103 | 0.00373 | 83.007 | 58.6195 | 85.5825 | 49.4312
2012-10-30 20:13 | 29.1791 | 97.6687 | 73.1909 | 100.2254 | 64.1085 | -1.4795 | 0.006751 | 0.32981 | 0.002227 | 48.8532 | 0.006965 | 0.21519 | 0.001499 | 30.8961 | 14.8354 | -0.102 | 0.003726 | 82.8333 | 58.3555 | 85.39 | 49.2731
2012-10-30 20:13 | 29.1783 | 97.4712 | 72.9947 | 100.0809 | 63.9417 | -1.4934 | 0.006754 | 0.32969 | 0.002227 | 48.8145 | 0.006958 | 0.21518 | 0.001497 | 30.9251 | 14.8306 | -0.103 | 0.003724 | 82.6406 | 58.1641 | 85.2503 | 49.1111
2012-10-30 20:13 | 29.1806 | 97.2844 | 72.7326 | 99.8865 | 63.6389 | -1.4754 | 0.006751 | 0.32966 | 0.002226 | 48.8319 | 0.006971 | 0.21513 | 0.0015 | 30.8611 | 14.8492 | -0.1017 | 0.003726 | 82.4352 | 57.8834 | 85.0373 | 48.7897
2012-10-30 20:13 | 29.1813 | 96.9673 | 72.5072 | 99.6112 | 63.3112 | -1.4749 | 0.006754 | 0.32966 | 0.002227 | 48.8102 | 0.006955 | 0.21519 | 0.001497 | 30.9405 | 14.8417 | -0.1017 | 0.003724 | 82.1256 | 57.6655 | 84.7695 | 48.4695
2012-10-30 20:13 | 29.1821 | 96.8357 | 72.2857 | 99.4078 | 63.1845 | -1.483 | 0.006741 | 0.32966 | 0.002222 | 48.9043 | 0.006971 | 0.21515 | 0.0015 | 30.8633 | 14.8058 | -0.1022 | 0.003722 | 82.0299 | 57.4799 | 84.602 | 48.3787
2012-10-30 20:13 | 29.1783 | 96.6087 | 72.1286 | 99.2716 | 63.1305 | -1.4818 | 0.006754 | 0.32969 | 0.002227 | 48.8145 | 0.006961 | 0.2151 | 0.001497 | 30.9013 | 14.8265 | -0.1022 | 0.003724 | 81.7822 | 57.3021 | 84.4451 | 48.304
2012-10-30 20:13 | 29.1753 | 96.4089 | 71.8802 | 99.1173 | 62.9303 | -1.483 | 0.006757 | 0.32964 | 0.002227 | 48.7842 | 0.006961 | 0.21516 | 0.001498 | 30.9096 | 14.8316 | -0.1022 | 0.003725 | 81.5773 | 57.0486 | 84.2857 | 48.0987
2012-10-30 20:13 | 29.1828 | 96.2386 | 71.696 | 98.8979 | 62.6904 | -1.4916 | 0.00676 | 0.32958 | 0.002228 | 48.754 | 0.006971 | 0.21516 | 0.0015 | 30.8653 | 14.8106 | -0.1028 | 0.003728 | 81.428 | 56.8854 | 84.0873 | 47.8798
2012-10-30 20:13 | 29.1791 | 96.0268 | 71.5312 | 98.6938 | 62.5514 | -1.4876 | 0.006751 | 0.32958 | 0.002225 | 48.819 | 0.006965 | 0.21512 | 0.001498 | 30.8857 | 14.7702 | -0.1026 | 0.003723 | 81.2566 | 56.761 | 83.9236 | 47.7812
2012-10-30 20:13 | 29.1753 | 95.8867 | 71.3164 | 98.5803 | 62.3852 | -1.4824 | 0.006754 | 0.32964 | 0.002226 | 48.8059 | 0.006971 | 0.21509 | 0.001499 | 30.855 | 14.7786 | -0.1022 | 0.003725 | 81.1081 | 56.5378 | 83.8017 | 47.6066
2012-10-30 20:13 | 29.1656 | 95.6462 | 71.1111 | 98.3512 | 62.1473 | -1.465 | 0.006769 | 0.32953 | 0.002231 | 48.6828 | 0.006971 | 0.21507 | 0.001499 | 30.8528 | 14.7852 | -0.101 | 0.00373 | 80.861 | 56.3259 | 83.566 | 47.3621
2012-10-30 20:13 | 29.1686 | 95.4297 | 70.9491 | 98.1878 | 61.8679 | -1.4743 | 0.006757 | 0.32961 | 0.002227 | 48.7799 | 0.006965 | 0.21513 | 0.001498 | 30.8877 | 14.7841 | -0.1016 | 0.003725 | 80.6456 | 56.165 | 83.4037 | 47.0838
2012-10-30 20:13 | 29.1716 | 95.251 | 70.7295 | 97.9714 | 61.5614 | -1.4581 | 0.006747 | 0.32966 | 0.002224 | 48.8608 | 0.006974 | 0.21518 | 0.001501 | 30.8542 | 14.8116 | -0.1005 | 0.003725 | 80.4394 | 55.9179 | 83.1598 | 46.7498
2012-10-30 20:13 | 29.1656 | 95.0596 | 70.5546 | 97.8066 | 61.4548 | -1.4691 | 0.006757 | 0.32958 | 0.002227 | 48.7756 | 0.006961 | 0.21512 | 0.001497 | 30.9035 | 14.7485 | -0.1013 | 0.003724 | 80.3111 | 55.8061 | 83.0581 | 46.7063
2012-10-30 20:13 | 29.1634 | 94.8795 | 70.3681 | 97.653 | 61.29 | -1.4494 | 0.00676 | 0.32959 | 0.002228 | 48.7562 | 0.006971 | 0.21516 | 0.0015 | 30.8653 | 14.7259 | -0.0999 | 0.003728 | 80.1536 | 55.6422 | 82.9271 | 46.5641
2012-10-30 20:13 | 29.1581 | 94.6168 | 70.1555 | 97.4246 | 61.1531 | -1.465 | 0.006747 | 0.32961 | 0.002224 | 48.8522 | 0.006968 | 0.21506 | 0.001499 | 30.8641 | 14.7083 | -0.101 | 0.003723 | 79.9085 | 55.4472 | 82.7163 | 46.4448
2012-10-30 20:13 | 29.1641 | 94.4373 | 69.908 | 97.2187 | 60.9964 | -1.4633 | 0.00676 | 0.32965 | 0.002228 | 48.7646 | 0.006974 | 0.21519 | 0.001501 | 30.8562 | 14.6898 | -0.1009 | 0.003729 | 79.7475 | 55.2182 | 82.5289 | 46.3066
2012-10-30 20:13 | 29.1574 | 94.177 | 69.6978 | 96.9812 | 60.8271 | -1.4535 | 0.006766 | 0.32968 | 0.002231 | 48.7257 | 0.006961 | 0.2151 | 0.001497 | 30.9013 | 14.6791 | -0.1002 | 0.003728 | 79.4979 | 55.0187 | 82.3021 | 46.148
2012-10-30 20:13 | 29.1619 | 93.9808 | 69.5563 | 96.8344 | 60.6651 | -1.4593 | 0.006763 | 0.32965 | 0.002229 | 48.743 | 0.006968 | 0.21516 | 0.001499 | 30.8786 | 14.6672 | -0.1006 | 0.003728 | 79.3136 | 54.8891 | 82.1672 | 45.9979
2012-10-30 20:13 | 29.1679 | 93.771 | 69.4048 | 96.6701 | 60.5045 | -1.4639 | 0.006776 | 0.32955 | 0.002233 | 48.6346 | 0.00698 | 0.21519 | 0.001502 | 30.8297 | 14.6936 | -0.1009 | 0.003735 | 79.0774 | 54.7112 | 81.9765 | 45.8109
2012-10-30 20:13 | 29.1664 | 93.6029 | 69.1804 | 96.487 | 60.4023 | -1.4546 | 0.006757 | 0.32958 | 0.002227 | 48.7756 | 0.00698 | 0.21515 | 0.001502 | 30.8235 | 14.6595 | -0.1003 | 0.003729 | 78.9434 | 54.5209 | 81.8275 | 45.7428
2012-10-30 20:13 | 29.1686 | 93.4006 | 69.0173 | 96.2963 | 60.2495 | -1.4674 | 0.006766 | 0.32959 | 0.00223 | 48.713 | 0.006961 | 0.21516 | 0.001498 | 30.9096 | 14.6699 | -0.1012 | 0.003728 | 78.7307 | 54.3474 | 81.6264 | 45.5796
2012-10-30 20:13 | 29.1581 | 93.1743 | 68.8184 | 96.1344 | 60.1305 | -1.4575 | 0.00676 | 0.32956 | 0.002228 | 48.7519 | 0.006971 | 0.21522 | 0.0015 | 30.8736 | 14.643 | -0.1005 | 0.003728 | 78.5313 | 54.1754 | 81.4914 | 45.4875
2012-10-30 20:13 | 29.1686 | 93.0295 | 68.6944 | 95.9103 | 59.9394 | -1.4679 | 0.006757 | 0.32958 | 0.002227 | 48.7756 | 0.006987 | 0.21518 | 0.001503 | 30.7968 | 14.6658 | -0.1012 | 0.00373 | 78.3637 | 54.0286 | 81.2445 | 45.2736
2012-10-30 20:14 | 29.1671 | 92.7932 | 68.5506 | 95.7385 | 59.7751 | -1.4645 | 0.00676 | 0.32958 | 0.002228 | 48.754 | 0.006977 | 0.21515 | 0.001501 | 30.8367 | 14.6357 | -0.101 | 0.003729 | 78.1575 | 53.9149 | 81.1028 | 45.1394
2012-10-30 20:14 | 29.1619 | 92.6252 | 68.3685 | 95.5857 | 59.5567 | -1.4645 | 0.006754 | 0.32956 | 0.002226 | 48.7952 | 0.00698 | 0.21518 | 0.001502 | 30.8276 | 14.6183 | -0.101 | 0.003728 | 78.0069 | 53.7502 | 80.9674 | 44.9384
2012-10-30 20:14 | 29.1649 | 92.4426 | 68.1407 | 95.4071 | 59.3852 | -1.4622 | 0.006763 | 0.32948 | 0.002228 | 48.7174 | 0.00698 | 0.21515 | 0.001502 | 30.8235 | 14.6335 | -0.1008 | 0.00373 | 77.8091 | 53.5072 | 80.7736 | 44.7517
2012-10-30 20:14 | 29.1611 | 92.2304 | 67.9869 | 95.214 | 59.172 | -1.4622 | 0.006747 | 0.32946 | 0.002223 | 48.8309 | 0.006965 | 0.21518 | 0.001499 | 30.894 | 14.6215 | -0.1008 | 0.003722 | 77.6089 | 53.3654 | 80.5925 | 44.5505
2012-10-30 20:14 | 29.1596 | 92.0395 | 67.8315 | 95.0006 | 59.0192 | -1.4697 | 0.006754 | 0.32945 | 0.002225 | 48.7781 | 0.006965 | 0.2151 | 0.001498 | 30.8836 | 14.5751 | -0.1013 | 0.003723 | 77.4644 | 53.2564 | 80.4255 | 44.4441
2012-10-30 20:14 | 29.1664 | 91.8607 | 67.6613 | 94.856 | 58.8941 | -1.4575 | 0.006747 | 0.32946 | 0.002223 | 48.8309 | 0.006968 | 0.21509 | 0.001499 | 30.8683 | 14.5534 | -0.1005 | 0.003722 | 77.3073 | 53.1079 | 80.3026 | 44.3407
2012-10-30 20:14 | 29.1709 | 91.653 | 67.4965 | 94.6826 | 58.8093 | -1.4558 | 0.006754 | 0.32946 | 0.002225 | 48.7803 | 0.006958 | 0.21509 | 0.001497 | 30.9126 | 14.6028 | -0.1004 | 0.003722 | 77.0502 | 52.8937 | 80.0798 | 44.2065
2012-10-30 20:14 | 29.1701 | 91.4817 | 67.3301 | 94.4812 | 58.5718 | -1.4656 | 0.006751 | 0.32946 | 0.002224 | 48.802 | 0.006965 | 0.21518 | 0.001499 | 30.894 | 14.5775 | -0.101 | 0.003723 | 76.9042 | 52.7526 | 79.9037 | 43.9943
2012-10-30 20:14 | 29.1731 | 91.2877 | 67.1445 | 94.2381 | 58.4273 | -1.4627 | 0.006751 | 0.32958 | 0.002225 | 48.819 | 0.00698 | 0.21512 | 0.001502 | 30.8193 | 14.5437 | -0.1008 | 0.003727 | 76.744 | 52.6008 | 79.6944 | 43.8836
2012-10-30 20:14 | 29.1596 | 91.0165 | 66.9231 | 94.0504 | 58.2314 | -1.4674 | 0.006782 | 0.32952 | 0.002235 | 48.5873 | 0.006965 | 0.2151 | 0.001498 | 30.8836 | 14.5506 | -0.1012 | 0.003733 | 76.4659 | 52.3725 | 79.4998 | 43.6808
2012-10-30 20:14 | 29.1619 | 90.8217 | 66.7675 | 93.8748 | 58.1327 | -1.4691 | 0.006747 | 0.32955 | 0.002223 | 48.8436 | 0.006977 | 0.21512 | 0.001501 | 30.8326 | 14.557 | -0.1013 | 0.003724 | 76.2647 | 52.2105 | 79.3178 | 43.5757
2012-10-30 20:14 | 29.1664 | 90.6178 | 66.6102 | 93.6824 | 57.9625 | -1.4662 | 0.006757 | 0.3295 | 0.002226 | 48.765 | 0.006965 | 0.21523 | 0.001499 | 30.9022 | 14.5534 | -0.1011 | 0.003725 | 76.0644 | 52.0568 | 79.129 | 43.4091
2012-10-30 20:14 | 29.1664 | 90.4207 | 66.4407 | 93.5045 | 57.7489 | -1.4639 | 0.006751 | 0.32948 | 0.002224 | 48.804 | 0.006968 | 0.21515 | 0.001499 | 30.8766 | 14.5493 | -0.1009 | 0.003723 | 75.8714 | 51.8914 | 78.9552 | 43.1996
2012-10-30 20:14 | 29.1664 | 90.1781 | 66.2904 | 93.3227 | 57.5857 | -1.4714 | 0.006757 | 0.32946 | 0.002226 | 48.7586 | 0.006974 | 0.2151 | 0.0015 | 30.8437 | 14.5166 | -0.1014 | 0.003726 | 75.6615 | 51.7738 | 78.8061 | 43.0691
2012-10-30 20:14 | 29.1694 | 89.984 | 66.1008 | 93.125 | 57.4528 | -1.4587 | 0.006744 | 0.32952 | 0.002222 | 48.8611 | 0.006965 | 0.2151 | 0.001498 | 30.8836 | 14.5359 | -0.1006 | 0.00372 | 75.4481 | 51.5649 | 78.5891 | 42.9169
2012-10-30 20:14 | 29.1686 | 89.8012 | 65.9651 | 92.9651 | 57.3626 | -1.4783 | 0.006757 | 0.32948 | 0.002226 | 48.7607 | 0.006977 | 0.21513 | 0.001501 | 30.8346 | 14.576 | -0.1019 | 0.003727 | 75.2252 | 51.3891 | 78.3891 | 42.7866
2012-10-30 20:14 | 29.1694 | 89.6389 | 65.7847 | 92.716 | 57.1769 | -1.4685 | 0.006763 | 0.32949 | 0.002228 | 48.7195 | 0.006971 | 0.21516 | 0.0015 | 30.8653 | 14.5523 | -0.1012 | 0.003728 | 75.0866 | 51.2324 | 78.1637 | 42.6246
2012-10-30 20:14 | 29.1686 | 89.3955 | 65.6605 | 92.575 | 56.9001 | -1.4749 | 0.00676 | 0.32948 | 0.002227 | 48.7391 | 0.006968 | 0.21519 | 0.001499 | 30.8828 | 14.5311 | -0.1017 | 0.003726 | 74.8644 | 51.1294 | 78.0439 | 42.369
2012-10-30 20:14 | 29.1664 | 89.2037 | 65.4733 | 92.3683 | 56.8007 | -1.4697 | 0.006751 | 0.32937 | 0.002224 | 48.7891 | 0.006968 | 0.21519 | 0.001499 | 30.8828 | 14.5329 | -0.1013 | 0.003723 | 74.6708 | 50.9404 | 77.8354 | 42.2678
2012-10-30 20:14 | 29.1649 | 89.058 | 65.3137 | 92.2039 | 56.6943 | -1.4697 | 0.006757 | 0.3294 | 0.002226 | 48.75 | 0.006965 | 0.21516 | 0.001499 | 30.8919 | 14.5437 | -0.1013 | 0.003725 | 74.5143 | 50.77 | 77.6602 | 42.1506
2012-10-30 20:14 | 29.1701 | 88.8584 | 65.2032 | 92.0652 | 56.5828 | -1.4731 | 0.006744 | 0.32936 | 0.002221 | 48.8375 | 0.006968 | 0.21516 | 0.001499 | 30.8786 | 14.5244 | -0.1016 | 0.00372 | 74.334 | 50.6788 | 77.5408 | 42.0584
2012-10-30 20:14 | 29.1671 | 88.6468 | 64.9957 | 91.8614 | 56.4126 | -1.4645 | 0.006747 | 0.32943 | 0.002223 | 48.8266 | 0.006965 | 0.21515 | 0.001499 | 30.8899 | 14.5091 | -0.101 | 0.003722 | 74.1377 | 50.4866 | 77.3523 | 41.9035
2012-10-30 20:14 | 29.1604 | 88.4353 | 64.8231 | 91.669 | 56.2192 | -1.4639 | 0.006754 | 0.32939 | 0.002225 | 48.7695 | 0.006968 | 0.21516 | 0.001499 | 30.8786 | 14.5392 | -0.1009 | 0.003724 | 73.8961 | 50.2839 | 77.1298 | 41.68
2012-10-30 20:14 | 29.1724 | 88.269 | 64.6692 | 91.4045 | 56.0951 | -1.465 | 0.00676 | 0.32937 | 0.002227 | 48.7241 | 0.006968 | 0.2151 | 0.001499 | 30.8703 | 14.5471 | -0.101 | 0.003726 | 73.7219 | 50.1221 | 76.8574 | 41.548
2012-10-30 20:14 | 29.1701 | 88.0884 | 64.5087 | 91.2544 | 55.9566 | -1.476 | 0.00676 | 0.32942 | 0.002227 | 48.7305 | 0.006983 | 0.2151 | 0.001502 | 30.804 | 14.5489 | -0.1018 | 0.003729 | 73.5395 | 49.9598 | 76.7055 | 41.4077
2012-10-30 20:15 | 29.1694 | 87.9017 | 64.3999 | 91.14 | 55.8042 | -1.472 | 0.00676 | 0.3294 | 0.002227 | 48.7284 | 0.00698 | 0.21521 | 0.001502 | 30.8317 | 14.54 | -0.1015 | 0.003729 | 73.3617 | 49.8599 | 76.6 | 41.2642
2012-10-30 20:15 | 29.1589 | 87.6711 | 64.2196 | 90.8907 | 55.6808 | -1.4754 | 0.006757 | 0.32939 | 0.002226 | 48.7478 | 0.006974 | 0.21516 | 0.001501 | 30.852 | 14.5336 | -0.1017 | 0.003727 | 73.1375 | 49.686 | 76.3571 | 41.1472
2012-10-30 20:15 | 29.1656 | 87.4312 | 64.0448 | 90.7117 | 55.5318 | -1.4656 | 0.006757 | 0.32933 | 0.002225 | 48.7392 | 0.006958 | 0.21519 | 0.001497 | 30.9271 | 14.5893 | -0.101 | 0.003722 | 72.8419 | 49.4555 | 76.1224 | 40.9425
2012-10-30 20:15 | 29.1626 | 87.265 | 63.8951 | 90.5192 | 55.4237 | -1.4714 | 0.00676 | 0.3293 | 0.002226 | 48.7135 | 0.006955 | 0.21515 | 0.001496 | 30.9343 | 14.5332 | -0.1014 | 0.003722 | 72.7318 | 49.3619 | 75.986 | 40.8905
2012-10-30 20:15 | 29.1596 | 87.1101 | 63.7801 | 90.376 | 55.2105 | -1.4743 | 0.006757 | 0.3293 | 0.002225 | 48.7351 | 0.006968 | 0.21507 | 0.001499 | 30.8661 | 14.5466 | -0.1016 | 0.003724 | 72.5635 | 49.2335 | 75.8294 | 40.6639
2012-10-30 20:15 | 29.1596 | 86.8976 | 63.6913 | 90.2206 | 55.1404 | -1.4737 | 0.006757 | 0.32927 | 0.002225 | 48.7308 | 0.006961 | 0.21507 | 0.001497 | 30.8971 | 14.5506 | -0.1016 | 0.003722 | 72.347 | 49.1407 | 75.67 | 40.5898
2012-10-30 20:15 | 29.1656 | 86.7289 | 63.5118 | 90.0105 | 55.0374 | -1.4708 | 0.00676 | 0.32927 | 0.002226 | 48.7092 | 0.006955 | 0.2151 | 0.001496 | 30.928 | 14.5485 | -0.1014 | 0.003722 | 72.1804 | 48.9633 | 75.462 | 40.4889
2012-10-30 20:15 | 29.1664 | 86.6006 | 63.362 | 89.8748 | 54.8396 | -1.4668 | 0.00676 | 0.32923 | 0.002226 | 48.7027 | 0.006977 | 0.21513 | 0.001501 | 30.8346 | 14.6023 | -0.1011 | 0.003727 | 71.9983 | 48.7597 | 75.2725 | 40.2373
2012-10-30 20:15 | 29.1491 | 86.3745 | 63.1708 | 89.7059 | 54.6626 | -1.4691 | 0.006751 | 0.32923 | 0.002223 | 48.7676 | 0.00698 | 0.2151 | 0.001501 | 30.8172 | 14.5687 | -0.1013 | 0.003724 | 71.8058 | 48.6021 | 75.1372 | 40.0939
2012-10-30 20:15 | 29.1461 | 86.2348 | 63.0596 | 89.4791 | 54.5427 | -1.4546 | 0.006773 | 0.32926 | 0.00223 | 48.6135 | 0.006968 | 0.21509 | 0.001499 | 30.8683 | 14.529 | -0.1003 | 0.003729 | 71.7058 | 48.5306 | 74.9501 | 40.0137
2012-10-30 20:15 | 29.1581 | 86.0152 | 62.9479 | 89.3546 | 54.356 | -1.4697 | 0.00676 | 0.32923 | 0.002226 | 48.7027 | 0.006971 | 0.21499 | 0.001499 | 30.8405 | 14.5451 | -0.1013 | 0.003725 | 71.4701 | 48.4028 | 74.8095 | 39.8109
2012-10-30 20:15 | 29.1469 | 85.8103 | 62.7705 | 89.15 | 54.0487 | -1.4685 | 0.006757 | 0.32923 | 0.002225 | 48.7243 | 0.006961 | 0.21502 | 0.001497 | 30.889 | 14.5747 | -0.1012 | 0.003722 | 71.2356 | 48.1958 | 74.5753 | 39.474
2012-10-30 20:15 | 29.1424 | 85.631 | 62.6075 | 88.99 | 53.7479 | -1.4702 | 0.006754 | 0.32922 | 0.002224 | 48.7439 | 0.006955 | 0.21506 | 0.001496 | 30.9218 | 14.5579 | -0.1014 | 0.00372 | 71.0731 | 48.0496 | 74.4321 | 39.19
2012-10-30 20:15 | 29.1371 | 85.451 | 62.463 | 88.761 | 53.8089 | -1.4697 | 0.006757 | 0.32925 | 0.002225 | 48.7265 | 0.006949 | 0.21505 | 0.001494 | 30.9465 | 14.5363 | -0.1013 | 0.003719 | 70.9147 | 47.9267 | 74.2247 | 39.2726
2012-10-30 20:15 | 29.1371 | 85.2838 | 62.3664 | 88.6661 | 53.7465 | -1.4789 | 0.006744 | 0.32923 | 0.00222 | 48.8182 | 0.006965 | 0.21513 | 0.001498 | 30.8877 | 14.5526 | -0.102 | 0.003718 | 70.7312 | 47.8138 | 74.1135 | 39.1939
2012-10-30 20:15 | 29.1401 | 85.0969 | 62.1915 | 88.4946 | 53.6287 | -1.4535 | 0.006738 | 0.32925 | 0.002218 | 48.8639 | 0.006965 | 0.2151 | 0.001498 | 30.8836 | 14.5679 | -0.1002 | 0.003716 | 70.529 | 47.6236 | 73.9267 | 39.0608
2012-10-30 20:15 | 29.1401 | 84.9297 | 61.9903 | 88.286 | 53.4571 | -1.4697 | 0.00676 | 0.32922 | 0.002226 | 48.7006 | 0.006974 | 0.21513 | 0.0015 | 30.8479 | 14.5597 | -0.1013 | 0.003726 | 70.37 | 47.4306 | 73.7263 | 38.8974
2012-10-30 20:15 | 29.1304 | 84.6882 | 61.8104 | 88.079 | 53.2836 | -1.465 | 0.006751 | 0.3291 | 0.002222 | 48.7483 | 0.006961 | 0.21513 | 0.001498 | 30.9055 | 14.5092 | -0.101 | 0.00372 | 70.179 | 47.3012 | 73.5698 | 38.7744
2012-10-30 20:15 | 29.1424 | 84.5329 | 61.6986 | 87.8937 | 53.1981 | -1.4691 | 0.006741 | 0.32922 | 0.002219 | 48.8379 | 0.006968 | 0.21507 | 0.001499 | 30.8661 | 14.5375 | -0.1013 | 0.003718 | 69.9954 | 47.1611 | 73.3562 | 38.6606
2012-10-30 20:15 | 29.1304 | 84.3081 | 61.524 | 87.7147 | 53.0223 | -1.4783 | 0.006763 | 0.32916 | 0.002226 | 48.6704 | 0.006955 | 0.21506 | 0.001496 | 30.9218 | 14.5541 | -0.1019 | 0.003722 | 69.754 | 46.9699 | 73.1606 | 38.4682
2012-10-30 20:15 | 29.1267 | 84.1333 | 61.3888 | 87.5631 | 52.9796 | -1.4633 | 0.00676 | 0.32914 | 0.002225 | 48.6899 | 0.006961 | 0.2151 | 0.001497 | 30.9013 | 14.5137 | -0.1009 | 0.003722 | 69.6196 | 46.8751 | 73.0494 | 38.4659
2012-10-30 20:15 | 29.1326 | 83.9833 | 61.2592 | 87.3716 | 52.7436 | -1.4766 | 0.00676 | 0.32916 | 0.002225 | 48.692 | 0.006961 | 0.21506 | 0.001497 | 30.8951 | 14.5645 | -0.1018 | 0.003722 | 69.4188 | 46.6947 | 72.8071 | 38.1791
2012-10-30 20:15 | 29.1379 | 83.7947 | 61.1251 | 87.1947 | 52.4601 | -1.4818 | 0.006754 | 0.32916 | 0.002223 | 48.7353 | 0.006968 | 0.21503 | 0.001498 | 30.8599 | 14.533 | -0.1022 | 0.003721 | 69.2617 | 46.5921 | 72.6617 | 37.9271
2012-10-30 20:15 | 29.1386 | 83.5483 | 60.9593 | 86.9829 | 52.3749 | -1.4743 | 0.006751 | 0.32916 | 0.002222 | 48.7569 | 0.006965 | 0.2151 | 0.001498 | 30.8836 | 14.5664 | -0.1016 | 0.00372 | 68.9819 | 46.3929 | 72.4165 | 37.8085
2012-10-30 20:15 | 29.1244 | 83.3515 | 60.798 | 86.8358 | 52.3373 | -1.4708 | 0.006747 | 0.32913 | 0.002221 | 48.7815 | 0.006974 | 0.21502 | 0.0015 | 30.8314 | 14.5318 | -0.1014 | 0.003721 | 68.8197 | 46.2662 | 72.304 | 37.8055
2012-10-30 20:15 | 29.1304 | 83.175 | 60.6026 | 86.6216 | 52.191 | -1.4749 | 0.006754 | 0.32923 | 0.002224 | 48.7459 | 0.006961 | 0.21512 | 0.001497 | 30.9035 | 14.5092 | -0.1017 | 0.003721 | 68.6658 | 46.0934 | 72.1124 | 37.6818
2012-10-30 20:16 | 29.1304 | 82.981 | 60.4321 | 86.3862 | 52.07 | -1.472 | 0.006747 | 0.32909 | 0.00222 | 48.7752 | 0.006943 | 0.21506 | 0.001493 | 30.9752 | 14.5337 | -0.1015 | 0.003713 | 68.4473 | 45.8984 | 71.8525 | 37.5363
2012-10-30 20:16 | 29.1267 | 82.7414 | 60.3006 | 86.261 | 51.9609 | -1.4876 | 0.00676 | 0.32925 | 0.002226 | 48.7049 | 0.006965 | 0.21505 | 0.001498 | 30.8754 | 14.5626 | -0.1026 | 0.003724 | 68.1788 | 45.738 | 71.6984 | 37.3983
2012-10-30 20:16 | 29.1192 | 82.5703 | 60.1613 | 86.0446 | 51.8909 | -1.4789 | 0.006769 | 0.32911 | 0.002228 | 48.6209 | 0.006958 | 0.21506 | 0.001496 | 30.9085 | 14.5347 | -0.102 | 0.003724 | 68.0356 | 45.6266 | 71.5099 | 37.3562
2012-10-30 20:16 | 29.1207 | 82.412 | 60.0776 | 85.8752 | 51.7284 | -1.4864 | 0.006751 | 0.32911 | 0.002222 | 48.7506 | 0.006971 | 0.2151 | 0.001499 | 30.857 | 14.5362 | -0.1025 | 0.003721 | 67.8758 | 45.5414 | 71.339 | 37.1922
2012-10-30 20:16 | 29.1222 | 82.2156 | 59.8814 | 85.6602 | 51.5736 | -1.4754 | 0.006744 | 0.3291 | 0.002219 | 48.7989 | 0.006955 | 0.21505 | 0.001496 | 30.9198 | 14.55 | -0.1017 | 0.003715 | 67.6656 | 45.3314 | 71.1102 | 37.0236
2012-10-30 20:16 | 29.1124 | 81.9926 | 59.7476 | 85.5023 | 51.4896 | -1.4749 | 0.006751 | 0.3291 | 0.002222 | 48.7483 | 0.006971 | 0.21502 | 0.001499 | 30.8446 | 14.5769 | -0.1017 | 0.003721 | 67.4157 | 45.1707 | 70.9254 | 36.9127
2012-10-30 20:16 | 29.1177 | 81.8267 | 59.6211 | 85.3556 | 51.2957 | -1.4778 | 0.006735 | 0.3291 | 0.002216 | 48.8641 | 0.006971 | 0.215 | 0.001499 | 30.8425 | 14.5536 | -0.1019 | 0.003715 | 67.2731 | 45.0675 | 70.802 | 36.7421
2012-10-30 20:16 | 29.1132 | 81.6813 | 59.4731 | 85.2029 | 51.2248 | -1.4737 | 0.006744 | 0.32909 | 0.002219 | 48.7969 | 0.006958 | 0.21503 | 0.001496 | 30.9043 | 14.5206 | -0.1016 | 0.003715 | 67.1607 | 44.9525 | 70.6823 | 36.7042
2012-10-30 20:16 | 29.1177 | 81.5107 | 59.3651 | 85.0365 | 51.1394 | -1.4708 | 0.006763 | 0.32903 | 0.002225 | 48.6512 | 0.006971 | 0.21505 | 0.001499 | 30.8488 | 14.5577 | -0.1014 | 0.003724 | 66.953 | 44.8074 | 70.4788 | 36.5817
2012-10-30 20:16 | 29.1139 | 81.347 | 59.245 | 84.8427 | 50.991 | -1.4818 | 0.006763 | 0.32899 | 0.002225 | 48.6448 | 0.006961 | 0.21503 | 0.001497 | 30.891 | 14.558 | -0.1022 | 0.003722 | 66.789 | 44.687 | 70.2847 | 36.433
2012-10-30 20:16 | 29.1064 | 81.13 | 59.063 | 84.6756 | 50.8545 | -1.4743 | 0.006751 | 0.329 | 0.002221 | 48.7334 | 0.006974 | 0.21509 | 0.0015 | 30.8417 | 14.5628 | -0.1016 | 0.003721 | 66.5672 | 44.5002 | 70.1128 | 36.2917
2012-10-30 20:16 | 29.1109 | 80.986 | 58.9627 | 84.4976 | 50.7574 | -1.472 | 0.006754 | 0.32904 | 0.002222 | 48.7182 | 0.006961 | 0.21512 | 0.001497 | 30.9035 | 14.5264 | -0.1015 | 0.003719 | 66.4596 | 44.4363 | 69.9712 | 36.231
2012-10-30 20:16 | 29.1124 | 80.8123 | 58.8478 | 84.3243 | 50.7433 | -1.4824 | 0.006757 | 0.32904 | 0.002223 | 48.6966 | 0.006965 | 0.215 | 0.001497 | 30.8691 | 14.5402 | -0.1022 | 0.00372 | 66.2721 | 44.3076 | 69.7841 | 36.2031
2012-10-30 20:16 | 29.0997 | 80.6053 | 58.7071 | 84.1367 | 50.5703 | -1.4893 | 0.006763 | 0.32906 | 0.002225 | 48.6555 | 0.006958 | 0.21506 | 0.001496 | 30.9085 | 14.5112 | -0.1027 | 0.003721 | 66.0941 | 44.1959 | 69.6255 | 36.0591
2012-10-30 20:16 | 29.1004 | 80.4117 | 58.5564 | 84.0158 | 50.4459 | -1.4772 | 0.006757 | 0.32904 | 0.002223 | 48.6966 | 0.006968 | 0.215 | 0.001498 | 30.8558 | 14.5282 | -0.1018 | 0.003721 | 65.8835 | 44.0282 | 69.4876 | 35.9177
2012-10-30 20:16 | 29.1087 | 80.2409 | 58.425 | 83.8112 | 50.333 | -1.4766 | 0.006757 | 0.32906 | 0.002223 | 48.6987 | 0.006968 | 0.21499 | 0.001498 | 30.8538 | 14.5202 | -0.1018 | 0.003721 | 65.7207 | 43.9048 | 69.291 | 35.8128
2012-10-30 20:16 | 29.0937 | 80.104 | 58.3246 | 83.6935 | 50.2319 | -1.4714 | 0.006763 | 0.32899 | 0.002225 | 48.6448 | 0.006955 | 0.21502 | 0.001495 | 30.9156 | 14.5174 | -0.1014 | 0.00372 | 65.5866 | 43.8072 | 69.1761 | 35.7145
2012-10-30 20:16 | 29.0892 | 79.9432 | 58.2192 | 83.476 | 50.0201 | -1.4841 | 0.006747 | 0.32903 | 0.00222 | 48.7666 | 0.006952 | 0.21505 | 0.001495 | 30.9331 | 14.5251 | -0.1023 | 0.003715 | 65.4181 | 43.6941 | 68.9509 | 35.495
2012-10-30 20:16 | 29.0884 | 79.7671 | 58.0747 | 83.3155 | 49.9607 | -1.4778 | 0.006738 | 0.32906 | 0.002217 | 48.836 | 0.006955 | 0.215 | 0.001495 | 30.9134 | 14.5162 | -0.1019 | 0.003712 | 65.2509 | 43.5585 | 68.7993 | 35.4445
2012-10-30 20:16 | 29.081 | 79.5348 | 57.947 | 83.1598 | 49.8633 | -1.4853 | 0.00676 | 0.32899 | 0.002224 | 48.6664 | 0.006983 | 0.2151 | 0.001502 | 30.804 | 14.5741 | -0.1024 | 0.003726 | 64.9607 | 43.3729 | 68.5857 | 35.2892
2012-10-30 20:16 | 29.0735 | 79.3748 | 57.8269 | 82.9621 | 49.7306 | -1.4853 | 0.006751 | 0.32893 | 0.002221 | 48.7227 | 0.006965 | 0.21503 | 0.001498 | 30.8732 | 14.5299 | -0.1024 | 0.003719 | 64.8449 | 43.297 | 68.4322 | 35.2007
2012-10-30 20:16 | 29.0727 | 79.2749 | 57.7057 | 82.8053 | 49.5851 | -1.483 | 0.006754 | 0.32897 | 0.002222 | 48.7074 | 0.006968 | 0.21505 | 0.001498 | 30.8621 | 14.5372 | -0.1022 | 0.00372 | 64.7377 | 43.1685 | 68.2681 | 35.0479
2012-10-30 20:16 | 29.0765 | 79.0804 | 57.5774 | 82.6608 | 49.4597 | -1.4772 | 0.006763 | 0.32894 | 0.002225 | 48.6383 | 0.006968 | 0.215 | 0.001498 | 30.8558 | 14.5451 | -0.1018 | 0.003723 | 64.5353 | 43.0323 | 68.1157 | 34.9146
2012-10-30 20:16 | 29.0735 | 78.8906 | 57.4889 | 82.4865 | 49.398 | -1.476 | 0.006751 | 0.32899 | 0.002221 | 48.7313 | 0.006968 | 0.21507 | 0.001499 | 30.8661 | 14.5299 | -0.1018 | 0.00372 | 64.3607 | 42.959 | 67.9566 | 34.8681
2012-10-30 20:16 | 29.0712 | 78.7472 | 57.3429 | 82.3358 | 49.2822 | -1.4806 | 0.006773 | 0.32904 | 0.002229 | 48.5816 | 0.006983 | 0.2151 | 0.001502 | 30.804 | 14.552 | -0.1021 | 0.003731 | 64.1952 | 42.7909 | 67.7838 | 34.7302
2012-10-30 20:17 | 29.0667 | 78.5786 | 57.2257 | 82.1525 | 49.219 | -1.483 | 0.006744 | 0.329 | 0.002219 | 48.784 | 0.006987 | 0.21497 | 0.001502 | 30.7677 | 14.5231 | -0.1022 | 0.003721 | 64.0555 | 42.7026 | 67.6294 | 34.6959
2012-10-30 20:17 | 29.0787 | 78.438 | 57.1173 | 82.0046 | 49.1058 | -1.4847 | 0.006754 | 0.329 | 0.002222 | 48.7117 | 0.006961 | 0.21502 | 0.001497 | 30.889 | 14.5595 | -0.1024 | 0.003719 | 63.8785 | 42.5578 | 67.4451 | 34.5463
2012-10-30 20:17 | 29.0675 | 78.2361 | 57.0011 | 81.8564 | 49.0202 | -1.4841 | 0.00676 | 0.32896 | 0.002224 | 48.6621 | 0.006968 | 0.21503 | 0.001498 | 30.8599 | 14.532 | -0.1023 | 0.003722 | 63.7041 | 42.4691 | 67.3244 | 34.4882
2012-10-30 20:17 | 29.0802 | 78.0923 | 56.8856 | 81.6864 | 48.9272 | -1.4905 | 0.006751 | 0.32897 | 0.002221 | 48.7291 | 0.006971 | 0.21503 | 0.001499 | 30.8467 | 14.5366 | -0.1028 | 0.00372 | 63.5557 | 42.349 | 67.1498 | 34.3906
2012-10-30 20:17 | 29.063 | 77.8118 | 56.7673 | 81.5397 | 48.8121 | -1.4899 | 0.00676 | 0.3289 | 0.002223 | 48.6536 | 0.006958 | 0.215 | 0.001496 | 30.9001 | 14.5398 | -0.1027 | 0.003719 | 63.272 | 42.2275 | 66.9999 | 34.2723
2012-10-30 20:17 | 29.0697 | 77.7231 | 56.6302 | 81.3827 | 48.6387 | -1.476 | 0.006757 | 0.32901 | 0.002223 | 48.6923 | 0.006965 | 0.21506 | 0.001498 | 30.8774 | 14.5587 | -0.1018 | 0.003721 | 63.1644 | 42.0715 | 66.824 | 34.08
2012-10-30 20:17 | 29.0592 | 77.5217 | 56.5419 | 81.159 | 48.5499 | -1.4806 | 0.006766 | 0.32901 | 0.002226 | 48.6275 | 0.006958 | 0.21506 | 0.001496 | 30.9085 | 14.5441 | -0.1021 | 0.003722 | 62.9776 | 41.9978 | 66.6149 | 34.0058
2012-10-30 20:17 | 29.0547 | 77.3875 | 56.3935 | 81.025 | 48.518 | -1.4899 | 0.00676 | 0.32897 | 0.002224 | 48.6642 | 0.006971 | 0.21503 | 0.001499 | 30.8467 | 14.5437 | -0.1027 | 0.003723 | 62.8438 | 41.8498 | 66.4813 | 33.9743
2012-10-30 20:17 | 29.0607 | 77.2331 | 56.2673 | 80.8749 | 48.4104 | -1.4882 | 0.006754 | 0.32896 | 0.002222 | 48.7054 | 0.006961 | 0.21509 | 0.001497 | 30.8993 | 14.5783 | -0.1026 | 0.003719 | 62.6548 | 41.689 | 66.2966 | 33.8321
2012-10-30 20:17 | 29.0555 | 77.0255 | 56.196 | 80.6602 | 48.2603 | -1.4893 | 0.006751 | 0.32897 | 0.002221 | 48.7291 | 0.006965 | 0.21505 | 0.001498 | 30.8754 | 14.5731 | -0.1027 | 0.003719 | 62.4524 | 41.6229 | 66.0871 | 33.6872
2012-10-30 20:17 | 29.051 | 76.8377 | 56.0748 | 80.5223 | 48.1383 | -1.4841 | 0.006751 | 0.3289 | 0.00222 | 48.7184 | 0.006971 | 0.215 | 0.001499 | 30.8425 | 14.5645 | -0.1023 | 0.003719 | 62.2732 | 41.5103 | 65.9578 | 33.5738
2012-10-30 20:17 | 29.0585 | 76.7574 | 55.9539 | 80.4346 | 48.0282 | -1.4789 | 0.006763 | 0.32891 | 0.002224 | 48.634 | 0.006965 | 0.21505 | 0.001498 | 30.8754 | 14.5638 | -0.102 | 0.003722 | 62.1936 | 41.3901 | 65.8708 | 33.4644
2012-10-30 20:17 | 29.0487 | 76.5833 | 55.8624 | 80.2571 | 47.9597 | -1.4847 | 0.006754 | 0.32888 | 0.002221 | 48.6947 | 0.006955 | 0.21503 | 0.001496 | 30.9176 | 14.5214 | -0.1024 | 0.003717 | 62.0619 | 41.341 | 65.7357 | 33.4383
2012-10-30 20:17 | 29.0398 | 76.3719 | 55.6979 | 80.0844 | 47.8372 | -1.4911 | 0.006738 | 0.32884 | 0.002216 | 48.8038 | 0.006952 | 0.215 | 0.001495 | 30.9268 | 14.5247 | -0.1028 | 0.003711 | 61.8472 | 41.1732 | 65.5597 | 33.3125
2012-10-30 20:17 | 29.0472 | 76.2189 | 55.5886 | 79.8936 | 47.7074 | -1.4812 | 0.006766 | 0.32878 | 0.002225 | 48.5933 | 0.006946 | 0.21496 | 0.001493 | 30.9473 | 14.5444 | -0.1021 | 0.003718 | 61.6745 | 41.0442 | 65.3492 | 33.163
2012-10-30 20:17 | 29.0368 | 76.0786 | 55.4964 | 79.727 | 47.5794 | -1.4841 | 0.006732 | 0.3288 | 0.002213 | 48.8409 | 0.006968 | 0.21502 | 0.001498 | 30.8579 | 14.5095 | -0.1023 | 0.003711 | 61.5691 | 40.9869 | 65.2175 | 33.0699
2012-10-30 20:17 | 29.0405 | 75.9561 | 55.3445 | 79.6239 | 47.5087 | -1.4916 | 0.00676 | 0.32886 | 0.002223 | 48.6472 | 0.006974 | 0.21497 | 0.001499 | 30.8251 | 14.5091 | -0.1028 | 0.003722 | 61.447 | 40.8354 | 65.1148 | 32.9996
2012-10-30 20:17 | 29.039 | 75.7521 | 55.2379 | 79.4356 | 47.417 | -1.4858 | 0.006751 | 0.32884 | 0.00222 | 48.7098 | 0.006958 | 0.215 | 0.001496 | 30.9001 | 14.5158 | -0.1024 | 0.003716 | 61.2363 | 40.7221 | 64.9198 | 32.9012
2012-10-30 20:17 | 29.0323 | 75.6001 | 55.0754 | 79.2498 | 47.2888 | -1.4818 | 0.006751 | 0.32883 | 0.00222 | 48.7077 | 0.006955 | 0.21502 | 0.001495 | 30.9156 | 14.4846 | -0.1022 | 0.003715 | 61.1155 | 40.5908 | 64.7652 | 32.8042
2012-10-30 20:17 | 29.033 | 75.3715 | 54.9944 | 79.0446 | 47.1562 | -1.472 | 0.006744 | 0.32886 | 0.002218 | 48.7626 | 0.006961 | 0.21502 | 0.001497 | 30.889 | 14.5302 | -0.1015 | 0.003715 | 60.8413 | 40.4642 | 64.5144 | 32.626
2012-10-30 20:17 | 29.0368 | 75.2644 | 54.8385 | 78.9073 | 47.0659 | -1.4754 | 0.006766 | 0.32884 | 0.002225 | 48.6018 | 0.006971 | 0.21499 | 0.001499 | 30.8405 | 14.5381 | -0.1017 | 0.003724 | 60.7263 | 40.3004 | 64.3692 | 32.5278
2012-10-30 20:17 | 29.0308 | 75.1245 | 54.7624 | 78.7487 | 47.0168 | -1.4922 | 0.006744 | 0.32883 | 0.002218 | 48.7583 | 0.006971 | 0.215 | 0.001499 | 30.8425 | 14.5647 | -0.1029 | 0.003717 | 60.5598 | 40.1977 | 64.184 | 32.4521
2012-10-30 20:17 | 29.0263 | 74.8906 | 54.6294 | 78.5497 | 46.8476 | -1.476 | 0.006757 | 0.32881 | 0.002222 | 48.6623 | 0.006961 | 0.215 | 0.001497 | 30.8868 | 14.5398 | -0.1018 | 0.003719 | 60.3508 | 40.0896 | 64.0099 | 32.3078
2012-10-30 20:17 | 29.0225 | 74.6917 | 54.5048 | 78.4429 | 46.7497 | -1.4801 | 0.006741 | 0.3288 | 0.002216 | 48.7757 | 0.006968 | 0.21507 | 0.001499 | 30.8661 | 14.5156 | -0.102 | 0.003715 | 60.1761 | 39.9892 | 63.9273 | 32.2341
2012-10-30 20:17 | 29.0255 | 74.5378 | 54.3871 | 78.2666 | 46.6978 | -1.4679 | 0.006757 | 0.32884 | 0.002222 | 48.6666 | 0.006958 | 0.21502 | 0.001496 | 30.9023 | 14.5839 | -0.1012 | 0.003718 | 59.9539 | 39.8032 | 63.6827 | 32.1139
2012-10-30 20:17 | 29.018 | 74.3582 | 54.325 | 78.1522 | 46.5766 | -1.4702 | 0.006763 | 0.32878 | 0.002224 | 48.6148 | 0.006971 | 0.21502 | 0.001499 | 30.8446 | 14.605 | -0.1014 | 0.003723 | 59.7532 | 39.72 | 63.5472 | 31.9716
2012-10-30 20:18 | 29.0143 | 74.2091 | 54.2395 | 77.9578 | 46.463 | -1.4806 | 0.006751 | 0.32888 | 0.00222 | 48.7163 | 0.006952 | 0.21497 | 0.001494 | 30.9226 | 14.5686 | -0.1021 | 0.003714 | 59.6405 | 39.6709 | 63.3892 | 31.8944
2012-10-30 20:18 | 29.0188 | 74.0681 | 54.0998 | 77.7676 | 46.3852 | -1.4702 | 0.006757 | 0.32881 | 0.002222 | 48.6623 | 0.006965 | 0.21502 | 0.001498 | 30.8712 | 14.5976 | -0.1014 | 0.00372 | 59.4705 | 39.5022 | 63.17 | 31.7876
2012-10-30 20:18 | 29.0083 | 73.8433 | 53.9685 | 77.5663 | 46.2923 | -1.4766 | 0.006757 | 0.32886 | 0.002222 | 48.6688 | 0.006958 | 0.215 | 0.001496 | 30.9001 | 14.5585 | -0.1018 | 0.003718 | 59.2848 | 39.41 | 63.0078 | 31.7338
2012-10-30 20:18 | 29.003 | 73.7308 | 53.9048 | 77.4885 | 46.2321 | -1.4766 | 0.006751 | 0.32881 | 0.00222 | 48.7055 | 0.006968 | 0.21506 | 0.001499 | 30.8641 | 14.5328 | -0.1018 | 0.003719 | 59.198 | 39.372 | 62.9557 | 31.6993
2012-10-30 20:18 | 29.003 | 73.616 | 53.7489 | 77.2785 | 46.1418 | -1.4916 | 0.006766 | 0.32881 | 0.002225 | 48.5975 | 0.006968 | 0.21494 | 0.001498 | 30.8474 | 14.4961 | -0.1028 | 0.003723 | 59.1199 | 39.2528 | 62.7824 | 31.6457
2012-10-30 20:18 | 29.0053 | 73.4115 | 53.6771 | 77.1624 | 46.0225 | -1.4754 | 0.006751 | 0.32875 | 0.002219 | 48.6971 | 0.006977 | 0.21496 | 0.0015 | 30.8098 | 14.5678 | -0.1017 | 0.003719 | 58.8437 | 39.1093 | 62.5946 | 31.4547
2012-10-30 20:18 | 29.0068 | 73.2752 | 53.5616 | 77.0417 | 45.9416 | -1.4853 | 0.006754 | 0.32877 | 0.00222 | 48.6775 | 0.006971 | 0.215 | 0.001499 | 30.8425 | 14.5897 | -0.1024 | 0.003719 | 58.6855 | 38.9719 | 62.452 | 31.3519
2012-10-30 20:18 | 29.0038 | 73.1381 | 53.4768 | 76.886 | 45.8248 | -1.4882 | 0.006747 | 0.32887 | 0.002219 | 48.743 | 0.006965 | 0.21496 | 0.001497 | 30.8629 | 14.5336 | -0.1026 | 0.003716 | 58.6045 | 38.9432 | 62.3524 | 31.2912
2012-10-30 20:18 | 29.003 | 73.0148 | 53.3473 | 76.7515 | 45.8083 | -1.4818 | 0.006747 | 0.3288 | 0.002218 | 48.7323 | 0.006961 | 0.21497 | 0.001496 | 30.8826 | 14.5328 | -0.1022 | 0.003714 | 58.482 | 38.8145 | 62.2187 | 31.2755
2012-10-30 20:18 | 28.9963 | 72.8204 | 53.2236 | 76.5768 | 45.6681 | -1.4731 | 0.006751 | 0.32878 | 0.00222 | 48.7012 | 0.006958 | 0.21505 | 0.001496 | 30.9064 | 14.5465 | -0.1016 | 0.003716 | 58.2739 | 38.6771 | 62.0303 | 31.1216
2012-10-30 20:18 | 29.0038 | 72.6938 | 53.1648 | 76.42 | 45.5579 | -1.4783 | 0.006741 | 0.32874 | 0.002216 | 48.7671 | 0.006974 | 0.21496 | 0.001499 | 30.8231 | 14.5622 | -0.1019 | 0.003715 | 58.1316 | 38.6026 | 61.8578 | 30.9957
2012-10-30 20:18 | 28.9963 | 72.5368 | 53.0442 | 76.3017 | 45.4248 | -1.4679 | 0.006744 | 0.32872 | 0.002217 | 48.7433 | 0.006961 | 0.21502 | 0.001497 | 30.889 | 14.5384 | -0.1012 | 0.003714 | 57.9984 | 38.5058 | 61.7633 | 30.8864
2012-10-30 20:18 | 28.9941 | 72.3545 | 52.9211 | 76.1199 | 45.3598 | -1.4812 | 0.006747 | 0.3288 | 0.002218 | 48.7323 | 0.006958 | 0.21499 | 0.001496 | 30.8981 | 14.5362 | -0.1021 | 0.003714 | 57.8183 | 38.3849 | 61.5837 | 30.8236
2012-10-30 20:18 | 28.9888 | 72.2228 | 52.8221 | 76 | 45.2445 | -1.4772 | 0.006757 | 0.32878 | 0.002222 | 48.658 | 0.006968 | 0.21503 | 0.001498 | 30.8599 | 14.5309 | -0.1018 | 0.00372 | 57.6919 | 38.2912 | 61.4691 | 30.7136
2012-10-30 20:18 | 28.9896 | 72.0587 | 52.7059 | 75.8479 | 45.1472 | -1.4847 | 0.006747 | 0.32872 | 0.002218 | 48.7217 | 0.006974 | 0.215 | 0.001499 | 30.8292 | 14.5602 | -0.1024 | 0.003717 | 57.4985 | 38.1457 | 61.2877 | 30.587
2012-10-30 20:18 | 29.0001 | 71.9159 | 52.6735 | 75.6788 | 45.0949 | -1.483 | 0.006757 | 0.32874 | 0.002221 | 48.6516 | 0.006958 | 0.21496 | 0.001496 | 30.8939 | 14.5748 | -0.1022 | 0.003717 | 57.3411 | 38.0987 | 61.104 | 30.5201
2012-10-30 20:18 | 28.9941 | 71.7757 | 52.5387 | 75.5734 | 45.0221 | -1.4743 | 0.006757 | 0.32877 | 0.002221 | 48.6559 | 0.006971 | 0.21497 | 0.001499 | 30.8383 | 14.5647 | -0.1016 | 0.00372 | 57.211 | 37.974 | 61.0087 | 30.4574
2012-10-30 20:18 | 29.0015 | 71.6067 | 52.4564 | 75.4088 | 44.9785 | -1.4691 | 0.006763 | 0.32864 | 0.002223 | 48.5935 | 0.006961 | 0.215 | 0.001497 | 30.8868 | 14.6252 | -0.1013 | 0.00372 | 56.9815 | 37.8312 | 60.7836 | 30.3533
2012-10-30 20:18 | 28.9963 | 71.4673 | 52.3263 | 75.2506 | 44.8633 | -1.4806 | 0.006763 | 0.32868 | 0.002223 | 48.5999 | 0.006955 | 0.215 | 0.001495 | 30.9134 | 14.62 | -0.1021 | 0.003718 | 56.8473 | 37.7063 | 60.6306 | 30.2433
2012-10-30 20:18 | 28.9926 | 71.291 | 52.2874 | 75.0825 | 44.7693 | -1.4858 | 0.006763 | 0.32877 | 0.002223 | 48.6127 | 0.006965 | 0.21502 | 0.001498 | 30.8712 | 14.6204 | -0.1024 | 0.003721 | 56.6706 | 37.667 | 60.4621 | 30.1489
2012-10-30 20:18 | 28.9896 | 71.2036 | 52.1791 | 74.9801 | 44.676 | -1.4749 | 0.006754 | 0.32875 | 0.00222 | 48.6755 | 0.00698 | 0.21507 | 0.001501 | 30.813 | 14.5807 | -0.1017 | 0.003721 | 56.6229 | 37.5984 | 60.3994 | 30.0953
2012-10-30 20:18 | 28.9941 | 71.0202 | 52.0782 | 74.8086 | 44.5587 | -1.4731 | 0.00676 | 0.32871 | 0.002222 | 48.6257 | 0.006971 | 0.215 | 0.001499 | 30.8425 | 14.6137 | -0.1016 | 0.003721 | 56.4065 | 37.4645 | 60.1949 | 29.945
2012-10-30 20:18 | 28.9843 | 70.8607 | 51.9317 | 74.6726 | 44.5018 | -1.4778 | 0.00676 | 0.3287 | 0.002222 | 48.6237 | 0.00698 | 0.21502 | 0.001501 | 30.8049 | 14.6039 | -0.1019 | 0.003723 | 56.2568 | 37.3278 | 60.0687 | 29.8979
2012-10-30 20:18 | 28.9881 | 70.7034 | 51.8262 | 74.4851 | 44.3759 | -1.4789 | 0.006747 | 0.3287 | 0.002218 | 48.7174 | 0.006961 | 0.21497 | 0.001496 | 30.8826 | 14.6281 | -0.102 | 0.003714 | 56.0753 | 37.1981 | 59.857 | 29.7478
2012-10-30 20:18 | 28.9903 | 70.5597 | 51.7425 | 74.3879 | 44.2917 | -1.4864 | 0.006754 | 0.3287 | 0.00222 | 48.6669 | 0.006968 | 0.21505 | 0.001498 | 30.8621 | 14.5732 | -0.1025 | 0.003718 | 55.9865 | 37.1693 | 59.8147 | 29.7185
2012-10-30 20:19 | 28.9896 | 70.4132 | 51.6363 | 74.2303 | 44.236 | -1.4911 | 0.006741 | 0.32871 | 0.002216 | 48.7628 | 0.006977 | 0.21497 | 0.0015 | 30.8118 | 14.5766 | -0.1028 | 0.003716 | 55.8366 | 37.0597 | 59.6537 | 29.6594
2012-10-30 20:19 | 29.0008 | 70.2209 | 51.5265 | 74.031 | 44.1254 | -1.4853 | 0.00676 | 0.32865 | 0.002222 | 48.6172 | 0.006968 | 0.21499 | 0.001498 | 30.8538 | 14.6367 | -0.1024 | 0.00372 | 55.5842 | 36.8898 | 59.3943 | 29.4887
2012-10-30 20:19 | 28.9978 | 70.0874 | 51.4414 | 73.9017 | 44.0477 | -1.4766 | 0.006773 | 0.32872 | 0.002226 | 48.5346 | 0.006961 | 0.21493 | 0.001496 | 30.8765 | 14.6133 | -0.1018 | 0.003722 | 55.4741 | 36.8281 | 59.2884 | 29.4344
2012-10-30 20:19 | 28.9903 | 69.9531 | 51.348 | 73.7641 | 43.9538 | -1.4801 | 0.00676 | 0.32868 | 0.002222 | 48.6215 | 0.006977 | 0.21503 | 0.0015 | 30.8201 | 14.5936 | -0.102 | 0.003722 | 55.3595 | 36.7544 | 59.1705 | 29.3602
2012-10-30 20:19 | 28.9903 | 69.788 | 51.2191 | 73.6186 | 43.8752 | -1.4916 | 0.006751 | 0.32865 | 0.002219 | 48.682 | 0.006961 | 0.21492 | 0.001496 | 30.8743 | 14.6018 | -0.1028 | 0.003715 | 55.1862 | 36.6173 | 59.0168 | 29.2734
2012-10-30 20:19 | 28.9903 | 69.6421 | 51.1136 | 73.4463 | 43.7416 | -1.4847 | 0.006744 | 0.32868 | 0.002217 | 48.7368 | 0.006968 | 0.21492 | 0.001498 | 30.8433 | 14.5854 | -0.1024 | 0.003715 | 55.0567 | 36.5282 | 58.8609 | 29.1562
2012-10-30 20:19 | 28.9896 | 69.5069 | 51.0308 | 73.2924 | 43.6976 | -1.4806 | 0.006747 | 0.32864 | 0.002217 | 48.7088 | 0.006965 | 0.21499 | 0.001497 | 30.867 | 14.6255 | -0.1021 | 0.003714 | 54.8814 | 36.4053 | 58.6669 | 29.0721
2012-10-30 20:19 | 28.9866 | 69.3464 | 50.9145 | 73.1592 | 43.6042 | -1.4882 | 0.006754 | 0.32865 | 0.00222 | 48.6603 | 0.006965 | 0.215 | 0.001497 | 30.8691 | 14.5899 | -0.1026 | 0.003717 | 54.7565 | 36.3246 | 58.5693 | 29.0143
2012-10-30 20:19 | 28.9851 | 69.1682 | 50.8348 | 73.039 | 43.4612 | -1.4887 | 0.006754 | 0.32865 | 0.00222 | 48.6603 | 0.006974 | 0.21496 | 0.001499 | 30.8231 | 14.621 | -0.1026 | 0.003719 | 54.5472 | 36.2138 | 58.418 | 28.8402
2012-10-30 20:19 | 28.9896 | 69.0843 | 50.7455 | 72.8903 | 43.4461 | -1.4858 | 0.006747 | 0.32864 | 0.002217 | 48.7088 | 0.006977 | 0.215 | 0.0015 | 30.816 | 14.6255 | -0.1024 | 0.003717 | 54.4588 | 36.12 | 58.2648 | 28.8206
2012-10-30 20:19 | 28.9948 | 68.9281 | 50.6843 | 72.7614 | 43.3923 | -1.4801 | 0.006747 | 0.32862 | 0.002217 | 48.7067 | 0.006965 | 0.21505 | 0.001498 | 30.8754 | 14.6144 | -0.102 | 0.003715 | 54.3137 | 36.0699 | 58.147 | 28.7779
2012-10-30 20:19 | 28.9843 | 68.7369 | 50.5409 | 72.5785 | 43.3071 | -1.4893 | 0.006763 | 0.32862 | 0.002222 | 48.5915 | 0.006958 | 0.215 | 0.001496 | 30.9001 | 14.5876 | -0.1027 | 0.003718 | 54.1493 | 35.9533 | 57.9909 | 28.7195
2012-10-30 20:19 | 28.9881 | 68.5523 | 50.4587 | 72.4558 | 43.1812 | -1.4824 | 0.006747 | 0.32862 | 0.002217 | 48.7067 | 0.006955 | 0.215 | 0.001495 | 30.9134 | 14.5628 | -0.1022 | 0.003712 | 53.9895 | 35.8959 | 57.893 | 28.6184
2012-10-30 20:19 | 28.9896 | 68.4462 | 50.3155 | 72.2887 | 43.0686 | -1.4754 | 0.00676 | 0.32872 | 0.002222 | 48.628 | 0.006971 | 0.21496 | 0.001498 | 30.8363 | 14.5439 | -0.1017 | 0.00372 | 53.9023 | 35.7716 | 57.7448 | 28.5247
2012-10-30 20:19 | 28.9806 | 68.3026 | 50.2205 | 72.1149 | 42.9849 | -1.4812 | 0.006751 | 0.32868 | 0.002219 | 48.6863 | 0.006965 | 0.21503 | 0.001498 | 30.8732 | 14.5308 | -0.1021 | 0.003717 | 53.7718 | 35.6897 | 57.5841 | 28.4541
2012-10-30 20:19 | 28.9926 | 68.1492 | 50.1621 | 71.9927 | 42.9143 | -1.4778 | 0.006754 | 0.32861 | 0.002219 | 48.654 | 0.006971 | 0.21497 | 0.001499 | 30.8383 | 14.5306 | -0.1019 | 0.003718 | 53.6186 | 35.6315 | 57.4621 | 28.3837
2012-10-30 20:19 | 29.0008 | 67.9689 | 49.9982 | 71.8131 | 42.7692 | -1.4864 | 0.006754 | 0.32864 | 0.00222 | 48.6583 | 0.006974 | 0.21503 | 0.0015 | 30.8334 | 14.551 | -0.1025 | 0.00372 | 53.4179 | 35.4472 | 57.2621 | 28.2182
2012-10-30 20:19 | 28.9903 | 67.8623 | 49.886 | 71.6799 | 42.6957 | -1.4806 | 0.006751 | 0.32854 | 0.002218 | 48.6649 | 0.006971 | 0.21494 | 0.001498 | 30.8342 | 14.5936 | -0.1021 | 0.003716 | 53.2687 | 35.2924 | 57.0863 | 28.1021
2012-10-30 20:19 | 28.9933 | 67.669 | 49.7952 | 71.5372 | 42.5926 | -1.4766 | 0.00676 | 0.32852 | 0.002221 | 48.5979 | 0.006952 | 0.21494 | 0.001494 | 30.9184 | 14.5354 | -0.1018 | 0.003715 | 53.1336 | 35.2598 | 57.0018 | 28.0572
2012-10-30 20:19 | 28.9941 | 67.5506 | 49.7099 | 71.3692 | 42.4832 | -1.4789 | 0.006747 | 0.32865 | 0.002217 | 48.7108 | 0.006971 | 0.21502 | 0.001499 | 30.8446 | 14.5321 | -0.102 | 0.003716 | 53.0185 | 35.1778 | 56.8371 | 27.9511
2012-10-30 20:19 | 29.0001 | 67.418 | 49.5907 | 71.2602 | 42.442 | -1.4656 | 0.006757 | 0.32857 | 0.00222 | 48.626 | 0.006952 | 0.21497 | 0.001494 | 30.9226 | 14.5381 | -0.101 | 0.003714 | 52.8799 | 35.0526 | 56.7221 | 27.9039
2012-10-30 20:19 | 28.9941 | 67.2696 | 49.4947 | 71.1199 | 42.3534 | -1.4662 | 0.006754 | 0.32861 | 0.002219 | 48.654 | 0.006958 | 0.215 | 0.001496 | 30.9001 | 14.4668 | -0.1011 | 0.003715 | 52.8028 | 35.0279 | 56.6531 | 27.8866
2012-10-30 20:19 | 28.9993 | 67.0977 | 49.3864 | 70.9525 | 42.2327 | -1.461 | 0.006754 | 0.3286 | 0.002219 | 48.6519 | 0.006961 | 0.21497 | 0.001496 | 30.8826 | 14.5087 | -0.1007 | 0.003715 | 52.589 | 34.8777 | 56.4438 | 27.724
2012-10-30 20:19 | 28.9956 | 66.94 | 49.281 | 70.7339 | 42.1936 | -1.446 | 0.006754 | 0.3286 | 0.002219 | 48.6519 | 0.006955 | 0.21497 | 0.001495 | 30.9093 | 14.505 | -0.0997 | 0.003714 | 52.435 | 34.776 | 56.2289 | 27.6886
2012-10-30 20:20 | 28.9971 | 66.7952 | 49.1729 | 70.6126 | 42.081 | -1.4512 | 0.00676 | 0.32854 | 0.002221 | 48.6001 | 0.006952 | 0.21493 | 0.001494 | 30.9164 | 14.5188 | -0.1 | 0.003715 | 52.2764 | 34.6541 | 56.0938 | 27.5622
2012-10-30 20:20 | 28.9941 | 66.5958 | 49.0603 | 70.433 | 42.0111 | -1.4587 | 0.006751 | 0.32848 | 0.002218 | 48.6563 | 0.006971 | 0.21499 | 0.001499 | 30.8405 | 14.528 | -0.1006 | 0.003717 | 52.0678 | 34.5323 | 55.905 | 27.4831
2012-10-30 20:20 | 28.9978 | 66.4686 | 48.9505 | 70.3139 | 41.9007 | -1.4645 | 0.006754 | 0.32851 | 0.002219 | 48.639 | 0.006965 | 0.21493 | 0.001497 | 30.8587 | 14.5399 | -0.101 | 0.003716 | 51.9287 | 34.4106 | 55.774 | 27.3608
2012-10-30 20:20 | 29.006 | 66.3304 | 48.8687 | 70.1455 | 41.7672 | -1.45 | 0.006738 | 0.32849 | 0.002213 | 48.7523 | 0.006955 | 0.21497 | 0.001495 | 30.9093 | 14.5195 | -0.1 | 0.003708 | 51.8109 | 34.3492 | 55.626 | 27.2477
2012-10-30 20:20 | 28.9941 | 66.122 | 48.7393 | 69.98 | 41.716 | -1.4633 | 0.006744 | 0.32852 | 0.002216 | 48.7132 | 0.006958 | 0.21496 | 0.001496 | 30.8939 | 14.5484 | -0.1009 | 0.003712 | 51.5736 | 34.1909 | 55.4316 | 27.1676
2012-10-30 20:20 | 28.9963 | 65.9623 | 48.624 | 69.8055 | 41.6277 | -1.4471 | 0.006754 | 0.32855 | 0.002219 | 48.6454 | 0.006965 | 0.2149 | 0.001497 | 30.8546 | 14.5343 | -0.0998 | 0.003716 | 51.428 | 34.0897 | 55.2712 | 27.0934
2012-10-30 20:20 | 29.0023 | 65.8065 | 48.5635 | 69.6579 | 41.4999 | -1.4477 | 0.006751 | 0.32847 | 0.002217 | 48.6543 | 0.006955 | 0.21496 | 0.001495 | 30.9073 | 14.5199 | -0.0998 | 0.003712 | 51.2866 | 34.0436 | 55.138 | 26.98
2012-10-30 20:20 | 29.0015 | 65.6592 | 48.4648 | 69.538 | 41.4282 | -1.4668 | 0.006754 | 0.32852 | 0.002219 | 48.6411 | 0.006965 | 0.21484 | 0.001496 | 30.8462 | 14.5517 | -0.1011 | 0.003715 | 51.1075 | 33.9131 | 54.9863 | 26.8765
2012-10-30 20:20 | 29.0015 | 65.5474 | 48.3904 | 69.3613 | 41.3652 | -1.446 | 0.006754 | 0.32849 | 0.002219 | 48.6368 | 0.006965 | 0.2149 | 0.001497 | 30.8546 | 14.5272 | -0.0997 | 0.003716 | 51.0202 | 33.8632 | 54.8341 | 26.838
2012-10-30 20:20 | 29.0023 | 65.3901 | 48.2815 | 69.2123 | 41.2834 | -1.4587 | 0.006744 | 0.32848 | 0.002215 | 48.7069 | 0.006971 | 0.21493 | 0.001498 | 30.8322 | 14.5076 | -0.1006 | 0.003713 | 50.8825 | 33.7739 | 54.7047 | 26.7758
2012-10-30 20:20 | 29.006 | 65.2473 | 48.2147 | 69.1007 | 41.1611 | -1.4471 | 0.006754 | 0.32848 | 0.002219 | 48.6347 | 0.006952 | 0.21499 | 0.001495 | 30.9248 | 14.5032 | -0.0998 | 0.003714 | 50.7441 | 33.7115 | 54.5975 | 26.6579
2012-10-30 20:20 | 29.0023 | 65.0971 | 48.0778 | 68.9509 | 41.0944 | -1.4575 | 0.006744 | 0.32845 | 0.002215 | 48.7026 | 0.006952 | 0.21496 | 0.001494 | 30.9206 | 14.5852 | -0.1005 | 0.003709 | 50.5119 | 33.4926 | 54.3657 | 26.5092
2012-10-30 20:20 | 28.9971 | 64.9337 | 47.9864 | 68.8035 | 41.0498 | -1.446 | 0.006751 | 0.32851 | 0.002218 | 48.6606 | 0.006965 | 0.21493 | 0.001497 | 30.8587 | 14.531 | -0.0997 | 0.003715 | 50.4027 | 33.4554 | 54.2725 | 26.5188
2012-10-30 20:20 | 29.006 | 64.8153 | 47.8856 | 68.6932 | 40.9603 | -1.4662 | 0.006751 | 0.32851 | 0.002218 | 48.6606 | 0.006955 | 0.2149 | 0.001495 | 30.8989 | 14.5562 | -0.1011 | 0.003713 | 50.2591 | 33.3294 | 54.137 | 26.4041
2012-10-30 20:20 | 29.0008 | 64.6751 | 47.8138 | 68.542 | 40.9039 | -1.4529 | 0.006738 | 0.32845 | 0.002213 | 48.7459 | 0.006955 | 0.21486 | 0.001494 | 30.8927 | 14.5674 | -0.1002 | 0.003707 | 50.1077 | 33.2464 | 53.9746 | 26.3365
2012-10-30 20:20 | 29.0075 | 64.5197 | 47.6912 | 68.3872 | 40.8437 | -1.4471 | 0.006747 | 0.32854 | 0.002217 | 48.6938 | 0.006952 | 0.21496 | 0.001494 | 30.9206 | 14.5373 | -0.0998 | 0.003711 | 49.9824 | 33.1539 | 53.8499 | 26.3064
2012-10-30 20:20 | 29.009 | 64.3862 | 47.6183 | 68.2656 | 40.7389 | -1.4552 | 0.006757 | 0.32845 | 0.002219 | 48.6088 | 0.006958 | 0.21497 | 0.001496 | 30.8959 | 14.5062 | -0.1003 | 0.003715 | 49.88 | 33.1121 | 53.7594 | 26.2327
2012-10-30 20:20 | 29.0045 | 64.235 | 47.5158 | 68.1111 | 40.6753 | -1.4454 | 0.006741 | 0.32841 | 0.002214 | 48.7178 | 0.006939 | 0.2149 | 0.001491 | 30.9702 | 14.5302 | -0.0996 | 0.003705 | 49.7048 | 32.9856 | 53.5809 | 26.1451
2012-10-30 20:20 | 29.0068 | 64.1061 | 47.4397 | 67.9711 | 40.5792 | -1.4529 | 0.00676 | 0.32844 | 0.00222 | 48.5852 | 0.006958 | 0.21499 | 0.001496 | 30.8981 | 14.5244 | -0.1002 | 0.003716 | 49.5817 | 32.9153 | 53.4467 | 26.0548
2012-10-30 20:20 | 29.0083 | 63.903 | 47.355 | 67.7725 | 40.4665 | -1.4477 | 0.006747 | 0.32851 | 0.002216 | 48.6895 | 0.006958 | 0.21497 | 0.001496 | 30.8959 | 14.5953 | -0.0998 | 0.003712 | 49.3077 | 32.7597 | 53.1772 | 25.8712
2012-10-30 20:20 | 29.0098 | 63.777 | 47.2624 | 67.6586 | 40.4128 | -1.4587 | 0.006747 | 0.32838 | 0.002216 | 48.6702 | 0.006958 | 0.21489 | 0.001495 | 30.8836 | 14.5804 | -0.1006 | 0.003711 | 49.1966 | 32.682 | 53.0782 | 25.8324
2012-10-30 20:20 | 29.0113 | 63.6897 | 47.2012 | 67.5216 | 40.3474 | -1.4593 | 0.006738 | 0.32839 | 0.002213 | 48.7373 | 0.006974 | 0.21487 | 0.001499 | 30.8106 | 14.5942 | -0.1006 | 0.003712 | 49.0955 | 32.607 | 52.9274 | 25.7532
2012-10-30 20:20 | 29.0008 | 63.5402 | 47.1044 | 67.4033 | 40.199 | -1.4604 | 0.006741 | 0.32842 | 0.002214 | 48.7199 | 0.006961 | 0.21493 | 0.001496 | 30.8765 | 14.5837 | -0.1007 | 0.00371 | 48.9565 | 32.5207 | 52.8196 | 25.6153
2012-10-30 20:20 | 29.0038 | 63.3616 | 46.9584 | 67.2947 | 40.1429 | -1.4604 | 0.006747 | 0.32847 | 0.002216 | 48.6831 | 0.006952 | 0.2149 | 0.001494 | 30.9123 | 14.5704 | -0.1007 | 0.00371 | 48.7912 | 32.388 | 52.7243 | 25.5725
2012-10-30 20:21 | 28.9918 | 63.2376 | 46.9621 | 67.121 | 40.0246 | -1.4674 | 0.006738 | 0.32849 | 0.002213 | 48.7523 | 0.006971 | 0.21496 | 0.001498 | 30.8363 | 14.5992 | -0.1012 | 0.003711 | 48.6384 | 32.3629 | 52.5218 | 25.4254
2012-10-30 20:21 | 28.9978 | 63.1238 | 46.8622 | 66.9961 | 39.9872 | -1.4575 | 0.006747 | 0.32838 | 0.002216 | 48.6702 | 0.006965 | 0.21496 | 0.001497 | 30.8629 | 14.5766 | -0.1005 | 0.003713 | 48.5472 | 32.2856 | 52.4195 | 25.4106
2012-10-30 20:21 | 29.0023 | 62.9892 | 46.7648 | 66.8543 | 39.9208 | -1.4483 | 0.006741 | 0.32838 | 0.002214 | 48.7135 | 0.006961 | 0.215 | 0.001497 | 30.8868 | 14.6097 | -0.0998 | 0.003711 | 48.3795 | 32.1551 | 52.2446 | 25.3111
2012-10-30 20:21 | 28.9993 | 62.8239 | 46.6598 | 66.7243 | 39.8981 | -1.4645 | 0.006747 | 0.32835 | 0.002215 | 48.6659 | 0.006974 | 0.21499 | 0.001499 | 30.8272 | 14.6148 | -0.101 | 0.003714 | 48.2091 | 32.045 | 52.1095 | 25.2833
2012-10-30 20:21 | 28.9911 | 62.7384 | 46.5653 | 66.5967 | 39.8151 | -1.4604 | 0.006763 | 0.32839 | 0.002221 | 48.5571 | 0.006977 | 0.21492 | 0.001499 | 30.8035 | 14.5781 | -0.1007 | 0.00372 | 48.1603 | 31.9872 | 52.0186 | 25.237
2012-10-30 20:21 | 28.9903 | 62.5443 | 46.4508 | 66.4187 | 39.6843 | -1.4575 | 0.006754 | 0.32849 | 0.002219 | 48.6368 | 0.006971 | 0.215 | 0.001499 | 30.8425 | 14.6099 | -0.1005 | 0.003718 | 47.9344 | 31.8409 | 51.8088 | 25.0744
2012-10-30 20:21 | 28.9821 | 62.4666 | 46.3838 | 66.2872 | 39.5855 | -1.4489 | 0.00676 | 0.32841 | 0.00222 | 48.5809 | 0.006943 | 0.21486 | 0.001492 | 30.9461 | 14.5895 | -0.0999 | 0.003712 | 47.8771 | 31.7943 | 51.6977 | 24.996
2012-10-30 20:21 | 28.9746 | 62.316 | 46.2939 | 66.1795 | 39.5228 | -1.4656 | 0.006751 | 0.32842 | 0.002217 | 48.6478 | 0.006952 | 0.21489 | 0.001494 | 30.9102 | 14.5697 | -0.101 | 0.003711 | 47.7463 | 31.7242 | 51.6098 | 24.9531
2012-10-30 20:21 | 28.9858 | 62.1841 | 46.2423 | 66.0366 | 39.4276 | -1.4627 | 0.006747 | 0.32838 | 0.002216 | 48.6702 | 0.006946 | 0.21486 | 0.001492 | 30.9328 | 14.5646 | -0.1008 | 0.003708 | 47.6195 | 31.6777 | 51.472 | 24.863
2012-10-30 20:21 | 28.9806 | 61.9855 | 46.0527 | 65.8657 | 39.3318 | -1.4598 | 0.006744 | 0.32844 | 0.002215 | 48.7005 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 14.588 | -0.1006 | 0.003705 | 47.3975 | 31.4647 | 51.2777 | 24.7438
2012-10-30 20:21 | 28.9708 | 61.8404 | 45.9801 | 65.748 | 39.2905 | -1.4708 | 0.006747 | 0.32835 | 0.002215 | 48.6659 | 0.006965 | 0.21486 | 0.001496 | 30.8484 | 14.5374 | -0.1014 | 0.003711 | 47.303 | 31.4427 | 51.2106 | 24.7531
2012-10-30 20:21 | 28.9693 | 61.7189 | 45.9237 | 65.6116 | 39.0959 | -1.4587 | 0.006738 | 0.32836 | 0.002213 | 48.733 | 0.006961 | 0.21494 | 0.001496 | 30.8785 | 14.5522 | -0.1006 | 0.003709 | 47.1667 | 31.3715 | 51.0594 | 24.5437
2012-10-30 20:21 | 28.9708 | 61.5812 | 45.7996 | 65.4358 | 39.0934 | -1.4668 | 0.006751 | 0.32832 | 0.002216 | 48.6328 | 0.006958 | 0.21493 | 0.001495 | 30.8898 | 14.5537 | -0.1011 | 0.003711 | 47.0275 | 31.2459 | 50.8821 | 24.5397
2012-10-30 20:21 | 28.9768 | 61.4672 | 45.6957 | 65.3261 | 39.0167 | -1.4552 | 0.006741 | 0.32836 | 0.002213 | 48.7113 | 0.006968 | 0.21492 | 0.001498 | 30.8433 | 14.5842 | -0.1003 | 0.003711 | 46.883 | 31.1115 | 50.7419 | 24.4325
2012-10-30 20:21 | 28.9746 | 61.2947 | 45.619 | 65.1619 | 39.0145 | -1.457 | 0.006747 | 0.32838 | 0.002216 | 48.6702 | 0.006968 | 0.21493 | 0.001498 | 30.8454 | 14.5452 | -0.1004 | 0.003714 | 46.7495 | 31.0738 | 50.6167 | 24.4693
2012-10-30 20:21 | 28.9753 | 61.1677 | 45.4941 | 65.0778 | 38.9166 | -1.4645 | 0.00676 | 0.32831 | 0.002219 | 48.566 | 0.006961 | 0.21497 | 0.001496 | 30.8826 | 14.5623 | -0.101 | 0.003715 | 46.6054 | 30.9318 | 50.5155 | 24.3543
2012-10-30 20:21 | 28.9731 | 61.0726 | 45.4448 | 64.9213 | 38.808 | -1.4598 | 0.006751 | 0.32833 | 0.002217 | 48.635 | 0.006965 | 0.21492 | 0.001497 | 30.8566 | 14.5152 | -0.1006 | 0.003714 | 46.5574 | 30.9296 | 50.4061 | 24.2928
2012-10-30 20:21 | 28.9723 | 60.8898 | 45.334 | 64.8125 | 38.756 | -1.4535 | 0.006738 | 0.32832 | 0.002212 | 48.7266 | 0.006958 | 0.21487 | 0.001495 | 30.8814 | 14.5144 | -0.1002 | 0.003707 | 46.3754 | 30.8196 | 50.2981 | 24.2416
2012-10-30 20:21 | 28.9761 | 60.7581 | 45.2593 | 64.6388 | 38.6888 | -1.4575 | 0.006757 | 0.32836 | 0.002219 | 48.596 | 0.006955 | 0.21483 | 0.001494 | 30.8886 | 14.5549 | -0.1005 | 0.003713 | 46.2032 | 30.7044 | 50.0839 | 24.1339
2012-10-30 20:21 | 28.9746 | 60.6443 | 45.1715 | 64.5331 | 38.5809 | -1.4564 | 0.00676 | 0.32833 | 0.00222 | 48.5703 | 0.006949 | 0.21493 | 0.001494 | 30.9298 | 14.5616 | -0.1004 | 0.003714 | 46.0827 | 30.6099 | 49.9715 | 24.0193
2012-10-30 20:21 | 28.9708 | 60.4972 | 45.097 | 64.3981 | 38.5101 | -1.4622 | 0.00676 | 0.32833 | 0.00222 | 48.5703 | 0.006955 | 0.21489 | 0.001495 | 30.8969 | 14.5578 | -0.1008 | 0.003715 | 45.9394 | 30.5392 | 49.8403 | 23.9523
2012-10-30 20:21 | 28.9768 | 60.3444 | 45.0087 | 64.2382 | 38.4648 | -1.461 | 0.00676 | 0.32833 | 0.00222 | 48.5703 | 0.006955 | 0.21493 | 0.001495 | 30.9031 | 14.5515 | -0.1007 | 0.003715 | 45.7929 | 30.4572 | 49.6867 | 23.9133
2012-10-30 20:21 | 28.9738 | 60.2135 | 44.9036 | 64.0499 | 38.3909 | -1.4622 | 0.006751 | 0.32825 | 0.002216 | 48.6221 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.5444 | -0.1008 | 0.003711 | 45.6691 | 30.3592 | 49.5055 | 23.8465
2012-10-30 20:22 | 28.9738 | 60.1012 | 44.8408 | 63.9573 | 38.2411 | -1.4535 | 0.006747 | 0.32831 | 0.002215 | 48.6596 | 0.006955 | 0.21483 | 0.001494 | 30.8886 | 14.5608 | -0.1002 | 0.003709 | 45.5404 | 30.28 | 49.3965 | 23.6803
2012-10-30 20:22 | 28.9843 | 59.9489 | 44.7531 | 63.8597 | 38.2752 | -1.4685 | 0.006751 | 0.32829 | 0.002216 | 48.6285 | 0.006958 | 0.21492 | 0.001495 | 30.8876 | 14.5305 | -0.1012 | 0.003711 | 45.4184 | 30.2226 | 49.3292 | 23.7447
2012-10-30 20:22 | 28.9693 | 59.837 | 44.6753 | 63.6902 | 38.1419 | -1.4633 | 0.006751 | 0.32828 | 0.002216 | 48.6264 | 0.006952 | 0.21487 | 0.001494 | 30.9081 | 14.5604 | -0.1009 | 0.00371 | 45.2766 | 30.1149 | 49.1298 | 23.5815
2012-10-30 20:22 | 28.9791 | 59.7189 | 44.5673 | 63.5648 | 38.1162 | -1.4685 | 0.00676 | 0.32833 | 0.00222 | 48.5703 | 0.006952 | 0.21489 | 0.001494 | 30.9102 | 14.5824 | -0.1012 | 0.003714 | 45.1365 | 29.9849 | 48.9824 | 23.5338
2012-10-30 20:22 | 28.9738 | 59.5741 | 44.4716 | 63.4707 | 38.0124 | -1.4783 | 0.006754 | 0.32826 | 0.002217 | 48.6026 | 0.006952 | 0.2148 | 0.001493 | 30.8977 | 14.5363 | -0.1019 | 0.00371 | 45.0378 | 29.9353 | 48.9344 | 23.4761
2012-10-30 20:22 | 28.9753 | 59.4399 | 44.4456 | 63.31 | 37.9429 | -1.4702 | 0.006738 | 0.32828 | 0.002212 | 48.7202 | 0.006958 | 0.2149 | 0.001495 | 30.8856 | 14.5827 | -0.1014 | 0.003707 | 44.8572 | 29.8629 | 48.7273 | 23.3602
2012-10-30 20:22 | 28.9761 | 59.3282 | 44.3246 | 63.2605 | 37.9161 | -1.4674 | 0.006763 | 0.32821 | 0.00222 | 48.5295 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.559 | -0.1012 | 0.003715 | 44.7692 | 29.7656 | 48.7015 | 23.3571
2012-10-30 20:22 | 28.9701 | 59.171 | 44.2676 | 63.0691 | 37.8115 | -1.4679 | 0.006744 | 0.32826 | 0.002214 | 48.6747 | 0.006961 | 0.21487 | 0.001496 | 30.8681 | 14.553 | -0.1012 | 0.00371 | 44.618 | 29.7146 | 48.5161 | 23.2585
2012-10-30 20:22 | 28.9686 | 59.0454 | 44.1914 | 62.9208 | 37.7469 | -1.4702 | 0.006751 | 0.32823 | 0.002216 | 48.6199 | 0.006955 | 0.2149 | 0.001495 | 30.8989 | 14.5392 | -0.1014 | 0.003711 | 44.5062 | 29.6522 | 48.3816 | 23.2077
2012-10-30 20:22 | 28.9678 | 58.9321 | 44.112 | 62.8272 | 37.758 | -1.461 | 0.006744 | 0.32828 | 0.002214 | 48.6769 | 0.006965 | 0.21493 | 0.001497 | 30.8587 | 14.5548 | -0.1007 | 0.003711 | 44.3773 | 29.5572 | 48.2724 | 23.2032
2012-10-30 20:22 | 28.9641 | 58.7926 | 44.014 | 62.6882 | 37.6951 | -1.4598 | 0.006744 | 0.32826 | 0.002214 | 48.6747 | 0.006958 | 0.21494 | 0.001496 | 30.8918 | 14.5756 | -0.1006 | 0.00371 | 44.217 | 29.4384 | 48.1126 | 23.1195
2012-10-30 20:22 | 28.9708 | 58.6985 | 43.9696 | 62.5442 | 37.5796 | -1.4662 | 0.006744 | 0.32822 | 0.002214 | 48.6683 | 0.006952 | 0.21486 | 0.001494 | 30.9061 | 14.5863 | -0.1011 | 0.003708 | 44.1122 | 29.3833 | 47.9579 | 22.9933
2012-10-30 20:22 | 28.9723 | 58.5642 | 43.8533 | 62.4336 | 37.5061 | -1.476 | 0.006754 | 0.32822 | 0.002217 | 48.5962 | 0.006968 | 0.21493 | 0.001498 | 30.8454 | 14.6001 | -0.1018 | 0.003715 | 43.9641 | 29.2532 | 47.8335 | 22.906
2012-10-30 20:22 | 28.9648 | 58.3781 | 43.775 | 62.3255 | 37.4237 | -1.472 | 0.006751 | 0.32826 | 0.002216 | 48.6242 | 0.006961 | 0.21483 | 0.001495 | 30.8619 | 14.5681 | -0.1015 | 0.003711 | 43.81 | 29.2069 | 47.7574 | 22.8556
2012-10-30 20:22 | 28.9686 | 58.2966 | 43.6648 | 62.2172 | 37.3013 | -1.4674 | 0.006738 | 0.32822 | 0.002212 | 48.7116 | 0.006965 | 0.21483 | 0.001496 | 30.8442 | 14.5229 | -0.1012 | 0.003708 | 43.7737 | 29.1419 | 47.6943 | 22.7784
2012-10-30 20:22 | 28.9708 | 58.1978 | 43.6199 | 62.0454 | 37.2838 | -1.4772 | 0.006735 | 0.32822 | 0.002211 | 48.7333 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 14.5945 | -0.1018 | 0.003702 | 43.6033 | 29.0254 | 47.4509 | 22.6893
2012-10-30 20:22 | 28.9738 | 58.0456 | 43.5443 | 61.9362 | 37.2079 | -1.476 | 0.006766 | 0.32825 | 0.002221 | 48.5143 | 0.006961 | 0.21482 | 0.001495 | 30.8598 | 14.5649 | -0.1018 | 0.003716 | 43.4807 | 28.9794 | 47.3713 | 22.643
2012-10-30 20:22 | 28.9701 | 57.8866 | 43.399 | 61.8397 | 37.0977 | -1.4674 | 0.006757 | 0.32829 | 0.002218 | 48.5853 | 0.006961 | 0.21489 | 0.001496 | 30.8703 | 14.553 | -0.1012 | 0.003714 | 43.3336 | 28.846 | 47.2867 | 22.5447
2012-10-30 20:22 | 28.9641 | 57.8184 | 43.3813 | 61.6557 | 37.072 | -1.4783 | 0.006751 | 0.32821 | 0.002216 | 48.6158 | 0.006952 | 0.21486 | 0.001494 | 30.9061 | 14.5674 | -0.1019 | 0.00371 | 43.251 | 28.8139 | 47.0883 | 22.5046
2012-10-30 20:22 | 28.9678 | 57.6474 | 43.3142 | 61.5627 | 37.0283 | -1.483 | 0.006741 | 0.32818 | 0.002212 | 48.6836 | 0.006946 | 0.21489 | 0.001493 | 30.9369 | 14.5833 | -0.1022 | 0.003705 | 43.0641 | 28.7309 | 46.9794 | 22.445
2012-10-30 20:22 | 28.9648 | 57.5589 | 43.2208 | 61.4397 | 36.9859 | -1.4864 | 0.006738 | 0.32821 | 0.002211 | 48.7096 | 0.006952 | 0.21489 | 0.001494 | 30.9102 | 14.5967 | -0.1025 | 0.003705 | 42.9622 | 28.6241 | 46.843 | 22.3892
2012-10-30 20:22 | 28.9671 | 57.4175 | 43.1837 | 61.3259 | 36.8541 | -1.4737 | 0.006754 | 0.32823 | 0.002217 | 48.5983 | 0.006958 | 0.21486 | 0.001495 | 30.8794 | 14.5337 | -0.1016 | 0.003712 | 42.8838 | 28.65 | 46.7922 | 22.3204
2012-10-30 20:22 | 28.9589 | 57.2733 | 43.0694 | 61.1706 | 36.7985 | -1.4743 | 0.006747 | 0.32821 | 0.002214 | 48.6446 | 0.006955 | 0.21486 | 0.001494 | 30.8927 | 14.5785 | -0.1016 | 0.003708 | 42.6948 | 28.4909 | 46.5921 | 22.22
2012-10-30 20:22 | 28.9708 | 57.1493 | 43.0105 | 61.0896 | 36.7158 | -1.4691 | 0.006763 | 0.32815 | 0.002219 | 48.5209 | 0.006946 | 0.21486 | 0.001492 | 30.9328 | 14.57 | -0.1013 | 0.003711 | 42.5793 | 28.4405 | 46.5196 | 22.1458
2012-10-30 20:22 | 28.9708 | 57.0832 | 42.8886 | 60.9503 | 36.7237 | -1.4737 | 0.006754 | 0.32822 | 0.002217 | 48.5962 | 0.006946 | 0.21482 | 0.001492 | 30.9264 | 14.5823 | -0.1016 | 0.003709 | 42.5009 | 28.3063 | 46.368 | 22.1414
2012-10-30 20:23 | 28.9559 | 56.9051 | 42.8108 | 60.7766 | 36.6338 | -1.4772 | 0.006747 | 0.32816 | 0.002214 | 48.6381 | 0.006952 | 0.21486 | 0.001494 | 30.9061 | 14.5633 | -0.1018 | 0.003708 | 42.3418 | 28.2475 | 46.2133 | 22.0705
2012-10-30 20:23 | 28.9581 | 56.7674 | 42.7107 | 60.651 | 36.5255 | -1.4772 | 0.006751 | 0.32823 | 0.002216 | 48.6199 | 0.006949 | 0.21486 | 0.001493 | 30.9194 | 14.5328 | -0.1018 | 0.003709 | 42.2346 | 28.1779 | 46.1182 | 21.9927
2012-10-30 20:23 | 28.9544 | 56.7092 | 42.6441 | 60.5699 | 36.5218 | -1.4778 | 0.006741 | 0.32812 | 0.002212 | 48.675 | 0.006971 | 0.21487 | 0.001498 | 30.8238 | 14.5659 | -0.1019 | 0.00371 | 42.1433 | 28.0782 | 46.004 | 21.9559
2012-10-30 20:23 | 28.9581 | 56.5497 | 42.5574 | 60.4419 | 36.4743 | -1.4778 | 0.006735 | 0.32816 | 0.00221 | 48.7247 | 0.006955 | 0.21486 | 0.001494 | 30.8927 | 14.5165 | -0.1019 | 0.003704 | 42.0332 | 28.0409 | 45.9254 | 21.9578
2012-10-30 20:23 | 28.9551 | 56.4417 | 42.5032 | 60.3575 | 36.3687 | -1.4911 | 0.006751 | 0.32818 | 0.002216 | 48.6115 | 0.006952 | 0.21487 | 0.001494 | 30.9081 | 14.5013 | -0.1028 | 0.00371 | 41.9404 | 28.0019 | 45.8562 | 21.8674
2012-10-30 20:23 | 28.9529 | 56.3112 | 42.3987 | 60.181 | 36.3073 | -1.4749 | 0.006729 | 0.32813 | 0.002208 | 48.7639 | 0.006958 | 0.21482 | 0.001495 | 30.8731 | 14.5317 | -0.1017 | 0.003703 | 41.7795 | 27.867 | 45.6493 | 21.7756
2012-10-30 20:23 | 28.9476 | 56.1776 | 42.2833 | 60.0517 | 36.1757 | -1.476 | 0.006747 | 0.32816 | 0.002214 | 48.6381 | 0.006955 | 0.21487 | 0.001494 | 30.8948 | 14.506 | -0.1018 | 0.003708 | 41.6716 | 27.7773 | 45.5457 | 21.6697
2012-10-30 20:23 | 28.9589 | 56.0955 | 42.2513 | 59.9816 | 36.1515 | -1.483 | 0.006738 | 0.32809 | 0.002211 | 48.6923 | 0.006955 | 0.21492 | 0.001495 | 30.9009 | 14.5418 | -0.1022 | 0.003706 | 41.5537 | 27.7095 | 45.4398 | 21.6097
2012-10-30 20:23 | 28.9499 | 55.962 | 42.144 | 59.8564 | 36.0991 | -1.4662 | 0.006729 | 0.32812 | 0.002208 | 48.7618 | 0.006961 | 0.21489 | 0.001496 | 30.8703 | 14.5042 | -0.1011 | 0.003704 | 41.4578 | 27.6398 | 45.3522 | 21.5949
2012-10-30 20:23 | 28.9521 | 55.8125 | 42.0832 | 59.7268 | 36.0658 | -1.4864 | 0.006738 | 0.32813 | 0.002211 | 48.6987 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.5105 | -0.1025 | 0.003706 | 41.302 | 27.5727 | 45.2163 | 21.5553
2012-10-30 20:23 | 28.9416 | 55.7242 | 42.0098 | 59.5031 | 35.9646 | -1.4691 | 0.006747 | 0.3281 | 0.002214 | 48.6295 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.4714 | -0.1013 | 0.003709 | 41.2528 | 27.5384 | 45.0317 | 21.4932
2012-10-30 20:23 | 28.9476 | 55.6446 | 41.9331 | 59.4626 | 35.9193 | -1.4697 | 0.006754 | 0.32806 | 0.002216 | 48.5727 | 0.006952 | 0.21482 | 0.001493 | 30.8997 | 14.5223 | -0.1013 | 0.003709 | 41.1223 | 27.4108 | 44.9403 | 21.397
2012-10-30 20:23 | 28.9379 | 55.5104 | 41.8802 | 59.3443 | 35.7991 | -1.4737 | 0.006747 | 0.32808 | 0.002214 | 48.6253 | 0.006955 | 0.21483 | 0.001494 | 30.8886 | 14.4881 | -0.1016 | 0.003708 | 41.0223 | 27.3921 | 44.8562 | 21.311
2012-10-30 20:23 | 28.9416 | 55.3701 | 41.7776 | 59.2201 | 35.7712 | -1.4749 | 0.006741 | 0.328 | 0.002211 | 48.6578 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.5163 | -0.1017 | 0.003703 | 40.8538 | 27.2613 | 44.7038 | 21.2549
2012-10-30 20:23 | 28.9401 | 55.3024 | 41.7171 | 59.1139 | 35.6671 | -1.4766 | 0.006738 | 0.32806 | 0.00221 | 48.688 | 0.006968 | 0.21479 | 0.001497 | 30.8246 | 14.4985 | -0.1018 | 0.003707 | 40.8039 | 27.2186 | 44.6154 | 21.1686
2012-10-30 20:23 | 28.9386 | 55.153 | 41.629 | 58.9727 | 35.6262 | -1.4783 | 0.006738 | 0.32786 | 0.002209 | 48.6581 | 0.006955 | 0.21484 | 0.001494 | 30.8906 | 14.4888 | -0.1019 | 0.003703 | 40.6642 | 27.1402 | 44.4839 | 21.1374
2012-10-30 20:23 | 28.9386 | 55.0674 | 41.5424 | 58.9184 | 35.5827 | -1.4754 | 0.006735 | 0.32796 | 0.002209 | 48.6947 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.4807 | -0.1017 | 0.003701 | 40.5867 | 27.0617 | 44.4377 | 21.102
2012-10-30 20:23 | 28.9356 | 54.9281 | 41.4489 | 58.7719 | 35.5245 | -1.4749 | 0.006741 | 0.32803 | 0.002211 | 48.6621 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.4818 | -0.1017 | 0.003704 | 40.4463 | 26.9671 | 44.2901 | 21.0427
2012-10-30 20:23 | 28.9281 | 54.7999 | 41.3981 | 58.6286 | 35.4381 | -1.4806 | 0.006744 | 0.32797 | 0.002212 | 48.6318 | 0.006955 | 0.21486 | 0.001494 | 30.8927 | 14.4906 | -0.1021 | 0.003706 | 40.3093 | 26.9075 | 44.138 | 20.9475
2012-10-30 20:23 | 28.9296 | 54.6651 | 41.3051 | 58.4904 | 35.4159 | -1.4754 | 0.006751 | 0.32796 | 0.002214 | 48.5793 | 0.006949 | 0.21483 | 0.001493 | 30.9152 | 14.4962 | -0.1017 | 0.003707 | 40.1689 | 26.8089 | 43.9942 | 20.9197
2012-10-30 20:23 | 28.9236 | 54.5656 | 41.2558 | 58.4301 | 35.327 | -1.4887 | 0.006741 | 0.32803 | 0.002211 | 48.6621 | 0.006965 | 0.21483 | 0.001496 | 30.8442 | 14.4698 | -0.1026 | 0.003707 | 40.0958 | 26.786 | 43.9603 | 20.8572
2012-10-30 20:23 | 28.9266 | 54.5063 | 41.2116 | 58.2856 | 35.2827 | -1.4801 | 0.006732 | 0.32797 | 0.002208 | 48.7185 | 0.006952 | 0.21483 | 0.001493 | 30.9019 | 14.4728 | -0.102 | 0.003701 | 40.0335 | 26.7388 | 43.8128 | 20.8099
2012-10-30 20:23 | 28.9162 | 54.4063 | 41.1185 | 58.1782 | 35.217 | -1.4858 | 0.006741 | 0.32805 | 0.002211 | 48.6643 | 0.006952 | 0.21479 | 0.001493 | 30.8956 | 14.4705 | -0.1024 | 0.003704 | 39.9358 | 26.648 | 43.7077 | 20.7465
2012-10-30 20:24 | 28.9132 | 54.2358 | 41.0447 | 58.0975 | 35.1509 | -1.4754 | 0.006744 | 0.32805 | 0.002212 | 48.6426 | 0.006949 | 0.21483 | 0.001493 | 30.9152 | 14.4879 | -0.1017 | 0.003705 | 39.7479 | 26.5568 | 43.6096 | 20.663
2012-10-30 20:24 | 28.9109 | 54.1283 | 40.9912 | 57.9438 | 35.0854 | -1.4882 | 0.006738 | 0.32803 | 0.00221 | 48.6837 | 0.006952 | 0.21484 | 0.001494 | 30.9039 | 14.4856 | -0.1026 | 0.003704 | 39.6427 | 26.5056 | 43.4582 | 20.5998
2012-10-30 20:24 | 28.9147 | 54.0464 | 40.9123 | 57.8545 | 35.0142 | -1.4945 | 0.006754 | 0.32803 | 0.002216 | 48.5684 | 0.006952 | 0.21487 | 0.001494 | 30.9081 | 14.4976 | -0.103 | 0.00371 | 39.5488 | 26.4147 | 43.3569 | 20.5166
2012-10-30 20:24 | 28.9094 | 53.9164 | 40.8558 | 57.7249 | 34.9616 | -1.4922 | 0.006741 | 0.32799 | 0.002211 | 48.6557 | 0.006952 | 0.21482 | 0.001493 | 30.8997 | 14.5045 | -0.1029 | 0.003704 | 39.4119 | 26.3513 | 43.2204 | 20.4571
2012-10-30 20:24 | 28.9087 | 53.7676 | 40.7528 | 57.5922 | 34.9411 | -1.4795 | 0.006751 | 0.32803 | 0.002215 | 48.59 | 0.006949 | 0.21483 | 0.001493 | 30.9152 | 14.4875 | -0.102 | 0.003708 | 39.2801 | 26.2653 | 43.1047 | 20.4536
2012-10-30 20:24 | 28.9079 | 53.6343 | 40.7047 | 57.4593 | 34.8219 | -1.4893 | 0.006738 | 0.328 | 0.00221 | 48.6794 | 0.006958 | 0.2148 | 0.001495 | 30.8711 | 14.4867 | -0.1027 | 0.003705 | 39.1476 | 26.218 | 42.9726 | 20.3352
2012-10-30 20:24 | 28.9034 | 53.5752 | 40.6293 | 57.3888 | 34.778 | -1.476 | 0.006738 | 0.328 | 0.00221 | 48.6794 | 0.006965 | 0.2148 | 0.001496 | 30.8401 | 14.4822 | -0.1018 | 0.003706 | 39.093 | 26.1471 | 42.9066 | 20.2958
2012-10-30 20:24 | 28.9072 | 53.4698 | 40.5505 | 57.2721 | 34.7423 | -1.502 | 0.006744 | 0.32794 | 0.002212 | 48.6275 | 0.006968 | 0.21483 | 0.001497 | 30.8309 | 14.5227 | -0.1035 | 0.003709 | 38.9471 | 26.0278 | 42.7494 | 20.2196
2012-10-30 20:24 | 28.9034 | 53.3413 | 40.4718 | 57.1052 | 34.6635 | -1.4789 | 0.006751 | 0.328 | 0.002214 | 48.5857 | 0.006949 | 0.21486 | 0.001493 | 30.9194 | 14.5149 | -0.102 | 0.003707 | 38.8264 | 25.9569 | 42.5903 | 20.1486
2012-10-30 20:24 | 28.9004 | 53.2252 | 40.4294 | 57.0011 | 34.6329 | -1.4853 | 0.006747 | 0.32797 | 0.002213 | 48.6102 | 0.006955 | 0.21482 | 0.001494 | 30.8864 | 14.4955 | -0.1024 | 0.003707 | 38.7297 | 25.9339 | 42.5056 | 20.1374
2012-10-30 20:24 | 28.9027 | 53.13 | 40.3254 | 56.9374 | 34.5957 | -1.4974 | 0.006747 | 0.32794 | 0.002213 | 48.6059 | 0.006958 | 0.21483 | 0.001495 | 30.8753 | 14.5386 | -0.1032 | 0.003708 | 38.5914 | 25.7868 | 42.3988 | 20.0571
2012-10-30 20:24 | 28.8982 | 53.0202 | 40.2815 | 56.8202 | 34.532 | -1.4893 | 0.006747 | 0.32808 | 0.002214 | 48.6253 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.5382 | -0.1027 | 0.003709 | 38.482 | 25.7433 | 42.282 | 19.9938
2012-10-30 20:24 | 28.8967 | 52.851 | 40.2091 | 56.6788 | 34.4713 | -1.4916 | 0.006744 | 0.32796 | 0.002212 | 48.6297 | 0.006952 | 0.21482 | 0.001493 | 30.8997 | 14.5204 | -0.1028 | 0.003705 | 38.3306 | 25.6887 | 42.1584 | 19.9509
2012-10-30 20:24 | 28.8997 | 52.7721 | 40.1294 | 56.569 | 34.4032 | -1.4853 | 0.006744 | 0.32802 | 0.002212 | 48.6383 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.503 | -0.1024 | 0.003707 | 38.2691 | 25.6264 | 42.066 | 19.9002
2012-10-30 20:24 | 28.8974 | 52.6683 | 40.0877 | 56.4812 | 34.3575 | -1.4928 | 0.006738 | 0.32797 | 0.00221 | 48.6751 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.5578 | -0.1029 | 0.003702 | 38.1105 | 25.5299 | 41.9234 | 19.7997
2012-10-30 20:24 | 28.8982 | 52.517 | 39.9822 | 56.342 | 34.2794 | -1.4899 | 0.006735 | 0.32793 | 0.002209 | 48.6904 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 14.5464 | -0.1027 | 0.003702 | 37.9706 | 25.4358 | 41.7956 | 19.733
2012-10-30 20:24 | 28.8922 | 52.4758 | 39.9565 | 56.2544 | 34.2615 | -1.4876 | 0.006729 | 0.32796 | 0.002207 | 48.7381 | 0.006961 | 0.21476 | 0.001495 | 30.8516 | 14.5118 | -0.1026 | 0.003702 | 37.964 | 25.4447 | 41.7426 | 19.7497
2012-10-30 20:24 | 28.8967 | 52.3399 | 39.8901 | 56.1928 | 34.1792 | -1.4934 | 0.006744 | 0.32797 | 0.002212 | 48.6318 | 0.006946 | 0.21477 | 0.001492 | 30.9202 | 14.5286 | -0.103 | 0.003704 | 37.8113 | 25.3615 | 41.6642 | 19.6506
2012-10-30 20:24 | 28.8967 | 52.2774 | 39.7916 | 56.0644 | 34.1436 | -1.4911 | 0.006751 | 0.32796 | 0.002214 | 48.5793 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.5286 | -0.1028 | 0.003707 | 37.7488 | 25.263 | 41.5358 | 19.615
2012-10-30 20:24 | 28.8877 | 52.1357 | 39.7432 | 55.9426 | 34.0241 | -1.4841 | 0.006735 | 0.32797 | 0.002209 | 48.6968 | 0.006955 | 0.21487 | 0.001494 | 30.8948 | 14.5236 | -0.1023 | 0.003703 | 37.6121 | 25.2196 | 41.419 | 19.5005
2012-10-30 20:24 | 28.8892 | 52.0591 | 39.6383 | 55.8508 | 33.9901 | -1.487 | 0.006741 | 0.32797 | 0.002211 | 48.6535 | 0.006955 | 0.21484 | 0.001494 | 30.8906 | 14.517 | -0.1025 | 0.003705 | 37.5421 | 25.1213 | 41.3338 | 19.4731
2012-10-30 20:24 | 28.8937 | 51.9426 | 39.6034 | 55.7541 | 34.0025 | -1.4905 | 0.006725 | 0.32796 | 0.002206 | 48.7671 | 0.006958 | 0.21482 | 0.001495 | 30.8731 | 14.4929 | -0.1028 | 0.003701 | 37.4497 | 25.1105 | 41.2612 | 19.5096
2012-10-30 20:24 | 28.8899 | 51.8412 | 39.5208 | 55.6103 | 33.8842 | -1.4922 | 0.006751 | 0.328 | 0.002214 | 48.5857 | 0.006955 | 0.2148 | 0.001494 | 30.8844 | 14.4932 | -0.1029 | 0.003708 | 37.348 | 25.0276 | 41.1171 | 19.391
2012-10-30 20:24 | 28.8869 | 51.7172 | 39.4706 | 55.5061 | 33.8338 | -1.502 | 0.006741 | 0.328 | 0.002211 | 48.6578 | 0.006961 | 0.21477 | 0.001495 | 30.8536 | 14.5106 | -0.1035 | 0.003706 | 37.2066 | 24.96 | 40.9955 | 19.3232
2012-10-30 20:25 | 28.8862 | 51.6227 | 39.3911 | 55.4354 | 33.7621 | -1.4986 | 0.006732 | 0.32805 | 0.002208 | 48.7294 | 0.006961 | 0.2148 | 0.001495 | 30.8578 | 14.5221 | -0.1033 | 0.003703 | 37.1006 | 24.869 | 40.9133 | 19.24
2012-10-30 20:25 | 28.8824 | 51.5408 | 39.3676 | 55.3382 | 33.6754 | -1.4911 | 0.006741 | 0.3279 | 0.00221 | 48.6428 | 0.006955 | 0.21482 | 0.001494 | 30.8864 | 14.4979 | -0.1028 | 0.003704 | 37.0429 | 24.8697 | 40.8403 | 19.1775
2012-10-30 20:25 | 28.8959 | 51.4645 | 39.3298 | 55.2115 | 33.681 | -1.5038 | 0.006744 | 0.32806 | 0.002212 | 48.6447 | 0.006946 | 0.21487 | 0.001493 | 30.9348 | 14.5033 | -0.1037 | 0.003705 | 36.9612 | 24.8265 | 40.7082 | 19.1777
2012-10-30 20:25 | 28.8869 | 51.3501 | 39.2066 | 55.1325 | 33.5614 | -1.4945 | 0.006763 | 0.32799 | 0.002218 | 48.4974 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.4861 | -0.103 | 0.003713 | 36.864 | 24.7205 | 40.6464 | 19.0753
2012-10-30 20:25 | 28.8817 | 51.1925 | 39.1777 | 55.0455 | 33.5365 | -1.5026 | 0.006732 | 0.32802 | 0.002208 | 48.725 | 0.006958 | 0.2148 | 0.001495 | 30.8711 | 14.4809 | -0.1036 | 0.003703 | 36.7116 | 24.6968 | 40.5646 | 19.0556
2012-10-30 20:25 | 28.8772 | 51.1372 | 39.1141 | 54.897 | 33.4925 | -1.4991 | 0.006747 | 0.32797 | 0.002213 | 48.6102 | 0.006958 | 0.21483 | 0.001495 | 30.8753 | 14.5335 | -0.1033 | 0.003708 | 36.6037 | 24.5806 | 40.3635 | 18.959
2012-10-30 20:25 | 28.8817 | 51.0167 | 39.0477 | 54.8392 | 33.4773 | -1.487 | 0.006747 | 0.32796 | 0.002213 | 48.6081 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.5258 | -0.1025 | 0.003708 | 36.4909 | 24.5219 | 40.3134 | 18.9515
2012-10-30 20:25 | 28.8742 | 50.9545 | 38.9456 | 54.7149 | 33.3947 | -1.4934 | 0.006744 | 0.32794 | 0.002212 | 48.6275 | 0.006952 | 0.21483 | 0.001493 | 30.9019 | 14.5224 | -0.103 | 0.003705 | 36.4321 | 24.4232 | 40.1925 | 18.8723
2012-10-30 20:25 | 28.8735 | 50.8717 | 38.9094 | 54.6285 | 33.3348 | -1.4887 | 0.006747 | 0.32799 | 0.002213 | 48.6124 | 0.006952 | 0.21486 | 0.001494 | 30.9061 | 14.5462 | -0.1026 | 0.003707 | 36.3255 | 24.3632 | 40.0823 | 18.7886
2012-10-30 20:25 | 28.8802 | 50.7807 | 38.8452 | 54.5456 | 33.3494 | -1.4997 | 0.006738 | 0.32796 | 0.00221 | 48.673 | 0.006965 | 0.21479 | 0.001496 | 30.8379 | 14.4835 | -0.1034 | 0.003706 | 36.2972 | 24.3617 | 40.0621 | 18.8659
2012-10-30 20:25 | 28.866 | 50.6571 | 38.7994 | 54.4379 | 33.2957 | -1.4986 | 0.006757 | 0.32794 | 0.002216 | 48.534 | 0.006955 | 0.21486 | 0.001494 | 30.8927 | 14.5019 | -0.1033 | 0.00371 | 36.1552 | 24.2975 | 39.936 | 18.7938
2012-10-30 20:25 | 28.8652 | 50.5664 | 38.7316 | 54.3437 | 33.1923 | -1.487 | 0.006741 | 0.32793 | 0.002211 | 48.6471 | 0.006961 | 0.21479 | 0.001495 | 30.8556 | 14.5011 | -0.1025 | 0.003706 | 36.0653 | 24.2305 | 39.8426 | 18.6912
2012-10-30 20:25 | 28.8727 | 50.4918 | 38.6642 | 54.2343 | 33.1958 | -1.5055 | 0.006738 | 0.32787 | 0.002209 | 48.6601 | 0.006955 | 0.21479 | 0.001494 | 30.8822 | 14.5331 | -0.1038 | 0.003703 | 35.9587 | 24.1311 | 39.7012 | 18.6627
2012-10-30 20:25 | 28.8735 | 50.3832 | 38.598 | 54.1455 | 33.1216 | -1.4991 | 0.006754 | 0.32787 | 0.002214 | 48.5449 | 0.006965 | 0.21477 | 0.001496 | 30.8359 | 14.4809 | -0.1033 | 0.00371 | 35.9023 | 24.1171 | 39.6646 | 18.6407
2012-10-30 20:25 | 28.866 | 50.2545 | 38.5038 | 54.0172 | 33.0944 | -1.4928 | 0.006757 | 0.32789 | 0.002216 | 48.5255 | 0.006961 | 0.21483 | 0.001495 | 30.8619 | 14.4979 | -0.1029 | 0.003711 | 35.7566 | 24.0059 | 39.5193 | 18.5965
2012-10-30 20:25 | 28.872 | 50.1627 | 38.4704 | 53.918 | 33.0135 | -1.4928 | 0.006744 | 0.32792 | 0.002211 | 48.6234 | 0.006961 | 0.21474 | 0.001495 | 30.8494 | 14.5161 | -0.1029 | 0.003706 | 35.6466 | 23.9543 | 39.4019 | 18.4974
2012-10-30 20:25 | 28.86 | 50.0686 | 38.3796 | 53.828 | 32.9502 | -1.4939 | 0.006747 | 0.32792 | 0.002212 | 48.6017 | 0.006961 | 0.21482 | 0.001495 | 30.8598 | 14.4674 | -0.103 | 0.003707 | 35.6012 | 23.9122 | 39.3606 | 18.4828
2012-10-30 20:25 | 28.866 | 49.9612 | 38.3264 | 53.7405 | 32.9128 | -1.4922 | 0.006751 | 0.32794 | 0.002214 | 48.5771 | 0.006955 | 0.21479 | 0.001494 | 30.8822 | 14.4652 | -0.1029 | 0.003708 | 35.496 | 23.8612 | 39.2753 | 18.4476
2012-10-30 20:25 | 28.8615 | 49.8589 | 38.2549 | 53.6074 | 32.8254 | -1.4997 | 0.00676 | 0.32792 | 0.002217 | 48.5083 | 0.006949 | 0.21483 | 0.001493 | 30.9152 | 14.473 | -0.1034 | 0.00371 | 35.3859 | 23.7819 | 39.1344 | 18.3524
2012-10-30 20:25 | 28.8607 | 49.776 | 38.1989 | 53.5208 | 32.7811 | -1.4928 | 0.006751 | 0.32787 | 0.002213 | 48.5664 | 0.006965 | 0.21471 | 0.001495 | 30.8276 | 14.513 | -0.1029 | 0.003708 | 35.263 | 23.6859 | 39.0078 | 18.2681
2012-10-30 20:25 | 28.8645 | 49.6585 | 38.183 | 53.4155 | 32.7455 | -1.4916 | 0.006735 | 0.32779 | 0.002208 | 48.669 | 0.006958 | 0.21476 | 0.001494 | 30.8649 | 14.4923 | -0.1028 | 0.003702 | 35.1662 | 23.6907 | 38.9232 | 18.2532
2012-10-30 20:25 | 28.8682 | 49.6114 | 38.0881 | 53.3451 | 32.6899 | -1.4991 | 0.006735 | 0.32783 | 0.002208 | 48.6754 | 0.006961 | 0.21474 | 0.001495 | 30.8494 | 14.4878 | -0.1033 | 0.003703 | 35.1236 | 23.6003 | 38.8573 | 18.2021
2012-10-30 20:25 | 28.8577 | 49.464 | 38.0106 | 53.1942 | 32.6597 | -1.4922 | 0.006747 | 0.32779 | 0.002212 | 48.5825 | 0.006955 | 0.21474 | 0.001494 | 30.8761 | 14.4692 | -0.1029 | 0.003706 | 34.9948 | 23.5414 | 38.725 | 18.1905
2012-10-30 20:26 | 28.8592 | 49.3324 | 37.9451 | 53.0904 | 32.5664 | -1.4853 | 0.006747 | 0.32787 | 0.002212 | 48.5952 | 0.006965 | 0.21476 | 0.001496 | 30.8339 | 14.4625 | -0.1024 | 0.003708 | 34.8699 | 23.4826 | 38.6279 | 18.1039
2012-10-30 20:26 | 28.8532 | 49.2873 | 37.8681 | 52.983 | 32.5525 | -1.4899 | 0.006732 | 0.32786 | 0.002207 | 48.7014 | 0.006968 | 0.21477 | 0.001497 | 30.8226 | 14.4565 | -0.1027 | 0.003704 | 34.8308 | 23.4116 | 38.5265 | 18.096
2012-10-30 20:26 | 28.8592 | 49.2229 | 37.8031 | 52.9149 | 32.4954 | -1.4957 | 0.006741 | 0.32787 | 0.00221 | 48.6385 | 0.006949 | 0.21489 | 0.001493 | 30.9236 | 14.4951 | -0.1031 | 0.003703 | 34.7278 | 23.308 | 38.4198 | 18.0003
2012-10-30 20:26 | 28.848 | 49.059 | 37.721 | 52.7789 | 32.417 | -1.4876 | 0.006744 | 0.32781 | 0.002211 | 48.6082 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.4595 | -0.1026 | 0.003703 | 34.5995 | 23.2615 | 38.3194 | 17.9575
2012-10-30 20:26 | 28.8457 | 48.9824 | 37.7187 | 52.6986 | 32.3595 | -1.4945 | 0.006735 | 0.32783 | 0.002208 | 48.6754 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 14.4327 | -0.103 | 0.0037 | 34.5497 | 23.286 | 38.2659 | 17.9268
2012-10-30 20:26 | 28.8427 | 48.8737 | 37.6329 | 52.6136 | 32.3091 | -1.4963 | 0.006738 | 0.32773 | 0.002208 | 48.6388 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.4256 | -0.1032 | 0.003701 | 34.4481 | 23.2073 | 38.188 | 17.8835
2012-10-30 20:26 | 28.8375 | 48.7785 | 37.6395 | 52.5499 | 32.2407 | -1.4997 | 0.006729 | 0.32777 | 0.002206 | 48.7102 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.4122 | -0.1034 | 0.003696 | 34.3663 | 23.2273 | 38.1377 | 17.8285
2012-10-30 20:26 | 28.842 | 48.7321 | 37.5493 | 52.4334 | 32.2215 | -1.5078 | 0.006751 | 0.32777 | 0.002213 | 48.5515 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4575 | -0.1039 | 0.003704 | 34.2746 | 23.0918 | 37.9759 | 17.764
2012-10-30 20:26 | 28.8368 | 48.6212 | 37.4495 | 52.3268 | 32.1531 | -1.4939 | 0.006751 | 0.32774 | 0.002213 | 48.5472 | 0.006955 | 0.21474 | 0.001494 | 30.8761 | 14.4197 | -0.103 | 0.003707 | 34.2015 | 23.0298 | 37.9071 | 17.7334
2012-10-30 20:26 | 28.8375 | 48.5279 | 37.3477 | 52.2026 | 32.0946 | -1.5038 | 0.006741 | 0.32777 | 0.00221 | 48.6235 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 14.4367 | -0.1037 | 0.003702 | 34.0912 | 22.911 | 37.7659 | 17.6579
2012-10-30 20:26 | 28.8285 | 48.4367 | 37.3466 | 52.1624 | 32.0698 | -1.4841 | 0.006725 | 0.32781 | 0.002205 | 48.7456 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.3991 | -0.1023 | 0.003697 | 34.0376 | 22.9475 | 37.7633 | 17.6707
2012-10-30 20:26 | 28.8375 | 48.3752 | 37.2609 | 52.0387 | 31.9841 | -1.4766 | 0.006754 | 0.32784 | 0.002214 | 48.5406 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.4449 | -0.1018 | 0.003704 | 33.9303 | 22.816 | 37.5938 | 17.5392
2012-10-30 20:26 | 28.8308 | 48.2471 | 37.2227 | 51.95 | 31.9142 | -1.5009 | 0.006735 | 0.32781 | 0.002208 | 48.6732 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.4545 | -0.1035 | 0.0037 | 33.7926 | 22.7682 | 37.4955 | 17.4597
2012-10-30 20:26 | 28.827 | 48.1415 | 37.1439 | 51.8759 | 31.8472 | -1.4928 | 0.006738 | 0.32777 | 0.002209 | 48.6451 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.4303 | -0.1029 | 0.003702 | 33.7112 | 22.7136 | 37.4456 | 17.4169
2012-10-30 20:26 | 28.8368 | 48.0964 | 37.0906 | 51.7608 | 31.8136 | -1.4945 | 0.006754 | 0.32783 | 0.002214 | 48.5385 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.4483 | -0.103 | 0.003707 | 33.6481 | 22.6423 | 37.3125 | 17.3653
2012-10-30 20:26 | 28.8382 | 47.9999 | 37.025 | 51.6568 | 31.7794 | -1.4841 | 0.006747 | 0.32777 | 0.002211 | 48.5803 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.4415 | -0.1023 | 0.003703 | 33.5584 | 22.5835 | 37.2153 | 17.3379
2012-10-30 20:26 | 28.8338 | 47.9093 | 36.9969 | 51.586 | 31.7987 | -1.4876 | 0.006747 | 0.32779 | 0.002212 | 48.5825 | 0.006952 | 0.21486 | 0.001494 | 30.9061 | 14.4616 | -0.1026 | 0.003706 | 33.4477 | 22.5353 | 37.1244 | 17.3371
2012-10-30 20:26 | 28.8308 | 47.7966 | 36.9308 | 51.4971 | 31.697 | -1.4731 | 0.00676 | 0.32781 | 0.002216 | 48.4932 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5076 | -0.1016 | 0.003706 | 33.289 | 22.4232 | 36.9895 | 17.1894
2012-10-30 20:26 | 28.8323 | 47.5944 | 36.7982 | 51.3189 | 31.588 | -1.4882 | 0.006741 | 0.32783 | 0.00221 | 48.6321 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 14.456 | -0.1026 | 0.003703 | 33.1384 | 22.3422 | 36.8629 | 17.132
2012-10-30 20:26 | 28.8323 | 47.4102 | 36.7074 | 51.0884 | 31.4695 | -1.4899 | 0.006741 | 0.32781 | 0.00221 | 48.6299 | 0.006958 | 0.2148 | 0.001495 | 30.8711 | 14.4519 | -0.1027 | 0.003705 | 32.9583 | 22.2555 | 36.6365 | 17.0176
2012-10-30 20:26 | 28.8338 | 47.2353 | 36.5669 | 50.9297 | 31.4315 | -1.4743 | 0.006747 | 0.32781 | 0.002212 | 48.5866 | 0.006955 | 0.21477 | 0.001494 | 30.8802 | 14.4167 | -0.1016 | 0.003706 | 32.8186 | 22.1502 | 36.513 | 17.0148
2012-10-30 20:27 | 28.842 | 47.0554 | 36.4647 | 50.7347 | 31.3094 | -1.4737 | 0.006751 | 0.32781 | 0.002213 | 48.5578 | 0.006955 | 0.2148 | 0.001494 | 30.8844 | 14.4412 | -0.1016 | 0.003707 | 32.6142 | 22.0235 | 36.2935 | 16.8682
2012-10-30 20:27 | 28.8382 | 46.9536 | 36.4175 | 50.6137 | 31.2701 | -1.4812 | 0.006732 | 0.32776 | 0.002206 | 48.6864 | 0.006952 | 0.2148 | 0.001493 | 30.8977 | 14.4333 | -0.1021 | 0.003699 | 32.5203 | 21.9842 | 36.1804 | 16.8368
2012-10-30 20:27 | 28.8293 | 46.831 | 36.3178 | 50.5501 | 31.2257 | -1.4737 | 0.006729 | 0.32774 | 0.002205 | 48.7059 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.4489 | -0.1016 | 0.003698 | 32.3821 | 21.8689 | 36.1012 | 16.7768
2012-10-30 20:27 | 28.8315 | 46.7469 | 36.3319 | 50.4115 | 31.0936 | -1.4812 | 0.006729 | 0.32774 | 0.002205 | 48.7059 | 0.006958 | 0.21484 | 0.001495 | 30.8773 | 14.4389 | -0.1021 | 0.0037 | 32.308 | 21.893 | 35.9726 | 16.6547
2012-10-30 20:27 | 28.8353 | 46.6958 | 36.2923 | 50.4036 | 31.1093 | -1.4853 | 0.006754 | 0.32776 | 0.002214 | 48.5278 | 0.006958 | 0.21476 | 0.001494 | 30.8649 | 14.3978 | -0.1024 | 0.003708 | 32.298 | 21.8945 | 36.0058 | 16.7115
2012-10-30 20:27 | 28.8323 | 46.5948 | 36.1946 | 50.2951 | 31.0549 | -1.4835 | 0.00676 | 0.32777 | 0.002216 | 48.4868 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.4274 | -0.1023 | 0.003707 | 32.1674 | 21.7672 | 35.8677 | 16.6275
2012-10-30 20:27 | 28.836 | 46.4769 | 36.1352 | 50.1541 | 30.9994 | -1.4812 | 0.006747 | 0.32776 | 0.002211 | 48.5782 | 0.006955 | 0.21477 | 0.001494 | 30.8802 | 14.3781 | -0.1021 | 0.003705 | 32.0988 | 21.7571 | 35.776 | 16.6213
2012-10-30 20:27 | 28.8435 | 46.4609 | 36.0874 | 50.0873 | 30.999 | -1.4743 | 0.006747 | 0.32774 | 0.002211 | 48.576 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.4386 | -0.1016 | 0.003705 | 32.0223 | 21.6488 | 35.6487 | 16.5604
2012-10-30 20:27 | 28.8405 | 46.3245 | 36.0686 | 50.0061 | 30.9447 | -1.4645 | 0.006754 | 0.32783 | 0.002214 | 48.5385 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.4316 | -0.101 | 0.003706 | 31.8929 | 21.637 | 35.5745 | 16.5131
2012-10-30 20:27 | 28.83 | 46.2827 | 35.9713 | 49.933 | 30.8947 | -1.4824 | 0.006738 | 0.3278 | 0.002209 | 48.6495 | 0.006968 | 0.21483 | 0.001497 | 30.8309 | 14.417 | -0.1022 | 0.003706 | 31.8657 | 21.5543 | 35.516 | 16.4777
2012-10-30 20:27 | 28.8412 | 46.2154 | 35.9273 | 49.8543 | 30.8427 | -1.4824 | 0.006747 | 0.32781 | 0.002212 | 48.5866 | 0.006965 | 0.2148 | 0.001496 | 30.8401 | 14.4241 | -0.1022 | 0.003708 | 31.7913 | 21.5032 | 35.4302 | 16.4186
2012-10-30 20:27 | 28.839 | 46.1112 | 35.862 | 49.7504 | 30.7932 | -1.4864 | 0.006741 | 0.32771 | 0.002209 | 48.6149 | 0.006958 | 0.21477 | 0.001494 | 30.8669 | 14.3974 | -0.1025 | 0.003703 | 31.7138 | 21.4646 | 35.353 | 16.3958
2012-10-30 20:27 | 28.8397 | 46.0374 | 35.8193 | 49.665 | 30.786 | -1.4772 | 0.006744 | 0.32776 | 0.00221 | 48.5998 | 0.006958 | 0.21477 | 0.001494 | 30.8669 | 14.3818 | -0.1018 | 0.003704 | 31.6556 | 21.4375 | 35.2832 | 16.4042
2012-10-30 20:27 | 28.8315 | 45.9311 | 35.74 | 49.5551 | 30.7264 | -1.4783 | 0.006744 | 0.32783 | 0.002211 | 48.6105 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.4185 | -0.1019 | 0.003704 | 31.5126 | 21.3215 | 35.1366 | 16.3079
2012-10-30 20:27 | 28.8412 | 45.8623 | 35.6945 | 49.5453 | 30.7519 | -1.4766 | 0.006766 | 0.3277 | 0.002217 | 48.4332 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.4159 | -0.1018 | 0.00371 | 31.4464 | 21.2786 | 35.1294 | 16.336
2012-10-30 20:27 | 28.8375 | 45.7801 | 35.6513 | 49.4085 | 30.7008 | -1.4702 | 0.006747 | 0.32773 | 0.002211 | 48.5739 | 0.006958 | 0.21477 | 0.001494 | 30.8669 | 14.4 | -0.1014 | 0.003705 | 31.3801 | 21.2513 | 35.0085 | 16.3008
2012-10-30 20:27 | 28.8353 | 45.7034 | 35.6136 | 49.332 | 30.596 | -1.4662 | 0.006738 | 0.32774 | 0.002208 | 48.6408 | 0.006955 | 0.21479 | 0.001494 | 30.8822 | 14.4059 | -0.1011 | 0.003702 | 31.2975 | 21.2077 | 34.9261 | 16.1901
2012-10-30 20:27 | 28.8338 | 45.6352 | 35.5253 | 49.2679 | 30.5471 | -1.4662 | 0.006754 | 0.3277 | 0.002213 | 48.5192 | 0.006958 | 0.21474 | 0.001494 | 30.8627 | 14.4085 | -0.1011 | 0.003707 | 31.2267 | 21.1168 | 34.8594 | 16.1386
2012-10-30 20:27 | 28.8427 | 45.5145 | 35.5224 | 49.1751 | 30.5481 | -1.4841 | 0.006747 | 0.32779 | 0.002212 | 48.5825 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.4133 | -0.1023 | 0.003704 | 31.1012 | 21.1091 | 34.7618 | 16.1348
2012-10-30 20:27 | 28.833 | 45.4734 | 35.4495 | 49.0753 | 30.495 | -1.4818 | 0.006754 | 0.3277 | 0.002213 | 48.5192 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.42 | -0.1022 | 0.003706 | 31.0534 | 21.0295 | 34.6553 | 16.075
2012-10-30 20:27 | 28.8315 | 45.3777 | 35.4362 | 48.9995 | 30.4343 | -1.483 | 0.006741 | 0.32773 | 0.002209 | 48.6171 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4021 | -0.1022 | 0.003699 | 30.9756 | 21.0341 | 34.5974 | 16.0322
2012-10-30 20:27 | 28.8382 | 45.2667 | 35.4034 | 48.9201 | 30.3896 | -1.4853 | 0.006747 | 0.32777 | 0.002211 | 48.5803 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.4088 | -0.1024 | 0.003704 | 30.8579 | 20.9946 | 34.5113 | 15.9808
2012-10-30 20:27 | 28.8442 | 45.1745 | 35.3029 | 48.8321 | 30.3088 | -1.4864 | 0.006747 | 0.32771 | 0.002211 | 48.5717 | 0.006955 | 0.21479 | 0.001494 | 30.8822 | 14.4312 | -0.1025 | 0.003705 | 30.7433 | 20.8717 | 34.4009 | 15.8776
2012-10-30 20:28 | 28.8442 | 45.0725 | 35.2674 | 48.7069 | 30.2101 | -1.4824 | 0.006747 | 0.32771 | 0.002211 | 48.5717 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.4189 | -0.1022 | 0.003705 | 30.6536 | 20.8485 | 34.288 | 15.7912
2012-10-30 20:28 | 28.842 | 44.9878 | 35.1783 | 48.5833 | 30.2355 | -1.4974 | 0.006732 | 0.32774 | 0.002206 | 48.6842 | 0.006965 | 0.21474 | 0.001496 | 30.8317 | 14.4004 | -0.1032 | 0.003702 | 30.5874 | 20.7779 | 34.1829 | 15.8351
2012-10-30 20:28 | 28.842 | 44.925 | 35.1783 | 48.5363 | 30.1605 | -1.4864 | 0.006747 | 0.32774 | 0.002211 | 48.576 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.4331 | -0.1025 | 0.003704 | 30.4919 | 20.7452 | 34.1032 | 15.7274
2012-10-30 20:28 | 28.848 | 44.8407 | 35.062 | 48.4601 | 30.1783 | -1.483 | 0.006754 | 0.32774 | 0.002214 | 48.5256 | 0.006958 | 0.21482 | 0.001495 | 30.8731 | 14.4431 | -0.1022 | 0.003709 | 30.3976 | 20.6189 | 34.017 | 15.7352
2012-10-30 20:28 | 28.8555 | 44.7579 | 35.0576 | 48.3932 | 30.1819 | -1.4899 | 0.006754 | 0.32774 | 0.002214 | 48.5256 | 0.006977 | 0.21477 | 0.001498 | 30.7829 | 14.4506 | -0.1027 | 0.003712 | 30.3073 | 20.607 | 33.9426 | 15.7313
2012-10-30 20:28 | 28.8487 | 44.6686 | 35.0232 | 48.312 | 30.0685 | -1.483 | 0.006747 | 0.32777 | 0.002211 | 48.5803 | 0.006955 | 0.21477 | 0.001494 | 30.8802 | 14.3949 | -0.1022 | 0.003705 | 30.2737 | 20.6283 | 33.9171 | 15.6736
2012-10-30 20:28 | 28.8592 | 44.573 | 34.9942 | 48.2167 | 30.0395 | -1.4922 | 0.006741 | 0.32766 | 0.002209 | 48.6064 | 0.006955 | 0.2148 | 0.001494 | 30.8844 | 14.4298 | -0.1029 | 0.003703 | 30.1432 | 20.5644 | 33.7869 | 15.6097
2012-10-30 20:28 | 28.8517 | 44.4948 | 34.9512 | 48.1427 | 30.0004 | -1.4945 | 0.006741 | 0.32767 | 0.002209 | 48.6085 | 0.006955 | 0.21473 | 0.001493 | 30.874 | 14.4387 | -0.103 | 0.003702 | 30.0561 | 20.5125 | 33.704 | 15.5617
2012-10-30 20:28 | 28.857 | 44.4098 | 34.8894 | 48.0696 | 29.9189 | -1.483 | 0.006754 | 0.32774 | 0.002214 | 48.5256 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.4889 | -0.1022 | 0.003704 | 29.9209 | 20.4005 | 33.5807 | 15.43
2012-10-30 20:28 | 28.8457 | 44.3985 | 34.8505 | 47.9604 | 29.8602 | -1.4922 | 0.006757 | 0.32771 | 0.002214 | 48.4998 | 0.006961 | 0.2148 | 0.001495 | 30.8578 | 14.4408 | -0.1029 | 0.003709 | 29.9577 | 20.4097 | 33.5196 | 15.4194
2012-10-30 20:28 | 28.8577 | 44.2848 | 34.8467 | 47.9215 | 29.8525 | -1.4754 | 0.006754 | 0.3277 | 0.002213 | 48.5192 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.4651 | -0.1017 | 0.003705 | 29.8197 | 20.3816 | 33.4564 | 15.3874
2012-10-30 20:28 | 28.8577 | 44.2023 | 34.7204 | 47.8079 | 29.8011 | -1.4887 | 0.006741 | 0.3277 | 0.002209 | 48.6128 | 0.006955 | 0.2148 | 0.001494 | 30.8844 | 14.4732 | -0.1026 | 0.003703 | 29.7291 | 20.2472 | 33.3347 | 15.3279
2012-10-30 20:28 | 28.854 | 44.124 | 34.697 | 47.7258 | 29.758 | -1.4882 | 0.006747 | 0.32773 | 0.002211 | 48.5739 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.4777 | -0.1026 | 0.003704 | 29.6463 | 20.2193 | 33.2481 | 15.2803
2012-10-30 20:28 | 28.8525 | 44.0517 | 34.656 | 47.6616 | 29.7407 | -1.4853 | 0.006744 | 0.32764 | 0.00221 | 48.5826 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.4558 | -0.1024 | 0.003702 | 29.5959 | 20.2002 | 33.2058 | 15.2849
2012-10-30 20:28 | 28.8532 | 43.9817 | 34.5936 | 47.5761 | 29.6703 | -1.4945 | 0.006766 | 0.32763 | 0.002217 | 48.4226 | 0.006955 | 0.21474 | 0.001494 | 30.8761 | 14.5259 | -0.103 | 0.003711 | 29.4558 | 20.0677 | 33.0502 | 15.1444
2012-10-30 20:28 | 28.8555 | 43.8622 | 34.5564 | 47.5 | 29.5937 | -1.4818 | 0.006751 | 0.32776 | 0.002213 | 48.5494 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 14.4588 | -0.1022 | 0.003706 | 29.4034 | 20.0976 | 33.0412 | 15.1349
2012-10-30 20:28 | 28.8585 | 43.818 | 34.4804 | 47.4247 | 29.6164 | -1.4934 | 0.006747 | 0.32773 | 0.002211 | 48.5739 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.4985 | -0.103 | 0.003703 | 29.3195 | 19.9819 | 32.9262 | 15.1179
2012-10-30 20:28 | 28.8532 | 43.7184 | 34.4515 | 47.341 | 29.5993 | -1.483 | 0.006757 | 0.32774 | 0.002215 | 48.5041 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4728 | -0.1022 | 0.003707 | 29.2456 | 19.9787 | 32.8682 | 15.1265
2012-10-30 20:28 | 28.8547 | 43.6531 | 34.4332 | 47.2915 | 29.5455 | -1.4876 | 0.006747 | 0.32771 | 0.002211 | 48.5717 | 0.006955 | 0.21479 | 0.001494 | 30.8822 | 14.4825 | -0.1026 | 0.003705 | 29.1706 | 19.9507 | 32.809 | 15.063
2012-10-30 20:28 | 28.8562 | 43.5799 | 34.3716 | 47.1872 | 29.5312 | -1.4893 | 0.006741 | 0.32773 | 0.002209 | 48.6171 | 0.006961 | 0.21477 | 0.001495 | 30.8536 | 14.5452 | -0.1027 | 0.003704 | 29.0347 | 19.8264 | 32.642 | 14.986
2012-10-30 20:28 | 28.8495 | 43.4986 | 34.2978 | 47.1256 | 29.4219 | -1.483 | 0.006747 | 0.32776 | 0.002211 | 48.5782 | 0.006955 | 0.21482 | 0.001494 | 30.8864 | 14.5263 | -0.1022 | 0.003705 | 28.9723 | 19.7715 | 32.5993 | 14.8956
2012-10-30 20:28 | 28.851 | 43.4097 | 34.2795 | 47.0448 | 29.4392 | -1.4963 | 0.006741 | 0.3277 | 0.002209 | 48.6128 | 0.006968 | 0.21482 | 0.001497 | 30.8288 | 14.5073 | -0.1032 | 0.003706 | 28.9024 | 19.7722 | 32.5375 | 14.9319
2012-10-30 20:28 | 28.8517 | 43.3357 | 34.2329 | 46.9632 | 29.3806 | -1.4789 | 0.006757 | 0.32773 | 0.002214 | 48.502 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.5407 | -0.102 | 0.003704 | 28.795 | 19.6922 | 32.4225 | 14.8399
2012-10-30 20:28 | 28.8457 | 43.2904 | 34.1716 | 46.8278 | 29.3588 | -1.4801 | 0.006744 | 0.32767 | 0.00221 | 48.5869 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.5388 | -0.102 | 0.003703 | 28.7516 | 19.6328 | 32.289 | 14.82
2012-10-30 20:28 | 28.8555 | 43.194 | 34.134 | 46.818 | 29.3213 | -1.4835 | 0.006751 | 0.32768 | 0.002212 | 48.5386 | 0.006965 | 0.21474 | 0.001496 | 30.8317 | 14.5241 | -0.1023 | 0.003708 | 28.6699 | 19.6099 | 32.2939 | 14.7972
2012-10-30 20:29 | 28.8525 | 43.0967 | 34.0955 | 46.7248 | 29.3183 | -1.4754 | 0.006757 | 0.32773 | 0.002214 | 48.502 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.5129 | -0.1017 | 0.003708 | 28.5838 | 19.5826 | 32.2119 | 14.8054
2012-10-30 20:29 | 28.8577 | 43.0586 | 34.081 | 46.6555 | 29.2564 | -1.4847 | 0.006763 | 0.32771 | 0.002216 | 48.4567 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.4814 | -0.1024 | 0.003707 | 28.5772 | 19.5996 | 32.1741 | 14.775
2012-10-30 20:29 | 28.854 | 42.9999 | 33.9825 | 46.6008 | 29.2132 | -1.4858 | 0.006747 | 0.32773 | 0.002211 | 48.5739 | 0.006958 | 0.21479 | 0.001494 | 30.8689 | 14.5185 | -0.1024 | 0.003705 | 28.4814 | 19.464 | 32.0823 | 14.6947
2012-10-30 20:29 | 28.848 | 42.8681 | 33.9647 | 46.536 | 29.1875 | -1.4858 | 0.006741 | 0.32763 | 0.002209 | 48.6021 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.5166 | -0.1024 | 0.003701 | 28.3515 | 19.4481 | 32.0194 | 14.6709
2012-10-30 20:29 | 28.8525 | 42.8254 | 33.8863 | 46.4424 | 29.1683 | -1.4951 | 0.006747 | 0.32764 | 0.002211 | 48.561 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5619 | -0.1031 | 0.003703 | 28.2635 | 19.3244 | 31.8805 | 14.6064
2012-10-30 20:29 | 28.854 | 42.7246 | 33.872 | 46.3694 | 29.1382 | -1.4887 | 0.006751 | 0.32753 | 0.002211 | 48.515 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.5103 | -0.1026 | 0.003705 | 28.2143 | 19.3617 | 31.8591 | 14.6279
2012-10-30 20:29 | 28.8465 | 42.6857 | 33.8526 | 46.2874 | 29.0833 | -1.483 | 0.006754 | 0.32763 | 0.002213 | 48.5086 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.511 | -0.1022 | 0.003704 | 28.1747 | 19.3416 | 31.7764 | 14.5723
2012-10-30 20:29 | 28.8427 | 42.5835 | 33.7975 | 46.2051 | 29.0914 | -1.4939 | 0.006744 | 0.32763 | 0.00221 | 48.5805 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5358 | -0.103 | 0.003701 | 28.0477 | 19.2617 | 31.6693 | 14.5556
2012-10-30 20:29 | 28.842 | 42.5514 | 33.7771 | 46.1613 | 29.0591 | -1.4812 | 0.006735 | 0.32761 | 0.002206 | 48.6432 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5228 | -0.1021 | 0.003695 | 28.0286 | 19.2543 | 31.6385 | 14.5363
2012-10-30 20:29 | 28.8427 | 42.4498 | 33.6751 | 46.0835 | 29.0203 | -1.4939 | 0.006744 | 0.32754 | 0.002209 | 48.5676 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5439 | -0.103 | 0.003701 | 27.9059 | 19.1312 | 31.5396 | 14.4764
2012-10-30 20:29 | 28.8427 | 42.3948 | 33.6278 | 45.9776 | 28.9808 | -1.4795 | 0.006738 | 0.32754 | 0.002207 | 48.6109 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5562 | -0.102 | 0.003696 | 27.8386 | 19.0716 | 31.4214 | 14.4246
2012-10-30 20:29 | 28.8405 | 42.3178 | 33.6769 | 45.9322 | 28.9668 | -1.4916 | 0.006725 | 0.32745 | 0.002202 | 48.6919 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.5295 | -0.1028 | 0.003694 | 27.7883 | 19.1474 | 31.4027 | 14.4373
2012-10-30 20:29 | 28.836 | 42.2661 | 33.5776 | 45.861 | 28.8873 | -1.4853 | 0.006741 | 0.3276 | 0.002208 | 48.5978 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4923 | -0.1024 | 0.003699 | 27.7738 | 19.0853 | 31.3687 | 14.395
2012-10-30 20:29 | 28.8368 | 42.1607 | 33.5153 | 45.8108 |  | -1.4864 | 0.006732 | 0.32755 | 0.002205 | 48.6563 | 0.006961 | 0.21473 | 0.001495 | 30.8474 | 14.5421 | -0.1025 | 0.0037 | 27.6186 | 18.9732 | 31.2687 | -14.5421
2012-10-30 20:29 | 28.839 | 42.1196 | 33.4859 | 45.7306 |  | -1.4818 | 0.006741 | 0.32758 | 0.002208 | 48.5956 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.5321 | -0.1022 | 0.0037 | 27.5875 | 18.9538 | 31.1985 | -14.5321
2012-10-30 20:29 | 28.8338 | 42.0633 | 33.4412 | 45.6587 | 28.8812 | -1.4772 | 0.006735 | 0.32758 | 0.002206 | 48.6389 | 0.006958 | 0.21476 | 0.001494 | 30.8649 | 14.4289 | -0.1018 | 0.0037 | 27.6344 | 19.0123 | 31.2298 | 14.4523
2012-10-30 20:29 | 28.8353 | 41.9979 | 33.3953 | 45.6053 |  | -1.4905 | 0.006741 | 0.3276 | 0.002208 | 48.5978 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.4712 | -0.1028 | 0.003699 | 27.5267 | 18.9241 | 31.1341 | -14.4712
2012-10-30 20:29 | 28.8382 | 41.9024 | 33.3114 | 45.5218 |  | -1.4899 | 0.006747 | 0.32758 | 0.00221 | 48.5524 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.4741 | -0.1027 | 0.003701 | 27.4283 | 18.8373 | 31.0477 | -14.4741
2012-10-30 20:29 | 28.839 | 41.8245 | 33.2885 | 45.452 | 28.74 | -1.4772 | 0.006741 | 0.32754 | 0.002208 | 48.5892 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.4709 | -0.1018 | 0.0037 | 27.3536 | 18.8176 | 30.9811 | 14.2691
2012-10-30 20:29 | 28.8368 | 41.8027 | 33.2586 | 45.3909 | 28.71 | -1.4934 | 0.006757 | 0.32754 | 0.002213 | 48.4742 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5054 | -0.103 | 0.003702 | 27.2973 | 18.7532 | 30.8855 | 14.2046
2012-10-30 20:29 | 28.827 | 41.6827 | 33.2133 | 45.279 | 28.6566 | -1.4835 | 0.006744 | 0.32766 | 0.00221 | 48.5848 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.4548 | -0.1023 | 0.003703 | 27.2279 | 18.7585 | 30.8242 | 14.2018
2012-10-30 20:29 | 28.8345 | 41.6469 | 33.1734 | 45.2355 | 28.6562 | -1.4812 | 0.006744 | 0.32755 | 0.002209 | 48.5697 | 0.006971 | 0.21473 | 0.001497 | 30.8032 | 14.4745 | -0.1021 | 0.003706 | 27.1724 | 18.6989 | 30.761 | 14.1817
2012-10-30 20:29 | 28.8338 | 41.5321 | 33.1333 | 45.1759 | 28.6079 | -1.4812 | 0.006747 | 0.32754 | 0.00221 | 48.546 | 0.006958 | 0.21474 | 0.001494 | 30.8627 | 14.4534 | -0.1021 | 0.003704 | 27.0787 | 18.6799 | 30.7225 | 14.1545
2012-10-30 20:29 | 28.848 | 41.4754 | 33.1119 | 45.1116 | 28.5864 | -1.4905 | 0.006725 | 0.32754 | 0.002203 | 48.7048 | 0.006958 | 0.21474 | 0.001494 | 30.8627 | 14.5003 | -0.1028 | 0.003697 | 26.9751 | 18.6116 | 30.6113 | 14.0861
2012-10-30 20:29 | 28.8368 | 41.3895 | 32.9902 | 45.0101 | 28.5435 | -1.5015 | 0.006729 | 0.32754 | 0.002204 | 48.6759 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.5176 | -0.1035 | 0.003695 | 26.8719 | 18.4726 | 30.4925 | 14.0259
2012-10-30 20:29 | 28.8382 | 41.3791 | 32.9718 | 44.9487 | 28.5568 | -1.4916 | 0.006754 | 0.32751 | 0.002212 | 48.4914 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4864 | -0.1028 | 0.003703 | 26.8927 | 18.4854 | 30.4623 | 14.0704
2012-10-30 20:30 | 28.8412 | 41.2482 | 32.9985 | 44.9242 | 28.4567 | -1.4916 | 0.006738 | 0.32751 | 0.002207 | 48.6066 | 0.006965 | 0.21473 | 0.001496 | 30.8297 | 14.5139 | -0.1028 | 0.003703 | 26.7343 | 18.4846 | 30.4103 | 13.9428
2012-10-30 20:30 | 28.8353 | 41.1794 | 32.9097 | 44.8123 | 28.4548 | -1.5009 | 0.006732 | 0.32754 | 0.002205 | 48.6542 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.508 | -0.1035 | 0.003696 | 26.6714 | 18.4017 | 30.3043 | 13.9468
2012-10-30 20:30 | 28.8323 | 41.1331 | 32.9067 | 44.7621 | 28.4121 | -1.4939 | 0.006744 | 0.32754 | 0.002209 | 48.5676 | 0.006961 | 0.21474 | 0.001495 | 30.8494 | 14.4846 | -0.103 | 0.003704 | 26.6485 | 18.4221 | 30.2775 | 13.9275
2012-10-30 20:30 | 28.8368 | 41.0824 | 32.8086 | 44.6724 | 28.2777 | -1.4766 | 0.006738 | 0.32754 | 0.002207 | 48.6109 | 0.006958 | 0.2147 | 0.001494 | 30.8566 | 14.5054 | -0.1018 | 0.003701 | 26.577 | 18.3032 | 30.167 | 13.7723
2012-10-30 20:30 | 28.8382 | 41.0445 | 32.7705 | 44.5953 | 28.3585 | -1.4858 | 0.006744 | 0.32751 | 0.002209 | 48.5633 | 0.006955 | 0.21473 | 0.001493 | 30.874 | 14.4782 | -0.1024 | 0.003702 | 26.5663 | 18.2923 | 30.1171 | 13.8803
2012-10-30 20:30 | 28.8368 | 40.9762 | 32.7335 | 44.5349 | 28.2856 | -1.4847 | 0.006744 | 0.32755 | 0.002209 | 48.5697 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5176 | -0.1024 | 0.003699 | 26.4586 | 18.2159 | 30.0173 | 13.768
2012-10-30 20:30 | 28.8338 | 40.8905 | 32.7148 | 44.5005 | 28.2787 | -1.4893 | 0.006735 | 0.32751 | 0.002206 | 48.6282 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.4697 | -0.1027 | 0.003697 | 26.4208 | 18.2451 | 30.0308 | 13.809
2012-10-30 20:30 | 28.8308 | 40.8206 | 32.6604 | 44.4032 | 28.24 | -1.4934 | 0.006751 | 0.32751 | 0.002211 | 48.513 | 0.006965 | 0.21474 | 0.001496 | 30.8317 | 14.4749 | -0.103 | 0.003707 | 26.3457 | 18.1855 | 29.9283 | 13.7651
2012-10-30 20:30 | 28.827 | 40.7301 | 32.5856 | 44.3445 | 28.2123 | -1.498 | 0.006763 | 0.3276 | 0.002216 | 48.4397 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.4507 | -0.1033 | 0.003707 | 26.2794 | 18.1349 | 29.8938 | 13.7616
2012-10-30 20:30 | 28.833 | 40.7204 | 32.5797 | 44.2444 | 28.127 | -1.5038 | 0.006744 | 0.32755 | 0.002209 | 48.5697 | 0.006958 | 0.21468 | 0.001494 | 30.8544 | 14.522 | -0.1037 | 0.003703 | 26.1984 | 18.0577 | 29.7224 | 13.605
2012-10-30 20:30 | 28.83 | 40.6426 | 32.5649 | 44.21 | 28.0803 | -1.4922 | 0.006738 | 0.32763 | 0.002208 | 48.6238 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.519 | -0.1029 | 0.0037 | 26.1236 | 18.0459 | 29.691 | 13.5613
2012-10-30 20:30 | 28.8293 | 40.5434 | 32.4892 | 44.115 | 28.0558 | -1.4939 | 0.006751 | 0.32755 | 0.002211 | 48.5193 | 0.006961 | 0.21477 | 0.001495 | 30.8536 | 14.5061 | -0.103 | 0.003706 | 26.0373 | 17.9831 | 29.6089 | 13.5497
2012-10-30 20:30 | 28.8368 | 40.4485 | 32.4137 | 44.036 | 28.107 | -1.4928 | 0.006757 | 0.32761 | 0.002214 | 48.4848 | 0.006974 | 0.21471 | 0.001497 | 30.7878 | 14.5013 | -0.1029 | 0.003711 | 25.9472 | 17.9124 | 29.5347 | 13.6057
2012-10-30 20:30 | 28.8255 | 40.4215 | 32.4143 | 44.0365 | 28.0639 | -1.4818 | 0.006732 | 0.32754 | 0.002205 | 48.6542 | 0.006958 | 0.2147 | 0.001494 | 30.8566 | 14.4982 | -0.1022 | 0.003699 | 25.9233 | 17.9161 | 29.5383 | 13.5657
2012-10-30 20:30 | 28.8293 | 40.3623 | 32.347 | 43.9382 | 27.9644 | -1.4939 | 0.006722 | 0.32761 | 0.002202 | 48.7373 | 0.006965 | 0.21473 | 0.001496 | 30.8297 | 14.4734 | -0.103 | 0.003698 | 25.8889 | 17.8736 | 29.4648 | 13.491
2012-10-30 20:30 | 28.8255 | 40.3349 | 32.3748 | 43.8833 | 27.9288 | -1.498 | 0.006744 | 0.32757 | 0.002209 | 48.5719 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 14.4941 | -0.1033 | 0.003701 | 25.8408 | 17.8807 | 29.3892 | 13.4347
2012-10-30 20:30 | 28.833 | 40.236 | 32.2836 | 43.8004 | 27.9284 | -1.4893 | 0.006741 | 0.32754 | 0.002208 | 48.5892 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.5016 | -0.1027 | 0.003701 | 25.7344 | 17.782 | 29.2988 | 13.4268
2012-10-30 20:30 | 28.8285 | 40.1449 | 32.2396 | 43.7684 | 27.9318 | -1.4887 | 0.006738 | 0.32758 | 0.002207 | 48.6172 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.5012 | -0.1026 | 0.003699 | 25.6437 | 17.7384 | 29.2672 | 13.4306
2012-10-30 20:30 | 28.8263 | 40.0757 | 32.1466 | 43.6837 | 27.8263 | -1.4934 | 0.006735 | 0.32751 | 0.002206 | 48.6282 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5071 | -0.103 | 0.003697 | 25.5686 | 17.6395 | 29.1766 | 13.3192
2012-10-30 20:30 | 28.8353 | 40.0572 | 32.1437 | 43.6141 | 27.8393 | -1.4939 | 0.006738 | 0.32754 | 0.002207 | 48.6109 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4753 | -0.103 | 0.003699 | 25.5819 | 17.6684 | 29.1388 | 13.364
2012-10-30 20:30 | 28.8338 | 39.9808 | 32.0909 | 43.5929 | 27.8457 | -1.4916 | 0.006751 | 0.32757 | 0.002211 | 48.5216 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.4983 | -0.1028 | 0.003704 | 25.4825 | 17.5926 | 29.0946 | 13.3474
2012-10-30 20:30 | 28.8278 | 39.9275 | 32.0849 | 43.4729 | 27.8 | -1.4934 | 0.006744 | 0.32748 | 0.002209 | 48.559 | 0.006952 | 0.21482 | 0.001493 | 30.8997 | 14.5005 | -0.103 | 0.003702 | 25.427 | 17.5844 | 28.9724 | 13.2995
2012-10-30 20:30 | 28.8293 | 39.8345 | 32.0469 | 43.4037 | 27.7617 | -1.4968 | 0.006754 | 0.32754 | 0.002212 | 48.4957 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.502 | -0.1032 | 0.003704 | 25.3325 | 17.5449 | 28.9017 | 13.2597
2012-10-30 20:30 | 28.821 | 39.7908 | 31.9794 | 43.3522 | 27.7892 | -1.4887 | 0.006744 | 0.32755 | 0.002209 | 48.5697 | 0.006965 | 0.21474 | 0.001496 | 30.8317 | 14.4937 | -0.1026 | 0.003705 | 25.2971 | 17.4857 | 28.8585 | 13.2955
2012-10-30 20:30 | 28.8218 | 39.7364 | 31.9407 | 43.2979 | 27.6628 | -1.4951 | 0.006744 | 0.3275 | 0.002209 | 48.5612 | 0.006952 | 0.21477 | 0.001493 | 30.8936 | 14.4945 | -0.1031 | 0.003702 | 25.2419 | 17.4462 | 28.8034 | 13.1683
2012-10-30 20:31 | 28.8293 | 39.6651 | 31.9364 | 43.2543 | 27.6782 | -1.4934 | 0.006741 | 0.32748 | 0.002208 | 48.5806 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5183 | -0.103 | 0.0037 | 25.1468 | 17.4181 | 28.736 | 13.1599
2012-10-30 20:31 | 28.824 | 39.5968 | 31.8876 | 43.1507 | 27.6133 | -1.4968 | 0.006738 | 0.3275 | 0.002207 | 48.6045 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.5171 | -0.1032 | 0.0037 | 25.0797 | 17.3705 | 28.6336 | 13.0962
2012-10-30 20:31 | 28.8248 | 39.5464 | 31.8529 | 43.0926 | 27.5902 | -1.5003 | 0.006744 | 0.32754 | 0.002209 | 48.5676 | 0.006958 | 0.21474 | 0.001494 | 30.8627 | 14.526 | -0.1034 | 0.003703 | 25.0204 | 17.3269 | 28.5666 | 13.0642
2012-10-30 20:31 | 28.83 | 39.4806 | 31.8068 | 43.0191 | 27.5755 | -1.4968 | 0.006744 | 0.32747 | 0.002208 | 48.5569 | 0.006965 | 0.21479 | 0.001496 | 30.8379 | 14.5108 | -0.1032 | 0.003704 | 24.9698 | 17.296 | 28.5083 | 13.0647
2012-10-30 20:31 | 28.8353 | 39.4426 | 31.7923 | 43.0205 | 27.5251 | -1.4945 | 0.006732 | 0.32753 | 0.002205 | 48.652 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.5243 | -0.103 | 0.003692 | 24.9183 | 17.268 | 28.4962 | 13.0008
2012-10-30 20:31 | 28.824 | 39.3683 | 31.7218 | 42.9227 | 27.4302 | -1.4905 | 0.006744 | 0.32755 | 0.002209 | 48.5697 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.5538 | -0.1028 | 0.003702 | 24.8145 | 17.168 | 28.3689 | 12.8764
2012-10-30 20:31 | 28.8218 | 39.3227 | 31.6722 | 42.8694 | 27.4121 | -1.4945 | 0.006738 | 0.32755 | 0.002207 | 48.6129 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.519 | -0.103 | 0.0037 | 24.8037 | 17.1532 | 28.3504 | 12.8931
2012-10-30 20:31 | 28.8218 | 39.2321 | 31.6446 | 42.7671 | 27.432 | -1.4997 | 0.006738 | 0.3275 | 0.002207 | 48.6045 | 0.006961 | 0.21473 | 0.001495 | 30.8474 | 14.5067 | -0.1034 | 0.003702 | 24.7254 | 17.1379 | 28.2604 | 12.9253
2012-10-30 20:31 | 28.8233 | 39.1942 | 31.5671 | 42.7057 | 27.4256 | -1.5032 | 0.006744 | 0.32745 | 0.002208 | 48.5547 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.4837 | -0.1036 | 0.003698 | 24.7105 | 17.0834 | 28.222 | 12.9419
2012-10-30 20:31 | 28.824 | 39.1437 | 31.556 | 42.6828 | 27.4024 | -1.5009 | 0.006744 | 0.32754 | 0.002209 | 48.5676 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5048 | -0.1035 | 0.0037 | 24.6389 | 17.0512 | 28.178 | 12.8976
2012-10-30 20:31 | 28.8285 | 39.0772 | 31.4973 | 42.5615 | 27.387 | -1.4928 | 0.006751 | 0.32747 | 0.002211 | 48.5066 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5338 | -0.1029 | 0.003702 | 24.5434 | 16.9635 | 28.0277 | 12.8532
2012-10-30 20:31 | 28.8255 | 38.9796 | 31.4746 | 42.4798 | 27.3521 | -1.5032 | 0.006747 | 0.32744 | 0.002209 | 48.5311 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.49 | -0.1036 | 0.003699 | 24.4896 | 16.9846 | 27.9898 | 12.8621
2012-10-30 20:31 | 28.8218 | 38.9996 | 31.463 | 42.4801 | 27.2648 | -1.4911 | 0.006751 | 0.32747 | 0.002211 | 48.5066 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.5026 | -0.1028 | 0.003703 | 24.497 | 16.9604 | 27.9775 | 12.7622
2012-10-30 20:31 | 28.836 | 38.864 | 31.4337 | 42.4195 | 27.2511 | -1.5032 | 0.006747 | 0.32751 | 0.00221 | 48.5417 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4923 | -0.1036 | 0.003702 | 24.3717 | 16.9414 | 27.9272 | 12.7588
2012-10-30 20:31 | 28.83 | 38.8501 | 31.3883 | 42.3349 | 27.281 | -1.498 | 0.006722 | 0.32751 | 0.002202 | 48.7223 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5149 | -0.1033 | 0.003691 | 24.3352 | 16.8734 | 27.82 | 12.7661
2012-10-30 20:31 | 28.8323 | 38.8052 | 31.2958 | 42.3175 | 27.2156 | -1.4957 | 0.006747 | 0.32747 | 0.002209 | 48.5353 | 0.006961 | 0.21471 | 0.001495 | 30.8453 | 14.5295 | -0.1031 | 0.003704 | 24.2757 | 16.7663 | 27.788 | 12.6861
2012-10-30 20:31 | 28.8255 | 38.7077 | 31.285 | 42.2359 | 27.1928 | -1.4945 | 0.006735 | 0.3275 | 0.002206 | 48.6261 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5023 | -0.103 | 0.003696 | 24.2054 | 16.7827 | 27.7336 | 12.6905
2012-10-30 20:31 | 28.8278 | 38.6469 | 31.2834 | 42.1556 | 27.1792 | -1.4974 | 0.006738 | 0.32748 | 0.002207 | 48.6023 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5168 | -0.1032 | 0.003696 | 24.1301 | 16.7666 | 27.6388 | 12.6624
2012-10-30 20:31 | 28.8323 | 38.5765 | 31.2208 | 42.1169 | 27.1638 | -1.5049 | 0.006747 | 0.32745 | 0.002209 | 48.5331 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5254 | -0.1037 | 0.003701 | 24.0511 | 16.6954 | 27.5915 | 12.6384
2012-10-30 20:31 | 28.83 | 38.5506 | 31.2066 | 42.0713 | 27.1495 | -1.4818 | 0.006754 | 0.32747 | 0.002212 | 48.485 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.5476 | -0.1022 | 0.003705 | 24.003 | 16.659 | 27.5237 | 12.6019
2012-10-30 20:31 | 28.8323 | 38.4583 | 31.1576 | 41.9988 | 27.1518 | -1.502 | 0.006738 | 0.32745 | 0.002206 | 48.598 | 0.006965 | 0.21477 | 0.001496 | 30.8359 | 14.5376 | -0.1035 | 0.003702 | 23.9207 | 16.62 | 27.4612 | 12.6142
2012-10-30 20:31 | 28.8353 | 38.4455 | 31.1369 | 41.9468 | 27.107 | -1.4951 | 0.006747 | 0.32757 | 0.00221 | 48.5503 | 0.006955 | 0.21473 | 0.001493 | 30.874 | 14.5121 | -0.1031 | 0.003703 | 23.9334 | 16.6248 | 27.4347 | 12.5949
2012-10-30 20:31 | 28.8338 | 38.4007 | 31.0406 | 41.8823 | 27.1414 | -1.4957 | 0.006747 | 0.32744 | 0.002209 | 48.5311 | 0.006961 | 0.21477 | 0.001495 | 30.8536 | 14.5106 | -0.1031 | 0.003704 | 23.8901 | 16.53 | 27.3717 | 12.6308
2012-10-30 20:31 | 28.8225 | 38.2711 | 31.0609 | 41.8041 | 27.0106 | -1.502 | 0.006735 | 0.32751 | 0.002206 | 48.6282 | 0.006955 | 0.21473 | 0.001493 | 30.874 | 14.5197 | -0.1035 | 0.003699 | 23.7514 | 16.5412 | 27.2844 | 12.4909
2012-10-30 20:31 | 28.8345 | 38.2121 | 31.0058 | 41.7453 | 26.9987 | -1.5043 | 0.006741 | 0.32758 | 0.002208 | 48.5956 | 0.006965 | 0.21471 | 0.001495 | 30.8276 | 14.5521 | -0.1037 | 0.003703 | 23.66 | 16.4537 | 27.1932 | 12.4466
2012-10-30 20:31 | 28.8375 | 38.1875 | 30.9733 | 41.705 | 26.9658 | -1.498 | 0.006744 | 0.32751 | 0.002209 | 48.5633 | 0.006958 | 0.21479 | 0.001494 | 30.8689 | 14.5143 | -0.1033 | 0.003703 | 23.6732 | 16.459 | 27.1907 | 12.4515
2012-10-30 20:32 | 28.833 | 38.0884 | 30.9135 | 41.6021 | 26.9892 | -1.5078 | 0.006757 | 0.32748 | 0.002213 | 48.4656 | 0.006955 | 0.21483 | 0.001494 | 30.8886 | 14.5342 | -0.1039 | 0.003707 | 23.5542 | 16.3793 | 27.0679 | 12.455
2012-10-30 20:32 | 28.827 | 38.0272 | 30.8878 | 41.5292 | 26.8835 | -1.5032 | 0.006741 | 0.32739 | 0.002207 | 48.5677 | 0.006961 | 0.21473 | 0.001495 | 30.8474 | 14.5527 | -0.1036 | 0.003702 | 23.4745 | 16.3351 | 26.9765 | 12.3308
2012-10-30 20:32 | 28.8345 | 37.9795 | 30.8597 | 41.4698 | 26.8551 | -1.5124 | 0.006744 | 0.32742 | 0.002208 | 48.5504 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.5643 | -0.1043 | 0.0037 | 23.4152 | 16.2954 | 26.9055 | 12.2908
2012-10-30 20:32 | 28.8285 | 37.9341 | 30.8182 | 41.4402 | 26.8292 | -1.4928 | 0.006738 | 0.32739 | 0.002206 | 48.5893 | 0.006958 | 0.21477 | 0.001494 | 30.8669 | 14.5624 | -0.1029 | 0.0037 | 23.3717 | 16.2558 | 26.8778 | 12.2668
2012-10-30 20:32 | 28.827 | 37.8695 | 30.7969 | 41.3679 | 26.8197 | -1.5015 | 0.006744 | 0.3275 | 0.002209 | 48.5612 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.5894 | -0.1035 | 0.003702 | 23.2801 | 16.2075 | 26.7785 | 12.2303
2012-10-30 20:32 | 28.8255 | 37.8404 | 30.7875 | 41.327 | 26.7704 | -1.5078 | 0.006741 | 0.32748 | 0.002208 | 48.5806 | 0.006958 | 0.21479 | 0.001494 | 30.8689 | 14.5145 | -0.1039 | 0.003702 | 23.3259 | 16.273 | 26.8125 | 12.2559
2012-10-30 20:32 | 28.8233 | 37.7554 | 30.6945 | 41.2618 | 26.7682 | -1.5119 | 0.006735 | 0.32744 | 0.002205 | 48.6175 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.5409 | -0.1042 | 0.003698 | 23.2145 | 16.1536 | 26.7209 | 12.2273
2012-10-30 20:32 | 28.8203 | 37.4646 | 30.5415 | 40.9675 | 26.6176 | -1.4991 | 0.006747 | 0.32744 | 0.002209 | 48.5311 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.5379 | -0.1033 | 0.003702 | 22.9267 | 16.0036 | 26.4296 | 12.0797
2012-10-30 20:32 | 28.8075 | 37.2033 | 30.355 | 40.6555 | 26.4771 | -1.4974 | 0.006738 | 0.32738 | 0.002206 | 48.5873 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4883 | -0.1032 | 0.003694 | 22.715 | 15.8667 | 26.1672 | 11.9888
2012-10-30 20:32 | 28.8075 | 37.1481 | 30.2681 | 40.5728 | 26.4811 | -1.4893 | 0.006735 | 0.32734 | 0.002205 | 48.6025 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4883 | -0.1027 | 0.003697 | 22.6598 | 15.7798 | 26.0845 | 11.9928
2012-10-30 20:32 | 28.8128 | 37.0627 | 30.3208 | 40.6017 | 26.3706 | -1.5067 | 0.006735 | 0.32735 | 0.002205 | 48.6046 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5059 | -0.1039 | 0.003695 | 22.5568 | 15.8149 | 26.0958 | 11.8647
2012-10-30 20:32 | 28.812 | 37.0461 | 30.2213 | 40.5143 | 26.4177 | -1.4974 | 0.006738 | 0.32737 | 0.002206 | 48.5852 | 0.006955 | 0.21468 | 0.001493 | 30.8677 | 14.5459 | -0.1032 | 0.003699 | 22.5002 | 15.6754 | 25.9684 | 11.8718
2012-10-30 20:32 | 28.803 | 36.9148 | 30.1649 | 40.4029 | 26.3289 | -1.5003 | 0.006725 | 0.32731 | 0.002201 | 48.6705 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4879 | -0.1034 | 0.003692 | 22.4269 | 15.677 | 25.915 | 11.841
2012-10-30 20:32 | 28.8098 | 36.9058 | 30.1599 | 40.3192 | 26.3597 | -1.4997 | 0.006747 | 0.32738 | 0.002209 | 48.5225 | 0.006961 | 0.21467 | 0.001494 | 30.8391 | 14.511 | -0.1034 | 0.003703 | 22.3948 | 15.6489 | 25.8082 | 11.8487
2012-10-30 20:32 | 28.8 | 36.8527 | 30.158 | 40.333 | 26.3139 | -1.4986 | 0.006738 | 0.32738 | 0.002206 | 48.5873 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.5135 | -0.1033 | 0.0037 | 22.3392 | 15.6445 | 25.8195 | 11.8004
2012-10-30 20:32 | 28.7993 | 36.7573 | 30.102 | 40.2023 | 26.2493 | -1.5015 | 0.006751 | 0.32739 | 0.00221 | 48.4958 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5087 | -0.1035 | 0.003699 | 22.2486 | 15.5933 | 25.6936 | 11.7406
2012-10-30 20:32 | 28.8 | 36.7304 | 30.0514 | 40.1636 | 26.238 | -1.4911 | 0.006738 | 0.32735 | 0.002206 | 48.583 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4931 | -0.1028 | 0.003697 | 22.2373 | 15.5583 | 25.6705 | 11.7449
2012-10-30 20:33 | 28.7993 | 36.6153 | 30.0467 | 40.096 | 26.2094 | -1.5072 | 0.006747 | 0.32739 | 0.002209 | 48.5245 | 0.006971 | 0.21467 | 0.001496 | 30.7949 | 14.5209 | -0.1039 | 0.003705 | 22.0944 | 15.5258 | 25.5751 | 11.6885
2012-10-30 20:33 | 28.8 | 36.6278 | 29.9645 | 40.0415 | 26.1662 | -1.4934 | 0.006744 | 0.32737 | 0.002208 | 48.542 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.5053 | -0.103 | 0.003699 | 22.1225 | 15.4592 | 25.5362 | 11.6609
2012-10-30 20:33 | 28.7933 | 36.5383 | 29.9578 | 39.9994 | 26.1195 | -1.4928 | 0.006741 | 0.32742 | 0.002207 | 48.572 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.4905 | -0.1029 | 0.0037 | 22.0478 | 15.4673 | 25.5089 | 11.629
2012-10-30 20:33 | 28.7933 | 36.4554 | 29.8986 | 39.9324 | 26.0716 | -1.4858 | 0.006738 | 0.32735 | 0.002206 | 48.583 | 0.006952 | 0.21464 | 0.001492 | 30.8749 | 14.4619 | -0.1024 | 0.003698 | 21.9935 | 15.4367 | 25.4705 | 11.6097
2012-10-30 20:33 | 28.7978 | 36.3968 | 29.8913 | 39.9054 | 26.0801 | -1.4893 | 0.006741 | 0.32734 | 0.002207 | 48.5593 | 0.006955 | 0.21468 | 0.001493 | 30.8677 | 14.5031 | -0.1027 | 0.0037 | 21.8937 | 15.3882 | 25.4023 | 11.577
2012-10-30 20:33 | 28.7948 | 36.3465 | 29.8488 | 39.8118 | 26.0411 | -1.5003 | 0.006754 | 0.32738 | 0.002211 | 48.4722 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.5083 | -0.1034 | 0.003702 | 21.8382 | 15.3405 | 25.3035 | 11.5328
2012-10-30 20:33 | 28.7926 | 36.3285 | 29.8308 | 39.7427 | 26.0309 | -1.4847 | 0.006741 | 0.32741 | 0.002207 | 48.5699 | 0.006958 | 0.21468 | 0.001494 | 30.8544 | 14.4857 | -0.1024 | 0.003701 | 21.8428 | 15.3451 | 25.257 | 11.5452
2012-10-30 20:33 | 28.7941 | 36.2787 | 29.8086 | 39.7245 | 26.0524 | -1.4882 | 0.006744 | 0.32739 | 0.002208 | 48.5461 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4545 | -0.1026 | 0.003699 | 21.8242 | 15.3541 | 25.27 | 11.5979
2012-10-30 20:33 | 28.7956 | 36.2526 | 29.7588 | 39.6787 | 26.0299 | -1.4853 | 0.006738 | 0.32741 | 0.002206 | 48.5916 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4846 | -0.1024 | 0.003696 | 21.768 | 15.2742 | 25.1941 | 11.5453
2012-10-30 20:33 | 28.7933 | 36.1674 | 29.6894 | 39.5621 | 25.8957 | -1.4922 | 0.006735 | 0.32735 | 0.002205 | 48.6046 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4782 | -0.1029 | 0.003695 | 21.6892 | 15.2112 | 25.0839 | 11.4175
2012-10-30 20:33 | 28.8 | 36.1149 | 29.7 | 39.5452 | 25.9024 | -1.4789 | 0.006725 | 0.32737 | 0.002202 | 48.6791 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.5176 | -0.102 | 0.003695 | 21.5973 | 15.1824 | 25.0276 | 11.3848
2012-10-30 20:33 | 28.7963 | 36.0994 | 29.6766 | 39.5336 | 25.9187 | -1.4697 | 0.006744 | 0.32741 | 0.002208 | 48.5483 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4731 | -0.1013 | 0.003699 | 21.6263 | 15.2035 | 25.0605 | 11.4456
2012-10-30 20:33 | 28.7836 | 36.0039 | 29.6165 | 39.4382 | 25.862 | -1.4806 | 0.006744 | 0.32738 | 0.002208 | 48.544 | 0.006955 | 0.21473 | 0.001493 | 30.874 | 14.4399 | -0.1021 | 0.003701 | 21.564 | 15.1766 | 24.9983 | 11.4221
2012-10-30 20:33 | 28.7843 | 35.9533 | 29.5856 | 39.3719 | 25.7948 | -1.4795 | 0.006741 | 0.32741 | 0.002207 | 48.5699 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4815 | -0.102 | 0.003698 | 21.4718 | 15.1041 | 24.8904 | 11.3133
2012-10-30 20:33 | 28.7866 | 35.9437 | 29.5603 | 39.3309 | 25.7771 | -1.4841 | 0.006741 | 0.32738 | 0.002207 | 48.5656 | 0.006965 | 0.2147 | 0.001495 | 30.8256 | 14.5001 | -0.1023 | 0.003702 | 21.4436 | 15.0602 | 24.8308 | 11.277
2012-10-30 20:33 | 28.7746 | 35.8923 | 29.497 | 39.2952 | 25.7491 | -1.4841 | 0.006735 | 0.32738 | 0.002205 | 48.6089 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4636 | -0.1023 | 0.003695 | 21.4287 | 15.0334 | 24.8316 | 11.2855
2012-10-30 20:33 | 28.7738 | 35.8441 | 29.512 | 39.2117 | 25.6923 | -1.4835 | 0.006741 | 0.32738 | 0.002207 | 48.5656 | 0.006924 | 0.21471 | 0.001487 | 31.0101 | 14.4587 | -0.1023 | 0.003694 | 21.3854 | 15.0533 | 24.753 | 11.2336
2012-10-30 20:33 | 28.7768 | 35.7643 | 29.44 | 39.1674 | 25.7073 | -1.4766 | 0.006744 | 0.32735 | 0.002208 | 48.5397 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.474 | -0.1018 | 0.003696 | 21.2903 | 14.966 | 24.6934 | 11.2333
2012-10-30 20:33 | 28.7693 | 35.7133 | 29.3575 | 39.0968 | 25.5598 | -1.4801 | 0.006757 | 0.32742 | 0.002212 | 48.457 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4787 | -0.102 | 0.003701 | 21.2346 | 14.8788 | 24.6181 | 11.0811
2012-10-30 20:33 | 28.7678 | 35.6645 | 29.3638 | 39.0835 | 25.4743 | -1.4899 | 0.006738 | 0.32734 | 0.002206 | 48.5809 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4486 | -0.1027 | 0.003697 | 21.2159 | 14.9152 | 24.6349 | 11.0257
2012-10-30 20:33 | 28.7738 | 35.5995 | 29.3343 | 38.9752 | 25.5243 | -1.4951 | 0.006744 | 0.32731 | 0.002207 | 48.5334 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4546 | -0.1031 | 0.003696 | 21.1449 | 14.8797 | 24.5206 | 11.0697
2012-10-30 20:33 | 28.7731 | 35.5593 | 29.3178 | 38.9194 | 25.4716 | -1.4812 | 0.006741 | 0.32741 | 0.002207 | 48.5699 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4784 | -0.1021 | 0.003699 | 21.0809 | 14.8394 | 24.441 | 10.9932
2012-10-30 20:33 | 28.7746 | 35.5174 | 29.2838 | 38.9051 | 25.4411 | -1.4864 | 0.006747 | 0.32734 | 0.002209 | 48.5161 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.5003 | -0.1025 | 0.003702 | 21.0171 | 14.7835 | 24.4048 | 10.9408
2012-10-30 20:33 | 28.7753 | 35.4629 | 29.2174 | 38.8033 | 25.4938 | -1.4916 | 0.006744 | 0.32734 | 0.002208 | 48.5377 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.448 | -0.1028 | 0.003701 | 21.0149 | 14.7694 | 24.3553 | 11.0458
2012-10-30 20:33 | 28.7768 | 35.3933 | 29.2505 | 38.7615 | 25.4553 | -1.4824 | 0.006757 | 0.32742 | 0.002212 | 48.457 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4576 | -0.1022 | 0.003702 | 20.9357 | 14.7929 | 24.3039 | 10.9977
2012-10-30 20:34 | 28.7671 | 35.3521 | 29.1263 | 38.7124 | 25.4456 | -1.487 | 0.00676 | 0.32726 | 0.002212 | 48.412 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.4887 | -0.1025 | 0.003704 | 20.8634 | 14.6376 | 24.2237 | 10.9569
2012-10-30 20:34 | 28.7753 | 35.305 | 29.1305 | 38.689 | 25.4218 | -1.4945 | 0.006744 | 0.32735 | 0.002208 | 48.5397 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.4561 | -0.103 | 0.003701 | 20.8489 | 14.6744 | 24.2329 | 10.9657
2012-10-30 20:34 | 28.7738 | 35.2561 | 29.1054 | 38.6166 | 25.3842 | -1.483 | 0.006766 | 0.32738 | 0.002215 | 48.3862 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.4832 | -0.1022 | 0.003706 | 20.7729 | 14.6222 | 24.1334 | 10.901
2012-10-30 20:34 | 28.7791 | 35.1746 | 29.0515 | 38.5628 | 25.3175 | -1.4853 | 0.006729 | 0.32729 | 0.002202 | 48.6393 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.5048 | -0.1024 | 0.00369 | 20.6698 | 14.5467 | 24.058 | 10.8127
2012-10-30 20:34 | 28.7843 | 35.168 | 29.0961 | 38.5167 | 25.2546 | -1.4893 | 0.006751 | 0.32735 | 0.00221 | 48.4894 | 0.006977 | 0.21461 | 0.001497 | 30.7601 | 14.4896 | -0.1027 | 0.003707 | 20.6784 | 14.6065 | 24.0271 | 10.765
2012-10-30 20:34 | 28.7813 | 35.0939 | 29.0418 | 38.498 | 25.2596 | -1.476 | 0.006747 | 0.32731 | 0.002208 | 48.5118 | 0.006958 | 0.21471 | 0.001494 | 30.8586 | 14.4825 | -0.1018 | 0.003702 | 20.6114 | 14.5593 | 24.0155 | 10.7771
2012-10-30 20:34 | 28.7918 | 35.0571 | 29.001 | 38.402 | 25.2741 | -1.4928 | 0.006741 | 0.32737 | 0.002207 | 48.5636 | 0.006949 | 0.2146 | 0.001491 | 30.882 | 14.4686 | -0.1029 | 0.003698 | 20.5885 | 14.5324 | 23.9334 | 10.8055
2012-10-30 20:34 | 28.7881 | 35.0258 | 28.9934 | 38.3274 | 25.2104 | -1.4853 | 0.006735 | 0.32734 | 0.002205 | 48.6025 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.5016 | -0.1024 | 0.003698 | 20.5242 | 14.4918 | 23.8258 | 10.7088
2012-10-30 20:34 | 28.7828 | 34.9415 | 28.9209 | 38.3063 | 25.153 | -1.4806 | 0.006757 | 0.32737 | 0.002212 | 48.4486 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4636 | -0.1021 | 0.003702 | 20.4779 | 14.4573 | 23.8427 | 10.6894
2012-10-30 20:34 | 28.7798 | 34.8793 | 28.9101 | 38.2245 | 25.106 | -1.4864 | 0.006751 | 0.32735 | 0.00221 | 48.4894 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5137 | -0.1025 | 0.003701 | 20.3656 | 14.3964 | 23.7108 | 10.5923
2012-10-30 20:34 | 28.7813 | 34.8019 |  | 38.1708 | 25.1515 | -1.4818 | 0.006732 | 0.32735 | 0.002204 | 48.6263 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.507 | -0.1022 | 0.003696 | 20.2949 | -14.507 | 23.6638 | 10.6445
2012-10-30 20:34 | 28.7873 | 34.8118 |  | 38.1373 | 25.1175 | -1.4986 | 0.006744 | 0.32737 | 0.002208 | 48.542 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4681 | -0.1033 | 0.003697 | 20.3437 | -14.4681 | 23.6692 | 10.6494
2012-10-30 20:34 | 28.7873 | 34.7408 |  | 38.1137 | 25.0774 | -1.4911 | 0.006744 | 0.32728 | 0.002207 | 48.5291 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.5416 | -0.1028 | 0.0037 | 20.1992 | -14.5416 | 23.5721 | 10.5358
2012-10-30 20:34 | 28.7828 | 34.7126 |  | 38.0461 | 25.0969 | -1.4795 | 0.006751 | 0.32722 | 0.002209 | 48.4702 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.5085 | -0.102 | 0.003702 | 20.2041 | -14.5085 | 23.5376 | 10.5884
2012-10-30 20:34 | 28.7783 | 34.6844 |  | 37.9706 | 25.0844 | -1.4772 | 0.006729 | 0.32729 | 0.002202 | 48.6393 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.5203 | -0.1018 | 0.003695 | 20.1641 | -14.5203 | 23.4503 | 10.5641
2012-10-30 20:34 | 28.7813 | 34.6045 |  | 37.95 | 25.0514 | -1.4858 | 0.006747 | 0.32738 | 0.002209 | 48.5225 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4907 | -0.1024 | 0.0037 | 20.1138 | -14.4907 | 23.4593 | 10.5607
2012-10-30 20:34 | 28.7783 | 34.5739 |  | 37.9154 | 24.9683 | -1.4812 | 0.006741 | 0.32737 | 0.002207 | 48.5636 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.4591 | -0.1021 | 0.003699 | 20.1148 | -14.4591 | 23.4563 | 10.5092
2012-10-30 20:34 | 28.7753 | 34.5117 | 28.708 | 37.8691 | 24.9332 | -1.4841 | 0.006747 | 0.32732 | 0.002208 | 48.5139 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.4969 | -0.1023 | 0.003701 | 20.0148 | 14.2111 | 23.3722 | 10.4363
2012-10-30 20:34 | 28.7761 | 34.4651 | 28.681 | 37.7792 | 24.914 | -1.4847 | 0.006744 | 0.32726 | 0.002207 | 48.5268 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.4733 | -0.1024 | 0.003699 | 19.9918 | 14.2077 | 23.3059 | 10.4407
2012-10-30 20:34 | 28.7768 | 34.4382 | 28.6659 | 37.7286 | 24.8385 | -1.4847 | 0.006738 | 0.32728 | 0.002205 | 48.5723 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.4372 | -0.1024 | 0.003698 | 20.001 | 14.2287 | 23.2914 | 10.4013
2012-10-30 20:34 | 28.7836 | 34.3937 | 28.6013 | 37.6802 | 24.8573 | -1.4934 | 0.006744 | 0.32731 | 0.002207 | 48.5334 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.5175 | -0.103 | 0.0037 | 19.8762 | 14.0838 | 23.1627 | 10.3398
2012-10-30 20:34 | 28.7693 | 34.3399 | 28.591 | 37.662 | 24.8591 | -1.4847 | 0.006735 | 0.32734 | 0.002205 | 48.6025 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4869 | -0.1024 | 0.003694 | 19.853 | 14.1041 | 23.1751 | 10.3722
2012-10-30 20:34 | 28.7641 | 34.2953 | 28.5461 | 37.6095 | 24.7978 | -1.4841 | 0.006735 | 0.32731 | 0.002204 | 48.5982 | 0.006936 | 0.2146 | 0.001488 | 30.9399 | 14.4817 | -0.1023 | 0.003692 | 19.8136 | 14.0644 | 23.1278 | 10.3161
2012-10-30 20:34 | 28.7633 | 34.2273 | 28.5136 | 37.5653 | 24.805 | -1.4812 | 0.006735 | 0.32729 | 0.002204 | 48.596 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4809 | -0.1021 | 0.003695 | 19.7464 | 14.0327 | 23.0844 | 10.3241
2012-10-30 20:34 | 28.7581 | 34.1866 | 28.4886 | 37.5325 | 24.7677 | -1.476 | 0.006738 | 0.32728 | 0.002205 | 48.5723 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4675 | -0.1018 | 0.003696 | 19.7191 | 14.0211 | 23.065 | 10.3002
2012-10-30 20:34 | 28.7656 | 34.1625 | 28.5001 | 37.5006 | 24.7472 | -1.4916 | 0.006741 | 0.32726 | 0.002206 | 48.5484 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.5362 | -0.1028 | 0.003697 | 19.6263 | 13.9639 | 22.9644 | 10.211
2012-10-30 20:35 | 28.7581 | 34.084 | 28.4252 | 37.4378 | 24.7717 | -1.4864 | 0.006725 | 0.32722 | 0.002201 | 48.6575 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.5287 | -0.1025 | 0.003694 | 19.5553 | 13.8965 | 22.9091 | 10.243
2012-10-30 20:35 | 28.7469 | 34.0609 | 28.4021 | 37.3359 | 24.7605 | -1.4887 | 0.006732 | 0.32724 | 0.002203 | 48.6092 | 0.006955 | 0.21468 | 0.001493 | 30.8677 | 14.493 | -0.1026 | 0.003696 | 19.5679 | 13.9091 | 22.8429 | 10.2675
2012-10-30 20:35 | 28.7506 | 34.0331 | 28.3899 | 37.312 | 24.7081 | -1.4887 | 0.006732 | 0.32722 | 0.002203 | 48.607 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4926 | -0.1026 | 0.003692 | 19.5405 | 13.8973 | 22.8194 | 10.2155
2012-10-30 20:35 | 28.7484 | 33.9756 | 28.4036 | 37.2822 | 24.6778 | -1.4858 | 0.006738 | 0.32722 | 0.002205 | 48.5637 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5027 | -0.1024 | 0.003695 | 19.4729 | 13.9009 | 22.7795 | 10.1751
2012-10-30 20:35 | 28.7491 | 33.9013 | 28.3408 | 37.2435 | 24.6465 | -1.4812 | 0.006738 | 0.32731 | 0.002205 | 48.5766 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.4503 | -0.1021 | 0.003698 | 19.451 | 13.8905 | 22.7932 | 10.1962
2012-10-30 20:35 | 28.7394 | 33.86 | 28.3351 | 37.1786 | 24.6247 | -1.4806 | 0.006741 | 0.32725 | 0.002206 | 48.5464 | 0.006958 | 0.21467 | 0.001494 | 30.8524 | 14.5018 | -0.1021 | 0.0037 | 19.3582 | 13.8333 | 22.6768 | 10.1229
2012-10-30 20:35 | 28.7416 | 33.8267 | 28.2539 | 37.1058 | 24.6309 | -1.4749 | 0.006735 | 0.32726 | 0.002204 | 48.5917 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.5489 | -0.1017 | 0.003694 | 19.2778 | 13.705 | 22.5569 | 10.082
2012-10-30 20:35 | 28.7431 | 33.7966 | 28.2515 | 37.064 | 24.5924 | -1.4795 | 0.006722 | 0.32722 | 0.0022 | 48.6793 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.5259 | -0.102 | 0.003689 | 19.2707 | 13.7256 | 22.5381 | 10.0665
2012-10-30 20:35 | 28.7349 | 33.6976 | 28.2274 | 37.0084 | 24.5962 | -1.4945 | 0.006735 | 0.32726 | 0.002204 | 48.5917 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.4933 | -0.103 | 0.003695 | 19.2043 | 13.7341 | 22.5151 | 10.1029
2012-10-30 20:35 | 28.7356 | 33.7023 | 28.2122 | 36.9894 | 24.6089 | -1.4691 | 0.006732 | 0.32725 | 0.002203 | 48.6113 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5103 | -0.1013 | 0.003694 | 19.192 | 13.7019 | 22.4791 | 10.0986
2012-10-30 20:35 | 28.7229 | 33.6106 | 28.1916 | 36.882 | 24.5481 | -1.4778 | 0.006738 | 0.32725 | 0.002205 | 48.568 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.5057 | -0.1019 | 0.003694 | 19.1049 | 13.6859 | 22.3763 | 10.0424
2012-10-30 20:35 | 28.7229 | 33.579 | 28.1003 | 36.8663 | 24.4679 | -1.4737 | 0.006732 | 0.32728 | 0.002203 | 48.6156 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5261 | -0.1016 | 0.003694 | 19.0529 | 13.5742 | 22.3402 | 9.9418
2012-10-30 20:35 | 28.7244 | 33.545 | 28.1137 | 36.8638 | 24.4934 | -1.4749 | 0.006732 | 0.32728 | 0.002203 | 48.6156 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.5154 | -0.1017 | 0.003694 | 19.0296 | 13.5983 | 22.3484 | 9.978
2012-10-30 20:35 | 28.7244 | 33.4858 | 28.0939 | 36.7771 | 24.4373 | -1.476 | 0.006751 | 0.32721 | 0.002209 | 48.4681 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.5358 | -0.1018 | 0.003698 | 18.95 | 13.5581 | 22.2413 | 9.9015
2012-10-30 20:35 | 28.7221 | 33.4479 | 28.0717 | 36.7866 | 24.3909 | -1.4847 | 0.006741 | 0.32731 | 0.002206 | 48.555 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.5049 | -0.1024 | 0.003696 | 18.943 | 13.5668 | 22.2817 | 9.886
2012-10-30 20:35 | 28.7206 | 33.3714 | 28.0424 | 36.6707 | 24.4014 | -1.4737 | 0.006754 | 0.32732 | 0.002211 | 48.4636 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.5116 | -0.1016 | 0.003704 | 18.8598 | 13.5308 | 22.1591 | 9.8898
2012-10-30 20:35 | 28.7154 | 33.3386 | 28.0134 | 36.6497 | 24.316 | -1.4714 | 0.006751 | 0.32724 | 0.002209 | 48.4724 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5472 | -0.1014 | 0.003699 | 18.7914 | 13.4662 | 22.1025 | 9.7688
2012-10-30 20:35 | 28.7251 | 33.3207 | 27.9873 | 36.6121 | 24.2936 | -1.4749 | 0.006722 | 0.32724 | 0.0022 | 48.6815 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.5406 | -0.1017 | 0.003692 | 18.7801 | 13.4467 | 22.0715 | 9.753
2012-10-30 20:35 | 28.7169 | 33.2493 | 27.9553 | 36.5802 | 24.2974 | -1.4783 | 0.006754 | 0.32724 | 0.00221 | 48.4508 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4834 | -0.1019 | 0.003702 | 18.7659 | 13.4719 | 22.0968 | 9.814
2012-10-30 20:35 | 28.7154 | 33.2004 | 27.9459 | 36.5353 | 24.2919 | -1.4749 | 0.006735 | 0.32725 | 0.002204 | 48.5896 | 0.006946 | 0.21461 | 0.001491 | 30.8974 | 14.486 | -0.1017 | 0.003695 | 18.7144 | 13.4599 | 22.0493 | 9.8059
2012-10-30 20:35 | 28.7094 | 33.1234 | 27.8803 | 36.4504 | 24.2056 | -1.472 | 0.006738 | 0.32728 | 0.002205 | 48.5723 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4718 | -0.1015 | 0.003696 | 18.6516 | 13.4085 | 21.9786 | 9.7338
2012-10-30 20:35 | 28.7079 | 33.1061 | 27.8589 | 36.3858 | 24.176 | -1.4789 | 0.006719 | 0.32724 | 0.002199 | 48.7032 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.5071 | -0.102 | 0.00369 | 18.599 | 13.3518 | 21.8787 | 9.6689
2012-10-30 20:35 | 28.7139 | 33.0568 | 27.8212 | 36.376 | 24.1178 | -1.4963 | 0.006725 | 0.32726 | 0.002201 | 48.6639 | 0.006955 | 0.21463 | 0.001493 | 30.8595 | 14.5049 | -0.1032 | 0.003694 | 18.5519 | 13.3163 | 21.8711 | 9.6129
2012-10-30 20:35 | 28.7154 | 33.0425 | 27.7909 | 36.2947 | 24.1434 | -1.4743 | 0.006729 | 0.32726 | 0.002202 | 48.635 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5472 | -0.1016 | 0.003693 | 18.4953 | 13.2437 | 21.7475 | 9.5962
2012-10-30 20:35 | 28.7064 | 32.9822 | 27.774 | 36.262 | 24.1384 | -1.4818 | 0.006735 | 0.32716 | 0.002203 | 48.5767 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.5219 | -0.1022 | 0.003695 | 18.4603 | 13.2521 | 21.7401 | 9.6165
2012-10-30 20:36 | 28.7117 | 32.9085 | 27.7475 | 36.2239 | 24.1437 | -1.4835 | 0.006735 | 0.32731 | 0.002204 | 48.5982 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5353 | -0.1023 | 0.003693 | 18.3732 | 13.2122 | 21.6886 | 9.6084
2012-10-30 20:36 | 28.7012 | 32.8704 | 27.6853 | 36.1858 | 24.1132 | -1.4835 | 0.006732 | 0.32732 | 0.002204 | 48.622 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.4963 | -0.1023 | 0.003696 | 18.3741 | 13.189 | 21.6895 | 9.6169
2012-10-30 20:36 | 28.7057 | 32.8946 | 27.6779 | 36.143 | 24.1337 | -1.4639 | 0.006732 | 0.32722 | 0.002203 | 48.607 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.513 | -0.1009 | 0.003693 | 18.3816 | 13.1649 | 21.63 | 9.6207
2012-10-30 20:36 | 28.7094 | 32.8075 | 27.6617 | 36.0914 | 24.1053 | -1.4818 | 0.006738 | 0.32719 | 0.002205 | 48.5594 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.533 | -0.1022 | 0.003695 | 18.2745 | 13.1287 | 21.5584 | 9.5723
2012-10-30 20:36 | 28.7094 | 32.768 | 27.5941 | 36.0283 | 24.1013 | -1.4691 | 0.006732 | 0.32722 | 0.002203 | 48.607 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5045 | -0.1013 | 0.003694 | 18.2635 | 13.0896 | 21.5238 | 9.5968
2012-10-30 20:36 | 28.6982 | 32.7173 | 27.5909 | 35.9579 | 24.062 | -1.4853 | 0.006738 | 0.32719 | 0.002205 | 48.5594 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4647 | -0.1024 | 0.003695 | 18.2526 | 13.1262 | 21.4932 | 9.5973
2012-10-30 20:36 | 28.7027 | 32.6705 | 27.5834 | 35.9111 | 24.0263 | -1.4853 | 0.006738 | 0.32722 | 0.002205 | 48.5637 | 0.006961 | 0.21461 | 0.001494 | 30.8308 | 14.5141 | -0.1024 | 0.003699 | 18.1564 | 13.0693 | 21.397 | 9.5122
2012-10-30 20:36 | 28.7049 | 32.6648 | 27.5618 | 35.8817 | 24.0245 | -1.4835 | 0.006732 | 0.32724 | 0.002203 | 48.6092 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4796 | -0.1023 | 0.003694 | 18.1852 | 13.0822 | 21.4021 | 9.5449
2012-10-30 20:36 | 28.7027 | 32.5521 | 27.488 | 35.8204 | 23.9541 | -1.4801 | 0.006729 | 0.32721 | 0.002202 | 48.6265 | 0.006955 | 0.21466 | 0.001493 | 30.8636 | 14.4692 | -0.102 | 0.003695 | 18.0829 | 13.0188 | 21.3512 | 9.4849
2012-10-30 20:36 | 28.6952 | 32.4459 | 27.4725 | 35.7103 | 23.8984 | -1.4795 | 0.006722 | 0.32724 | 0.0022 | 48.6815 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4699 | -0.102 | 0.003691 | 17.976 | 13.0026 | 21.2404 | 9.4285
2012-10-30 20:36 | 28.6952 | 32.4498 | 27.393 | 35.6984 | 23.8944 | -1.4714 | 0.006735 | 0.32726 | 0.002204 | 48.5917 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.4862 | -0.1014 | 0.003694 | 17.9636 | 12.9068 | 21.2122 | 9.4082
2012-10-30 20:36 | 28.6929 | 32.4278 | 27.4026 | 35.633 | 23.8961 | -1.4743 | 0.006735 | 0.32722 | 0.002204 | 48.5853 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4513 | -0.1016 | 0.003692 | 17.9765 | 12.9513 | 21.1817 | 9.4448
2012-10-30 20:36 | 28.6862 | 32.3895 | 27.38 | 35.6066 | 23.9215 | -1.476 | 0.006747 | 0.32722 | 0.002208 | 48.4989 | 0.006952 | 0.21466 | 0.001492 | 30.8769 | 14.4609 | -0.1018 | 0.0037 | 17.9286 | 12.9191 | 21.1457 | 9.4606
2012-10-30 20:36 | 28.6937 | 32.2983 | 27.3159 | 35.5707 | 23.9089 | -1.4743 | 0.006738 | 0.32729 | 0.002205 | 48.5744 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4725 | -0.1016 | 0.003694 | 17.8258 | 12.8434 | 21.0982 | 9.4364
2012-10-30 20:36 | 28.6847 | 32.2498 | 27.279 | 35.4867 | 23.7955 | -1.4772 | 0.006754 | 0.32728 | 0.00221 | 48.4572 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.4267 | -0.1018 | 0.003701 | 17.8231 | 12.8523 | 21.06 | 9.3688
2012-10-30 20:36 | 28.6974 | 32.2585 | 27.2917 | 35.4797 | 23.756 | -1.4778 | 0.006747 | 0.32719 | 0.002208 | 48.4946 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.4925 | -0.1019 | 0.0037 | 17.766 | 12.7992 | 20.9872 | 9.2635
2012-10-30 20:36 | 28.6944 | 32.1687 | 27.3325 | 35.4727 | 23.7691 | -1.4835 | 0.006751 | 0.32719 | 0.002209 | 48.4659 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.465 | -0.1023 | 0.003702 | 17.7037 | 12.8675 | 21.0077 | 9.3041
2012-10-30 20:36 | 28.6959 | 32.1583 | 27.2663 | 35.4071 | 23.6862 | -1.4772 | 0.006732 | 0.32726 | 0.002203 | 48.6133 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4461 | -0.1018 | 0.003695 | 17.7122 | 12.8202 | 20.961 | 9.2401
2012-10-30 20:36 | 28.6937 | 32.1245 | 27.1925 | 35.3773 | 23.688 | -1.4864 | 0.006751 | 0.32716 | 0.002209 | 48.4616 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.4806 | -0.1025 | 0.003701 | 17.6439 | 12.7119 | 20.8967 | 9.2074
2012-10-30 20:36 | 28.6877 | 32.0988 | 27.1785 | 35.3082 | 23.5494 | -1.4789 | 0.006735 | 0.32731 | 0.002204 | 48.5982 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.4542 | -0.102 | 0.003696 | 17.6446 | 12.7243 | 20.854 | 9.0952
2012-10-30 20:36 | 28.6884 | 32.0561 | 27.1354 | 35.297 | 23.5381 | -1.4824 | 0.006744 | 0.32728 | 0.002207 | 48.5291 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4386 | -0.1022 | 0.003697 | 17.6175 | 12.6968 | 20.8584 | 9.0995
2012-10-30 20:36 | 28.6839 | 31.9489 | 27.111 | 35.2057 | 23.5818 | -1.4795 | 0.006738 | 0.32722 | 0.002205 | 48.5637 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4749 | -0.102 | 0.003694 | 17.474 | 12.6361 | 20.7308 | 9.1069
2012-10-30 20:36 | 28.6929 | 31.954 | 27.0602 | 35.1792 | 23.5827 | -1.4899 | 0.006747 | 0.32722 | 0.002208 | 48.4989 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.439 | -0.1027 | 0.003696 | 17.515 | 12.6212 | 20.7402 | 9.1437
2012-10-30 20:36 | 28.6929 | 31.8987 | 27.0682 | 35.1318 | 23.5908 | -1.4772 | 0.006735 | 0.32726 | 0.002204 | 48.5917 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4553 | -0.1018 | 0.003695 | 17.4434 | 12.6129 | 20.6765 | 9.1355
2012-10-30 20:36 | 28.6959 | 31.878 | 27.0672 | 35.0875 | 23.6259 | -1.4812 | 0.006735 | 0.32722 | 0.002204 | 48.5853 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.4951 | -0.1021 | 0.003696 | 17.3829 | 12.5721 | 20.5924 | 9.1308
2012-10-30 20:37 | 28.6922 | 31.8151 | 27.0356 | 35.0759 | 23.6102 | -1.4737 | 0.006738 | 0.32722 | 0.002205 | 48.5637 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4791 | -0.1016 | 0.003695 | 17.336 | 12.5565 | 20.5968 | 9.1311
2012-10-30 20:37 | 28.6922 | 31.7637 | 26.948 | 35.0325 | 23.5539 | -1.4801 | 0.006757 | 0.32719 | 0.002211 | 48.4228 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4587 | -0.102 | 0.003703 | 17.305 | 12.4893 | 20.5738 | 9.0952
2012-10-30 20:37 | 28.7042 | 31.7481 | 27.0158 | 34.9734 | 23.5619 | -1.4824 | 0.006732 | 0.32724 | 0.002203 | 48.6092 | 0.006958 | 0.21467 | 0.001494 | 30.8524 | 14.483 | -0.1022 | 0.003697 | 17.2651 | 12.5328 | 20.4904 | 9.0789
2012-10-30 20:37 | 28.6982 | 31.6908 | 26.9859 | 34.9161 | 23.4795 | -1.483 | 0.006744 | 0.32721 | 0.002207 | 48.5184 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4566 | -0.1022 | 0.003696 | 17.2342 | 12.5293 | 20.4595 | 9.0229
2012-10-30 20:37 | 28.6952 | 31.672 | 26.8912 | 34.8697 | 23.4323 | -1.4858 | 0.006754 | 0.32724 | 0.00221 | 48.4508 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4862 | -0.1024 | 0.0037 | 17.1858 | 12.405 | 20.3835 | 8.9461
2012-10-30 20:37 | 28.7027 | 31.6282 | 26.9107 | 34.8693 | 23.4318 | -1.4806 | 0.006729 | 0.32722 | 0.002202 | 48.6286 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.4733 | -0.1021 | 0.003693 | 17.1549 | 12.4374 | 20.396 | 8.9585
2012-10-30 20:37 | 28.6974 | 31.5478 | 26.8655 | 34.8443 | 23.3863 | -1.4754 | 0.006732 | 0.32725 | 0.002203 | 48.6113 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.4802 | -0.1017 | 0.003691 | 17.0676 | 12.3853 | 20.3641 | 8.9061
2012-10-30 20:37 | 28.7019 | 31.5207 | 26.87 | 34.7896 | 23.4189 | -1.4963 | 0.006744 | 0.32722 | 0.002207 | 48.5205 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4807 | -0.1032 | 0.003698 | 17.04 | 12.3893 | 20.3089 | 8.9382
2012-10-30 20:37 | 28.6997 | 31.4633 | 26.828 | 34.744 | 23.3886 | -1.4783 | 0.006732 | 0.32719 | 0.002203 | 48.6026 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.5029 | -0.1019 | 0.003695 | 16.9604 | 12.3251 | 20.2411 | 8.8857
2012-10-30 20:37 | 28.7019 | 31.4497 | 26.8023 | 34.6988 | 23.3264 | -1.4812 | 0.006754 | 0.32719 | 0.00221 | 48.4443 | 0.006927 | 0.21463 | 0.001487 | 30.9843 | 14.5215 | -0.1021 | 0.003697 | 16.9282 | 12.2808 | 20.1773 | 8.8049
2012-10-30 20:37 | 28.6944 | 31.4066 | 26.7589 | 34.6361 | 23.2828 | -1.4778 | 0.006744 | 0.32714 | 0.002206 | 48.5076 | 0.006949 | 0.21458 | 0.001491 | 30.8798 | 14.4324 | -0.1019 | 0.003697 | 16.9742 | 12.3265 | 20.2037 | 8.8504
2012-10-30 20:37 | 28.6989 | 31.344 | 26.7554 | 34.6366 | 23.2993 | -1.4702 | 0.006751 | 0.32722 | 0.002209 | 48.4702 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.5021 | -0.1014 | 0.003698 | 16.8419 | 12.2533 | 20.1345 | 8.7972
2012-10-30 20:37 | 28.6982 | 31.3157 | 26.7547 | 34.6004 | 23.3147 | -1.4835 | 0.006741 | 0.32716 | 0.002205 | 48.5335 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4484 | -0.1023 | 0.003693 | 16.8673 | 12.3063 | 20.152 | 8.8663
2012-10-30 20:37 | 28.7079 | 31.2898 | 26.6887 | 34.5746 | 23.2601 | -1.4841 | 0.006741 | 0.32724 | 0.002206 | 48.5443 | 0.006952 | 0.21466 | 0.001492 | 30.8769 | 14.4744 | -0.1023 | 0.003698 | 16.8154 | 12.2143 | 20.1002 | 8.7857
2012-10-30 20:37 | 28.6877 | 31.2381 | 26.6525 | 34.4952 | 23.2439 | -1.4691 | 0.006751 | 0.32726 | 0.002209 | 48.4765 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4828 | -0.1013 | 0.003698 | 16.7553 | 12.1697 | 20.0124 | 8.7611
2012-10-30 20:37 | 28.7072 | 31.2062 | 26.6441 | 34.4712 | 23.2674 | -1.4754 | 0.006738 | 0.32724 | 0.002205 | 48.5659 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4778 | -0.1017 | 0.003694 | 16.7284 | 12.1663 | 19.9934 | 8.7896
2012-10-30 20:37 | 28.6997 | 31.1592 | 26.6047 | 34.4401 | 23.2519 | -1.4841 | 0.006754 | 0.32719 | 0.00221 | 48.4443 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.5233 | -0.1023 | 0.0037 | 16.6359 | 12.0814 | 19.9168 | 8.7286
2012-10-30 20:37 | 28.6944 | 31.1184 | 26.5834 | 34.3913 | 23.2023 | -1.4789 | 0.006735 | 0.32719 | 0.002204 | 48.581 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5221 | -0.102 | 0.003695 | 16.5963 | 12.0613 | 19.8692 | 8.6802
2012-10-30 20:37 | 28.6952 | 31.0955 | 26.5842 | 34.3724 | 23.1991 | -1.4812 | 0.006738 | 0.32718 | 0.002205 | 48.5573 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.4658 | -0.1021 | 0.003698 | 16.6297 | 12.1184 | 19.9066 | 8.7333
2012-10-30 20:37 | 28.6937 | 31.0229 | 26.5748 | 34.3393 | 23.2217 | -1.4783 | 0.006729 | 0.32719 | 0.002202 | 48.6243 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.4643 | -0.1019 | 0.003694 | 16.5586 | 12.1105 | 19.875 | 8.7574
2012-10-30 20:37 | 28.6967 | 31.0773 | 26.5339 | 34.3028 | 23.1926 | -1.4801 | 0.006732 | 0.32718 | 0.002203 | 48.6006 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.451 | -0.102 | 0.003695 | 16.6263 | 12.0829 | 19.8518 | 8.7416
2012-10-30 20:37 | 28.6922 | 30.978 | 26.4456 | 34.2391 | 23.1478 | -1.4818 | 0.006741 | 0.32724 | 0.002206 | 48.5443 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.471 | -0.1022 | 0.003697 | 16.507 | 11.9746 | 19.7681 | 8.6768
2012-10-30 20:37 | 28.6922 | 30.978 | 26.4336 | 34.1957 | 23.1559 | -1.4766 | 0.006738 | 0.32726 | 0.002205 | 48.5701 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.4954 | -0.1018 | 0.003698 | 16.4826 | 11.9382 | 19.7003 | 8.6605
2012-10-30 20:37 | 28.6892 | 30.9355 | 26.4306 | 34.1651 | 23.1328 | -1.4754 | 0.006741 | 0.32724 | 0.002206 | 48.5443 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.472 | -0.1017 | 0.003696 | 16.4635 | 11.9586 | 19.6931 | 8.6608
2012-10-30 20:37 | 28.6952 | 30.8863 | 26.4366 | 34.1277 | 23.0945 | -1.4795 | 0.006747 | 0.32722 | 0.002208 | 48.4989 | 0.006965 | 0.21471 | 0.001495 | 30.8276 | 14.4658 | -0.102 | 0.003703 | 16.4205 | 11.9708 | 19.6619 | 8.6287
2012-10-30 20:37 | 28.7004 | 30.8599 | 26.394 | 34.0855 | 23.0796 | -1.4685 | 0.006735 | 0.32724 | 0.002204 | 48.5875 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4996 | -0.1012 | 0.003693 | 16.3603 | 11.8944 | 19.5859 | 8.58
2012-10-30 20:37 | 28.6959 | 30.7803 | 26.3815 | 34.0336 | 23.051 | -1.4766 | 0.006732 | 0.32721 | 0.002203 | 48.6049 | 0.006965 | 0.21468 | 0.001495 | 30.8234 | 14.4665 | -0.1018 | 0.003698 | 16.3138 | 11.915 | 19.5671 | 8.5845
2012-10-30 20:38 | 28.6952 | 30.7994 | 26.3169 | 34.0369 | 23.0382 | -1.476 | 0.006735 | 0.32719 | 0.002204 | 48.581 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4699 | -0.1018 | 0.003695 | 16.3295 | 11.847 | 19.567 | 8.5683
2012-10-30 20:38 | 28.7049 | 30.7459 | 26.3146 | 33.9518 | 23.072 | -1.4737 | 0.006751 | 0.32722 | 0.002209 | 48.4702 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.451 | -0.1016 | 0.003699 | 16.2949 | 11.8636 | 19.5008 | 8.621
2012-10-30 20:38 | 28.6982 | 30.6997 | 26.268 | 33.9136 | 23.0372 | -1.4685 | 0.006735 | 0.32721 | 0.002204 | 48.5832 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.428 | -0.1012 | 0.003697 | 16.2717 | 11.84 | 19.4856 | 8.6092
2012-10-30 20:38 | 28.6989 | 30.6491 | 26.2687 | 33.9024 | 22.9775 | -1.4656 | 0.006725 | 0.32716 | 0.0022 | 48.6489 | 0.006958 | 0.21474 | 0.001494 | 30.8627 | 14.4409 | -0.101 | 0.003694 | 16.2082 | 11.8278 | 19.4615 | 8.5366
2012-10-30 20:38 | 28.7034 | 30.5825 | 26.2413 | 33.8674 | 22.9096 | -1.4702 | 0.006735 | 0.32721 | 0.002204 | 48.5832 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4046 | -0.1014 | 0.003695 | 16.1779 | 11.8367 | 19.4628 | 8.505
2012-10-30 20:38 | 28.6914 | 30.6021 | 26.2173 | 33.816 | 22.8895 | -1.4835 | 0.006735 | 0.32711 | 0.002203 | 48.5681 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.4375 | -0.1023 | 0.003695 | 16.1646 | 11.7798 | 19.3785 | 8.452
2012-10-30 20:38 | 28.6997 | 30.5196 | 26.1697 | 33.7493 | 22.9099 | -1.4708 | 0.006732 | 0.32714 | 0.002202 | 48.594 | 0.006955 | 0.21463 | 0.001493 | 30.8595 | 14.405 | -0.1014 | 0.003695 | 16.1146 | 11.7647 | 19.3443 | 8.5049
2012-10-30 20:38 | 28.6899 | 30.498 | 26.1559 | 33.7237 | 22.9001 | -1.4726 | 0.006744 | 0.32708 | 0.002206 | 48.4991 | 0.006939 | 0.21458 | 0.001489 | 30.9243 | 14.4115 | -0.1015 | 0.003695 | 16.0865 | 11.7444 | 19.3122 | 8.4886
2012-10-30 20:38 | 28.6914 | 30.5034 | 26.1534 | 33.7607 | 22.8895 | -1.476 | 0.006738 | 0.32705 | 0.002204 | 48.538 | 0.006952 | 0.21457 | 0.001492 | 30.8645 | 14.413 | -0.1018 | 0.003696 | 16.0904 | 11.7404 | 19.3477 | 8.4765
2012-10-30 20:38 | 28.6884 | 30.4293 | 26.1105 | 33.6511 | 22.8946 | -1.4899 | 0.006735 | 0.32714 | 0.002203 | 48.5724 | 0.006936 | 0.21461 | 0.001489 | 30.9419 | 14.4304 | -0.1027 | 0.003692 | 15.9989 | 11.6801 | 19.2207 | 8.4642
2012-10-30 20:38 | 28.6982 | 30.447 | 26.1123 | 33.6491 | 22.9205 | -1.4882 | 0.006744 | 0.32706 | 0.002206 | 48.4969 | 0.006939 | 0.2146 | 0.001489 | 30.9265 | 14.4239 | -0.1026 | 0.003695 | 16.0231 | 11.6884 | 19.2252 | 8.4966
2012-10-30 20:38 | 28.6982 | 30.4155 | 26.0763 | 33.6017 | 22.8642 | -1.4835 | 0.006732 | 0.32711 | 0.002202 | 48.5897 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.3913 | -0.1023 | 0.003695 | 16.0242 | 11.685 | 19.2104 | 8.4729
2012-10-30 20:38 | 28.6959 | 30.3658 | 26.0381 | 33.5599 | 22.7975 | -1.4766 | 0.006735 | 0.32716 | 0.002203 | 48.5767 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4298 | -0.1018 | 0.003695 | 15.936 | 11.6083 | 19.1301 | 8.3677
2012-10-30 20:38 | 28.6967 | 30.331 | 26.0429 | 33.5449 | 22.7499 | -1.4783 | 0.006744 | 0.32714 | 0.002206 | 48.5076 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.4347 | -0.1019 | 0.003694 | 15.8963 | 11.6082 | 19.1102 | 8.3152
2012-10-30 20:38 | 28.6944 | 30.2656 | 25.9886 | 33.5071 | 22.7436 | -1.4853 | 0.006744 | 0.32706 | 0.002206 | 48.4969 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4324 | -0.1024 | 0.003697 | 15.8332 | 11.5562 | 19.0747 | 8.3112
2012-10-30 20:38 | 28.6869 | 30.2462 | 25.9172 | 33.4443 | 22.7522 | -1.4887 | 0.006741 | 0.32715 | 0.002205 | 48.5314 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4493 | -0.1026 | 0.003695 | 15.7969 | 11.4679 | 18.995 | 8.3029
2012-10-30 20:38 | 28.6839 | 30.2195 | 25.9462 | 33.3979 | 22.7291 | -1.4864 | 0.006738 | 0.32718 | 0.002205 | 48.5573 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4341 | -0.1025 | 0.003697 | 15.7854 | 11.5121 | 18.9638 | 8.295
2012-10-30 20:38 | 28.6884 | 30.1372 | 25.8508 | 33.3511 | 22.7336 | -1.487 | 0.006735 | 0.32712 | 0.002203 | 48.5703 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.4264 | -0.1025 | 0.003695 | 15.7108 | 11.4244 | 18.9247 | 8.3072
2012-10-30 20:38 | 28.6854 | 30.0987 | 25.8678 | 33.3046 | 22.7306 | -1.487 | 0.006735 | 0.32714 | 0.002203 | 48.5724 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4601 | -0.1025 | 0.003694 | 15.6386 | 11.4077 | 18.8445 | 8.2705
2012-10-30 20:38 | 28.6877 | 30.1128 | 25.906 | 33.303 | 22.6725 | -1.4835 | 0.006741 | 0.32712 | 0.002205 | 48.5271 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4134 | -0.1023 | 0.003695 | 15.6994 | 11.4926 | 18.8896 | 8.2591
2012-10-30 20:38 | 28.6869 | 30.0804 | 25.8533 | 33.2746 | 22.6878 | -1.498 | 0.006738 | 0.32715 | 0.002204 | 48.553 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4167 | -0.1033 | 0.003693 | 15.6637 | 11.4366 | 18.8579 | 8.2711
2012-10-30 20:38 | 28.6839 | 30.034 | 25.8063 | 33.2123 | 22.6888 | -1.498 | 0.006738 | 0.32709 | 0.002204 | 48.5444 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.4137 | -0.1033 | 0.003694 | 15.6203 | 11.3926 | 18.7986 | 8.2751
2012-10-30 20:38 | 28.6817 | 29.9884 | 25.7881 | 33.1588 | 22.6826 | -1.5009 | 0.006729 | 0.327 | 0.0022 | 48.5964 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.4319 | -0.1035 | 0.003691 | 15.5565 | 11.3562 | 18.7269 | 8.2507
2012-10-30 20:38 | 28.6742 | 29.9256 | 25.7646 | 33.1434 | 22.6187 | -1.4853 | 0.006732 | 0.32705 | 0.002202 | 48.5813 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.4326 | -0.1024 | 0.003694 | 15.493 | 11.332 | 18.7108 | 8.1861
2012-10-30 20:38 | 28.675 | 29.9106 | 25.7734 | 33.0692 | 22.6155 | -1.4876 | 0.006741 | 0.32706 | 0.002205 | 48.5185 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.4293 | -0.1026 | 0.003696 | 15.4813 | 11.3441 | 18.6399 | 8.1862
2012-10-30 20:38 | 28.6577 | 29.8972 | 25.7281 | 33.048 | 22.6022 | -1.4887 | 0.006738 | 0.32699 | 0.002203 | 48.5294 | 0.006946 | 0.21461 | 0.001491 | 30.8974 | 14.4201 | -0.1026 | 0.003694 | 15.4771 | 11.308 | 18.6279 | 8.1821
2012-10-30 20:39 | 28.6787 | 29.859 | 25.7132 | 33.0137 | 22.6152 | -1.4876 | 0.006732 | 0.32699 | 0.002201 | 48.5726 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4575 | -0.1026 | 0.003689 | 15.4015 | 11.2557 | 18.5562 | 8.1577
2012-10-30 20:39 | 28.6735 | 29.8143 | 25.676 | 32.9809 | 22.5899 | -1.4806 | 0.006744 | 0.32709 | 0.002206 | 48.5012 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4278 | -0.1021 | 0.003696 | 15.3865 | 11.2482 | 18.5531 | 8.1621
2012-10-30 20:39 | 28.6697 | 29.7632 | 25.7002 | 32.906 | 22.5095 | -1.4899 | 0.006754 | 0.32706 | 0.002209 | 48.4251 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4485 | -0.1027 | 0.003698 | 15.3147 | 11.2517 | 18.4575 | 8.061
2012-10-30 20:39 | 28.666 | 29.7516 | 25.6445 | 32.8825 | 22.522 | -1.4893 | 0.006738 | 0.32716 | 0.002204 | 48.5551 | 0.006955 | 0.21468 | 0.001493 | 30.8677 | 14.4244 | -0.1027 | 0.003697 | 15.3272 | 11.2201 | 18.4581 | 8.0976
2012-10-30 20:39 | 28.6757 | 29.7218 | 25.6502 | 32.8765 | 22.5035 | -1.4778 | 0.006735 | 0.32714 | 0.002203 | 48.5724 | 0.006952 | 0.21466 | 0.001492 | 30.8769 | 14.4626 | -0.1019 | 0.003695 | 15.2592 | 11.1876 | 18.4139 | 8.0409
2012-10-30 20:39 | 28.6787 | 29.6695 | 25.5692 | 32.8044 | 22.4984 | -1.4772 | 0.006744 | 0.32709 | 0.002206 | 48.5012 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4615 | -0.1018 | 0.003696 | 15.208 | 11.1077 | 18.3429 | 8.0369
2012-10-30 20:39 | 28.678 | 29.6254 | 25.5845 | 32.7998 | 22.4937 | -1.498 | 0.006738 | 0.32709 | 0.002204 | 48.5444 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.3996 | -0.1033 | 0.003695 | 15.2258 | 11.1849 | 18.4002 | 8.0941
2012-10-30 20:39 | 28.663 | 29.567 | 25.5295 | 32.69 | 22.4585 | -1.483 | 0.006747 | 0.32716 | 0.002207 | 48.4903 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.4295 | -0.1022 | 0.003698 | 15.1375 | 11.1 | 18.2605 | 8.029
2012-10-30 20:39 | 28.681 | 29.5889 | 25.4955 | 32.6843 | 22.4202 | -1.4702 | 0.006747 | 0.32714 | 0.002207 | 48.486 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4149 | -0.1014 | 0.003699 | 15.174 | 11.0806 | 18.2694 | 8.0053
2012-10-30 20:39 | 28.672 | 29.5089 | 25.4985 | 32.6635 | 22.4152 | -1.4824 | 0.006741 | 0.32711 | 0.002205 | 48.5248 | 0.006955 | 0.21466 | 0.001493 | 30.8636 | 14.3936 | -0.1022 | 0.003698 | 15.1153 | 11.1049 | 18.2699 | 8.0216
2012-10-30 20:39 | 28.6712 | 29.4923 | 25.4217 | 32.6311 | 22.366 | -1.4835 | 0.006732 | 0.32714 | 0.002202 | 48.594 | 0.006955 | 0.21466 | 0.001493 | 30.8636 | 14.4418 | -0.1023 | 0.003695 | 15.0505 | 10.9799 | 18.1893 | 7.9242
2012-10-30 20:39 | 28.6697 | 29.4276 | 25.4522 | 32.5822 | 22.3202 | -1.4858 | 0.006741 | 0.32709 | 0.002205 | 48.5228 | 0.006955 | 0.21463 | 0.001493 | 30.8595 | 14.4199 | -0.1024 | 0.003698 | 15.0077 | 11.0323 | 18.1623 | 7.9003
2012-10-30 20:39 | 28.6682 | 29.3906 | 25.4027 | 32.5649 | 22.3429 | -1.4824 | 0.006744 | 0.32718 | 0.002206 | 48.5141 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.447 | -0.1022 | 0.003697 | 14.9436 | 10.9557 | 18.1179 | 7.8959
2012-10-30 20:39 | 28.675 | 29.4171 | 25.3895 | 32.4612 | 22.3497 | -1.4841 | 0.006732 | 0.32712 | 0.002202 | 48.5919 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.4048 | -0.1023 | 0.003694 | 15.0123 | 10.9847 | 18.0564 | 7.9449
2012-10-30 20:39 | 28.681 | 29.3126 | 25.4035 | 32.4751 | 22.3194 | -1.4968 | 0.006741 | 0.32716 | 0.002205 | 48.5335 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.472 | -0.1032 | 0.003696 | 14.8406 | 10.9315 | 18.0031 | 7.8474
2012-10-30 20:39 | 28.6735 | 29.3169 | 25.328 | 32.3926 | 22.2636 | -1.5003 | 0.006744 | 0.32715 | 0.002206 | 48.5098 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4033 | -0.1034 | 0.003697 | 14.9136 | 10.9247 | 17.9893 | 7.8603
2012-10-30 20:39 | 28.6854 | 29.2538 | 25.2958 | 32.3729 | 22.2956 | -1.4876 | 0.006725 | 0.32706 | 0.002199 | 48.6339 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.4397 | -0.1026 | 0.003688 | 14.8141 | 10.8561 | 17.9332 | 7.8559
2012-10-30 20:39 | 28.6697 | 29.2263 | 25.2121 | 32.3532 | 22.2638 | -1.4778 | 0.006751 | 0.32716 | 0.002209 | 48.4616 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4158 | -0.1019 | 0.003698 | 14.8105 | 10.7963 | 17.9374 | 7.848
2012-10-30 20:39 | 28.6742 | 29.211 | 25.2606 | 32.3064 | 22.2603 | -1.4731 | 0.006744 | 0.32711 | 0.002206 | 48.5033 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4693 | -0.1016 | 0.003698 | 14.7417 | 10.7913 | 17.8371 | 7.791
2012-10-30 20:39 | 28.6667 | 29.1325 | 25.2091 | 32.2357 | 22.2487 | -1.4795 | 0.006744 | 0.32716 | 0.002206 | 48.5119 | 0.006958 | 0.21468 | 0.001494 | 30.8544 | 14.4169 | -0.102 | 0.0037 | 14.7156 | 10.7922 | 17.8188 | 7.8318
2012-10-30 20:39 | 28.6727 | 29.1424 | 25.2271 | 32.218 | 22.2185 | -1.4778 | 0.006729 | 0.32711 | 0.002201 | 48.6114 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4188 | -0.1019 | 0.00369 | 14.7236 | 10.8083 | 17.7992 | 7.7997
2012-10-30 20:39 | 28.6712 | 29.1172 | 25.1655 | 32.1534 | 22.225 | -1.4772 | 0.006735 | 0.32718 | 0.002204 | 48.5789 | 0.006958 | 0.2146 | 0.001493 | 30.8421 | 14.4214 | -0.1018 | 0.003697 | 14.6958 | 10.7441 | 17.732 | 7.8036
2012-10-30 20:39 | 28.681 | 29.0402 | 25.1994 | 32.1434 | 22.2227 | -1.4945 | 0.006729 | 0.32709 | 0.002201 | 48.6093 | 0.006946 | 0.21461 | 0.001491 | 30.8974 | 14.4434 | -0.103 | 0.003692 | 14.5968 | 10.756 | 17.7 | 7.7793
2012-10-30 20:39 | 28.6735 | 29.009 | 25.1518 | 32.1083 | 22.1467 | -1.4911 | 0.006738 | 0.32712 | 0.002204 | 48.5487 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4319 | -0.1028 | 0.003694 | 14.5771 | 10.7199 | 17.6764 | 7.7148
2012-10-30 20:39 | 28.6772 | 29.0009 | 25.0955 | 32.0686 | 22.1585 | -1.4818 | 0.006735 | 0.32718 | 0.002204 | 48.5789 | 0.006955 | 0.21468 | 0.001493 | 30.8677 | 14.4233 | -0.1022 | 0.003697 | 14.5776 | 10.6722 | 17.6453 | 7.7352
2012-10-30 20:40 | 28.678 | 28.9583 | 25.0562 | 32.022 | 22.1432 | -1.4893 | 0.006744 | 0.32705 | 0.002206 | 48.4948 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.4078 | -0.1027 | 0.003699 | 14.5505 | 10.6484 | 17.6142 | 7.7354
2012-10-30 20:40 | 28.6727 | 28.9451 | 25.087 | 32.0127 | 22.0895 | -1.4934 | 0.006751 | 0.32709 | 0.002208 | 48.4509 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4188 | -0.103 | 0.003697 | 14.5263 | 10.6682 | 17.5939 | 7.6707
2012-10-30 20:40 | 28.6757 | 28.8967 | 25.0299 | 31.917 | 22.1328 | -1.4847 | 0.006738 | 0.32702 | 0.002203 | 48.5337 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4014 | -0.1024 | 0.003694 | 14.4953 | 10.6285 | 17.5156 | 7.7314
2012-10-30 20:40 | 28.6765 | 28.8265 | 24.9826 | 31.8704 | 22.1054 | -1.4841 | 0.006735 | 0.32705 | 0.002203 | 48.5596 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.39 | -0.1023 | 0.003695 | 14.4365 | 10.5926 | 17.4804 | 7.7154
2012-10-30 20:40 | 28.6787 | 28.8326 | 24.9768 | 31.849 | 22.0875 | -1.4789 | 0.006735 | 0.32703 | 0.002203 | 48.5574 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.433 | -0.102 | 0.003693 | 14.3996 | 10.5438 | 17.416 | 7.6545
2012-10-30 20:40 | 28.6824 | 28.8008 | 24.9645 | 31.8487 | 22.0992 | -1.4818 | 0.006738 | 0.32706 | 0.002204 | 48.5401 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.3877 | -0.1022 | 0.003695 | 14.4131 | 10.5768 | 17.461 | 7.7115
2012-10-30 20:40 | 28.6787 |  | 24.9528 | 31.77 | 22.023 | -1.4824 | 0.006738 | 0.32708 | 0.002204 | 48.5423 | 0.006943 | 0.21458 | 0.00149 | 30.9065 | 14.3922 | -0.1022 | 0.003694 | -14.3922 | 10.5606 | 17.3778 | 7.6308
2012-10-30 20:40 | 28.675 |  | 24.889 | 31.7426 | 22.0031 | -1.4662 | 0.006738 | 0.32699 | 0.002203 | 48.5294 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.3558 | -0.1011 | 0.003694 | -14.3558 | 10.5332 | 17.3868 | 7.6473
2012-10-30 20:40 | 28.675 |  | 24.8569 | 31.6636 | 21.9991 | -1.4887 | 0.006735 | 0.32703 | 0.002203 | 48.5574 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.3558 | -0.1026 | 0.003691 | -14.3558 | 10.5011 | 17.3078 | 7.6433
2012-10-30 20:40 | 28.6757 |  | 24.8777 | 31.6762 | 22.0119 | -1.4812 | 0.006754 | 0.32705 | 0.002209 | 48.423 | 0.006955 | 0.21466 | 0.001493 | 30.8636 | 14.3933 | -0.1021 | 0.003702 | -14.3933 | 10.4844 | 17.2829 | 7.6186
2012-10-30 20:40 | 28.6735 |  | 24.8354 | 31.674 | 21.9452 | -1.4749 | 0.006725 | 0.32708 | 0.0022 | 48.6361 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.4074 | -0.1017 | 0.003693 | -14.4074 | 10.428 | 17.2666 | 7.5378
2012-10-30 20:40 | 28.6795 | 28.5963 | 24.8134 | 31.5892 | 21.9633 | -1.4887 | 0.006741 | 0.32706 | 0.002205 | 48.5185 | 0.006933 | 0.21461 | 0.001488 | 30.9553 | 14.4052 | -0.1026 | 0.003693 | 14.1911 | 10.4082 | 17.184 | 7.5581
2012-10-30 20:40 | 28.6854 | 28.5824 | 24.7992 | 31.5161 | 21.8845 | -1.4772 | 0.006741 | 0.32703 | 0.002205 | 48.5142 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4315 | -0.1018 | 0.003697 | 14.1509 | 10.3677 | 17.0846 | 7.453
2012-10-30 20:40 | 28.6735 | 28.515 | 24.7392 | 31.4963 | 21.9008 | -1.4858 | 0.006747 | 0.32711 | 0.002207 | 48.4817 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.3462 | -0.1024 | 0.003699 | 14.1688 | 10.393 | 17.1501 | 7.5546
2012-10-30 20:40 | 28.6862 | 28.5079 | 24.772 | 31.4419 | 21.9498 | -1.4824 | 0.006732 | 0.32706 | 0.002202 | 48.5833 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.3874 | -0.1022 | 0.003693 | 14.1205 | 10.3846 | 17.0545 | 7.5624
2012-10-30 20:40 | 28.681 | 28.4987 | 24.7507 | 31.4327 | 21.9325 | -1.4668 | 0.006741 | 0.32714 | 0.002205 | 48.5292 | 0.006927 | 0.21461 | 0.001487 | 30.9821 | 14.4108 | -0.1011 | 0.003692 | 14.0879 | 10.3399 | 17.0219 | 7.5217
2012-10-30 20:40 | 28.6765 | 28.4387 | 24.6861 | 31.3769 | 21.8958 | -1.476 | 0.006735 | 0.32702 | 0.002202 | 48.5553 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.3941 | -0.1018 | 0.003694 | 14.0446 | 10.292 | 16.9828 | 7.5017
2012-10-30 20:40 | 28.678 | 28.4085 | 24.6756 | 31.3784 | 21.8691 | -1.4824 | 0.006751 | 0.32708 | 0.002208 | 48.4488 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.3548 | -0.1022 | 0.0037 | 14.0537 | 10.3208 | 17.0236 | 7.5143
2012-10-30 20:40 | 28.6869 | 28.3976 | 24.6644 | 31.3044 | 21.882 | -1.4783 | 0.006751 | 0.32702 | 0.002208 | 48.4402 | 0.006952 | 0.21458 | 0.001492 | 30.8665 | 14.3963 | -0.1019 | 0.0037 | 14.0013 | 10.2681 | 16.9081 | 7.4857
2012-10-30 20:40 | 28.6899 | 28.3411 | 24.6474 | 31.256 | 21.8447 | -1.4772 | 0.006741 | 0.32705 | 0.002205 | 48.5164 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.4523 | -0.1018 | 0.003696 | 13.8888 | 10.1951 | 16.8037 | 7.3924
2012-10-30 20:40 | 28.6907 | 28.2903 | 24.6442 | 31.2134 | 21.8092 | -1.4887 | 0.006744 | 0.32706 | 0.002206 | 48.4969 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4123 | -0.1026 | 0.003696 | 13.878 | 10.2319 | 16.8011 | 7.3969
2012-10-30 20:40 | 28.6899 | 28.2459 | 24.5672 | 31.2205 | 21.76 | -1.4876 | 0.006744 | 0.32703 | 0.002206 | 48.4926 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4442 | -0.1026 | 0.003694 | 13.8017 | 10.123 | 16.7763 | 7.3158
2012-10-30 20:40 | 28.6937 | 28.2338 | 24.571 | 31.1848 | 21.8202 | -1.4876 | 0.006741 | 0.32711 | 0.002205 | 48.5248 | 0.00693 | 0.21463 | 0.001487 | 30.9709 | 14.4235 | -0.1026 | 0.003692 | 13.8103 | 10.1475 | 16.7613 | 7.3967
2012-10-30 20:40 | 28.6877 | 28.204 | 24.569 | 31.1236 | 21.6892 | -1.4656 | 0.006729 | 0.32706 | 0.002201 | 48.605 | 0.006965 | 0.21455 | 0.001494 | 30.8047 | 14.4379 | -0.101 | 0.003695 | 13.7661 | 10.1311 | 16.6857 | 7.2513
2012-10-30 20:41 | 28.6892 | 28.1619 | 24.5304 | 31.1251 | 21.7109 | -1.4841 | 0.006738 | 0.32711 | 0.002204 | 48.5465 | 0.006955 | 0.21463 | 0.001493 | 30.8595 | 14.4435 | -0.1023 | 0.003697 | 13.7184 | 10.0869 | 16.6816 | 7.2674
2012-10-30 20:41 | 28.6892 | 28.138 | 24.5425 | 31.0579 | 21.7028 | -1.476 | 0.006744 | 0.32712 | 0.002206 | 48.5055 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.4353 | -0.1018 | 0.003698 | 13.7027 | 10.1072 | 16.6226 | 7.2675
2012-10-30 20:41 | 28.6854 | 28.0866 | 24.5186 | 31.0107 | 21.7232 | -1.4772 | 0.006722 | 0.32715 | 0.002199 | 48.6686 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.4478 | -0.1018 | 0.003692 | 13.6388 | 10.0708 | 16.5629 | 7.2754
2012-10-30 20:41 | 28.6944 | 28.0559 | 24.5276 | 30.9684 | 21.7161 | -1.4795 | 0.006735 | 0.32705 | 0.002203 | 48.5596 | 0.006955 | 0.21461 | 0.001493 | 30.8574 | 14.4446 | -0.102 | 0.003696 | 13.6113 | 10.083 | 16.5238 | 7.2715
2012-10-30 20:41 | 28.6854 | 27.9874 | 24.4785 | 30.9752 | 21.6587 | -1.4835 | 0.006744 | 0.32709 | 0.002206 | 48.5012 | 0.006958 | 0.2146 | 0.001493 | 30.8421 | 14.4519 | -0.1023 | 0.003699 | 13.5355 | 10.0266 | 16.5233 | 7.2068
2012-10-30 20:41 | 28.6847 | 28.0065 | 24.4738 | 30.9468 | 21.6862 | -1.4789 | 0.006754 | 0.32708 | 0.002209 | 48.4273 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4553 | -0.102 | 0.003699 | 13.5512 | 10.0185 | 16.4915 | 7.2309
2012-10-30 20:41 | 28.6854 | 27.9715 | 24.4264 | 30.8922 | 21.6304 | -1.4812 | 0.006722 | 0.32703 | 0.002198 | 48.6513 | 0.006952 | 0.21463 | 0.001492 | 30.8728 | 14.456 | -0.1021 | 0.00369 | 13.5155 | 9.9704 | 16.4362 | 7.1744
2012-10-30 20:41 | 28.7019 | 27.9761 | 24.4669 | 30.8574 | 21.655 | -1.487 | 0.006732 | 0.32714 | 0.002202 | 48.594 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4929 | -0.1025 | 0.003693 | 13.4832 | 9.974 | 16.3645 | 7.1621
2012-10-30 20:41 | 28.6899 | 27.9323 | 24.4228 | 30.8415 | 21.6067 | -1.472 | 0.006735 | 0.32711 | 0.002203 | 48.5681 | 0.006952 | 0.21464 | 0.001492 | 30.8749 | 14.4523 | -0.1015 | 0.003695 | 13.48 | 9.9705 | 16.3892 | 7.1544
2012-10-30 20:41 | 28.6967 | 27.8914 | 24.3655 | 30.7811 | 21.6014 | -1.4708 | 0.006732 | 0.32709 | 0.002202 | 48.5876 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.4755 | -0.1014 | 0.003694 | 13.4159 | 9.89 | 16.3056 | 7.1259
2012-10-30 20:41 | 28.6997 | 27.8348 | 24.3925 | 30.7644 | 21.5842 | -1.4789 | 0.006747 | 0.32708 | 0.002207 | 48.4775 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4458 | -0.102 | 0.003699 | 13.389 | 9.9467 | 16.3186 | 7.1384
2012-10-30 20:41 | 28.6884 | 27.8354 | 24.3732 | 30.7255 | 21.5769 | -1.4795 | 0.006747 | 0.32711 | 0.002207 | 48.4817 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.459 | -0.102 | 0.003698 | 13.3764 | 9.9142 | 16.2665 | 7.1179
2012-10-30 20:41 | 28.6944 | 27.8414 | 24.3471 | 30.7394 | 21.5265 | -1.4812 | 0.006744 | 0.32705 | 0.002206 | 48.4948 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.5058 | -0.1021 | 0.003696 | 13.3356 | 9.8413 | 16.2336 | 7.0207
2012-10-30 20:41 | 28.6862 | 27.7935 | 24.3068 | 30.6364 | 21.4941 | -1.4743 | 0.006741 | 0.32706 | 0.002205 | 48.5185 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.465 | -0.1016 | 0.003697 | 13.3285 | 9.8418 | 16.1714 | 7.0291
2012-10-30 20:41 | 28.6937 | 27.7255 | 24.3103 | 30.6321 | 21.5056 | -1.4858 | 0.006747 | 0.32712 | 0.002207 | 48.4839 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4602 | -0.1024 | 0.003696 | 13.2653 | 9.8501 | 16.1719 | 7.0454
2012-10-30 20:41 | 28.6959 | 27.6919 | 24.2604 | 30.5908 | 21.4634 | -1.472 | 0.006744 | 0.32712 | 0.002206 | 48.5055 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.5277 | -0.1015 | 0.003699 | 13.1642 | 9.7327 | 16.0631 | 6.9357
2012-10-30 20:41 | 28.6787 | 27.631 | 24.2712 | 30.542 | 21.4744 | -1.4749 | 0.006732 | 0.32711 | 0.002202 | 48.5897 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.4575 | -0.1017 | 0.003694 | 13.1735 | 9.8137 | 16.0845 | 7.0169
2012-10-30 20:41 | 28.6884 | 27.6447 | 24.2569 | 30.5241 | 21.4761 | -1.4801 | 0.006754 | 0.32708 | 0.002209 | 48.4273 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.5039 | -0.102 | 0.0037 | 13.1408 | 9.753 | 16.0202 | 6.9722
2012-10-30 20:41 | 28.6795 | 27.5761 | 24.2038 | 30.4757 | 21.4631 | -1.4645 | 0.006751 | 0.32708 | 0.002208 | 48.4488 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4664 | -0.101 | 0.003696 | 13.1097 | 9.7374 | 16.0093 | 6.9967
2012-10-30 20:41 | 28.6877 | 27.5923 | 24.192 | 30.4405 | 21.4552 | -1.4812 | 0.006738 | 0.32714 | 0.002204 | 48.5508 | 0.006961 | 0.21466 | 0.001494 | 30.8369 | 14.4869 | -0.1021 | 0.003698 | 13.1054 | 9.7051 | 15.9536 | 6.9683
2012-10-30 20:41 | 28.6854 | 27.5741 | 24.2017 | 30.4145 | 21.4085 | -1.4737 | 0.006732 | 0.32712 | 0.002202 | 48.5919 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4927 | -0.1016 | 0.003691 | 13.0814 | 9.709 | 15.9218 | 6.9158
2012-10-30 20:41 | 28.6787 | 27.5276 | 24.1549 | 30.3802 | 21.3897 | -1.476 | 0.006725 | 0.32711 | 0.0022 | 48.6403 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.4615 | -0.1018 | 0.003693 | 13.0661 | 9.6934 | 15.9187 | 6.9282
2012-10-30 20:41 | 28.672 | 27.4732 | 24.1602 | 30.3774 | 21.3669 | -1.483 | 0.006741 | 0.32708 | 0.002205 | 48.5207 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4712 | -0.1022 | 0.003696 | 13.002 | 9.689 | 15.9062 | 6.8957
2012-10-30 20:41 | 28.6765 | 27.4658 | 24.1286 | 30.3582 | 21.3794 | -1.4772 | 0.006741 | 0.32703 | 0.002205 | 48.5142 | 0.006958 | 0.2147 | 0.001494 | 30.8566 | 14.4757 | -0.1018 | 0.003699 | 12.9901 | 9.6529 | 15.8825 | 6.9037
2012-10-30 20:41 | 28.6787 | 27.468 | 24.1107 | 30.317 | 21.3978 | -1.4708 | 0.006738 | 0.32711 | 0.002204 | 48.5465 | 0.006961 | 0.21467 | 0.001494 | 30.8391 | 14.5227 | -0.1014 | 0.003698 | 12.9453 | 9.588 | 15.7943 | 6.8751
2012-10-30 20:42 | 28.6712 | 27.4247 | 24.1072 | 30.2621 | 21.3338 | -1.4668 | 0.006747 | 0.32702 | 0.002206 | 48.4689 | 0.006965 | 0.21473 | 0.001496 | 30.8297 | 14.5234 | -0.1011 | 0.003702 | 12.9013 | 9.5838 | 15.7387 | 6.8104
2012-10-30 20:42 | 28.6757 | 27.4013 | 24.0716 | 30.2192 | 21.3584 | -1.4783 | 0.006735 | 0.32708 | 0.002203 | 48.5639 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5524 | -0.1019 | 0.003693 | 12.8489 | 9.5192 | 15.6668 | 6.806
2012-10-30 20:42 | 28.672 | 27.3578 | 24.0478 | 30.2195 | 21.3063 | -1.4754 | 0.006744 | 0.32706 | 0.002206 | 48.4969 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4956 | -0.1017 | 0.003694 | 12.8622 | 9.5522 | 15.7239 | 6.8107
2012-10-30 20:42 | 28.6795 | 27.3773 | 24.0072 | 30.2112 | 21.3219 | -1.4795 | 0.006751 | 0.32709 | 0.002208 | 48.4509 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4583 | -0.102 | 0.003699 | 12.919 | 9.5489 | 15.7529 | 6.8636
2012-10-30 20:42 | 28.669 | 27.2832 | 23.9685 | 30.1415 | 21.2952 | -1.4766 | 0.006732 | 0.32705 | 0.002202 | 48.5813 | 0.006955 | 0.21473 | 0.001493 | 30.874 | 14.4804 | -0.1018 | 0.003695 | 12.8028 | 9.4881 | 15.6611 | 6.8148
2012-10-30 20:42 | 28.681 | 27.2992 | 24.0247 | 30.1179 | 21.2709 | -1.4887 | 0.006732 | 0.32706 | 0.002202 | 48.5833 | 0.006961 | 0.21461 | 0.001494 | 30.8308 | 14.4842 | -0.1026 | 0.003696 | 12.815 | 9.5405 | 15.6337 | 6.7867
2012-10-30 20:42 | 28.6675 | 27.2578 | 23.9791 | 30.061 | 21.1807 | -1.483 | 0.006751 | 0.327 | 0.002208 | 48.438 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.5034 | -0.1022 | 0.003696 | 12.7544 | 9.4757 | 15.5576 | 6.6773
2012-10-30 20:42 | 28.6795 | 27.222 | 23.955 | 30.0612 | 21.1604 | -1.4639 | 0.006738 | 0.32698 | 0.002203 | 48.5272 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.5195 | -0.1009 | 0.003694 | 12.7025 | 9.4355 | 15.5417 | 6.6409
2012-10-30 20:42 | 28.6667 | 27.2013 | 23.9381 | 30.0049 | 21.2082 | -1.4766 | 0.006741 | 0.32703 | 0.002205 | 48.5142 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5067 | -0.1018 | 0.003694 | 12.6946 | 9.4314 | 15.4982 | 6.7015
2012-10-30 20:42 | 28.6652 | 27.152 | 23.9005 | 29.9955 | 21.1986 | -1.4743 | 0.006735 | 0.32696 | 0.002202 | 48.5467 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4807 | -0.1016 | 0.003693 | 12.6713 | 9.4198 | 15.5148 | 6.7179
2012-10-30 20:42 | 28.6772 | 27.1521 | 23.8563 | 29.9996 | 21.1743 | -1.4778 | 0.006747 | 0.32702 | 0.002206 | 48.4689 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4641 | -0.1019 | 0.003696 | 12.688 | 9.3922 | 15.5355 | 6.7102
2012-10-30 20:42 | 28.6697 | 27.0928 | 23.8809 | 29.9369 | 21.1748 | -1.4749 | 0.006744 | 0.32708 | 0.002206 | 48.4991 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.5015 | -0.1017 | 0.003698 | 12.5913 | 9.3794 | 15.4354 | 6.6733
2012-10-30 20:42 | 28.6705 | 27.0896 | 23.8536 | 29.91 | 21.1676 | -1.4754 | 0.006719 | 0.32702 | 0.002197 | 48.6709 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.4452 | -0.1017 | 0.00369 | 12.6444 | 9.4084 | 15.4648 | 6.7224
2012-10-30 20:42 | 28.6607 | 27.0241 | 23.7795 | 29.8489 | 21.1214 | -1.4656 | 0.006732 | 0.32705 | 0.002202 | 48.5813 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.468 | -0.101 | 0.003692 | 12.5561 | 9.3115 | 15.3809 | 6.6534
2012-10-30 20:42 | 28.6637 | 27.0231 | 23.7825 | 29.8401 | 21.1567 | -1.4783 | 0.006741 | 0.32702 | 0.002204 | 48.5121 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.4873 | -0.1019 | 0.003696 | 12.5358 | 9.2952 | 15.3528 | 6.6694
2012-10-30 20:42 | 28.6675 | 26.987 | 23.7944 | 29.8202 | 21.1323 | -1.4778 | 0.006751 | 0.32709 | 0.002208 | 48.4509 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5075 | -0.1019 | 0.003697 | 12.4795 | 9.2869 | 15.3127 | 6.6248
2012-10-30 20:42 | 28.6675 | 26.987 | 23.7703 | 29.8202 | 21.1524 | -1.4743 | 0.006744 | 0.32708 | 0.002206 | 48.4991 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4422 | -0.1016 | 0.003696 | 12.5448 | 9.3281 | 15.378 | 6.7102
2012-10-30 20:42 | 28.6622 | 26.9539 | 23.7529 | 29.7517 | 21.1148 | -1.4754 | 0.006747 | 0.32703 | 0.002206 | 48.471 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4777 | -0.1017 | 0.003697 | 12.4762 | 9.2752 | 15.274 | 6.6371
2012-10-30 20:42 | 28.663 | 26.8869 | 23.7055 | 29.7288 | 21.0874 | -1.472 | 0.006738 | 0.32703 | 0.002204 | 48.5358 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.454 | -0.1015 | 0.003692 | 12.4329 | 9.2515 | 15.2748 | 6.6334
2012-10-30 20:42 | 28.6585 | 26.8705 | 23.697 | 29.6809 | 21.0748 | -1.4847 | 0.006735 | 0.32703 | 0.002203 | 48.5574 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4617 | -0.1024 | 0.003694 | 12.4088 | 9.2353 | 15.2192 | 6.6131
2012-10-30 20:42 | 28.6622 | 26.8542 | 23.7007 | 29.6846 | 21.0704 | -1.4743 | 0.006751 | 0.32708 | 0.002208 | 48.4488 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.5226 | -0.1016 | 0.003701 | 12.3316 | 9.1781 | 15.162 | 6.5478
2012-10-30 20:42 | 28.6705 | 26.7988 | 23.6688 | 29.6218 | 21.0787 | -1.472 | 0.006735 | 0.32702 | 0.002202 | 48.5553 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5105 | -0.1015 | 0.003692 | 12.2883 | 9.1583 | 15.1113 | 6.5682
2012-10-30 20:42 | 28.6637 | 26.7681 | 23.6299 | 29.6111 | 21.0477 | -1.4841 | 0.006741 | 0.32703 | 0.002205 | 48.5142 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.5404 | -0.1023 | 0.003697 | 12.2277 | 9.0895 | 15.0707 | 6.5073
2012-10-30 20:42 | 28.6675 | 26.7559 | 23.6618 | 29.5833 | 21.0353 | -1.4708 | 0.006738 | 0.32698 | 0.002203 | 48.5272 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4789 | -0.1014 | 0.003695 | 12.277 | 9.1829 | 15.1044 | 6.5564
2012-10-30 20:42 | 28.6645 | 26.741 | 23.6186 | 29.5171 | 21 | -1.4731 | 0.006738 | 0.32702 | 0.002203 | 48.5337 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.5004 | -0.1016 | 0.003696 | 12.2406 | 9.1182 | 15.0167 | 6.4996
2012-10-30 20:42 | 28.666 | 26.7146 | 23.6201 | 29.5305 | 21.0338 | -1.4674 | 0.006741 | 0.32702 | 0.002204 | 48.5121 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4937 | -0.1012 | 0.003695 | 12.2209 | 9.1264 | 15.0368 | 6.5401
2012-10-30 20:43 | 28.6667 | 26.6674 | 23.5646 | 29.4878 | 20.9901 | -1.4691 | 0.006738 | 0.32706 | 0.002204 | 48.5401 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.474 | -0.1013 | 0.003696 | 12.1934 | 9.0906 | 15.0138 | 6.5161
2012-10-30 20:43 | 28.663 | 26.5959 | 23.581 | 29.4446 | 20.9945 | -1.4679 | 0.006725 | 0.327 | 0.002199 | 48.6253 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4826 | -0.1012 | 0.003689 | 12.1133 | 9.0984 | 14.962 | 6.5119
2012-10-30 20:43 | 28.6735 | 26.6144 | 23.5272 | 29.4196 | 20.9808 | -1.4702 | 0.006725 | 0.32702 | 0.002199 | 48.6275 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.5012 | -0.1014 | 0.00369 | 12.1132 | 9.026 | 14.9184 | 6.4796
2012-10-30 20:43 | 28.6682 | 26.5812 | 23.55 | 29.4222 | 20.9755 | -1.4731 | 0.006735 | 0.327 | 0.002202 | 48.5531 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.4633 | -0.1016 | 0.003692 | 12.1179 | 9.0867 | 14.9589 | 6.5122
2012-10-30 20:43 | 28.6727 | 26.5219 | 23.4982 | 29.3359 | 20.9234 | -1.476 | 0.006738 | 0.32692 | 0.002203 | 48.5187 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4678 | -0.1018 | 0.003692 | 12.0541 | 9.0304 | 14.8681 | 6.4556
2012-10-30 20:43 | 28.6712 | 26.5204 | 23.4847 | 29.3501 | 20.8856 | -1.4818 | 0.006732 | 0.32689 | 0.002201 | 48.5576 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.5071 | -0.1022 | 0.003692 | 12.0133 | 8.9776 | 14.843 | 6.3785
2012-10-30 20:43 | 28.6735 | 26.5266 | 23.4468 | 29.3051 | 20.8879 | -1.4778 | 0.006735 | 0.32693 | 0.002202 | 48.5424 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.4441 | -0.1019 | 0.003691 | 12.0825 | 9.0027 | 14.861 | 6.4438
2012-10-30 20:43 | 28.6742 | 26.4595 | 23.4354 | 29.3255 | 20.8643 | -1.4749 | 0.006722 | 0.32696 | 0.002198 | 48.6406 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4734 | -0.1017 | 0.00369 | 11.9861 | 8.962 | 14.8521 | 6.3909
2012-10-30 20:43 | 28.669 | 26.4344 | 23.3981 | 29.2611 | 20.8753 | -1.4916 | 0.006722 | 0.32705 | 0.002198 | 48.6535 | 0.006949 | 0.21458 | 0.001491 | 30.8798 | 14.4559 | -0.1028 | 0.003689 | 11.9785 | 8.9422 | 14.8052 | 6.4194
2012-10-30 20:43 | 28.6757 | 26.4411 | 23.449 | 29.2362 | 20.8456 | -1.4841 | 0.006729 | 0.32699 | 0.0022 | 48.5943 | 0.006939 | 0.21457 | 0.001489 | 30.9223 | 14.4912 | -0.1023 | 0.003689 | 11.9499 | 8.9578 | 14.745 | 6.3544
2012-10-30 20:43 | 28.678 | 26.3716 | 23.3669 | 29.2543 | 20.8439 | -1.4743 | 0.006738 | 0.32696 | 0.002203 | 48.5251 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4853 | -0.1016 | 0.003692 | 11.8863 | 8.8816 | 14.769 | 6.3586
2012-10-30 20:43 | 28.669 | 26.3665 | 23.3659 | 29.1664 | 20.8713 | -1.4685 | 0.006738 | 0.327 | 0.002203 | 48.5315 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4886 | -0.1012 | 0.003693 | 11.8779 | 8.8773 | 14.6778 | 6.3827
2012-10-30 20:43 | 28.669 | 26.3147 | 23.3579 | 29.1664 | 20.847 | -1.4801 | 0.006747 | 0.32702 | 0.002206 | 48.4689 | 0.006924 | 0.21458 | 0.001486 | 30.9913 | 14.4926 | -0.102 | 0.003692 | 11.8221 | 8.8653 | 14.6738 | 6.3544
2012-10-30 20:43 | 28.669 | 26.2867 | 23.3579 | 29.1229 | 20.8389 | -1.4726 | 0.006735 | 0.32698 | 0.002202 | 48.5488 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.5294 | -0.1015 | 0.003692 | 11.7573 | 8.8285 | 14.5935 | 6.3095
2012-10-30 20:43 | 28.6645 | 26.2902 | 23.3252 | 29.154 | 20.8385 | -1.4731 | 0.006735 | 0.32696 | 0.002202 | 48.5467 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5004 | -0.1016 | 0.003691 | 11.7898 | 8.8248 | 14.6536 | 6.3381
2012-10-30 20:43 | 28.6682 | 26.262 | 23.3329 | 29.0866 | 20.8018 | -1.4847 | 0.006725 | 0.3269 | 0.002198 | 48.6103 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.5163 | -0.1024 | 0.003687 | 11.7457 | 8.8166 | 14.5703 | 6.2855
2012-10-30 20:43 | 28.6735 | 26.2273 | 23.3262 | 29.0682 | 20.7788 | -1.4824 | 0.006735 | 0.32699 | 0.002202 | 48.551 | 0.006939 | 0.21461 | 0.001489 | 30.9285 | 14.4727 | -0.1022 | 0.003691 | 11.7546 | 8.8535 | 14.5955 | 6.3061
2012-10-30 20:43 | 28.6757 | 26.2096 | 23.3002 | 29.0033 | 20.7689 | -1.4858 | 0.006735 | 0.32702 | 0.002202 | 48.5553 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.5238 | -0.1024 | 0.003691 | 11.6858 | 8.7764 | 14.4795 | 6.2451
2012-10-30 20:43 | 28.6682 | 26.1542 | 23.2686 | 28.9998 | 20.7371 | -1.4887 | 0.006732 | 0.32698 | 0.002201 | 48.5704 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4796 | -0.1026 | 0.003693 | 11.6746 | 8.789 | 14.5202 | 6.2575
2012-10-30 20:43 | 28.6675 | 26.1375 | 23.2518 | 28.9872 | 20.7486 | -1.4714 | 0.006738 | 0.32698 | 0.002203 | 48.5272 | 0.006955 | 0.21468 | 0.001493 | 30.8677 | 14.4993 | -0.1014 | 0.003696 | 11.6382 | 8.7525 | 14.4879 | 6.2493
2012-10-30 20:43 | 28.6705 | 26.1485 | 23.2347 | 28.9665 | 20.7798 | -1.476 | 0.006729 | 0.32696 | 0.0022 | 48.59 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.5227 | -0.1018 | 0.003692 | 11.6258 | 8.712 | 14.4438 | 6.2571
2012-10-30 20:43 | 28.6697 | 26.0878 | 23.2661 | 28.9539 | 20.7386 | -1.4783 | 0.006729 | 0.32702 | 0.002201 | 48.5986 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4933 | -0.1019 | 0.003691 | 11.5945 | 8.7728 | 14.4606 | 6.2453
2012-10-30 20:43 | 28.6772 | 26.0513 | 23.2173 | 28.9219 | 20.7057 | -1.4778 | 0.006732 | 0.327 | 0.002201 | 48.5747 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.509 | -0.1019 | 0.003694 | 11.5423 | 8.7083 | 14.4129 | 6.1967
2012-10-30 20:43 | 28.675 | 26.0172 | 23.195 | 28.9118 | 20.6752 | -1.4783 | 0.006744 | 0.32706 | 0.002206 | 48.4969 | 0.006955 | 0.21464 | 0.001493 | 30.8615 | 14.4374 | -0.1019 | 0.003699 | 11.5798 | 8.7576 | 14.4744 | 6.2378
2012-10-30 20:43 | 28.6765 | 26.0027 | 23.1764 | 28.8699 | 20.6565 | -1.4824 | 0.006738 | 0.327 | 0.002203 | 48.5315 | 0.006946 | 0.21461 | 0.001491 | 30.8974 | 14.4512 | -0.1022 | 0.003694 | 11.5515 | 8.7252 | 14.4187 | 6.2053
2012-10-30 20:44 | 28.6742 | 25.9764 | 23.1459 | 28.8202 | 20.6138 | -1.4789 | 0.00676 | 0.32696 | 0.00221 | 48.3672 | 0.006952 | 0.21463 | 0.001492 | 30.8728 | 14.4897 | -0.102 | 0.003702 | 11.4867 | 8.6562 | 14.3305 | 6.1241
2012-10-30 20:44 | 28.681 | 25.9992 | 23.1809 | 28.831 | 20.6489 | -1.4835 | 0.006719 | 0.32696 | 0.002197 | 48.6623 | 0.006958 | 0.21467 | 0.001494 | 30.8524 | 14.4883 | -0.1023 | 0.003691 | 11.5109 | 8.6926 | 14.3427 | 6.1606
2012-10-30 20:44 | 28.6802 | 25.9185 | 23.1399 | 28.8065 | 20.5511 | -1.4749 | 0.006747 | 0.32702 | 0.002206 | 48.4689 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4549 | -0.1017 | 0.003694 | 11.4636 | 8.685 | 14.3516 | 6.0962
2012-10-30 20:44 | 28.6727 | 25.927 | 23.1122 | 28.7674 | 20.5719 | -1.4806 | 0.006738 | 0.32702 | 0.002203 | 48.5337 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4637 | -0.1021 | 0.003693 | 11.4633 | 8.6485 | 14.3037 | 6.1082
2012-10-30 20:44 | 28.669 | 25.8913 | 23.0563 | 28.7361 | 20.5844 | -1.4766 | 0.006738 | 0.32696 | 0.002203 | 48.5251 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.46 | -0.1018 | 0.003691 | 11.4313 | 8.5963 | 14.2761 | 6.1244
2012-10-30 20:44 | 28.6765 | 25.8589 | 23.1 |  | 20.5838 | -1.4772 | 0.006744 | 0.32695 | 0.002205 | 48.4798 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.4961 | -0.1018 | 0.003697 | 11.3628 | 8.6039 | -14.4961 | 6.0877
2012-10-30 20:44 | 28.6817 | 25.8281 | 23.073 |  | 20.5445 | -1.4812 | 0.006725 | 0.32699 | 0.002199 | 48.6232 | 0.006955 | 0.21461 | 0.001493 | 30.8574 | 14.4441 | -0.1021 | 0.003692 | 11.384 | 8.6289 | -14.4441 | 6.1004
2012-10-30 20:44 | 28.6832 | 25.8336 | 23.0584 |  | 20.5662 | -1.4795 | 0.006729 | 0.32696 | 0.0022 | 48.59 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4864 | -0.102 | 0.003691 | 11.3472 | 8.572 | -14.4864 | 6.0798
2012-10-30 20:44 | 28.6772 | 25.7556 | 23.0443 |  | 20.5683 | -1.4772 | 0.006744 | 0.32702 | 0.002205 | 48.4905 | 0.006955 | 0.21463 | 0.001493 | 30.8595 | 14.5008 | -0.1018 | 0.003698 | 11.2548 | 8.5435 | -14.5008 | 6.0675
2012-10-30 20:44 | 28.6854 | 25.7678 | 23.0123 |  | 20.5604 | -1.4731 | 0.006735 | 0.32706 | 0.002203 | 48.5617 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.5131 | -0.1016 | 0.003692 | 11.2547 | 8.4992 | -14.5131 | 6.0473
2012-10-30 20:44 | 28.6847 | 25.7192 | 23.0478 | 28.5936 | 20.5799 | -1.4754 | 0.006725 | 0.32699 | 0.002199 | 48.6232 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.439 | -0.1017 | 0.003692 | 11.2802 | 8.6088 | 14.1546 | 6.1409
2012-10-30 20:44 | 28.6824 | 25.7528 | 22.9892 | 28.5635 | 20.521 | -1.4789 | 0.006751 | 0.327 | 0.002208 | 48.438 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4938 | -0.102 | 0.003699 | 11.259 | 8.4954 | 14.0697 | 6.0272
2012-10-30 20:44 | 28.681 | 25.6715 | 22.9797 | 28.5265 | 20.5115 | -1.4737 | 0.006735 | 0.32693 | 0.002202 | 48.5424 | 0.006946 | 0.21455 | 0.00149 | 30.889 | 14.5006 | -0.1016 | 0.003692 | 11.1709 | 8.4791 | 14.0259 | 6.0109
2012-10-30 20:44 | 28.6914 | 25.6739 | 22.8936 | 28.5289 | 20.5017 | -1.4668 | 0.006725 | 0.32695 | 0.002199 | 48.6168 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.5232 | -0.1011 | 0.003691 | 11.1507 | 8.3704 | 14.0057 | 5.9785
2012-10-30 20:44 | 28.6839 | 25.6864 | 22.9746 | 28.4977 | 20.5063 | -1.4749 | 0.006722 | 0.32695 | 0.002198 | 48.6385 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5035 | -0.1017 | 0.003685 | 11.1829 | 8.4711 | 13.9942 | 6.0028
2012-10-30 20:44 | 28.681 | 25.5715 | 22.9154 | 28.4551 | 20.4347 | -1.476 | 0.006735 | 0.32698 | 0.002202 | 48.5488 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.5414 | -0.1018 | 0.003691 | 11.0301 | 8.374 | 13.9137 | 5.8933
2012-10-30 20:44 | 28.6877 | 25.5822 | 22.914 | 28.438 | 20.4454 | -1.4818 | 0.006722 | 0.32689 | 0.002197 | 48.6299 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4869 | -0.1022 | 0.003689 | 11.0953 | 8.4271 | 13.9511 | 5.9585
2012-10-30 20:44 | 28.6907 | 25.5332 | 22.917 | 28.4331 | 20.505 | -1.4714 | 0.006732 | 0.32702 | 0.002201 | 48.5769 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.5143 | -0.1014 | 0.003693 | 11.0189 | 8.4027 | 13.9188 | 5.9907
2012-10-30 20:44 | 28.6907 | 25.5412 | 22.8687 | 28.4093 | 20.4686 | -1.4783 | 0.006729 | 0.32699 | 0.0022 | 48.5943 | 0.006952 | 0.21461 | 0.001492 | 30.8707 | 14.5103 | -0.1019 | 0.003692 | 11.0309 | 8.3584 | 13.899 | 5.9583
2012-10-30 20:44 | 28.6839 | 25.4744 | 22.8338 | 28.347 | 20.4093 | -1.4835 | 0.006741 | 0.32695 | 0.002204 | 48.5014 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4749 | -0.1023 | 0.003696 | 10.9995 | 8.3589 | 13.8721 | 5.9344
2012-10-30 20:44 | 28.6854 | 25.4559 | 22.8755 | 28.3326 | 20.4229 | -1.4841 | 0.006751 | 0.32702 | 0.002208 | 48.4402 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.505 | -0.1023 | 0.003698 | 10.9509 | 8.3705 | 13.8276 | 5.9179
2012-10-30 20:44 | 28.6899 | 25.4484 | 22.8076 | 28.3371 | 20.4395 | -1.4749 | 0.006751 | 0.32702 | 0.002208 | 48.4402 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5135 | -0.1017 | 0.003699 | 10.9349 | 8.2941 | 13.8236 | 5.926
2012-10-30 20:44 | 28.6869 | 25.4014 | 22.8529 | 28.2786 | 20.4163 | -1.4772 | 0.006747 | 0.327 | 0.002206 | 48.4667 | 0.006952 | 0.21464 | 0.001492 | 30.8749 | 14.5065 | -0.1018 | 0.003698 | 10.8949 | 8.3464 | 13.7721 | 5.9098
2012-10-30 20:44 | 28.6832 | 25.3817 | 22.8009 | 28.263 | 20.3843 | -1.4853 | 0.006735 | 0.327 | 0.002202 | 48.5531 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.5068 | -0.1024 | 0.003692 | 10.8749 | 8.2941 | 13.7562 | 5.8775
2012-10-30 20:44 | 28.6884 | 25.4109 | 22.79 | 28.2484 | 20.3855 | -1.4853 | 0.006741 | 0.32699 | 0.002204 | 48.5078 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.561 | -0.1024 | 0.003693 | 10.8499 | 8.229 | 13.6874 | 5.8245
2012-10-30 20:44 | 28.6862 | 25.3567 | 22.7878 | 28.2224 | 20.3752 | -1.4905 | 0.006744 | 0.32696 | 0.002205 | 48.4819 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4935 | -0.1028 | 0.003696 | 10.8632 | 8.2943 | 13.7289 | 5.8817
2012-10-30 20:44 | 28.6824 | 25.2848 | 22.7598 | 28.2424 | 20.3188 | -1.483 | 0.006732 | 0.32696 | 0.002201 | 48.5683 | 0.006952 | 0.21463 | 0.001492 | 30.8728 | 14.5101 | -0.1022 | 0.003693 | 10.7747 | 8.2497 | 13.7323 | 5.8087
2012-10-30 20:45 | 28.6839 | 25.2823 | 22.7653 | 28.1843 | 20.3486 | -1.4911 | 0.006738 | 0.32693 | 0.002203 | 48.5208 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4749 | -0.1028 | 0.003693 | 10.8074 | 8.2904 | 13.7094 | 5.8737
2012-10-30 20:45 | 28.6817 | 25.2841 | 22.7511 | 28.1663 | 20.3545 | -1.4795 | 0.006747 | 0.32689 | 0.002206 | 48.4497 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.4931 | -0.102 | 0.003698 | 10.791 | 8.258 | 13.6732 | 5.8614
2012-10-30 20:45 | 28.6907 | 25.2931 | 22.7762 | 28.1435 | 20.3393 | -1.4726 | 0.006729 | 0.32699 | 0.0022 | 48.5943 | 0.006939 | 0.21458 | 0.001489 | 30.9243 | 14.4899 | -0.1015 | 0.003689 | 10.8032 | 8.2863 | 13.6536 | 5.8494
2012-10-30 20:45 | 28.6839 | 25.2263 | 22.7009 | 28.1169 | 20.2759 | -1.4818 | 0.006735 | 0.327 | 0.002202 | 48.5531 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.5035 | -0.1022 | 0.003693 | 10.7228 | 8.1974 | 13.6134 | 5.7724
2012-10-30 20:45 | 28.6765 | 25.1949 | 22.7016 | 28.0936 | 20.2482 | -1.4766 | 0.006738 | 0.32695 | 0.002203 | 48.523 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.5124 | -0.1018 | 0.003693 | 10.6825 | 8.1892 | 13.5812 | 5.7358
2012-10-30 20:45 | 28.6787 | 25.173 | 22.6716 | 28.0958 | 20.303 | -1.4858 | 0.006741 | 0.32695 | 0.002204 | 48.5014 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.5717 | -0.1024 | 0.003693 | 10.6013 | 8.0999 | 13.5241 | 5.7313
2012-10-30 20:45 | 28.6757 | 25.154 | 22.6605 | 28.0293 | 20.207 | -1.4899 | 0.006754 | 0.32696 | 0.002208 | 48.4101 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5238 | -0.1027 | 0.003697 | 10.6302 | 8.1367 | 13.5055 | 5.6832
2012-10-30 20:45 | 28.6757 | 25.09 | 22.6806 | 27.9657 | 20.2272 | -1.4911 | 0.006735 | 0.32696 | 0.002202 | 48.5467 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.5238 | -0.1028 | 0.00369 | 10.5662 | 8.1568 | 13.4419 | 5.7034
2012-10-30 20:45 | 28.6772 | 25.1035 | 22.6137 | 28.003 | 20.253 | -1.4887 | 0.006729 | 0.32699 | 0.0022 | 48.5943 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5131 | -0.1026 | 0.00369 | 10.5904 | 8.1006 | 13.4899 | 5.7399
2012-10-30 20:45 | 28.6705 | 25.0367 | 22.6392 | 27.9288 | 20.218 | -1.4968 | 0.006741 | 0.32702 | 0.002204 | 48.5121 | 0.006965 | 0.21467 | 0.001495 | 30.8214 | 14.4982 | -0.1032 | 0.003699 | 10.5385 | 8.141 | 13.4306 | 5.7198
2012-10-30 20:45 | 28.6712 | 25.0454 | 22.5835 | 27.9493 | 20.1823 | -1.4922 | 0.006747 | 0.32699 | 0.002206 | 48.4647 | 0.006936 | 0.21458 | 0.001488 | 30.9377 | 14.5112 | -0.1029 | 0.003694 | 10.5342 | 8.0723 | 13.4381 | 5.6711
2012-10-30 20:45 | 28.6712 | 25.0454 | 22.5634 | 27.9295 | 20.1782 | -1.4853 | 0.006735 | 0.32699 | 0.002202 | 48.551 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.503 | -0.1024 | 0.003692 | 10.5424 | 8.0604 | 13.4265 | 5.6752
2012-10-30 20:45 | 28.6712 | 25.0134 | 22.5191 | 27.9056 | 20.1621 | -1.4737 | 0.006744 | 0.32695 | 0.002205 | 48.4798 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.552 | -0.1016 | 0.003697 | 10.4614 | 7.9671 | 13.3536 | 5.6101
2012-10-30 20:45 | 28.6697 | 24.9438 | 22.5498 | 27.8883 | 20.1363 | -1.4853 | 0.006741 | 0.32696 | 0.002204 | 48.5035 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.5015 | -0.1024 | 0.003693 | 10.4423 | 8.0483 | 13.3868 | 5.6348
2012-10-30 20:45 | 28.675 | 24.9811 | 22.5511 | 27.8578 | 20.1254 | -1.4818 | 0.006751 | 0.32692 | 0.002207 | 48.4253 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.5027 | -0.1022 | 0.003698 | 10.4784 | 8.0484 | 13.3551 | 5.6227
2012-10-30 20:45 | 28.6765 | 24.9746 | 22.4841 | 27.8712 | 20.1148 | -1.4893 | 0.006738 | 0.32693 | 0.002203 | 48.5208 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5001 | -0.1027 | 0.003694 | 10.4745 | 7.984 | 13.3711 | 5.6147
2012-10-30 20:45 | 28.6765 | 24.8985 | 22.5163 | 27.8235 | 20.1148 | -1.4905 | 0.006735 | 0.32693 | 0.002202 | 48.5424 | 0.006927 | 0.21463 | 0.001487 | 30.9843 | 14.5205 | -0.1028 | 0.003689 | 10.378 | 7.9958 | 13.303 | 5.5943
2012-10-30 20:45 | 28.6839 | 24.946 | 22.5318 | 27.7713 | 20.102 | -1.483 | 0.006747 | 0.32698 | 0.002206 | 48.4624 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.5116 | -0.1022 | 0.003698 | 10.4344 | 8.0202 | 13.2597 | 5.5904
2012-10-30 20:45 | 28.672 | 24.874 | 22.4313 | 27.7118 | 20.086 | -1.4876 | 0.006741 | 0.32698 | 0.002204 | 48.5056 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4793 | -0.1026 | 0.003693 | 10.3947 | 7.952 | 13.2325 | 5.6067
2012-10-30 20:45 | 28.6705 | 24.8084 | 22.4379 | 27.75 | 20.1128 | -1.4876 | 0.006735 | 0.32699 | 0.002202 | 48.551 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.486 | -0.1026 | 0.003691 | 10.3224 | 7.9519 | 13.264 | 5.6268
2012-10-30 20:45 | 28.675 | 24.8089 | 22.4222 | 27.7307 | 20.0809 | -1.4801 | 0.006732 | 0.32693 | 0.002201 | 48.564 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4986 | -0.102 | 0.00369 | 10.3103 | 7.9236 | 13.2321 | 5.5823
2012-10-30 20:45 | 28.681 | 24.8389 | 22.4121 | 27.7009 | 20.0829 | -1.4766 | 0.006754 | 0.32692 | 0.002208 | 48.4038 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5006 | -0.1018 | 0.003697 | 10.3383 | 7.9115 | 13.2003 | 5.5823
2012-10-30 20:45 | 28.6787 | 24.7685 | 22.4138 | 27.6787 | 20.0361 | -1.4957 | 0.006747 | 0.32695 | 0.002206 | 48.4583 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5513 | -0.1031 | 0.003695 | 10.2172 | 7.8625 | 13.1274 | 5.4848
2012-10-30 20:45 | 28.6742 | 24.7279 | 22.4134 | 27.6384 | 20.0154 | -1.4835 | 0.006722 | 0.3269 | 0.002197 | 48.632 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5101 | -0.1023 | 0.003688 | 10.2178 | 7.9033 | 13.1283 | 5.5053
2012-10-30 20:45 | 28.6607 | 24.7505 | 22.3717 | 27.6249 | 20.0383 | -1.4841 | 0.006738 | 0.327 | 0.002203 | 48.5315 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5129 | -0.1023 | 0.003693 | 10.2376 | 7.8588 | 13.112 | 5.5254
2012-10-30 20:45 | 28.6622 | 24.7039 | 22.3691 | 27.5986 | 20.0196 | -1.4887 | 0.006744 | 0.32693 | 0.002205 | 48.4776 | 0.006927 | 0.21463 | 0.001487 | 30.9843 | 14.5022 | -0.1026 | 0.003692 | 10.2017 | 7.8669 | 13.0964 | 5.5174
2012-10-30 20:45 | 28.6607 | 24.6783 | 22.3435 | 27.5534 | 20.0181 | -1.5084 | 0.006732 | 0.32698 | 0.002201 | 48.5704 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4884 | -0.104 | 0.003691 | 10.1899 | 7.8551 | 13.065 | 5.5297
2012-10-30 20:46 | 28.6675 | 24.6731 | 22.3583 | 27.5124 | 20.0289 | -1.4858 | 0.006735 | 0.32698 | 0.002202 | 48.5488 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4952 | -0.1024 | 0.003693 | 10.1779 | 7.8631 | 13.0172 | 5.5337
2012-10-30 20:46 | 28.6645 | 24.6661 | 22.3271 | 27.5254 | 20.0016 | -1.4812 | 0.006716 | 0.32696 | 0.002196 | 48.684 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.5289 | -0.1021 | 0.003688 | 10.1372 | 7.7982 | 12.9965 | 5.4727
2012-10-30 20:46 | 28.6697 | 24.6633 | 22.2961 | 27.5027 | 20.0028 | -1.4934 | 0.006732 | 0.32698 | 0.002201 | 48.5704 | 0.006927 | 0.21461 | 0.001487 | 30.9821 | 14.5015 | -0.103 | 0.003688 | 10.1618 | 7.7946 | 13.0012 | 5.5013
2012-10-30 20:46 | 28.654 | 24.6315 | 22.3529 | 27.4671 | 19.9143 | -1.4887 | 0.006744 | 0.32698 | 0.002205 | 48.484 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.4817 | -0.1026 | 0.003697 | 10.1498 | 7.8712 | 12.9854 | 5.4326
2012-10-30 20:46 | 28.6652 | 24.6147 | 22.2513 | 27.4743 | 19.9498 | -1.4899 | 0.006751 | 0.32693 | 0.002207 | 48.4273 | 0.006917 | 0.2147 | 0.001485 | 31.0395 | 14.5011 | -0.1027 | 0.003692 | 10.1136 | 7.7502 | 12.9732 | 5.4487
2012-10-30 20:46 | 28.6645 | 24.5939 | 22.2667 | 27.4339 | 19.9612 | -1.4876 | 0.006744 | 0.32699 | 0.002205 | 48.4862 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.5004 | -0.1026 | 0.003698 | 10.0935 | 7.7663 | 12.9335 | 5.4608
2012-10-30 20:46 | 28.6577 | 24.5551 | 22.28 | 27.4112 | 19.8694 | -1.483 | 0.006725 | 0.327 | 0.002199 | 48.6253 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4569 | -0.1022 | 0.00369 | 10.0982 | 7.8231 | 12.9543 | 5.4125
2012-10-30 20:46 | 28.6592 | 24.5405 | 22.2332 | 27.3888 | 19.9276 | -1.4893 | 0.006747 | 0.32693 | 0.002206 | 48.4561 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.4665 | -0.1027 | 0.003698 | 10.074 | 7.7667 | 12.9223 | 5.4611
2012-10-30 20:46 | 28.6577 | 24.515 | 22.1833 | 27.4231 | 19.8977 | -1.4801 | 0.006741 | 0.3269 | 0.002204 | 48.4949 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.4977 | -0.102 | 0.003697 | 10.0173 | 7.6856 | 12.9254 | 5.4
2012-10-30 20:46 | 28.6645 | 24.5218 | 22.2506 | 27.3782 | 19.9207 | -1.4749 | 0.006732 | 0.3269 | 0.002201 | 48.5597 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.48 | -0.1017 | 0.003692 | 10.0418 | 7.7706 | 12.8982 | 5.4407
2012-10-30 20:46 | 28.6495 | 24.4506 | 22.2074 | 27.2955 | 19.8895 | -1.4847 | 0.006735 | 0.32689 | 0.002202 | 48.536 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.4976 | -0.1024 | 0.003693 | 9.953 | 7.7098 | 12.7979 | 5.3919
2012-10-30 20:46 | 28.657 | 24.4822 | 22.1746 | 27.3349 | 19.9011 | -1.4853 | 0.006744 | 0.32698 | 0.002205 | 48.484 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.4521 | -0.1024 | 0.003692 | 10.0301 | 7.7225 | 12.8828 | 5.449
2012-10-30 20:46 | 28.651 | 24.396 | 22.1847 | 27.2771 | 19.8951 | -1.4772 | 0.006741 | 0.32696 | 0.002204 | 48.5035 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4869 | -0.1018 | 0.003694 | 9.9091 | 7.6978 | 12.7902 | 5.4082
2012-10-30 20:46 | 28.6592 | 24.4282 | 22.1647 | 27.2694 | 19.879 | -1.4835 | 0.006751 | 0.32698 | 0.002207 | 48.4337 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4788 | -0.1023 | 0.003696 | 9.9494 | 7.6859 | 12.7906 | 5.4002
2012-10-30 20:46 | 28.6615 | 24.3985 | 22.1871 | 27.2478 | 19.8489 | -1.4824 | 0.006716 | 0.32695 | 0.002196 | 48.682 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4811 | -0.1022 | 0.003687 | 9.9174 | 7.706 | 12.7667 | 5.3678
2012-10-30 20:46 | 28.6645 | 24.3573 | 22.0813 | 27.227 | 19.8358 | -1.4945 | 0.006751 | 0.32692 | 0.002207 | 48.4253 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4555 | -0.103 | 0.003697 | 9.9018 | 7.6258 | 12.7715 | 5.3803
2012-10-30 20:46 | 28.66 | 24.3488 | 22.0728 | 27.1986 | 19.8434 | -1.4951 | 0.006735 | 0.32698 | 0.002202 | 48.5488 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4714 | -0.1031 | 0.00369 | 9.8774 | 7.6014 | 12.7272 | 5.372
2012-10-30 20:46 | 28.6705 | 24.3433 | 22.0994 | 27.213 | 19.8296 | -1.4806 | 0.006732 | 0.32692 | 0.002201 | 48.5619 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4778 | -0.1021 | 0.003692 | 9.8655 | 7.6216 | 12.7352 | 5.3518
2012-10-30 20:46 | 28.6645 | 24.3333 | 22.0733 | 27.1513 | 19.8277 | -1.487 | 0.006747 | 0.3269 | 0.002206 | 48.4518 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4473 | -0.1025 | 0.003697 | 9.886 | 7.626 | 12.704 | 5.3804
2012-10-30 20:46 | 28.6705 | 24.2871 | 22.0994 | 27.0936 | 19.8579 | -1.4864 | 0.006735 | 0.32693 | 0.002202 | 48.5424 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4615 | -0.1025 | 0.003691 | 9.8256 | 7.6379 | 12.6321 | 5.3964
2012-10-30 20:46 | 28.6682 | 24.2648 | 22.0487 | 27.1271 | 19.7949 | -1.4893 | 0.006735 | 0.32692 | 0.002202 | 48.5403 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.4714 | -0.1027 | 0.003694 | 9.7934 | 7.5773 | 12.6557 | 5.3235
2012-10-30 20:46 | 28.675 | 24.2555 | 22.0112 | 27.0901 | 19.8503 | -1.4911 | 0.006744 | 0.32698 | 0.002205 | 48.484 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.5272 | -0.1028 | 0.003696 | 9.7283 | 7.484 | 12.5629 | 5.3231
2012-10-30 20:46 | 28.6682 | 24.2086 | 22.0004 | 27.0316 | 19.7585 | -1.4899 | 0.006744 | 0.32695 | 0.002205 | 48.4798 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4551 | -0.1027 | 0.003696 | 9.7535 | 7.5453 | 12.5765 | 5.3034
2012-10-30 20:46 | 28.672 | 24.2405 | 22.0082 | 26.9995 | 19.7745 | -1.4934 | 0.006751 | 0.32696 | 0.002207 | 48.4316 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4426 | -0.103 | 0.003696 | 9.7979 | 7.5656 | 12.5569 | 5.3319
2012-10-30 20:47 | 28.672 | 24.112 | 21.9477 | 26.9836 | 19.7259 | -1.4968 | 0.006732 | 0.32689 | 0.002201 | 48.5576 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.4752 | -0.1032 | 0.003693 | 9.6368 | 7.4725 | 12.5084 | 5.2507
2012-10-30 20:47 | 28.681 | 24.097 | 21.9406 | 26.9328 | 19.7228 | -1.5043 | 0.006754 | 0.32695 | 0.002208 | 48.4081 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4924 | -0.1037 | 0.003697 | 9.6046 | 7.4482 | 12.4404 | 5.2304
2012-10-30 20:47 | 28.6765 | 24.0122 | 21.9401 | 26.9084 | 19.7061 | -1.4899 | 0.006741 | 0.32698 | 0.002204 | 48.5056 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4512 | -0.1027 | 0.003696 | 9.561 | 7.4889 | 12.4572 | 5.2549
2012-10-30 20:47 | 28.6824 | 24.0783 | 21.8936 | 26.9063 | 19.6958 | -1.4911 | 0.006735 | 0.32699 | 0.002202 | 48.551 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.4693 | -0.1028 | 0.003694 | 9.609 | 7.4243 | 12.437 | 5.2265
2012-10-30 20:47 | 28.6802 | 24.0319 | 21.9358 | 26.8762 | 19.7017 | -1.4928 | 0.006747 | 0.327 | 0.002206 | 48.4667 | 0.006933 | 0.2146 | 0.001488 | 30.9533 | 14.4834 | -0.1029 | 0.003694 | 9.5485 | 7.4524 | 12.3928 | 5.2183
2012-10-30 20:47 | 28.6765 | 23.976 | 21.8917 | 26.8287 | 19.694 | -1.4899 | 0.006732 | 0.32698 | 0.002201 | 48.5704 | 0.006958 | 0.21467 | 0.001494 | 30.8524 | 14.5083 | -0.1027 | 0.003695 | 9.4677 | 7.3834 | 12.3204 | 5.1857
2012-10-30 20:47 | 28.6787 | 23.9542 | 21.8778 | 26.803 | 19.6395 | -1.4858 | 0.006735 | 0.32695 | 0.002202 | 48.5446 | 0.006955 | 0.21466 | 0.001493 | 30.8636 | 14.5105 | -0.1024 | 0.003695 | 9.4437 | 7.3673 | 12.2925 | 5.129
2012-10-30 20:47 | 28.6862 | 23.9255 | 21.8813 | 26.7826 | 19.7077 | -1.483 | 0.006747 | 0.32688 | 0.002205 | 48.4475 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.5017 | -0.1022 | 0.003696 | 9.4238 | 7.3796 | 12.2809 | 5.206
2012-10-30 20:47 | 28.681 | 23.9444 | 21.8438 | 26.7575 | 19.6661 | -1.5055 | 0.006729 | 0.32692 | 0.0022 | 48.5836 | 0.006933 | 0.21461 | 0.001488 | 30.9553 | 14.5332 | -0.1038 | 0.003688 | 9.4112 | 7.3106 | 12.2243 | 5.1329
2012-10-30 20:47 | 28.6772 | 23.9125 | 21.832 | 26.7377 | 19.642 | -1.4997 | 0.006744 | 0.32692 | 0.002205 | 48.4755 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4764 | -0.1034 | 0.003696 | 9.4361 | 7.3556 | 12.2613 | 5.1656
2012-10-30 20:47 | 28.681 | 23.8721 | 21.7954 | 26.6857 | 19.6701 | -1.487 | 0.006744 | 0.32692 | 0.002205 | 48.4755 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5128 | -0.1025 | 0.003695 | 9.3593 | 7.2826 | 12.1729 | 5.1573
2012-10-30 20:47 | 28.681 | 23.848 | 21.8438 | 26.7136 | 19.6054 | -1.4963 | 0.006747 | 0.32699 | 0.002206 | 48.4647 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.5006 | -0.1032 | 0.003697 | 9.3474 | 7.3432 | 12.213 | 5.1048
2012-10-30 20:47 | 28.6869 | 23.8218 | 21.8255 | 26.6756 | 19.6234 | -1.4997 | 0.006744 | 0.32695 | 0.002205 | 48.4798 | 0.006961 | 0.21471 | 0.001495 | 30.8453 | 14.5024 | -0.1034 | 0.0037 | 9.3194 | 7.3231 | 12.1732 | 5.121
2012-10-30 20:47 | 28.6817 | 23.8005 | 21.7881 | 26.6186 | 19.6223 | -1.4986 | 0.006744 | 0.32689 | 0.002205 | 48.4712 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.5013 | -0.1033 | 0.003695 | 9.2992 | 7.2868 | 12.1173 | 5.121
2012-10-30 20:47 | 28.6847 | 23.7995 | 21.7749 | 26.6256 | 19.5564 | -1.4963 | 0.006732 | 0.32695 | 0.002201 | 48.5663 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.492 | -0.1032 | 0.003692 | 9.3075 | 7.2829 | 12.1336 | 5.0644
2012-10-30 20:47 | 28.6929 | 23.8037 | 21.8678 | 26.6418 | 19.593 | -1.5026 | 0.006738 | 0.32686 | 0.002202 | 48.5101 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.5206 | -0.1036 | 0.003692 | 9.2831 | 7.3472 | 12.1212 | 5.0724
2012-10-30 20:47 | 28.6877 | 23.7543 | 21.7699 | 26.5528 | 19.5716 | -1.4963 | 0.006729 | 0.3269 | 0.0022 | 48.5814 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4828 | -0.1032 | 0.003688 | 9.2715 | 7.2871 | 12.07 | 5.0888
2012-10-30 20:47 | 28.6847 | 23.7594 | 21.7467 | 26.5458 | 19.601 | -1.4945 | 0.006738 | 0.32693 | 0.002203 | 48.5208 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.5083 | -0.103 | 0.003693 | 9.2511 | 7.2384 | 12.0375 | 5.0927
2012-10-30 20:47 | 28.6824 | 23.745 | 21.7081 | 26.5595 | 19.5663 | -1.4986 | 0.006738 | 0.32689 | 0.002203 | 48.5144 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.502 | -0.1033 | 0.003694 | 9.243 | 7.2061 | 12.0575 | 5.0643
2012-10-30 20:47 | 28.6839 | 23.6983 | 21.7056 | 26.537 | 19.5799 | -1.4934 | 0.006732 | 0.32699 | 0.002201 | 48.5726 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5075 | -0.103 | 0.00369 | 9.1908 | 7.1981 | 12.0295 | 5.0724
2012-10-30 20:47 | 28.6907 | 23.6408 | 21.6882 | 26.5159 | 19.4774 | -1.483 | 0.006725 | 0.32699 | 0.002199 | 48.6232 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5266 | -0.1022 | 0.003689 | 9.1142 | 7.1616 | 11.9893 | 4.9508
2012-10-30 20:47 | 28.6877 | 23.6579 | 21.7053 | 26.5129 | 19.523 | -1.4974 | 0.006735 | 0.32692 | 0.002202 | 48.5403 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5154 | -0.1032 | 0.003691 | 9.1425 | 7.1899 | 11.9975 | 5.0076
2012-10-30 20:47 | 28.6907 | 23.6328 | 21.6922 | 26.4361 | 19.5179 | -1.4882 | 0.006741 | 0.32692 | 0.002204 | 48.4971 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5388 | -0.1026 | 0.003695 | 9.094 | 7.1534 | 11.8973 | 4.9791
2012-10-30 20:47 | 28.6869 | 23.641 | 21.6561 | 26.4243 | 19.4655 | -1.5107 | 0.006732 | 0.32692 | 0.002201 | 48.5619 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4901 | -0.1041 | 0.003693 | 9.1509 | 7.166 | 11.9342 | 4.9754
2012-10-30 20:47 | 28.6922 | 23.6102 | 21.6534 | 26.4137 | 19.4951 | -1.4957 | 0.006738 | 0.32692 | 0.002203 | 48.5187 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4628 | -0.1031 | 0.003693 | 9.1474 | 7.1906 | 11.9509 | 5.0323
2012-10-30 20:47 | 28.6869 | 23.5687 | 21.6279 | 26.4164 | 19.4534 | -1.5003 | 0.006722 | 0.32692 | 0.002198 | 48.6342 | 0.006955 | 0.21466 | 0.001493 | 30.8636 | 14.4412 | -0.1034 | 0.003691 | 9.1275 | 7.1867 | 11.9752 | 5.0122
2012-10-30 20:47 | 28.678 | 23.5558 | 21.5827 | 26.3516 | 19.4688 | -1.4928 | 0.006738 | 0.32696 | 0.002203 | 48.5251 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4486 | -0.1029 | 0.003693 | 9.1072 | 7.1341 | 11.903 | 5.0202
2012-10-30 20:48 | 28.6884 | 23.5099 | 21.6213 | 26.3101 | 19.4427 | -1.4922 | 0.006744 | 0.3269 | 0.002205 | 48.4733 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.508 | -0.1029 | 0.003695 | 9.0019 | 7.1133 | 11.8021 | 4.9347
2012-10-30 20:48 | 28.6989 | 23.5043 | 21.5874 | 26.3446 | 19.4451 | -1.4928 | 0.006735 | 0.32692 | 0.002202 | 48.5403 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.5144 | -0.1029 | 0.00369 | 8.9899 | 7.073 | 11.8302 | 4.9307
2012-10-30 20:48 | 28.6847 | 23.5102 | 21.5571 | 26.2785 | 19.4309 | -1.4991 | 0.006751 | 0.32693 | 0.002207 | 48.4273 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4879 | -0.1033 | 0.003697 | 9.0223 | 7.0692 | 11.7906 | 4.943
2012-10-30 20:48 | 28.6944 | 23.4556 | 21.5829 | 26.2682 | 19.4406 | -1.4916 | 0.006732 | 0.32688 | 0.002201 | 48.5554 | 0.006946 | 0.2146 | 0.001491 | 30.8953 | 14.4895 | -0.1028 | 0.003692 | 8.9661 | 7.0934 | 11.7787 | 4.9511
2012-10-30 20:48 | 28.6862 | 23.4756 | 21.5707 | 26.244 | 19.396 | -1.4905 | 0.006738 | 0.32692 | 0.002203 | 48.5187 | 0.006949 | 0.21463 | 0.001491 | 30.8862 | 14.469 | -0.1028 | 0.003694 | 9.0066 | 7.1017 | 11.775 | 4.927
2012-10-30 20:48 | 28.6989 | 23.4923 | 21.5431 | 26.2727 | 19.437 | -1.4974 | 0.006725 | 0.32689 | 0.002198 | 48.6082 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4695 | -0.1032 | 0.003688 | 9.0228 | 7.0736 | 11.8032 | 4.9675
2012-10-30 20:48 | 28.6967 | 23.466 | 21.5005 | 26.2545 | 19.4186 | -1.4997 | 0.006735 | 0.32689 | 0.002202 | 48.536 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4795 | -0.1034 | 0.003691 | 8.9865 | 7.021 | 11.775 | 4.9391
2012-10-30 20:48 | 28.6884 | 23.4175 | 21.5527 | 26.2023 | 19.4063 | -1.4997 | 0.006725 | 0.32688 | 0.002198 | 48.6059 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.508 | -0.1034 | 0.003689 | 8.9095 | 7.0447 | 11.6943 | 4.8983
2012-10-30 20:48 | 28.6839 | 23.4411 | 21.52 | 26.1579 | 19.422 | -1.4911 | 0.006741 | 0.3269 | 0.002204 | 48.4949 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.4994 | -0.1028 | 0.003696 | 8.9417 | 7.0206 | 11.6585 | 4.9226
2012-10-30 20:48 | 28.6854 | 23.3863 | 21.4892 | 26.1514 | 19.3952 | -1.4963 | 0.006732 | 0.32698 | 0.002201 | 48.5704 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4886 | -0.1032 | 0.003693 | 8.8977 | 7.0006 | 11.6628 | 4.9066
2012-10-30 20:48 | 28.6862 | 23.3228 | 21.5021 | 26.1322 | 19.396 | -1.4991 | 0.006741 | 0.32679 | 0.002203 | 48.4778 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.465 | -0.1033 | 0.00369 | 8.8578 | 7.0371 | 11.6672 | 4.931
2012-10-30 20:48 | 28.678 | 23.3226 | 21.4173 | 26.1041 | 19.3554 | -1.4928 | 0.006735 | 0.32688 | 0.002202 | 48.5338 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4853 | -0.1029 | 0.003693 | 8.8373 | 6.932 | 11.6188 | 4.8701
2012-10-30 20:48 | 28.6839 | 23.3044 | 21.4272 | 26.074 | 19.337 | -1.4997 | 0.006735 | 0.32692 | 0.002202 | 48.5403 | 0.006943 | 0.2146 | 0.00149 | 30.9087 | 14.4953 | -0.1034 | 0.003692 | 8.8091 | 6.9319 | 11.5787 | 4.8417
2012-10-30 20:48 | 28.6877 | 23.3122 | 21.4391 | 26.0698 | 19.3448 | -1.5009 | 0.006744 | 0.32692 | 0.002205 | 48.4755 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4787 | -0.1035 | 0.003694 | 8.8335 | 6.9604 | 11.5911 | 4.8661
2012-10-30 20:48 | 28.6967 | 23.3172 | 21.3956 | 26.0349 | 19.3498 | -1.498 | 0.006744 | 0.32695 | 0.002205 | 48.4798 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.5203 | -0.1033 | 0.003697 | 8.7969 | 6.8753 | 11.5146 | 4.8295
2012-10-30 20:48 | 28.6922 | 23.2403 | 21.4193 | 26.0024 | 19.3169 | -1.4853 | 0.006729 | 0.32688 | 0.0022 | 48.5771 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.5036 | -0.1024 | 0.003692 | 8.7367 | 6.9157 | 11.4988 | 4.8133
2012-10-30 20:48 | 28.681 | 23.2412 | 21.3678 | 25.9593 | 19.33 | -1.4928 | 0.006729 | 0.32682 | 0.002199 | 48.5686 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4434 | -0.1029 | 0.003687 | 8.7978 | 6.9244 | 11.5159 | 4.8866
2012-10-30 20:48 | 28.6839 | 23.224 | 21.3788 | 25.9821 | 19.3329 | -1.5015 | 0.006741 | 0.32692 | 0.002204 | 48.4971 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4953 | -0.1035 | 0.003694 | 8.7287 | 6.8835 | 11.4868 | 4.8376
2012-10-30 20:48 | 28.6884 | 23.1923 | 21.3469 | 25.9747 | 19.2929 | -1.4916 | 0.006732 | 0.32695 | 0.002201 | 48.5663 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4794 | -0.1028 | 0.003692 | 8.7129 | 6.8675 | 11.4953 | 4.8135
2012-10-30 20:48 | 28.6892 | 23.181 | 21.3598 | 25.9395 | 19.2937 | -1.4945 | 0.006719 | 0.32692 | 0.002197 | 48.6559 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.4598 | -0.103 | 0.003689 | 8.7212 | 6.9 | 11.4797 | 4.8339
2012-10-30 20:48 | 28.6862 | 23.1821 | 21.3447 | 25.9445 | 19.2826 | -1.4997 | 0.006738 | 0.3269 | 0.002203 | 48.5165 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4201 | -0.1034 | 0.003693 | 8.762 | 6.9246 | 11.5244 | 4.8625
2012-10-30 20:48 | 28.6899 | 23.0973 | 21.3484 | 25.8603 | 19.2944 | -1.5003 | 0.006738 | 0.32688 | 0.002202 | 48.5122 | 0.006927 | 0.21466 | 0.001487 | 30.9883 | 14.4646 | -0.1034 | 0.003689 | 8.6327 | 6.8838 | 11.3957 | 4.8298
2012-10-30 20:48 | 28.6854 | 23.1491 | 21.3197 | 25.8797 | 19.2332 | -1.4893 | 0.006751 | 0.32688 | 0.002207 | 48.4188 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4642 | -0.1027 | 0.003699 | 8.6849 | 6.8555 | 11.4155 | 4.769
2012-10-30 20:48 | 28.6877 | 23.1031 | 21.3382 | 25.8341 | 19.276 | -1.4957 | 0.006729 | 0.32685 | 0.002199 | 48.5727 | 0.006936 | 0.2146 | 0.001488 | 30.9399 | 14.442 | -0.1031 | 0.003687 | 8.6611 | 6.8962 | 11.3921 | 4.834
2012-10-30 20:48 | 28.6922 | 23.0996 | 21.2983 | 25.8746 | 19.2724 | -1.4899 | 0.006719 | 0.32689 | 0.002196 | 48.6516 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.4587 | -0.1027 | 0.003686 | 8.6409 | 6.8396 | 11.4159 | 4.8137
2012-10-30 20:48 | 28.6802 | 23.0876 | 21.3064 | 25.8026 | 19.2118 | -1.5049 | 0.006738 | 0.32688 | 0.002202 | 48.5122 | 0.006952 | 0.21461 | 0.001492 | 30.8707 | 14.4671 | -0.1037 | 0.003694 | 8.6205 | 6.8393 | 11.3355 | 4.7447
2012-10-30 20:49 | 28.6884 | 23.0917 | 21.2864 | 25.8308 | 19.2362 | -1.5095 | 0.006738 | 0.32686 | 0.002202 | 48.5101 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5324 | -0.1041 | 0.003691 | 8.5593 | 6.754 | 11.2984 | 4.7038
2012-10-30 20:49 | 28.6847 | 23.0478 | 21.2746 | 25.7551 | 19.2163 | -1.4911 | 0.006729 | 0.32686 | 0.002199 | 48.575 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.5043 | -0.1028 | 0.00369 | 8.5435 | 6.7703 | 11.2508 | 4.712
2012-10-30 20:49 | 28.6869 | 22.9977 | 21.2607 | 25.7294 | 19.2387 | -1.4905 | 0.006744 | 0.32682 | 0.002204 | 48.4606 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4779 | -0.1028 | 0.003693 | 8.5198 | 6.7828 | 11.2515 | 4.7608
2012-10-30 20:49 | 28.6817 | 22.9925 | 21.2635 | 25.7282 | 19.2295 | -1.4951 | 0.006735 | 0.32685 | 0.002201 | 48.5295 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4605 | -0.1031 | 0.003691 | 8.532 | 6.803 | 11.2677 | 4.769
2012-10-30 20:49 | 28.6772 | 22.9599 | 21.2025 | 25.6717 | 19.2169 | -1.4795 | 0.006735 | 0.3269 | 0.002202 | 48.5381 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4927 | -0.102 | 0.003694 | 8.4672 | 6.7098 | 11.179 | 4.7242
2012-10-30 20:49 | 28.678 | 22.9727 | 21.2114 | 25.6605 | 19.2217 | -1.4847 | 0.006725 | 0.32692 | 0.002199 | 48.6125 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4527 | -0.1024 | 0.00369 | 8.52 | 6.7587 | 11.2078 | 4.769
2012-10-30 20:49 | 28.681 | 22.9999 | 21.2225 | 25.7075 | 19.1923 | -1.476 | 0.006741 | 0.32685 | 0.002203 | 48.4863 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.472 | -0.1018 | 0.003692 | 8.5279 | 6.7505 | 11.2355 | 4.7203
2012-10-30 20:49 | 28.6787 | 22.9855 | 21.1959 | 25.6532 | 19.1819 | -1.4702 | 0.006722 | 0.32686 | 0.002197 | 48.6256 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.4901 | -0.1014 | 0.003684 | 8.4954 | 6.7058 | 11.1631 | 4.6918
2012-10-30 20:49 | 28.6802 | 22.9065 | 21.2055 | 25.6267 | 19.151 | -1.4818 | 0.006735 | 0.3269 | 0.002202 | 48.5381 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4182 | -0.1022 | 0.003691 | 8.4883 | 6.7873 | 11.2085 | 4.7328
2012-10-30 20:49 | 28.6824 | 22.8644 | 21.131 | 25.5929 | 19.1613 | -1.4853 | 0.006735 | 0.32688 | 0.002202 | 48.5338 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.4285 | -0.1024 | 0.003691 | 8.4359 | 6.7025 | 11.1644 | 4.7328
2012-10-30 20:49 | 28.6862 | 22.9085 | 21.159 | 25.5727 | 19.1772 | -1.472 | 0.006741 | 0.32695 | 0.002204 | 48.5014 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4201 | -0.1015 | 0.003693 | 8.4884 | 6.7389 | 11.1526 | 4.7571
2012-10-30 20:49 | 28.6802 | 22.8824 | 21.1288 | 25.5667 | 19.1591 | -1.4645 | 0.006741 | 0.32695 | 0.002204 | 48.5014 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4753 | -0.101 | 0.003695 | 8.4071 | 6.6535 | 11.0914 | 4.6838
2012-10-30 20:49 | 28.6869 | 22.8931 | 21.1274 | 25.5334 | 19.1334 | -1.4737 | 0.006751 | 0.3269 | 0.002207 | 48.423 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4371 | -0.1016 | 0.003698 | 8.456 | 6.6903 | 11.0963 | 4.6963
2012-10-30 20:49 | 28.6847 | 22.8185 | 21.101 | 25.5152 | 19.119 | -1.4789 | 0.006747 | 0.32686 | 0.002205 | 48.4454 | 0.006946 | 0.2146 | 0.001491 | 30.8953 | 14.4798 | -0.102 | 0.003696 | 8.3387 | 6.6212 | 11.0354 | 4.6392
2012-10-30 20:49 | 28.6877 | 22.8255 | 21.104 | 25.4902 | 19.122 | -1.4749 | 0.006735 | 0.32686 | 0.002201 | 48.5317 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4665 | -0.1017 | 0.003691 | 8.359 | 6.6375 | 11.0237 | 4.6555
2012-10-30 20:49 | 28.6862 | 22.8401 | 21.0662 | 25.4407 | 19.0881 | -1.4743 | 0.006738 | 0.3268 | 0.002202 | 48.5015 | 0.00693 | 0.21466 | 0.001488 | 30.9749 | 14.4568 | -0.1016 | 0.00369 | 8.3833 | 6.6094 | 10.9839 | 4.6313
2012-10-30 20:49 | 28.6742 | 22.7758 | 21.0784 | 25.4607 | 19.1207 | -1.4858 | 0.006735 | 0.32688 | 0.002202 | 48.5338 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.3795 | -0.1024 | 0.003691 | 8.3963 | 6.6989 | 11.0812 | 4.7412
2012-10-30 20:49 | 28.6854 | 22.7225 | 21.0694 | 25.4559 | 19.1157 | -1.4806 | 0.006732 | 0.32683 | 0.0022 | 48.549 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4397 | -0.1021 | 0.00369 | 8.2828 | 6.6297 | 11.0162 | 4.676
2012-10-30 20:49 | 28.6839 | 22.7533 | 21.0719 | 25.4304 | 19.098 | -1.4806 | 0.006735 | 0.32682 | 0.002201 | 48.5253 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4504 | -0.1021 | 0.003692 | 8.3029 | 6.6215 | 10.98 | 4.6476
2012-10-30 20:49 | 28.6757 | 22.7491 | 21.0314 | 25.3782 | 19.0371 | -1.4795 | 0.006741 | 0.32685 | 0.002203 | 48.4863 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4463 | -0.102 | 0.003692 | 8.3028 | 6.5851 | 10.9319 | 4.5908
2012-10-30 20:49 | 28.6847 | 22.7621 | 21.0243 | 25.3752 | 19.0177 | -1.4766 | 0.006729 | 0.32686 | 0.002199 | 48.575 | 0.006952 | 0.21466 | 0.001492 | 30.8769 | 14.4471 | -0.1018 | 0.003691 | 8.315 | 6.5772 | 10.9281 | 4.5706
2012-10-30 20:49 | 28.6824 | 22.6994 | 20.9776 | 25.3409 | 19.0397 | -1.4708 | 0.006744 | 0.32686 | 0.002204 | 48.4669 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4163 | -0.1014 | 0.003695 | 8.2831 | 6.5613 | 10.9246 | 4.6234
2012-10-30 20:49 | 28.6832 | 22.6519 | 21.0147 | 25.3377 | 19.0243 | -1.4835 | 0.006735 | 0.32686 | 0.002201 | 48.5317 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4416 | -0.1023 | 0.003688 | 8.2103 | 6.5731 | 10.8961 | 4.5827
2012-10-30 20:49 | 28.6817 | 22.6383 | 20.9849 | 25.2921 | 19.0228 | -1.4691 | 0.006732 | 0.32688 | 0.002201 | 48.5554 | 0.006958 | 0.21467 | 0.001494 | 30.8524 | 14.4278 | -0.1013 | 0.003695 | 8.2105 | 6.5571 | 10.8643 | 4.595
2012-10-30 20:49 | 28.6847 | 22.6494 | 20.9637 | 25.2631 | 18.9732 | -1.476 | 0.006744 | 0.32688 | 0.002204 | 48.469 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4594 | -0.1018 | 0.003695 | 8.19 | 6.5043 | 10.8037 | 4.5138
2012-10-30 20:49 | 28.6884 | 22.6329 | 20.9755 | 25.2588 | 18.9728 | -1.4795 | 0.006735 | 0.32685 | 0.002201 | 48.5295 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4631 | -0.102 | 0.003691 | 8.1698 | 6.5124 | 10.7957 | 4.5097
2012-10-30 20:49 | 28.681 | 22.6255 | 20.9236 | 25.2274 | 18.9897 | -1.4731 | 0.006729 | 0.32685 | 0.002199 | 48.5727 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4598 | -0.1016 | 0.003688 | 8.1657 | 6.4638 | 10.7676 | 4.5299
2012-10-30 20:50 | 28.6787 | 22.5628 | 20.9213 | 25.177 | 18.959 | -1.4824 | 0.006732 | 0.32686 | 0.0022 | 48.5533 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4452 | -0.1022 | 0.003689 | 8.1176 | 6.4761 | 10.7318 | 4.5138
2012-10-30 20:50 | 28.6847 | 22.5487 | 20.9314 | 25.2031 | 18.9205 | -1.4847 | 0.006747 | 0.32685 | 0.002205 | 48.4432 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4431 | -0.1024 | 0.003695 | 8.1056 | 6.4883 | 10.76 | 4.4774
2012-10-30 20:50 | 28.6795 | 22.5636 | 20.8818 | 25.1939 | 18.9517 | -1.4795 | 0.006729 | 0.32685 | 0.002199 | 48.5727 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4623 | -0.102 | 0.003688 | 8.1013 | 6.4195 | 10.7316 | 4.4894
2012-10-30 20:50 | 28.6877 | 22.4913 | 20.9101 | 25.182 | 18.9397 | -1.4783 | 0.006722 | 0.32685 | 0.002197 | 48.6233 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4461 | -0.1019 | 0.003688 | 8.0452 | 6.464 | 10.7359 | 4.4936
2012-10-30 20:50 | 28.6832 | 22.5029 | 20.8612 | 25.1655 | 18.9595 | -1.4818 | 0.006741 | 0.32682 | 0.002203 | 48.4821 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4456 | -0.1022 | 0.003692 | 8.0573 | 6.4156 | 10.7199 | 4.5139
2012-10-30 20:50 | 28.6854 | 22.5011 | 20.8715 | 25.1317 | 18.9374 | -1.483 | 0.006735 | 0.32686 | 0.002201 | 48.5317 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4438 | -0.1022 | 0.00369 | 8.0573 | 6.4277 | 10.6879 | 4.4936
2012-10-30 20:50 | 28.6922 | 22.4716 | 20.8298 | 25.0624 | 18.932 | -1.4783 | 0.006729 | 0.32679 | 0.002199 | 48.5643 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4465 | -0.1019 | 0.003689 | 8.0251 | 6.3833 | 10.6159 | 4.4855
2012-10-30 20:50 | 28.6892 | 22.4122 | 20.8591 | 25.0674 | 18.8763 | -1.4783 | 0.006741 | 0.32688 | 0.002203 | 48.4906 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4639 | -0.1019 | 0.003692 | 7.9483 | 6.3952 | 10.6035 | 4.4124
2012-10-30 20:50 | 28.6869 | 22.4462 | 20.8043 | 25.0611 | 18.8659 | -1.4835 | 0.006738 | 0.32695 | 0.002203 | 48.523 | 0.006961 | 0.21471 | 0.001495 | 30.8453 | 14.4249 | -0.1023 | 0.003698 | 8.0213 | 6.3794 | 10.6362 | 4.441
2012-10-30 20:50 | 28.6944 | 22.4134 | 20.8199 | 25.0606 | 18.9018 | -1.472 | 0.006719 | 0.32688 | 0.002196 | 48.6494 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4364 | -0.1015 | 0.003688 | 7.977 | 6.3835 | 10.6242 | 4.4654
2012-10-30 20:50 | 28.6922 | 22.4152 | 20.8258 | 24.9983 | 18.9118 | -1.4708 | 0.006741 | 0.3269 | 0.002204 | 48.4949 | 0.006924 | 0.21467 | 0.001486 | 31.0039 | 14.4628 | -0.1014 | 0.00369 | 7.9524 | 6.363 | 10.5355 | 4.449
2012-10-30 20:50 | 28.6989 | 22.3736 | 20.8204 | 25.0371 | 18.886 | -1.4754 | 0.006719 | 0.32692 | 0.002197 | 48.6559 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4613 | -0.1017 | 0.003687 | 7.9123 | 6.3591 | 10.5758 | 4.4247
2012-10-30 20:50 | 28.6952 | 22.386 | 20.8005 | 24.9653 | 18.8783 | -1.4645 | 0.006735 | 0.32679 | 0.002201 | 48.521 | 0.006946 | 0.2146 | 0.001491 | 30.8953 | 14.5148 | -0.101 | 0.003692 | 7.8712 | 6.2857 | 10.4505 | 4.3635
2012-10-30 20:50 | 28.7049 | 22.3715 | 20.8021 | 24.979 | 18.9082 | -1.4679 | 0.006729 | 0.32685 | 0.002199 | 48.5727 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5204 | -0.1012 | 0.003688 | 7.8511 | 6.2817 | 10.4586 | 4.3878
2012-10-30 20:50 | 28.6944 | 22.3449 | 20.8118 | 24.9444 | 18.8815 | -1.4789 | 0.006744 | 0.32677 | 0.002204 | 48.454 | 0.00693 | 0.21466 | 0.001488 | 30.9749 | 14.4936 | -0.102 | 0.003692 | 7.8513 | 6.3182 | 10.4508 | 4.3879
2012-10-30 20:50 | 28.6997 | 22.2616 | 20.7686 | 24.9217 | 18.8868 | -1.4766 | 0.006729 | 0.32686 | 0.002199 | 48.575 | 0.006961 | 0.21467 | 0.001494 | 30.8391 | 14.507 | -0.1018 | 0.003693 | 7.7546 | 6.2616 | 10.4147 | 4.3798
2012-10-30 20:50 | 28.6952 | 22.249 | 20.7884 | 24.9212 | 18.854 | -1.4812 | 0.006732 | 0.32683 | 0.0022 | 48.549 | 0.006961 | 0.21466 | 0.001494 | 30.8369 | 14.4862 | -0.1021 | 0.003694 | 7.7628 | 6.3022 | 10.435 | 4.3678
2012-10-30 20:50 | 28.7027 | 22.2807 | 20.7636 | 24.8766 | 18.8939 | -1.4835 | 0.006729 | 0.32683 | 0.002199 | 48.5707 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4978 | -0.1023 | 0.003688 | 7.7829 | 6.2658 | 10.3788 | 4.3961
2012-10-30 20:50 | 28.6944 | 22.2644 | 20.7229 | 24.8243 | 18.8613 | -1.4806 | 0.006744 | 0.32689 | 0.002205 | 48.4712 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.4936 | -0.1021 | 0.003692 | 7.7708 | 6.2293 | 10.3307 | 4.3677
2012-10-30 20:50 | 28.6974 | 22.2714 | 20.738 | 24.8473 | 18.844 | -1.4702 | 0.006738 | 0.32686 | 0.002202 | 48.5101 | 0.006943 | 0.2146 | 0.00149 | 30.9087 | 14.517 | -0.1014 | 0.003692 | 7.7544 | 6.221 | 10.3303 | 4.327
2012-10-30 20:50 | 28.6997 | 22.2374 | 20.6999 | 24.8095 | 18.8422 | -1.4674 | 0.006732 | 0.32685 | 0.0022 | 48.5511 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.4866 | -0.1012 | 0.003691 | 7.7508 | 6.2133 | 10.3229 | 4.3556
2012-10-30 20:50 | 28.6989 | 22.2487 | 20.7436 | 24.7807 | 18.7766 | -1.4737 | 0.006757 | 0.32685 | 0.002208 | 48.3715 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5266 | -0.1016 | 0.003697 | 7.7221 | 6.217 | 10.2541 | 4.25
2012-10-30 20:50 | 28.6959 | 22.149 | 20.7002 | 24.7817 | 18.7898 | -1.4934 | 0.006738 | 0.32676 | 0.002202 | 48.4951 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5277 | -0.103 | 0.003691 | 7.6213 | 6.1725 | 10.254 | 4.2621
2012-10-30 20:50 | 28.7027 | 22.1638 | 20.6706 | 24.7965 | 18.7885 | -1.4772 | 0.006732 | 0.3268 | 0.0022 | 48.5447 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5019 | -0.1018 | 0.003691 | 7.6619 | 6.1687 | 10.2946 | 4.2866
2012-10-30 20:50 | 28.6907 | 22.1277 | 20.6707 | 24.7604 | 18.7927 | -1.4737 | 0.006741 | 0.32688 | 0.002203 | 48.4906 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4695 | -0.1016 | 0.003694 | 7.6582 | 6.2012 | 10.2909 | 4.3232
2012-10-30 20:50 | 28.6937 | 22.0823 | 20.6454 | 24.7394 | 18.7673 | -1.4697 | 0.006738 | 0.32677 | 0.002202 | 48.4972 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5051 | -0.1013 | 0.003691 | 7.5772 | 6.1403 | 10.2343 | 4.2622
2012-10-30 20:51 | 28.6974 | 22.1142 | 20.6855 | 24.707 | 18.7427 | -1.4841 | 0.006738 | 0.32683 | 0.002202 | 48.5058 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4966 | -0.1023 | 0.003692 | 7.6176 | 6.1889 | 10.2104 | 4.2461
2012-10-30 20:51 | 28.6937 | 22.0904 | 20.6333 | 24.6552 | 18.739 | -1.476 | 0.006751 | 0.32677 | 0.002206 | 48.4038 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4725 | -0.1018 | 0.003696 | 7.6179 | 6.1608 | 10.1827 | 4.2665
2012-10-30 20:51 | 28.6974 | 22.0497 | 20.5966 | 24.6669 | 18.7427 | -1.4743 | 0.006738 | 0.32683 | 0.002202 | 48.5058 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.521 | -0.1016 | 0.003692 | 7.5287 | 6.0756 | 10.1459 | 4.2217
2012-10-30 20:51 | 28.6937 | 22.0621 | 20.6293 | 24.6392 | 18.739 | -1.4685 | 0.006722 | 0.32675 | 0.002196 | 48.6083 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.4521 | -0.1012 | 0.003686 | 7.61 | 6.1772 | 10.1871 | 4.2869
2012-10-30 20:51 | 28.6892 | 22.0133 | 20.6207 | 24.6387 | 18.7669 | -1.4766 | 0.006735 | 0.32672 | 0.0022 | 48.5102 | 0.006933 | 0.21454 | 0.001487 | 30.9449 | 14.4353 | -0.1018 | 0.003687 | 7.578 | 6.1854 | 10.2034 | 4.3316
2012-10-30 20:51 | 28.6959 | 22.0321 | 20.6072 | 24.6334 | 18.7087 | -1.461 | 0.006729 | 0.32673 | 0.002199 | 48.5557 | 0.00693 | 0.21461 | 0.001487 | 30.9687 | 14.4624 | -0.1007 | 0.003686 | 7.5697 | 6.1448 | 10.171 | 4.2463
2012-10-30 20:51 | 28.6877 | 22.0441 | 20.6112 | 24.5931 | 18.7167 | -1.4656 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.006949 | 0.21466 | 0.001492 | 30.8902 | 14.4746 | -0.101 | 0.003689 | 7.5695 | 6.1366 | 10.1185 | 4.2421
2012-10-30 20:51 | 28.6929 | 21.9928 | 20.5921 | 24.5822 | 18.7057 | -1.4783 | 0.006738 | 0.32676 | 0.002202 | 48.4951 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.5165 | -0.1019 | 0.003694 | 7.4763 | 6.0756 | 10.0657 | 4.1892
2012-10-30 20:51 | 28.6869 | 21.9828 | 20.5619 | 24.5522 | 18.6551 | -1.4749 | 0.006722 | 0.32679 | 0.002197 | 48.6148 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4861 | -0.1017 | 0.003686 | 7.4967 | 6.0758 | 10.0661 | 4.169
2012-10-30 20:51 | 28.6854 | 21.9692 | 20.5604 | 24.5106 | 18.6455 | -1.4702 | 0.006722 | 0.32672 | 0.002196 | 48.604 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4764 | -0.1014 | 0.003685 | 7.4928 | 6.084 | 10.0342 | 4.1691
2012-10-30 20:51 | 28.6817 | 22.0018 | 20.5445 | 24.4748 | 18.6864 | -1.4754 | 0.006735 | 0.3268 | 0.002201 | 48.5231 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.4482 | -0.1017 | 0.00369 | 7.5536 | 6.0963 | 10.0266 | 4.2382
2012-10-30 20:51 | 28.6802 | 21.964 | 20.5188 | 24.4853 | 18.6646 | -1.476 | 0.006735 | 0.32675 | 0.002201 | 48.5145 | 0.006952 | 0.21464 | 0.001492 | 30.8749 | 14.5079 | -0.1018 | 0.003693 | 7.4561 | 6.0109 | 9.9774 | 4.1567
2012-10-30 20:51 | 28.6862 | 21.9055 | 20.5288 | 24.4994 | 18.7112 | -1.4731 | 0.006725 | 0.32676 | 0.002197 | 48.5888 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4854 | -0.1016 | 0.003687 | 7.4201 | 6.0434 | 10.014 | 4.2258
2012-10-30 20:51 | 28.6802 | 21.9035 | 20.5026 | 24.4573 | 18.6444 | -1.4766 | 0.006741 | 0.32689 | 0.002204 | 48.4928 | 0.006952 | 0.21457 | 0.001492 | 30.8645 | 14.459 | -0.1018 | 0.003696 | 7.4445 | 6.0436 | 9.9983 | 4.1854
2012-10-30 20:51 | 28.6765 | 21.9038 | 20.4949 | 24.4014 | 18.6082 | -1.4893 | 0.006729 | 0.32685 | 0.002199 | 48.5727 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4716 | -0.1027 | 0.003688 | 7.4322 | 6.0233 | 9.9298 | 4.1366
2012-10-30 20:51 | 28.6817 | 21.8647 | 20.5082 | 24.4628 | 18.6256 | -1.461 | 0.006725 | 0.32682 | 0.002198 | 48.5975 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4849 | -0.1007 | 0.003688 | 7.3798 | 6.0233 | 9.9779 | 4.1407
2012-10-30 20:51 | 28.6869 | 21.8941 | 20.477 | 24.4078 | 18.574 | -1.4754 | 0.006741 | 0.32685 | 0.002203 | 48.4863 | 0.006936 | 0.21461 | 0.001489 | 30.9419 | 14.535 | -0.1017 | 0.003692 | 7.3591 | 5.942 | 9.8728 | 4.039
2012-10-30 20:51 | 28.6832 | 21.842 | 20.4652 | 24.4121 | 18.619 | -1.4749 | 0.006729 | 0.32682 | 0.002199 | 48.5686 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.4905 | -0.1017 | 0.003687 | 7.3515 | 5.9747 | 9.9216 | 4.1285
2012-10-30 20:51 | 28.6862 | 21.8047 | 20.4561 | 24.3911 | 18.6058 | -1.4702 | 0.006735 | 0.32679 | 0.002201 | 48.521 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.4568 | -0.1014 | 0.003693 | 7.3479 | 5.9993 | 9.9343 | 4.149
2012-10-30 20:51 | 28.6772 | 21.7674 | 20.4228 | 24.354 | 18.5887 | -1.4697 | 0.006729 | 0.32682 | 0.002199 | 48.5686 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4886 | -0.1013 | 0.00369 | 7.2788 | 5.9342 | 9.8654 | 4.1001
2012-10-30 20:51 | 28.6795 | 21.7738 | 20.4453 | 24.3082 | 18.5261 | -1.4818 | 0.006735 | 0.3268 | 0.002201 | 48.5231 | 0.006955 | 0.21463 | 0.001493 | 30.8595 | 14.4379 | -0.1022 | 0.003694 | 7.3359 | 6.0074 | 9.8703 | 4.0882
2012-10-30 20:51 | 28.6735 | 21.7637 | 20.4191 | 24.3022 | 18.5687 | -1.472 | 0.006751 | 0.32682 | 0.002206 | 48.4103 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.4727 | -0.1015 | 0.003699 | 7.291 | 5.9464 | 9.8295 | 4.096
2012-10-30 20:51 | 28.6772 | 21.6666 | 20.4228 | 24.2858 | 18.54 | -1.4679 | 0.006732 | 0.32686 | 0.0022 | 48.5533 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4764 | -0.1012 | 0.00369 | 7.1902 | 5.9464 | 9.8094 | 4.0636
2012-10-30 20:51 | 28.678 | 21.74 | 20.371 | 24.2826 | 18.5611 | -1.4766 | 0.006735 | 0.32688 | 0.002202 | 48.5338 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4568 | -0.1018 | 0.003693 | 7.2832 | 5.9142 | 9.8258 | 4.1043
2012-10-30 20:51 | 28.6765 | 21.7183 | 20.4059 | 24.261 | 18.5677 | -1.4749 | 0.006738 | 0.32673 | 0.002202 | 48.4908 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4797 | -0.1017 | 0.003694 | 7.2386 | 5.9262 | 9.7813 | 4.088
2012-10-30 20:51 | 28.672 | 21.7058 | 20.357 | 24.2084 | 18.5307 | -1.4772 | 0.006751 | 0.32673 | 0.002206 | 48.3974 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4467 | -0.1018 | 0.003695 | 7.2591 | 5.9103 | 9.7617 | 4.084
2012-10-30 20:51 | 28.6757 | 21.6974 | 20.3485 | 24.2161 | 18.5263 | -1.4789 | 0.006738 | 0.3267 | 0.002201 | 48.4865 | 0.006939 | 0.21455 | 0.001489 | 30.9202 | 14.4218 | -0.102 | 0.00369 | 7.2756 | 5.9267 | 9.7943 | 4.1045
2012-10-30 20:52 | 28.6824 | 21.7041 | 20.3431 | 24.1786 | 18.5006 | -1.4824 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4571 | -0.1022 | 0.003688 | 7.247 | 5.886 | 9.7215 | 4.0435
2012-10-30 20:52 | 28.678 | 21.6472 | 20.3427 | 24.2304 | 18.5043 | -1.4685 | 0.006738 | 0.32672 | 0.002201 | 48.4886 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4364 | -0.1012 | 0.00369 | 7.2108 | 5.9063 | 9.794 | 4.0679
2012-10-30 20:52 | 28.672 | 21.6533 | 20.3408 | 24.1602 | 18.5145 | -1.4679 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4712 | -0.1012 | 0.00369 | 7.1821 | 5.8696 | 9.689 | 4.0433
2012-10-30 20:52 | 28.6727 | 21.6016 | 20.3536 | 24.1449 | 18.5031 | -1.4645 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4678 | -0.101 | 0.003686 | 7.1338 | 5.8858 | 9.6771 | 4.0353
2012-10-30 20:52 | 28.6697 | 21.5704 | 20.2859 | 24.1057 | 18.4757 | -1.4662 | 0.006707 | 0.3268 | 0.002192 | 48.7257 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5097 | -0.1011 | 0.003683 | 7.0607 | 5.7762 | 9.596 | 3.966
2012-10-30 20:52 | 28.6735 | 21.6306 | 20.2978 | 24.1296 | 18.5404 | -1.4697 | 0.006732 | 0.3268 | 0.0022 | 48.5447 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4441 | -0.1013 | 0.003692 | 7.1865 | 5.8537 | 9.6855 | 4.0963
2012-10-30 20:52 | 28.6697 | 21.5986 | 20.2617 | 24.0857 | 18.4919 | -1.4662 | 0.006732 | 0.32688 | 0.002201 | 48.5554 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4444 | -0.1011 | 0.003688 | 7.1542 | 5.8173 | 9.6413 | 4.0475
2012-10-30 20:52 | 28.6667 | 21.5875 | 20.2384 | 24.0827 | 18.4849 | -1.4708 | 0.006722 | 0.32676 | 0.002196 | 48.6105 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4577 | -0.1014 | 0.003685 | 7.1298 | 5.7807 | 9.625 | 4.0272
2012-10-30 20:52 | 28.663 | 21.5354 | 20.3116 | 24.079 | 18.4487 | -1.4731 | 0.006732 | 0.32677 | 0.0022 | 48.5404 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4703 | -0.1016 | 0.003692 | 7.0651 | 5.8413 | 9.6087 | 3.9784
2012-10-30 20:52 | 28.6555 | 21.4876 | 20.2192 | 23.9832 | 18.425 | -1.4714 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4587 | -0.1014 | 0.003689 | 7.0289 | 5.7605 | 9.5245 | 3.9663
2012-10-30 20:52 | 28.648 | 21.4397 | 20.2117 | 23.8913 | 18.4216 | -1.4778 | 0.006738 | 0.32673 | 0.002202 | 48.4908 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4757 | -0.1019 | 0.003693 | 6.964 | 5.736 | 9.4156 | 3.9459
2012-10-30 20:52 | 28.6368 | 21.3599 | 20.1924 | 23.852 | 18.4104 | -1.4726 | 0.006716 | 0.3268 | 0.002195 | 48.6604 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.5257 | -0.1015 | 0.003684 | 6.8342 | 5.6667 | 9.3263 | 3.8847
2012-10-30 20:52 | 28.6442 | 21.3471 | 20.1512 | 23.8634 | 18.4137 | -1.4714 | 0.006747 | 0.32679 | 0.002205 | 48.4347 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4597 | -0.1014 | 0.003695 | 6.8874 | 5.6915 | 9.4037 | 3.954
2012-10-30 20:52 | 28.6428 | 21.3578 | 20.162 | 23.8018 | 18.4407 | -1.4737 | 0.006744 | 0.32672 | 0.002203 | 48.4454 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4991 | -0.1016 | 0.003695 | 6.8587 | 5.6629 | 9.3027 | 3.9416
2012-10-30 20:52 | 28.6435 | 21.3303 | 20.1101 | 23.8427 | 18.3725 | -1.4812 | 0.006747 | 0.3268 | 0.002205 | 48.4368 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5487 | -0.1021 | 0.003696 | 6.7816 | 5.5614 | 9.294 | 3.8238
2012-10-30 20:52 | 28.6383 | 21.2928 | 20.1777 | 23.7893 | 18.3875 | -1.4853 | 0.006719 | 0.3268 | 0.002196 | 48.6386 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.5027 | -0.1024 | 0.003684 | 6.7901 | 5.675 | 9.2866 | 3.8848
2012-10-30 20:53 | 28.639 | 21.2935 | 20.1218 | 23.7819 | 18.3477 | -1.4668 | 0.006754 | 0.3268 | 0.002207 | 48.3866 | 0.006924 | 0.21466 | 0.001486 | 31.0017 | 14.5157 | -0.1011 | 0.003693 | 6.7778 | 5.6061 | 9.2662 | 3.832
2012-10-30 20:53 | 28.633 | 21.2673 | 20.0875 | 23.7518 | 18.3336 | -1.4679 | 0.006735 | 0.32675 | 0.002201 | 48.5145 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.477 | -0.1012 | 0.003692 | 6.7903 | 5.6105 | 9.2748 | 3.8566
2012-10-30 20:53 | 28.6345 | 21.2607 | 20.0485 | 23.7373 | 18.3675 | -1.4697 | 0.006722 | 0.3268 | 0.002197 | 48.6169 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4908 | -0.1013 | 0.003688 | 6.7699 | 5.5577 | 9.2465 | 3.8767
2012-10-30 20:53 | 28.6345 | 21.289 | 20.097 | 23.7413 | 18.3229 | -1.4662 | 0.006738 | 0.32683 | 0.002202 | 48.5058 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4908 | -0.1011 | 0.00369 | 6.7982 | 5.6062 | 9.2505 | 3.8321
2012-10-30 20:53 | 28.642 | 21.2561 | 20.1248 | 23.7408 | 18.2817 | -1.4726 | 0.006744 | 0.32677 | 0.002204 | 48.454 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5432 | -0.1015 | 0.003693 | 6.7129 | 5.5816 | 9.1976 | 3.7385
2012-10-30 20:53 | 28.627 | 21.2129 | 20.0653 | 23.7097 | 18.2992 | -1.4731 | 0.006747 | 0.32688 | 0.002205 | 48.4475 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.4792 | -0.1016 | 0.003697 | 6.7337 | 5.5861 | 9.2305 | 3.82
2012-10-30 20:53 | 28.6353 | 21.2252 | 20.0372 | 23.6738 | 18.3196 | -1.4662 | 0.006729 | 0.32682 | 0.002199 | 48.5686 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.5161 | -0.1011 | 0.003687 | 6.7091 | 5.5211 | 9.1577 | 3.8035
2012-10-30 20:53 | 28.6323 | 21.1495 | 20.0827 | 23.6788 | 18.3126 | -1.461 | 0.006732 | 0.32682 | 0.0022 | 48.5469 | 0.006921 | 0.21466 | 0.001486 | 31.0152 | 14.5049 | -0.1007 | 0.003686 | 6.6446 | 5.5778 | 9.1739 | 3.8077
2012-10-30 20:53 | 28.6308 | 21.1642 | 20.0489 | 23.6412 | 18.3314 | -1.4737 | 0.006744 | 0.32677 | 0.002204 | 48.454 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5075 | -0.1016 | 0.003693 | 6.6567 | 5.5414 | 9.1337 | 3.8239
2012-10-30 20:53 | 28.6345 | 21.2002 | 20.0566 | 23.6288 | 18.3391 | -1.4911 | 0.006741 | 0.32682 | 0.002203 | 48.4821 | 0.006949 | 0.21457 | 0.001491 | 30.8778 | 14.5397 | -0.1028 | 0.003694 | 6.6605 | 5.5169 | 9.0891 | 3.7994
2012-10-30 20:53 | 28.6278 | 21.1491 | 20.0175 | 23.6101 | 18.2472 | -1.4812 | 0.006747 | 0.32682 | 0.002205 | 48.439 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.4922 | -0.1021 | 0.003698 | 6.6569 | 5.5253 | 9.1179 | 3.755
2012-10-30 20:53 | 28.6225 | 21.1317 | 20.0365 | 23.5847 | 18.2217 | -1.487 | 0.006747 | 0.32676 | 0.002205 | 48.4304 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.5359 | -0.1025 | 0.003692 | 6.5958 | 5.5006 | 9.0488 | 3.6858
2012-10-30 20:53 | 28.6285 | 21.0933 | 20.0304 | 23.5666 | 18.3088 | -1.4806 | 0.006738 | 0.32689 | 0.002203 | 48.5144 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4766 | -0.1021 | 0.003693 | 6.6167 | 5.5538 | 9.09 | 3.8322
2012-10-30 20:53 | 28.6323 | 21.0849 | 20.0342 | 23.5744 | 18.2761 | -1.4708 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.5008 | -0.1014 | 0.003693 | 6.5841 | 5.5334 | 9.0736 | 3.7753
2012-10-30 20:53 | 28.6315 | 21.0963 | 20.0172 | 23.5615 | 18.2469 | -1.476 | 0.006719 | 0.3268 | 0.002196 | 48.6386 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.5531 | -0.1018 | 0.003686 | 6.5432 | 5.4641 | 9.0084 | 3.6938
2012-10-30 20:53 | 28.6308 | 21.0754 | 19.9356 | 23.5608 | 18.2462 | -1.4766 | 0.006747 | 0.32679 | 0.002205 | 48.4347 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.536 | -0.1018 | 0.003696 | 6.5394 | 5.3996 | 9.0248 | 3.7102
2012-10-30 20:53 | 28.624 | 21.0605 | 19.949 | 23.5138 | 18.2921 | -1.4997 | 0.006735 | 0.32679 | 0.002201 | 48.521 | 0.006952 | 0.21468 | 0.001492 | 30.881 | 14.5333 | -0.1034 | 0.003693 | 6.5272 | 5.4157 | 8.9805 | 3.7588
2012-10-30 20:53 | 28.63 | 21.0423 | 20.0278 | 23.5239 | 18.2494 | -1.4801 | 0.006744 | 0.3268 | 0.002204 | 48.4583 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5108 | -0.102 | 0.003694 | 6.5315 | 5.517 | 9.0131 | 3.7386
2012-10-30 20:53 | 28.6338 | 21.0582 | 19.9871 | 23.5035 | 18.2248 | -1.4806 | 0.006741 | 0.32685 | 0.002203 | 48.4863 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4656 | -0.1021 | 0.003694 | 6.5926 | 5.5215 | 9.0379 | 3.7592
2012-10-30 20:53 | 28.63 | 21.0221 | 19.9429 | 23.4797 | 18.2535 | -1.4754 | 0.006735 | 0.32689 | 0.002202 | 48.536 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5271 | -0.1017 | 0.003692 | 6.495 | 5.4158 | 8.9526 | 3.7264
2012-10-30 20:53 | 28.6233 | 20.979 | 19.9402 | 23.5091 | 18.2427 | -1.4743 | 0.006741 | 0.32682 | 0.002203 | 48.4821 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.4837 | -0.1016 | 0.003692 | 6.4953 | 5.4565 | 9.0254 | 3.759
2012-10-30 20:53 | 28.6285 | 21.0044 | 19.909 | 23.438 | 18.1546 | -1.4789 | 0.006732 | 0.32685 | 0.0022 | 48.5511 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.497 | -0.102 | 0.003692 | 6.5074 | 5.412 | 8.941 | 3.6576
2012-10-30 20:53 | 28.63 | 20.9898 | 19.9226 | 23.4113 | 18.1642 | -1.487 | 0.006722 | 0.32686 | 0.002197 | 48.6256 | 0.006952 | 0.21467 | 0.001492 | 30.879 | 14.5067 | -0.1025 | 0.003689 | 6.4831 | 5.4159 | 8.9046 | 3.6575
2012-10-30 20:53 | 28.63 | 21.0019 | 19.9186 | 23.4354 | 18.1724 | -1.4766 | 0.006744 | 0.32677 | 0.002204 | 48.454 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.5067 | -0.1018 | 0.003692 | 6.4952 | 5.4119 | 8.9287 | 3.6657
2012-10-30 20:53 | 28.624 | 20.9959 | 19.864 | 23.3732 | 18.1826 | -1.4795 | 0.006751 | 0.32679 | 0.002206 | 48.406 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.5048 | -0.102 | 0.003693 | 6.4911 | 5.3592 | 8.8684 | 3.6778
2012-10-30 20:53 | 28.6278 | 20.9755 | 19.9083 | 23.373 | 18.2026 | -1.4997 | 0.006713 | 0.32677 | 0.002194 | 48.6778 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.4841 | -0.1034 | 0.003685 | 6.4914 | 5.4242 | 8.8889 | 3.7185
2012-10-30 20:53 | 28.6225 | 20.9459 | 19.9151 | 23.4038 | 18.1283 | -1.4795 | 0.006722 | 0.3268 | 0.002197 | 48.6169 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4502 | -0.102 | 0.003686 | 6.4957 | 5.4649 | 8.9536 | 3.6781
2012-10-30 20:53 | 28.6225 | 21.0025 | 19.903 | 23.3596 | 18.1567 | -1.4835 | 0.006725 | 0.32682 | 0.002198 | 48.5975 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.4665 | -0.1023 | 0.003691 | 6.536 | 5.4365 | 8.8931 | 3.6902
2012-10-30 20:54 | 28.6173 | 20.9326 | 19.8412 | 23.3142 | 18.1799 | -1.4801 | 0.006741 | 0.32676 | 0.002203 | 48.4735 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4573 | -0.102 | 0.003691 | 6.4753 | 5.3839 | 8.8569 | 3.7226
2012-10-30 20:54 | 28.6195 | 20.8864 | 19.8676 | 23.3647 | 18.17 | -1.4778 | 0.006732 | 0.32679 | 0.0022 | 48.5426 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.488 | -0.1019 | 0.003691 | 6.3984 | 5.3796 | 8.8767 | 3.682
2012-10-30 20:54 | 28.6173 | 20.9125 | 19.829 | 23.3142 | 18.184 | -1.4754 | 0.006744 | 0.32686 | 0.002204 | 48.4669 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4205 | -0.1017 | 0.003693 | 6.492 | 5.4085 | 8.8937 | 3.7635
2012-10-30 20:54 | 28.6218 | 20.8806 | 19.7566 | 23.3227 | 18.1601 | -1.4911 | 0.006741 | 0.32682 | 0.002203 | 48.4821 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4536 | -0.1028 | 0.003693 | 6.427 | 5.303 | 8.8691 | 3.7065
2012-10-30 20:54 | 28.6128 | 20.8797 | 19.78 | 23.3338 | 18.1592 | -1.4864 | 0.006741 | 0.3268 | 0.002203 | 48.4799 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.4528 | -0.1025 | 0.003696 | 6.4269 | 5.3272 | 8.881 | 3.7064
2012-10-30 20:54 | 28.618 | 20.8566 | 19.7812 | 23.2707 | 18.1685 | -1.4766 | 0.006735 | 0.32683 | 0.002201 | 48.5274 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4702 | -0.1018 | 0.003688 | 6.3864 | 5.311 | 8.8005 | 3.6983
2012-10-30 20:54 | 28.6068 | 20.8414 | 19.8064 | 23.2595 | 18.1532 | -1.487 | 0.006735 | 0.32683 | 0.002201 | 48.5274 | 0.006927 | 0.21461 | 0.001487 | 30.9821 | 14.4631 | -0.1025 | 0.003688 | 6.3783 | 5.3433 | 8.7964 | 3.6901
2012-10-30 20:54 | 28.6158 | 20.8584 | 19.7952 | 23.2765 | 18.1379 | -1.487 | 0.006751 | 0.32679 | 0.002206 | 48.406 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4762 | -0.1025 | 0.003694 | 6.3822 | 5.319 | 8.8003 | 3.6617
2012-10-30 20:54 | 28.6165 | 20.8632 | 19.7918 | 23.2491 | 18.1345 | -1.4876 | 0.006725 | 0.32682 | 0.002198 | 48.5975 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4646 | -0.1026 | 0.003687 | 6.3986 | 5.3272 | 8.7845 | 3.6699
2012-10-30 20:54 | 28.615 | 20.8253 | 19.7701 | 23.2034 | 18.1492 | -1.476 | 0.006747 | 0.32682 | 0.002205 | 48.439 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4876 | -0.1018 | 0.003694 | 6.3377 | 5.2825 | 8.7158 | 3.6616
2012-10-30 20:54 | 28.612 | 20.8344 | 19.7994 | 23.2245 | 18.1097 | -1.4656 | 0.006741 | 0.32683 | 0.002203 | 48.4842 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4642 | -0.101 | 0.003694 | 6.3702 | 5.3352 | 8.7603 | 3.6455
2012-10-30 20:54 | 28.6083 | 20.7863 | 19.7957 | 23.2248 | 18.0979 | -1.4864 | 0.006738 | 0.32682 | 0.002202 | 48.5037 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4972 | -0.1025 | 0.003692 | 6.2891 | 5.2985 | 8.7276 | 3.6007
2012-10-30 20:54 | 28.606 | 20.7759 | 19.7732 | 23.1944 | 18.0875 | -1.4778 | 0.006722 | 0.32686 | 0.002197 | 48.6256 | 0.006949 | 0.21464 | 0.001492 | 30.8882 | 14.446 | -0.1019 | 0.003689 | 6.3299 | 5.3272 | 8.7484 | 3.6415
2012-10-30 20:54 | 28.6083 | 20.8065 | 19.7027 | 23.1967 | 18.0776 | -1.4801 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4931 | -0.102 | 0.003691 | 6.3134 | 5.2096 | 8.7036 | 3.5845
2012-10-30 20:54 | 28.6075 | 20.729 | 19.7181 | 23.1597 | 18.0403 | -1.4772 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006921 | 0.21467 | 0.001486 | 31.0173 | 14.4923 | -0.1018 | 0.003686 | 6.2367 | 5.2258 | 8.6674 | 3.548
2012-10-30 20:54 | 28.6053 | 20.747 | 19.7199 | 23.0931 | 18.0787 | -1.4749 | 0.006738 | 0.32685 | 0.002202 | 48.5079 | 0.006952 | 0.21466 | 0.001492 | 30.8769 | 14.4942 | -0.1017 | 0.003694 | 6.2528 | 5.2257 | 8.5989 | 3.5845
2012-10-30 20:54 | 28.6068 | 20.6757 | 19.7295 | 23.0986 | 18.0924 | -1.4772 | 0.006722 | 0.32676 | 0.002196 | 48.6105 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4468 | -0.1018 | 0.003688 | 6.2289 | 5.2827 | 8.6518 | 3.6456
2012-10-30 20:54 | 28.609 | 20.6739 | 19.7196 | 23.0727 | 18.0175 | -1.4749 | 0.006738 | 0.3268 | 0.002202 | 48.5015 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4938 | -0.1017 | 0.003694 | 6.1801 | 5.2258 | 8.5789 | 3.5237
2012-10-30 20:54 | 28.609 | 20.7507 | 19.671 | 23.1129 | 18.0743 | -1.4783 | 0.006738 | 0.32683 | 0.002202 | 48.5058 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5142 | -0.1019 | 0.003693 | 6.2365 | 5.1568 | 8.5987 | 3.5601
2012-10-30 20:54 | 28.6098 | 20.6949 | 19.6637 | 23.0695 | 18.0304 | -1.4806 | 0.006735 | 0.3267 | 0.0022 | 48.5081 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4538 | -0.1021 | 0.00369 | 6.2411 | 5.2099 | 8.6157 | 3.5766
2012-10-30 20:54 | 28.612 | 20.6688 | 19.6983 | 23.0958 | 18.0123 | -1.4841 | 0.006725 | 0.32676 | 0.002197 | 48.5888 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.4642 | -0.1023 | 0.00369 | 6.2046 | 5.2341 | 8.6316 | 3.5481
2012-10-30 20:54 | 28.612 | 20.6284 | 19.6699 | 23.0395 | 17.9636 | -1.4905 | 0.006747 | 0.32685 | 0.002205 | 48.4432 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.5009 | -0.1028 | 0.003694 | 6.1275 | 5.169 | 8.5386 | 3.4627
2012-10-30 20:54 | 28.6083 | 20.6005 | 19.6743 | 23.0157 | 18.0208 | -1.4847 | 0.006725 | 0.32683 | 0.002198 | 48.5996 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.485 | -0.1024 | 0.003688 | 6.1155 | 5.1893 | 8.5307 | 3.5358
2012-10-30 20:54 | 28.6016 | 20.5695 | 19.6231 | 23.0251 | 18.0263 | -1.4882 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4375 | -0.1026 | 0.003691 | 6.132 | 5.1856 | 8.5876 | 3.5888
2012-10-30 20:54 | 28.6188 | 20.6352 | 19.612 | 22.99 | 18.0232 | -1.4812 | 0.006732 | 0.32682 | 0.0022 | 48.5469 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5404 | -0.1021 | 0.00369 | 6.0948 | 5.0716 | 8.4496 | 3.4828
2012-10-30 20:55 | 28.6105 | 20.5703 | 19.6077 | 22.9857 | 17.9824 | -1.4818 | 0.006735 | 0.32683 | 0.002201 | 48.5274 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4627 | -0.1022 | 0.00369 | 6.1076 | 5.145 | 8.523 | 3.5197
2012-10-30 20:55 | 28.6128 | 20.5686 | 19.61 | 22.9156 | 17.9482 | -1.4847 | 0.006741 | 0.32682 | 0.002203 | 48.4821 | 0.00693 | 0.21461 | 0.001487 | 30.9687 | 14.4691 | -0.1024 | 0.00369 | 6.0995 | 5.1409 | 8.4465 | 3.4791
2012-10-30 20:55 | 28.6158 | 20.5352 | 19.6049 | 22.9548 | 17.9756 | -1.4887 | 0.006732 | 0.32677 | 0.0022 | 48.5404 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.468 | -0.1026 | 0.00369 | 6.0672 | 5.1369 | 8.4868 | 3.5076
2012-10-30 20:55 | 28.6203 | 20.5195 | 19.5811 | 22.9311 | 17.9801 | -1.4714 | 0.006747 | 0.32682 | 0.002205 | 48.439 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5133 | -0.1014 | 0.003696 | 6.0062 | 5.0678 | 8.4178 | 3.4668
2012-10-30 20:55 | 28.6158 | 20.5312 | 19.5766 | 22.9266 | 17.9187 | -1.4899 | 0.006741 | 0.32677 | 0.002203 | 48.4756 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4598 | -0.1027 | 0.003695 | 6.0714 | 5.1168 | 8.4668 | 3.4589
2012-10-30 20:55 | 28.6195 | 20.4904 | 19.5682 | 22.9464 | 17.9468 | -1.487 | 0.006741 | 0.32679 | 0.002203 | 48.4778 | 0.006939 | 0.21461 | 0.001489 | 30.9285 | 14.4513 | -0.1025 | 0.003692 | 6.0391 | 5.1169 | 8.4951 | 3.4955
2012-10-30 20:55 | 28.6248 | 20.4755 | 19.5654 | 22.8954 | 17.9643 | -1.4899 | 0.006735 | 0.32679 | 0.002201 | 48.521 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.5178 | -0.1027 | 0.003692 | 5.9577 | 5.0476 | 8.3776 | 3.4465
2012-10-30 20:55 | 28.6188 | 20.4857 | 19.5432 | 22.8572 | 17.9461 | -1.4818 | 0.006722 | 0.32672 | 0.002196 | 48.604 | 0.006933 | 0.21458 | 0.001488 | 30.9511 | 14.4506 | -0.1022 | 0.003684 | 6.0351 | 5.0926 | 8.4066 | 3.4955
2012-10-30 20:55 | 28.6053 | 20.4479 | 19.5742 | 22.8517 | 17.9488 | -1.4824 | 0.006744 | 0.32672 | 0.002203 | 48.4454 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4738 | -0.1022 | 0.003693 | 5.9741 | 5.1004 | 8.3779 | 3.475
2012-10-30 20:55 | 28.6143 | 20.469 | 19.5346 | 22.8446 | 17.9172 | -1.4899 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4747 | -0.1027 | 0.003691 | 5.9943 | 5.0599 | 8.3699 | 3.4425
2012-10-30 20:55 | 28.615 | 20.4253 | 19.5475 | 22.789 | 17.9423 | -1.4876 | 0.006725 | 0.3267 | 0.002197 | 48.5802 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4754 | -0.1026 | 0.003688 | 5.9499 | 5.0721 | 8.3136 | 3.4669
2012-10-30 20:55 | 28.6158 | 20.414 | 19.528 | 22.8019 | 17.8863 | -1.4893 | 0.006725 | 0.32677 | 0.002198 | 48.5909 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.415 | -0.1027 | 0.003688 | 5.999 | 5.113 | 8.3869 | 3.4713
2012-10-30 20:55 | 28.6105 | 20.4531 | 19.5308 | 22.7845 | 17.9175 | -1.4893 | 0.006747 | 0.32676 | 0.002205 | 48.4304 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4097 | -0.1027 | 0.003696 | 6.0434 | 5.1211 | 8.3748 | 3.5078
2012-10-30 20:55 | 28.615 | 20.4132 | 19.5434 | 22.8011 | 17.9585 | -1.4864 | 0.006738 | 0.32679 | 0.002202 | 48.4994 | 0.006924 | 0.21461 | 0.001486 | 30.9955 | 14.4223 | -0.1025 | 0.003688 | 5.9909 | 5.1211 | 8.3788 | 3.5362
2012-10-30 20:55 | 28.6068 | 20.409 | 19.4826 | 22.8009 | 17.8894 | -1.4968 | 0.006732 | 0.32683 | 0.0022 | 48.549 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4549 | -0.1032 | 0.003688 | 5.9541 | 5.0277 | 8.346 | 3.4345
2012-10-30 20:55 | 28.609 | 20.395 | 19.4645 | 22.7508 | 17.8957 | -1.487 | 0.006741 | 0.32679 | 0.002203 | 48.4778 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4408 | -0.1025 | 0.003693 | 5.9542 | 5.0237 | 8.31 | 3.4549
2012-10-30 20:55 | 28.6053 | 20.3469 | 19.4689 | 22.7672 | 17.9082 | -1.483 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.482 | -0.1022 | 0.003687 | 5.8649 | 4.9869 | 8.2852 | 3.4262
2012-10-30 20:55 | 28.5993 | 20.3287 | 19.4872 | 22.7612 | 17.9185 | -1.4974 | 0.006722 | 0.32672 | 0.002196 | 48.604 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4515 | -0.1032 | 0.003684 | 5.8772 | 5.0357 | 8.3097 | 3.467
2012-10-30 20:55 | 28.5986 | 20.3563 | 19.5068 | 22.7243 | 17.8569 | -1.4847 | 0.006729 | 0.32676 | 0.002199 | 48.56 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4794 | -0.1024 | 0.003689 | 5.8769 | 5.0274 | 8.2449 | 3.3775
2012-10-30 20:55 | 28.6023 | 20.3398 | 19.4619 | 22.72 | 17.8768 | -1.483 | 0.006738 | 0.32677 | 0.002202 | 48.4972 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4831 | -0.1022 | 0.003691 | 5.8567 | 4.9788 | 8.2369 | 3.3937
2012-10-30 20:55 | 28.606 | 20.2869 | 19.4615 | 22.6955 | 17.8846 | -1.4882 | 0.006738 | 0.3267 | 0.002201 | 48.4865 | 0.006949 | 0.21458 | 0.001491 | 30.8798 | 14.5112 | -0.1026 | 0.003692 | 5.7757 | 4.9503 | 8.1843 | 3.3734
2012-10-30 20:55 | 28.6098 | 20.319 | 19.4329 | 22.6671 | 17.8843 | -1.4951 | 0.006725 | 0.32673 | 0.002197 | 48.5845 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.5314 | -0.1031 | 0.003688 | 5.7876 | 4.9015 | 8.1357 | 3.3529
2012-10-30 20:55 | 28.609 | 20.294 | 19.4119 | 22.6582 | 17.8754 | -1.4911 | 0.006738 | 0.32677 | 0.002202 | 48.4972 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4979 | -0.1028 | 0.003692 | 5.7961 | 4.914 | 8.1603 | 3.3775
2012-10-30 20:55 | 28.5986 | 20.2714 | 19.361 | 22.6317 | 17.8447 | -1.4864 | 0.006754 | 0.32679 | 0.002207 | 48.3845 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4345 | -0.1025 | 0.003696 | 5.8369 | 4.9265 | 8.1972 | 3.4102
2012-10-30 20:55 | 28.612 | 20.297 | 19.4028 | 22.6612 | 17.8459 | -1.4928 | 0.006722 | 0.32676 | 0.002196 | 48.6105 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4724 | -0.1029 | 0.003687 | 5.8246 | 4.9304 | 8.1888 | 3.3735
2012-10-30 20:55 | 28.6128 | 20.2331 | 19.44 | 22.6177 | 17.8346 | -1.4968 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.4772 | -0.1032 | 0.00369 | 5.7559 | 4.9628 | 8.1405 | 3.3574
2012-10-30 20:55 | 28.6083 | 20.2326 | 19.3667 | 22.577 | 17.8098 | -1.4911 | 0.006725 | 0.32672 | 0.002197 | 48.5823 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.4646 | -0.1028 | 0.003684 | 5.768 | 4.9021 | 8.1124 | 3.3452
2012-10-30 20:56 | 28.6046 | 20.2046 | 19.3954 | 22.533 | 17.7939 | -1.4847 | 0.006732 | 0.32675 | 0.0022 | 48.5361 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5017 | -0.1024 | 0.003692 | 5.7029 | 4.8937 | 8.0313 | 3.2922
2012-10-30 20:56 | 28.612 | 20.208 | 19.3744 | 22.5686 | 17.781 | -1.4922 | 0.006729 | 0.32676 | 0.002199 | 48.56 | 0.006955 | 0.21474 | 0.001494 | 30.8761 | 14.4683 | -0.1029 | 0.003693 | 5.7397 | 4.9061 | 8.1003 | 3.3127
2012-10-30 20:56 | 28.6038 | 20.1998 | 19.3784 | 22.5202 | 17.7444 | -1.5038 | 0.006735 | 0.32675 | 0.002201 | 48.5145 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.4397 | -0.1037 | 0.003693 | 5.7601 | 4.9387 | 8.0805 | 3.3047
2012-10-30 20:56 | 28.6158 | 20.1916 | 19.362 | 22.5523 | 17.7685 | -1.4905 | 0.006735 | 0.32672 | 0.0022 | 48.5102 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.5047 | -0.1028 | 0.003688 | 5.6869 | 4.8573 | 8.0476 | 3.2638
2012-10-30 20:56 | 28.6135 | 20.1448 | 19.3111 | 22.5379 | 17.7622 | -1.4963 | 0.006747 | 0.32676 | 0.002205 | 48.4304 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4779 | -0.1032 | 0.003692 | 5.6669 | 4.8332 | 8.06 | 3.2843
2012-10-30 20:56 | 28.6105 | 20.1944 | 19.3324 | 22.4987 | 17.743 | -1.4882 | 0.006741 | 0.32676 | 0.002203 | 48.4735 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.4627 | -0.1026 | 0.00369 | 5.7317 | 4.8697 | 8.036 | 3.2803
2012-10-30 20:56 | 28.6135 | 20.1812 | 19.3152 | 22.4815 | 17.7622 | -1.4905 | 0.006732 | 0.3267 | 0.002199 | 48.5297 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.5228 | -0.1028 | 0.003686 | 5.6584 | 4.7924 | 7.9587 | 3.2394
2012-10-30 20:56 | 28.6105 | 20.1458 | 19.3162 | 22.4665 | 17.7592 | -1.487 | 0.006719 | 0.32676 | 0.002196 | 48.6322 | 0.006933 | 0.21463 | 0.001488 | 30.9575 | 14.4627 | -0.1025 | 0.003684 | 5.6831 | 4.8535 | 8.0038 | 3.2965
2012-10-30 20:56 | 28.615 | 20.0775 | 19.2843 | 22.4508 | 17.7393 | -1.4905 | 0.006735 | 0.32663 | 0.0022 | 48.4974 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4794 | -0.1028 | 0.003687 | 5.5981 | 4.8049 | 7.9714 | 3.2599
2012-10-30 20:56 | 28.6143 | 20.1254 | 19.2633 | 22.43 | 17.7427 | -1.4997 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4543 | -0.1034 | 0.003689 | 5.6711 | 4.809 | 7.9757 | 3.2884
2012-10-30 20:56 | 28.6143 | 20.0849 | 19.2957 | 22.4018 | 17.7224 | -1.4963 | 0.006729 | 0.32675 | 0.002199 | 48.5577 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4624 | -0.1032 | 0.003688 | 5.6225 | 4.8333 | 7.9394 | 3.26
2012-10-30 20:56 | 28.6233 | 20.0535 | 19.3007 | 22.4068 | 17.7679 | -1.4876 | 0.006735 | 0.32669 | 0.0022 | 48.5059 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.5 | -0.1026 | 0.003689 | 5.5535 | 4.8007 | 7.9068 | 3.2679
2012-10-30 20:56 | 28.618 | 20.0765 | 19.2468 | 22.4297 | 17.7342 | -1.4957 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.462 | -0.1031 | 0.003688 | 5.6145 | 4.7848 | 7.9677 | 3.2722
2012-10-30 20:56 | 28.6195 | 20.0699 | 19.224 | 22.411 | 17.7032 | -1.4905 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.5166 | -0.1028 | 0.003688 | 5.5533 | 4.7074 | 7.8944 | 3.1866
2012-10-30 20:56 | 28.6143 | 20.0081 | 19.1945 | 22.3575 | 17.6899 | -1.4899 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4543 | -0.1027 | 0.00369 | 5.5538 | 4.7402 | 7.9032 | 3.2356
2012-10-30 20:56 | 28.6263 | 20.0282 | 19.2794 | 22.3493 | 17.7506 | -1.4911 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.4744 | -0.1028 | 0.003687 | 5.5538 | 4.805 | 7.8749 | 3.2762
2012-10-30 20:56 | 28.6278 | 20.0418 | 19.2687 | 22.3508 | 17.6994 | -1.487 | 0.006741 | 0.32682 | 0.002203 | 48.4821 | 0.006955 | 0.21474 | 0.001494 | 30.8761 | 14.4718 | -0.1025 | 0.003697 | 5.57 | 4.7969 | 7.879 | 3.2276
2012-10-30 20:56 | 28.624 | 20.0178 | 19.2244 | 22.3752 | 17.6834 | -1.4858 | 0.006729 | 0.32673 | 0.002199 | 48.5557 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4966 | -0.1024 | 0.003691 | 5.5212 | 4.7278 | 7.8786 | 3.1868
2012-10-30 20:56 | 28.6195 | 19.9607 | 19.2321 | 22.3305 | 17.6829 | -1.4899 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4472 | -0.1027 | 0.003689 | 5.5135 | 4.7849 | 7.8833 | 3.2357
2012-10-30 20:56 | 28.6225 | 19.9232 | 19.2472 | 22.3012 | 17.6859 | -1.4922 | 0.006735 | 0.32667 | 0.0022 | 48.5038 | 0.006933 | 0.2146 | 0.001488 | 30.9533 | 14.4421 | -0.1029 | 0.003688 | 5.4811 | 4.8051 | 7.8591 | 3.2438
2012-10-30 20:56 | 28.624 | 19.949 | 19.1758 | 22.2826 | 17.6793 | -1.4847 | 0.006738 | 0.32676 | 0.002202 | 48.4951 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4884 | -0.1024 | 0.00369 | 5.4606 | 4.6874 | 7.7942 | 3.1909
2012-10-30 20:56 | 28.6278 | 19.9852 | 19.2201 | 22.2582 | 17.6506 | -1.4905 | 0.006735 | 0.32675 | 0.002201 | 48.5145 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4963 | -0.1028 | 0.00369 | 5.4889 | 4.7238 | 7.7619 | 3.1543
2012-10-30 20:56 | 28.624 | 19.9531 | 19.1312 | 22.2745 | 17.6265 | -1.4824 | 0.006735 | 0.32673 | 0.002201 | 48.5124 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4558 | -0.1022 | 0.00369 | 5.4973 | 4.6754 | 7.8187 | 3.1707
2012-10-30 20:56 | 28.618 | 19.943 | 19.1577 | 22.2403 | 17.6165 | -1.498 | 0.006741 | 0.32669 | 0.002202 | 48.4627 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4376 | -0.1033 | 0.003692 | 5.5054 | 4.7201 | 7.8027 | 3.1789
2012-10-30 20:57 | 28.624 | 19.9166 | 19.1718 | 22.2383 | 17.6549 | -1.476 | 0.006738 | 0.3267 | 0.002201 | 48.4865 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4517 | -0.1018 | 0.003691 | 5.4649 | 4.7201 | 7.7866 | 3.2032
2012-10-30 20:57 | 28.6353 | 19.9279 | 19.1831 | 22.2335 | 17.6419 | -1.4899 | 0.006741 | 0.32669 | 0.002202 | 48.4627 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4671 | -0.1027 | 0.00369 | 5.4608 | 4.716 | 7.7664 | 3.1748
2012-10-30 20:57 | 28.6323 | 19.9007 | 19.1882 | 22.1821 | 17.6836 | -1.4899 | 0.006741 | 0.32675 | 0.002203 | 48.4713 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4763 | -0.1027 | 0.003693 | 5.4244 | 4.7119 | 7.7058 | 3.2073
2012-10-30 20:57 | 28.6315 | 19.8675 | 19.1266 | 22.2095 | 17.6543 | -1.4934 | 0.006738 | 0.32669 | 0.002201 | 48.4843 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4674 | -0.103 | 0.003693 | 5.4001 | 4.6592 | 7.7421 | 3.1869
2012-10-30 20:57 | 28.645 | 19.8689 | 19.1644 | 22.1787 | 17.6191 | -1.4806 | 0.006735 | 0.3267 | 0.0022 | 48.5081 | 0.006921 | 0.21471 | 0.001486 | 31.0236 | 14.5094 | -0.1021 | 0.003686 | 5.3595 | 4.655 | 7.6693 | 3.1097
2012-10-30 20:57 | 28.6405 | 19.8603 | 19.1113 | 22.1984 | 17.6268 | -1.4945 | 0.006738 | 0.32667 | 0.002201 | 48.4822 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4805 | -0.103 | 0.003689 | 5.3798 | 4.6308 | 7.7179 | 3.1463
2012-10-30 20:57 | 28.6405 | 19.836 | 19.0829 | 22.1621 | 17.6146 | -1.5026 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4927 | -0.1036 | 0.003691 | 5.3433 | 4.5902 | 7.6694 | 3.1219
2012-10-30 20:57 | 28.6435 | 19.7824 | 19.0819 | 22.157 | 17.6135 | -1.4991 | 0.006747 | 0.3268 | 0.002205 | 48.4368 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.4875 | -0.1033 | 0.003697 | 5.2949 | 4.5944 | 7.6695 | 3.126
2012-10-30 20:57 | 28.6413 | 19.8004 | 19.0918 | 22.1226 | 17.6154 | -1.5078 | 0.006732 | 0.32675 | 0.0022 | 48.5361 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.4813 | -0.1039 | 0.003691 | 5.3191 | 4.6105 | 7.6413 | 3.1341
2012-10-30 20:57 | 28.651 | 19.798 | 19.0732 | 22.1645 | 17.5561 | -1.4899 | 0.006741 | 0.32679 | 0.002203 | 48.4778 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.491 | -0.1027 | 0.003693 | 5.307 | 4.5822 | 7.6735 | 3.0651
2012-10-30 20:57 | 28.6487 | 19.7309 | 19.0871 | 22.0776 | 17.57 | -1.4951 | 0.006725 | 0.32667 | 0.002197 | 48.5759 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5131 | -0.1031 | 0.003687 | 5.2178 | 4.574 | 7.5645 | 3.0569
2012-10-30 20:57 | 28.6562 | 19.7425 | 19.0581 | 22.1214 | 17.5531 | -1.4882 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.439 | -0.1026 | 0.003687 | 5.3035 | 4.6191 | 7.6824 | 3.1141
2012-10-30 20:57 | 28.6472 | 19.7456 | 19.0775 | 22.064 | 17.532 | -1.4864 | 0.006725 | 0.32669 | 0.002197 | 48.578 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4749 | -0.1025 | 0.003686 | 5.2707 | 4.6026 | 7.5891 | 3.0571
2012-10-30 20:57 | 28.6487 | 19.739 | 19.0304 | 22.0333 | 17.5213 | -1.5015 | 0.006729 | 0.32667 | 0.002198 | 48.547 | 0.00693 | 0.21458 | 0.001487 | 30.9645 | 14.4764 | -0.1035 | 0.003685 | 5.2626 | 4.554 | 7.5569 | 3.0449
2012-10-30 20:57 | 28.6435 | 19.7662 | 19.0859 | 22.0442 | 17.512 | -1.5009 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.006924 | 0.21464 | 0.001486 | 30.9997 | 14.459 | -0.1035 | 0.003684 | 5.3072 | 4.6269 | 7.5852 | 3.053
2012-10-30 20:57 | 28.639 | 19.7131 | 19.0328 | 22.0115 | 17.5359 | -1.502 | 0.006719 | 0.32676 | 0.002196 | 48.6322 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4504 | -0.1035 | 0.003684 | 5.2627 | 4.5824 | 7.5611 | 3.0855
2012-10-30 20:57 | 28.6368 | 19.6745 | 19.0103 | 22.0214 | 17.5297 | -1.4945 | 0.006735 | 0.32664 | 0.0022 | 48.4995 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.4604 | -0.103 | 0.003687 | 5.2141 | 4.5499 | 7.561 | 3.0693
2012-10-30 20:57 | 28.6465 | 19.7125 | 19.0403 | 22.0109 | 17.4988 | -1.5067 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4579 | -0.1039 | 0.003689 | 5.2546 | 4.5824 | 7.553 | 3.0409
2012-10-30 20:57 | 28.6495 | 19.7034 | 19.0312 | 22.0018 | 17.5018 | -1.487 | 0.006722 | 0.32663 | 0.002196 | 48.5912 | 0.006924 | 0.21467 | 0.001486 | 31.0039 | 14.4527 | -0.1025 | 0.003682 | 5.2507 | 4.5785 | 7.5491 | 3.0491
2012-10-30 20:57 | 28.6413 | 19.6628 | 19.0148 | 21.9856 | 17.5139 | -1.4974 | 0.006725 | 0.32672 | 0.002197 | 48.5823 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4405 | -0.1032 | 0.003687 | 5.2223 | 4.5743 | 7.5451 | 3.0734
2012-10-30 20:57 | 28.642 | 19.619 | 18.9872 | 21.9661 | 17.4943 | -1.4974 | 0.006729 | 0.32667 | 0.002198 | 48.547 | 0.006955 | 0.21464 | 0.001493 | 30.8615 | 14.4616 | -0.1032 | 0.003691 | 5.1574 | 4.5256 | 7.5045 | 3.0327
2012-10-30 20:57 | 28.6472 | 19.6525 | 18.9681 | 21.9391 | 17.5157 | -1.5055 | 0.006719 | 0.32667 | 0.002195 | 48.6193 | 0.00693 | 0.21463 | 0.001487 | 30.9709 | 14.4586 | -0.1038 | 0.003682 | 5.1939 | 4.5095 | 7.4805 | 3.0571
2012-10-30 20:57 | 28.6435 | 19.6002 | 18.9563 | 21.9031 | 17.4917 | -1.5038 | 0.006722 | 0.32675 | 0.002196 | 48.6083 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4386 | -0.1037 | 0.003685 | 5.1616 | 4.5177 | 7.4645 | 3.0531
2012-10-30 20:57 | 28.6428 | 19.6562 | 18.9353 | 21.9266 | 17.4504 | -1.5009 | 0.006738 | 0.32675 | 0.002202 | 48.4929 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.446 | -0.1035 | 0.00369 | 5.2102 | 4.4893 | 7.4806 | 3.0044
2012-10-30 20:57 | 28.6472 | 19.6039 | 18.9559 | 21.9149 | 17.4954 | -1.5095 | 0.006725 | 0.32673 | 0.002197 | 48.5845 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4341 | -0.1041 | 0.003688 | 5.1698 | 4.5218 | 7.4808 | 3.0613
2012-10-30 20:58 | 28.6375 | 19.6023 | 18.9543 | 21.9455 | 17.4857 | -1.5032 | 0.006722 | 0.32675 | 0.002196 | 48.6083 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.4244 | -0.1036 | 0.003687 | 5.1779 | 4.5299 | 7.5211 | 3.0613
2012-10-30 20:58 | 28.6502 | 19.611 | 18.9224 | 21.8937 | 17.4537 | -1.5015 | 0.006719 | 0.32675 | 0.002195 | 48.63 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4534 | -0.1035 | 0.003683 | 5.1576 | 4.469 | 7.4403 | 3.0003
2012-10-30 20:58 | 28.6502 | 19.5543 | 18.8941 | 21.8574 | 17.4822 | -1.4911 | 0.006732 | 0.32675 | 0.0022 | 48.5361 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4698 | -0.1028 | 0.003688 | 5.0845 | 4.4243 | 7.3876 | 3.0124
2012-10-30 20:58 | 28.651 | 19.5592 | 18.9516 | 21.8622 | 17.4464 | -1.5159 | 0.006735 | 0.32675 | 0.002201 | 48.5145 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.4869 | -0.1045 | 0.003688 | 5.0723 | 4.4647 | 7.3753 | 2.9595
2012-10-30 20:58 | 28.6383 | 19.5343 | 18.9065 | 21.8374 | 17.4337 | -1.509 | 0.006741 | 0.32673 | 0.002202 | 48.4692 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4252 | -0.104 | 0.003691 | 5.1091 | 4.4813 | 7.4122 | 3.0085
2012-10-30 20:58 | 28.645 | 19.5491 | 18.897 | 21.8038 | 17.4688 | -1.4986 | 0.006738 | 0.32673 | 0.002202 | 48.4908 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.436 | -0.1033 | 0.003693 | 5.1131 | 4.461 | 7.3678 | 3.0328
2012-10-30 20:58 | 28.6457 | 19.5336 | 18.9017 | 21.8126 | 17.4614 | -1.4928 | 0.006729 | 0.32673 | 0.002199 | 48.5557 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4979 | -0.1029 | 0.00369 | 5.0357 | 4.4038 | 7.3147 | 2.9635
2012-10-30 20:58 | 28.6435 | 19.5233 | 18.8955 | 21.8023 | 17.508 | -1.5009 | 0.006725 | 0.32669 | 0.002197 | 48.578 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4835 | -0.1035 | 0.003688 | 5.0398 | 4.412 | 7.3188 | 3.0245
2012-10-30 20:58 | 28.6435 | 19.4788 | 18.8793 | 21.7862 | 17.4348 | -1.4893 | 0.006747 | 0.32667 | 0.002204 | 48.4175 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.4345 | -0.1027 | 0.003697 | 5.0443 | 4.4448 | 7.3517 | 3.0003
2012-10-30 20:58 | 28.642 | 19.5056 | 18.8899 | 21.7766 | 17.4333 | -1.4922 | 0.006738 | 0.32675 | 0.002202 | 48.4929 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4779 | -0.1029 | 0.003693 | 5.0277 | 4.412 | 7.2987 | 2.9554
2012-10-30 20:58 | 28.6405 | 19.4677 | 18.8965 | 21.7832 | 17.4603 | -1.487 | 0.006735 | 0.32677 | 0.002201 | 48.5188 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.4764 | -0.1025 | 0.003693 | 4.9913 | 4.4201 | 7.3068 | 2.9839
2012-10-30 20:58 | 28.6428 | 19.4578 | 18.9029 | 21.7895 | 17.4098 | -1.498 | 0.006732 | 0.32676 | 0.0022 | 48.5383 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5072 | -0.1033 | 0.003692 | 4.9506 | 4.3957 | 7.2823 | 2.9026
2012-10-30 20:58 | 28.6338 | 19.4408 | 18.8534 | 21.7402 | 17.4373 | -1.5078 | 0.006741 | 0.32673 | 0.002202 | 48.4692 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5227 | -0.1039 | 0.003692 | 4.9181 | 4.3307 | 7.2175 | 2.9146
2012-10-30 20:58 | 28.6442 | 19.4876 | 18.8678 | 21.7344 | 17.4315 | -1.4974 | 0.006747 | 0.32672 | 0.002204 | 48.4239 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4964 | -0.1032 | 0.003694 | 4.9912 | 4.3714 | 7.238 | 2.9351
2012-10-30 20:58 | 28.6398 | 19.4387 | 18.831 | 21.6978 | 17.3986 | -1.4939 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.4879 | -0.103 | 0.003689 | 4.9508 | 4.3431 | 7.2099 | 2.9107
2012-10-30 20:58 | 28.639 | 19.446 | 18.8707 | 21.697 | 17.41 | -1.498 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.4626 | -0.1033 | 0.003687 | 4.9834 | 4.4081 | 7.2344 | 2.9474
2012-10-30 20:58 | 28.636 | 19.4146 | 18.7867 | 21.7343 | 17.4192 | -1.498 | 0.006735 | 0.32669 | 0.0022 | 48.5059 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4596 | -0.1033 | 0.003691 | 4.955 | 4.3271 | 7.2747 | 2.9596
2012-10-30 20:58 | 28.639 | 19.4055 | 18.8221 | 21.6849 | 17.4222 | -1.4899 | 0.006741 | 0.32669 | 0.002202 | 48.4627 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.483 | -0.1027 | 0.003692 | 4.9225 | 4.3391 | 7.2019 | 2.9392
2012-10-30 20:58 | 28.639 | 19.3933 | 18.8383 | 21.6768 | 17.4141 | -1.5055 | 0.006754 | 0.32677 | 0.002207 | 48.3823 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4708 | -0.1038 | 0.003698 | 4.9225 | 4.3675 | 7.206 | 2.9433
2012-10-30 20:58 | 28.6293 | 19.4201 | 18.8205 | 21.6429 | 17.3719 | -1.4945 | 0.006732 | 0.32667 | 0.002199 | 48.5254 | 0.006939 | 0.2146 | 0.001489 | 30.9265 | 14.4937 | -0.103 | 0.003688 | 4.9264 | 4.3268 | 7.1492 | 2.8782
2012-10-30 20:58 | 28.6323 | 19.3664 | 18.787 | 21.6338 | 17.4033 | -1.4991 | 0.006738 | 0.32672 | 0.002201 | 48.4886 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4478 | -0.1033 | 0.00369 | 4.9186 | 4.3392 | 7.186 | 2.9555
2012-10-30 20:58 | 28.6353 | 19.4301 | 18.7697 | 21.661 | 17.3901 | -1.4997 | 0.006725 | 0.32673 | 0.002197 | 48.5845 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4671 | -0.1034 | 0.003687 | 4.963 | 4.3026 | 7.1939 | 2.923
2012-10-30 20:58 | 28.624 | 19.35 | 18.7787 | 21.6094 | 17.3747 | -1.4934 | 0.006722 | 0.32661 | 0.002196 | 48.589 | 0.006946 | 0.21461 | 0.001491 | 30.8974 | 14.4476 | -0.103 | 0.003687 | 4.9024 | 4.3311 | 7.1618 | 2.9271
2012-10-30 20:58 | 28.6255 | 19.3596 | 18.7762 | 21.5665 | 17.3965 | -1.4911 | 0.006738 | 0.32673 | 0.002202 | 48.4908 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4328 | -0.1028 | 0.003689 | 4.9268 | 4.3434 | 7.1337 | 2.9637
2012-10-30 20:58 | 28.6315 | 19.3534 | 18.77 | 21.6088 | 17.37 | -1.5003 | 0.006732 | 0.32682 | 0.0022 | 48.5469 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4429 | -0.1034 | 0.003689 | 4.9105 | 4.3271 | 7.1659 | 2.9271
2012-10-30 20:58 | 28.6323 | 19.3299 | 18.7505 | 21.6096 | 17.4196 | -1.4928 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4355 | -0.1029 | 0.003689 | 4.8944 | 4.315 | 7.1741 | 2.9841
2012-10-30 20:59 | 28.6255 | 19.3231 | 18.7316 | 21.6109 | 17.3478 | -1.5015 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4695 | -0.1035 | 0.003689 | 4.8536 | 4.2621 | 7.1414 | 2.8783
2012-10-30 20:59 | 28.6255 | 19.3393 | 18.6991 | 21.5665 | 17.3153 | -1.5009 | 0.006738 | 0.32667 | 0.002201 | 48.4822 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.4532 | -0.1035 | 0.003692 | 4.8861 | 4.2459 | 7.1133 | 2.8621
2012-10-30 20:59 | 28.6285 | 19.3464 | 18.7346 | 21.5251 | 17.3264 | -1.4974 | 0.006735 | 0.32675 | 0.002201 | 48.5145 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4603 | -0.1032 | 0.00369 | 4.8861 | 4.2743 | 7.0648 | 2.8661
2012-10-30 20:59 | 28.6225 | 19.3404 | 18.688 | 21.5272 | 17.3448 | -1.4974 | 0.006725 | 0.32672 | 0.002197 | 48.5823 | 0.006914 | 0.21474 | 0.001485 | 31.0592 | 14.4298 | -0.1032 | 0.003682 | 4.9106 | 4.2582 | 7.0974 | 2.915
2012-10-30 20:59 | 28.6113 | 19.3049 | 18.6687 | 21.5281 | 17.3214 | -1.4991 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4064 | -0.1033 | 0.003687 | 4.8985 | 4.2623 | 7.1217 | 2.915
2012-10-30 20:59 | 28.618 | 19.3035 | 18.6754 | 21.5308 | 17.3403 | -1.4905 | 0.006725 | 0.32676 | 0.002197 | 48.5888 | 0.00693 | 0.21466 | 0.001488 | 30.9749 | 14.4376 | -0.1028 | 0.003685 | 4.8659 | 4.2378 | 7.0932 | 2.9027
2012-10-30 20:59 | 28.6105 | 19.2757 | 18.6963 | 21.491 | 17.3368 | -1.5078 | 0.006725 | 0.32667 | 0.002197 | 48.5759 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4464 | -0.1039 | 0.003685 | 4.8293 | 4.2499 | 7.0446 | 2.8904
2012-10-30 20:59 | 28.6188 | 19.284 | 18.6884 | 21.4952 | 17.3411 | -1.4916 | 0.006738 | 0.32675 | 0.002202 | 48.4929 | 0.006952 | 0.21464 | 0.001492 | 30.8749 | 14.4955 | -0.1028 | 0.003694 | 4.7885 | 4.1929 | 6.9997 | 2.8456
2012-10-30 20:59 | 28.6188 | 19.2354 | 18.6478 | 21.4468 | 17.3614 | -1.4974 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.006927 | 0.21466 | 0.001487 | 30.9883 | 14.4669 | -0.1032 | 0.003686 | 4.7685 | 4.1809 | 6.9799 | 2.8945
2012-10-30 20:59 | 28.6083 | 19.2695 | 18.6576 | 21.4726 | 17.3509 | -1.5026 | 0.00671 | 0.32672 | 0.002192 | 48.6909 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4605 | -0.1036 | 0.003683 | 4.809 | 4.1971 | 7.0121 | 2.8904
2012-10-30 20:59 | 28.6068 | 19.268 | 18.6358 | 21.4631 | 17.3291 | -1.4934 | 0.006741 | 0.3267 | 0.002202 | 48.4649 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4549 | -0.103 | 0.003694 | 4.8131 | 4.1809 | 7.0082 | 2.8742
2012-10-30 20:59 | 28.6023 | 19.1946 | 18.6435 | 21.4061 | 17.3246 | -1.5078 | 0.006735 | 0.32673 | 0.002201 | 48.5124 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4178 | -0.1039 | 0.00369 | 4.7768 | 4.2257 | 6.9883 | 2.9068
2012-10-30 20:59 | 28.6023 | 19.2311 | 18.6394 | 21.4424 | 17.3083 | -1.5078 | 0.006722 | 0.32677 | 0.002197 | 48.6126 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4627 | -0.1039 | 0.003688 | 4.7684 | 4.1767 | 6.9797 | 2.8456
2012-10-30 20:59 | 28.6016 | 19.1899 | 18.6509 | 21.4135 | 17.3036 | -1.4974 | 0.006716 | 0.32676 | 0.002195 | 48.654 | 0.006952 | 0.21471 | 0.001493 | 30.8852 | 14.4456 | -0.1032 | 0.003688 | 4.7443 | 4.2053 | 6.9679 | 2.858
2012-10-30 20:59 | 28.5978 | 19.1861 | 18.6309 | 21.4299 | 17.312 | -1.498 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.00693 | 0.21461 | 0.001487 | 30.9687 | 14.4786 | -0.1033 | 0.003688 | 4.7075 | 4.1523 | 6.9513 | 2.8334
2012-10-30 20:59 | 28.6046 | 19.2091 | 18.6255 | 21.4326 | 17.2741 | -1.4997 | 0.006735 | 0.32673 | 0.002201 | 48.5124 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.5058 | -0.1034 | 0.00369 | 4.7033 | 4.1197 | 6.9268 | 2.7683
2012-10-30 20:59 | 28.6046 | 19.1645 | 18.6053 | 21.364 | 17.3269 | -1.4974 | 0.006744 | 0.32675 | 0.002204 | 48.4497 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4935 | -0.1032 | 0.003693 | 4.671 | 4.1118 | 6.8705 | 2.8334
2012-10-30 20:59 | 28.5956 | 19.1353 | 18.653 | 21.4115 | 17.2732 | -1.509 | 0.006725 | 0.32669 | 0.002197 | 48.578 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5049 | -0.104 | 0.003686 | 4.6304 | 4.1481 | 6.9066 | 2.7683
2012-10-30 20:59 | 28.5881 | 19.1318 | 18.5928 | 21.3556 | 17.2901 | -1.4922 | 0.006738 | 0.32669 | 0.002201 | 48.4843 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4811 | -0.1029 | 0.003691 | 4.6507 | 4.1117 | 6.8745 | 2.809
2012-10-30 20:59 | 28.6038 | 19.1678 | 18.6247 | 21.3592 | 17.3139 | -1.5026 | 0.006722 | 0.32666 | 0.002196 | 48.5955 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4846 | -0.1036 | 0.003685 | 4.6832 | 4.1401 | 6.8746 | 2.8293
2012-10-30 20:59 | 28.5963 | 19.1319 | 18.6172 | 21.3194 | 17.2861 | -1.4951 | 0.006738 | 0.32676 | 0.002202 | 48.4951 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.4689 | -0.1031 | 0.003692 | 4.663 | 4.1483 | 6.8505 | 2.8172
2012-10-30 20:59 | 28.5941 | 19.1297 | 18.6434 | 21.3374 | 17.2839 | -1.5032 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4912 | -0.1036 | 0.003691 | 4.6385 | 4.1522 | 6.8462 | 2.7927
2012-10-30 20:59 | 28.5993 | 19.0782 | 18.5959 | 21.3103 | 17.2688 | -1.5015 | 0.006744 | 0.32667 | 0.002203 | 48.4391 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.5045 | -0.1035 | 0.003694 | 4.5737 | 4.0914 | 6.8058 | 2.7643
2012-10-30 20:59 | 28.5986 | 19.1139 | 18.6195 | 21.2733 | 17.2234 | -1.4934 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5202 | -0.103 | 0.003686 | 4.5937 | 4.0993 | 6.7531 | 2.7032
2012-10-30 20:59 | 28.5956 | 19.0988 | 18.5598 | 21.2985 | 17.2813 | -1.5032 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.456 | -0.1036 | 0.003693 | 4.6428 | 4.1038 | 6.8425 | 2.8253
2012-10-30 20:59 | 28.5881 | 19.0913 | 18.5888 | 21.291 | 17.2251 | -1.5055 | 0.006738 | 0.32672 | 0.002201 | 48.4886 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4281 | -0.1038 | 0.00369 | 4.6632 | 4.1607 | 6.8629 | 2.797
2012-10-30 21:00 | 28.5911 | 19.0821 | 18.5796 | 21.2577 | 17.2403 | -1.4974 | 0.006751 | 0.32669 | 0.002205 | 48.3909 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.48 | -0.1032 | 0.003693 | 4.6021 | 4.0996 | 6.7777 | 2.7603
2012-10-30 21:00 | 28.5881 | 19.0994 | 18.5685 | 21.2668 | 17.2373 | -1.4905 | 0.006725 | 0.32663 | 0.002197 | 48.5695 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4485 | -0.1028 | 0.003686 | 4.6509 | 4.12 | 6.8183 | 2.7888
2012-10-30 21:00 | 28.5888 | 19.0272 | 18.5611 | 21.2352 | 17.2867 | -1.4789 | 0.006729 | 0.32673 | 0.002199 | 48.5557 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4492 | -0.102 | 0.003687 | 4.578 | 4.1119 | 6.786 | 2.8375
2012-10-30 21:00 | 28.5768 | 19.0233 | 18.5775 | 21.2192 | 17.2097 | -1.4847 | 0.006735 | 0.3266 | 0.0022 | 48.4931 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4168 | -0.1024 | 0.003689 | 4.6065 | 4.1607 | 6.8024 | 2.7929
2012-10-30 21:00 | 28.5843 | 19.0308 | 18.5728 | 21.2227 | 17.2375 | -1.4893 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4447 | -0.1027 | 0.003692 | 4.5861 | 4.1281 | 6.778 | 2.7928
2012-10-30 21:00 | 28.5873 | 19.0297 | 18.5312 | 21.1974 | 17.2365 | -1.4858 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4517 | -0.1024 | 0.003686 | 4.578 | 4.0795 | 6.7457 | 2.7848
2012-10-30 21:00 | 28.5783 | 18.96 | 18.5263 | 21.1763 | 17.2234 | -1.4824 | 0.006725 | 0.32669 | 0.002197 | 48.578 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4183 | -0.1022 | 0.003689 | 4.5417 | 4.108 | 6.758 | 2.8051
2012-10-30 21:00 | 28.5836 | 18.9855 | 18.5153 | 21.1977 | 17.2206 | -1.476 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4317 | -0.1018 | 0.003686 | 4.5538 | 4.0836 | 6.766 | 2.7889
2012-10-30 21:00 | 28.5753 | 19.0056 | 18.503 | 21.1612 | 17.1879 | -1.4858 | 0.006741 | 0.32677 | 0.002203 | 48.4756 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4601 | -0.1024 | 0.003692 | 4.5455 | 4.0429 | 6.7011 | 2.7278
2012-10-30 21:00 | 28.5791 | 19.0013 | 18.5108 | 21.1771 | 17.1998 | -1.4911 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006946 | 0.21458 | 0.001491 | 30.8932 | 14.4476 | -0.1028 | 0.003689 | 4.5537 | 4.0632 | 6.7295 | 2.7522
2012-10-30 21:00 | 28.5738 | 19 | 18.5136 | 21.1476 | 17.1823 | -1.476 | 0.006738 | 0.32663 | 0.002201 | 48.4758 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4219 | -0.1018 | 0.003689 | 4.5781 | 4.0917 | 6.7257 | 2.7604
2012-10-30 21:00 | 28.5768 | 18.9828 | 18.4923 | 21.1344 | 17.2016 | -1.4778 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006952 | 0.21461 | 0.001492 | 30.8707 | 14.4494 | -0.1019 | 0.00369 | 4.5334 | 4.0429 | 6.685 | 2.7522
2012-10-30 21:00 | 28.5768 | 18.9301 | 18.4964 | 21.1304 | 17.226 | -1.4835 | 0.006719 | 0.32672 | 0.002195 | 48.6257 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4249 | -0.1023 | 0.003683 | 4.5052 | 4.0715 | 6.7055 | 2.8011
2012-10-30 21:00 | 28.5746 | 18.9846 | 18.4252 | 21.1443 | 17.2481 | -1.4778 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.4635 | -0.1019 | 0.003688 | 4.5211 | 3.9617 | 6.6808 | 2.7846
2012-10-30 21:00 | 28.5806 | 18.9623 | 18.5083 | 21.1423 | 17.2013 | -1.5015 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.4287 | -0.1035 | 0.00369 | 4.5336 | 4.0796 | 6.7136 | 2.7726
2012-10-30 21:00 | 28.5678 | 18.913 | 18.4225 | 21.1052 | 17.1723 | -1.4916 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4567 | -0.1028 | 0.003688 | 4.4563 | 3.9658 | 6.6485 | 2.7156
2012-10-30 21:00 | 28.5761 | 18.901 | 18.4754 | 21.0853 | 17.1968 | -1.476 | 0.006738 | 0.32669 | 0.002201 | 48.4843 | 0.006927 | 0.21463 | 0.001487 | 30.9843 | 14.4487 | -0.1018 | 0.003688 | 4.4523 | 4.0267 | 6.6366 | 2.7481
2012-10-30 21:00 | 28.5813 | 18.886 | 18.4603 | 21.0703 | 17.1695 | -1.4743 | 0.006735 | 0.3267 | 0.0022 | 48.5081 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4661 | -0.1016 | 0.003691 | 4.4199 | 3.9942 | 6.6042 | 2.7034
2012-10-30 21:00 | 28.5664 | 18.9116 | 18.4211 | 21.0796 | 17.1912 | -1.4916 | 0.006754 | 0.32664 | 0.002206 | 48.363 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 14.4431 | -0.1028 | 0.003693 | 4.4685 | 3.978 | 6.6365 | 2.7481
2012-10-30 21:00 | 28.5678 | 18.9049 | 18.4387 | 21.0406 | 17.1154 | -1.4922 | 0.006738 | 0.32669 | 0.002201 | 48.4843 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.473 | -0.1029 | 0.003689 | 4.4319 | 3.9657 | 6.5676 | 2.6424
2012-10-30 21:00 | 28.5701 | 18.9193 | 18.4288 | 21.0227 | 17.1868 | -1.4818 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.4345 | -0.1022 | 0.00369 | 4.4848 | 3.9943 | 6.5882 | 2.7523
2012-10-30 21:00 | 28.5776 | 18.8539 | 18.4282 | 21.0222 | 17.1252 | -1.4858 | 0.006729 | 0.32673 | 0.002199 | 48.5557 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4543 | -0.1024 | 0.003689 | 4.3996 | 3.9739 | 6.5679 | 2.6709
2012-10-30 21:00 | 28.5701 | 18.8586 | 18.4491 | 20.9824 | 17.1543 | -1.4824 | 0.006741 | 0.32666 | 0.002202 | 48.4585 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.4713 | -0.1022 | 0.003694 | 4.3873 | 3.9778 | 6.5111 | 2.683
2012-10-30 21:00 | 28.5641 | 18.8282 | 18.4269 | 21.0046 | 17.1401 | -1.487 | 0.006741 | 0.3266 | 0.002202 | 48.4499 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4693 | -0.1025 | 0.003689 | 4.3589 | 3.9576 | 6.5353 | 2.6708
2012-10-30 21:00 | 28.5678 | 18.8036 | 18.4549 | 20.9518 | 17.1438 | -1.4905 | 0.006729 | 0.32669 | 0.002198 | 48.5491 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.469 | -0.1028 | 0.003687 | 4.3346 | 3.9859 | 6.4828 | 2.6748
2012-10-30 21:00 | 28.5641 | 18.8161 | 18.3945 | 20.9804 | 17.1604 | -1.4772 | 0.006738 | 0.32663 | 0.002201 | 48.4758 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4571 | -0.1018 | 0.003688 | 4.359 | 3.9374 | 6.5233 | 2.7033
2012-10-30 21:00 | 28.5664 | 18.8224 | 18.3968 | 20.9544 | 17.1627 | -1.4835 | 0.006754 | 0.32661 | 0.002206 | 48.3588 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4594 | -0.1023 | 0.003696 | 4.363 | 3.9374 | 6.495 | 2.7033
2012-10-30 21:01 | 28.5701 | 18.8099 | 18.3599 | 20.9541 | 17.1055 | -1.4853 | 0.006725 | 0.32667 | 0.002197 | 48.5759 | 0.006914 | 0.21463 | 0.001484 | 31.0425 | 14.4345 | -0.1024 | 0.003681 | 4.3754 | 3.9254 | 6.5196 | 2.671
2012-10-30 21:01 | 28.5828 | 18.8388 | 18.4213 | 20.991 | 17.1019 | -1.4662 | 0.006722 | 0.3267 | 0.002196 | 48.6019 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4799 | -0.1011 | 0.003687 | 4.3589 | 3.9414 | 6.5111 | 2.622
2012-10-30 21:01 | 28.5738 | 18.7569 | 18.3879 | 20.9295 | 17.1173 | -1.4835 | 0.006741 | 0.32667 | 0.002202 | 48.4606 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.475 | -0.1023 | 0.003689 | 4.2819 | 3.9129 | 6.4545 | 2.6423
2012-10-30 21:01 | 28.5716 | 18.7749 | 18.406 | 20.9031 | 17.1192 | -1.4858 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4932 | -0.1024 | 0.003686 | 4.2817 | 3.9128 | 6.4099 | 2.626
2012-10-30 21:01 | 28.5649 | 18.7926 | 18.4034 | 20.9126 | 17.1044 | -1.4864 | 0.006729 | 0.32669 | 0.002198 | 48.5491 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.4212 | -0.1025 | 0.003685 | 4.3714 | 3.9822 | 6.4914 | 2.6832
2012-10-30 21:01 | 28.5686 | 18.776 | 18.3868 | 20.9203 | 17.0918 | -1.4778 | 0.006722 | 0.32663 | 0.002196 | 48.5912 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4494 | -0.1019 | 0.003686 | 4.3266 | 3.9374 | 6.4709 | 2.6424
2012-10-30 21:01 | 28.5731 | 18.7238 | 18.3629 | 20.9208 | 17.0841 | -1.4691 | 0.006744 | 0.32675 | 0.002204 | 48.4497 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4375 | -0.1013 | 0.003692 | 4.2863 | 3.9254 | 6.4833 | 2.6466
2012-10-30 21:01 | 28.5723 | 18.7148 | 18.3175 | 20.9038 | 17.0955 | -1.4795 | 0.006747 | 0.32667 | 0.002204 | 48.4175 | 0.00693 | 0.2146 | 0.001487 | 30.9667 | 14.4449 | -0.102 | 0.003691 | 4.2699 | 3.8726 | 6.4589 | 2.6506
2012-10-30 21:01 | 28.5761 | 18.7389 | 18.3618 | 20.8672 | 17.1156 | -1.4795 | 0.006719 | 0.32676 | 0.002196 | 48.6322 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.4569 | -0.102 | 0.003685 | 4.282 | 3.9049 | 6.4103 | 2.6587
2012-10-30 21:01 | 28.5716 | 18.7141 | 18.3371 | 20.8466 | 17.0785 | -1.4824 | 0.006729 | 0.3267 | 0.002198 | 48.5513 | 0.006958 | 0.21468 | 0.001494 | 30.8544 | 14.4687 | -0.1022 | 0.003692 | 4.2454 | 3.8684 | 6.3779 | 2.6098
2012-10-30 21:01 | 28.5693 | 18.6875 | 18.355 | 20.8806 | 17.0681 | -1.4818 | 0.006738 | 0.32667 | 0.002201 | 48.4822 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4705 | -0.1022 | 0.00369 | 4.217 | 3.8845 | 6.4101 | 2.5976
2012-10-30 21:01 | 28.5723 | 18.6662 | 18.3256 | 20.8553 | 17.0874 | -1.4801 | 0.006744 | 0.32672 | 0.002203 | 48.4454 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4653 | -0.102 | 0.003694 | 4.2009 | 3.8603 | 6.39 | 2.6221
2012-10-30 21:01 | 28.5641 | 18.658 | 18.3458 | 20.7865 | 17.0995 | -1.4737 | 0.006744 | 0.3267 | 0.002203 | 48.4434 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.4408 | -0.1016 | 0.003693 | 4.2172 | 3.905 | 6.3457 | 2.6587
2012-10-30 21:01 | 28.5708 | 18.6647 | 18.3525 | 20.8579 | 17.0656 | -1.4714 | 0.006747 | 0.32663 | 0.002204 | 48.4111 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4842 | -0.1014 | 0.003695 | 4.1805 | 3.8683 | 6.3737 | 2.5814
2012-10-30 21:01 | 28.5776 | 18.6553 | 18.3471 | 20.8849 | 17.1049 | -1.4916 | 0.006729 | 0.32669 | 0.002198 | 48.5491 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.4502 | -0.1028 | 0.003688 | 4.2051 | 3.8969 | 6.4347 | 2.6547
2012-10-30 21:01 | 28.5731 | 18.6467 | 18.3305 | 20.7794 | 17.0557 | -1.4743 | 0.006741 | 0.32661 | 0.002202 | 48.452 | 0.00693 | 0.21457 | 0.001487 | 30.9625 | 14.4783 | -0.1016 | 0.003689 | 4.1684 | 3.8522 | 6.3011 | 2.5774
2012-10-30 21:01 | 28.5589 | 18.6042 | 18.296 | 20.7692 | 17.074 | -1.4806 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4641 | -0.1021 | 0.003687 | 4.1401 | 3.8319 | 6.3051 | 2.6099
2012-10-30 21:01 | 28.5611 | 18.6104 | 18.3063 | 20.7634 | 17.0518 | -1.4853 | 0.006722 | 0.3267 | 0.002196 | 48.6019 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4623 | -0.1024 | 0.003687 | 4.1481 | 3.844 | 6.3011 | 2.5895
2012-10-30 21:01 | 28.5723 | 18.6378 | 18.3094 | 20.7665 | 17.0874 | -1.4882 | 0.006741 | 0.32664 | 0.002202 | 48.4563 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.4857 | -0.1026 | 0.003693 | 4.1521 | 3.8237 | 6.2808 | 2.6017
2012-10-30 21:01 | 28.5634 | 18.6087 | 18.268 | 20.7818 | 17.0825 | -1.4743 | 0.006719 | 0.32673 | 0.002195 | 48.6279 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.4686 | -0.1016 | 0.003686 | 4.1401 | 3.7994 | 6.3132 | 2.6139
2012-10-30 21:01 | 28.5611 | 18.6347 | 18.3144 | 20.723 | 17.0396 | -1.4818 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4786 | -0.1022 | 0.003684 | 4.1561 | 3.8358 | 6.2444 | 2.561
2012-10-30 21:01 | 28.5671 | 18.6002 | 18.2555 | 20.7532 | 17.0253 | -1.4858 | 0.006744 | 0.32673 | 0.002203 | 48.4477 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4683 | -0.1024 | 0.003694 | 4.1319 | 3.7872 | 6.2849 | 2.557
2012-10-30 21:01 | 28.5693 | 18.5578 | 18.3145 | 20.7958 | 17.0397 | -1.4708 | 0.006725 | 0.32666 | 0.002197 | 48.5738 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4623 | -0.1014 | 0.003686 | 4.0955 | 3.8522 | 6.3335 | 2.5774
2012-10-30 21:01 | 28.5708 | 18.5796 | 18.3038 | 20.7125 | 17.0615 | -1.4801 | 0.006725 | 0.32666 | 0.002197 | 48.5738 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.476 | -0.102 | 0.003687 | 4.1036 | 3.8278 | 6.2365 | 2.5855
2012-10-30 21:01 | 28.5701 | 18.5829 | 18.2788 | 20.732 | 17.0202 | -1.4876 | 0.006725 | 0.32673 | 0.002197 | 48.5845 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4917 | -0.1026 | 0.003686 | 4.0912 | 3.7871 | 6.2403 | 2.5285
2012-10-30 21:01 | 28.5686 | 18.5449 | 18.2489 | 20.6901 | 16.9739 | -1.4824 | 0.006741 | 0.32667 | 0.002202 | 48.4606 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4779 | -0.1022 | 0.003691 | 4.067 | 3.771 | 6.2122 | 2.496
2012-10-30 21:01 | 28.5716 | 18.5155 | 18.2803 | 20.6971 | 16.9973 | -1.4882 | 0.006744 | 0.32666 | 0.002203 | 48.437 | 0.006936 | 0.21464 | 0.001489 | 30.9461 | 14.4768 | -0.1026 | 0.003692 | 4.0387 | 3.8035 | 6.2203 | 2.5205
2012-10-30 21:02 | 28.5791 | 18.5068 | 18.2148 | 20.7006 | 17.0048 | -1.4876 | 0.006735 | 0.32669 | 0.0022 | 48.5059 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.4762 | -0.1026 | 0.003692 | 4.0306 | 3.7386 | 6.2244 | 2.5286
2012-10-30 21:02 | 28.5783 | 18.506 | 18.2383 | 20.6755 | 17.0162 | -1.4749 | 0.006732 | 0.32666 | 0.002199 | 48.5233 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4999 | -0.1017 | 0.003689 | 4.0061 | 3.7384 | 6.1756 | 2.5163
2012-10-30 21:02 | 28.5708 | 18.535 | 18.2268 | 20.664 | 16.9355 | -1.4812 | 0.006751 | 0.32669 | 0.002205 | 48.3909 | 0.006921 | 0.21467 | 0.001486 | 31.0173 | 14.4597 | -0.1021 | 0.003691 | 4.0753 | 3.7671 | 6.2043 | 2.4758
2012-10-30 21:02 | 28.5783 | 18.5303 | 18.2383 | 20.6634 | 16.9877 | -1.4812 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.4713 | -0.1021 | 0.00369 | 4.059 | 3.767 | 6.1921 | 2.5164
2012-10-30 21:02 | 28.5783 | 18.5263 | 18.2383 | 20.6594 | 17.0162 | -1.4679 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4876 | -0.1012 | 0.003687 | 4.0387 | 3.7507 | 6.1718 | 2.5286
2012-10-30 21:02 | 28.5738 | 18.4772 | 18.246 | 20.6104 | 17.0076 | -1.483 | 0.006741 | 0.32673 | 0.002202 | 48.4692 | 0.006939 | 0.21461 | 0.001489 | 30.9285 | 14.4831 | -0.1022 | 0.003691 | 3.9941 | 3.7629 | 6.1273 | 2.5245
2012-10-30 21:02 | 28.5761 | 18.4511 | 18.2199 | 20.645 | 16.9611 | -1.4882 | 0.006738 | 0.32667 | 0.002201 | 48.4822 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4732 | -0.1026 | 0.003689 | 3.9779 | 3.7467 | 6.1718 | 2.4879
2012-10-30 21:02 | 28.5761 | 18.4551 | 18.1712 | 20.6248 | 16.9489 | -1.4743 | 0.006735 | 0.32667 | 0.0022 | 48.5038 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4895 | -0.1016 | 0.003689 | 3.9656 | 3.6817 | 6.1353 | 2.4594
2012-10-30 21:02 | 28.5791 | 18.4743 | 18.1985 | 20.6278 | 16.9641 | -1.4818 | 0.006744 | 0.3267 | 0.002203 | 48.4434 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4721 | -0.1022 | 0.003693 | 4.0022 | 3.7264 | 6.1557 | 2.492
2012-10-30 21:02 | 28.5664 | 18.4333 | 18.194 | 20.5949 | 16.9636 | -1.4766 | 0.006725 | 0.32667 | 0.002197 | 48.5759 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4594 | -0.1018 | 0.003687 | 3.9739 | 3.7346 | 6.1355 | 2.5042
2012-10-30 21:02 | 28.5768 | 18.3788 | 18.1881 | 20.5932 | 17.0065 | -1.4789 | 0.006725 | 0.32675 | 0.002197 | 48.5866 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4494 | -0.102 | 0.003685 | 3.9294 | 3.7387 | 6.1438 | 2.5571
2012-10-30 21:02 | 28.5753 | 18.4138 | 18.2029 | 20.5432 | 16.9725 | -1.4858 | 0.006744 | 0.3267 | 0.002203 | 48.4434 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.4846 | -0.1024 | 0.003694 | 3.9292 | 3.7183 | 6.0586 | 2.4879
2012-10-30 21:02 | 28.5851 | 18.4236 | 18.1843 | 20.5692 | 16.9742 | -1.4795 | 0.006741 | 0.32673 | 0.002202 | 48.4692 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.5148 | -0.102 | 0.00369 | 3.9088 | 3.6695 | 6.0544 | 2.4594
2012-10-30 21:02 | 28.5806 | 18.3947 | 18.2203 | 20.5404 | 16.9616 | -1.4841 | 0.006744 | 0.32667 | 0.002203 | 48.4391 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.494 | -0.1023 | 0.003694 | 3.9007 | 3.7263 | 6.0464 | 2.4676
2012-10-30 21:02 | 28.5723 | 18.3986 | 18.1674 | 20.52 | 16.9329 | -1.4864 | 0.006725 | 0.32676 | 0.002197 | 48.5888 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.449 | -0.1025 | 0.003686 | 3.9496 | 3.7184 | 6.071 | 2.4839
2012-10-30 21:02 | 28.5768 | 18.3625 | 18.1678 | 20.5124 | 16.9293 | -1.4818 | 0.006751 | 0.32669 | 0.002205 | 48.3909 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4943 | -0.1022 | 0.003694 | 3.8682 | 3.6735 | 6.0181 | 2.435
2012-10-30 21:02 | 28.5768 | 18.3707 | 18.1638 | 20.5407 | 16.9578 | -1.4812 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4494 | -0.1021 | 0.003686 | 3.9213 | 3.7144 | 6.0913 | 2.5084
2012-10-30 21:02 | 28.5678 | 18.3495 | 18.1264 | 20.4993 | 16.9203 | -1.4893 | 0.006735 | 0.3267 | 0.0022 | 48.5081 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4445 | -0.1027 | 0.003691 | 3.905 | 3.6819 | 6.0548 | 2.4758
2012-10-30 21:02 | 28.5798 | 18.3209 | 18.1384 | 20.5194 | 16.9161 | -1.4847 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.432 | -0.1024 | 0.003692 | 3.8889 | 3.7064 | 6.0874 | 2.4841
2012-10-30 21:02 | 28.5828 | 18.332 | 18.1333 | 20.486 | 16.9028 | -1.4945 | 0.006741 | 0.32673 | 0.002202 | 48.4692 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.4554 | -0.103 | 0.00369 | 3.8766 | 3.6779 | 6.0306 | 2.4474
2012-10-30 21:02 | 28.5828 | 18.3158 | 18.117 | 20.4982 | 16.9556 | -1.4858 | 0.006741 | 0.32673 | 0.002202 | 48.4692 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4758 | -0.1024 | 0.003692 | 3.84 | 3.6412 | 6.0224 | 2.4798
2012-10-30 21:02 | 28.5798 | 18.398 | 18.1059 | 20.4669 | 16.9079 | -1.4934 | 0.006735 | 0.32664 | 0.0022 | 48.4995 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.5054 | -0.103 | 0.003688 | 3.8926 | 3.6005 | 5.9615 | 2.4025
2012-10-30 21:02 | 28.5828 | 18.3523 | 18.1414 | 20.4618 | 16.9272 | -1.4911 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.4799 | -0.1028 | 0.003691 | 3.8724 | 3.6615 | 5.9819 | 2.4473
2012-10-30 21:02 | 28.5858 | 18.331 | 18.1322 | 20.489 | 16.8692 | -1.4818 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4502 | -0.1022 | 0.003689 | 3.8808 | 3.682 | 6.0388 | 2.419
2012-10-30 21:02 | 28.5881 | 18.3049 | 18.0858 | 20.4752 | 16.9162 | -1.4806 | 0.006732 | 0.3266 | 0.002199 | 48.5147 | 0.006949 | 0.21461 | 0.001491 | 30.884 | 14.4648 | -0.1021 | 0.00369 | 3.8401 | 3.621 | 6.0104 | 2.4514
2012-10-30 21:02 | 28.5843 | 18.3295 | 18.1348 | 20.4027 | 16.8921 | -1.4922 | 0.006722 | 0.32669 | 0.002196 | 48.5997 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4487 | -0.1029 | 0.003684 | 3.8808 | 3.6861 | 5.954 | 2.4434
2012-10-30 21:02 | 28.5873 | 18.3203 | 18.1337 | 20.442 | 16.9154 | -1.4806 | 0.006744 | 0.32667 | 0.002203 | 48.4391 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.4681 | -0.1021 | 0.00369 | 3.8522 | 3.6656 | 5.9739 | 2.4473
2012-10-30 21:03 | 28.5836 | 18.2639 | 18.1422 | 20.3696 | 16.867 | -1.4783 | 0.006722 | 0.3266 | 0.002195 | 48.5869 | 0.006946 | 0.21463 | 0.001491 | 30.8995 | 14.5255 | -0.1019 | 0.003686 | 3.7384 | 3.6167 | 5.8441 | 2.3415
2012-10-30 21:03 | 28.5866 | 18.2912 | 18.1006 | 20.4413 | 16.9269 | -1.4754 | 0.006713 | 0.32664 | 0.002193 | 48.6584 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5204 | -0.1017 | 0.003683 | 3.7708 | 3.5802 | 5.9209 | 2.4065
2012-10-30 21:03 | 28.5798 | 18.2682 | 18.0978 | 20.3941 | 16.9201 | -1.4766 | 0.006751 | 0.32667 | 0.002205 | 48.3888 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.481 | -0.1018 | 0.003693 | 3.7872 | 3.6168 | 5.9131 | 2.4391
2012-10-30 21:03 | 28.5873 | 18.2676 | 18.1013 | 20.3652 | 16.891 | -1.4743 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.5415 | -0.1016 | 0.003693 | 3.7261 | 3.5598 | 5.8237 | 2.3495
2012-10-30 21:03 | 28.5783 | 18.2261 | 18.1247 | 20.3926 | 16.8576 | -1.4824 | 0.006751 | 0.32672 | 0.002206 | 48.3952 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.5447 | -0.1022 | 0.003694 | 3.6814 | 3.58 | 5.8479 | 2.3129
2012-10-30 21:03 | 28.5761 | 18.2686 | 18.0617 | 20.3743 | 16.8473 | -1.4743 | 0.006732 | 0.32672 | 0.002199 | 48.5318 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5303 | -0.1016 | 0.003689 | 3.7383 | 3.5314 | 5.844 | 2.317
2012-10-30 21:03 | 28.5813 | 18.221 | 18.0871 | 20.3795 | 16.8769 | -1.4858 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4947 | -0.1024 | 0.003689 | 3.7263 | 3.5924 | 5.8848 | 2.3822
2012-10-30 21:03 | 28.5791 | 18.2026 | 18.1215 | 20.353 | 16.8341 | -1.4812 | 0.006729 | 0.32672 | 0.002198 | 48.5534 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.5333 | -0.1021 | 0.003686 | 3.6693 | 3.5882 | 5.8197 | 2.3008
2012-10-30 21:03 | 28.5768 | 18.2125 | 18.1029 | 20.3467 | 16.909 | -1.4818 | 0.006741 | 0.32661 | 0.002202 | 48.452 | 0.006952 | 0.21464 | 0.001492 | 30.8749 | 14.5595 | -0.1022 | 0.003694 | 3.653 | 3.5434 | 5.7872 | 2.3495
2012-10-30 21:03 | 28.5738 | 18.2135 | 18.1162 | 20.3437 | 16.8816 | -1.4812 | 0.006744 | 0.3267 | 0.002203 | 48.4434 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.4954 | -0.1021 | 0.00369 | 3.7181 | 3.6208 | 5.8483 | 2.3862
2012-10-30 21:03 | 28.5761 | 18.2118 | 18.086 | 20.3298 | 16.8758 | -1.4743 | 0.006747 | 0.32669 | 0.002204 | 48.4196 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5548 | -0.1016 | 0.003694 | 3.657 | 3.5312 | 5.775 | 2.321
2012-10-30 21:03 | 28.5813 | 18.1845 | 18.0709 | 20.3592 | 16.8403 | -1.4766 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.5273 | -0.1018 | 0.003693 | 3.6572 | 3.5436 | 5.8319 | 2.313
2012-10-30 21:03 | 28.5731 | 18.2007 | 18.0384 | 20.3268 | 16.8362 | -1.4772 | 0.006729 | 0.32675 | 0.002199 | 48.5577 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.5314 | -0.1018 | 0.003687 | 3.6693 | 3.507 | 5.7954 | 2.3048
2012-10-30 21:03 | 28.5836 | 18.1503 | 18.0489 | 20.3049 | 16.8751 | -1.4801 | 0.006725 | 0.32669 | 0.002197 | 48.578 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5133 | -0.102 | 0.003688 | 3.637 | 3.5356 | 5.7916 | 2.3618
2012-10-30 21:03 | 28.5768 | 18.1597 | 18.0177 | 20.2779 | 16.8358 | -1.4789 | 0.006729 | 0.32673 | 0.002199 | 48.5557 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4902 | -0.102 | 0.003691 | 3.6695 | 3.5275 | 5.7877 | 2.3456
2012-10-30 21:03 | 28.5783 | 18.1734 | 18.0639 | 20.2996 | 16.882 | -1.4876 | 0.006735 | 0.32664 | 0.0022 | 48.4995 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.5366 | -0.1026 | 0.003688 | 3.6368 | 3.5273 | 5.763 | 2.3454
2012-10-30 21:03 | 28.5806 | 18.1757 | 18.0378 | 20.31 | 16.8315 | -1.483 | 0.006741 | 0.32664 | 0.002202 | 48.4563 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.5022 | -0.1022 | 0.003691 | 3.6735 | 3.5356 | 5.8078 | 2.3293
2012-10-30 21:03 | 28.5806 | 18.1473 | 18.0215 | 20.2777 | 16.8274 | -1.4835 | 0.006732 | 0.32669 | 0.002199 | 48.5275 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4981 | -0.1023 | 0.003687 | 3.6492 | 3.5234 | 5.7796 | 2.3293
2012-10-30 21:03 | 28.5678 | 18.1264 | 18.025 | 20.2325 | 16.839 | -1.4835 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5057 | -0.1023 | 0.003693 | 3.6207 | 3.5193 | 5.7268 | 2.3333
2012-10-30 21:03 | 28.5731 | 18.1236 | 18.0384 | 20.25 | 16.8159 | -1.4847 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5314 | -0.1024 | 0.003688 | 3.5922 | 3.507 | 5.7186 | 2.2845
2012-10-30 21:03 | 28.5761 | 18.1185 | 18.0089 | 20.253 | 16.827 | -1.4928 | 0.006741 | 0.3267 | 0.002202 | 48.4649 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4609 | -0.1029 | 0.003693 | 3.6576 | 3.548 | 5.7921 | 2.3661
2012-10-30 21:03 | 28.5753 | 18.0852 | 17.9675 | 20.232 | 16.7774 | -1.4778 | 0.006722 | 0.32667 | 0.002196 | 48.5976 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4928 | -0.1019 | 0.003685 | 3.5924 | 3.4747 | 5.7392 | 2.2846
2012-10-30 21:03 | 28.5708 | 18.0604 | 17.9671 | 20.2032 | 16.8054 | -1.4772 | 0.006735 | 0.32673 | 0.002201 | 48.5124 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4801 | -0.1018 | 0.00369 | 3.5803 | 3.487 | 5.7231 | 2.3253
2012-10-30 21:03 | 28.5753 | 18.0609 | 17.9838 | 20.1956 | 16.7855 | -1.5009 | 0.006732 | 0.3267 | 0.002199 | 48.5297 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4887 | -0.1035 | 0.003689 | 3.5722 | 3.4951 | 5.7069 | 2.2968
2012-10-30 21:03 | 28.5708 | 18.0401 | 17.9793 | 20.1668 | 16.7892 | -1.5003 | 0.006732 | 0.32673 | 0.0022 | 48.534 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4842 | -0.1034 | 0.003689 | 3.5559 | 3.4951 | 5.6826 | 2.305
2012-10-30 21:04 | 28.5821 | 18.0514 | 17.9987 | 20.2064 | 16.8005 | -1.4899 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5077 | -0.1027 | 0.00369 | 3.5437 | 3.491 | 5.6987 | 2.2928
2012-10-30 21:04 | 28.5746 | 18.0439 | 17.9871 | 20.1827 | 16.8377 | -1.4847 | 0.006738 | 0.32669 | 0.002201 | 48.4843 | 0.006955 | 0.2147 | 0.001493 | 30.8699 | 14.488 | -0.1024 | 0.003694 | 3.5559 | 3.4991 | 5.6947 | 2.3497
2012-10-30 21:04 | 28.5836 | 18.0164 | 17.9637 | 20.1796 | 16.7979 | -1.4934 | 0.006732 | 0.3266 | 0.002199 | 48.5147 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.5378 | -0.103 | 0.00369 | 3.4786 | 3.4259 | 5.6418 | 2.2601
2012-10-30 21:04 | 28.5813 | 18.0019 | 18.0141 | 20.1571 | 16.7915 | -1.4899 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.511 | -0.1027 | 0.003688 | 3.4909 | 3.5031 | 5.6461 | 2.2805
2012-10-30 21:04 | 28.5806 | 18.0459 | 17.9809 | 20.1766 | 16.7624 | -1.4864 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5593 | -0.1025 | 0.00369 | 3.4866 | 3.4216 | 5.6173 | 2.2031
2012-10-30 21:04 | 28.5791 | 18.0647 | 17.9997 | 20.167 | 16.8137 | -1.4824 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.006936 | 0.21461 | 0.001489 | 30.9419 | 14.5618 | -0.1022 | 0.003687 | 3.5029 | 3.4379 | 5.6052 | 2.2519
2012-10-30 21:04 | 28.5761 | 17.9967 | 17.9846 | 20.1357 | 16.7701 | -1.4934 | 0.006735 | 0.32663 | 0.0022 | 48.4974 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5221 | -0.103 | 0.003689 | 3.4746 | 3.4625 | 5.6136 | 2.248
2012-10-30 21:04 | 28.5746 | 18.0358 | 18.0236 | 20.1099 | 16.793 | -1.4806 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.5206 | -0.1021 | 0.003685 | 3.5152 | 3.503 | 5.5893 | 2.2724
2012-10-30 21:04 | 28.5851 | 17.9773 | 18.0017 | 20.1285 | 16.8157 | -1.4853 | 0.006729 | 0.32661 | 0.002198 | 48.5384 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5107 | -0.1024 | 0.003688 | 3.4666 | 3.491 | 5.6178 | 2.305
2012-10-30 21:04 | 28.5828 | 17.9994 | 17.9547 | 20.0939 | 16.789 | -1.4824 | 0.006744 | 0.32664 | 0.002203 | 48.4348 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5533 | -0.1022 | 0.003692 | 3.4461 | 3.4014 | 5.5406 | 2.2357
2012-10-30 21:04 | 28.5836 | 17.988 | 17.9758 | 20.0704 | 16.7572 | -1.4754 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5459 | -0.1017 | 0.003689 | 3.4421 | 3.4299 | 5.5245 | 2.2113
2012-10-30 21:04 | 28.5866 | 18.0072 | 17.9382 | 20.0734 | 16.7562 | -1.4945 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.5612 | -0.103 | 0.003688 | 3.446 | 3.377 | 5.5122 | 2.195
2012-10-30 21:04 | 28.5806 | 17.9607 | 17.9282 | 20.0755 | 16.7298 | -1.4818 | 0.006741 | 0.32672 | 0.002202 | 48.467 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.547 | -0.1022 | 0.00369 | 3.4137 | 3.3812 | 5.5285 | 2.1828
2012-10-30 21:04 | 28.5903 | 17.9541 | 17.9257 | 20.1014 | 16.7639 | -1.483 | 0.006716 | 0.32663 | 0.002194 | 48.6346 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.5241 | -0.1022 | 0.003686 | 3.43 | 3.4016 | 5.5773 | 2.2398
2012-10-30 21:04 | 28.5926 | 17.9767 | 17.9158 | 20.0754 | 16.7093 | -1.487 | 0.006719 | 0.32669 | 0.002195 | 48.6214 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5427 | -0.1025 | 0.003684 | 3.434 | 3.3731 | 5.5327 | 2.1666
2012-10-30 21:04 | 28.5881 | 17.96 | 17.9154 | 20.0506 | 16.7252 | -1.4934 | 0.006729 | 0.32667 | 0.002198 | 48.547 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.53 | -0.103 | 0.003686 | 3.43 | 3.3854 | 5.5206 | 2.1952
2012-10-30 21:04 | 28.5851 | 17.9449 | 17.8921 | 20.0517 | 16.7628 | -1.4893 | 0.006719 | 0.32666 | 0.002195 | 48.6172 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.5474 | -0.1027 | 0.003687 | 3.3975 | 3.3447 | 5.5043 | 2.2154
2012-10-30 21:04 | 28.5798 | 17.9314 | 17.9193 | 19.9898 | 16.7087 | -1.4806 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5054 | -0.1021 | 0.003688 | 3.426 | 3.4139 | 5.4844 | 2.2033
2012-10-30 21:04 | 28.5888 | 17.9283 | 17.8917 | 20.0797 | 16.7665 | -1.4893 | 0.006725 | 0.3266 | 0.002196 | 48.5652 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.5511 | -0.1027 | 0.003684 | 3.3772 | 3.3406 | 5.5286 | 2.2154
2012-10-30 21:04 | 28.5813 | 17.9208 | 17.8518 | 20.0074 | 16.7184 | -1.4841 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5355 | -0.1023 | 0.00369 | 3.3853 | 3.3163 | 5.4719 | 2.1829
2012-10-30 21:04 | 28.5806 | 17.9201 | 17.9079 | 19.9865 | 16.7217 | -1.4801 | 0.006725 | 0.32666 | 0.002197 | 48.5738 | 0.00693 | 0.21466 | 0.001488 | 30.9749 | 14.5022 | -0.102 | 0.003685 | 3.4179 | 3.4057 | 5.4843 | 2.2195
2012-10-30 21:04 | 28.5933 | 17.9044 | 17.8962 | 19.9871 | 16.6897 | -1.4864 | 0.00671 | 0.32664 | 0.002192 | 48.6802 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.5312 | -0.1025 | 0.00368 | 3.3732 | 3.365 | 5.4559 | 2.1585
2012-10-30 21:04 | 28.5911 | 17.8981 | 17.8453 | 20.0011 | 16.72 | -1.4893 | 0.006735 | 0.3266 | 0.0022 | 48.4931 | 0.006924 | 0.2147 | 0.001487 | 31.0081 | 14.529 | -0.1027 | 0.003687 | 3.3691 | 3.3163 | 5.4721 | 2.191
2012-10-30 21:04 | 28.5941 | 17.8849 | 17.8524 | 19.9919 | 16.7027 | -1.4899 | 0.006738 | 0.32663 | 0.002201 | 48.4758 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.532 | -0.1027 | 0.00369 | 3.3529 | 3.3204 | 5.4599 | 2.1707
2012-10-30 21:04 | 28.5963 | 17.9033 | 17.8302 | 19.9577 | 16.6927 | -1.4934 | 0.006741 | 0.32669 | 0.002202 | 48.4627 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4934 | -0.103 | 0.003691 | 3.4099 | 3.3368 | 5.4643 | 2.1993
2012-10-30 21:05 | 28.5888 | 17.8512 | 17.8268 | 19.9704 | 16.6893 | -1.4957 | 0.006747 | 0.3267 | 0.002204 | 48.4218 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4859 | -0.1031 | 0.003693 | 3.3653 | 3.3409 | 5.4845 | 2.2034
2012-10-30 21:05 | 28.6001 | 17.8259 | 17.8381 | 19.9858 | 16.6558 | -1.4783 | 0.006741 | 0.32666 | 0.002202 | 48.4585 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5257 | -0.1019 | 0.003693 | 3.3002 | 3.3124 | 5.4601 | 2.1301
2012-10-30 21:05 | 28.5926 | 17.8631 | 17.8022 | 19.9055 | 16.6239 | -1.4922 | 0.006729 | 0.32667 | 0.002198 | 48.547 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5101 | -0.1029 | 0.003688 | 3.353 | 3.2921 | 5.3954 | 2.1138
2012-10-30 21:05 | 28.5971 | 17.8595 | 17.8229 | 19.8817 | 16.6813 | -1.476 | 0.006747 | 0.32669 | 0.002204 | 48.4196 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4942 | -0.1018 | 0.003694 | 3.3653 | 3.3287 | 5.3875 | 2.1871
2012-10-30 21:05 | 28.5948 | 17.7922 | 17.8287 | 19.9117 | 16.6261 | -1.4939 | 0.006751 | 0.32661 | 0.002205 | 48.3802 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.5123 | -0.103 | 0.003695 | 3.2799 | 3.3164 | 5.3994 | 2.1138
2012-10-30 21:05 | 28.6105 | 17.7876 | 17.816 | 19.8748 | 16.6418 | -1.4847 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5035 | -0.1024 | 0.003689 | 3.2841 | 3.3125 | 5.3713 | 2.1383
2012-10-30 21:05 | 28.5971 | 17.7904 | 17.8229 | 19.8817 | 16.6122 | -1.4916 | 0.006722 | 0.32673 | 0.002196 | 48.6062 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5309 | -0.1028 | 0.003686 | 3.2595 | 3.292 | 5.3508 | 2.0813
2012-10-30 21:05 | 28.5986 | 17.7838 | 17.7879 | 19.8751 | 16.634 | -1.4841 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4957 | -0.1023 | 0.003692 | 3.2881 | 3.2922 | 5.3794 | 2.1383
2012-10-30 21:05 | 28.609 | 17.7983 | 17.8267 | 19.8571 | 16.6403 | -1.4841 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.006946 | 0.21464 | 0.001491 | 30.9015 | 14.5061 | -0.1023 | 0.003691 | 3.2922 | 3.3206 | 5.351 | 2.1342
2012-10-30 21:05 | 28.5918 | 17.777 | 17.7973 | 19.8521 | 16.6109 | -1.4876 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.4726 | -0.1026 | 0.003691 | 3.3044 | 3.3247 | 5.3795 | 2.1383
2012-10-30 21:05 | 28.6023 | 17.7997 | 17.7672 | 19.8464 | 16.6336 | -1.4864 | 0.006725 | 0.32663 | 0.002197 | 48.5695 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4831 | -0.1025 | 0.003686 | 3.3166 | 3.2841 | 5.3633 | 2.1505
2012-10-30 21:05 | 28.6031 | 17.8411 | 17.7761 | 19.8391 | 16.6466 | -1.476 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006939 | 0.21455 | 0.001489 | 30.9202 | 14.5002 | -0.1018 | 0.00369 | 3.3409 | 3.2759 | 5.3389 | 2.1464
2012-10-30 21:05 | 28.5993 | 17.752 | 17.7845 | 19.8272 | 16.6266 | -1.4824 | 0.006732 | 0.3266 | 0.002199 | 48.5147 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.4964 | -0.1022 | 0.003688 | 3.2556 | 3.2881 | 5.3308 | 2.1302
2012-10-30 21:05 | 28.6016 | 17.7422 | 17.7665 | 19.8619 | 16.6207 | -1.4916 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006917 | 0.21464 | 0.001485 | 31.0311 | 14.4946 | -0.1028 | 0.003683 | 3.2476 | 3.2719 | 5.3673 | 2.1261
2012-10-30 21:05 | 28.5993 | 17.7439 | 17.7358 | 19.8313 | 16.651 | -1.4887 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.5127 | -0.1026 | 0.003688 | 3.2312 | 3.2231 | 5.3186 | 2.1383
2012-10-30 21:05 | 28.6031 | 17.7558 | 17.7437 | 19.8431 | 16.6019 | -1.4876 | 0.006738 | 0.32672 | 0.002201 | 48.4886 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4839 | -0.1026 | 0.003691 | 3.2719 | 3.2598 | 5.3592 | 2.118
2012-10-30 21:05 | 28.5978 | 17.7221 | 17.7587 | 19.8217 | 16.6495 | -1.4853 | 0.006719 | 0.32666 | 0.002195 | 48.6172 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4704 | -0.1024 | 0.003684 | 3.2517 | 3.2883 | 5.3513 | 2.1791
2012-10-30 21:05 | 28.6038 | 17.7038 | 17.7687 | 19.7791 | 16.5985 | -1.4853 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5335 | -0.1024 | 0.003693 | 3.1703 | 3.2352 | 5.2456 | 2.065
2012-10-30 21:05 | 28.6083 | 17.7489 | 17.7651 | 19.7836 | 16.5664 | -1.4772 | 0.006725 | 0.3266 | 0.002196 | 48.5652 | 0.006927 | 0.21466 | 0.001487 | 30.9883 | 14.5054 | -0.1018 | 0.003683 | 3.2435 | 3.2597 | 5.2782 | 2.061
2012-10-30 21:05 | 28.6053 | 17.689 | 17.7175 | 19.7927 | 16.6204 | -1.4922 | 0.006732 | 0.32666 | 0.002199 | 48.5233 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5269 | -0.1029 | 0.003689 | 3.1621 | 3.1906 | 5.2658 | 2.0935
2012-10-30 21:05 | 28.606 | 17.641 | 17.7141 | 19.7813 | 16.556 | -1.4818 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5194 | -0.1022 | 0.003685 | 3.1216 | 3.1947 | 5.2619 | 2.0366
2012-10-30 21:05 | 28.6038 | 17.6875 | 17.7565 | 19.7184 | 16.6026 | -1.4864 | 0.006735 | 0.32667 | 0.0022 | 48.5038 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5661 | -0.1025 | 0.003691 | 3.1214 | 3.1904 | 5.1523 | 2.0365
2012-10-30 21:05 | 28.6008 | 17.7008 | 17.6886 | 19.7356 | 16.5833 | -1.4991 | 0.006735 | 0.32663 | 0.0022 | 48.4974 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5224 | -0.1033 | 0.003689 | 3.1784 | 3.1662 | 5.2132 | 2.0609
2012-10-30 21:05 | 28.6038 | 17.6632 | 17.7363 | 19.7184 | 16.6189 | -1.4991 | 0.006741 | 0.3266 | 0.002202 | 48.4499 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5213 | -0.1033 | 0.003691 | 3.1419 | 3.215 | 5.1971 | 2.0976
2012-10-30 21:06 | 28.6158 | 17.6589 | 17.732 | 19.7264 | 16.5902 | -1.4905 | 0.006725 | 0.32669 | 0.002197 | 48.578 | 0.006955 | 0.21471 | 0.001493 | 30.8719 | 14.5414 | -0.1028 | 0.00369 | 3.1175 | 3.1906 | 5.185 | 2.0488
2012-10-30 21:06 | 28.615 | 17.6866 | 17.715 | 19.7134 | 16.565 | -1.4916 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5569 | -0.1028 | 0.003692 | 3.1297 | 3.1581 | 5.1565 | 2.0081
2012-10-30 21:06 | 28.6135 | 17.5998 | 17.6688 | 19.6876 | 16.5716 | -1.4934 | 0.006725 | 0.32676 | 0.002197 | 48.5888 | 0.006921 | 0.21473 | 0.001486 | 31.0257 | 14.5432 | -0.103 | 0.003683 | 3.0566 | 3.1256 | 5.1444 | 2.0284
2012-10-30 21:06 | 28.609 | 17.6196 | 17.6603 | 19.6548 | 16.5509 | -1.4905 | 0.006735 | 0.3267 | 0.0022 | 48.5081 | 0.006933 | 0.21463 | 0.001488 | 30.9575 | 14.4694 | -0.1028 | 0.003688 | 3.1502 | 3.1909 | 5.1854 | 2.0815
2012-10-30 21:06 | 28.6075 | 17.6141 | 17.6141 | 19.6331 | 16.5534 | -1.4905 | 0.006722 | 0.32664 | 0.002196 | 48.5933 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5005 | -0.1028 | 0.003685 | 3.1136 | 3.1136 | 5.1326 | 2.0529
2012-10-30 21:06 | 28.618 | 17.584 | 17.6449 | 19.6395 | 16.507 | -1.498 | 0.006732 | 0.3267 | 0.002199 | 48.5297 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5436 | -0.1033 | 0.003688 | 3.0404 | 3.1013 | 5.0959 | 1.9634
2012-10-30 21:06 | 28.6098 | 17.5839 | 17.6489 | 19.6556 | 16.5313 | -1.4939 | 0.006732 | 0.32667 | 0.002199 | 48.5254 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.511 | -0.103 | 0.003688 | 3.0729 | 3.1379 | 5.1446 | 2.0203
2012-10-30 21:06 | 28.6128 | 17.5788 | 17.6641 | 19.6424 | 16.5303 | -1.4882 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5221 | -0.1026 | 0.00369 | 3.0567 | 3.142 | 5.1203 | 2.0082
2012-10-30 21:06 | 28.6143 | 17.56 | 17.6331 | 19.6763 | 16.5602 | -1.4916 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.5114 | -0.1028 | 0.003687 | 3.0486 | 3.1217 | 5.1649 | 2.0488
2012-10-30 21:06 | 28.6158 | 17.6021 | 17.6305 | 19.6373 | 16.5821 | -1.4916 | 0.006722 | 0.32661 | 0.002196 | 48.589 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.4843 | -0.1028 | 0.003684 | 3.1178 | 3.1462 | 5.153 | 2.0978
2012-10-30 21:06 | 28.6098 | 17.592 | 17.657 | 19.6394 | 16.5232 | -1.5009 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4946 | -0.1035 | 0.003688 | 3.0974 | 3.1624 | 5.1448 | 2.0286
2012-10-30 21:06 | 28.612 | 17.5617 | 17.6511 | 19.6173 | 16.5335 | -1.4939 | 0.006719 | 0.3266 | 0.002194 | 48.6086 | 0.006924 | 0.21468 | 0.001486 | 31.0059 | 14.5254 | -0.103 | 0.00368 | 3.0363 | 3.1257 | 5.0919 | 2.0081
2012-10-30 21:06 | 28.6143 | 17.5275 | 17.6128 | 19.6034 | 16.5196 | -1.4899 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006933 | 0.21463 | 0.001488 | 30.9575 | 14.5195 | -0.1027 | 0.003687 | 3.008 | 3.0933 | 5.0839 | 2.0001
2012-10-30 21:06 | 28.6165 | 17.5622 | 17.6393 | 19.6502 | 16.5462 | -1.4749 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.5054 | -0.1017 | 0.003685 | 3.0568 | 3.1339 | 5.1448 | 2.0408
2012-10-30 21:06 | 28.6113 | 17.5407 | 17.6219 | 19.5923 | 16.5084 | -1.5055 | 0.006722 | 0.32659 | 0.002195 | 48.5846 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.5206 | -0.1038 | 0.003683 | 3.0201 | 3.1013 | 5.0717 | 1.9878
2012-10-30 21:06 | 28.6203 | 17.5051 | 17.6431 | 19.5811 | 16.5256 | -1.4754 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006943 | 0.2146 | 0.00149 | 30.9087 | 14.5174 | -0.1017 | 0.003691 | 2.9877 | 3.1257 | 5.0637 | 2.0082
2012-10-30 21:06 | 28.6098 | 17.5067 | 17.6123 | 19.5746 | 16.511 | -1.4876 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.511 | -0.1026 | 0.003693 | 2.9957 | 3.1013 | 5.0636 | 2
2012-10-30 21:06 | 28.615 | 17.5607 | 17.6053 | 19.5556 | 16.504 | -1.4824 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4794 | -0.1022 | 0.003688 | 3.0813 | 3.1259 | 5.0762 | 2.0246
2012-10-30 21:06 | 28.6188 | 17.5036 | 17.5767 | 19.5715 | 16.5241 | -1.4887 | 0.006738 | 0.32661 | 0.002201 | 48.4736 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.471 | -0.1026 | 0.003692 | 3.0326 | 3.1057 | 5.1005 | 2.0531
2012-10-30 21:06 | 28.6135 | 17.5064 | 17.5673 | 19.5622 | 16.4781 | -1.4922 | 0.006725 | 0.32667 | 0.002197 | 48.5759 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5269 | -0.1029 | 0.003686 | 2.9795 | 3.0404 | 5.0353 | 1.9512
2012-10-30 21:06 | 28.612 | 17.5049 | 17.5739 | 19.5323 | 16.4847 | -1.4882 | 0.006713 | 0.32656 | 0.002192 | 48.6455 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.505 | -0.1026 | 0.00368 | 2.9999 | 3.0689 | 5.0273 | 1.9797
2012-10-30 21:07 | 28.609 | 17.4897 | 17.5425 | 19.4929 | 16.4695 | -1.4939 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.449 | -0.103 | 0.003689 | 3.0407 | 3.0935 | 5.0439 | 2.0205
2012-10-30 21:07 | 28.6188 | 17.4467 | 17.532 | 19.4946 | 16.4549 | -1.4934 | 0.006707 | 0.32654 | 0.00219 | 48.6869 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4465 | -0.103 | 0.003678 | 3.0002 | 3.0855 | 5.0481 | 2.0084
2012-10-30 21:07 | 28.612 | 17.5211 | 17.5008 | 19.5161 | 16.4481 | -1.4939 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4438 | -0.103 | 0.00369 | 3.0773 | 3.057 | 5.0723 | 2.0043
2012-10-30 21:07 | 28.6195 | 17.4515 | 17.5733 | 19.5155 | 16.4434 | -1.4882 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.4758 | -0.1026 | 0.003684 | 2.9757 | 3.0975 | 5.0397 | 1.9676
2012-10-30 21:07 | 28.624 | 17.456 | 17.5331 | 19.4188 | 16.4642 | -1.4939 | 0.006744 | 0.32661 | 0.002203 | 48.4305 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.5374 | -0.103 | 0.003691 | 2.9186 | 2.9957 | 4.8814 | 1.9268
2012-10-30 21:07 | 28.6203 | 17.4441 | 17.4848 | 19.5163 | 16.4849 | -1.4876 | 0.006741 | 0.32664 | 0.002202 | 48.4563 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4725 | -0.1026 | 0.003691 | 2.9716 | 3.0123 | 5.0438 | 2.0124
2012-10-30 21:07 | 28.6188 | 17.4751 | 17.4914 | 19.4338 | 16.463 | -1.4887 | 0.006732 | 0.32663 | 0.002199 | 48.519 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4669 | -0.1026 | 0.003688 | 3.0082 | 3.0245 | 4.9669 | 1.9961
2012-10-30 21:07 | 28.6203 | 17.4401 | 17.5172 | 19.4556 | 16.4605 | -1.5009 | 0.006719 | 0.32654 | 0.002194 | 48.5999 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5092 | -0.1035 | 0.003683 | 2.9309 | 3.008 | 4.9464 | 1.9513
2012-10-30 21:07 | 28.615 | 17.4348 | 17.4713 | 19.4503 | 16.4552 | -1.5009 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5039 | -0.1035 | 0.003692 | 2.9309 | 2.9674 | 4.9464 | 1.9513
2012-10-30 21:07 | 28.615 | 17.4348 | 17.4835 | 19.4098 | 16.4186 | -1.4928 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5202 | -0.1029 | 0.003691 | 2.9146 | 2.9633 | 4.8896 | 1.8984
2012-10-30 21:07 | 28.6128 | 17.4285 | 17.5057 | 19.4157 | 16.4082 | -1.483 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006946 | 0.21461 | 0.001491 | 30.8974 | 14.4772 | -0.1022 | 0.003689 | 2.9513 | 3.0285 | 4.9385 | 1.931
2012-10-30 21:07 | 28.6173 | 17.4208 | 17.5021 | 19.4081 | 16.4493 | -1.502 | 0.006744 | 0.32653 | 0.002202 | 48.4177 | 0.006946 | 0.2146 | 0.001491 | 30.8953 | 14.4777 | -0.1035 | 0.003693 | 2.9431 | 3.0244 | 4.9304 | 1.9716
2012-10-30 21:07 | 28.6143 | 17.4056 | 17.4869 | 19.4172 | 16.4301 | -1.5032 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.491 | -0.1036 | 0.003684 | 2.9146 | 2.9959 | 4.9262 | 1.9391
2012-10-30 21:07 | 28.6128 | 17.4001 | 17.4976 | 19.3752 | 16.4123 | -1.4916 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.006936 | 0.21461 | 0.001489 | 30.9419 | 14.4936 | -0.1028 | 0.003691 | 2.9065 | 3.004 | 4.8816 | 1.9187
2012-10-30 21:07 | 28.6128 | 17.3798 | 17.4976 | 19.4036 | 16.4042 | -1.4934 | 0.006729 | 0.32659 | 0.002198 | 48.5341 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.514 | -0.103 | 0.003688 | 2.8658 | 2.9836 | 4.8896 | 1.8902
2012-10-30 21:07 | 28.6113 | 17.4189 | 17.5001 | 19.398 | 16.3905 | -1.498 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006927 | 0.21466 | 0.001487 | 30.9883 | 14.4921 | -0.1033 | 0.003685 | 2.9268 | 3.008 | 4.9059 | 1.8984
2012-10-30 21:07 | 28.6053 | 17.3804 | 17.4819 | 19.3839 | 16.4211 | -1.4945 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.482 | -0.103 | 0.003689 | 2.8984 | 2.9999 | 4.9019 | 1.9391
2012-10-30 21:07 | 28.6001 | 17.3793 | 17.4524 | 19.3666 | 16.3955 | -1.5026 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.00693 | 0.21466 | 0.001488 | 30.9749 | 14.4972 | -0.1036 | 0.003689 | 2.8821 | 2.9552 | 4.8694 | 1.8983
2012-10-30 21:07 | 28.606 | 17.3648 | 17.442 | 19.4008 | 16.3852 | -1.4934 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4664 | -0.103 | 0.003688 | 2.8984 | 2.9756 | 4.9344 | 1.9188
2012-10-30 21:07 | 28.5911 | 17.3946 | 17.4556 | 19.3414 | 16.3581 | -1.5032 | 0.006719 | 0.32661 | 0.002195 | 48.6107 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4923 | -0.1036 | 0.003686 | 2.9023 | 2.9633 | 4.8491 | 1.8658
2012-10-30 21:07 | 28.5941 | 17.3854 | 17.4423 | 19.3363 | 16.357 | -1.498 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006927 | 0.21466 | 0.001487 | 30.9883 | 14.5116 | -0.1033 | 0.003684 | 2.8738 | 2.9307 | 4.8247 | 1.8454
2012-10-30 21:07 | 28.5978 | 17.3932 | 17.4663 | 19.3521 | 16.4217 | -1.5015 | 0.006716 | 0.32656 | 0.002193 | 48.6237 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.4786 | -0.1035 | 0.00368 | 2.9146 | 2.9877 | 4.8735 | 1.9431
2012-10-30 21:07 | 28.5918 | 17.3628 | 17.4278 | 19.3704 | 16.3994 | -1.4951 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006933 | 0.21461 | 0.001488 | 30.9553 | 14.4848 | -0.1031 | 0.003684 | 2.878 | 2.943 | 4.8856 | 1.9146
2012-10-30 21:07 | 28.5978 | 17.381 | 17.4744 | 19.3278 | 16.3851 | -1.5015 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.4255 | -0.1035 | 0.003683 | 2.9555 | 3.0489 | 4.9023 | 1.9596
2012-10-30 21:07 | 28.5903 | 17.3695 | 17.4101 | 19.292 | 16.3573 | -1.4928 | 0.006722 | 0.3266 | 0.002195 | 48.5869 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4547 | -0.1029 | 0.003687 | 2.9148 | 2.9554 | 4.8373 | 1.9026
2012-10-30 21:07 | 28.5933 | 17.2993 | 17.4131 | 19.2909 | 16.3643 | -1.4997 | 0.006725 | 0.3266 | 0.002196 | 48.5652 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.5271 | -0.1034 | 0.003684 | 2.7722 | 2.886 | 4.7638 | 1.8372
2012-10-30 21:08 | 28.5978 | 17.3241 | 17.4013 | 19.2914 | 16.3932 | -1.4916 | 0.006719 | 0.32661 | 0.002195 | 48.6107 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.499 | -0.1028 | 0.003683 | 2.8251 | 2.9023 | 4.7924 | 1.8942
2012-10-30 21:08 | 28.5881 | 17.3226 | 17.4201 | 19.3181 | 16.4039 | -1.5043 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.477 | -0.1037 | 0.003687 | 2.8456 | 2.9431 | 4.8411 | 1.9269
2012-10-30 21:08 | 28.5903 | 17.3085 | 17.4182 | 19.2555 | 16.3491 | -1.4916 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4996 | -0.1028 | 0.003691 | 2.8089 | 2.9186 | 4.7559 | 1.8495
2012-10-30 21:08 | 28.5821 | 17.3288 | 17.3978 | 19.2797 | 16.3694 | -1.4997 | 0.006725 | 0.32666 | 0.002197 | 48.5738 | 0.006921 | 0.21464 | 0.001486 | 31.0131 | 14.4425 | -0.1034 | 0.003683 | 2.8863 | 2.9553 | 4.8372 | 1.9269
2012-10-30 21:08 | 28.5888 | 17.2745 | 17.3598 | 19.2257 | 16.3598 | -1.4934 | 0.006735 | 0.3266 | 0.0022 | 48.4931 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4777 | -0.103 | 0.003691 | 2.7968 | 2.8821 | 4.748 | 1.8821
2012-10-30 21:08 | 28.5821 | 17.3003 | 17.3613 | 19.2392 | 16.3491 | -1.4864 | 0.006719 | 0.32663 | 0.002195 | 48.6129 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4547 | -0.1025 | 0.003684 | 2.8456 | 2.9066 | 4.7845 | 1.8944
2012-10-30 21:08 | 28.5941 | 17.2758 | 17.3651 | 19.2593 | 16.3366 | -1.4945 | 0.006719 | 0.32667 | 0.002195 | 48.6193 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4545 | -0.103 | 0.003682 | 2.8213 | 2.9106 | 4.8048 | 1.8821
2012-10-30 21:08 | 28.5873 | 17.269 | 17.3461 | 19.2647 | 16.342 | -1.4968 | 0.006722 | 0.32663 | 0.002196 | 48.5912 | 0.006921 | 0.21467 | 0.001486 | 31.0173 | 14.4599 | -0.1032 | 0.003682 | 2.8091 | 2.8862 | 4.8048 | 1.8821
2012-10-30 21:08 | 28.5851 | 17.2383 | 17.3724 | 19.2179 | 16.3114 | -1.4835 | 0.006725 | 0.32667 | 0.002197 | 48.5759 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4577 | -0.1023 | 0.003688 | 2.7806 | 2.9147 | 4.7602 | 1.8537
2012-10-30 21:08 | 28.5881 | 17.2982 | 17.3348 | 19.2412 | 16.3428 | -1.4939 | 0.006725 | 0.3266 | 0.002196 | 48.5652 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4729 | -0.103 | 0.003684 | 2.8253 | 2.8619 | 4.7683 | 1.8699
2012-10-30 21:08 | 28.5821 | 17.2841 | 17.3775 | 19.219 | 16.3531 | -1.4801 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.006924 | 0.21467 | 0.001486 | 31.0039 | 14.4139 | -0.102 | 0.003686 | 2.8702 | 2.9636 | 4.8051 | 1.9392
2012-10-30 21:08 | 28.5806 | 17.2094 | 17.3313 | 19.1972 | 16.2987 | -1.4916 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.494 | -0.1028 | 0.003689 | 2.7154 | 2.8373 | 4.7032 | 1.8047
2012-10-30 21:08 | 28.5926 | 17.2255 | 17.3514 | 19.2133 | 16.3555 | -1.509 | 0.006713 | 0.32667 | 0.002193 | 48.6627 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4978 | -0.104 | 0.003683 | 2.7277 | 2.8536 | 4.7155 | 1.8577
2012-10-30 21:08 | 28.5821 | 17.2272 | 17.3328 | 19.223 | 16.3409 | -1.498 | 0.006732 | 0.32663 | 0.002199 | 48.519 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.4792 | -0.1033 | 0.003688 | 2.748 | 2.8536 | 4.7438 | 1.8617
2012-10-30 21:08 | 28.5798 | 17.2046 | 17.3508 | 19.1721 | 16.3061 | -1.4951 | 0.006735 | 0.32664 | 0.0022 | 48.4995 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.5014 | -0.1031 | 0.003688 | 2.7032 | 2.8494 | 4.6707 | 1.8047
2012-10-30 21:08 | 28.5933 | 17.1856 | 17.3237 | 19.1856 | 16.3521 | -1.5038 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5026 | -0.1037 | 0.003685 | 2.683 | 2.8211 | 4.683 | 1.8495
2012-10-30 21:08 | 28.5896 | 17.1738 | 17.3484 | 19.1495 | 16.3077 | -1.4997 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006943 | 0.2146 | 0.00149 | 30.9087 | 14.5112 | -0.1034 | 0.003686 | 2.6626 | 2.8372 | 4.6383 | 1.7965
2012-10-30 21:08 | 28.5828 | 17.1873 | 17.3173 | 19.1913 | 16.3335 | -1.4986 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.4921 | -0.1033 | 0.003687 | 2.6952 | 2.8252 | 4.6992 | 1.8414
2012-10-30 21:08 | 28.5768 | 17.1894 | 17.2666 | 19.1813 | 16.3153 | -1.5067 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4861 | -0.1039 | 0.003684 | 2.7033 | 2.7805 | 4.6952 | 1.8292
2012-10-30 21:08 | 28.5806 | 17.1648 | 17.2826 | 19.1243 | 16.3313 | -1.498 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.4736 | -0.1033 | 0.003685 | 2.6912 | 2.809 | 4.6507 | 1.8577
2012-10-30 21:08 | 28.5836 | 17.1718 | 17.3343 | 19.1435 | 16.3424 | -1.502 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4603 | -0.1035 | 0.003689 | 2.7115 | 2.874 | 4.6832 | 1.8821
2012-10-30 21:08 | 28.5828 | 17.1629 | 17.3132 | 19.167 | 16.3294 | -1.5015 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006939 | 0.21457 | 0.001489 | 30.9223 | 14.5247 | -0.1035 | 0.003686 | 2.6382 | 2.7885 | 4.6423 | 1.8047
2012-10-30 21:08 | 28.5911 | 17.1834 | 17.3337 | 19.1145 | 16.3174 | -1.4968 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5657 | -0.1032 | 0.003692 | 2.6177 | 2.768 | 4.5488 | 1.7517
2012-10-30 21:08 | 28.5768 | 17.1894 | 17.3072 | 19.1286 | 16.3234 | -1.487 | 0.006725 | 0.32663 | 0.002197 | 48.5695 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4657 | -0.1025 | 0.003686 | 2.7237 | 2.8415 | 4.6629 | 1.8577
2012-10-30 21:08 | 28.5873 | 17.1552 | 17.3258 | 19.1189 | 16.3217 | -1.498 | 0.006716 | 0.3266 | 0.002193 | 48.6303 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4966 | -0.1033 | 0.003683 | 2.6586 | 2.8292 | 4.6223 | 1.8251
2012-10-30 21:08 | 28.5873 | 17.1471 | 17.3218 | 19.1148 | 16.281 | -1.5038 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006939 | 0.21463 | 0.001489 | 30.9307 | 14.4966 | -0.1037 | 0.003685 | 2.6505 | 2.8252 | 4.6182 | 1.7844
2012-10-30 21:09 | 28.5768 | 17.1813 | 17.3316 | 19.1002 | 16.2909 | -1.487 | 0.006719 | 0.32661 | 0.002195 | 48.6107 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5024 | -0.1025 | 0.003684 | 2.6789 | 2.8292 | 4.5978 | 1.7885
2012-10-30 21:09 | 28.5761 | 17.144 | 17.2821 | 19.0955 | 16.2983 | -1.4968 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.5221 | -0.1032 | 0.003688 | 2.6219 | 2.76 | 4.5734 | 1.7762
2012-10-30 21:09 | 28.5806 | 17.1566 | 17.2582 | 19.1081 | 16.3313 | -1.4951 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.494 | -0.1031 | 0.003691 | 2.6626 | 2.7642 | 4.6141 | 1.8373
2012-10-30 21:09 | 28.5806 | 17.1485 | 17.2704 | 19.0797 | 16.3069 | -1.4951 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4981 | -0.1031 | 0.00369 | 2.6504 | 2.7723 | 4.5816 | 1.8088
2012-10-30 21:09 | 28.5858 | 17.109 | 17.2634 | 19.0647 | 16.3161 | -1.4974 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.5277 | -0.1032 | 0.003686 | 2.5813 | 2.7357 | 4.537 | 1.7884
2012-10-30 21:09 | 28.5761 | 17.1034 | 17.2821 | 19.0914 | 16.2861 | -1.4916 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5017 | -0.1028 | 0.003687 | 2.6017 | 2.7804 | 4.5897 | 1.7844
2012-10-30 21:09 | 28.5791 | 17.1104 | 17.2283 | 19.0661 | 16.2525 | -1.4882 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006927 | 0.21466 | 0.001487 | 30.9883 | 14.5333 | -0.1026 | 0.003685 | 2.5771 | 2.695 | 4.5328 | 1.7192
2012-10-30 21:09 | 28.5776 | 17.0845 | 17.2714 | 19.0484 | 16.2673 | -1.4945 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006933 | 0.2146 | 0.001488 | 30.9533 | 14.5032 | -0.103 | 0.003684 | 2.5813 | 2.7682 | 4.5452 | 1.7641
2012-10-30 21:09 | 28.5843 | 17.0953 | 17.2741 | 19.0753 | 16.2902 | -1.4934 | 0.006747 | 0.32657 | 0.002203 | 48.4025 | 0.00693 | 0.21466 | 0.001488 | 30.9749 | 14.5426 | -0.103 | 0.003691 | 2.5527 | 2.7315 | 4.5327 | 1.7476
2012-10-30 21:09 | 28.5791 | 17.0657 | 17.1998 | 19.0701 | 16.281 | -1.4858 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.5047 | -0.1024 | 0.00369 | 2.561 | 2.6951 | 4.5654 | 1.7763
2012-10-30 21:09 | 28.5783 | 17.1096 | 17.2843 | 19.0369 | 16.2395 | -1.4864 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006939 | 0.21461 | 0.001489 | 30.9285 | 14.4387 | -0.1025 | 0.003685 | 2.6709 | 2.8456 | 4.5982 | 1.8008
2012-10-30 21:09 | 28.5753 | 17.066 | 17.2691 | 19.0258 | 16.2527 | -1.4922 | 0.006719 | 0.32653 | 0.002194 | 48.5979 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4969 | -0.1029 | 0.003682 | 2.5691 | 2.7722 | 4.5289 | 1.7558
2012-10-30 21:09 | 28.5836 | 17.0662 | 17.2328 | 19.022 | 16.3099 | -1.498 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006943 | 0.21461 | 0.00149 | 30.9107 | 14.444 | -0.1033 | 0.003689 | 2.6222 | 2.7888 | 4.578 | 1.8659
2012-10-30 21:09 | 28.5821 | 17.0565 | 17.2516 | 19.0245 | 16.2473 | -1.5009 | 0.006744 | 0.3266 | 0.002203 | 48.4284 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5037 | -0.1035 | 0.003692 | 2.5528 | 2.7479 | 4.5208 | 1.7436
2012-10-30 21:09 | 28.5821 | 17.1012 | 17.2313 | 18.9597 | 16.2433 | -1.4934 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006936 | 0.21463 | 0.001489 | 30.9441 | 14.4914 | -0.103 | 0.003689 | 2.6098 | 2.7399 | 4.4683 | 1.7519
2012-10-30 21:09 | 28.5753 | 17.0579 | 17.2123 | 18.9651 | 16.2568 | -1.4939 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.4561 | -0.103 | 0.003686 | 2.6018 | 2.7562 | 4.509 | 1.8007
2012-10-30 21:09 | 28.5731 | 17.0435 | 17.2304 | 18.9629 | 16.218 | -1.5015 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006924 | 0.21463 | 0.001486 | 30.9977 | 14.4987 | -0.1035 | 0.003687 | 2.5448 | 2.7317 | 4.4642 | 1.7193
2012-10-30 21:09 | 28.5693 | 17.0112 | 17.2063 | 18.9915 | 16.2305 | -1.483 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4337 | -0.1022 | 0.003687 | 2.5775 | 2.7726 | 4.5578 | 1.7968
2012-10-30 21:09 | 28.5723 | 17.0508 | 17.2174 | 18.9864 | 16.2579 | -1.4841 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.4571 | -0.1023 | 0.00369 | 2.5937 | 2.7603 | 4.5293 | 1.8008
2012-10-30 21:09 | 28.5783 | 17.004 | 17.199 | 18.9356 | 16.2517 | -1.4916 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4795 | -0.1028 | 0.003688 | 2.5245 | 2.7195 | 4.4561 | 1.7722
2012-10-30 21:09 | 28.5761 | 17.0952 | 17.2049 | 18.9699 | 16.2129 | -1.4986 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006927 | 0.21464 | 0.001487 | 30.9863 | 14.4609 | -0.1033 | 0.003689 | 2.6343 | 2.744 | 4.509 | 1.752
2012-10-30 21:09 | 28.5723 | 17.0264 | 17.2093 | 18.9661 | 16.2701 | -1.4934 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4979 | -0.103 | 0.003684 | 2.5285 | 2.7114 | 4.4682 | 1.7722
2012-10-30 21:09 | 28.5559 | 16.9856 | 17.1766 | 18.9862 | 16.2578 | -1.4986 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4815 | -0.1033 | 0.003686 | 2.5041 | 2.6951 | 4.5047 | 1.7763
2012-10-30 21:09 | 28.5693 | 17.0275 | 17.1738 | 18.9347 | 16.2386 | -1.4876 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4705 | -0.1026 | 0.003685 | 2.557 | 2.7033 | 4.4642 | 1.7681
2012-10-30 21:09 | 28.5596 | 17.0097 | 17.1763 | 18.9534 | 16.2126 | -1.4887 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4771 | -0.1026 | 0.003681 | 2.5326 | 2.6992 | 4.4763 | 1.7355
2012-10-30 21:09 | 28.5604 | 17.0145 | 17.1933 | 18.9096 | 16.246 | -1.4934 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.4575 | -0.103 | 0.003691 | 2.557 | 2.7358 | 4.4521 | 1.7885
2012-10-30 21:09 | 28.5566 | 16.9701 | 17.1895 | 18.9545 | 16.2056 | -1.4818 | 0.006719 | 0.32667 | 0.002195 | 48.6193 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4251 | -0.1022 | 0.003683 | 2.545 | 2.7644 | 4.5294 | 1.7805
2012-10-30 21:10 | 28.5581 | 17.0163 | 17.1991 | 18.9276 | 16.2111 | -1.5084 | 0.006707 | 0.32657 | 0.00219 | 48.6912 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4307 | -0.104 | 0.003681 | 2.5856 | 2.7684 | 4.4969 | 1.7804
2012-10-30 21:10 | 28.5484 | 16.9822 | 17.161 | 18.8814 | 16.2421 | -1.4991 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4618 | -0.1033 | 0.003687 | 2.5204 | 2.6992 | 4.4196 | 1.7803
2012-10-30 21:10 | 28.5551 | 16.993 | 17.1393 | 18.9246 | 16.2203 | -1.4939 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006955 | 0.21467 | 0.001493 | 30.8657 | 14.4359 | -0.103 | 0.003689 | 2.5571 | 2.7034 | 4.4887 | 1.7844
2012-10-30 21:10 | 28.5506 | 16.96 | 17.1754 | 18.8877 | 16.228 | -1.5032 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4681 | -0.1036 | 0.003685 | 2.4919 | 2.7073 | 4.4196 | 1.7599
2012-10-30 21:10 | 28.5634 | 16.985 | 17.1597 | 18.8843 | 16.1961 | -1.4974 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4686 | -0.1032 | 0.003687 | 2.5164 | 2.6911 | 4.4157 | 1.7275
2012-10-30 21:10 | 28.5514 | 16.9771 | 17.1721 | 18.8723 | 16.2044 | -1.4899 | 0.006735 | 0.3266 | 0.0022 | 48.4931 | 0.006917 | 0.21467 | 0.001485 | 31.0353 | 14.4199 | -0.1027 | 0.003685 | 2.5572 | 2.7522 | 4.4524 | 1.7845
2012-10-30 21:10 | 28.5506 | 16.9722 | 17.1185 | 18.9039 | 16.2036 | -1.4963 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4925 | -0.1032 | 0.003687 | 2.4797 | 2.626 | 4.4114 | 1.7111
2012-10-30 21:10 | 28.5514 | 16.973 | 17.1396 | 18.8601 | 16.2085 | -1.4882 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4444 | -0.1026 | 0.003684 | 2.5286 | 2.6952 | 4.4157 | 1.7641
2012-10-30 21:10 | 28.5446 | 16.9662 | 17.1125 | 18.8452 | 16.222 | -1.4899 | 0.006722 | 0.32666 | 0.002196 | 48.5955 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4784 | -0.1027 | 0.003685 | 2.4878 | 2.6341 | 4.3668 | 1.7436
2012-10-30 21:10 | 28.5371 | 16.9384 | 17.1294 | 18.8296 | 16.1983 | -1.4905 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4056 | -0.1028 | 0.003687 | 2.5328 | 2.7238 | 4.424 | 1.7927
2012-10-30 21:10 | 28.5461 | 16.9514 | 17.0652 | 18.8427 | 16.2032 | -1.4939 | 0.006713 | 0.32661 | 0.002193 | 48.6541 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4432 | -0.103 | 0.003684 | 2.5082 | 2.622 | 4.3995 | 1.76
2012-10-30 21:10 | 28.5431 | 16.9484 | 17.1191 | 18.8234 | 16.1921 | -1.5055 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006921 | 0.21471 | 0.001486 | 31.0236 | 14.4402 | -0.1038 | 0.003682 | 2.5082 | 2.6789 | 4.3832 | 1.7519
2012-10-30 21:10 | 28.5499 | 16.9552 | 17.1341 | 18.8546 | 16.1948 | -1.4934 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.4347 | -0.103 | 0.003685 | 2.5205 | 2.6994 | 4.4199 | 1.7601
2012-10-30 21:10 | 28.5454 | 16.9264 | 17.0727 | 18.8217 | 16.1537 | -1.4968 | 0.006722 | 0.32661 | 0.002196 | 48.589 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.418 | -0.1032 | 0.003685 | 2.5084 | 2.6547 | 4.4037 | 1.7357
2012-10-30 21:10 | 28.5431 | 16.9484 | 17.1029 | 18.8032 | 16.1676 | -1.4911 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4116 | -0.1028 | 0.003688 | 2.5368 | 2.6913 | 4.3916 | 1.756
2012-10-30 21:10 | 28.5484 | 16.9416 | 17.1204 | 18.8328 | 16.1648 | -1.5107 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4332 | -0.1041 | 0.003684 | 2.5084 | 2.6872 | 4.3996 | 1.7316
2012-10-30 21:10 | 28.5416 | 16.9063 | 17.1176 | 18.8138 | 16.1702 | -1.5009 | 0.006732 | 0.3266 | 0.002199 | 48.5147 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4917 | -0.1035 | 0.003686 | 2.4146 | 2.6259 | 4.3221 | 1.6785
2012-10-30 21:10 | 28.5371 | 16.8571 | 17.0847 | 18.7688 | 16.1494 | -1.4963 | 0.006722 | 0.32659 | 0.002195 | 48.5846 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4301 | -0.1032 | 0.003686 | 2.427 | 2.6546 | 4.3387 | 1.7193
2012-10-30 21:10 | 28.5289 | 16.8895 | 17.048 | 18.7606 | 16.1372 | -1.5055 | 0.006716 | 0.3266 | 0.002193 | 48.6303 | 0.006943 | 0.21464 | 0.00149 | 30.9149 | 14.4056 | -0.1038 | 0.003683 | 2.4839 | 2.6424 | 4.355 | 1.7316
2012-10-30 21:10 | 28.5454 | 16.9182 | 17.0889 | 18.765 | 16.1333 | -1.4887 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.4506 | -0.1026 | 0.003687 | 2.4676 | 2.6383 | 4.3144 | 1.6827
2012-10-30 21:10 | 28.5349 | 16.8996 | 17.0987 | 18.7504 | 16.1472 | -1.4968 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4524 | -0.1032 | 0.00369 | 2.4472 | 2.6463 | 4.298 | 1.6948
2012-10-30 21:10 | 28.5379 | 16.866 | 17.1058 | 18.7575 | 16.1624 | -1.4939 | 0.006732 | 0.32666 | 0.002199 | 48.5233 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.4554 | -0.103 | 0.003688 | 2.4106 | 2.6504 | 4.3021 | 1.707
2012-10-30 21:10 | 28.5379 | 16.9026 | 17.0489 | 18.7453 | 16.1258 | -1.4928 | 0.006738 | 0.32663 | 0.002201 | 48.4758 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.4431 | -0.1029 | 0.003689 | 2.4595 | 2.6058 | 4.3022 | 1.6827
2012-10-30 21:10 | 28.5454 | 16.8369 | 17.0645 | 18.7731 | 16.1903 | -1.4939 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.006943 | 0.2146 | 0.00149 | 30.9087 | 14.4139 | -0.103 | 0.003692 | 2.423 | 2.6506 | 4.3592 | 1.7764
2012-10-30 21:10 | 28.5401 | 16.8398 | 17.0633 | 18.7434 | 16.124 | -1.4916 | 0.006719 | 0.3266 | 0.002194 | 48.6086 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.482 | -0.1028 | 0.003682 | 2.3578 | 2.5813 | 4.2614 | 1.642
2012-10-30 21:11 | 28.5267 | 16.8345 | 17.0458 | 18.7057 | 16.135 | -1.5095 | 0.006729 | 0.32663 | 0.002198 | 48.5406 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4156 | -0.1041 | 0.00369 | 2.4189 | 2.6302 | 4.2901 | 1.7194
2012-10-30 21:11 | 28.5409 | 16.8365 | 17.0682 | 18.7524 | 16.1492 | -1.4986 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4461 | -0.1033 | 0.003689 | 2.3904 | 2.6221 | 4.3063 | 1.7031
2012-10-30 21:11 | 28.5282 | 16.8075 | 17.0473 | 18.6951 | 16.108 | -1.4991 | 0.006716 | 0.32664 | 0.002194 | 48.6367 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4538 | -0.1033 | 0.003682 | 2.3537 | 2.5935 | 4.2413 | 1.6542
2012-10-30 21:11 | 28.5326 | 16.816 | 17.0233 | 18.7197 | 16.1002 | -1.4928 | 0.006725 | 0.32663 | 0.002197 | 48.5695 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.4378 | -0.1029 | 0.003686 | 2.3782 | 2.5855 | 4.2819 | 1.6624
2012-10-30 21:11 | 28.5274 | 16.823 | 17.0343 | 18.674 | 16.0828 | -1.4991 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006924 | 0.21471 | 0.001487 | 31.0101 | 14.3918 | -0.1033 | 0.003685 | 2.4312 | 2.6425 | 4.2822 | 1.691
2012-10-30 21:11 | 28.5371 | 16.8205 | 17.0197 | 18.704 | 16.0884 | -1.4991 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.4219 | -0.1033 | 0.003685 | 2.3986 | 2.5978 | 4.2821 | 1.6665
2012-10-30 21:11 | 28.5334 | 16.7924 | 16.9794 | 18.6557 | 16.0928 | -1.4968 | 0.006729 | 0.32666 | 0.002198 | 48.545 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4631 | -0.1032 | 0.003689 | 2.3293 | 2.5163 | 4.1926 | 1.6297
2012-10-30 21:11 | 28.5304 | 16.7813 | 17.0089 | 18.6851 | 16.0858 | -1.4986 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.00693 | 0.21467 | 0.001488 | 30.9771 | 14.4275 | -0.1033 | 0.003687 | 2.3538 | 2.5814 | 4.2576 | 1.6583
2012-10-30 21:11 | 28.5364 | 16.8483 | 17.023 | 18.6384 | 16.0836 | -1.5038 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.4539 | -0.1037 | 0.003691 | 2.3944 | 2.5691 | 4.1845 | 1.6297
2012-10-30 21:11 | 28.5326 | 16.8079 | 17.0395 | 18.6792 | 16.1083 | -1.5026 | 0.006719 | 0.32657 | 0.002194 | 48.6043 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4501 | -0.1036 | 0.003685 | 2.3578 | 2.5894 | 4.2291 | 1.6582
2012-10-30 21:11 | 28.5394 | 16.8228 | 16.961 | 18.6536 | 16.1395 | -1.4963 | 0.006725 | 0.3266 | 0.002196 | 48.5652 | 0.006955 | 0.21474 | 0.001494 | 30.8761 | 14.4691 | -0.1032 | 0.00369 | 2.3537 | 2.4919 | 4.1845 | 1.6704
2012-10-30 21:11 | 28.5311 | 16.7739 | 17.0502 | 18.6331 | 16.0865 | -1.498 | 0.006722 | 0.32659 | 0.002195 | 48.5846 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.4282 | -0.1033 | 0.003686 | 2.3457 | 2.622 | 4.2049 | 1.6583
2012-10-30 21:11 | 28.5319 | 16.799 | 17.0267 | 18.6461 | 16.0588 | -1.4887 | 0.006744 | 0.3266 | 0.002203 | 48.4284 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.4167 | -0.1026 | 0.003692 | 2.3823 | 2.61 | 4.2294 | 1.6421
2012-10-30 21:11 | 28.5289 | 16.7595 | 16.9993 | 18.6431 | 16.0924 | -1.4876 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4627 | -0.1026 | 0.003689 | 2.2968 | 2.5366 | 4.1804 | 1.6297
2012-10-30 21:11 | 28.5401 | 16.7422 | 17.0023 | 18.6664 | 16.1036 | -1.502 | 0.006738 | 0.32661 | 0.002201 | 48.4736 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.5065 | -0.1035 | 0.003689 | 2.2357 | 2.4958 | 4.1599 | 1.5971
2012-10-30 21:11 | 28.5304 | 16.7406 | 17.0333 | 18.6081 | 16.0654 | -1.4991 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006943 | 0.21466 | 0.00149 | 30.9169 | 14.4764 | -0.1033 | 0.003689 | 2.2642 | 2.5569 | 4.1317 | 1.589
2012-10-30 21:11 | 28.5289 | 16.731 | 16.9505 | 18.6552 | 16.0843 | -1.5009 | 0.006744 | 0.32664 | 0.002203 | 48.4348 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.4912 | -0.1035 | 0.003695 | 2.2398 | 2.4593 | 4.164 | 1.5931
2012-10-30 21:11 | 28.5274 | 16.7092 | 16.9815 | 18.6254 | 16.095 | -1.4934 | 0.006741 | 0.32661 | 0.002202 | 48.452 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.4775 | -0.103 | 0.003694 | 2.2317 | 2.504 | 4.1479 | 1.6175
2012-10-30 21:11 | 28.5297 | 16.7643 | 16.9676 | 18.6277 | 16.0362 | -1.5015 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.4594 | -0.1035 | 0.003691 | 2.3049 | 2.5082 | 4.1683 | 1.5768
2012-10-30 21:11 | 28.5319 | 16.734 | 16.9535 | 18.6055 | 16.1117 | -1.4957 | 0.006722 | 0.3266 | 0.002195 | 48.5869 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4698 | -0.1031 | 0.003686 | 2.2642 | 2.4837 | 4.1357 | 1.6419
2012-10-30 21:11 | 28.5394 | 16.7252 | 16.9732 | 18.6049 | 16.0256 | -1.4858 | 0.006722 | 0.3266 | 0.002195 | 48.5869 | 0.006927 | 0.21463 | 0.001487 | 30.9843 | 14.4813 | -0.1024 | 0.003682 | 2.2439 | 2.4919 | 4.1236 | 1.5443
2012-10-30 21:11 | 28.5334 | 16.703 | 17.0078 | 18.5746 | 16.0562 | -1.4968 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.4672 | -0.1032 | 0.003688 | 2.2358 | 2.5406 | 4.1074 | 1.589
2012-10-30 21:11 | 28.5364 | 16.7385 | 16.9417 | 18.5492 | 16.0633 | -1.5026 | 0.006732 | 0.3267 | 0.002199 | 48.5297 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4824 | -0.1036 | 0.003688 | 2.2561 | 2.4593 | 4.0668 | 1.5809
2012-10-30 21:11 | 28.5282 | 16.6896 | 17.0067 | 18.5613 | 16.0673 | -1.5067 | 0.006741 | 0.32661 | 0.002202 | 48.452 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.4498 | -0.1039 | 0.003691 | 2.2398 | 2.5569 | 4.1115 | 1.6175
2012-10-30 21:11 | 28.5356 | 16.7174 | 16.9491 | 18.5687 | 16.0747 | -1.4957 | 0.006735 | 0.32666 | 0.0022 | 48.5017 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.4653 | -0.1031 | 0.003689 | 2.2521 | 2.4838 | 4.1034 | 1.6094
2012-10-30 21:12 | 28.5304 | 16.6553 | 16.9317 | 18.5311 | 16.0369 | -1.4974 | 0.006747 | 0.32657 | 0.002203 | 48.4025 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.4887 | -0.1032 | 0.003694 | 2.1666 | 2.443 | 4.0424 | 1.5482
2012-10-30 21:12 | 28.5356 | 16.6645 | 16.9531 | 18.5241 | 16.0625 | -1.4951 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006921 | 0.21473 | 0.001486 | 31.0257 | 14.4939 | -0.1031 | 0.003687 | 2.1706 | 2.4592 | 4.0302 | 1.5686
2012-10-30 21:12 | 28.5341 | 16.663 | 16.9435 | 18.5631 | 16.0366 | -1.5072 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.4801 | -0.1039 | 0.00369 | 2.1829 | 2.4634 | 4.083 | 1.5565
2012-10-30 21:12 | 28.5304 | 16.6918 | 16.9439 | 18.5108 | 16.0654 | -1.5072 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.452 | -0.1039 | 0.003687 | 2.2398 | 2.4919 | 4.0588 | 1.6134
2012-10-30 21:12 | 28.5326 | 16.6697 | 16.9136 | 18.4927 | 16.0066 | -1.4957 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4378 | -0.1031 | 0.003691 | 2.2319 | 2.4758 | 4.0549 | 1.5688
2012-10-30 21:12 | 28.5349 | 16.6638 | 16.9484 | 18.5031 | 16.0089 | -1.4997 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.485 | -0.1034 | 0.003692 | 2.1788 | 2.4634 | 4.0181 | 1.5239
2012-10-30 21:12 | 28.5409 | 16.6048 | 16.9259 | 18.4686 | 16.0393 | -1.5032 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.4788 | -0.1036 | 0.003686 | 2.126 | 2.4471 | 3.9898 | 1.5605
2012-10-30 21:12 | 28.5431 | 16.6232 | 16.9241 | 18.4951 | 15.9967 | -1.4963 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4647 | -0.1032 | 0.00369 | 2.1585 | 2.4594 | 4.0304 | 1.532
2012-10-30 21:12 | 28.5409 | 16.6292 | 16.9015 | 18.5132 | 16.0149 | -1.5072 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4339 | -0.1039 | 0.00369 | 2.1953 | 2.4676 | 4.0793 | 1.581
2012-10-30 21:12 | 28.5424 | 16.6347 | 16.9234 | 18.4741 | 16.0001 | -1.4858 | 0.006744 | 0.32664 | 0.002203 | 48.4348 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.4476 | -0.1024 | 0.003693 | 2.1871 | 2.4758 | 4.0265 | 1.5525
2012-10-30 21:12 | 28.5536 | 16.6134 | 16.9427 | 18.5056 | 15.9747 | -1.4876 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4711 | -0.1026 | 0.003687 | 2.1423 | 2.4716 | 4.0345 | 1.5036
2012-10-30 21:12 | 28.5461 | 16.5774 | 16.9108 | 18.4616 | 15.9916 | -1.4951 | 0.006729 | 0.32663 | 0.002198 | 48.5406 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.4554 | -0.1031 | 0.003688 | 2.122 | 2.4554 | 4.0062 | 1.5362
2012-10-30 21:12 | 28.5589 | 16.5821 | 16.9073 | 18.446 | 15.984 | -1.4968 | 0.006729 | 0.32661 | 0.002198 | 48.5384 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4805 | -0.1032 | 0.003688 | 2.1016 | 2.4268 | 3.9655 | 1.5035
2012-10-30 21:12 | 28.5589 | 16.5414 | 16.9439 | 18.3771 | 16.0044 | -1.5032 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.5049 | -0.1036 | 0.003687 | 2.0365 | 2.439 | 3.8722 | 1.4995
2012-10-30 21:12 | 28.5634 | 16.5541 | 16.9078 | 18.3816 | 15.956 | -1.5084 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.489 | -0.104 | 0.003689 | 2.0651 | 2.4188 | 3.8926 | 1.467
2012-10-30 21:13 | 28.5604 | 16.4941 | 16.9088 | 18.3867 | 15.957 | -1.4997 | 0.006725 | 0.32663 | 0.002197 | 48.5695 | 0.006924 | 0.21468 | 0.001486 | 31.0059 | 14.5064 | -0.1034 | 0.003683 | 1.9877 | 2.4024 | 3.8803 | 1.4506
2012-10-30 21:13 | 28.5596 | 16.5706 | 16.9243 | 18.4224 | 15.9847 | -1.5032 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.526 | -0.1036 | 0.003692 | 2.0446 | 2.3983 | 3.8964 | 1.4587
2012-10-30 21:13 | 28.5596 | 16.4771 | 16.8552 | 18.3778 | 16.0173 | -1.5032 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.006917 | 0.21474 | 0.001485 | 31.0457 | 14.4812 | -0.1036 | 0.003684 | 1.9959 | 2.374 | 3.8966 | 1.5361
2012-10-30 21:13 | 28.5589 | 16.4926 | 16.8626 | 18.3933 | 16.0003 | -1.4986 | 0.006744 | 0.32666 | 0.002203 | 48.437 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4968 | -0.1033 | 0.003695 | 1.9958 | 2.3658 | 3.8965 | 1.5035
2012-10-30 21:13 | 28.5574 | 16.5196 | 16.8937 | 18.3999 | 15.9703 | -1.4968 | 0.006725 | 0.3266 | 0.002196 | 48.5652 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.4708 | -0.1032 | 0.003688 | 2.0488 | 2.4229 | 3.9291 | 1.4995
2012-10-30 21:13 | 28.5619 | 16.4997 | 16.9103 | 18.3598 | 15.983 | -1.5026 | 0.006738 | 0.32669 | 0.002201 | 48.4843 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5446 | -0.1036 | 0.003691 | 1.9551 | 2.3657 | 3.8152 | 1.4384
2012-10-30 21:13 | 28.5589 | 16.5089 | 16.8708 | 18.3284 | 15.9678 | -1.4922 | 0.006738 | 0.32666 | 0.002201 | 48.4801 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.5294 | -0.1029 | 0.003689 | 1.9795 | 2.3414 | 3.799 | 1.4384
2012-10-30 21:13 | 28.5566 | 16.4619 | 16.8888 | 18.3707 | 15.9655 | -1.4951 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006946 | 0.21477 | 0.001492 | 30.9202 | 14.4863 | -0.1031 | 0.00369 | 1.9756 | 2.4025 | 3.8844 | 1.4792
2012-10-30 21:13 | 28.5611 | 16.4664 | 16.8404 | 18.3387 | 15.9333 | -1.5049 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.4704 | -0.1037 | 0.003686 | 1.996 | 2.37 | 3.8683 | 1.4629
2012-10-30 21:13 | 28.5551 | 16.4563 | 16.8588 | 18.3733 | 15.9436 | -1.498 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4848 | -0.1033 | 0.003687 | 1.9715 | 2.374 | 3.8885 | 1.4588
2012-10-30 21:13 | 28.5641 | 16.4449 | 16.8516 | 18.3133 | 15.9892 | -1.4974 | 0.006741 | 0.3266 | 0.002202 | 48.4499 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5346 | -0.1032 | 0.003692 | 1.9103 | 2.317 | 3.7787 | 1.4546
2012-10-30 21:13 | 28.5604 | 16.4657 | 16.8519 | 18.3137 | 15.9611 | -1.4887 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006946 | 0.21466 | 0.001491 | 30.9035 | 14.4779 | -0.1026 | 0.003687 | 1.9878 | 2.374 | 3.8358 | 1.4832
2012-10-30 21:13 | 28.5656 | 16.5115 | 16.8327 | 18.3473 | 15.9745 | -1.4945 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.4953 | -0.103 | 0.003687 | 2.0162 | 2.3374 | 3.852 | 1.4792
2012-10-30 21:13 | 28.5678 | 16.4486 | 16.8187 | 18.3292 | 15.9278 | -1.4968 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 14.469 | -0.1032 | 0.003685 | 1.9796 | 2.3497 | 3.8602 | 1.4588
2012-10-30 21:13 | 28.5686 | 16.4047 | 16.852 | 18.3097 | 15.949 | -1.4991 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 14.4942 | -0.1033 | 0.003691 | 1.9105 | 2.3578 | 3.8155 | 1.4548
2012-10-30 21:13 | 28.5693 | 16.4298 | 16.8446 | 18.2983 | 15.9375 | -1.4997 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5276 | -0.1034 | 0.00369 | 1.9022 | 2.317 | 3.7707 | 1.4099
2012-10-30 21:13 | 28.5753 | 16.4195 | 16.8546 | 18.288 | 15.9516 | -1.4945 | 0.006747 | 0.32653 | 0.002203 | 48.3962 | 0.006924 | 0.2147 | 0.001487 | 31.0081 | 14.5417 | -0.103 | 0.00369 | 1.8778 | 2.3129 | 3.7463 | 1.4099
2012-10-30 21:13 | 28.5768 | 16.3966 | 16.8236 | 18.2652 | 15.9531 | -1.5009 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006943 | 0.21467 | 0.00149 | 30.9191 | 14.5432 | -0.1035 | 0.003688 | 1.8534 | 2.2804 | 3.722 | 1.4099
2012-10-30 21:13 | 28.5686 | 16.3844 | 16.8236 | 18.2732 | 15.9449 | -1.4991 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.4983 | -0.1033 | 0.003687 | 1.8861 | 2.3253 | 3.7749 | 1.4466
2012-10-30 21:13 | 28.5731 | 16.4458 | 16.8281 | 18.3021 | 15.982 | -1.4951 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5232 | -0.1031 | 0.003689 | 1.9226 | 2.3049 | 3.7789 | 1.4588
2012-10-30 21:13 | 28.5768 | 16.417 | 16.848 | 18.3139 | 15.9409 | -1.5159 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5391 | -0.1045 | 0.003687 | 1.8779 | 2.3089 | 3.7748 | 1.4018
2012-10-30 21:13 | 28.5753 | 16.4114 | 16.8709 | 18.2434 | 15.9353 | -1.4911 | 0.006719 | 0.32651 | 0.002194 | 48.5956 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5336 | -0.1028 | 0.003683 | 1.8778 | 2.3373 | 3.7098 | 1.4017
2012-10-30 21:13 | 28.5738 | 16.3977 | 16.8409 | 18.2744 | 15.9582 | -1.5038 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5198 | -0.1037 | 0.003689 | 1.8779 | 2.3211 | 3.7546 | 1.4384
2012-10-30 21:13 | 28.5678 | 16.351 | 16.8024 | 18.24 | 15.936 | -1.4986 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5016 | -0.1033 | 0.003685 | 1.8494 | 2.3008 | 3.7384 | 1.4344
2012-10-30 21:13 | 28.5821 | 16.3979 | 16.8411 | 18.2948 | 15.938 | -1.4951 | 0.006738 | 0.32659 | 0.002201 | 48.4693 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5077 | -0.1031 | 0.003691 | 1.8902 | 2.3334 | 3.7871 | 1.4303
2012-10-30 21:14 | 28.5783 | 16.4063 | 16.8048 | 18.2586 | 15.9302 | -1.4905 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4958 | -0.1028 | 0.003687 | 1.9105 | 2.309 | 3.7628 | 1.4344
2012-10-30 21:14 | 28.5723 | 16.3921 | 16.7663 | 18.2648 | 15.9282 | -1.4974 | 0.006716 | 0.3266 | 0.002193 | 48.6303 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.5387 | -0.1032 | 0.00368 | 1.8534 | 2.2276 | 3.7261 | 1.3895
2012-10-30 21:14 | 28.5723 | 16.3393 | 16.7988 | 18.2323 | 15.9445 | -1.513 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.551 | -0.1043 | 0.003688 | 1.7883 | 2.2478 | 3.6813 | 1.3935
2012-10-30 21:14 | 28.5686 | 16.4169 | 16.7788 | 18.2205 | 15.9123 | -1.5038 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5024 | -0.1037 | 0.00369 | 1.9145 | 2.2764 | 3.7181 | 1.4099
2012-10-30 21:14 | 28.5761 | 16.3837 | 16.7619 | 18.2442 | 15.9158 | -1.4916 | 0.006744 | 0.32656 | 0.002202 | 48.4219 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.5303 | -0.1028 | 0.003692 | 1.8534 | 2.2316 | 3.7139 | 1.3855
2012-10-30 21:14 | 28.5731 | 16.3767 | 16.8524 | 18.2209 | 15.9372 | -1.4939 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.5844 | -0.103 | 0.003685 | 1.7923 | 2.268 | 3.6365 | 1.3528
2012-10-30 21:14 | 28.5686 | 16.3396 | 16.7829 | 18.2043 | 15.9205 | -1.4864 | 0.006735 | 0.32663 | 0.0022 | 48.4974 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.5513 | -0.1025 | 0.003691 | 1.7883 | 2.2316 | 3.653 | 1.3692
2012-10-30 21:14 | 28.5731 | 16.336 | 16.8443 | 18.2088 | 15.925 | -1.5026 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.5925 | -0.1036 | 0.003688 | 1.7435 | 2.2518 | 3.6163 | 1.3325
2012-10-30 21:14 | 28.5806 | 16.3679 | 16.8234 | 18.2244 | 15.9121 | -1.4876 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.547 | -0.1026 | 0.003688 | 1.8209 | 2.2764 | 3.6774 | 1.3651
2012-10-30 21:14 | 28.5768 | 16.356 | 16.7992 | 18.2246 | 15.9124 | -1.4986 | 0.006729 | 0.32667 | 0.002198 | 48.547 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5718 | -0.1033 | 0.003688 | 1.7842 | 2.2274 | 3.6528 | 1.3406
2012-10-30 21:14 | 28.5723 | 16.3718 | 16.7825 | 18.1431 | 15.9201 | -1.4957 | 0.006719 | 0.32656 | 0.002194 | 48.602 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5632 | -0.1031 | 0.003683 | 1.8086 | 2.2193 | 3.5799 | 1.3569
2012-10-30 21:14 | 28.5783 | 16.3737 | 16.8333 | 18.1815 | 15.918 | -1.5084 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.6141 | -0.104 | 0.003692 | 1.7596 | 2.2192 | 3.5674 | 1.3039
2012-10-30 21:14 | 28.5686 | 16.3071 | 16.8032 | 18.1434 | 15.8838 | -1.4951 | 0.006732 | 0.3266 | 0.002199 | 48.5147 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5432 | -0.1031 | 0.003688 | 1.7639 | 2.26 | 3.6002 | 1.3406
2012-10-30 21:14 | 28.5753 | 16.33 | 16.7408 | 18.1663 | 15.919 | -1.498 | 0.006738 | 0.32663 | 0.002201 | 48.4758 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.554 | -0.1033 | 0.003693 | 1.776 | 2.1868 | 3.6123 | 1.365
2012-10-30 21:14 | 28.5693 | 16.2793 | 16.7877 | 18.1969 | 15.9171 | -1.4864 | 0.006722 | 0.32664 | 0.002196 | 48.5933 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.5928 | -0.1025 | 0.003687 | 1.6865 | 2.1949 | 3.6041 | 1.3243
2012-10-30 21:14 | 28.5731 | 16.3116 | 16.7833 | 18.1195 | 15.872 | -1.4824 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.564 | -0.1022 | 0.003689 | 1.7476 | 2.2193 | 3.5555 | 1.308
2012-10-30 21:14 | 28.5701 | 16.2923 | 16.8169 | 18.149 | 15.8609 | -1.5009 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5406 | -0.1035 | 0.00369 | 1.7517 | 2.2763 | 3.6084 | 1.3203
2012-10-30 21:14 | 28.5723 | 16.2742 | 16.7947 | 18.1147 | 15.8957 | -1.487 | 0.006729 | 0.32663 | 0.002198 | 48.5406 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.5346 | -0.1025 | 0.003689 | 1.7396 | 2.2601 | 3.5801 | 1.3611
2012-10-30 21:14 | 28.5738 | 16.2553 | 16.7393 | 18.1364 | 15.8972 | -1.5032 | 0.006751 | 0.32656 | 0.002205 | 48.3716 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5892 | -0.1036 | 0.003694 | 1.6661 | 2.1501 | 3.5472 | 1.308
2012-10-30 21:14 | 28.5761 | 16.2861 | 16.7741 | 18.1671 | 15.9117 | -1.4783 | 0.006738 | 0.32659 | 0.002201 | 48.4693 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5507 | -0.1019 | 0.003691 | 1.7354 | 2.2234 | 3.6164 | 1.361
2012-10-30 21:14 | 28.5753 | 16.2894 | 16.7733 | 18.1379 | 15.9068 | -1.4928 | 0.006747 | 0.32667 | 0.002204 | 48.4175 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.5376 | -0.1029 | 0.003695 | 1.7518 | 2.2357 | 3.6003 | 1.3692
2012-10-30 21:14 | 28.5791 | 16.2891 | 16.7649 | 18.1052 | 15.935 | -1.4789 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5659 | -0.102 | 0.003692 | 1.7232 | 2.199 | 3.5393 | 1.3691
2012-10-30 21:14 | 28.5686 | 16.2338 | 16.7178 | 18.115 | 15.8838 | -1.4783 | 0.006747 | 0.32656 | 0.002203 | 48.4003 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.5554 | -0.1019 | 0.003691 | 1.6784 | 2.1624 | 3.5596 | 1.3284
2012-10-30 21:14 | 28.5686 | 16.2379 | 16.7626 | 18.0907 | 15.8431 | -1.4795 | 0.006741 | 0.32667 | 0.002202 | 48.4606 | 0.006949 | 0.2147 | 0.001492 | 30.8965 | 14.5187 | -0.102 | 0.003694 | 1.7192 | 2.2439 | 3.572 | 1.3244
2012-10-30 21:14 | 28.5768 | 16.2136 | 16.7586 | 18.111 | 15.8554 | -1.4812 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4984 | -0.1021 | 0.003688 | 1.7152 | 2.2602 | 3.6126 | 1.357
2012-10-30 21:14 | 28.5753 | 16.2527 | 16.7652 | 18.1014 | 15.8173 | -1.4702 | 0.006744 | 0.32659 | 0.002202 | 48.4262 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5091 | -0.1014 | 0.00369 | 1.7436 | 2.2561 | 3.5923 | 1.3082
2012-10-30 21:15 | 28.5753 | 16.2243 | 16.7083 | 18.1095 | 15.8376 | -1.4847 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 14.5499 | -0.1024 | 0.00369 | 1.6744 | 2.1584 | 3.5596 | 1.2877
2012-10-30 21:15 | 28.5806 | 16.2336 | 16.742 | 18.123 | 15.8185 | -1.4824 | 0.006747 | 0.3266 | 0.002204 | 48.4068 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.5266 | -0.1022 | 0.003692 | 1.707 | 2.2154 | 3.5964 | 1.2919
2012-10-30 21:15 | 28.5783 | 16.2191 | 16.7519 | 18.1004 | 15.8284 | -1.4789 | 0.006754 | 0.32656 | 0.002206 | 48.3502 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.5366 | -0.102 | 0.003699 | 1.6825 | 2.2153 | 3.5638 | 1.2918
2012-10-30 21:15 | 28.5761 | 16.1966 | 16.7335 | 18.1063 | 15.8669 | -1.4789 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.5711 | -0.102 | 0.003688 | 1.6255 | 2.1624 | 3.5352 | 1.2958
2012-10-30 21:15 | 28.5708 | 16.1831 | 16.7322 | 18.0726 | 15.8697 | -1.4818 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5617 | -0.1022 | 0.003693 | 1.6214 | 2.1705 | 3.5109 | 1.308
2012-10-30 21:15 | 28.5813 | 16.2303 | 16.7143 | 18.0547 | 15.8477 | -1.4858 | 0.006747 | 0.32659 | 0.002203 | 48.4046 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5355 | -0.1024 | 0.003693 | 1.6948 | 2.1788 | 3.5192 | 1.3122
2012-10-30 21:15 | 28.5731 | 16.1976 | 16.7386 | 18.0668 | 15.8517 | -1.4679 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5558 | -0.1012 | 0.003687 | 1.6418 | 2.1828 | 3.511 | 1.2959
2012-10-30 21:15 | 28.5873 | 16.2241 | 16.7244 | 18.0445 | 15.9107 | -1.4772 | 0.006729 | 0.32667 | 0.002198 | 48.547 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.566 | -0.1018 | 0.003688 | 1.6581 | 2.1584 | 3.4785 | 1.3447
2012-10-30 21:15 | 28.5843 | 16.2292 | 16.701 | 18.0496 | 15.8588 | -1.4702 | 0.006744 | 0.32653 | 0.002202 | 48.4177 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.4977 | -0.1014 | 0.003692 | 1.7315 | 2.2033 | 3.5519 | 1.3611
2012-10-30 21:15 | 28.5836 | 16.2041 | 16.7085 | 18.0408 | 15.9029 | -1.4916 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006946 | 0.21482 | 0.001492 | 30.9264 | 14.5215 | -0.1028 | 0.003692 | 1.6826 | 2.187 | 3.5193 | 1.3814
2012-10-30 21:15 | 28.5723 | 16.1684 | 16.6931 | 18.0538 | 15.8468 | -1.4754 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5183 | -0.1017 | 0.003692 | 1.6501 | 2.1748 | 3.5355 | 1.3285
2012-10-30 21:15 | 28.5843 | 16.1966 | 16.6888 | 18.0415 | 15.8385 | -1.4766 | 0.006744 | 0.32659 | 0.002202 | 48.4262 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5793 | -0.1018 | 0.003692 | 1.6173 | 2.1095 | 3.4622 | 1.2592
2012-10-30 21:15 | 28.5798 | 16.1352 | 16.6884 | 18.0451 | 15.8503 | -1.4743 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.5544 | -0.1016 | 0.003685 | 1.5808 | 2.134 | 3.4907 | 1.2959
2012-10-30 21:15 | 28.5791 | 16.1711 | 16.704 | 18.0525 | 15.8577 | -1.487 | 0.006729 | 0.32663 | 0.002198 | 48.5406 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.5455 | -0.1025 | 0.00369 | 1.6256 | 2.1585 | 3.507 | 1.3122
2012-10-30 21:15 | 28.5753 | 16.1225 | 16.6636 | 18.0325 | 15.8458 | -1.4835 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.554 | -0.1023 | 0.003691 | 1.5685 | 2.1096 | 3.4785 | 1.2918
2012-10-30 21:15 | 28.5836 | 16.2 | 16.6922 | 18.0042 | 15.8256 | -1.4778 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.5704 | -0.1019 | 0.00369 | 1.6296 | 2.1218 | 3.4338 | 1.2552
2012-10-30 21:15 | 28.5843 | 16.1763 | 16.701 | 18.0293 | 15.8263 | -1.4801 | 0.006735 | 0.3266 | 0.0022 | 48.4931 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.6201 | -0.102 | 0.00369 | 1.5562 | 2.0809 | 3.4092 | 1.2062
2012-10-30 21:15 | 28.5768 | 16.1322 | 16.6529 | 17.9528 | 15.8391 | -1.487 | 0.006747 | 0.32653 | 0.002203 | 48.3962 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5677 | -0.1025 | 0.003692 | 1.5645 | 2.0852 | 3.3851 | 1.2714
2012-10-30 21:15 | 28.5768 | 16.1566 | 16.6732 | 18.0299 | 15.8228 | -1.4806 | 0.006744 | 0.3266 | 0.002203 | 48.4284 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5187 | -0.1021 | 0.003693 | 1.6379 | 2.1545 | 3.5112 | 1.3041
2012-10-30 21:15 | 28.5798 | 16.1718 | 16.6843 | 17.9761 | 15.8055 | -1.4835 | 0.006747 | 0.3266 | 0.002204 | 48.4068 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5258 | -0.1023 | 0.003693 | 1.646 | 2.1585 | 3.4503 | 1.2797
2012-10-30 21:15 | 28.5903 | 16.1497 | 16.6664 | 17.9866 | 15.8445 | -1.4899 | 0.006744 | 0.3266 | 0.002203 | 48.4284 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.52 | -0.1027 | 0.003694 | 1.6297 | 2.1464 | 3.4666 | 1.3245
2012-10-30 21:15 | 28.5851 | 16.1445 | 16.6368 | 17.9611 | 15.8433 | -1.4847 | 0.006722 | 0.32661 | 0.002196 | 48.589 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5189 | -0.1024 | 0.003685 | 1.6256 | 2.1179 | 3.4422 | 1.3244
2012-10-30 21:15 | 28.5836 | 16.1146 | 16.6353 | 17.9758 | 15.8337 | -1.4749 | 0.006732 | 0.32667 | 0.002199 | 48.5254 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5215 | -0.1017 | 0.003689 | 1.5931 | 2.1138 | 3.4543 | 1.3122
2012-10-30 21:15 | 28.5821 | 16.1131 | 16.6175 | 17.9987 | 15.8241 | -1.4737 | 0.006722 | 0.32666 | 0.002196 | 48.5955 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5322 | -0.1016 | 0.003686 | 1.5809 | 2.0853 | 3.4665 | 1.2919
2012-10-30 21:15 | 28.5693 | 16.1165 | 16.6291 | 17.9453 | 15.7746 | -1.4714 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.5276 | -0.1014 | 0.003692 | 1.5889 | 2.1015 | 3.4177 | 1.247
2012-10-30 21:16 | 28.5873 | 16.1386 | 16.6471 | 17.9958 | 15.7926 | -1.4789 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5007 | -0.102 | 0.003692 | 1.6379 | 2.1464 | 3.4951 | 1.2919
2012-10-30 21:16 | 28.5903 | 16.109 | 16.6379 | 17.9338 | 15.8282 | -1.4697 | 0.006741 | 0.32661 | 0.002202 | 48.452 | 0.006924 | 0.2147 | 0.001487 | 31.0081 | 14.4996 | -0.1013 | 0.003689 | 1.6094 | 2.1383 | 3.4342 | 1.3286
2012-10-30 21:16 | 28.5903 | 16.1253 | 16.6135 | 17.9744 | 15.816 | -1.4731 | 0.006735 | 0.3266 | 0.0022 | 48.4931 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.4996 | -0.1016 | 0.003688 | 1.6257 | 2.1139 | 3.4748 | 1.3164
2012-10-30 21:16 | 28.5836 | 16.1227 | 16.6434 | 17.915 | 15.8418 | -1.4726 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006921 | 0.21471 | 0.001486 | 31.0236 | 14.5296 | -0.1015 | 0.003684 | 1.5931 | 2.1138 | 3.3854 | 1.3122
2012-10-30 21:16 | 28.5881 | 16.0784 | 16.6479 | 17.9235 | 15.8056 | -1.472 | 0.006751 | 0.32663 | 0.002205 | 48.3825 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.53 | -0.1015 | 0.003697 | 1.5484 | 2.1179 | 3.3935 | 1.2756
2012-10-30 21:16 | 28.5888 | 16.0628 | 16.6283 | 17.981 | 15.8348 | -1.4818 | 0.006735 | 0.32661 | 0.0022 | 48.4952 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5348 | -0.1022 | 0.003689 | 1.528 | 2.0935 | 3.4462 | 1.3
2012-10-30 21:16 | 28.5873 | 16.0938 | 16.5983 | 17.9268 | 15.8455 | -1.4668 | 0.006747 | 0.32653 | 0.002203 | 48.3962 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.5333 | -0.1011 | 0.003691 | 1.5605 | 2.065 | 3.3935 | 1.3122
2012-10-30 21:16 | 28.5978 | 16.0881 | 16.6169 | 17.8886 | 15.7705 | -1.4668 | 0.006725 | 0.32664 | 0.002197 | 48.5716 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.5357 | -0.1011 | 0.003689 | 1.5524 | 2.0812 | 3.3529 | 1.2348
2012-10-30 21:16 | 28.5911 | 16.0773 | 16.6468 | 17.9103 | 15.8086 | -1.4783 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 14.5575 | -0.1019 | 0.00369 | 1.5198 | 2.0893 | 3.3528 | 1.2511
2012-10-30 21:16 | 28.5918 | 16.0983 | 16.615 | 17.9272 | 15.7768 | -1.4783 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006921 | 0.21483 | 0.001487 | 31.0403 | 14.5297 | -0.1019 | 0.003686 | 1.5686 | 2.0853 | 3.3975 | 1.2471
2012-10-30 21:16 | 28.5956 | 16.0533 | 16.6188 | 17.9148 | 15.7683 | -1.472 | 0.006725 | 0.32661 | 0.002196 | 48.5673 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4927 | -0.1015 | 0.003685 | 1.5606 | 2.1261 | 3.4221 | 1.2756
2012-10-30 21:16 | 28.5903 | 16.0236 | 16.642 | 17.9135 | 15.8119 | -1.4731 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5445 | -0.1016 | 0.003689 | 1.4791 | 2.0975 | 3.369 | 1.2674
2012-10-30 21:16 | 28.5896 | 16.0717 | 16.625 | 17.8479 | 15.799 | -1.4754 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5275 | -0.1017 | 0.003688 | 1.5442 | 2.0975 | 3.3204 | 1.2715
2012-10-30 21:16 | 28.6001 | 16.0822 | 16.6152 | 17.9112 | 15.8095 | -1.4731 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5584 | -0.1016 | 0.003691 | 1.5238 | 2.0568 | 3.3528 | 1.2511
2012-10-30 21:16 | 28.5918 | 16.0291 | 16.6313 | 17.915 | 15.7686 | -1.472 | 0.006729 | 0.32659 | 0.002198 | 48.5341 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5215 | -0.1015 | 0.003688 | 1.5076 | 2.1098 | 3.3935 | 1.2471
2012-10-30 21:16 | 28.5956 | 16.0492 | 16.6147 | 17.858 | 15.7724 | -1.4858 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.4968 | -0.1024 | 0.003688 | 1.5524 | 2.1179 | 3.3612 | 1.2756
2012-10-30 21:16 | 28.5926 | 16.0666 | 16.5995 | 17.8955 | 15.7816 | -1.4616 | 0.006744 | 0.32659 | 0.002202 | 48.4262 | 0.006917 | 0.21474 | 0.001485 | 31.0457 | 14.5264 | -0.1008 | 0.003687 | 1.5402 | 2.0731 | 3.3691 | 1.2552
2012-10-30 21:16 | 28.5948 | 16.0566 | 16.6017 | 17.849 | 15.7879 | -1.4778 | 0.006732 | 0.32664 | 0.002199 | 48.5211 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4633 | -0.1019 | 0.003688 | 1.5933 | 2.1384 | 3.3857 | 1.3246
2012-10-30 21:16 | 28.5918 | 16.0332 | 16.5947 | 17.8582 | 15.7808 | -1.4691 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006917 | 0.21473 | 0.001485 | 31.0437 | 14.5011 | -0.1013 | 0.003682 | 1.5321 | 2.0936 | 3.3571 | 1.2797
2012-10-30 21:16 | 28.5948 | 16.0118 | 16.5936 | 17.8572 | 15.7716 | -1.4824 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.5286 | -0.1022 | 0.003684 | 1.4832 | 2.065 | 3.3286 | 1.243
2012-10-30 21:16 | 28.6008 | 16.0259 | 16.5874 | 17.8307 | 15.7572 | -1.4754 | 0.006722 | 0.32656 | 0.002195 | 48.5803 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5387 | -0.1017 | 0.003686 | 1.4872 | 2.0487 | 3.292 | 1.2185
2012-10-30 21:16 | 28.5986 | 16.0604 | 16.5445 | 17.8447 | 15.755 | -1.4772 | 0.006729 | 0.32663 | 0.002198 | 48.5406 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4794 | -0.1018 | 0.003687 | 1.581 | 2.0651 | 3.3653 | 1.2756
2012-10-30 21:16 | 28.5978 | 16.0311 | 16.6007 | 17.7668 | 15.7746 | -1.4541 | 0.006741 | 0.3266 | 0.002202 | 48.4499 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5234 | -0.1002 | 0.003691 | 1.5077 | 2.0773 | 3.2434 | 1.2512
2012-10-30 21:16 | 28.5873 | 16.0491 | 16.582 | 17.8294 | 15.7519 | -1.4841 | 0.006741 | 0.32663 | 0.002202 | 48.4542 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4844 | -0.1023 | 0.00369 | 1.5647 | 2.0976 | 3.345 | 1.2675
2012-10-30 21:16 | 28.5933 | 15.994 | 16.5718 | 17.811 | 15.7823 | -1.4708 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006946 | 0.21482 | 0.001492 | 30.9264 | 14.4863 | -0.1014 | 0.003693 | 1.5077 | 2.0855 | 3.3247 | 1.296
2012-10-30 21:16 | 28.5956 | 16.0004 | 16.5293 | 17.8255 | 15.7235 | -1.4668 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006946 | 0.21468 | 0.001491 | 30.9077 | 14.4968 | -0.1011 | 0.003689 | 1.5036 | 2.0325 | 3.3287 | 1.2267
2012-10-30 21:16 | 28.5971 | 15.9978 | 16.5878 | 17.8432 | 15.7861 | -1.4685 | 0.006722 | 0.32657 | 0.002195 | 48.5826 | 0.006936 | 0.21466 | 0.001489 | 30.9481 | 14.5187 | -0.1012 | 0.003684 | 1.4791 | 2.0691 | 3.3245 | 1.2674
2012-10-30 21:16 | 28.6031 | 16.0242 | 16.5612 | 17.8046 | 15.7921 | -1.465 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.5165 | -0.101 | 0.003689 | 1.5077 | 2.0447 | 3.2881 | 1.2756
2012-10-30 21:17 | 28.5963 | 16.0174 | 16.5666 | 17.8059 | 15.7405 | -1.4749 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5423 | -0.1017 | 0.00369 | 1.4751 | 2.0243 | 3.2636 | 1.1982
2012-10-30 21:17 | 28.6001 | 16.0415 | 16.5826 | 17.7934 | 15.7728 | -1.4685 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5094 | -0.1012 | 0.00369 | 1.5321 | 2.0732 | 3.284 | 1.2634
2012-10-30 21:17 | 28.5978 | 15.9822 | 16.5478 | 17.8439 | 15.7665 | -1.4674 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5194 | -0.1012 | 0.003687 | 1.4628 | 2.0284 | 3.3245 | 1.2471
2012-10-30 21:17 | 28.5926 | 16.0015 | 16.5711 | 17.7819 | 15.749 | -1.4662 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.5468 | -0.1011 | 0.003689 | 1.4547 | 2.0243 | 3.2351 | 1.2022
2012-10-30 21:17 | 28.5926 | 15.9974 | 16.5914 | 17.8103 | 15.745 | -1.472 | 0.006754 | 0.32656 | 0.002206 | 48.3502 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.5345 | -0.1015 | 0.003696 | 1.4629 | 2.0569 | 3.2758 | 1.2105
2012-10-30 21:17 | 28.5911 | 16.0325 | 16.5208 | 17.7479 | 15.7435 | -1.4656 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.4963 | -0.101 | 0.003685 | 1.5362 | 2.0245 | 3.2516 | 1.2472
2012-10-30 21:17 | 28.5851 | 15.9858 | 16.5839 | 17.7947 | 15.7334 | -1.4743 | 0.006716 | 0.3265 | 0.002193 | 48.6152 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4944 | -0.1016 | 0.003685 | 1.4914 | 2.0895 | 3.3003 | 1.239
2012-10-30 21:17 | 28.5956 | 15.98 | 16.5659 | 17.8092 | 15.7398 | -1.465 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5294 | -0.101 | 0.003688 | 1.4506 | 2.0365 | 3.2798 | 1.2104
2012-10-30 21:17 | 28.5933 | 15.9818 | 16.527 | 17.8151 | 15.7335 | -1.4702 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5393 | -0.1014 | 0.003686 | 1.4425 | 1.9877 | 3.2758 | 1.1942
2012-10-30 21:17 | 28.5843 | 15.9809 | 16.5506 | 17.7573 | 15.7407 | -1.4668 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.5426 | -0.1011 | 0.003685 | 1.4383 | 2.008 | 3.2147 | 1.1981
2012-10-30 21:17 | 28.5948 | 15.9833 | 16.5489 | 17.7638 | 15.739 | -1.4708 | 0.00671 | 0.32656 | 0.002191 | 48.6672 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.549 | -0.1014 | 0.003682 | 1.4343 | 1.9999 | 3.2148 | 1.19
2012-10-30 21:17 | 28.5828 | 15.9184 | 16.545 | 17.7761 | 15.7352 | -1.4957 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5411 | -0.1031 | 0.003687 | 1.3773 | 2.0039 | 3.235 | 1.1941
2012-10-30 21:17 | 28.5873 | 15.931 | 16.5088 | 17.74 | 15.7315 | -1.4749 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.57 | -0.1017 | 0.003687 | 1.361 | 1.9388 | 3.17 | 1.1615
2012-10-30 21:17 | 28.5926 | 15.9608 | 16.5345 | 17.7453 | 15.7368 | -1.4743 | 0.006738 | 0.32657 | 0.0022 | 48.4672 | 0.006921 | 0.21477 | 0.001486 | 31.0319 | 14.5345 | -0.1016 | 0.003686 | 1.4263 | 2 | 3.2108 | 1.2023
2012-10-30 21:17 | 28.5843 | 15.9199 | 16.5546 | 17.802 | 15.7163 | -1.465 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.4895 | -0.101 | 0.00369 | 1.4304 | 2.0651 | 3.3125 | 1.2268
2012-10-30 21:17 | 28.5888 | 15.9325 | 16.5103 | 17.7375 | 15.7452 | -1.4743 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.4981 | -0.1016 | 0.003688 | 1.4344 | 2.0122 | 3.2394 | 1.2471
2012-10-30 21:17 | 28.5866 | 15.9303 | 16.5366 | 17.7272 | 15.7308 | -1.4743 | 0.006729 | 0.32664 | 0.002198 | 48.5427 | 0.006961 | 0.21467 | 0.001494 | 30.8391 | 14.4878 | -0.1016 | 0.003692 | 1.4425 | 2.0488 | 3.2394 | 1.243
2012-10-30 21:17 | 28.5896 | 15.9048 | 16.4827 | 17.7383 | 15.7135 | -1.4806 | 0.006738 | 0.32657 | 0.0022 | 48.4672 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5479 | -0.1021 | 0.00369 | 1.3569 | 1.9348 | 3.1904 | 1.1656
2012-10-30 21:17 | 28.5836 | 15.9314 | 16.4889 | 17.7079 | 15.7238 | -1.4726 | 0.006738 | 0.32661 | 0.002201 | 48.4736 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.497 | -0.1015 | 0.00369 | 1.4344 | 1.9919 | 3.2109 | 1.2268
2012-10-30 21:17 | 28.5933 | 15.9207 | 16.4741 | 17.7298 | 15.7742 | -1.4789 | 0.006735 | 0.32663 | 0.0022 | 48.4974 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5271 | -0.102 | 0.00369 | 1.3936 | 1.947 | 3.2027 | 1.2471
2012-10-30 21:17 | 28.5896 | 15.9862 | 16.4908 | 17.6977 | 15.7501 | -1.4812 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4867 | -0.1021 | 0.003685 | 1.4995 | 2.0041 | 3.211 | 1.2634
2012-10-30 21:17 | 28.5888 | 15.8715 | 16.5144 | 17.7375 | 15.7575 | -1.4789 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5185 | -0.102 | 0.003689 | 1.353 | 1.9959 | 3.219 | 1.239
2012-10-30 21:17 | 28.5918 | 15.9437 | 16.4971 | 17.7161 | 15.679 | -1.4749 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006946 | 0.21467 | 0.001491 | 30.9057 | 14.5052 | -0.1017 | 0.003693 | 1.4385 | 1.9919 | 3.2109 | 1.1738
2012-10-30 21:17 | 28.5956 | 15.923 | 16.4805 | 17.6996 | 15.7032 | -1.483 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006949 | 0.21467 | 0.001492 | 30.8924 | 14.4764 | -0.1022 | 0.00369 | 1.4466 | 2.0041 | 3.2232 | 1.2268
2012-10-30 21:17 | 28.5978 | 15.9008 | 16.5071 | 17.714 | 15.7135 | -1.4662 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4663 | -0.1011 | 0.003684 | 1.4345 | 2.0408 | 3.2477 | 1.2472
2012-10-30 21:17 | 28.5836 | 15.8785 | 16.4726 | 17.6755 | 15.6586 | -1.4702 | 0.006719 | 0.32659 | 0.002194 | 48.6063 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.4848 | -0.1014 | 0.003684 | 1.3937 | 1.9878 | 3.1907 | 1.1738
2012-10-30 21:17 | 28.5956 | 15.8783 | 16.4724 | 17.6915 | 15.7235 | -1.4835 | 0.006713 | 0.32654 | 0.002192 | 48.6434 | 0.006908 | 0.21467 | 0.001483 | 31.0757 | 14.4682 | -0.1023 | 0.003675 | 1.4101 | 2.0042 | 3.2233 | 1.2553
2012-10-30 21:18 | 28.5903 | 15.9055 | 16.463 | 17.67 | 15.7549 | -1.4726 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.4792 | -0.1015 | 0.003692 | 1.4263 | 1.9838 | 3.1908 | 1.2757
2012-10-30 21:18 | 28.5903 | 15.9137 | 16.4467 | 17.6659 | 15.6816 | -1.483 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.4833 | -0.1022 | 0.003691 | 1.4304 | 1.9634 | 3.1826 | 1.1983
2012-10-30 21:18 | 28.5933 | 15.9207 | 16.4701 | 17.6933 | 15.6887 | -1.4697 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.4781 | -0.1013 | 0.003687 | 1.4426 | 1.992 | 3.2152 | 1.2106
2012-10-30 21:18 | 28.5836 | 15.8866 | 16.4482 | 17.6511 | 15.6912 | -1.4841 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006949 | 0.21471 | 0.001492 | 30.8985 | 14.497 | -0.1023 | 0.003687 | 1.3896 | 1.9512 | 3.1541 | 1.1942
2012-10-30 21:18 | 28.6008 | 15.9201 | 16.4288 | 17.6724 | 15.688 | -1.4697 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.4856 | -0.1013 | 0.003686 | 1.4345 | 1.9432 | 3.1868 | 1.2024
2012-10-30 21:18 | 28.6023 | 15.9175 | 16.4099 | 17.6779 | 15.6936 | -1.4789 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.4871 | -0.102 | 0.003686 | 1.4304 | 1.9228 | 3.1908 | 1.2065
2012-10-30 21:18 | 28.5963 | 15.9034 | 16.469 | 17.6638 | 15.7161 | -1.4737 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4934 | -0.1016 | 0.003688 | 1.41 | 1.9756 | 3.1704 | 1.2227
2012-10-30 21:18 | 28.5881 | 15.9115 | 16.473 | 17.6353 | 15.6835 | -1.4691 | 0.006722 | 0.32657 | 0.002195 | 48.5826 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.5015 | -0.1013 | 0.003683 | 1.41 | 1.9715 | 3.1338 | 1.182
2012-10-30 21:18 | 28.5896 | 15.8926 | 16.4338 | 17.6571 | 15.7216 | -1.4939 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4581 | -0.103 | 0.003686 | 1.4345 | 1.9757 | 3.199 | 1.2635
2012-10-30 21:18 | 28.5858 | 15.8563 | 16.3934 | 17.6249 | 15.6893 | -1.476 | 0.006719 | 0.3265 | 0.002194 | 48.5935 | 0.006933 | 0.21461 | 0.001488 | 30.9553 | 14.4298 | -0.1018 | 0.003682 | 1.4265 | 1.9636 | 3.1951 | 1.2595
2012-10-30 21:18 | 28.5896 | 15.8845 | 16.4094 | 17.653 | 15.6361 | -1.4714 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.4377 | -0.1014 | 0.003686 | 1.4468 | 1.9717 | 3.2153 | 1.1984
2012-10-30 21:18 | 28.5896 | 15.8682 | 16.4338 | 17.6774 | 15.6646 | -1.4899 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.4092 | -0.1027 | 0.003688 | 1.459 | 2.0246 | 3.2682 | 1.2554
2012-10-30 21:18 | 28.6001 | 15.8868 | 16.4159 | 17.6554 | 15.6873 | -1.4668 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4237 | -0.1011 | 0.003684 | 1.4631 | 1.9922 | 3.2317 | 1.2636
2012-10-30 21:18 | 28.6031 | 15.8817 | 16.4433 | 17.6584 | 15.7066 | -1.4934 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006933 | 0.21467 | 0.001488 | 30.9637 | 14.4023 | -0.103 | 0.003685 | 1.4794 | 2.041 | 3.2561 | 1.3043
2012-10-30 21:18 | 28.5986 | 15.8568 | 16.39 | 17.5971 | 15.6288 | -1.4714 | 0.006747 | 0.32643 | 0.002202 | 48.3811 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.4794 | -0.1014 | 0.00369 | 1.3774 | 1.9106 | 3.1177 | 1.1494
2012-10-30 21:18 | 28.5978 | 15.8601 | 16.4014 | 17.6328 | 15.6565 | -1.476 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006943 | 0.21463 | 0.00149 | 30.9129 | 14.4867 | -0.1018 | 0.003688 | 1.3734 | 1.9147 | 3.1461 | 1.1698
2012-10-30 21:18 | 28.5888 | 15.8348 | 16.3883 | 17.5954 | 15.6679 | -1.4841 | 0.006719 | 0.32653 | 0.002194 | 48.5979 | 0.006936 | 0.21458 | 0.001488 | 30.9377 | 14.4573 | -0.1023 | 0.003682 | 1.3775 | 1.931 | 3.1381 | 1.2106
2012-10-30 21:18 | 28.6008 | 15.855 | 16.4125 | 17.5993 | 15.684 | -1.4743 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.502 | -0.1016 | 0.003686 | 1.353 | 1.9105 | 3.0973 | 1.182
2012-10-30 21:18 | 28.5948 | 15.8327 | 16.378 | 17.573 | 15.6372 | -1.4876 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006921 | 0.2147 | 0.001486 | 31.0215 | 14.4878 | -0.1026 | 0.003688 | 1.3449 | 1.8902 | 3.0852 | 1.1494
2012-10-30 21:18 | 28.5941 | 15.8605 | 16.3733 | 17.6129 | 15.6773 | -1.4841 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.4749 | -0.1023 | 0.003687 | 1.3856 | 1.8984 | 3.138 | 1.2024
2012-10-30 21:18 | 28.6038 | 15.8254 | 16.4114 | 17.582 | 15.6666 | -1.4749 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5009 | -0.1017 | 0.003689 | 1.3245 | 1.9105 | 3.0811 | 1.1657
2012-10-30 21:18 | 28.5993 | 15.8168 | 16.3866 | 17.5531 | 15.658 | -1.4772 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.476 | -0.1018 | 0.00369 | 1.3408 | 1.9106 | 3.0771 | 1.182
2012-10-30 21:18 | 28.6031 | 15.7921 | 16.4067 | 17.5853 | 15.6618 | -1.4818 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5206 | -0.1022 | 0.003688 | 1.2715 | 1.8861 | 3.0647 | 1.1412
2012-10-30 21:18 | 28.6068 | 15.8162 | 16.4226 | 17.5647 | 15.6859 | -1.4749 | 0.006738 | 0.32659 | 0.002201 | 48.4693 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.4998 | -0.1017 | 0.003688 | 1.3164 | 1.9228 | 3.0649 | 1.1861
2012-10-30 21:18 | 28.6038 | 15.748 | 16.3992 | 17.5698 | 15.6544 | -1.4795 | 0.006716 | 0.32653 | 0.002193 | 48.6196 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5131 | -0.102 | 0.003682 | 1.2349 | 1.8861 | 3.0567 | 1.1413
2012-10-30 21:18 | 28.6038 | 15.748 | 16.4277 | 17.5495 | 15.6299 | -1.4801 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4846 | -0.102 | 0.003685 | 1.2634 | 1.9431 | 3.0649 | 1.1453
2012-10-30 21:18 | 28.6031 | 15.7636 | 16.3985 | 17.561 | 15.6659 | -1.4772 | 0.006722 | 0.32653 | 0.002195 | 48.5762 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.492 | -0.1018 | 0.003684 | 1.2716 | 1.9065 | 3.069 | 1.1739
2012-10-30 21:18 | 28.6001 | 15.7891 | 16.3874 | 17.5214 | 15.667 | -1.4853 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4931 | -0.1024 | 0.003685 | 1.296 | 1.8943 | 3.0283 | 1.1739
2012-10-30 21:19 | 28.6031 | 15.7677 | 16.3904 | 17.565 | 15.6292 | -1.4818 | 0.006729 | 0.32663 | 0.002198 | 48.5406 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4675 | -0.1022 | 0.003687 | 1.3002 | 1.9229 | 3.0975 | 1.1617
2012-10-30 21:19 | 28.6016 | 15.7866 | 16.397 | 17.5351 | 15.6359 | -1.4847 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.4864 | -0.1024 | 0.003689 | 1.3002 | 1.9106 | 3.0487 | 1.1495
2012-10-30 21:19 | 28.6008 | 15.802 | 16.3962 | 17.5587 | 15.6514 | -1.472 | 0.006716 | 0.3266 | 0.002193 | 48.6303 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.4734 | -0.1015 | 0.003684 | 1.3286 | 1.9228 | 3.0853 | 1.178
2012-10-30 21:19 | 28.5956 | 15.748 | 16.387 | 17.5372 | 15.5932 | -1.4841 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.5212 | -0.1023 | 0.003687 | 1.2268 | 1.8658 | 3.016 | 1.072
2012-10-30 21:19 | 28.5948 | 15.735 | 16.3658 | 17.5283 | 15.6006 | -1.4749 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.4796 | -0.1017 | 0.003686 | 1.2554 | 1.8862 | 3.0487 | 1.121
2012-10-30 21:19 | 28.5926 | 15.692 | 16.3962 | 17.522 | 15.6147 | -1.4731 | 0.006722 | 0.3266 | 0.002195 | 48.5869 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4734 | -0.1016 | 0.003685 | 1.2186 | 1.9228 | 3.0486 | 1.1413
2012-10-30 21:19 | 28.5948 | 15.7594 | 16.3984 | 17.5648 | 15.5965 | -1.4824 | 0.006719 | 0.3265 | 0.002194 | 48.5935 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.4796 | -0.1022 | 0.003686 | 1.2798 | 1.9188 | 3.0852 | 1.1169
2012-10-30 21:19 | 28.5918 | 15.736 | 16.4035 | 17.5375 | 15.6587 | -1.4795 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006943 | 0.21468 | 0.001491 | 30.9211 | 14.4766 | -0.102 | 0.003687 | 1.2594 | 1.9269 | 3.0609 | 1.1821
2012-10-30 21:19 | 28.5963 | 15.7853 | 16.3795 | 17.5176 | 15.655 | -1.4766 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5097 | -0.1018 | 0.003689 | 1.2756 | 1.8698 | 3.0079 | 1.1453
2012-10-30 21:19 | 28.6031 | 15.7025 | 16.3904 | 17.4838 | 15.6577 | -1.4887 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.4879 | -0.1026 | 0.003687 | 1.2146 | 1.9025 | 2.9959 | 1.1698
2012-10-30 21:19 | 28.5926 | 15.7409 | 16.388 | 17.4895 | 15.635 | -1.4806 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5019 | -0.1021 | 0.003693 | 1.239 | 1.8861 | 2.9876 | 1.1331
2012-10-30 21:19 | 28.6023 | 15.7262 | 16.3774 | 17.4952 | 15.6732 | -1.4853 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 14.5483 | -0.1024 | 0.003686 | 1.1779 | 1.8291 | 2.9469 | 1.1249
2012-10-30 21:19 | 28.5993 | 15.7639 | 16.3581 | 17.4962 | 15.6214 | -1.4766 | 0.006732 | 0.32666 | 0.002199 | 48.5233 | 0.006946 | 0.21477 | 0.001492 | 30.9202 | 14.5168 | -0.1018 | 0.003691 | 1.2471 | 1.8413 | 2.9794 | 1.1046
2012-10-30 21:19 | 28.5933 | 15.7579 | 16.3277 | 17.5024 | 15.6235 | -1.4801 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5149 | -0.102 | 0.003694 | 1.243 | 1.8128 | 2.9875 | 1.1086
2012-10-30 21:19 | 28.6008 | 15.7165 | 16.3352 | 17.4734 | 15.6269 | -1.4934 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4979 | -0.103 | 0.003686 | 1.2186 | 1.8373 | 2.9755 | 1.129
2012-10-30 21:19 | 28.6023 | 15.7343 | 16.3652 | 17.5195 | 15.6203 | -1.4772 | 0.006713 | 0.3265 | 0.002192 | 48.637 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5442 | -0.1018 | 0.003681 | 1.1901 | 1.821 | 2.9753 | 1.0761
2012-10-30 21:19 | 28.5971 | 15.6965 | 16.3518 | 17.5103 | 15.6069 | -1.4864 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5105 | -0.1025 | 0.003687 | 1.186 | 1.8413 | 2.9998 | 1.0964
2012-10-30 21:19 | 28.5963 | 15.6672 | 16.347 | 17.477 | 15.5613 | -1.4754 | 0.006729 | 0.32659 | 0.002198 | 48.5341 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.5138 | -0.1017 | 0.003687 | 1.1534 | 1.8332 | 2.9632 | 1.0475
2012-10-30 21:19 | 28.5956 | 15.6991 | 16.387 | 17.5088 | 15.6177 | -1.4812 | 0.006719 | 0.3265 | 0.002194 | 48.5935 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5416 | -0.1021 | 0.003685 | 1.1575 | 1.8454 | 2.9672 | 1.0761
2012-10-30 21:19 | 28.6046 | 15.7651 | 16.3634 | 17.4731 | 15.647 | -1.4754 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.5465 | -0.1017 | 0.003685 | 1.2186 | 1.8169 | 2.9266 | 1.1005
2012-10-30 21:19 | 28.6008 | 15.7043 | 16.3555 | 17.4653 | 15.5984 | -1.4818 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.502 | -0.1022 | 0.003691 | 1.2023 | 1.8535 | 2.9633 | 1.0964
2012-10-30 21:19 | 28.6001 | 15.7036 | 16.3345 | 17.4483 | 15.5896 | -1.4882 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.538 | -0.1026 | 0.003693 | 1.1656 | 1.7965 | 2.9103 | 1.0516
2012-10-30 21:19 | 28.6053 | 15.7251 | 16.3397 | 17.4738 | 15.5826 | -1.4899 | 0.006732 | 0.3266 | 0.002199 | 48.5147 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5065 | -0.1027 | 0.003688 | 1.2186 | 1.8332 | 2.9673 | 1.0761
2012-10-30 21:19 | 28.6083 | 15.7159 | 16.3386 | 17.4484 | 15.5815 | -1.487 | 0.006751 | 0.32659 | 0.002205 | 48.3759 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5054 | -0.1025 | 0.003694 | 1.2105 | 1.8332 | 2.943 | 1.0761
2012-10-30 21:19 | 28.6098 | 15.697 | 16.3361 | 17.5027 | 15.6278 | -1.4806 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.5477 | -0.1021 | 0.00369 | 1.1493 | 1.7884 | 2.955 | 1.0801
2012-10-30 21:19 | 28.609 | 15.7003 | 16.3556 | 17.4531 | 15.5985 | -1.487 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.502 | -0.1025 | 0.003689 | 1.1983 | 1.8536 | 2.9511 | 1.0965
2012-10-30 21:19 | 28.6031 | 15.6985 | 16.3497 | 17.4879 | 15.5559 | -1.4824 | 0.006722 | 0.3266 | 0.002195 | 48.5869 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5206 | -0.1022 | 0.003685 | 1.1779 | 1.8291 | 2.9673 | 1.0353
2012-10-30 21:20 | 28.6038 | 15.7032 | 16.3667 | 17.4398 | 15.5892 | -1.4847 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.509 | -0.1024 | 0.003689 | 1.1942 | 1.8577 | 2.9308 | 1.0802
2012-10-30 21:20 | 28.5978 | 15.6606 | 16.3037 | 17.4623 | 15.5628 | -1.4737 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5316 | -0.1016 | 0.003686 | 1.129 | 1.7721 | 2.9307 | 1.0312
2012-10-30 21:20 | 28.6046 | 15.6552 | 16.3675 | 17.4366 | 15.5941 | -1.4737 | 0.006738 | 0.32661 | 0.002201 | 48.4736 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5384 | -0.1016 | 0.003689 | 1.1168 | 1.8291 | 2.8982 | 1.0557
2012-10-30 21:20 | 28.6053 | 15.6559 | 16.3478 | 17.4251 | 15.5866 | -1.476 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5472 | -0.1018 | 0.003692 | 1.1087 | 1.8006 | 2.8779 | 1.0394
2012-10-30 21:20 | 28.6098 | 15.6644 | 16.3117 | 17.4418 | 15.6237 | -1.4783 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.5232 | -0.1019 | 0.00369 | 1.1412 | 1.7885 | 2.9186 | 1.1005
2012-10-30 21:20 | 28.6023 | 15.6244 | 16.3408 | 17.3977 | 15.5836 | -1.4737 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006952 | 0.2147 | 0.001493 | 30.8832 | 14.5116 | -0.1016 | 0.003688 | 1.1128 | 1.8292 | 2.8861 | 1.072
2012-10-30 21:20 | 28.612 | 15.6463 | 16.3464 | 17.3912 | 15.6056 | -1.4806 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5539 | -0.1021 | 0.003687 | 1.0924 | 1.7925 | 2.8373 | 1.0517
2012-10-30 21:20 | 28.6083 | 15.6874 | 16.3549 | 17.3915 | 15.5856 | -1.4916 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006958 | 0.2147 | 0.001494 | 30.8566 | 14.5054 | -0.1028 | 0.003692 | 1.182 | 1.8495 | 2.8861 | 1.0802
2012-10-30 21:20 | 28.609 | 15.6881 | 16.3231 | 17.4044 | 15.5618 | -1.4754 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5306 | -0.1017 | 0.003684 | 1.1575 | 1.7925 | 2.8738 | 1.0312
2012-10-30 21:20 | 28.6083 | 15.6385 | 16.3142 | 17.3875 | 15.6019 | -1.4754 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006939 | 0.21464 | 0.001489 | 30.9327 | 14.5135 | -0.1017 | 0.003688 | 1.125 | 1.8007 | 2.874 | 1.0884
2012-10-30 21:20 | 28.609 | 15.6474 | 16.3271 | 17.4044 | 15.5374 | -1.4806 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.4449 | -0.1021 | 0.003688 | 1.2025 | 1.8822 | 2.9595 | 1.0925
2012-10-30 21:20 | 28.6113 | 15.6619 | 16.2969 | 17.3498 | 15.6008 | -1.465 | 0.006732 | 0.32661 | 0.002199 | 48.5168 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5492 | -0.101 | 0.003689 | 1.1127 | 1.7477 | 2.8006 | 1.0516
2012-10-30 21:20 | 28.6016 | 15.6033 | 16.2994 | 17.3848 | 15.5666 | -1.472 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.515 | -0.1015 | 0.003686 | 1.0883 | 1.7844 | 2.8698 | 1.0516
2012-10-30 21:20 | 28.6143 | 15.6323 | 16.3243 | 17.3406 | 15.5793 | -1.4818 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5032 | -0.1022 | 0.003685 | 1.1291 | 1.8211 | 2.8374 | 1.0761
2012-10-30 21:20 | 28.6128 | 15.6064 | 16.2821 | 17.4366 | 15.5534 | -1.4812 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.5466 | -0.1021 | 0.00369 | 1.0598 | 1.7355 | 2.89 | 1.0068
2012-10-30 21:20 | 28.6083 | 15.5693 | 16.3061 | 17.3387 | 15.5978 | -1.4743 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5013 | -0.1016 | 0.003691 | 1.068 | 1.8048 | 2.8374 | 1.0965
2012-10-30 21:20 | 28.6075 | 15.597 | 16.3094 | 17.4029 | 15.5888 | -1.4853 | 0.006741 | 0.3266 | 0.002202 | 48.4499 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5291 | -0.1024 | 0.003691 | 1.0679 | 1.7803 | 2.8738 | 1.0597
2012-10-30 21:20 | 28.618 | 15.6238 | 16.271 | 17.3809 | 15.5871 | -1.487 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5151 | -0.1025 | 0.003687 | 1.1087 | 1.7559 | 2.8658 | 1.072
2012-10-30 21:20 | 28.615 | 15.5963 | 16.2762 | 17.4023 | 15.5963 | -1.4789 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.5325 | -0.102 | 0.00368 | 1.0638 | 1.7437 | 2.8698 | 1.0638
2012-10-30 21:20 | 28.6083 | 15.6059 | 16.2817 | 17.3834 | 15.5285 | -1.4778 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.538 | -0.1019 | 0.003683 | 1.0679 | 1.7437 | 2.8454 | 0.9905
2012-10-30 21:20 | 28.6113 | 15.6171 | 16.301 | 17.3539 | 15.5967 | -1.4754 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5614 | -0.1017 | 0.003685 | 1.0557 | 1.7396 | 2.7925 | 1.0353
2012-10-30 21:20 | 28.6046 | 15.6185 | 16.3105 | 17.3675 | 15.5167 | -1.4887 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5343 | -0.1026 | 0.003687 | 1.0842 | 1.7762 | 2.8332 | 0.9824
2012-10-30 21:20 | 28.6143 | 15.6201 | 16.2795 | 17.3406 | 15.6079 | -1.4847 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5562 | -0.1024 | 0.003686 | 1.0639 | 1.7233 | 2.7844 | 1.0517
2012-10-30 21:20 | 28.6053 | 15.55 | 16.2624 | 17.3316 | 15.5541 | -1.4841 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006933 | 0.21466 | 0.001488 | 30.9615 | 14.5636 | -0.1023 | 0.003687 | 0.9864 | 1.6988 | 2.768 | 0.9905
2012-10-30 21:20 | 28.6113 | 15.6049 | 16.3132 | 17.3376 | 15.5682 | -1.4783 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.5125 | -0.1019 | 0.003689 | 1.0924 | 1.8007 | 2.8251 | 1.0557
2012-10-30 21:20 | 28.6113 | 15.5886 | 16.3294 | 17.3661 | 15.5601 | -1.4835 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006917 | 0.21473 | 0.001485 | 31.0437 | 14.4717 | -0.1023 | 0.003684 | 1.1169 | 1.8577 | 2.8944 | 1.0884
2012-10-30 21:20 | 28.6046 | 15.6022 | 16.282 | 17.3756 | 15.59 | -1.4841 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5221 | -0.1023 | 0.00369 | 1.0801 | 1.7599 | 2.8535 | 1.0679
2012-10-30 21:20 | 28.6023 | 15.5755 | 16.2838 | 17.3286 | 15.5592 | -1.4726 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.5116 | -0.1015 | 0.003686 | 1.0639 | 1.7722 | 2.817 | 1.0476
2012-10-30 21:21 | 28.6038 | 15.5933 | 16.265 | 17.3464 | 15.5729 | -1.4991 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.4968 | -0.1033 | 0.003687 | 1.0965 | 1.7682 | 2.8496 | 1.0761
2012-10-30 21:21 | 28.606 | 15.5507 | 16.255 | 17.3242 | 15.5751 | -1.4812 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006921 | 0.21474 | 0.001486 | 31.0277 | 14.5112 | -0.1021 | 0.003684 | 1.0395 | 1.7438 | 2.813 | 1.0639
2012-10-30 21:21 | 28.6046 | 15.5615 | 16.2617 | 17.2903 | 15.5452 | -1.4668 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.465 | -0.1011 | 0.003685 | 1.0965 | 1.7967 | 2.8253 | 1.0802
2012-10-30 21:21 | 28.6046 | 15.6063 | 16.2576 | 17.3188 | 15.5859 | -1.4778 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5058 | -0.1019 | 0.003685 | 1.1005 | 1.7518 | 2.813 | 1.0801
2012-10-30 21:21 | 28.6053 | 15.5541 | 16.2787 | 17.3398 | 15.5663 | -1.4812 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.5432 | -0.1021 | 0.003688 | 1.0109 | 1.7355 | 2.7966 | 1.0231
2012-10-30 21:21 | 28.6075 | 15.5563 | 16.2605 | 17.3135 | 15.5725 | -1.4818 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5331 | -0.1022 | 0.003687 | 1.0232 | 1.7274 | 2.7804 | 1.0394
2012-10-30 21:21 | 28.6098 | 15.5423 | 16.2384 | 17.3036 | 15.5137 | -1.4847 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.4906 | -0.1024 | 0.003687 | 1.0517 | 1.7478 | 2.813 | 1.0231
2012-10-30 21:21 | 28.6008 | 15.5333 | 16.2579 | 17.3068 | 15.5658 | -1.483 | 0.006741 | 0.32661 | 0.002202 | 48.452 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4285 | -0.1022 | 0.003691 | 1.1048 | 1.8294 | 2.8783 | 1.1373
2012-10-30 21:21 | 28.5993 | 15.5318 | 16.2361 | 17.2566 | 15.491 | -1.4887 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.5086 | -0.1026 | 0.003689 | 1.0232 | 1.7275 | 2.748 | 0.9824
2012-10-30 21:21 | 28.606 | 15.5548 | 16.2753 | 17.2633 | 15.5466 | -1.4795 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5194 | -0.102 | 0.003691 | 1.0354 | 1.7559 | 2.7439 | 1.0272
2012-10-30 21:21 | 28.5978 | 15.5588 | 16.2305 | 17.2916 | 15.5384 | -1.476 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 14.4582 | -0.1018 | 0.003688 | 1.1006 | 1.7723 | 2.8334 | 1.0802
2012-10-30 21:21 | 28.606 | 15.5548 | 16.2672 | 17.312 | 15.5222 | -1.4778 | 0.006744 | 0.32659 | 0.002202 | 48.4262 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4786 | -0.1019 | 0.003692 | 1.0762 | 1.7886 | 2.8334 | 1.0436
2012-10-30 21:21 | 28.5978 | 15.5384 | 16.2305 | 17.2632 | 15.5099 | -1.4743 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.503 | -0.1016 | 0.003688 | 1.0354 | 1.7275 | 2.7602 | 1.0069
2012-10-30 21:21 | 28.6038 | 15.577 | 16.2446 | 17.2367 | 15.5322 | -1.4674 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006949 | 0.21473 | 0.001492 | 30.9007 | 14.4805 | -0.1012 | 0.00369 | 1.0965 | 1.7641 | 2.7562 | 1.0517
2012-10-30 21:21 | 28.6016 | 15.5463 | 16.2384 | 17.3076 | 15.5341 | -1.4853 | 0.006735 | 0.32661 | 0.0022 | 48.4952 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.4864 | -0.1024 | 0.00369 | 1.0599 | 1.752 | 2.8212 | 1.0477
2012-10-30 21:21 | 28.6023 | 15.5266 | 16.2431 | 17.2718 | 15.5307 | -1.4801 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.479 | -0.102 | 0.003689 | 1.0476 | 1.7641 | 2.7928 | 1.0517
2012-10-30 21:21 | 28.5986 | 15.5514 | 16.2313 | 17.2681 | 15.4822 | -1.4928 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5202 | -0.1029 | 0.003686 | 1.0312 | 1.7111 | 2.7479 | 0.962
2012-10-30 21:21 | 28.5993 | 15.5358 | 16.2198 | 17.2566 | 15.487 | -1.4795 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.5005 | -0.102 | 0.003691 | 1.0353 | 1.7193 | 2.7561 | 0.9865
2012-10-30 21:21 | 28.5956 | 15.5281 | 16.1795 | 17.2448 | 15.4995 | -1.4847 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.5131 | -0.1024 | 0.003687 | 1.015 | 1.6664 | 2.7317 | 0.9864
2012-10-30 21:21 | 28.5948 | 15.4987 | 16.2275 | 17.2399 | 15.5232 | -1.4812 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.5204 | -0.1021 | 0.003687 | 0.9783 | 1.7071 | 2.7195 | 1.0028
2012-10-30 21:21 | 28.5963 | 15.5206 | 16.1883 | 17.2414 | 15.5002 | -1.4731 | 0.006744 | 0.32651 | 0.002202 | 48.4155 | 0.006921 | 0.21473 | 0.001486 | 31.0257 | 14.5097 | -0.1016 | 0.003688 | 1.0109 | 1.6786 | 2.7317 | 0.9905
2012-10-30 21:21 | 28.606 | 15.5059 | 16.2224 | 17.2633 | 15.5262 | -1.4853 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.5112 | -0.1024 | 0.003686 | 0.9947 | 1.7112 | 2.7521 | 1.015
2012-10-30 21:21 | 28.6008 | 15.4966 | 16.2253 | 17.2012 | 15.4925 | -1.4766 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.506 | -0.1018 | 0.003687 | 0.9906 | 1.7193 | 2.6952 | 0.9865
2012-10-30 21:21 | 28.6046 | 15.5208 | 16.16 | 17.2456 | 15.5737 | -1.4864 | 0.006722 | 0.32653 | 0.002195 | 48.5762 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.518 | -0.1025 | 0.003685 | 1.0028 | 1.642 | 2.7276 | 1.0557
2012-10-30 21:21 | 28.6098 | 15.526 | 16.1977 | 17.2265 | 15.5178 | -1.4749 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006917 | 0.21477 | 0.001486 | 31.0499 | 14.5069 | -0.1017 | 0.003687 | 1.0191 | 1.6908 | 2.7196 | 1.0109
2012-10-30 21:21 | 28.6038 | 15.5077 | 16.1917 | 17.2489 | 15.5688 | -1.487 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.5294 | -0.1025 | 0.003686 | 0.9783 | 1.6623 | 2.7195 | 1.0394
2012-10-30 21:21 | 28.5993 | 15.4707 | 16.2116 | 17.22 | 15.4951 | -1.4847 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5168 | -0.1024 | 0.003687 | 0.9539 | 1.6948 | 2.7032 | 0.9783
2012-10-30 21:21 | 28.6053 | 15.5011 | 16.2217 | 17.2341 | 15.4807 | -1.4726 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.4901 | -0.1015 | 0.003688 | 1.011 | 1.7316 | 2.744 | 0.9906
2012-10-30 21:22 | 28.6053 | 15.493 | 16.1566 | 17.1976 | 15.5296 | -1.487 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5228 | -0.1025 | 0.003687 | 0.9702 | 1.6338 | 2.6748 | 1.0068
2012-10-30 21:22 | 28.5963 | 15.488 | 16.168 | 17.2008 | 15.4758 | -1.4835 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4975 | -0.1023 | 0.003682 | 0.9905 | 1.6705 | 2.7033 | 0.9783
2012-10-30 21:22 | 28.5993 | 15.5195 | 16.1791 | 17.2444 | 15.4992 | -1.472 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.5331 | -0.1015 | 0.003685 | 0.9864 | 1.646 | 2.7113 | 0.9661
2012-10-30 21:22 | 28.6001 | 15.5244 | 16.2043 | 17.2249 | 15.4959 | -1.4882 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.4809 | -0.1026 | 0.00368 | 1.0435 | 1.7234 | 2.744 | 1.015
2012-10-30 21:22 | 28.6031 | 15.4826 | 16.1625 | 17.1547 | 15.4948 | -1.4835 | 0.006713 | 0.32654 | 0.002192 | 48.6434 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.5002 | -0.1023 | 0.003681 | 0.9824 | 1.6623 | 2.6545 | 0.9946
2012-10-30 21:22 | 28.5963 | 15.4636 | 16.1354 | 17.217 | 15.4799 | -1.4916 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.5015 | -0.1028 | 0.003688 | 0.9621 | 1.6339 | 2.7155 | 0.9784
2012-10-30 21:22 | 28.6068 | 15.527 | 16.1785 | 17.2113 | 15.4619 | -1.483 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4672 | -0.1022 | 0.003687 | 1.0598 | 1.7113 | 2.7441 | 0.9947
2012-10-30 21:22 | 28.6023 | 15.4533 | 16.1658 | 17.1824 | 15.4981 | -1.4795 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.4708 | -0.102 | 0.003686 | 0.9825 | 1.695 | 2.7116 | 1.0273
2012-10-30 21:22 | 28.6016 | 15.4893 | 16.1366 | 17.1817 | 15.4567 | -1.4714 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.515 | -0.1014 | 0.00369 | 0.9743 | 1.6216 | 2.6667 | 0.9417
2012-10-30 21:22 | 28.5911 | 15.5358 | 16.1383 | 17.2078 | 15.4543 | -1.4697 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006911 | 0.21467 | 0.001484 | 31.0622 | 14.4882 | -0.1013 | 0.003684 | 1.0476 | 1.6501 | 2.7196 | 0.9661
2012-10-30 21:22 | 28.6046 | 15.4515 | 16.1803 | 17.2131 | 15.48 | -1.4702 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006933 | 0.21464 | 0.001488 | 30.9595 | 14.5139 | -0.1014 | 0.003689 | 0.9376 | 1.6664 | 2.6992 | 0.9661
2012-10-30 21:22 | 28.6001 | 15.447 | 16.1718 | 17.2005 | 15.4307 | -1.483 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.5013 | -0.1022 | 0.003686 | 0.9457 | 1.6705 | 2.6992 | 0.9294
2012-10-30 21:22 | 28.6038 | 15.4466 | 16.151 | 17.1514 | 15.4752 | -1.4639 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.00693 | 0.21464 | 0.001487 | 30.9729 | 14.4886 | -0.1009 | 0.003684 | 0.958 | 1.6624 | 2.6628 | 0.9866
2012-10-30 21:22 | 28.6075 | 15.5196 | 16.1669 | 17.2079 | 15.434 | -1.4783 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006939 | 0.21467 | 0.00149 | 30.9369 | 14.5413 | -0.1019 | 0.003687 | 0.9783 | 1.6256 | 2.6666 | 0.8927
2012-10-30 21:22 | 28.5971 | 15.444 | 16.1443 | 17.1813 | 15.444 | -1.4858 | 0.006741 | 0.32653 | 0.002201 | 48.4393 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5227 | -0.1024 | 0.00369 | 0.9213 | 1.6216 | 2.6586 | 0.9213
2012-10-30 21:22 | 28.6075 | 15.43 | 16.1507 | 17.147 | 15.4626 | -1.4853 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5046 | -0.1024 | 0.003687 | 0.9254 | 1.6461 | 2.6424 | 0.958
2012-10-30 21:22 | 28.6113 | 15.4297 | 16.1911 | 17.1711 | 15.4501 | -1.4887 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.5451 | -0.1026 | 0.003686 | 0.8846 | 1.646 | 2.626 | 0.905
2012-10-30 21:22 | 28.6008 | 15.4314 | 16.1602 | 17.1768 | 15.4722 | -1.483 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5754 | -0.1022 | 0.003688 | 0.856 | 1.5848 | 2.6014 | 0.8968
2012-10-30 21:22 | 28.6038 | 15.4466 | 16.1429 | 17.127 | 15.4507 | -1.4841 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 14.5335 | -0.1023 | 0.003683 | 0.9131 | 1.6094 | 2.5935 | 0.9172
2012-10-30 21:22 | 28.6075 | 15.4585 | 16.1262 | 17.1266 | 15.4503 | -1.4789 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 14.5698 | -0.102 | 0.003689 | 0.8887 | 1.5564 | 2.5568 | 0.8805
2012-10-30 21:22 | 28.612 | 15.4304 | 16.1389 | 17.2043 | 15.5363 | -1.4806 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5499 | -0.1021 | 0.003691 | 0.8805 | 1.589 | 2.6544 | 0.9864
2012-10-30 21:22 | 28.6113 | 15.4501 | 16.1545 | 17.1467 | 15.4256 | -1.4714 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006924 | 0.21467 | 0.001486 | 31.0039 | 14.5247 | -0.1014 | 0.003687 | 0.9254 | 1.6298 | 2.622 | 0.9009
2012-10-30 21:22 | 28.612 | 15.4467 | 16.1185 | 17.1636 | 15.4508 | -1.4841 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.5295 | -0.1023 | 0.003685 | 0.9172 | 1.589 | 2.6341 | 0.9213
2012-10-30 21:22 | 28.6113 | 15.393 | 16.1992 | 17.1304 | 15.4908 | -1.4806 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.5736 | -0.1021 | 0.003687 | 0.8194 | 1.6256 | 2.5568 | 0.9172
2012-10-30 21:22 | 28.609 | 15.4193 | 16.14 | 17.1322 | 15.4926 | -1.4876 | 0.006713 | 0.32653 | 0.002192 | 48.6413 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.5061 | -0.1026 | 0.00368 | 0.9132 | 1.6339 | 2.6261 | 0.9865
2012-10-30 21:22 | 28.6195 | 15.4949 | 16.1667 | 17.1427 | 15.4542 | -1.4778 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006939 | 0.21468 | 0.00149 | 30.9389 | 14.5411 | -0.1019 | 0.003687 | 0.9538 | 1.6256 | 2.6016 | 0.9131
2012-10-30 21:22 | 28.621 | 15.3783 | 16.1642 | 17.132 | 15.4679 | -1.4806 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5344 | -0.1021 | 0.003688 | 0.8439 | 1.6298 | 2.5976 | 0.9335
2012-10-30 21:22 | 28.6135 | 15.4034 | 16.12 | 17.153 | 15.4604 | -1.4685 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4983 | -0.1012 | 0.003687 | 0.9051 | 1.6217 | 2.6547 | 0.9621
2012-10-30 21:22 | 28.6113 | 15.393 | 16.1504 | 17.1061 | 15.4419 | -1.476 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.5002 | -0.1018 | 0.003684 | 0.8928 | 1.6502 | 2.6059 | 0.9417
2012-10-30 21:23 | 28.6255 | 15.4113 | 16.1687 | 17.1081 | 15.4887 | -1.461 | 0.006738 | 0.32659 | 0.002201 | 48.4693 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5511 | -0.1007 | 0.00369 | 0.8602 | 1.6176 | 2.557 | 0.9376
2012-10-30 21:23 | 28.6128 | 15.3782 | 16.1519 | 17.0913 | 15.4556 | -1.4656 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.5344 | -0.101 | 0.003685 | 0.8438 | 1.6175 | 2.5569 | 0.9212
2012-10-30 21:23 | 28.6248 | 15.3821 | 16.1598 | 17.1358 | 15.4595 | -1.4754 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.53 | -0.1017 | 0.003692 | 0.8521 | 1.6298 | 2.6058 | 0.9295
2012-10-30 21:23 | 28.612 | 15.3774 | 16.147 | 17.0986 | 15.4711 | -1.4726 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5132 | -0.1015 | 0.00369 | 0.8642 | 1.6338 | 2.5854 | 0.9579
2012-10-30 21:23 | 28.6165 | 15.3819 | 16.1556 | 17.0869 | 15.4634 | -1.4737 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.534 | -0.1016 | 0.003688 | 0.8479 | 1.6216 | 2.5529 | 0.9294
2012-10-30 21:23 | 28.609 | 15.3663 | 16.1155 | 17.0916 | 15.4315 | -1.4749 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 14.5061 | -0.1017 | 0.003688 | 0.8602 | 1.6094 | 2.5855 | 0.9254
2012-10-30 21:23 | 28.615 | 15.3927 | 16.1622 | 17.0772 | 15.4782 | -1.4754 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006924 | 0.2147 | 0.001487 | 31.0081 | 14.5243 | -0.1017 | 0.003684 | 0.8684 | 1.6379 | 2.5529 | 0.9539
2012-10-30 21:23 | 28.6046 | 15.3782 | 16.1315 | 17.075 | 15.4434 | -1.4778 | 0.006713 | 0.32653 | 0.002192 | 48.6413 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4935 | -0.1019 | 0.003683 | 0.8847 | 1.638 | 2.5815 | 0.9499
2012-10-30 21:23 | 28.6128 | 15.4149 | 16.1682 | 17.0954 | 15.4964 | -1.4772 | 0.006722 | 0.32657 | 0.002195 | 48.5826 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 14.5099 | -0.1018 | 0.003684 | 0.905 | 1.6583 | 2.5855 | 0.9865
2012-10-30 21:23 | 28.6068 | 15.4171 | 16.1337 | 17.13 | 15.4782 | -1.4662 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4916 | -0.1011 | 0.003687 | 0.9255 | 1.6421 | 2.6384 | 0.9866
2012-10-30 21:23 | 28.6128 | 15.4231 | 16.156 | 17.0669 | 15.4556 | -1.4876 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.518 | -0.1026 | 0.003691 | 0.9051 | 1.638 | 2.5489 | 0.9376
2012-10-30 21:23 | 28.6105 | 15.3596 | 16.0723 | 17.0524 | 15.4656 | -1.4795 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006921 | 0.21468 | 0.001486 | 31.0194 | 14.5239 | -0.102 | 0.003685 | 0.8357 | 1.5484 | 2.5285 | 0.9417
2012-10-30 21:23 | 28.6128 | 15.3619 | 16.0949 | 17.0913 | 15.4556 | -1.4702 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006921 | 0.21473 | 0.001486 | 31.0257 | 14.518 | -0.1014 | 0.003685 | 0.8439 | 1.5769 | 2.5733 | 0.9376
2012-10-30 21:23 | 28.606 | 15.3429 | 16.0718 | 17.0723 | 15.4611 | -1.4697 | 0.006735 | 0.32664 | 0.0022 | 48.4995 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5153 | -0.1013 | 0.00369 | 0.8276 | 1.5565 | 2.557 | 0.9458
2012-10-30 21:23 | 28.6105 | 15.4004 | 16.1008 | 17.089 | 15.4696 | -1.4766 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.5524 | -0.1018 | 0.00369 | 0.848 | 1.5484 | 2.5366 | 0.9172
2012-10-30 21:23 | 28.6135 | 15.3463 | 16.1404 | 17.0432 | 15.4889 | -1.4766 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5351 | -0.1018 | 0.003687 | 0.8112 | 1.6053 | 2.5081 | 0.9538
2012-10-30 21:23 | 28.6128 | 15.3375 | 16.1193 | 17.0547 | 15.4353 | -1.4749 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.465 | -0.1017 | 0.003694 | 0.8725 | 1.6543 | 2.5897 | 0.9703
2012-10-30 21:23 | 28.6188 | 15.3679 | 16.0968 | 17.0607 | 15.425 | -1.4766 | 0.006716 | 0.32656 | 0.002193 | 48.6237 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5077 | -0.1018 | 0.003682 | 0.8602 | 1.5891 | 2.553 | 0.9173
2012-10-30 21:23 | 28.6158 | 15.3405 | 16.1345 | 17.0618 | 15.4546 | -1.465 | 0.006722 | 0.32656 | 0.002195 | 48.5803 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5129 | -0.101 | 0.003687 | 0.8276 | 1.6216 | 2.5489 | 0.9417
2012-10-30 21:23 | 28.6218 | 15.3506 | 16.0958 | 17.019 | 15.4117 | -1.4789 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5352 | -0.102 | 0.003685 | 0.8154 | 1.5606 | 2.4838 | 0.8765
2012-10-30 21:23 | 28.615 | 15.3723 | 16.1053 | 17.0285 | 15.4375 | -1.472 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006946 | 0.21477 | 0.001492 | 30.9202 | 14.4917 | -0.1015 | 0.00369 | 0.8806 | 1.6136 | 2.5368 | 0.9458
2012-10-30 21:23 | 28.6143 | 15.3064 | 16.0883 | 17.0115 | 15.4368 | -1.4876 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006917 | 0.21468 | 0.001485 | 31.0373 | 14.5318 | -0.1026 | 0.003685 | 0.7746 | 1.5565 | 2.4797 | 0.905
2012-10-30 21:23 | 28.615 | 15.3152 | 16.093 | 17.0691 | 15.4171 | -1.4714 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 14.5651 | -0.1014 | 0.003689 | 0.7501 | 1.5279 | 2.504 | 0.852
2012-10-30 21:23 | 28.6128 | 15.3171 | 16.1315 | 17.006 | 15.4231 | -1.4726 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5344 | -0.1015 | 0.003685 | 0.7827 | 1.5971 | 2.4716 | 0.8887
2012-10-30 21:23 | 28.615 | 15.3152 | 16.1419 | 16.996 | 15.4375 | -1.4812 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5651 | -0.1021 | 0.003686 | 0.7501 | 1.5768 | 2.4309 | 0.8724
2012-10-30 21:23 | 28.6046 | 15.2967 | 16.0867 | 17.0425 | 15.4108 | -1.483 | 0.006713 | 0.32654 | 0.002192 | 48.6434 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5669 | -0.1022 | 0.003681 | 0.7298 | 1.5198 | 2.4756 | 0.8439
2012-10-30 21:23 | 28.612 | 15.3122 | 16.09 | 17.0133 | 15.41 | -1.4795 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5825 | -0.102 | 0.003689 | 0.7297 | 1.5075 | 2.4308 | 0.8275
2012-10-30 21:24 | 28.6068 | 15.3111 | 16.0808 | 16.9756 | 15.3845 | -1.472 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.561 | -0.1015 | 0.00369 | 0.7501 | 1.5198 | 2.4146 | 0.8235
2012-10-30 21:24 | 28.6128 | 15.3456 | 16.1315 | 17.006 | 15.3986 | -1.4818 | 0.006744 | 0.32651 | 0.002202 | 48.4155 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.5751 | -0.1022 | 0.003693 | 0.7705 | 1.5564 | 2.4309 | 0.8235
2012-10-30 21:24 | 28.6135 | 15.2934 | 16.0997 | 17.0107 | 15.44 | -1.4714 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5922 | -0.1014 | 0.00369 | 0.7012 | 1.5075 | 2.4185 | 0.8478
2012-10-30 21:24 | 28.6038 | 15.2715 | 16.0941 | 16.9848 | 15.3611 | -1.4801 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.558 | -0.102 | 0.003687 | 0.7135 | 1.5361 | 2.4268 | 0.8031
2012-10-30 21:24 | 28.6135 | 15.3137 | 16.0915 | 17.0107 | 15.4278 | -1.4697 | 0.006738 | 0.3266 | 0.002201 | 48.4715 | 0.006917 | 0.21471 | 0.001485 | 31.0415 | 14.531 | -0.1013 | 0.003686 | 0.7827 | 1.5605 | 2.4797 | 0.8968
2012-10-30 21:24 | 28.6135 | 15.2934 | 16.1241 | 16.9985 | 15.436 | -1.4806 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.006921 | 0.21476 | 0.001486 | 31.0299 | 14.5962 | -0.1021 | 0.003687 | 0.6972 | 1.5279 | 2.4023 | 0.8398
2012-10-30 21:24 | 28.6038 | 15.2837 | 16.0981 | 16.9929 | 15.4059 | -1.4772 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.5825 | -0.1018 | 0.003687 | 0.7012 | 1.5156 | 2.4104 | 0.8234
2012-10-30 21:24 | 28.6068 | 15.2867 | 16.0848 | 16.9796 | 15.4333 | -1.4812 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.5691 | -0.1021 | 0.003692 | 0.7176 | 1.5157 | 2.4105 | 0.8642
2012-10-30 21:24 | 28.6075 | 15.2752 | 16.0855 | 16.9844 | 15.3974 | -1.4818 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.5862 | -0.1022 | 0.003685 | 0.689 | 1.4993 | 2.3982 | 0.8112
2012-10-30 21:24 | 28.6083 | 15.2597 | 16.1189 | 16.9527 | 15.3697 | -1.4731 | 0.006744 | 0.32653 | 0.002202 | 48.4177 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5339 | -0.1016 | 0.003691 | 0.7258 | 1.585 | 2.4188 | 0.8358
2012-10-30 21:24 | 28.5971 | 15.2973 | 16.0914 | 16.9577 | 15.4236 | -1.4697 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5023 | -0.1013 | 0.003688 | 0.795 | 1.5891 | 2.4554 | 0.9213
2012-10-30 21:24 | 28.6023 | 15.2862 | 16.0437 | 16.9507 | 15.4411 | -1.483 | 0.006741 | 0.32657 | 0.002201 | 48.4456 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 14.5361 | -0.1022 | 0.00369 | 0.7501 | 1.5076 | 2.4146 | 0.905
2012-10-30 21:24 | 28.5993 | 15.2792 | 16.0651 | 16.9843 | 15.4136 | -1.4876 | 0.006713 | 0.32653 | 0.002192 | 48.6413 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5045 | -0.1026 | 0.003682 | 0.7747 | 1.5606 | 2.4798 | 0.9091
2012-10-30 21:24 | 28.6016 | 15.3263 | 16.0512 | 16.9663 | 15.3915 | -1.4905 | 0.006751 | 0.32654 | 0.002204 | 48.3696 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.5395 | -0.1028 | 0.003694 | 0.7868 | 1.5117 | 2.4268 | 0.852
2012-10-30 21:24 | 28.5933 | 15.2935 | 16.0713 | 16.9743 | 15.3669 | -1.4714 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006917 | 0.2147 | 0.001485 | 31.0395 | 14.4781 | -0.1014 | 0.003682 | 0.8154 | 1.5932 | 2.4962 | 0.8888
2012-10-30 21:24 | 28.5971 | 15.2892 | 16.0426 | 16.9374 | 15.387 | -1.4882 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006921 | 0.21477 | 0.001486 | 31.0319 | 14.5146 | -0.1026 | 0.003681 | 0.7746 | 1.528 | 2.4228 | 0.8724
2012-10-30 21:24 | 28.5926 | 15.2643 | 16.0462 | 16.9573 | 15.4069 | -1.4835 | 0.006741 | 0.3266 | 0.002202 | 48.4499 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5264 | -0.1023 | 0.003691 | 0.7379 | 1.5198 | 2.4309 | 0.8805
2012-10-30 21:24 | 28.606 | 15.2777 | 16.0718 | 16.9951 | 15.3714 | -1.4951 | 0.006738 | 0.32657 | 0.0022 | 48.4672 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 14.5724 | -0.1031 | 0.003687 | 0.7053 | 1.4994 | 2.4227 | 0.799
2012-10-30 21:24 | 28.5933 | 15.3139 | 16.0306 | 16.9539 | 15.4198 | -1.4778 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.5801 | -0.1019 | 0.003685 | 0.7338 | 1.4505 | 2.3738 | 0.8397
2012-10-30 21:24 | 28.5926 | 15.2847 | 15.9811 | 16.9654 | 15.3866 | -1.4934 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5142 | -0.103 | 0.003685 | 0.7705 | 1.4669 | 2.4512 | 0.8724
2012-10-30 21:24 | 28.6008 | 15.297 | 16.0951 | 16.9533 | 15.4029 | -1.4864 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.5835 | -0.1025 | 0.003686 | 0.7135 | 1.5116 | 2.3698 | 0.8194
2012-10-30 21:24 | 28.5933 | 15.2935 | 16.0306 | 16.9702 | 15.3669 | -1.5043 | 0.006744 | 0.32656 | 0.002202 | 48.4219 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.5026 | -0.1037 | 0.003691 | 0.7909 | 1.528 | 2.4676 | 0.8643
2012-10-30 21:24 | 28.5941 | 15.2495 | 16.0355 | 16.9791 | 15.3555 | -1.4824 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 14.5238 | -0.1022 | 0.00369 | 0.7257 | 1.5117 | 2.4553 | 0.8317
2012-10-30 21:24 | 28.6031 | 15.3074 | 16.0404 | 16.9394 | 15.3971 | -1.4899 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.5247 | -0.1027 | 0.003688 | 0.7827 | 1.5157 | 2.4147 | 0.8724
2012-10-30 21:24 | 28.5948 | 15.2217 | 16.0566 | 16.9311 | 15.3806 | -1.4887 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5123 | -0.1026 | 0.003689 | 0.7094 | 1.5443 | 2.4188 | 0.8683
2012-10-30 21:24 | 28.5888 | 15.2483 | 16.0099 | 16.8966 | 15.3746 | -1.4789 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.5226 | -0.102 | 0.003687 | 0.7257 | 1.4873 | 2.374 | 0.852
2012-10-30 21:24 | 28.5948 | 15.2339 | 16.024 | 16.9514 | 15.3969 | -1.4812 | 0.006747 | 0.3265 | 0.002203 | 48.3919 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5082 | -0.1021 | 0.003692 | 0.7257 | 1.5158 | 2.4432 | 0.8887
2012-10-30 21:25 | 28.5986 | 15.2459 | 16.0197 | 16.9105 | 15.3844 | -1.4905 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.5324 | -0.1028 | 0.003682 | 0.7135 | 1.4873 | 2.3781 | 0.852
2012-10-30 21:25 | 28.6023 | 15.2496 | 16.06 | 16.9345 | 15.3759 | -1.4876 | 0.006747 | 0.32654 | 0.002203 | 48.3983 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.5198 | -0.1026 | 0.00369 | 0.7298 | 1.5402 | 2.4147 | 0.8561
2012-10-30 21:25 | 28.5873 | 15.2183 | 16.0287 | 16.9195 | 15.3568 | -1.4963 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.4762 | -0.1032 | 0.003684 | 0.7421 | 1.5525 | 2.4433 | 0.8806
2012-10-30 21:25 | 28.5956 | 15.2388 | 16.0207 | 16.8912 | 15.3651 | -1.4887 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.5253 | -0.1026 | 0.003686 | 0.7135 | 1.4954 | 2.3659 | 0.8398
2012-10-30 21:25 | 28.5941 | 15.2414 | 16.0355 | 16.9222 | 15.3555 | -1.4887 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5197 | -0.1026 | 0.003687 | 0.7217 | 1.5158 | 2.4025 | 0.8358
2012-10-30 21:25 | 28.5926 | 15.2562 | 16.0096 | 16.9126 | 15.3417 | -1.4899 | 0.006757 | 0.32657 | 0.002207 | 48.3309 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 14.5223 | -0.1027 | 0.003694 | 0.7339 | 1.4873 | 2.3903 | 0.8194
2012-10-30 21:25 | 28.5851 | 15.2405 | 16.0102 | 16.8848 | 15.3994 | -1.4841 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006917 | 0.21474 | 0.001485 | 31.0457 | 14.5107 | -0.1023 | 0.003684 | 0.7298 | 1.4995 | 2.3741 | 0.8887
2012-10-30 21:25 | 28.5888 | 15.2442 | 15.9732 | 16.8803 | 15.3339 | -1.4841 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5348 | -0.1023 | 0.003684 | 0.7094 | 1.4384 | 2.3455 | 0.7991
2012-10-30 21:25 | 28.5933 | 15.2487 | 16.0347 | 16.9052 | 15.3873 | -1.4754 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.5108 | -0.1017 | 0.003686 | 0.7379 | 1.5239 | 2.3944 | 0.8765
2012-10-30 21:25 | 28.5806 | 15.2157 | 16.0505 | 16.8925 | 15.3623 | -1.4841 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.5225 | -0.1023 | 0.003689 | 0.6932 | 1.528 | 2.37 | 0.8398
2012-10-30 21:25 | 28.5731 | 15.2367 | 15.9982 | 16.9012 | 15.3548 | -1.4772 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006924 | 0.21471 | 0.001487 | 31.0101 | 14.4947 | -0.1018 | 0.003684 | 0.742 | 1.5035 | 2.4065 | 0.8601
2012-10-30 21:25 | 28.5828 | 15.2382 | 16.0201 | 16.9312 | 15.3034 | -1.4853 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.4962 | -0.1024 | 0.003686 | 0.742 | 1.5239 | 2.435 | 0.8072
2012-10-30 21:25 | 28.5918 | 15.2065 | 16.0088 | 16.8874 | 15.3613 | -1.4899 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5297 | -0.1027 | 0.003686 | 0.6768 | 1.4791 | 2.3577 | 0.8316
2012-10-30 21:25 | 28.5851 | 15.2487 | 15.9899 | 16.8726 | 15.3587 | -1.4974 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.5556 | -0.1032 | 0.003689 | 0.6931 | 1.4343 | 2.317 | 0.8031
2012-10-30 21:25 | 28.5828 | 15.2505 | 16.0364 | 16.8825 | 15.3605 | -1.4841 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5696 | -0.1023 | 0.003684 | 0.6809 | 1.4668 | 2.3129 | 0.7909
2012-10-30 21:25 | 28.5746 | 15.2504 | 16.0242 | 16.8377 | 15.3278 | -1.4847 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.4554 | -0.1024 | 0.003684 | 0.795 | 1.5688 | 2.3823 | 0.8724
2012-10-30 21:25 | 28.5701 | 15.2215 | 15.979 | 16.8657 | 15.3437 | -1.4824 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.4753 | -0.1022 | 0.003687 | 0.7462 | 1.5037 | 2.3904 | 0.8684
2012-10-30 21:25 | 28.5836 | 15.2553 | 16.0087 | 16.8264 | 15.3776 | -1.4801 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.497 | -0.102 | 0.003686 | 0.7583 | 1.5117 | 2.3294 | 0.8806
2012-10-30 21:25 | 28.5776 | 15.1719 | 15.9905 | 16.8773 | 15.3349 | -1.4934 | 0.006722 | 0.32656 | 0.002195 | 48.5803 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.5195 | -0.103 | 0.003683 | 0.6524 | 1.471 | 2.3578 | 0.8154
2012-10-30 21:25 | 28.5753 | 15.2389 | 16.0004 | 16.8912 | 15.3489 | -1.4824 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.4683 | -0.1022 | 0.00369 | 0.7706 | 1.5321 | 2.4229 | 0.8806
2012-10-30 21:25 | 28.5768 | 15.1956 | 15.9775 | 16.8439 | 15.3382 | -1.4899 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5106 | -0.1027 | 0.003684 | 0.685 | 1.4669 | 2.3333 | 0.8276
2012-10-30 21:25 | 28.5783 | 15.2582 | 15.979 | 16.8536 | 15.3682 | -1.4887 | 0.006719 | 0.32653 | 0.002194 | 48.5979 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 14.4835 | -0.1026 | 0.003681 | 0.7747 | 1.4955 | 2.3701 | 0.8847
2012-10-30 21:25 | 28.5866 | 15.2013 | 15.9181 | 16.8903 | 15.3154 | -1.4772 | 0.006716 | 0.3265 | 0.002193 | 48.6152 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.5 | -0.1018 | 0.003684 | 0.7013 | 1.4181 | 2.3903 | 0.8154
2012-10-30 21:25 | 28.5791 | 15.1571 | 15.9228 | 16.8097 | 15.2875 | -1.4882 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 14.5088 | -0.1026 | 0.003684 | 0.6483 | 1.414 | 2.3009 | 0.7787
2012-10-30 21:25 | 28.5746 | 15.1771 | 15.9346 | 16.8052 | 15.3523 | -1.4795 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.4921 | -0.102 | 0.00369 | 0.685 | 1.4425 | 2.3131 | 0.8602
2012-10-30 21:25 | 28.5723 | 15.1503 | 16.0178 | 16.8151 | 15.297 | -1.4858 | 0.006744 | 0.3266 | 0.002203 | 48.4284 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.5346 | -0.1024 | 0.003692 | 0.6157 | 1.4832 | 2.2805 | 0.7624
2012-10-30 21:25 | 28.5738 | 15.1681 | 15.9542 | 16.8003 | 15.2863 | -1.4789 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006924 | 0.21466 | 0.001486 | 31.0017 | 14.5239 | -0.102 | 0.003684 | 0.6442 | 1.4303 | 2.2764 | 0.7624
2012-10-30 21:25 | 28.5686 | 15.1751 | 15.9286 | 16.8073 | 15.3503 | -1.4887 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5391 | -0.1026 | 0.003686 | 0.636 | 1.3895 | 2.2682 | 0.8112
2012-10-30 21:25 | 28.5746 | 15.1771 | 16.0038 | 16.8783 | 15.3645 | -1.4899 | 0.006722 | 0.32653 | 0.002195 | 48.5762 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5329 | -0.1027 | 0.003683 | 0.6442 | 1.4709 | 2.3454 | 0.8316
2012-10-30 21:26 | 28.5746 | 15.1526 | 16.0079 | 16.8377 | 15.3482 | -1.4818 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.5329 | -0.1022 | 0.003687 | 0.6197 | 1.475 | 2.3048 | 0.8153
2012-10-30 21:26 | 28.5768 | 15.1833 | 15.9897 | 16.8277 | 15.33 | -1.4789 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.5147 | -0.102 | 0.00369 | 0.6686 | 1.475 | 2.313 | 0.8153
2012-10-30 21:26 | 28.5791 | 15.1612 | 15.9106 | 16.8219 | 15.3364 | -1.487 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.5292 | -0.1025 | 0.003689 | 0.632 | 1.3814 | 2.2927 | 0.8072
2012-10-30 21:26 | 28.5678 | 15.1662 | 15.9767 | 16.8106 | 15.2803 | -1.4899 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.469 | -0.1027 | 0.003687 | 0.6972 | 1.5077 | 2.3416 | 0.8113
2012-10-30 21:26 | 28.5723 | 15.1992 | 15.9689 | 16.8069 | 15.354 | -1.4853 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.5387 | -0.1024 | 0.003686 | 0.6605 | 1.4302 | 2.2682 | 0.8153
2012-10-30 21:26 | 28.5701 | 15.1929 | 15.9423 | 16.8047 | 15.307 | -1.5009 | 0.006719 | 0.32654 | 0.002194 | 48.5999 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5324 | -0.1035 | 0.003686 | 0.6605 | 1.4099 | 2.2723 | 0.7746
2012-10-30 21:26 | 28.5738 | 15.2089 | 15.942 | 16.8206 | 15.2944 | -1.4812 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5321 | -0.1021 | 0.00369 | 0.6768 | 1.4099 | 2.2885 | 0.7623
2012-10-30 21:26 | 28.5664 | 15.1729 | 15.9427 | 16.8214 | 15.3318 | -1.4858 | 0.006722 | 0.32657 | 0.002195 | 48.5826 | 0.006921 | 0.21477 | 0.001486 | 31.0319 | 14.5043 | -0.1024 | 0.003681 | 0.6686 | 1.4384 | 2.3171 | 0.8275
2012-10-30 21:26 | 28.5656 | 15.1803 | 15.9256 | 16.8002 | 15.3433 | -1.4864 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.5279 | -0.1025 | 0.00369 | 0.6524 | 1.3977 | 2.2723 | 0.8154
2012-10-30 21:26 | 28.5574 | 15.1517 | 15.95 | 16.7839 | 15.2739 | -1.4812 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.5361 | -0.1021 | 0.003687 | 0.6156 | 1.4139 | 2.2478 | 0.7378
2012-10-30 21:26 | 28.5611 | 15.1391 | 15.9537 | 16.7673 | 15.3265 | -1.4841 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.5194 | -0.1023 | 0.003687 | 0.6197 | 1.4343 | 2.2479 | 0.8071
2012-10-30 21:26 | 28.5641 | 15.1258 | 15.973 | 16.8312 | 15.2847 | -1.4945 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006924 | 0.2147 | 0.001487 | 31.0081 | 14.5387 | -0.103 | 0.003684 | 0.5871 | 1.4343 | 2.2925 | 0.746
2012-10-30 21:26 | 28.5626 | 15.1284 | 15.9755 | 16.8094 | 15.2832 | -1.483 | 0.006719 | 0.32653 | 0.002194 | 48.5979 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.5372 | -0.1022 | 0.003683 | 0.5912 | 1.4383 | 2.2722 | 0.746
2012-10-30 21:26 | 28.5559 | 15.138 | 15.9648 | 16.7783 | 15.2887 | -1.4899 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.506 | -0.1027 | 0.003684 | 0.632 | 1.4588 | 2.2723 | 0.7827
2012-10-30 21:26 | 28.5559 | 15.1665 | 15.9403 | 16.8393 | 15.3376 | -1.4934 | 0.006707 | 0.32643 | 0.002189 | 48.6696 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.5223 | -0.103 | 0.00368 | 0.6442 | 1.418 | 2.317 | 0.8153
2012-10-30 21:26 | 28.5521 | 15.1423 | 15.9203 | 16.8111 | 15.3216 | -1.4905 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4777 | -0.1028 | 0.003685 | 0.6646 | 1.4426 | 2.3334 | 0.8439
2012-10-30 21:26 | 28.5476 | 15.146 | 15.9687 | 16.77 | 15.313 | -1.4858 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006952 | 0.2148 | 0.001493 | 30.8977 | 14.4895 | -0.1024 | 0.003692 | 0.6565 | 1.4792 | 2.2805 | 0.8235
2012-10-30 21:26 | 28.5529 | 15.1105 | 15.9577 | 16.7631 | 15.3346 | -1.4858 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4908 | -0.1024 | 0.003687 | 0.6197 | 1.4669 | 2.2723 | 0.8438
2012-10-30 21:26 | 28.5379 | 15.1485 | 15.8938 | 16.7441 | 15.2911 | -1.4968 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.4717 | -0.1032 | 0.003684 | 0.6768 | 1.4221 | 2.2724 | 0.8194
2012-10-30 21:26 | 28.5536 | 15.1887 | 15.9299 | 16.7842 | 15.2742 | -1.4934 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.5078 | -0.103 | 0.003685 | 0.6809 | 1.4221 | 2.2764 | 0.7664
2012-10-30 21:26 | 28.5484 | 15.1835 | 15.884 | 16.7464 | 15.2975 | -1.4887 | 0.006719 | 0.32653 | 0.002194 | 48.5979 | 0.006924 | 0.21468 | 0.001486 | 31.0059 | 14.5026 | -0.1026 | 0.00368 | 0.6809 | 1.3814 | 2.2438 | 0.7949
2012-10-30 21:26 | 28.5454 | 15.1438 | 15.9461 | 16.776 | 15.2782 | -1.476 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4996 | -0.1018 | 0.003682 | 0.6442 | 1.4465 | 2.2764 | 0.7786
2012-10-30 21:26 | 28.5341 | 15.1162 | 15.9023 | 16.7606 | 15.2832 | -1.4858 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4964 | -0.1024 | 0.003685 | 0.6198 | 1.4059 | 2.2642 | 0.7868
2012-10-30 21:27 | 28.5341 | 15.0958 | 15.9023 | 16.7362 | 15.2751 | -1.4841 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5005 | -0.1023 | 0.003685 | 0.5953 | 1.4018 | 2.2357 | 0.7746
2012-10-30 21:27 | 28.5326 | 15.184 | 15.8885 | 16.7591 | 15.2858 | -1.4841 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5317 | -0.1023 | 0.003687 | 0.6523 | 1.3568 | 2.2274 | 0.7541
2012-10-30 21:27 | 28.5304 | 15.1002 | 15.8863 | 16.7406 | 15.2388 | -1.4795 | 0.006716 | 0.3265 | 0.002193 | 48.6152 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.505 | -0.102 | 0.003683 | 0.5952 | 1.3813 | 2.2356 | 0.7338
2012-10-30 21:27 | 28.5409 | 15.1148 | 15.9091 | 16.7105 | 15.3104 | -1.4905 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.5196 | -0.1028 | 0.00369 | 0.5952 | 1.3895 | 2.1909 | 0.7908
2012-10-30 21:27 | 28.5341 | 15.0876 | 15.9267 | 16.7484 | 15.2425 | -1.4818 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.4924 | -0.1022 | 0.003684 | 0.5952 | 1.4343 | 2.256 | 0.7501
2012-10-30 21:27 | 28.5267 | 15.1047 | 15.9437 | 16.6597 | 15.2555 | -1.4882 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.5258 | -0.1026 | 0.003689 | 0.5789 | 1.4179 | 2.1339 | 0.7297
2012-10-30 21:27 | 28.5401 | 15.0733 | 15.9164 | 16.7097 | 15.2607 | -1.4905 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.531 | -0.1028 | 0.003686 | 0.5423 | 1.3854 | 2.1787 | 0.7297
2012-10-30 21:27 | 28.5446 | 15.0859 | 15.9005 | 16.7467 | 15.2693 | -1.4858 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.5233 | -0.1024 | 0.003687 | 0.5626 | 1.3772 | 2.2234 | 0.746
2012-10-30 21:27 | 28.5341 | 15.0632 | 15.9267 | 16.6671 | 15.271 | -1.4882 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006933 | 0.21468 | 0.001488 | 30.9657 | 14.5291 | -0.1026 | 0.003685 | 0.5341 | 1.3976 | 2.138 | 0.7419
2012-10-30 21:27 | 28.5371 | 15.0743 | 15.9053 | 16.6985 | 15.2496 | -1.4968 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.5076 | -0.1032 | 0.003685 | 0.5667 | 1.3977 | 2.1909 | 0.742
2012-10-30 21:27 | 28.5454 | 15.0867 | 15.9217 | 16.7109 | 15.2864 | -1.4876 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.5159 | -0.1026 | 0.00369 | 0.5708 | 1.4058 | 2.195 | 0.7705
2012-10-30 21:27 | 28.5394 | 15.0277 | 15.8913 | 16.6927 | 15.2804 | -1.4899 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5058 | -0.1027 | 0.003689 | 0.5219 | 1.3855 | 2.1869 | 0.7746
2012-10-30 21:27 | 28.5371 | 15.0906 | 15.8686 | 16.6416 | 15.2292 | -1.4934 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.5321 | -0.103 | 0.003683 | 0.5585 | 1.3365 | 2.1095 | 0.6971
2012-10-30 21:27 | 28.5371 | 15.0988 | 15.8971 | 16.6863 | 15.2455 | -1.483 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.528 | -0.1022 | 0.003687 | 0.5708 | 1.3691 | 2.1583 | 0.7175
2012-10-30 21:27 | 28.5416 | 15.0544 | 15.8446 | 16.6908 | 15.2867 | -1.4963 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006921 | 0.21471 | 0.001486 | 31.0236 | 14.5243 | -0.1032 | 0.003683 | 0.5301 | 1.3203 | 2.1665 | 0.7624
2012-10-30 21:27 | 28.5364 | 15.0573 | 15.8557 | 16.6613 | 15.2896 | -1.4864 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.5314 | -0.1025 | 0.003685 | 0.5259 | 1.3243 | 2.1299 | 0.7582
2012-10-30 21:27 | 28.5394 | 15.0481 | 15.8668 | 16.6805 | 15.2682 | -1.4801 | 0.006719 | 0.32654 | 0.002194 | 48.5999 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 14.5099 | -0.102 | 0.003686 | 0.5382 | 1.3569 | 2.1706 | 0.7583
2012-10-30 21:27 | 28.5386 | 15.0269 | 15.8742 | 16.6919 | 15.2959 | -1.4737 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006943 | 0.21484 | 0.001492 | 30.944 | 14.505 | -0.1016 | 0.003688 | 0.5219 | 1.3692 | 2.1869 | 0.7909
2012-10-30 21:27 | 28.5439 | 15.0893 | 15.8917 | 16.6769 | 15.236 | -1.4887 | 0.006738 | 0.32664 | 0.002201 | 48.4779 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5226 | -0.1026 | 0.003689 | 0.5667 | 1.3691 | 2.1543 | 0.7134
2012-10-30 21:27 | 28.5267 | 15.011 | 15.8419 | 16.6597 | 15.2555 | -1.4841 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.4768 | -0.1023 | 0.003686 | 0.5342 | 1.3651 | 2.1829 | 0.7787
2012-10-30 21:27 | 28.5401 | 15.057 | 15.8635 | 16.669 | 15.2933 | -1.4951 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5106 | -0.1031 | 0.003688 | 0.5464 | 1.3529 | 2.1584 | 0.7827
2012-10-30 21:27 | 28.5326 | 15.0291 | 15.8845 | 16.6859 | 15.2654 | -1.487 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 14.548 | -0.1025 | 0.003687 | 0.4811 | 1.3365 | 2.1379 | 0.7174
2012-10-30 21:27 | 28.5282 | 15.0328 | 15.8882 | 16.6693 | 15.2651 | -1.4911 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006958 | 0.21476 | 0.001494 | 30.8649 | 14.515 | -0.1028 | 0.003695 | 0.5178 | 1.3732 | 2.1543 | 0.7501
2012-10-30 21:28 | 28.5274 | 15.0117 | 15.863 | 16.701 | 15.3091 | -1.476 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 14.5346 | -0.1018 | 0.003691 | 0.4771 | 1.3284 | 2.1664 | 0.7745
2012-10-30 21:28 | 28.5341 | 15.0428 | 15.8615 | 16.6712 | 15.2792 | -1.4882 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.5209 | -0.1026 | 0.003689 | 0.5219 | 1.3406 | 2.1503 | 0.7583
2012-10-30 21:28 | 28.5222 | 15.0268 | 15.8578 | 16.6227 | 15.2469 | -1.4835 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.5009 | -0.1023 | 0.003686 | 0.5259 | 1.3569 | 2.1218 | 0.746
2012-10-30 21:28 | 28.5199 | 15.049 | 15.8677 | 16.6488 | 15.265 | -1.4876 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 14.5434 | -0.1026 | 0.003686 | 0.5056 | 1.3243 | 2.1054 | 0.7216
2012-10-30 21:28 | 28.5229 | 15.0561 | 15.8707 | 16.6396 | 15.2109 | -1.4858 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006939 | 0.21466 | 0.001489 | 30.9347 | 14.526 | -0.1024 | 0.003685 | 0.5301 | 1.3447 | 2.1136 | 0.6849
2012-10-30 21:28 | 28.5222 | 15.0472 | 15.8904 | 16.6633 | 15.2469 | -1.472 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006921 | 0.21474 | 0.001486 | 31.0277 | 14.4927 | -0.1015 | 0.003683 | 0.5545 | 1.3977 | 2.1706 | 0.7542
2012-10-30 21:28 | 28.5184 | 15.0027 | 15.8296 | 16.6514 | 15.3042 | -1.4731 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006946 | 0.2147 | 0.001491 | 30.9099 | 14.493 | -0.1016 | 0.003686 | 0.5097 | 1.3366 | 2.1584 | 0.8112
2012-10-30 21:28 | 28.5199 | 15.0245 | 15.8636 | 16.6529 | 15.2609 | -1.4772 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.5108 | -0.1018 | 0.00369 | 0.5137 | 1.3528 | 2.1421 | 0.7501
2012-10-30 21:28 | 28.5192 | 15.0524 | 15.8385 | 16.6359 | 15.2154 | -1.4818 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006921 | 0.21474 | 0.001486 | 31.0277 | 14.4979 | -0.1022 | 0.003682 | 0.5545 | 1.3406 | 2.138 | 0.7175
2012-10-30 21:28 | 28.5139 | 15.0348 | 15.8495 | 16.6144 | 15.2345 | -1.4795 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 14.5252 | -0.102 | 0.003687 | 0.5096 | 1.3243 | 2.0892 | 0.7093
2012-10-30 21:28 | 28.5162 | 15.029 | 15.8233 | 16.6045 | 15.2368 | -1.4801 | 0.006719 | 0.32654 | 0.002194 | 48.5999 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5234 | -0.102 | 0.003683 | 0.5056 | 1.2999 | 2.0811 | 0.7134
2012-10-30 21:28 | 28.5139 | 14.9941 | 15.8373 | 16.5981 | 15.2508 | -1.4754 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.5374 | -0.1017 | 0.003685 | 0.4567 | 1.2999 | 2.0607 | 0.7134
2012-10-30 21:28 | 28.5139 | 15.0267 | 15.8291 | 16.6184 | 15.2427 | -1.4841 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.5497 | -0.1023 | 0.003689 | 0.477 | 1.2794 | 2.0687 | 0.693
2012-10-30 21:28 | 28.5207 | 14.9927 | 15.8237 | 16.6334 | 15.2821 | -1.4841 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.532 | -0.1023 | 0.003689 | 0.4607 | 1.2917 | 2.1014 | 0.7501
2012-10-30 21:28 | 28.5237 | 14.9917 | 15.8552 | 16.6364 | 15.2443 | -1.4882 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.535 | -0.1026 | 0.003688 | 0.4567 | 1.3202 | 2.1014 | 0.7093
2012-10-30 21:28 | 28.5199 | 14.996 | 15.827 | 16.6366 | 15.2242 | -1.4783 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5312 | -0.1019 | 0.003691 | 0.4648 | 1.2958 | 2.1054 | 0.693
2012-10-30 21:28 | 28.5132 | 14.9812 | 15.8121 | 16.6259 | 15.2338 | -1.4899 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 14.5163 | -0.1027 | 0.003687 | 0.4649 | 1.2958 | 2.1096 | 0.7175
2012-10-30 21:28 | 28.5139 | 15.0104 | 15.8169 | 16.6062 | 15.2671 | -1.4882 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.00693 | 0.2147 | 0.001488 | 30.9812 | 14.5333 | -0.1026 | 0.003685 | 0.4771 | 1.2836 | 2.0729 | 0.7338
2012-10-30 21:28 | 28.5154 | 15.0078 | 15.8103 | 16.6077 | 15.2238 | -1.4754 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006921 | 0.21474 | 0.001486 | 31.0277 | 14.4737 | -0.1017 | 0.003685 | 0.5341 | 1.3366 | 2.134 | 0.7501
2012-10-30 21:28 | 28.5124 | 14.9559 | 15.8276 | 16.5844 | 15.2534 | -1.4853 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 14.5359 | -0.1024 | 0.003691 | 0.42 | 1.2917 | 2.0485 | 0.7175
2012-10-30 21:28 | 28.5169 | 15.0093 | 15.7996 | 16.6377 | 15.266 | -1.4893 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006949 | 0.21482 | 0.001493 | 30.9131 | 14.4874 | -0.1027 | 0.003692 | 0.5219 | 1.3122 | 2.1503 | 0.7786
2012-10-30 21:28 | 28.5154 | 14.9874 | 15.8225 | 16.5752 | 15.2523 | -1.4818 | 0.006716 | 0.32657 | 0.002193 | 48.626 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.5471 | -0.1022 | 0.003682 | 0.4403 | 1.2754 | 2.0281 | 0.7052
2012-10-30 21:28 | 28.5147 | 14.9827 | 15.8259 | 16.6111 | 15.2312 | -1.4835 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.5056 | -0.1023 | 0.003687 | 0.4771 | 1.3203 | 2.1055 | 0.7256
2012-10-30 21:29 | 28.5102 | 14.9415 | 15.8132 | 16.5537 | 15.2023 | -1.4824 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 14.5011 | -0.1022 | 0.003689 | 0.4404 | 1.3121 | 2.0526 | 0.7012
2012-10-30 21:29 | 28.5102 | 14.9904 | 15.8091 | 16.5619 | 15.239 | -1.4853 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.4807 | -0.1024 | 0.003692 | 0.5097 | 1.3284 | 2.0812 | 0.7583
2012-10-30 21:29 | 28.5117 | 14.9919 | 15.8025 | 16.5674 | 15.2038 | -1.476 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 14.4904 | -0.1018 | 0.003689 | 0.5015 | 1.3121 | 2.077 | 0.7134
2012-10-30 21:29 | 28.5124 | 14.9518 | 15.7869 | 16.5559 | 15.1963 | -1.4864 | 0.006722 | 0.32659 | 0.002195 | 48.5846 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.4584 | -0.1025 | 0.003686 | 0.4934 | 1.3285 | 2.0975 | 0.7379
2012-10-30 21:29 | 28.5169 | 14.9563 | 15.8077 | 16.5604 | 15.2497 | -1.4795 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.4874 | -0.102 | 0.003692 | 0.4689 | 1.3203 | 2.073 | 0.7623
2012-10-30 21:29 | 28.5139 | 14.9615 | 15.7803 | 16.5452 | 15.2264 | -1.4968 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.4966 | -0.1032 | 0.003687 | 0.4649 | 1.2837 | 2.0486 | 0.7298
2012-10-30 21:29 | 28.5154 | 14.9467 | 15.7981 | 16.5874 | 15.2442 | -1.4905 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.5226 | -0.1028 | 0.003691 | 0.4241 | 1.2755 | 2.0648 | 0.7216
2012-10-30 21:29 | 28.5184 | 14.9374 | 15.8092 | 16.5253 | 15.239 | -1.476 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 14.4889 | -0.1018 | 0.003686 | 0.4485 | 1.3203 | 2.0364 | 0.7501
2012-10-30 21:29 | 28.5177 | 14.9204 | 15.7556 | 16.5572 | 15.1854 | -1.4835 | 0.006732 | 0.32657 | 0.002198 | 48.5104 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.5168 | -0.1023 | 0.003687 | 0.4036 | 1.2388 | 2.0404 | 0.6686
2012-10-30 21:29 | 28.5222 | 14.9779 | 15.7967 | 16.5535 | 15.2021 | -1.487 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 14.509 | -0.1025 | 0.003686 | 0.4689 | 1.2877 | 2.0445 | 0.6931
2012-10-30 21:29 | 28.5207 | 14.9438 | 15.7504 | 16.5154 | 15.2454 | -1.4858 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.5238 | -0.1024 | 0.003689 | 0.42 | 1.2266 | 1.9916 | 0.7216
2012-10-30 21:29 | 28.5109 | 14.9381 | 15.7691 | 16.53 | 15.1867 | -1.487 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.4977 | -0.1025 | 0.003685 | 0.4404 | 1.2714 | 2.0323 | 0.689
2012-10-30 21:29 | 28.5274 | 14.9464 | 15.7938 | 16.5465 | 15.1991 | -1.4882 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.5305 | -0.1026 | 0.003692 | 0.4159 | 1.2633 | 2.016 | 0.6686
2012-10-30 21:29 | 28.5237 | 14.9183 | 15.7779 | 16.5347 | 15.1954 | -1.487 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.5309 | -0.1025 | 0.003686 | 0.3874 | 1.247 | 2.0038 | 0.6645
2012-10-30 21:29 | 28.5192 | 14.9097 | 15.7734 | 16.4895 | 15.1624 | -1.4841 | 0.006719 | 0.32654 | 0.002194 | 48.5999 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5101 | -0.1023 | 0.003686 | 0.3996 | 1.2633 | 1.9794 | 0.6523
2012-10-30 21:29 | 28.5184 | 14.9782 | 15.7889 | 16.5416 | 15.1901 | -1.4841 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.4726 | -0.1023 | 0.003682 | 0.5056 | 1.3163 | 2.069 | 0.7175
2012-10-30 21:29 | 28.5169 | 14.9563 | 15.8036 | 16.5279 | 15.209 | -1.4922 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.4752 | -0.1029 | 0.003684 | 0.4811 | 1.3284 | 2.0527 | 0.7338
2012-10-30 21:29 | 28.5252 | 14.9687 | 15.7712 | 16.5769 | 15.1847 | -1.483 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4998 | -0.1022 | 0.003684 | 0.4689 | 1.2714 | 2.0771 | 0.6849
2012-10-30 21:29 | 28.5274 | 14.9138 | 15.8141 | 16.5425 | 15.2113 | -1.4766 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.5101 | -0.1018 | 0.003686 | 0.4037 | 1.304 | 2.0324 | 0.7012
2012-10-30 21:29 | 28.5222 | 14.9005 | 15.7519 | 16.4803 | 15.2102 | -1.4858 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.4805 | -0.1024 | 0.003686 | 0.42 | 1.2714 | 1.9998 | 0.7297
2012-10-30 21:29 | 28.5244 | 14.8823 | 15.7541 | 16.5313 | 15.2083 | -1.4841 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.499 | -0.1023 | 0.003686 | 0.3833 | 1.2551 | 2.0323 | 0.7093
2012-10-30 21:29 | 28.5252 | 14.9442 | 15.7509 | 16.5037 | 15.1725 | -1.4864 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.512 | -0.1025 | 0.003686 | 0.4322 | 1.2389 | 1.9917 | 0.6605
2012-10-30 21:29 | 28.5334 | 14.9198 | 15.7672 | 16.5159 | 15.1522 | -1.4847 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.5121 | -0.1024 | 0.003687 | 0.4077 | 1.2551 | 2.0038 | 0.6401
2012-10-30 21:29 | 28.5229 | 14.893 | 15.7526 | 16.4729 | 15.1498 | -1.4922 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 14.4608 | -0.1029 | 0.003686 | 0.4322 | 1.2918 | 2.0121 | 0.689
2012-10-30 21:29 | 28.5214 | 14.9282 | 15.7715 | 16.4999 | 15.1931 | -1.4922 | 0.006716 | 0.32653 | 0.002193 | 48.6196 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.4797 | -0.1029 | 0.003682 | 0.4485 | 1.2918 | 2.0202 | 0.7134
2012-10-30 21:30 | 28.5334 | 14.928 | 15.7876 | 16.5281 | 15.2173 | -1.4876 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 14.5284 | -0.1026 | 0.003688 | 0.3996 | 1.2592 | 1.9997 | 0.6889
2012-10-30 21:30 | 28.5259 | 14.8919 | 15.7679 | 16.484 | 15.1813 | -1.4905 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.00693 | 0.21483 | 0.001489 | 31 | 14.5168 | -0.1028 | 0.003689 | 0.3751 | 1.2511 | 1.9672 | 0.6645
2012-10-30 21:30 | 28.5364 | 14.8984 | 15.815 | 16.5189 | 15.1715 | -1.4905 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006924 | 0.2147 | 0.001487 | 31.0081 | 14.5028 | -0.1028 | 0.00368 | 0.3956 | 1.3122 | 2.0161 | 0.6687
2012-10-30 21:30 | 28.5259 | 14.8879 | 15.7719 | 16.4556 | 15.1406 | -1.4812 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 14.5453 | -0.1021 | 0.003688 | 0.3426 | 1.2266 | 1.9103 | 0.5953
2012-10-30 21:30 | 28.5319 | 14.8531 | 15.7861 | 16.486 | 15.1547 | -1.4882 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 14.5187 | -0.1026 | 0.003684 | 0.3344 | 1.2674 | 1.9673 | 0.636
2012-10-30 21:30 | 28.5274 | 14.8894 | 15.7694 | 16.4937 | 15.2073 | -1.4928 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4979 | -0.1029 | 0.003686 | 0.3915 | 1.2715 | 1.9958 | 0.7094
2012-10-30 21:30 | 28.5222 | 14.8964 | 15.756 | 16.4519 | 15.1654 | -1.4812 | 0.006716 | 0.32651 | 0.002193 | 48.6173 | 0.006917 | 0.21464 | 0.001485 | 31.0311 | 14.562 | -0.1021 | 0.003678 | 0.3344 | 1.194 | 1.8899 | 0.6034
2012-10-30 21:30 | 28.5207 | 14.8704 | 15.7504 | 16.4463 | 15.1354 | -1.4899 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006917 | 0.21468 | 0.001485 | 31.0373 | 14.5075 | -0.1027 | 0.003685 | 0.3629 | 1.2429 | 1.9388 | 0.6279
2012-10-30 21:30 | 28.5222 | 14.8842 | 15.7601 | 16.4681 | 15.1573 | -1.4853 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006917 | 0.21464 | 0.001485 | 31.0311 | 14.5009 | -0.1024 | 0.003684 | 0.3833 | 1.2592 | 1.9672 | 0.6564
2012-10-30 21:30 | 28.5289 | 14.8909 | 15.7546 | 16.4586 | 15.1966 | -1.4847 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006917 | 0.21473 | 0.001485 | 31.0437 | 14.5198 | -0.1024 | 0.003679 | 0.3711 | 1.2348 | 1.9388 | 0.6768
2012-10-30 21:30 | 28.5259 | 14.8593 | 15.7597 | 16.5003 | 15.165 | -1.4795 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 14.4719 | -0.102 | 0.003683 | 0.3874 | 1.2878 | 2.0284 | 0.6931
2012-10-30 21:30 | 28.5184 | 14.8315 | 15.74 | 16.5091 | 15.1738 | -1.4876 | 0.006747 | 0.3265 | 0.002203 | 48.3919 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.4889 | -0.1026 | 0.003692 | 0.3426 | 1.2511 | 2.0202 | 0.6849
2012-10-30 21:30 | 28.5192 | 14.9138 | 15.7612 | 16.4529 | 15.1706 | -1.4928 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006905 | 0.21473 | 0.001483 | 31.0976 | 14.506 | -0.1029 | 0.003681 | 0.4078 | 1.2552 | 1.9469 | 0.6646
2012-10-30 21:30 | 28.5124 | 14.8662 | 15.7259 | 16.4461 | 15.1312 | -1.4963 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006914 | 0.21477 | 0.001485 | 31.0633 | 14.487 | -0.1032 | 0.003682 | 0.3792 | 1.2389 | 1.9591 | 0.6442
2012-10-30 21:30 | 28.5154 | 14.8855 | 15.7574 | 16.4776 | 15.179 | -1.5009 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006927 | 0.21468 | 0.001487 | 30.9925 | 14.4941 | -0.1035 | 0.003686 | 0.3914 | 1.2633 | 1.9835 | 0.6849
2012-10-30 21:30 | 28.5192 | 14.8445 | 15.7449 | 16.4773 | 15.1787 | -1.4795 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 14.4897 | -0.102 | 0.003686 | 0.3548 | 1.2552 | 1.9876 | 0.689
2012-10-30 21:30 | 28.5109 | 14.8321 | 15.7488 | 16.4975 | 15.1663 | -1.487 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.4488 | -0.1025 | 0.003686 | 0.3833 | 1.3 | 2.0487 | 0.7175
2012-10-30 21:30 | 28.5102 | 14.8681 | 15.7196 | 16.4724 | 15.1779 | -1.4934 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.4807 | -0.103 | 0.003687 | 0.3874 | 1.2389 | 1.9917 | 0.6972
2012-10-30 21:30 | 28.5132 | 14.8589 | 15.7307 | 16.4673 | 15.1564 | -1.4864 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5163 | -0.1025 | 0.003687 | 0.3426 | 1.2144 | 1.951 | 0.6401
2012-10-30 21:30 | 28.5237 | 14.8979 | 15.7046 | 16.4534 | 15.1221 | -1.4951 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5146 | -0.1031 | 0.003688 | 0.3833 | 1.19 | 1.9388 | 0.6075
2012-10-30 21:30 | 28.5177 | 14.8797 | 15.7556 | 16.4352 | 15.1731 | -1.4824 | 0.006719 | 0.3265 | 0.002194 | 48.5935 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 14.5168 | -0.1022 | 0.003682 | 0.3629 | 1.2388 | 1.9184 | 0.6563
2012-10-30 21:30 | 28.5079 | 14.8373 | 15.7254 | 16.405 | 15.1307 | -1.4812 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.5151 | -0.1021 | 0.003686 | 0.3222 | 1.2103 | 1.8899 | 0.6156
2012-10-30 21:30 | 28.5102 | 14.8436 | 15.7196 | 16.4358 | 15.1249 | -1.4899 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4644 | -0.1027 | 0.003691 | 0.3792 | 1.2552 | 1.9714 | 0.6605
2012-10-30 21:30 | 28.5214 | 14.8711 | 15.7389 | 16.4592 | 15.132 | -1.4905 | 0.006707 | 0.32649 | 0.00219 | 48.6782 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4715 | -0.1028 | 0.00368 | 0.3996 | 1.2674 | 1.9877 | 0.6605
2012-10-30 21:30 | 28.5102 | 14.8273 | 15.7033 | 16.4358 | 15.1453 | -1.4905 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 14.4725 | -0.1028 | 0.003684 | 0.3548 | 1.2308 | 1.9633 | 0.6728
2012-10-30 21:31 | 28.5207 | 14.8623 | 15.7057 | 16.4422 | 15.1272 | -1.4847 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 14.4912 | -0.1024 | 0.003691 | 0.3711 | 1.2145 | 1.951 | 0.636
2012-10-30 21:31 | 28.5222 | 14.8434 | 15.6949 | 16.4315 | 15.1532 | -1.4876 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.4397 | -0.1026 | 0.003686 | 0.4037 | 1.2552 | 1.9918 | 0.7135
2012-10-30 21:31 | 28.5109 | 14.8566 | 15.6714 | 16.3795 | 15.1011 | -1.4853 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 14.4406 | -0.1024 | 0.003692 | 0.416 | 1.2308 | 1.9389 | 0.6605
2012-10-30 21:31 | 28.5147 | 14.8848 | 15.6874 | 16.4037 | 15.1375 | -1.4812 | 0.006719 | 0.3265 | 0.002194 | 48.5935 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5097 | -0.1021 | 0.003684 | 0.3751 | 1.1777 | 1.894 | 0.6278
2012-10-30 21:31 | 28.5109 | 14.8443 | 15.7162 | 16.4162 | 15.1215 | -1.4939 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.4855 | -0.103 | 0.003685 | 0.3588 | 1.2307 | 1.9307 | 0.636
2012-10-30 21:31 | 28.5087 | 14.8381 | 15.6814 | 16.4018 | 15.1601 | -1.4893 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 14.4955 | -0.1027 | 0.003691 | 0.3426 | 1.1859 | 1.9063 | 0.6646
2012-10-30 21:31 | 28.5139 | 14.831 | 15.7355 | 16.411 | 15.1775 | -1.4841 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.5007 | -0.1023 | 0.003691 | 0.3303 | 1.2348 | 1.9103 | 0.6768
2012-10-30 21:31 | 28.5169 | 14.83 | 15.7181 | 16.3774 | 15.1397 | -1.4795 | 0.006751 | 0.32647 | 0.002204 | 48.3589 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.4996 | -0.102 | 0.003692 | 0.3304 | 1.2185 | 1.8778 | 0.6401
2012-10-30 21:31 | 28.5199 | 14.8411 | 15.7334 | 16.3763 | 15.1224 | -1.4893 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.5026 | -0.1027 | 0.003682 | 0.3385 | 1.2308 | 1.8737 | 0.6198
2012-10-30 21:31 | 28.5154 | 14.8285 | 15.7166 | 16.4044 | 15.1586 | -1.4951 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.5389 | -0.1031 | 0.003687 | 0.2896 | 1.1777 | 1.8655 | 0.6197
2012-10-30 21:31 | 28.5199 | 14.8289 | 15.713 | 16.3926 | 15.1509 | -1.4824 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006924 | 0.21486 | 0.001488 | 31.0311 | 14.5597 | -0.1022 | 0.003683 | 0.2692 | 1.1533 | 1.8329 | 0.5912
2012-10-30 21:31 | 28.5177 | 14.8348 | 15.6619 | 16.3579 | 15.1568 | -1.4899 | 0.006744 | 0.32653 | 0.002202 | 48.4177 | 0.006921 | 0.21476 | 0.001486 | 31.0299 | 14.5127 | -0.1027 | 0.003688 | 0.3221 | 1.1492 | 1.8452 | 0.6441
2012-10-30 21:31 | 28.5117 | 14.8207 | 15.7007 | 16.3722 | 15.1345 | -1.4939 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 14.5271 | -0.103 | 0.003686 | 0.2936 | 1.1736 | 1.8451 | 0.6074
2012-10-30 21:31 | 28.5132 | 14.8385 | 15.7104 | 16.3737 | 15.132 | -1.4864 | 0.006725 | 0.32659 | 0.002196 | 48.563 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.5286 | -0.1025 | 0.003684 | 0.3099 | 1.1818 | 1.8451 | 0.6034
2012-10-30 21:31 | 28.5132 | 14.8059 | 15.7104 | 16.333 | 15.1279 | -1.4922 | 0.006713 | 0.32649 | 0.002192 | 48.6347 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.5204 | -0.1029 | 0.003682 | 0.2855 | 1.19 | 1.8126 | 0.6075
2012-10-30 21:31 | 28.5072 | 14.7754 | 15.6596 | 16.4043 | 15.0933 | -1.4887 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 14.5185 | -0.1026 | 0.003691 | 0.2569 | 1.1411 | 1.8858 | 0.5748
2012-10-30 21:31 | 28.5154 | 14.8121 | 15.7126 | 16.3881 | 15.1138 | -1.4928 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.543 | -0.1029 | 0.003691 | 0.2691 | 1.1696 | 1.8451 | 0.5708
2012-10-30 21:31 | 28.5072 | 14.8284 | 15.6718 | 16.3555 | 15.1586 | -1.4945 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006943 | 0.2147 | 0.001491 | 30.9232 | 14.5063 | -0.103 | 0.003692 | 0.3221 | 1.1655 | 1.8492 | 0.6523
2012-10-30 21:31 | 28.5124 | 14.7806 | 15.6933 | 16.3648 | 15.1352 | -1.4905 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006946 | 0.21477 | 0.001492 | 30.9202 | 14.5318 | -0.1028 | 0.003692 | 0.2488 | 1.1615 | 1.833 | 0.6034
2012-10-30 21:31 | 28.5162 | 14.8252 | 15.7052 | 16.3645 | 15.1431 | -1.4887 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5601 | -0.1026 | 0.003687 | 0.2651 | 1.1451 | 1.8044 | 0.583
2012-10-30 21:31 | 28.5139 | 14.7943 | 15.6744 | 16.3825 | 15.0715 | -1.4928 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.517 | -0.1029 | 0.003684 | 0.2773 | 1.1574 | 1.8655 | 0.5545
2012-10-30 21:31 | 28.5132 | 14.8059 | 15.6656 | 16.3859 | 15.1401 | -1.4847 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.5326 | -0.1024 | 0.003687 | 0.2733 | 1.133 | 1.8533 | 0.6075
2012-10-30 21:31 | 28.5192 | 14.7589 | 15.7245 | 16.3431 | 15.1543 | -1.4945 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4775 | -0.103 | 0.00369 | 0.2814 | 1.247 | 1.8656 | 0.6768
2012-10-30 21:31 | 28.5079 | 14.7639 | 15.6888 | 16.3643 | 15.09 | -1.4893 | 0.006738 | 0.32644 | 0.0022 | 48.4479 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.4702 | -0.1027 | 0.003689 | 0.2937 | 1.2186 | 1.8941 | 0.6198
2012-10-30 21:31 | 28.5102 | 14.811 | 15.6789 | 16.3382 | 15.1127 | -1.5003 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 14.5623 | -0.1034 | 0.003685 | 0.2487 | 1.1166 | 1.7759 | 0.5504
2012-10-30 21:31 | 28.5147 | 14.7299 | 15.6915 | 16.3386 | 15.0968 | -1.4939 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.5586 | -0.103 | 0.003691 | 0.1713 | 1.1329 | 1.78 | 0.5382
2012-10-30 21:32 | 28.5117 | 14.7514 | 15.6519 | 16.3356 | 15.0938 | -1.4731 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5352 | -0.1016 | 0.003689 | 0.2162 | 1.1167 | 1.8004 | 0.5586
2012-10-30 21:32 | 28.5079 | 14.7843 | 15.6888 | 16.3521 | 15.1226 | -1.4934 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.4988 | -0.103 | 0.003689 | 0.2855 | 1.19 | 1.8533 | 0.6238
2012-10-30 21:32 | 28.5237 | 14.743 | 15.7087 | 16.3517 | 15.0854 | -1.4893 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.5187 | -0.1027 | 0.00369 | 0.2243 | 1.19 | 1.833 | 0.5667
2012-10-30 21:32 | 28.5132 | 14.7447 | 15.6819 | 16.3493 | 15.1116 | -1.4876 | 0.006722 | 0.32654 | 0.002195 | 48.5783 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.5326 | -0.1026 | 0.003682 | 0.2121 | 1.1493 | 1.8167 | 0.579
2012-10-30 21:32 | 28.5147 | 14.7503 | 15.6915 | 16.3711 | 15.1131 | -1.4939 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 14.4852 | -0.103 | 0.003683 | 0.2651 | 1.2063 | 1.8859 | 0.6279
2012-10-30 21:32 | 28.5079 | 14.7557 | 15.6684 | 16.3399 | 15.0614 | -1.4963 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006921 | 0.21476 | 0.001486 | 31.0299 | 14.507 | -0.1032 | 0.003685 | 0.2487 | 1.1614 | 1.8329 | 0.5544
2012-10-30 21:32 | 28.5117 | 14.7473 | 15.7048 | 16.3681 | 15.1671 | -1.4864 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.523 | -0.1025 | 0.00369 | 0.2243 | 1.1818 | 1.8451 | 0.6441
2012-10-30 21:32 | 28.5079 | 14.7557 | 15.6603 | 16.3643 | 15.1104 | -1.4893 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.511 | -0.1027 | 0.003691 | 0.2447 | 1.1493 | 1.8533 | 0.5994
2012-10-30 21:32 | 28.5027 | 14.7546 | 15.6877 | 16.3347 | 15.1011 | -1.4893 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.5344 | -0.1027 | 0.003692 | 0.2202 | 1.1533 | 1.8003 | 0.5667
2012-10-30 21:32 | 28.5064 | 14.7175 | 15.6628 | 16.3262 | 15.0966 | -1.4893 | 0.006747 | 0.32651 | 0.002203 | 48.394 | 0.006936 | 0.21467 | 0.001489 | 30.9503 | 14.5177 | -0.1027 | 0.003692 | 0.1998 | 1.1451 | 1.8085 | 0.5789
2012-10-30 21:32 | 28.5012 | 14.7653 | 15.7106 | 16.3414 | 15.067 | -1.4789 | 0.006744 | 0.32657 | 0.002202 | 48.4241 | 0.006952 | 0.2148 | 0.001493 | 30.8977 | 14.5247 | -0.102 | 0.003695 | 0.2406 | 1.1859 | 1.8167 | 0.5423
2012-10-30 21:32 | 28.5109 | 14.7506 | 15.7081 | 16.3348 | 15.1256 | -1.4818 | 0.006747 | 0.32649 | 0.002203 | 48.3897 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.5385 | -0.1022 | 0.003694 | 0.2121 | 1.1696 | 1.7963 | 0.5871
2012-10-30 21:32 | 28.5064 | 14.7909 | 15.6954 | 16.3222 | 15.1048 | -1.4847 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5258 | -0.1024 | 0.00369 | 0.2651 | 1.1696 | 1.7964 | 0.579
2012-10-30 21:32 | 28.5057 | 14.7494 | 15.6418 | 16.3337 | 15.1204 | -1.4824 | 0.006719 | 0.32651 | 0.002194 | 48.5956 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.5374 | -0.1022 | 0.003686 | 0.212 | 1.1044 | 1.7963 | 0.583
2012-10-30 21:32 | 28.5027 | 14.7342 | 15.6795 | 16.3347 | 15.1337 | -1.4818 | 0.006713 | 0.32657 | 0.002192 | 48.6477 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.4814 | -0.1022 | 0.00368 | 0.2528 | 1.1981 | 1.8533 | 0.6523
2012-10-30 21:32 | 28.5004 | 14.7238 | 15.608 | 16.2999 | 15.1069 | -1.4841 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.5158 | -0.1023 | 0.003689 | 0.208 | 1.0922 | 1.7841 | 0.5911
2012-10-30 21:32 | 28.4997 | 14.7394 | 15.6399 | 16.2911 | 15.0981 | -1.487 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.511 | -0.1025 | 0.003686 | 0.2284 | 1.1289 | 1.7801 | 0.5871
2012-10-30 21:33 | 28.4952 | 14.7634 | 15.6476 | 16.2784 | 15.0732 | -1.4858 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 14.5432 | -0.1024 | 0.003692 | 0.2202 | 1.1044 | 1.7352 | 0.53
2012-10-30 21:33 | 28.4982 | 14.7379 | 15.6099 | 16.2977 | 15.1088 | -1.4789 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.538 | -0.102 | 0.003689 | 0.1999 | 1.0719 | 1.7597 | 0.5708
2012-10-30 21:33 | 28.493 | 14.7653 | 15.625 | 16.2844 | 15.0832 | -1.4812 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.5532 | -0.1021 | 0.003686 | 0.2121 | 1.0718 | 1.7312 | 0.53
2012-10-30 21:33 | 28.4922 | 14.7074 | 15.6527 | 16.2876 | 15.1069 | -1.4743 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.5239 | -0.1016 | 0.003687 | 0.1835 | 1.1288 | 1.7637 | 0.583
2012-10-30 21:33 | 28.4952 | 14.7389 | 15.6476 | 16.3069 | 15.1221 | -1.4789 | 0.006751 | 0.32659 | 0.002205 | 48.3759 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.4453 | -0.102 | 0.003693 | 0.2936 | 1.2023 | 1.8616 | 0.6768
2012-10-30 21:33 | 28.4944 | 14.73 | 15.6346 | 16.2939 | 15.0846 | -1.4882 | 0.006757 | 0.32651 | 0.002206 | 48.3223 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 14.5057 | -0.1026 | 0.003695 | 0.2243 | 1.1289 | 1.7882 | 0.5789
2012-10-30 21:33 | 28.4944 | 14.7259 | 15.602 | 16.2939 | 15.0805 | -1.4835 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.5057 | -0.1023 | 0.003686 | 0.2202 | 1.0963 | 1.7882 | 0.5748
2012-10-30 21:33 | 28.4877 | 14.7396 | 15.5831 | 16.2547 | 15.0657 | -1.4766 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 14.499 | -0.1018 | 0.003684 | 0.2406 | 1.0841 | 1.7557 | 0.5667
2012-10-30 21:33 | 28.493 | 14.6837 | 15.6372 | 16.2722 | 15.1281 | -1.4853 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.4921 | -0.1024 | 0.003686 | 0.1916 | 1.1451 | 1.7801 | 0.636
2012-10-30 21:33 | 28.4989 | 14.7386 | 15.6268 | 16.2374 | 15.0606 | -1.5003 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006949 | 0.21482 | 0.001493 | 30.9131 | 14.4857 | -0.1034 | 0.003691 | 0.2529 | 1.1411 | 1.7517 | 0.5749
2012-10-30 21:33 | 28.5034 | 14.7023 | 15.6028 | 16.2785 | 15.0855 | -1.4916 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.4739 | -0.1028 | 0.003687 | 0.2284 | 1.1289 | 1.8046 | 0.6116
2012-10-30 21:33 | 28.4892 | 14.6962 | 15.5968 | 16.2399 | 15.059 | -1.4928 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.4719 | -0.1029 | 0.003686 | 0.2243 | 1.1249 | 1.768 | 0.5871
2012-10-30 21:33 | 28.5019 | 14.6886 | 15.6054 | 16.2322 | 15.084 | -1.4824 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.4846 | -0.1022 | 0.003682 | 0.204 | 1.1208 | 1.7476 | 0.5994
2012-10-30 21:33 | 28.4959 | 14.707 | 15.6116 | 16.2506 | 15.0576 | -1.4882 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.5113 | -0.1026 | 0.003689 | 0.1957 | 1.1003 | 1.7393 | 0.5463
2012-10-30 21:33 | 28.4967 | 14.6426 | 15.5839 | 16.2637 | 15.0421 | -1.4876 | 0.006744 | 0.32651 | 0.002202 | 48.4155 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.4876 | -0.1026 | 0.003692 | 0.155 | 1.0963 | 1.7761 | 0.5545
2012-10-30 21:33 | 28.5049 | 14.6589 | 15.5881 | 16.2352 | 15.0951 | -1.4887 | 0.006744 | 0.32651 | 0.002202 | 48.4155 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5366 | -0.1026 | 0.00369 | 0.1223 | 1.0515 | 1.6986 | 0.5585
2012-10-30 21:33 | 28.4974 | 14.6922 | 15.5724 | 16.2237 | 15.0713 | -1.4708 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.5046 | -0.1014 | 0.003688 | 0.1876 | 1.0678 | 1.7191 | 0.5667
2012-10-30 21:33 | 28.4982 | 14.7052 | 15.5691 | 16.2163 | 15.0477 | -1.4911 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 14.4605 | -0.1028 | 0.003685 | 0.2447 | 1.1086 | 1.7558 | 0.5872
2012-10-30 21:33 | 28.5012 | 14.6756 | 15.5925 | 16.2112 | 15.0588 | -1.4876 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 14.4839 | -0.1026 | 0.00369 | 0.1917 | 1.1086 | 1.7273 | 0.5749
2012-10-30 21:33 | 28.5049 | 14.7079 | 15.6003 | 16.2637 | 15.0788 | -1.4731 | 0.006716 | 0.32654 | 0.002193 | 48.6217 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.4999 | -0.1016 | 0.003685 | 0.208 | 1.1004 | 1.7638 | 0.5789
2012-10-30 21:33 | 28.4967 | 14.6793 | 15.5636 | 16.2474 | 15.0339 | -1.4887 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.5284 | -0.1026 | 0.003684 | 0.1509 | 1.0352 | 1.719 | 0.5055
2012-10-30 21:33 | 28.4982 | 14.6849 | 15.5651 | 16.2326 | 15.0477 | -1.4887 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5584 | -0.1026 | 0.003688 | 0.1265 | 1.0067 | 1.6742 | 0.4893
2012-10-30 21:33 | 28.5012 | 14.7082 | 15.5762 | 16.2031 | 15.0181 | -1.4766 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.59 | -0.1018 | 0.003683 | 0.1182 | 0.9862 | 1.6131 | 0.4281
2012-10-30 21:33 | 28.4989 | 14.6529 | 15.6228 | 16.2374 | 15.0728 | -1.4847 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5591 | -0.1024 | 0.003686 | 0.0938 | 1.0637 | 1.6783 | 0.5137
2012-10-30 21:33 | 28.4952 | 14.69 | 15.5906 | 16.2459 | 15.0528 | -1.4882 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5432 | -0.1026 | 0.003687 | 0.1468 | 1.0474 | 1.7027 | 0.5096
2012-10-30 21:33 | 28.5019 | 14.7049 | 15.6298 | 16.2282 | 15.0147 | -1.487 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.5213 | -0.1025 | 0.003687 | 0.1836 | 1.1085 | 1.7069 | 0.4934
2012-10-30 21:34 | 28.5004 | 14.6993 | 15.5998 | 16.2551 | 15.0743 | -1.4806 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.5525 | -0.1021 | 0.003686 | 0.1468 | 1.0473 | 1.7026 | 0.5218
2012-10-30 21:34 | 28.5004 | 14.683 | 15.5713 | 16.2023 | 15.0702 | -1.4945 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006911 | 0.21474 | 0.001484 | 31.0726 | 14.5851 | -0.103 | 0.003682 | 0.0979 | 0.9862 | 1.6172 | 0.4851
2012-10-30 21:34 | 28.5057 | 14.6638 | 15.6011 | 16.1831 | 15.1245 | -1.4905 | 0.006713 | 0.32651 | 0.002192 | 48.6391 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.5904 | -0.1028 | 0.003683 | 0.0734 | 1.0107 | 1.5927 | 0.5341
2012-10-30 21:34 | 28.4997 | 14.6537 | 15.591 | 16.2341 | 15.0614 | -1.4887 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.6293 | -0.1026 | 0.003691 | 0.0244 | 0.9617 | 1.6048 | 0.4321
2012-10-30 21:34 | 28.5147 | 14.6728 | 15.5775 | 16.1759 | 15.0723 | -1.4778 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.5953 | -0.1019 | 0.003687 | 0.0775 | 0.9822 | 1.5806 | 0.477
2012-10-30 21:34 | 28.5012 | 14.6634 | 15.6006 | 16.1705 | 15.0466 | -1.4801 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.5737 | -0.102 | 0.003686 | 0.0897 | 1.0269 | 1.5968 | 0.4729
2012-10-30 21:34 | 28.5034 | 14.637 | 15.5784 | 16.2215 | 15.061 | -1.4824 | 0.006747 | 0.32647 | 0.002203 | 48.3876 | 0.006921 | 0.21477 | 0.001486 | 31.0319 | 14.5963 | -0.1022 | 0.003689 | 0.0407 | 0.9821 | 1.6252 | 0.4647
2012-10-30 21:34 | 28.5124 | 14.6501 | 15.5752 | 16.198 | 15.0904 | -1.4876 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.6297 | -0.1026 | 0.003685 | 0.0204 | 0.9455 | 1.5683 | 0.4607
2012-10-30 21:34 | 28.5087 | 14.6668 | 15.5471 | 16.178 | 15.0948 | -1.4963 | 0.006747 | 0.3265 | 0.002203 | 48.3919 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.6097 | -0.1032 | 0.003691 | 0.0571 | 0.9374 | 1.5683 | 0.4851
2012-10-30 21:34 | 28.5169 | 14.6546 | 15.5715 | 16.1903 | 15.0501 | -1.4864 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.6179 | -0.1025 | 0.003685 | 0.0367 | 0.9536 | 1.5724 | 0.4322
2012-10-30 21:34 | 28.5154 | 14.6368 | 15.57 | 16.2213 | 15.0445 | -1.4824 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.5838 | -0.1022 | 0.003688 | 0.053 | 0.9862 | 1.6375 | 0.4607
2012-10-30 21:34 | 28.5147 | 14.6443 | 15.5938 | 16.1637 | 15.0682 | -1.4882 | 0.006719 | 0.3265 | 0.002194 | 48.5935 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.5994 | -0.1026 | 0.003683 | 0.0449 | 0.9944 | 1.5643 | 0.4688
2012-10-30 21:34 | 28.5177 | 14.6391 | 15.5968 | 16.1626 | 15.0712 | -1.4882 | 0.006735 | 0.32657 | 0.002199 | 48.4888 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.6065 | -0.1026 | 0.003687 | 0.0326 | 0.9903 | 1.5561 | 0.4647
2012-10-30 21:34 | 28.5252 | 14.614 | 15.5961 | 16.1782 | 15.0461 | -1.4876 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 14.6181 | -0.1026 | 0.00369 | -0.0041 | 0.978 | 1.5601 | 0.428
2012-10-30 21:34 | 28.5267 | 14.6073 | 15.5854 | 16.1716 | 15.0273 | -1.4766 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.6277 | -0.1018 | 0.00369 | -0.0204 | 0.9577 | 1.5439 | 0.3996
2012-10-30 21:34 | 28.5267 | 14.591 | 15.5691 | 16.1838 | 15.0802 | -1.4934 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.6155 | -0.103 | 0.003685 | -0.0245 | 0.9536 | 1.5683 | 0.4647
2012-10-30 21:34 | 28.5341 | 14.5943 | 15.5806 | 16.1464 | 15.0347 | -1.4853 | 0.006744 | 0.32647 | 0.002202 | 48.4091 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.6188 | -0.1024 | 0.003692 | -0.0245 | 0.9618 | 1.5276 | 0.4159
2012-10-30 21:34 | 28.5371 | 14.6218 | 15.6121 | 16.1576 | 15.0866 | -1.483 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 14.6055 | -0.1022 | 0.003686 | 0.0163 | 1.0066 | 1.5521 | 0.4811
2012-10-30 21:34 | 28.5394 | 14.6282 | 15.5696 | 16.2128 | 15.0889 | -1.4899 | 0.006754 | 0.32649 | 0.002205 | 48.3395 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.6323 | -0.1027 | 0.003694 | -0.0041 | 0.9373 | 1.5805 | 0.4566
2012-10-30 21:34 | 28.5371 | 14.5973 | 15.5592 | 16.1576 | 15.0377 | -1.4783 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.6218 | -0.1019 | 0.003686 | -0.0245 | 0.9374 | 1.5358 | 0.4159
2012-10-30 21:34 | 28.5394 | 14.6363 | 15.6144 | 16.1639 | 15.04 | -1.4697 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.6486 | -0.1013 | 0.003689 | -0.0123 | 0.9658 | 1.5153 | 0.3914
2012-10-30 21:34 | 28.5356 | 14.6285 | 15.578 | 16.1846 | 15.0565 | -1.4812 | 0.00671 | 0.3266 | 0.002191 | 48.6738 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.6407 | -0.1021 | 0.00368 | -0.0122 | 0.9373 | 1.5439 | 0.4158
2012-10-30 21:34 | 28.5446 | 14.6456 | 15.5504 | 16.1895 | 15.1022 | -1.4749 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.6497 | -0.1017 | 0.003693 | -0.0041 | 0.9007 | 1.5398 | 0.4525
2012-10-30 21:34 | 28.5379 | 14.6308 | 15.5966 | 16.1706 | 15.067 | -1.4795 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006927 | 0.21467 | 0.001487 | 30.9905 | 14.6104 | -0.102 | 0.003687 | 0.0204 | 0.9862 | 1.5602 | 0.4566
2012-10-30 21:34 | 28.5304 | 14.6314 | 15.5321 | 16.1631 | 15.0513 | -1.4887 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.6069 | -0.1026 | 0.00369 | 0.0245 | 0.9252 | 1.5562 | 0.4444
2012-10-30 21:34 | 28.5326 | 14.654 | 15.5547 | 16.1775 | 15.0128 | -1.4795 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 14.601 | -0.102 | 0.003686 | 0.053 | 0.9537 | 1.5765 | 0.4118
2012-10-30 21:35 | 28.5416 | 14.5978 | 15.5881 | 16.158 | 15.0503 | -1.4905 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.6263 | -0.1028 | 0.00369 | -0.0285 | 0.9618 | 1.5317 | 0.424
2012-10-30 21:35 | 28.5364 | 14.617 | 15.5625 | 16.165 | 15.0573 | -1.4858 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.6456 | -0.1024 | 0.003683 | -0.0286 | 0.9169 | 1.5194 | 0.4117
2012-10-30 21:35 | 28.5394 | 14.5915 | 15.5411 | 16.1802 | 15.0563 | -1.4974 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 14.6608 | -0.1032 | 0.003688 | -0.0693 | 0.8803 | 1.5194 | 0.3955
2012-10-30 21:35 | 28.5364 | 14.6089 | 15.5462 | 16.1732 | 15.0696 | -1.4818 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.6211 | -0.1022 | 0.003686 | -0.0122 | 0.9251 | 1.5521 | 0.4485
2012-10-30 21:35 | 28.5409 | 14.5767 | 15.5304 | 16.1044 | 15.0496 | -1.4812 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 14.6297 | -0.1021 | 0.003691 | -0.053 | 0.9007 | 1.4747 | 0.4199
2012-10-30 21:35 | 28.5476 | 14.5997 | 15.5452 | 16.1884 | 15.0522 | -1.4795 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 14.6038 | -0.102 | 0.003689 | -0.0041 | 0.9414 | 1.5846 | 0.4484
2012-10-30 21:35 | 28.5424 | 14.6271 | 15.5685 | 16.1263 | 15.0348 | -1.4795 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.5822 | -0.102 | 0.00369 | 0.0449 | 0.9863 | 1.5441 | 0.4526
2012-10-30 21:35 | 28.5461 | 14.6063 | 15.5804 | 16.1259 | 15.0222 | -1.4789 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006917 | 0.2147 | 0.001485 | 31.0395 | 14.6104 | -0.102 | 0.003684 | -0.0041 | 0.97 | 1.5155 | 0.4118
2012-10-30 21:35 | 28.5544 | 14.6024 | 15.5683 | 16.1464 | 15.055 | -1.4882 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.6065 | -0.1026 | 0.00369 | -0.0041 | 0.9618 | 1.5399 | 0.4485
2012-10-30 21:35 | 28.5491 | 14.5889 | 15.5752 | 16.1004 | 15.0171 | -1.4905 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.6093 | -0.1028 | 0.003687 | -0.0204 | 0.9659 | 1.4911 | 0.4078
2012-10-30 21:35 | 28.5529 | 14.6091 | 15.5587 | 16.153 | 15.0535 | -1.4882 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.6458 | -0.1026 | 0.00369 | -0.0367 | 0.9129 | 1.5072 | 0.4077
2012-10-30 21:35 | 28.5626 | 14.582 | 15.5602 | 16.1546 | 14.9939 | -1.4939 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.6595 | -0.103 | 0.003685 | -0.0775 | 0.9007 | 1.4951 | 0.3344
2012-10-30 21:35 | 28.5611 | 14.6132 | 15.5343 | 16.1816 | 15.0005 | -1.4778 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.6458 | -0.1019 | 0.003689 | -0.0326 | 0.8885 | 1.5358 | 0.3547
2012-10-30 21:35 | 28.5559 | 14.608 | 15.5291 | 16.1316 | 15.032 | -1.472 | 0.006744 | 0.3264 | 0.002201 | 48.3984 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.5998 | -0.1015 | 0.003689 | 0.0082 | 0.9293 | 1.5318 | 0.4322
2012-10-30 21:35 | 28.5566 | 14.5883 | 15.5502 | 16.1323 | 15.0531 | -1.4708 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 14.625 | -0.1014 | 0.003685 | -0.0367 | 0.9252 | 1.5073 | 0.4281
2012-10-30 21:35 | 28.5566 | 14.6372 | 15.5787 | 16.112 | 15.0083 | -1.4841 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.6168 | -0.1023 | 0.003684 | 0.0204 | 0.9619 | 1.4952 | 0.3915
2012-10-30 21:35 | 28.5634 | 14.5828 | 15.5284 | 16.1798 | 15.011 | -1.476 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.693 | -0.1018 | 0.003687 | -0.1102 | 0.8354 | 1.4868 | 0.318
2012-10-30 21:35 | 28.5649 | 14.5762 | 15.5381 | 16.1162 | 15.0206 | -1.4812 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.6292 | -0.1021 | 0.003692 | -0.053 | 0.9089 | 1.487 | 0.3914
2012-10-30 21:35 | 28.5701 | 14.6059 | 15.5026 | 16.1377 | 15.0747 | -1.483 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.5732 | -0.1022 | 0.00369 | 0.0327 | 0.9294 | 1.5645 | 0.5015
2012-10-30 21:35 | 28.5671 | 14.6151 | 15.5159 | 16.094 | 15.0147 | -1.483 | 0.006732 | 0.32659 | 0.002199 | 48.5125 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.6151 | -0.1022 | 0.003688 | 0 | 0.9008 | 1.4789 | 0.3996
2012-10-30 21:35 | 28.5656 | 14.5402 | 15.5469 | 16.0721 | 15.0662 | -1.4963 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.6218 | -0.1032 | 0.003689 | -0.0816 | 0.9251 | 1.4503 | 0.4444
2012-10-30 21:35 | 28.5746 | 14.59 | 15.503 | 16.1625 | 15.0263 | -1.483 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 14.6348 | -0.1022 | 0.003685 | -0.0448 | 0.8682 | 1.5277 | 0.3915
2012-10-30 21:35 | 28.5664 | 14.5573 | 15.4866 | 16.1096 | 15.0058 | -1.4795 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.6511 | -0.102 | 0.00369 | -0.0938 | 0.8355 | 1.4585 | 0.3547
2012-10-30 21:35 | 28.5693 | 14.6132 | 15.4977 | 16.1125 | 15.0291 | -1.4731 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 14.6418 | -0.1016 | 0.003683 | -0.0286 | 0.8559 | 1.4707 | 0.3873
2012-10-30 21:35 | 28.5836 | 14.5582 | 15.5364 | 16.1308 | 15.0679 | -1.483 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006952 | 0.21473 | 0.001493 | 30.8874 | 14.6398 | -0.1022 | 0.003691 | -0.0816 | 0.8966 | 1.491 | 0.4281
2012-10-30 21:35 | 28.5738 | 14.5525 | 15.5551 | 16.1251 | 15.0051 | -1.5003 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.6381 | -0.1034 | 0.003687 | -0.0856 | 0.917 | 1.487 | 0.367
2012-10-30 21:36 | 28.5873 | 14.5496 | 15.5401 | 16.1264 | 14.9737 | -1.4887 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.6516 | -0.1026 | 0.003686 | -0.102 | 0.8885 | 1.4748 | 0.3221
2012-10-30 21:36 | 28.5723 | 14.5632 | 15.5414 | 16.0829 | 15.0199 | -1.4818 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.5999 | -0.1022 | 0.00368 | -0.0367 | 0.9415 | 1.483 | 0.42
2012-10-30 21:36 | 28.5761 | 14.5752 | 15.5004 | 16.1437 | 15.0155 | -1.472 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.006936 | 0.2147 | 0.001489 | 30.9544 | 14.6323 | -0.1015 | 0.003683 | -0.0571 | 0.8681 | 1.5114 | 0.3832
2012-10-30 21:36 | 28.5791 | 14.5904 | 15.536 | 16.106 | 15.0185 | -1.4749 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.6312 | -0.1017 | 0.00368 | -0.0408 | 0.9048 | 1.4748 | 0.3873
2012-10-30 21:36 | 28.5753 | 14.5458 | 15.5689 | 16.0981 | 15.0106 | -1.4818 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006908 | 0.21476 | 0.001484 | 31.0883 | 14.5866 | -0.1022 | 0.003683 | -0.0408 | 0.9823 | 1.5115 | 0.424
2012-10-30 21:36 | 28.5851 | 14.5923 | 15.4931 | 16.1242 | 15.0327 | -1.4882 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 14.6413 | -0.1026 | 0.003686 | -0.049 | 0.8518 | 1.4829 | 0.3914
2012-10-30 21:36 | 28.5776 | 14.5195 | 15.506 | 16.0801 | 15.0007 | -1.4812 | 0.006741 | 0.32644 | 0.002201 | 48.4263 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 14.6623 | -0.1021 | 0.003689 | -0.1428 | 0.8437 | 1.4178 | 0.3384
2012-10-30 21:36 | 28.5753 | 14.5417 | 15.5281 | 16.0981 | 15.0229 | -1.4905 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.6478 | -0.1028 | 0.003688 | -0.1061 | 0.8803 | 1.4503 | 0.3751
2012-10-30 21:36 | 28.5836 | 14.5704 | 15.4998 | 16.0657 | 15.0556 | -1.4754 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.6642 | -0.1017 | 0.003681 | -0.0938 | 0.8356 | 1.4015 | 0.3914
2012-10-30 21:36 | 28.5873 | 14.5741 | 15.5116 | 16.0735 | 15.0553 | -1.483 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006936 | 0.21468 | 0.001489 | 30.9523 | 14.6475 | -0.1022 | 0.003688 | -0.0734 | 0.8641 | 1.426 | 0.4078
2012-10-30 21:36 | 28.5873 | 14.566 | 15.5279 | 16.0613 | 15.0145 | -1.4795 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006917 | 0.21474 | 0.001485 | 31.0457 | 14.6924 | -0.102 | 0.003683 | -0.1264 | 0.8355 | 1.3689 | 0.3221
2012-10-30 21:36 | 28.5881 | 14.5586 | 15.4961 | 16.0865 | 15.0316 | -1.4754 | 0.006719 | 0.32657 | 0.002194 | 48.6043 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.6891 | -0.1017 | 0.003685 | -0.1305 | 0.807 | 1.3974 | 0.3425
2012-10-30 21:36 | 28.5866 | 14.553 | 15.515 | 16.0565 | 14.9608 | -1.4858 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.6754 | -0.1024 | 0.003683 | -0.1224 | 0.8396 | 1.3811 | 0.2854
2012-10-30 21:36 | 28.5888 | 14.5511 | 15.5131 | 16.0953 | 15.0119 | -1.4841 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006955 | 0.21476 | 0.001494 | 30.8782 | 14.7061 | -0.1023 | 0.003692 | -0.155 | 0.807 | 1.3892 | 0.3058
2012-10-30 21:36 | 28.5888 | 14.5185 | 15.5457 | 16.0099 | 14.9875 | -1.4864 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.645 | -0.1025 | 0.003688 | -0.1265 | 0.9007 | 1.3649 | 0.3425
2012-10-30 21:36 | 28.5963 | 14.5342 | 15.5206 | 16.0662 | 15.048 | -1.4922 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.6647 | -0.1029 | 0.003688 | -0.1305 | 0.8559 | 1.4015 | 0.3833
2012-10-30 21:36 | 28.5993 | 14.5535 | 15.5318 | 16.0896 | 14.9572 | -1.4783 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006952 | 0.21476 | 0.001493 | 30.8915 | 14.6677 | -0.1019 | 0.003691 | -0.1142 | 0.8641 | 1.4219 | 0.2895
2012-10-30 21:36 | 28.6083 | 14.5176 | 15.4552 | 16.0823 | 14.9784 | -1.4939 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.64 | -0.103 | 0.003689 | -0.1224 | 0.8152 | 1.4423 | 0.3384
2012-10-30 21:36 | 28.6105 | 14.5769 | 15.5307 | 16.1211 | 15.0418 | -1.4812 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.7034 | -0.1021 | 0.00369 | -0.1265 | 0.8273 | 1.4177 | 0.3384
2012-10-30 21:36 | 28.6128 | 14.5466 | 15.533 | 16.0908 | 15.04 | -1.4853 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006949 | 0.21468 | 0.001492 | 30.8944 | 14.6934 | -0.1024 | 0.003689 | -0.1468 | 0.8396 | 1.3974 | 0.3466
2012-10-30 21:36 | 28.6143 | 14.5114 | 15.5101 | 16.0883 | 15.0537 | -1.4864 | 0.006738 | 0.32638 | 0.002199 | 48.4393 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.6664 | -0.1025 | 0.003688 | -0.155 | 0.8437 | 1.4219 | 0.3873
2012-10-30 21:36 | 28.618 | 14.5028 | 15.4934 | 16.0676 | 15.0085 | -1.4916 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006965 | 0.21471 | 0.001495 | 30.8276 | 14.7109 | -0.1028 | 0.003694 | -0.2081 | 0.7825 | 1.3567 | 0.2976
2012-10-30 21:36 | 28.615 | 14.5325 | 15.5149 | 16.0564 | 15.034 | -1.4795 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 14.6956 | -0.102 | 0.003691 | -0.1631 | 0.8193 | 1.3608 | 0.3384
2012-10-30 21:36 | 28.615 | 14.5366 | 15.5189 | 16.0198 | 14.9974 | -1.4858 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 14.6671 | -0.1024 | 0.003691 | -0.1305 | 0.8518 | 1.3527 | 0.3303
2012-10-30 21:36 | 28.6143 | 14.5155 | 15.4694 | 16.0191 | 15.017 | -1.4778 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.7153 | -0.1019 | 0.003688 | -0.1998 | 0.7541 | 1.3038 | 0.3017
2012-10-30 21:36 | 28.6285 | 14.5297 | 15.4795 | 16.0781 | 15.0231 | -1.4778 | 0.006713 | 0.32649 | 0.002192 | 48.6347 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.6602 | -0.1019 | 0.003681 | -0.1305 | 0.8193 | 1.4179 | 0.3629
2012-10-30 21:36 | 28.6218 | 14.5066 | 15.4972 | 16.0307 | 15.0245 | -1.4737 | 0.006703 | 0.32646 | 0.002188 | 48.703 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 14.7024 | -0.1016 | 0.003677 | -0.1958 | 0.7948 | 1.3283 | 0.3221
2012-10-30 21:36 | 28.6315 | 14.5245 | 15.4906 | 16.0526 | 15.0179 | -1.4783 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.6795 | -0.1019 | 0.003692 | -0.155 | 0.8111 | 1.3731 | 0.3384
2012-10-30 21:36 | 28.6308 | 14.5279 | 15.5755 | 16.0926 | 14.9968 | -1.4789 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.7196 | -0.102 | 0.003691 | -0.1917 | 0.8559 | 1.373 | 0.2772
2012-10-30 21:36 | 28.6368 | 14.5339 | 15.5082 | 16.0619 | 14.9825 | -1.4864 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 14.6766 | -0.1025 | 0.003691 | -0.1427 | 0.8316 | 1.3853 | 0.3059
2012-10-30 21:37 | 28.636 | 14.529 | 15.5196 | 16.0652 | 14.9898 | -1.4737 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006921 | 0.21468 | 0.001486 | 31.0194 | 14.631 | -0.1016 | 0.003679 | -0.102 | 0.8886 | 1.4342 | 0.3588
2012-10-30 21:37 | 28.6315 | 14.5123 | 15.5395 | 16.0485 | 15.0098 | -1.4737 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.655 | -0.1016 | 0.003685 | -0.1427 | 0.8845 | 1.3935 | 0.3548
2012-10-30 21:37 | 28.6323 | 14.5131 | 15.5037 | 16.0452 | 14.9983 | -1.4662 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.006946 | 0.21473 | 0.001492 | 30.9141 | 14.5865 | -0.1011 | 0.003688 | -0.0734 | 0.9172 | 1.4587 | 0.4118
2012-10-30 21:37 | 28.6368 | 14.4523 | 15.4796 | 16.0334 | 15.011 | -1.4708 | 0.006751 | 0.32653 | 0.002204 | 48.3675 | 0.006952 | 0.21479 | 0.001493 | 30.8956 | 14.6399 | -0.1014 | 0.003697 | -0.1876 | 0.8397 | 1.3935 | 0.3711
2012-10-30 21:37 | 28.6435 | 14.5161 | 15.4863 | 16.0157 | 14.9484 | -1.4824 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 14.6507 | -0.1022 | 0.003692 | -0.1346 | 0.8356 | 1.365 | 0.2977
2012-10-30 21:37 | 28.6353 | 14.4712 | 15.4822 | 16.0319 | 15.0136 | -1.4754 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.6344 | -0.1017 | 0.003682 | -0.1632 | 0.8478 | 1.3975 | 0.3792
2012-10-30 21:37 | 28.6405 | 14.5213 | 15.4793 | 16.0371 | 14.978 | -1.4766 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 14.7007 | -0.1018 | 0.003684 | -0.1794 | 0.7786 | 1.3364 | 0.2773
2012-10-30 21:37 | 28.6383 | 14.515 | 15.4934 | 16.0431 | 15.0207 | -1.4627 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 14.6129 | -0.1008 | 0.003685 | -0.0979 | 0.8805 | 1.4302 | 0.4078
2012-10-30 21:37 | 28.6435 | 14.5161 | 15.4986 | 16.0198 | 15.0136 | -1.4743 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.6833 | -0.1016 | 0.003685 | -0.1672 | 0.8153 | 1.3365 | 0.3303
2012-10-30 21:37 | 28.6457 | 14.5509 | 15.513 | 16.0139 | 14.9832 | -1.4662 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.6733 | -0.1011 | 0.003682 | -0.1224 | 0.8397 | 1.3406 | 0.3099
2012-10-30 21:37 | 28.6428 | 14.4909 | 15.4734 | 16.011 | 15.017 | -1.4708 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 14.6908 | -0.1014 | 0.003686 | -0.1999 | 0.7826 | 1.3202 | 0.3262
2012-10-30 21:37 | 28.6383 | 14.5272 | 15.5178 | 16.0105 | 15.0166 | -1.4754 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 14.6863 | -0.1017 | 0.003685 | -0.1591 | 0.8315 | 1.3242 | 0.3303
2012-10-30 21:37 | 28.6435 | 14.4794 | 15.4904 | 16.0076 | 15.0055 | -1.4801 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006914 | 0.21474 | 0.001485 | 31.0592 | 14.6629 | -0.102 | 0.003683 | -0.1835 | 0.8275 | 1.3447 | 0.3426
2012-10-30 21:37 | 28.6472 | 14.4668 | 15.4493 | 16.0276 | 14.9847 | -1.4702 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.6666 | -0.1014 | 0.003689 | -0.1998 | 0.7827 | 1.361 | 0.3181
2012-10-30 21:37 | 28.645 | 14.5094 | 15.443 | 16.0376 | 14.9907 | -1.4749 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.6726 | -0.1017 | 0.003687 | -0.1632 | 0.7704 | 1.365 | 0.3181
2012-10-30 21:37 | 28.6547 | 14.5069 | 15.522 | 16.0636 | 14.9718 | -1.4633 | 0.006722 | 0.32653 | 0.002195 | 48.5762 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.7435 | -0.1009 | 0.003682 | -0.2366 | 0.7785 | 1.3201 | 0.2283
2012-10-30 21:37 | 28.6487 | 14.456 | 15.516 | 15.9802 | 14.9903 | -1.4714 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.713 | -0.1014 | 0.00369 | -0.257 | 0.803 | 1.2672 | 0.2773
2012-10-30 21:37 | 28.6532 | 14.5054 | 15.4594 | 16.0295 | 15.0111 | -1.4645 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 14.7053 | -0.101 | 0.003686 | -0.1999 | 0.7541 | 1.3242 | 0.3058
2012-10-30 21:37 | 28.6525 | 14.4353 | 15.4953 | 15.984 | 14.9656 | -1.4697 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 14.6964 | -0.1013 | 0.003685 | -0.2611 | 0.7989 | 1.2876 | 0.2692
2012-10-30 21:37 | 28.651 | 14.5154 | 15.449 | 15.9744 | 15.013 | -1.4737 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 14.6908 | -0.1016 | 0.003686 | -0.1754 | 0.7582 | 1.2836 | 0.3222
2012-10-30 21:37 | 28.6525 | 14.5251 | 15.4465 | 16.0044 | 15.0185 | -1.4708 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 14.7209 | -0.1014 | 0.003689 | -0.1958 | 0.7256 | 1.2835 | 0.2976
2012-10-30 21:37 | 28.6592 | 14.491 | 15.4817 | 16.007 | 15.013 | -1.4795 | 0.006738 | 0.32657 | 0.0022 | 48.4672 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 14.7154 | -0.102 | 0.00369 | -0.2244 | 0.7663 | 1.2916 | 0.2976
2012-10-30 21:37 | 28.6577 | 14.4691 | 15.4965 | 16.0136 | 15.0115 | -1.4795 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.7179 | -0.102 | 0.003683 | -0.2488 | 0.7786 | 1.2957 | 0.2936
2012-10-30 21:37 | 28.654 | 14.498 | 15.4683 | 16.0059 | 14.9874 | -1.4691 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.7387 | -0.1013 | 0.003691 | -0.2407 | 0.7296 | 1.2672 | 0.2487
2012-10-30 21:37 | 28.6525 | 14.4476 | 15.4505 | 15.9759 | 14.9982 | -1.4726 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.725 | -0.1015 | 0.003687 | -0.2774 | 0.7255 | 1.2509 | 0.2732
2012-10-30 21:37 | 28.6615 | 14.4403 | 15.5043 | 15.9889 | 15.0112 | -1.4679 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.6769 | -0.1012 | 0.003688 | -0.2366 | 0.8274 | 1.312 | 0.3343
2012-10-30 21:37 | 28.6577 | 14.5181 | 15.4639 | 16.0299 | 14.9993 | -1.4685 | 0.006738 | 0.32644 | 0.0022 | 48.4479 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.722 | -0.1012 | 0.003688 | -0.2039 | 0.7419 | 1.3079 | 0.2773
2012-10-30 21:37 | 28.66 | 14.4551 | 15.4906 | 16.0485 | 14.9445 | -1.476 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006949 | 0.21476 | 0.001492 | 30.9049 | 14.708 | -0.1018 | 0.00369 | -0.2529 | 0.7826 | 1.3405 | 0.2365
2012-10-30 21:38 | 28.6577 | 14.5058 | 15.4761 | 15.9892 | 14.9789 | -1.4726 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.6731 | -0.1015 | 0.00369 | -0.1673 | 0.803 | 1.3161 | 0.3058
2012-10-30 21:38 | 28.6615 | 14.4811 | 15.4636 | 16.0256 | 15.0153 | -1.4714 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.6606 | -0.1014 | 0.003687 | -0.1795 | 0.803 | 1.365 | 0.3547
2012-10-30 21:38 | 28.6645 | 14.4759 | 15.4666 | 15.9757 | 14.9979 | -1.4645 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.6513 | -0.101 | 0.003693 | -0.1754 | 0.8153 | 1.3244 | 0.3466
2012-10-30 21:38 | 28.6615 | 14.4566 | 15.488 | 16.0093 | 14.9786 | -1.4639 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.6769 | -0.1009 | 0.003687 | -0.2203 | 0.8111 | 1.3324 | 0.3017
2012-10-30 21:38 | 28.6637 | 14.4588 | 15.4454 | 15.9586 | 14.9604 | -1.4726 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006946 | 0.21471 | 0.001491 | 30.9119 | 14.6872 | -0.1015 | 0.003689 | -0.2284 | 0.7582 | 1.2714 | 0.2732
2012-10-30 21:38 | 28.6547 | 14.4702 | 15.4364 | 15.9496 | 14.9555 | -1.4876 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 14.6578 | -0.1026 | 0.003686 | -0.1876 | 0.7786 | 1.2918 | 0.2977
2012-10-30 21:38 | 28.678 | 14.5098 | 15.4964 | 15.9973 | 14.987 | -1.4731 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.7178 | -0.1016 | 0.003685 | -0.208 | 0.7786 | 1.2795 | 0.2692
2012-10-30 21:38 | 28.6667 | 14.4332 | 15.5136 | 15.9819 | 15.0001 | -1.4645 | 0.006747 | 0.32647 | 0.002203 | 48.3876 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.6984 | -0.101 | 0.003691 | -0.2652 | 0.8152 | 1.2835 | 0.3017
2012-10-30 21:38 | 28.6637 | 14.4996 | 15.4943 | 16.0115 | 14.9768 | -1.465 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 14.675 | -0.101 | 0.003687 | -0.1754 | 0.8193 | 1.3365 | 0.3018
2012-10-30 21:38 | 28.6667 | 14.4251 | 15.5095 | 15.9575 | 14.9757 | -1.4662 | 0.006741 | 0.32644 | 0.002201 | 48.4263 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.6861 | -0.1011 | 0.003691 | -0.261 | 0.8234 | 1.2714 | 0.2896
2012-10-30 21:38 | 28.669 | 14.4233 | 15.4955 | 15.9679 | 14.9617 | -1.4726 | 0.006738 | 0.32654 | 0.0022 | 48.4629 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 14.6884 | -0.1015 | 0.003688 | -0.2651 | 0.8071 | 1.2795 | 0.2733
2012-10-30 21:38 | 28.6765 | 14.5001 | 15.499 | 15.9673 | 14.9896 | -1.4674 | 0.006738 | 0.32657 | 0.0022 | 48.4672 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.6919 | -0.1012 | 0.003691 | -0.1918 | 0.8071 | 1.2754 | 0.2977
2012-10-30 21:38 | 28.6712 | 14.4867 | 15.514 | 15.9294 | 14.9639 | -1.4714 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006952 | 0.21479 | 0.001493 | 30.8956 | 14.707 | -0.1014 | 0.003687 | -0.2203 | 0.807 | 1.2224 | 0.2569
2012-10-30 21:38 | 28.6735 | 14.4237 | 15.4878 | 15.9928 | 15.0192 | -1.4656 | 0.006744 | 0.32654 | 0.002202 | 48.4198 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 14.7215 | -0.101 | 0.003687 | -0.2978 | 0.7663 | 1.2713 | 0.2977
2012-10-30 21:38 | 28.66 | 14.4306 | 15.4743 | 15.9508 | 14.9486 | -1.4778 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.7202 | -0.1019 | 0.003683 | -0.2896 | 0.7541 | 1.2306 | 0.2284
2012-10-30 21:38 | 28.651 | 14.4257 | 15.5061 | 15.9784 | 14.9192 | -1.4766 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.7275 | -0.1018 | 0.003687 | -0.3018 | 0.7786 | 1.2509 | 0.1917
2012-10-30 21:38 | 28.6562 | 14.4023 | 15.4787 | 15.9836 | 14.9693 | -1.4708 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.745 | -0.1014 | 0.003687 | -0.3427 | 0.7337 | 1.2386 | 0.2243
2012-10-30 21:38 | 28.6577 | 14.4365 | 15.4313 | 15.9282 | 14.9952 | -1.4691 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.7628 | -0.1013 | 0.003688 | -0.3263 | 0.6685 | 1.1654 | 0.2324
2012-10-30 21:38 | 28.6675 | 14.4871 | 15.4778 | 15.9664 | 14.9969 | -1.4772 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21486 | 0.00149 | 30.9908 | 14.7318 | -0.1018 | 0.003685 | -0.2447 | 0.746 | 1.2346 | 0.2651
2012-10-30 21:38 | 28.6622 | 14.4165 | 15.4643 | 15.9571 | 15.016 | -1.4766 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.751 | -0.1018 | 0.003689 | -0.3345 | 0.7133 | 1.2061 | 0.265
2012-10-30 21:38 | 28.6645 | 14.4025 | 15.4625 | 15.996 | 14.9857 | -1.4754 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.7574 | -0.1017 | 0.003692 | -0.3549 | 0.7051 | 1.2386 | 0.2283
2012-10-30 21:38 | 28.6697 | 14.4199 | 15.5288 | 15.9564 | 14.9542 | -1.4737 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 14.7789 | -0.1016 | 0.003688 | -0.359 | 0.7499 | 1.1775 | 0.1753
2012-10-30 21:38 | 28.663 | 14.4295 | 15.514 | 15.966 | 14.9761 | -1.4702 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006921 | 0.21473 | 0.001486 | 31.0257 | 14.7926 | -0.1014 | 0.00368 | -0.3631 | 0.7214 | 1.1734 | 0.1835
2012-10-30 21:38 | 28.669 | 14.411 | 15.5648 | 15.972 | 14.9861 | -1.4645 | 0.006747 | 0.3265 | 0.002203 | 48.3919 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.8149 | -0.101 | 0.003691 | -0.4039 | 0.7499 | 1.1571 | 0.1712
2012-10-30 21:38 | 28.6645 | 14.431 | 15.5603 | 16.0001 | 14.9572 | -1.4726 | 0.006729 | 0.32657 | 0.002198 | 48.532 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 14.7981 | -0.1015 | 0.003685 | -0.3671 | 0.7622 | 1.202 | 0.1591
2012-10-30 21:38 | 28.66 | 14.4306 | 15.5028 | 15.9671 | 15.0057 | -1.4766 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 14.8018 | -0.1018 | 0.003687 | -0.3712 | 0.701 | 1.1653 | 0.2039
2012-10-30 21:38 | 28.66 | 14.4184 | 15.5028 | 15.9508 | 15.0057 | -1.4818 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 14.7733 | -0.1022 | 0.003685 | -0.3549 | 0.7295 | 1.1775 | 0.2324
2012-10-30 21:38 | 28.6615 | 14.4239 | 15.5206 | 15.9686 | 14.9501 | -1.476 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 14.7544 | -0.1018 | 0.003691 | -0.3305 | 0.7662 | 1.2142 | 0.1957
2012-10-30 21:38 | 28.6555 | 14.4465 | 15.4698 | 15.93 | 15.0134 | -1.4737 | 0.006751 | 0.32654 | 0.002204 | 48.3696 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.7647 | -0.1016 | 0.003695 | -0.3182 | 0.7051 | 1.1653 | 0.2487
2012-10-30 21:38 | 28.6585 | 14.4658 | 15.481 | 15.9412 | 14.9716 | -1.4731 | 0.006729 | 0.32659 | 0.002198 | 48.5341 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.7636 | -0.1016 | 0.003687 | -0.2978 | 0.7174 | 1.1776 | 0.208
2012-10-30 21:39 | 28.6555 | 14.4547 | 15.5106 | 15.9626 | 14.9319 | -1.4772 | 0.006741 | 0.32641 | 0.0022 | 48.422 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.7035 | -0.1018 | 0.003689 | -0.2488 | 0.8071 | 1.2591 | 0.2284
2012-10-30 21:39 | 28.6645 | 14.4188 | 15.4544 | 15.9716 | 14.9979 | -1.4766 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 14.7451 | -0.1018 | 0.003685 | -0.3263 | 0.7093 | 1.2265 | 0.2528
2012-10-30 21:39 | 28.6607 | 14.4068 | 15.4954 | 15.9271 | 14.9615 | -1.4714 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.00693 | 0.21483 | 0.001489 | 31 | 14.7576 | -0.1014 | 0.003686 | -0.3508 | 0.7378 | 1.1695 | 0.2039
2012-10-30 21:39 | 28.657 | 14.395 | 15.5121 | 15.9559 | 14.9456 | -1.4674 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 14.7417 | -0.1012 | 0.003687 | -0.3467 | 0.7704 | 1.2142 | 0.2039
2012-10-30 21:39 | 28.6525 | 14.4394 | 15.4994 | 15.9392 | 14.9737 | -1.4674 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.7127 | -0.1012 | 0.003688 | -0.2733 | 0.7867 | 1.2265 | 0.261
2012-10-30 21:39 | 28.657 | 14.4358 | 15.4754 | 15.9519 | 14.9578 | -1.4674 | 0.006741 | 0.32656 | 0.002201 | 48.4434 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.7825 | -0.1012 | 0.003693 | -0.3467 | 0.6929 | 1.1694 | 0.1753
2012-10-30 21:39 | 28.66 | 14.4224 | 15.5151 | 15.8979 | 14.9853 | -1.4789 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.7488 | -0.102 | 0.00369 | -0.3264 | 0.7663 | 1.1491 | 0.2365
2012-10-30 21:39 | 28.6495 | 14.4568 | 15.4679 | 15.9403 | 14.9666 | -1.4679 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.726 | -0.1012 | 0.003689 | -0.2692 | 0.7419 | 1.2143 | 0.2406
2012-10-30 21:39 | 28.6652 | 14.4358 | 15.4958 | 15.9723 | 14.9619 | -1.476 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006921 | 0.2147 | 0.001486 | 31.0215 | 14.7214 | -0.1018 | 0.003686 | -0.2856 | 0.7744 | 1.2509 | 0.2405
2012-10-30 21:39 | 28.6517 | 14.4264 | 15.3968 | 15.8977 | 15.0055 | -1.4656 | 0.006716 | 0.3265 | 0.002193 | 48.6152 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.7813 | -0.101 | 0.003683 | -0.3549 | 0.6155 | 1.1164 | 0.2242
2012-10-30 21:39 | 28.6622 | 14.4042 | 15.4928 | 15.9408 | 14.9753 | -1.4806 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.7306 | -0.1021 | 0.003685 | -0.3264 | 0.7622 | 1.2102 | 0.2447
2012-10-30 21:39 | 28.6592 | 14.4135 | 15.4613 | 15.9419 | 14.9763 | -1.4778 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.7561 | -0.1019 | 0.003685 | -0.3426 | 0.7052 | 1.1858 | 0.2202
2012-10-30 21:39 | 28.6615 | 14.4321 | 15.4758 | 15.9523 | 14.9419 | -1.4627 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006927 | 0.2147 | 0.001487 | 30.9947 | 14.7136 | -0.1008 | 0.003681 | -0.2815 | 0.7622 | 1.2387 | 0.2283
2012-10-30 21:39 | 28.666 | 14.457 | 15.5088 | 15.9039 | 14.975 | -1.4726 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 14.7629 | -0.1015 | 0.003689 | -0.3059 | 0.7459 | 1.141 | 0.2121
2012-10-30 21:39 | 28.6577 | 14.3875 | 15.468 | 15.9037 | 14.9463 | -1.461 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.7628 | -0.1007 | 0.003687 | -0.3753 | 0.7052 | 1.1409 | 0.1835
2012-10-30 21:39 | 28.6645 | 14.4065 | 15.4666 | 15.9146 | 15.0102 | -1.4668 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.7125 | -0.1011 | 0.003688 | -0.306 | 0.7541 | 1.2021 | 0.2977
2012-10-30 21:39 | 28.657 | 14.3664 | 15.4795 | 15.8827 | 14.9904 | -1.4708 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.7254 | -0.1014 | 0.003685 | -0.359 | 0.7541 | 1.1573 | 0.265
2012-10-30 21:39 | 28.6592 | 14.3808 | 15.4287 | 15.9378 | 14.9641 | -1.4633 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 14.7317 | -0.1009 | 0.003682 | -0.3509 | 0.697 | 1.2061 | 0.2324
2012-10-30 21:39 | 28.672 | 14.4018 | 15.4537 | 15.9262 | 14.9769 | -1.4841 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006911 | 0.21476 | 0.001484 | 31.0748 | 14.7282 | -0.1023 | 0.003683 | -0.3264 | 0.7255 | 1.198 | 0.2487
2012-10-30 21:39 | 28.672 | 14.3936 | 15.47 | 15.914 | 14.9361 | -1.4743 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 14.6588 | -0.1016 | 0.003684 | -0.2652 | 0.8112 | 1.2552 | 0.2773
2012-10-30 21:39 | 28.6757 | 14.3933 | 15.4656 | 15.8851 | 14.9806 | -1.4731 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.6951 | -0.1016 | 0.003687 | -0.3018 | 0.7705 | 1.19 | 0.2855
2012-10-30 21:39 | 28.672 | 14.414 | 15.4334 | 15.8895 | 14.9973 | -1.4864 | 0.006703 | 0.32647 | 0.002188 | 48.7052 | 0.006933 | 0.21486 | 0.00149 | 30.9908 | 14.6751 | -0.1025 | 0.003678 | -0.2611 | 0.7583 | 1.2144 | 0.3222
2012-10-30 21:39 | 28.672 | 14.3855 | 15.4089 | 15.8692 | 14.928 | -1.4714 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.6833 | -0.1014 | 0.003685 | -0.2978 | 0.7256 | 1.1859 | 0.2447
2012-10-30 21:39 | 28.6795 | 14.3277 | 15.4205 | 15.9052 | 14.9926 | -1.476 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.6663 | -0.1018 | 0.003683 | -0.3386 | 0.7542 | 1.2389 | 0.3263
2012-10-30 21:39 | 28.6802 | 14.3814 | 15.4619 | 15.8814 | 14.9566 | -1.4801 | 0.006719 | 0.32657 | 0.002194 | 48.6043 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.6793 | -0.102 | 0.003684 | -0.2979 | 0.7826 | 1.2021 | 0.2773
2012-10-30 21:39 | 28.6899 | 14.3748 | 15.4431 | 15.9278 | 14.9499 | -1.4679 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.7093 | -0.1012 | 0.003684 | -0.3345 | 0.7338 | 1.2185 | 0.2406
2012-10-30 21:39 | 28.6869 | 14.4167 | 15.4401 | 15.9004 | 14.9673 | -1.4858 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 14.7227 | -0.1024 | 0.003686 | -0.306 | 0.7174 | 1.1777 | 0.2446
2012-10-30 21:39 | 28.6839 | 14.4015 | 15.4738 | 15.9136 | 14.9358 | -1.4806 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.6789 | -0.1021 | 0.003687 | -0.2774 | 0.7949 | 1.2347 | 0.2569
2012-10-30 21:39 | 28.6899 | 14.3626 | 15.4431 | 15.9278 | 14.9459 | -1.4812 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 14.7257 | -0.1021 | 0.003686 | -0.3631 | 0.7174 | 1.2021 | 0.2202
2012-10-30 21:40 | 28.6765 | 14.3247 | 15.389 | 15.8574 | 14.908 | -1.4645 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.6878 | -0.101 | 0.003687 | -0.3631 | 0.7012 | 1.1696 | 0.2202
2012-10-30 21:40 | 28.6914 | 14.3355 | 15.4772 | 15.9171 | 14.9555 | -1.4749 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 14.6701 | -0.1017 | 0.003689 | -0.3346 | 0.8071 | 1.247 | 0.2854
2012-10-30 21:40 | 28.6862 | 14.3425 | 15.4272 | 15.9037 | 14.9626 | -1.4668 | 0.006741 | 0.32644 | 0.002201 | 48.4263 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 14.7301 | -0.1011 | 0.003692 | -0.3876 | 0.6971 | 1.1736 | 0.2325
2012-10-30 21:40 | 28.7019 | 14.4072 | 15.4307 | 15.9479 | 14.9619 | -1.4847 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 14.7377 | -0.1024 | 0.003687 | -0.3305 | 0.693 | 1.2102 | 0.2242
2012-10-30 21:40 | 28.6854 | 14.3458 | 15.4223 | 15.8785 | 14.9495 | -1.4726 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006939 | 0.21473 | 0.00149 | 30.9452 | 14.7048 | -0.1015 | 0.003689 | -0.359 | 0.7175 | 1.1737 | 0.2447
2012-10-30 21:40 | 28.6839 | 14.3158 | 15.3882 | 15.8933 | 14.9766 | -1.4702 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 14.6993 | -0.1014 | 0.003687 | -0.3835 | 0.6889 | 1.194 | 0.2773
2012-10-30 21:40 | 28.6967 | 14.3939 | 15.4336 | 15.9102 | 14.9853 | -1.4835 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.7365 | -0.1023 | 0.003687 | -0.3426 | 0.6971 | 1.1737 | 0.2488
2012-10-30 21:40 | 28.6929 | 14.3248 | 15.4502 | 15.9064 | 14.9693 | -1.4708 | 0.006741 | 0.32653 | 0.002201 | 48.4393 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.7083 | -0.1014 | 0.003691 | -0.3835 | 0.7419 | 1.1981 | 0.261
2012-10-30 21:40 | 28.6922 | 14.3853 | 15.4536 | 15.8975 | 14.9319 | -1.4737 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.7443 | -0.1016 | 0.003691 | -0.359 | 0.7093 | 1.1532 | 0.1876
2012-10-30 21:40 | 28.6884 | 14.3611 | 15.4335 | 15.91 | 14.9444 | -1.4593 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.7078 | -0.1006 | 0.00369 | -0.3467 | 0.7257 | 1.2022 | 0.2366
2012-10-30 21:40 | 28.6884 | 14.3447 | 15.4131 | 15.9019 | 14.8995 | -1.4656 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.6956 | -0.101 | 0.003686 | -0.3509 | 0.7175 | 1.2063 | 0.2039
2012-10-30 21:40 | 28.7027 | 14.3999 | 15.4192 | 15.8917 | 14.9342 | -1.4731 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 14.7262 | -0.1016 | 0.003683 | -0.3263 | 0.693 | 1.1655 | 0.208
2012-10-30 21:40 | 28.7012 | 14.3861 | 15.4544 | 15.8984 | 14.9368 | -1.4662 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.006908 | 0.21482 | 0.001484 | 31.0966 | 14.7492 | -0.1011 | 0.003685 | -0.3631 | 0.7052 | 1.1492 | 0.1876
2012-10-30 21:40 | 28.6952 | 14.3515 | 15.4036 | 15.9168 | 14.9104 | -1.4754 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.7146 | -0.1017 | 0.003693 | -0.3631 | 0.689 | 1.2022 | 0.1958
2012-10-30 21:40 | 28.6997 | 14.3683 | 15.4285 | 15.8724 | 14.9394 | -1.472 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006927 | 0.21473 | 0.001487 | 30.9988 | 14.7028 | -0.1015 | 0.003684 | -0.3345 | 0.7257 | 1.1696 | 0.2366
2012-10-30 21:40 | 28.7034 | 14.3638 | 15.3996 | 15.9006 | 14.9023 | -1.472 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006952 | 0.21479 | 0.001493 | 30.8956 | 14.7514 | -0.1015 | 0.003695 | -0.3876 | 0.6482 | 1.1492 | 0.1509
2012-10-30 21:40 | 28.7012 | 14.3371 | 15.4137 | 15.8373 | 14.9857 | -1.4714 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 14.6962 | -0.1014 | 0.003688 | -0.3591 | 0.7175 | 1.1411 | 0.2895
2012-10-30 21:40 | 28.7072 | 14.3717 | 15.4115 | 15.9084 | 14.9469 | -1.4766 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.7389 | -0.1018 | 0.003686 | -0.3672 | 0.6726 | 1.1695 | 0.208
2012-10-30 21:40 | 28.6989 | 14.4165 | 15.4114 | 15.8757 | 14.9182 | -1.4749 | 0.006716 | 0.32651 | 0.002193 | 48.6173 | 0.006943 | 0.21487 | 0.001492 | 30.9481 | 14.6653 | -0.1017 | 0.003685 | -0.2488 | 0.7461 | 1.2104 | 0.2529
2012-10-30 21:40 | 28.7079 | 14.3683 | 15.4041 | 15.8847 | 14.9149 | -1.4708 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.7273 | -0.1014 | 0.003687 | -0.359 | 0.6768 | 1.1574 | 0.1876
2012-10-30 21:40 | 28.7057 | 14.3294 | 15.4059 | 15.8825 | 14.9494 | -1.4749 | 0.006744 | 0.32641 | 0.002201 | 48.4005 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 14.7211 | -0.1017 | 0.003688 | -0.3917 | 0.6848 | 1.1614 | 0.2283
2012-10-30 21:40 | 28.6974 | 14.3497 | 15.3773 | 15.8783 | 14.9534 | -1.4633 | 0.006729 | 0.32656 | 0.002197 | 48.5298 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.6638 | -0.1009 | 0.003687 | -0.3141 | 0.7135 | 1.2145 | 0.2896
2012-10-30 21:40 | 28.7034 | 14.3353 | 15.3833 | 15.8884 | 14.9431 | -1.4714 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 14.6902 | -0.1014 | 0.003688 | -0.3549 | 0.6931 | 1.1982 | 0.2529
2012-10-30 21:40 | 28.7049 | 14.3164 | 15.4051 | 15.8736 | 14.9405 | -1.4754 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.7407 | -0.1017 | 0.003688 | -0.4243 | 0.6644 | 1.1329 | 0.1998
2012-10-30 21:40 | 28.7057 | 14.3253 | 15.3815 | 15.8296 | 14.9413 | -1.4656 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 14.7129 | -0.101 | 0.003688 | -0.3876 | 0.6686 | 1.1167 | 0.2284
2012-10-30 21:40 | 28.7042 | 14.3565 | 15.3922 | 15.8444 | 14.9316 | -1.4598 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006961 | 0.21476 | 0.001495 | 30.8516 | 14.6788 | -0.1006 | 0.003695 | -0.3223 | 0.7134 | 1.1656 | 0.2528
2012-10-30 21:40 | 28.7117 | 14.3313 | 15.4242 | 15.8885 | 14.9432 | -1.4697 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.7067 | -0.1013 | 0.003691 | -0.3754 | 0.7175 | 1.1818 | 0.2365
2012-10-30 21:40 | 28.7079 | 14.3357 | 15.3756 | 15.8358 | 14.9557 | -1.4639 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.6906 | -0.1009 | 0.003684 | -0.3549 | 0.685 | 1.1452 | 0.2651
2012-10-30 21:41 | 28.7117 | 14.364 | 15.4038 | 15.8519 | 14.9514 | -1.4685 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.7556 | -0.1012 | 0.003688 | -0.3916 | 0.6482 | 1.0963 | 0.1958
2012-10-30 21:41 | 28.7117 | 14.3354 | 15.3956 | 15.8641 | 14.9595 | -1.457 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.7393 | -0.1004 | 0.003686 | -0.4039 | 0.6563 | 1.1248 | 0.2202
2012-10-30 21:41 | 28.7109 | 14.3632 | 15.4478 | 15.8185 | 14.9465 | -1.4702 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 14.7263 | -0.1014 | 0.003683 | -0.3631 | 0.7215 | 1.0922 | 0.2202
2012-10-30 21:41 | 28.7072 | 14.3391 | 15.383 | 15.8596 | 14.9183 | -1.4731 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.7552 | -0.1016 | 0.003685 | -0.4161 | 0.6278 | 1.1044 | 0.1631
2012-10-30 21:41 | 28.7064 | 14.3301 | 15.4107 | 15.7977 | 14.942 | -1.4772 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.7585 | -0.1018 | 0.003688 | -0.4284 | 0.6522 | 1.0392 | 0.1835
2012-10-30 21:41 | 28.7094 | 14.3331 | 15.4015 | 15.8618 | 14.9001 | -1.465 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.737 | -0.101 | 0.00369 | -0.4039 | 0.6645 | 1.1248 | 0.1631
2012-10-30 21:41 | 28.7064 | 14.3179 | 15.3822 | 15.8425 | 14.9379 | -1.4778 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 14.7422 | -0.1019 | 0.003691 | -0.4243 | 0.64 | 1.1003 | 0.1957
2012-10-30 21:41 | 28.7087 | 14.3365 | 15.4008 | 15.8326 | 14.9769 | -1.4726 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.8016 | -0.1015 | 0.003687 | -0.4651 | 0.5992 | 1.031 | 0.1753
2012-10-30 21:41 | 28.7147 | 14.2976 | 15.4068 | 15.8467 | 14.8932 | -1.4743 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 14.7953 | -0.1016 | 0.00369 | -0.4977 | 0.6115 | 1.0514 | 0.0979
2012-10-30 21:41 | 28.7124 | 14.2994 | 15.3882 | 15.8118 | 14.9398 | -1.4633 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.7849 | -0.1009 | 0.00369 | -0.4855 | 0.6033 | 1.0269 | 0.1549
2012-10-30 21:41 | 28.7094 | 14.2923 | 15.3852 | 15.874 | 14.9328 | -1.4662 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006917 | 0.21479 | 0.001486 | 31.0519 | 14.7737 | -0.1011 | 0.003684 | -0.4814 | 0.6115 | 1.1003 | 0.1591
2012-10-30 21:41 | 28.7012 | 14.3249 | 15.377 | 15.8129 | 14.9205 | -1.465 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 14.7777 | -0.101 | 0.003687 | -0.4528 | 0.5993 | 1.0352 | 0.1428
2012-10-30 21:41 | 28.6974 | 14.2844 | 15.4221 | 15.8172 | 14.9289 | -1.4835 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 14.7291 | -0.1023 | 0.003685 | -0.4447 | 0.693 | 1.0881 | 0.1998
2012-10-30 21:41 | 28.7034 | 14.3108 | 15.3792 | 15.8151 | 14.9594 | -1.4679 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.7432 | -0.1012 | 0.003684 | -0.4324 | 0.636 | 1.0719 | 0.2162
2012-10-30 21:41 | 28.7034 | 14.3189 | 15.3751 | 15.8313 | 14.8941 | -1.4691 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 14.7473 | -0.1013 | 0.003685 | -0.4284 | 0.6278 | 1.084 | 0.1468
2012-10-30 21:41 | 28.7079 | 14.3071 | 15.4041 | 15.8481 | 14.9516 | -1.457 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.7273 | -0.1004 | 0.003685 | -0.4202 | 0.6768 | 1.1208 | 0.2243
2012-10-30 21:41 | 28.7102 | 14.3176 | 15.4227 | 15.8259 | 14.9213 | -1.4702 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 14.7296 | -0.1014 | 0.003685 | -0.412 | 0.6931 | 1.0963 | 0.1917
2012-10-30 21:41 | 28.7087 | 14.3324 | 15.4375 | 15.8407 | 14.879 | -1.4593 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.7771 | -0.1006 | 0.003688 | -0.4447 | 0.6604 | 1.0636 | 0.1019
2012-10-30 21:41 | 28.7027 | 14.3101 | 15.3663 | 15.7818 | 14.9057 | -1.4668 | 0.006713 | 0.3265 | 0.002192 | 48.637 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 14.7181 | -0.1011 | 0.003683 | -0.408 | 0.6482 | 1.0637 | 0.1876
2012-10-30 21:41 | 28.7124 | 14.3035 | 15.4086 | 15.8281 | 14.9154 | -1.4795 | 0.006719 | 0.32651 | 0.002194 | 48.5956 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.7645 | -0.102 | 0.003681 | -0.461 | 0.6441 | 1.0636 | 0.1509
2012-10-30 21:41 | 28.7139 | 14.2805 | 15.349 | 15.7889 | 14.8883 | -1.465 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.7537 | -0.101 | 0.003688 | -0.4732 | 0.5953 | 1.0352 | 0.1346
2012-10-30 21:41 | 28.7169 | 14.3202 | 15.3683 | 15.8123 | 14.928 | -1.4662 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.00693 | 0.21468 | 0.001488 | 30.9791 | 14.7608 | -0.1011 | 0.00369 | -0.4406 | 0.6075 | 1.0515 | 0.1672
2012-10-30 21:41 | 28.7169 | 14.3365 | 15.3642 | 15.8041 | 14.9362 | -1.4743 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006917 | 0.21474 | 0.001485 | 31.0457 | 14.7037 | -0.1016 | 0.003687 | -0.3672 | 0.6605 | 1.1004 | 0.2325
2012-10-30 21:41 | 28.7087 | 14.2671 | 15.3478 | 15.8041 | 14.8709 | -1.4737 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 14.7404 | -0.1016 | 0.003687 | -0.4733 | 0.6074 | 1.0637 | 0.1305
2012-10-30 21:41 | 28.7236 | 14.3024 | 15.3464 | 15.823 | 14.8817 | -1.4731 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 14.739 | -0.1016 | 0.003687 | -0.4366 | 0.6074 | 1.084 | 0.1427
2012-10-30 21:41 | 28.7229 | 14.318 | 15.3702 | 15.8183 | 14.9218 | -1.4622 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006949 | 0.21477 | 0.001492 | 30.9069 | 14.722 | -0.1008 | 0.003692 | -0.404 | 0.6482 | 1.0963 | 0.1998
2012-10-30 21:41 | 28.7214 | 14.288 | 15.3972 | 15.7964 | 14.904 | -1.4656 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 14.7449 | -0.101 | 0.003688 | -0.4569 | 0.6523 | 1.0515 | 0.1591
2012-10-30 21:41 | 28.7259 | 14.2721 | 15.3447 | 15.7846 | 14.9207 | -1.4731 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 14.7617 | -0.1016 | 0.003685 | -0.4896 | 0.583 | 1.0229 | 0.159
2012-10-30 21:42 | 28.7274 | 14.294 | 15.3706 | 15.8146 | 14.91 | -1.4622 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 14.7468 | -0.1008 | 0.003692 | -0.4528 | 0.6238 | 1.0678 | 0.1632
2012-10-30 21:42 | 28.7281 | 14.2824 | 15.3958 | 15.7827 | 14.9025 | -1.4726 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 14.7516 | -0.1015 | 0.003689 | -0.4692 | 0.6442 | 1.0311 | 0.1509
2012-10-30 21:42 | 28.7289 | 14.3322 | 15.3599 | 15.7673 | 14.8829 | -1.4656 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006914 | 0.21479 | 0.001485 | 31.0654 | 14.781 | -0.101 | 0.003686 | -0.4488 | 0.5789 | 0.9863 | 0.1019
2012-10-30 21:42 | 28.7296 | 14.2513 | 15.3484 | 15.7965 | 14.9448 | -1.4656 | 0.006747 | 0.32651 | 0.002203 | 48.394 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 14.7817 | -0.101 | 0.003692 | -0.5304 | 0.5667 | 1.0148 | 0.1631
2012-10-30 21:42 | 28.7229 | 14.265 | 15.3661 | 15.8101 | 14.9218 | -1.4708 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006939 | 0.21489 | 0.001491 | 30.9682 | 14.7097 | -0.1014 | 0.003687 | -0.4447 | 0.6564 | 1.1004 | 0.2121
2012-10-30 21:42 | 28.7259 | 14.2312 | 15.3406 | 15.7968 | 14.9452 | -1.4616 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.7168 | -0.1008 | 0.003685 | -0.4856 | 0.6238 | 1.08 | 0.2284
2012-10-30 21:42 | 28.7304 | 14.2766 | 15.3777 | 15.8095 | 14.9048 | -1.4627 | 0.006722 | 0.32653 | 0.002195 | 48.5762 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 14.7417 | -0.1008 | 0.003685 | -0.4651 | 0.636 | 1.0678 | 0.1631
2012-10-30 21:42 | 28.7304 | 14.2806 | 15.3899 | 15.8136 | 14.9089 | -1.465 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.7458 | -0.101 | 0.003685 | -0.4652 | 0.6441 | 1.0678 | 0.1631
2012-10-30 21:42 | 28.7289 | 14.2506 | 15.3925 | 15.751 | 14.9156 | -1.4627 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 14.7402 | -0.1008 | 0.003683 | -0.4896 | 0.6523 | 1.0108 | 0.1754
2012-10-30 21:42 | 28.7394 | 14.2529 | 15.3704 | 15.8103 | 14.9097 | -1.472 | 0.006722 | 0.32657 | 0.002195 | 48.5826 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 14.7466 | -0.1015 | 0.003687 | -0.4937 | 0.6238 | 1.0637 | 0.1631
2012-10-30 21:42 | 28.7296 | 14.2513 | 15.3769 | 15.7598 | 14.9122 | -1.472 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.7246 | -0.1015 | 0.003684 | -0.4733 | 0.6523 | 1.0352 | 0.1876
2012-10-30 21:42 | 28.7409 | 14.2871 | 15.38 | 15.7752 | 14.9072 | -1.4697 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 14.74 | -0.1013 | 0.003687 | -0.4529 | 0.64 | 1.0352 | 0.1672
2012-10-30 21:42 | 28.7401 | 14.2781 | 15.3589 | 15.8029 | 14.9431 | -1.4645 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 14.7147 | -0.101 | 0.003686 | -0.4366 | 0.6442 | 1.0882 | 0.2284
2012-10-30 21:42 | 28.7476 | 14.2693 | 15.3827 | 15.79 | 14.8772 | -1.4778 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 14.7385 | -0.1019 | 0.003691 | -0.4692 | 0.6442 | 1.0515 | 0.1387
2012-10-30 21:42 | 28.7461 | 14.2596 | 15.3567 | 15.7682 | 14.9287 | -1.4604 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 14.7288 | -0.1007 | 0.003691 | -0.4692 | 0.6279 | 1.0394 | 0.1999
2012-10-30 21:42 | 28.7319 | 14.2454 | 15.3384 | 15.7295 | 14.9145 | -1.4754 | 0.006741 | 0.32641 | 0.0022 | 48.422 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 14.7269 | -0.1017 | 0.00369 | -0.4815 | 0.6115 | 1.0026 | 0.1876
2012-10-30 21:42 | 28.7446 | 14.2785 | 15.3593 | 15.7748 | 14.8986 | -1.483 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 14.7192 | -0.1022 | 0.003679 | -0.4407 | 0.6401 | 1.0556 | 0.1794
2012-10-30 21:42 | 28.7431 | 14.2729 | 15.3456 | 15.7855 | 14.889 | -1.4645 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 14.7136 | -0.101 | 0.003688 | -0.4407 | 0.632 | 1.0719 | 0.1754
2012-10-30 21:42 | 28.7454 | 14.2711 | 15.3886 | 15.7919 | 14.8831 | -1.4697 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 14.7648 | -0.1013 | 0.003691 | -0.4937 | 0.6238 | 1.0271 | 0.1183
2012-10-30 21:42 | 28.7491 | 14.3075 | 15.3719 | 15.816 | 14.895 | -1.4691 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 14.7074 | -0.1013 | 0.003688 | -0.3999 | 0.6645 | 1.1086 | 0.1876
2012-10-30 21:42 | 28.7364 | 14.2866 | 15.3389 | 15.7422 | 14.87 | -1.4622 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.00693 | 0.21483 | 0.001489 | 31 | 14.7273 | -0.1008 | 0.003685 | -0.4407 | 0.6116 | 1.0149 | 0.1427
2012-10-30 21:42 | 28.7409 | 14.3197 | 15.3474 | 15.7385 | 14.9357 | -1.4558 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 14.7155 | -0.1004 | 0.003684 | -0.3958 | 0.6319 | 1.023 | 0.2202
2012-10-30 21:42 | 28.7311 | 14.3181 | 15.3499 | 15.7613 | 14.9341 | -1.4685 | 0.006729 | 0.32654 | 0.002197 | 48.5277 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 14.7424 | -0.1012 | 0.003686 | -0.4243 | 0.6075 | 1.0189 | 0.1917
2012-10-30 21:42 | 28.7349 | 14.2933 | 15.3455 | 15.7814 | 14.9012 | -1.472 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.7299 | -0.1015 | 0.003686 | -0.4366 | 0.6156 | 1.0515 | 0.1713
2012-10-30 21:42 | 28.7386 | 14.3215 | 15.3737 | 15.8055 | 14.9253 | -1.4627 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 14.6805 | -0.1008 | 0.003687 | -0.359 | 0.6932 | 1.125 | 0.2448
2012-10-30 21:42 | 28.7371 | 14.3282 | 15.3355 | 15.7632 | 14.8707 | -1.4772 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 14.7525 | -0.1018 | 0.003686 | -0.4243 | 0.583 | 1.0107 | 0.1182
2012-10-30 21:43 | 28.7401 | 14.2985 | 15.3181 | 15.7866 | 14.8656 | -1.4702 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.7392 | -0.1014 | 0.003683 | -0.4407 | 0.5789 | 1.0474 | 0.1264
2012-10-30 21:43 | 28.7289 | 14.2791 | 15.3191 | 15.7387 | 14.887 | -1.4656 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006943 | 0.21471 | 0.001491 | 30.9252 | 14.6994 | -0.101 | 0.003688 | -0.4203 | 0.6197 | 1.0393 | 0.1876
2012-10-30 21:43 | 28.7296 | 14.2921 | 15.3321 | 15.768 | 14.9081 | -1.4679 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 14.7205 | -0.1012 | 0.003688 | -0.4284 | 0.6116 | 1.0475 | 0.1876
2012-10-30 21:43 | 28.7304 | 14.3255 | 15.3247 | 15.7484 | 14.9252 | -1.4697 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 14.6805 | -0.1013 | 0.003688 | -0.355 | 0.6442 | 1.0679 | 0.2447
2012-10-30 21:43 | 28.7177 | 14.3088 | 15.312 | 15.7438 | 14.8677 | -1.4731 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006946 | 0.21482 | 0.001492 | 30.9264 | 14.7004 | -0.1016 | 0.00369 | -0.3916 | 0.6116 | 1.0434 | 0.1673
2012-10-30 21:43 | 28.7184 | 14.3421 | 15.276 | 15.7486 | 14.901 | -1.4714 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.6563 | -0.1014 | 0.003685 | -0.3142 | 0.6197 | 1.0923 | 0.2447
2012-10-30 21:43 | 28.7251 | 14.2998 | 15.2909 | 15.7431 | 14.8832 | -1.4737 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 14.7119 | -0.1016 | 0.003686 | -0.4121 | 0.579 | 1.0312 | 0.1713
2012-10-30 21:43 | 28.7251 | 14.2917 | 15.3683 | 15.7268 | 14.8628 | -1.465 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.6874 | -0.101 | 0.003688 | -0.3957 | 0.6809 | 1.0394 | 0.1754
2012-10-30 21:43 | 28.7274 | 14.3225 | 15.3013 | 15.725 | 14.857 | -1.4685 | 0.006735 | 0.32656 | 0.002199 | 48.4866 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 14.7142 | -0.1012 | 0.00369 | -0.3917 | 0.5871 | 1.0108 | 0.1428
2012-10-30 21:43 | 28.7334 | 14.2918 | 15.3236 | 15.7432 | 14.8793 | -1.465 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 14.7365 | -0.101 | 0.003688 | -0.4447 | 0.5871 | 1.0067 | 0.1428
2012-10-30 21:43 | 28.7177 | 14.2761 | 15.2753 | 15.7153 | 14.884 | -1.4691 | 0.006744 | 0.32656 | 0.002202 | 48.4219 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 14.7535 | -0.1013 | 0.003693 | -0.4774 | 0.5218 | 0.9618 | 0.1305
2012-10-30 21:43 | 28.7177 | 14.2679 | 15.2794 | 15.7479 | 14.9125 | -1.4639 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 14.7004 | -0.1009 | 0.00368 | -0.4325 | 0.579 | 1.0475 | 0.2121
2012-10-30 21:43 | 28.7206 | 14.2504 | 15.2782 | 15.7427 | 14.842 | -1.4662 | 0.006722 | 0.32651 | 0.002195 | 48.5739 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.6911 | -0.1011 | 0.003688 | -0.4407 | 0.5871 | 1.0516 | 0.1509
2012-10-30 21:43 | 28.7214 | 14.2716 | 15.3483 | 15.7068 | 14.855 | -1.4714 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 14.696 | -0.1014 | 0.003682 | -0.4244 | 0.6523 | 1.0108 | 0.159
2012-10-30 21:43 | 28.7274 | 14.2491 | 15.3095 | 15.7617 | 14.857 | -1.4662 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006946 | 0.21486 | 0.001492 | 30.9328 | 14.7061 | -0.1011 | 0.003687 | -0.457 | 0.6034 | 1.0556 | 0.1509
2012-10-30 21:43 | 28.7192 | 14.2531 | 15.2768 | 15.7413 | 14.8365 | -1.4702 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 14.6897 | -0.1014 | 0.003683 | -0.4366 | 0.5871 | 1.0516 | 0.1468
2012-10-30 21:43 | 28.7184 | 14.2482 | 15.2801 | 15.7079 | 14.8806 | -1.472 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 14.6726 | -0.1015 | 0.003684 | -0.4244 | 0.6075 | 1.0353 | 0.208
2012-10-30 21:43 | 28.7251 | 14.2835 | 15.299 | 15.7227 | 14.8832 | -1.4535 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.7201 | -0.1002 | 0.003688 | -0.4366 | 0.5789 | 1.0026 | 0.1631
2012-10-30 21:43 | 28.7281 | 14.2947 | 15.3265 | 15.6687 | 14.8536 | -1.4662 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.7516 | -0.1011 | 0.003685 | -0.4569 | 0.5749 | 0.9171 | 0.102
2012-10-30 21:43 | 28.7236 | 14.2412 | 15.2853 | 15.6805 | 14.8899 | -1.4668 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 14.7594 | -0.1011 | 0.003683 | -0.5182 | 0.5259 | 0.9211 | 0.1305
2012-10-30 21:43 | 28.7184 | 14.2482 | 15.276 | 15.6834 | 14.8317 | -1.4801 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006952 | 0.21474 | 0.001493 | 30.8894 | 14.7052 | -0.102 | 0.003687 | -0.457 | 0.5708 | 0.9782 | 0.1265
2012-10-30 21:43 | 28.7087 | 14.263 | 15.2337 | 15.6778 | 14.8301 | -1.4691 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 14.72 | -0.1013 | 0.003684 | -0.457 | 0.5137 | 0.9578 | 0.1101
2012-10-30 21:43 | 28.7162 | 14.2379 | 15.2616 | 15.722 | 14.858 | -1.4558 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 14.7193 | -0.1004 | 0.003683 | -0.4814 | 0.5423 | 1.0027 | 0.1387
2012-10-30 21:43 | 28.7177 | 14.2557 | 15.2794 | 15.699 | 14.8391 | -1.4708 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 14.6882 | -0.1014 | 0.003688 | -0.4325 | 0.5912 | 1.0108 | 0.1509
2012-10-30 21:43 | 28.7244 | 14.2501 | 15.2779 | 15.7342 | 14.7928 | -1.4662 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 14.7724 | -0.1011 | 0.003688 | -0.5223 | 0.5055 | 0.9618 | 0.0204
2012-10-30 21:43 | 28.7169 | 14.2059 | 15.3194 | 15.6901 | 14.8342 | -1.461 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 14.7363 | -0.1007 | 0.003687 | -0.5304 | 0.5831 | 0.9538 | 0.0979
2012-10-30 21:43 | 28.7162 | 14.2542 | 15.2494 | 15.6853 | 14.8417 | -1.4679 | 0.006738 | 0.32644 | 0.0022 | 48.4479 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 14.6989 | -0.1012 | 0.00369 | -0.4447 | 0.5505 | 0.9864 | 0.1428
2012-10-30 21:43 | 28.7259 | 14.2272 | 15.2713 | 15.6909 | 14.9044 | -1.4622 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 14.7046 | -0.1008 | 0.00369 | -0.4774 | 0.5667 | 0.9863 | 0.1998
2012-10-30 21:44 | 28.7259 | 14.2639 | 15.2794 | 15.6176 | 14.8718 | -1.4668 | 0.006719 | 0.32656 | 0.002194 | 48.602 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 14.6964 | -0.1011 | 0.003683 | -0.4325 | 0.583 | 0.9212 | 0.1754
2012-10-30 21:44 | 28.7184 | 14.2482 | 15.2679 | 15.6712 | 14.8724 | -1.4743 | 0.006744 | 0.32643 | 0.002201 | 48.4026 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 14.693 | -0.1016 | 0.003692 | -0.4448 | 0.5749 | 0.9782 | 0.1794
2012-10-30 21:44 | 28.7206 | 14.2545 | 15.2619 | 15.6653 | 14.8502 | -1.4633 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 14.6993 | -0.1009 | 0.003682 | -0.4448 | 0.5626 | 0.966 | 0.1509
2012-10-30 21:44 | 28.7221 | 14.2642 | 15.2553 | 15.6627 | 14.8476 | -1.4674 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 14.7171 | -0.1012 | 0.003689 | -0.4529 | 0.5382 | 0.9456 | 0.1305
2012-10-30 21:44 | 28.7162 | 14.2338 | 15.2127 | 15.6487 | 14.8539 | -1.4691 | 0.006735 | 0.32659 | 0.0022 | 48.4909 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 14.6296 | -0.1013 | 0.00369 | -0.3958 | 0.5831 | 1.0191 | 0.2243
2012-10-30 21:44 | 28.7214 | 14.2839 | 15.2872 | 15.6702 | 14.8714 | -1.4772 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 14.6797 | -0.1018 | 0.003685 | -0.3958 | 0.6075 | 0.9905 | 0.1917
2012-10-30 21:44 | 28.7162 | 14.2542 | 15.2738 | 15.6609 | 14.8376 | -1.4702 | 0.006738 | 0.32644 | 0.0022 | 48.4479 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 14.6826 | -0.1014 | 0.003689 | -0.4284 | 0.5912 | 0.9783 | 0.155
2012-10-30 21:44 | 28.7221 | 14.2927 | 15.2838 | 15.7075 | 14.8639 | -1.4633 | 0.006751 | 0.32644 | 0.002204 | 48.3546 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 14.7089 | -0.1009 | 0.003693 | -0.4162 | 0.5749 | 0.9986 | 0.155
2012-10-30 21:44 | 28.7192 | 14.2531 | 15.285 | 15.6842 | 14.8365 | -1.4564 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 14.6285 | -0.1004 | 0.003692 | -0.3754 | 0.6565 | 1.0557 | 0.208
2012-10-30 21:44 | 28.7132 | 14.2675 | 15.2423 | 15.6497 | 14.8265 | -1.4726 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 14.6674 | -0.1015 | 0.003686 | -0.3999 | 0.5749 | 0.9823 | 0.1591
2012-10-30 21:44 | 28.7206 | 14.2504 | 15.2782 | 15.6856 | 14.8624 | -1.4604 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 14.6911 | -0.1007 | 0.003685 | -0.4407 | 0.5871 | 0.9945 | 0.1713
2012-10-30 21:44 | 28.7206 | 14.2668 | 15.2375 | 15.6816 | 14.8542 | -1.4622 | 0.006741 | 0.32653 | 0.002201 | 48.4393 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 14.6911 | -0.1008 | 0.003689 | -0.4243 | 0.5464 | 0.9905 | 0.1631
2012-10-30 21:44 | 28.7199 | 14.2579 | 15.2408 | 15.6564 | 14.8576 | -1.4645 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 14.6496 | -0.101 | 0.003688 | -0.3917 | 0.5912 | 1.0068 | 0.208
2012-10-30 21:44 | 28.7221 | 14.2683 | 15.2797 | 15.6505 | 14.8313 | -1.4593 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 14.6844 | -0.1006 | 0.003689 | -0.4161 | 0.5953 | 0.9661 | 0.1469
2012-10-30 21:44 | 28.7236 | 14.2616 | 15.269 | 15.6683 | 14.8124 | -1.4731 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 14.6655 | -0.1016 | 0.003689 | -0.4039 | 0.6035 | 1.0028 | 0.1469
2012-10-30 21:44 | 28.7244 | 14.2461 | 15.2413 | 15.6609 | 14.805 | -1.4616 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006943 | 0.21486 | 0.001492 | 30.9461 | 14.6827 | -0.1008 | 0.003687 | -0.4366 | 0.5586 | 0.9782 | 0.1223
2012-10-30 21:44 | 0 |  |  |  |  | 1241028.7267 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006936 | 0.21479 | 0.00149 | 30.9668 |  | 85556.5204 | 0.00369 | 0 | 0 | 0 | 0
2012-10-30 21:44 | 27.193 | 23.7514 | 23.383 | 23.7674 | 23.0182 | 0.1144 | 0.006725 | 0.32656 | 0.002196 | 48.5587 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 22.8297 | 0.0079 | 0.003688 | 0.9217 | 0.5533 | 0.9377 | 0.1885
2012-10-30 21:44 | 26.678 | 23.2164 | 22.868 | 23.2684 | 22.4551 | 0.1121 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.2625 | 0.0077 | 0.003686 | 0.9539 | 0.6055 | 1.0059 | 0.1926
2012-10-30 21:44 | 26.6772 | 23.2276 | 22.8712 | 23.2596 | 22.4422 | 0.1202 | 0.006744 | 0.32646 | 0.002202 | 48.4069 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.3379 | 0.0083 | 0.003691 | 0.8897 | 0.5333 | 0.9217 | 0.1043
2012-10-30 21:44 | 26.6855 | 23.2119 | 22.8714 | 23.2599 | 22.4786 | 0.0994 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.3342 | 0.0069 | 0.003684 | 0.8777 | 0.5372 | 0.9257 | 0.1444
2012-10-30 21:44 | 26.684 | 23.2344 | 22.882 | 23.2664 | 22.4731 | 0.1219 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.2966 | 0.0084 | 0.00369 | 0.9378 | 0.5854 | 0.9698 | 0.1765
2012-10-30 21:44 | 26.687 | 23.2174 | 22.905 | 23.2614 | 22.4681 | 0.115 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3838 | 0.0079 | 0.003682 | 0.8336 | 0.5212 | 0.8776 | 0.0843
2012-10-30 21:44 | 26.684 | 23.1944 | 22.89 | 23.2784 | 22.4731 | 0.1017 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 22.3407 | 0.007 | 0.003686 | 0.8537 | 0.5493 | 0.9377 | 0.1324
2012-10-30 21:44 | 26.6847 | 23.2391 | 22.8747 | 23.3071 | 22.5019 | 0.115 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3254 | 0.0079 | 0.003687 | 0.9137 | 0.5493 | 0.9817 | 0.1765
2012-10-30 21:44 | 26.6877 | 23.2061 | 22.8696 | 23.3021 | 22.4808 | 0.1011 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 22.3083 | 0.007 | 0.003691 | 0.8978 | 0.5613 | 0.9938 | 0.1725
2012-10-30 21:44 | 26.69 | 23.2204 | 22.928 | 23.2884 | 22.431 | 0.1156 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3026 | 0.008 | 0.003681 | 0.9178 | 0.6254 | 0.9858 | 0.1284
2012-10-30 21:45 | 26.6765 | 23.2069 | 22.8705 | 23.2509 | 22.4656 | 0.093 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 22.3052 | 0.0064 | 0.003688 | 0.9017 | 0.5653 | 0.9457 | 0.1604
2012-10-30 21:45 | 26.6817 | 23.2161 | 22.8436 | 23.2561 | 22.4828 | 0.1069 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 22.3304 | 0.0074 | 0.003687 | 0.8857 | 0.5132 | 0.9257 | 0.1524
2012-10-30 21:45 | 26.6847 | 23.183 | 22.9147 | 23.2711 | 22.4738 | 0.1081 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3334 | 0.0075 | 0.003684 | 0.8496 | 0.5813 | 0.9377 | 0.1404
2012-10-30 21:45 | 26.684 | 23.2224 | 22.886 | 23.2744 | 22.449 | 0.0988 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.2886 | 0.0068 | 0.003688 | 0.9338 | 0.5974 | 0.9858 | 0.1604
2012-10-30 21:45 | 26.6847 | 23.1911 | 22.8626 | 23.2951 | 22.4457 | 0.1139 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.3254 | 0.0079 | 0.003683 | 0.8657 | 0.5372 | 0.9697 | 0.1203
2012-10-30 21:45 | 26.687 | 23.2174 | 22.917 | 23.2614 | 22.4962 | 0.1 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006943 | 0.21473 | 0.001491 | 30.9274 | 22.3157 | 0.0069 | 0.003686 | 0.9017 | 0.6013 | 0.9457 | 0.1805
2012-10-30 21:45 | 26.6922 | 23.1705 | 22.8942 | 23.2586 | 22.4893 | 0.1104 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.3329 | 0.0076 | 0.003687 | 0.8376 | 0.5613 | 0.9257 | 0.1564
2012-10-30 21:45 | 26.6832 | 23.1655 | 22.8691 | 23.2416 | 22.4643 | 0.1011 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 22.3319 | 0.007 | 0.003691 | 0.8336 | 0.5372 | 0.9097 | 0.1324
2012-10-30 21:45 | 26.6877 | 23.17 | 22.9017 | 23.2341 | 22.4407 | 0.111 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.3083 | 0.0077 | 0.003685 | 0.8617 | 0.5934 | 0.9258 | 0.1324
2012-10-30 21:45 | 26.6862 | 23.1685 | 22.8842 | 23.2326 | 22.4111 | 0.1046 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 22.3189 | 0.0072 | 0.003693 | 0.8496 | 0.5653 | 0.9137 | 0.0922
2012-10-30 21:45 | 26.6967 | 23.2391 | 22.8786 | 23.2271 | 22.4577 | 0.1121 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006939 | 0.21474 | 0.00149 | 30.9473 | 22.3173 | 0.0077 | 0.00369 | 0.9218 | 0.5613 | 0.9098 | 0.1404
2012-10-30 21:45 | 26.6892 | 23.1956 | 22.8671 | 23.2396 | 22.4342 | 0.1034 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.2777 | 0.0071 | 0.003688 | 0.9179 | 0.5894 | 0.9619 | 0.1565
2012-10-30 21:45 | 26.69 | 23.1803 | 22.888 | 23.2604 | 22.4791 | 0.1029 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.2825 | 0.0071 | 0.003687 | 0.8978 | 0.6055 | 0.9779 | 0.1966
2012-10-30 21:45 | 26.6847 | 23.147 | 22.8706 | 23.2911 | 22.4578 | 0.0919 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3013 | 0.0063 | 0.003686 | 0.8457 | 0.5693 | 0.9898 | 0.1565
2012-10-30 21:45 | 26.6959 | 23.1502 | 22.9379 | 23.2503 | 22.4449 | 0.1058 | 0.006747 | 0.32646 | 0.002203 | 48.3854 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 22.2965 | 0.0073 | 0.003695 | 0.8537 | 0.6414 | 0.9538 | 0.1484
2012-10-30 21:45 | 26.6847 | 23.167 | 22.8827 | 23.2191 | 22.4497 | 0.1133 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3093 | 0.0078 | 0.003691 | 0.8577 | 0.5734 | 0.9098 | 0.1404
2012-10-30 21:45 | 26.6885 | 23.1188 | 22.8785 | 23.2629 | 22.4696 | 0.1133 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3212 | 0.0078 | 0.003686 | 0.7976 | 0.5573 | 0.9417 | 0.1484
2012-10-30 21:45 | 26.693 | 23.2034 | 22.891 | 23.2634 | 22.434 | 0.111 | 0.006738 | 0.32644 | 0.0022 | 48.4479 | 0.006952 | 0.21482 | 0.001493 | 30.8997 | 22.3257 | 0.0077 | 0.003693 | 0.8777 | 0.5653 | 0.9377 | 0.1083
2012-10-30 21:45 | 26.6922 | 23.2106 | 22.8822 | 23.2146 | 22.4412 | 0.1075 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006936 | 0.21487 | 0.00149 | 30.9794 | 22.2767 | 0.0074 | 0.003683 | 0.9339 | 0.6055 | 0.9379 | 0.1645
2012-10-30 21:45 | 26.6855 | 23.1959 | 22.8474 | 23.2519 | 22.4144 | 0.1063 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 22.2058 | 0.0073 | 0.003682 | 0.9901 | 0.6416 | 1.0461 | 0.2086
2012-10-30 21:45 | 26.6885 | 23.1468 | 22.8584 | 23.2269 | 22.4295 | 0.1219 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.281 | 0.0084 | 0.003684 | 0.8658 | 0.5774 | 0.9459 | 0.1485
2012-10-30 21:45 | 26.6832 | 23.1455 | 22.8812 | 23.1815 | 22.4402 | 0.1058 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 22.2236 | 0.0073 | 0.003682 | 0.9219 | 0.6576 | 0.9579 | 0.2166
2012-10-30 21:45 | 26.687 | 23.1373 | 22.8569 | 23.2414 | 22.3999 | 0.1011 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006917 | 0.21479 | 0.001486 | 31.0519 | 22.2474 | 0.007 | 0.003685 | 0.8899 | 0.6095 | 0.994 | 0.1525
2012-10-30 21:45 | 26.6877 | 23.174 | 22.8456 | 23.2221 | 22.4608 | 0.1087 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.2963 | 0.0075 | 0.003691 | 0.8777 | 0.5493 | 0.9258 | 0.1645
2012-10-30 21:45 | 26.6862 | 23.1685 | 22.8441 | 23.2326 | 22.4392 | 0.1098 | 0.006747 | 0.32649 | 0.002203 | 48.3897 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.2707 | 0.0076 | 0.003693 | 0.8978 | 0.5734 | 0.9619 | 0.1685
2012-10-30 21:45 | 26.6825 | 23.1568 | 22.8925 | 23.2209 | 22.4515 | 0.1023 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006939 | 0.2147 | 0.00149 | 30.9411 | 22.271 | 0.0071 | 0.003686 | 0.8858 | 0.6215 | 0.9499 | 0.1805
2012-10-30 21:45 | 26.684 | 23.1703 | 22.8299 | 23.1864 | 22.4731 | 0.115 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.2645 | 0.0079 | 0.003685 | 0.9058 | 0.5654 | 0.9219 | 0.2086
2012-10-30 21:45 | 26.6855 | 23.1678 | 22.8154 | 23.2079 | 22.4385 | 0.1115 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 22.3061 | 0.0077 | 0.003684 | 0.8617 | 0.5093 | 0.9018 | 0.1324
2012-10-30 21:45 | 26.6855 | 23.1919 | 22.8634 | 23.1919 | 22.4626 | 0.1092 | 0.006741 | 0.32659 | 0.002202 | 48.4477 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.2861 | 0.0075 | 0.003692 | 0.9058 | 0.5773 | 0.9058 | 0.1765
2012-10-30 21:46 | 26.687 | 23.1773 | 22.8689 | 23.2494 | 22.4079 | 0.1121 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.3277 | 0.0077 | 0.003687 | 0.8496 | 0.5412 | 0.9217 | 0.0802
2012-10-30 21:46 | 26.6937 | 23.152 | 22.9077 | 23.188 | 22.4708 | 0.1214 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.2862 | 0.0084 | 0.003687 | 0.8658 | 0.6215 | 0.9018 | 0.1846
2012-10-30 21:46 | 26.6832 | 23.1295 | 22.8251 | 23.1896 | 22.3841 | 0.1052 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3038 | 0.0073 | 0.003688 | 0.8257 | 0.5213 | 0.8858 | 0.0803
2012-10-30 21:46 | 26.6922 | 23.1385 | 22.8421 | 23.2266 | 22.4332 | 0.1115 | 0.00671 | 0.32654 | 0.002191 | 48.6651 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.3048 | 0.0077 | 0.003678 | 0.8337 | 0.5373 | 0.9218 | 0.1284
2012-10-30 21:46 | 26.6832 | 23.1415 | 22.9012 | 23.1896 | 22.4482 | 0.1104 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.2677 | 0.0076 | 0.003684 | 0.8738 | 0.6335 | 0.9219 | 0.1805
2012-10-30 21:46 | 26.6967 | 23.159 | 22.9147 | 23.2311 | 22.4377 | 0.1133 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.2692 | 0.0078 | 0.00369 | 0.8898 | 0.6455 | 0.9619 | 0.1685
2012-10-30 21:46 | 26.684 | 23.1543 | 22.8419 | 23.2184 | 22.4009 | 0.1034 | 0.006747 | 0.32653 | 0.002203 | 48.3962 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 22.3006 | 0.0071 | 0.00369 | 0.8537 | 0.5413 | 0.9178 | 0.1003
2012-10-30 21:46 | 26.6825 | 23.1368 | 22.8765 | 23.1969 | 22.4074 | 0.1127 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.251 | 0.0078 | 0.003684 | 0.8858 | 0.6255 | 0.9459 | 0.1564
2012-10-30 21:46 | 26.69 | 23.1803 | 22.8599 | 23.2124 | 22.447 | 0.1092 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.3146 | 0.0075 | 0.003691 | 0.8657 | 0.5453 | 0.8978 | 0.1324
2012-10-30 21:46 | 26.6922 | 23.1465 | 22.8661 | 23.2466 | 22.4412 | 0.1081 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.006921 | 0.2148 | 0.001487 | 31.0361 | 22.2767 | 0.0075 | 0.00368 | 0.8698 | 0.5894 | 0.9699 | 0.1645
2012-10-30 21:46 | 26.6915 | 23.2019 | 22.8374 | 23.2499 | 22.4445 | 0.1029 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.2961 | 0.0071 | 0.003682 | 0.9058 | 0.5413 | 0.9538 | 0.1484
2012-10-30 21:46 | 26.684 | 23.1503 | 22.894 | 23.1944 | 22.4611 | 0.1087 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 22.2926 | 0.0075 | 0.003686 | 0.8577 | 0.6014 | 0.9018 | 0.1685
2012-10-30 21:46 | 26.6847 | 23.147 | 22.8626 | 23.183 | 22.4377 | 0.1092 | 0.006747 | 0.32646 | 0.002203 | 48.3854 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 22.3374 | 0.0075 | 0.003691 | 0.8096 | 0.5252 | 0.8456 | 0.1003
2012-10-30 21:46 | 26.6817 | 23.136 | 22.8436 | 23.1961 | 22.4146 | 0.0948 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.3264 | 0.0065 | 0.003685 | 0.8096 | 0.5172 | 0.8697 | 0.0882
2012-10-30 21:46 | 26.6862 | 23.1485 | 22.8641 | 23.1765 | 22.4713 | 0.1127 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.375 | 0.0078 | 0.00369 | 0.7735 | 0.4891 | 0.8015 | 0.0963
2012-10-30 21:46 | 26.6832 | 23.1215 | 22.9052 | 23.2056 | 22.4763 | 0.1133 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.36 | 0.0078 | 0.003683 | 0.7615 | 0.5452 | 0.8456 | 0.1163
2012-10-30 21:46 | 26.684 | 23.1583 | 22.894 | 23.1904 | 22.453 | 0.1098 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3086 | 0.0076 | 0.003688 | 0.8497 | 0.5854 | 0.8818 | 0.1444
2012-10-30 21:46 | 26.6817 | 23.0799 | 22.8997 | 23.148 | 22.4267 | 0.1087 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006914 | 0.21482 | 0.001485 | 31.0696 | 22.2943 | 0.0075 | 0.003683 | 0.7856 | 0.6054 | 0.8537 | 0.1324
2012-10-30 21:46 | 26.6915 | 23.1178 | 22.8694 | 23.2019 | 22.4565 | 0.104 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3402 | 0.0072 | 0.003691 | 0.7776 | 0.5292 | 0.8617 | 0.1163
2012-10-30 21:46 | 26.6959 | 23.1462 | 22.8899 | 23.1862 | 22.4248 | 0.1029 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3246 | 0.0071 | 0.003684 | 0.8216 | 0.5653 | 0.8616 | 0.1002
2012-10-30 21:46 | 26.6855 | 23.1158 | 22.8955 | 23.1999 | 22.4305 | 0.1208 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.278 | 0.0083 | 0.003687 | 0.8378 | 0.6175 | 0.9219 | 0.1525
2012-10-30 21:46 | 26.6922 | 23.1585 | 22.8902 | 23.1665 | 22.4372 | 0.111 | 0.006735 | 0.32654 | 0.002199 | 48.4845 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.2928 | 0.0077 | 0.003689 | 0.8657 | 0.5974 | 0.8737 | 0.1444
2012-10-30 21:46 | 26.687 | 23.1253 | 22.8729 | 23.1733 | 22.436 | 0.1063 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.3036 | 0.0073 | 0.003687 | 0.8217 | 0.5693 | 0.8697 | 0.1324
2012-10-30 21:47 | 26.6952 | 23.0814 | 22.8811 | 23.1735 | 22.4442 | 0.1058 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3158 | 0.0073 | 0.003688 | 0.7656 | 0.5653 | 0.8577 | 0.1284
2012-10-30 21:47 | 26.6974 | 23.1197 | 22.8593 | 23.1677 | 22.4303 | 0.1069 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.306 | 0.0074 | 0.003686 | 0.8137 | 0.5533 | 0.8617 | 0.1243
2012-10-30 21:47 | 26.7012 | 23.1034 | 22.8591 | 23.2076 | 22.4341 | 0.111 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3459 | 0.0077 | 0.003684 | 0.7575 | 0.5132 | 0.8617 | 0.0882
2012-10-30 21:47 | 26.6974 | 23.0836 | 22.8513 | 23.1877 | 22.4023 | 0.1156 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 22.3421 | 0.008 | 0.003686 | 0.7415 | 0.5092 | 0.8456 | 0.0602
2012-10-30 21:47 | 26.6974 | 23.0756 | 22.8833 | 23.1797 | 22.4584 | 0.0977 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 22.3581 | 0.0067 | 0.003688 | 0.7175 | 0.5252 | 0.8216 | 0.1003
2012-10-30 21:47 | 26.7012 | 23.1155 | 22.8791 | 23.2436 | 22.4422 | 0.0965 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.3339 | 0.0067 | 0.003684 | 0.7816 | 0.5452 | 0.9097 | 0.1083
2012-10-30 21:47 | 26.6885 | 23.1028 | 22.8825 | 23.2069 | 22.4335 | 0.1058 | 0.006725 | 0.32651 | 0.002196 | 48.5523 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.3573 | 0.0073 | 0.003686 | 0.7455 | 0.5252 | 0.8496 | 0.0762
2012-10-30 21:47 | 26.7004 | 23.0586 | 22.8783 | 23.1667 | 22.3852 | 0.0994 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 22.3411 | 0.0069 | 0.003682 | 0.7175 | 0.5372 | 0.8256 | 0.0441
2012-10-30 21:47 | 26.6959 | 23.0741 | 22.8578 | 23.1862 | 22.4449 | 0.1104 | 0.006729 | 0.3266 | 0.002198 | 48.5363 | 0.006927 | 0.21487 | 0.001488 | 31.0196 | 22.3486 | 0.0076 | 0.003686 | 0.7255 | 0.5092 | 0.8376 | 0.0963
2012-10-30 21:47 | 26.6937 | 23.0439 | 22.8556 | 23.172 | 22.4507 | 0.104 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 22.3143 | 0.0072 | 0.00369 | 0.7296 | 0.5413 | 0.8577 | 0.1364
2012-10-30 21:47 | 26.6922 | 23.1025 | 22.8822 | 23.1905 | 22.4251 | 0.1237 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3369 | 0.0085 | 0.003689 | 0.7656 | 0.5453 | 0.8536 | 0.0882
2012-10-30 21:47 | 26.6959 | 23.1022 | 22.9019 | 23.2103 | 22.4569 | 0.1069 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.3607 | 0.0074 | 0.003688 | 0.7415 | 0.5412 | 0.8496 | 0.0962
2012-10-30 21:47 | 26.6915 | 23.0897 | 22.8734 | 23.1658 | 22.4646 | 0.1196 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006933 | 0.21487 | 0.00149 | 30.9928 | 22.3803 | 0.0082 | 0.003688 | 0.7094 | 0.4931 | 0.7855 | 0.0843
2012-10-30 21:47 | 26.687 | 23.0612 | 22.933 | 23.1733 | 22.428 | 0.1121 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.3678 | 0.0077 | 0.003683 | 0.6934 | 0.5652 | 0.8055 | 0.0602
2012-10-30 21:47 | 26.6862 | 23.0604 | 22.8962 | 23.1565 | 22.4673 | 0.1052 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.3309 | 0.0073 | 0.003685 | 0.7295 | 0.5653 | 0.8256 | 0.1364
2012-10-30 21:47 | 26.6757 | 23.0739 | 22.9097 | 23.138 | 22.4447 | 0.1069 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 22.3324 | 0.0074 | 0.003693 | 0.7415 | 0.5773 | 0.8056 | 0.1123
2012-10-30 21:47 | 26.6847 | 23.111 | 22.9067 | 23.2031 | 22.4016 | 0.115 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 22.3695 | 0.0079 | 0.003689 | 0.7415 | 0.5372 | 0.8336 | 0.0321
2012-10-30 21:47 | 26.6877 | 23.11 | 22.8777 | 23.178 | 22.4327 | 0.1058 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.3725 | 0.0073 | 0.003687 | 0.7375 | 0.5052 | 0.8055 | 0.0602
2012-10-30 21:47 | 26.6787 | 23.0249 | 22.8366 | 23.117 | 22.4157 | 0.1023 | 0.006713 | 0.32649 | 0.002192 | 48.6347 | 0.006949 | 0.2148 | 0.001493 | 30.9111 | 22.3354 | 0.0071 | 0.003685 | 0.6895 | 0.5012 | 0.7816 | 0.0803
2012-10-30 21:47 | 26.6907 | 23.105 | 22.8646 | 23.185 | 22.4036 | 0.1179 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 22.3434 | 0.0081 | 0.003688 | 0.7616 | 0.5212 | 0.8416 | 0.0602
2012-10-30 21:47 | 26.6877 | 23.0539 | 22.9217 | 23.182 | 22.4327 | 0.1011 | 0.006713 | 0.3265 | 0.002192 | 48.637 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 22.3685 | 0.007 | 0.003682 | 0.6854 | 0.5532 | 0.8135 | 0.0642
2012-10-30 21:47 | 26.6945 | 23.0607 | 22.8684 | 23.1648 | 22.4274 | 0.0988 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3633 | 0.0068 | 0.003687 | 0.6974 | 0.5051 | 0.8015 | 0.0641
2012-10-30 21:47 | 26.6855 | 23.0717 | 22.8594 | 23.1478 | 22.3984 | 0.1144 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3422 | 0.0079 | 0.003681 | 0.7295 | 0.5172 | 0.8056 | 0.0562
2012-10-30 21:47 | 26.6885 | 23.0307 | 22.8504 | 23.1468 | 22.4415 | 0.1063 | 0.006722 | 0.32653 | 0.002195 | 48.5762 | 0.006946 | 0.21477 | 0.001492 | 30.9202 | 22.3412 | 0.0073 | 0.003687 | 0.6895 | 0.5092 | 0.8056 | 0.1003
2012-10-30 21:47 | 26.684 | 23.0182 | 22.8499 | 23.1343 | 22.4049 | 0.111 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.3447 | 0.0077 | 0.003685 | 0.6735 | 0.5052 | 0.7896 | 0.0602
2012-10-30 21:47 | 26.6952 | 23.0494 | 22.8771 | 23.1895 | 22.4362 | 0.1156 | 0.006741 | 0.32651 | 0.002201 | 48.437 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.368 | 0.008 | 0.00369 | 0.6814 | 0.5091 | 0.8215 | 0.0682
2012-10-30 21:47 | 26.6922 | 23.0664 | 22.8822 | 23.1946 | 22.4292 | 0.1075 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006933 | 0.2147 | 0.001489 | 30.9678 | 22.357 | 0.0074 | 0.003685 | 0.7094 | 0.5252 | 0.8376 | 0.0722
2012-10-30 21:48 | 26.693 | 23.0792 | 22.8709 | 23.1633 | 22.462 | 0.1087 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.3898 | 0.0075 | 0.003691 | 0.6894 | 0.4811 | 0.7735 | 0.0722
2012-10-30 21:48 | 26.6817 | 22.9958 | 22.8877 | 23.132 | 22.4146 | 0.0982 | 0.006725 | 0.32657 | 0.002196 | 48.5609 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3184 | 0.0068 | 0.003685 | 0.6774 | 0.5693 | 0.8136 | 0.0962
2012-10-30 21:48 | 26.6937 | 23.0399 | 22.8837 | 23.12 | 22.4668 | 0.1034 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 22.3544 | 0.0071 | 0.003687 | 0.6855 | 0.5293 | 0.7656 | 0.1124
2012-10-30 21:48 | 26.6945 | 23.0167 | 22.8484 | 23.1888 | 22.3954 | 0.1011 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.3352 | 0.007 | 0.003689 | 0.6815 | 0.5132 | 0.8536 | 0.0602
2012-10-30 21:48 | 26.6855 | 23.0557 | 22.8634 | 23.1398 | 22.3984 | 0.1063 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.3101 | 0.0073 | 0.003687 | 0.7456 | 0.5533 | 0.8297 | 0.0883
2012-10-30 21:48 | 26.6937 | 23.0319 | 22.8957 | 23.148 | 22.4226 | 0.1081 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.3344 | 0.0075 | 0.003685 | 0.6975 | 0.5613 | 0.8136 | 0.0882
2012-10-30 21:48 | 26.6967 | 23.0509 | 22.8546 | 23.119 | 22.4337 | 0.1185 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 22.3534 | 0.0082 | 0.003688 | 0.6975 | 0.5012 | 0.7656 | 0.0803
2012-10-30 21:48 | 26.7019 | 23.0601 | 22.8999 | 23.1682 | 22.4469 | 0.1023 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3185 | 0.0071 | 0.003684 | 0.7416 | 0.5814 | 0.8497 | 0.1284
2012-10-30 21:48 | 26.6952 | 23.0614 | 22.8852 | 23.1015 | 22.4201 | 0.1058 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3399 | 0.0073 | 0.003688 | 0.7215 | 0.5453 | 0.7616 | 0.0802
2012-10-30 21:48 | 26.69 | 23.0362 | 22.8719 | 23.1683 | 22.4189 | 0.1196 | 0.006738 | 0.32657 | 0.0022 | 48.4672 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3628 | 0.0082 | 0.003689 | 0.6734 | 0.5091 | 0.8055 | 0.0561
2012-10-30 21:48 | 26.6952 | 23.0694 | 22.8251 | 23.1455 | 22.4442 | 0.1115 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006933 | 0.21471 | 0.001489 | 30.9699 | 22.3239 | 0.0077 | 0.003689 | 0.7455 | 0.5012 | 0.8216 | 0.1203
2012-10-30 21:48 | 26.6982 | 23.0123 | 22.8922 | 23.1565 | 22.387 | 0.1058 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.3429 | 0.0073 | 0.003687 | 0.6694 | 0.5493 | 0.8136 | 0.0441
2012-10-30 21:48 | 26.6997 | 23.0178 | 22.8616 | 23.114 | 22.4487 | 0.1214 | 0.006741 | 0.32653 | 0.002201 | 48.4393 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3885 | 0.0084 | 0.003691 | 0.6293 | 0.4731 | 0.7255 | 0.0602
2012-10-30 21:48 | 26.6952 | 22.9973 | 22.8731 | 23.1455 | 22.4642 | 0.1075 | 0.006729 | 0.32651 | 0.002197 | 48.5234 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.364 | 0.0074 | 0.003688 | 0.6333 | 0.5091 | 0.7815 | 0.1002
2012-10-30 21:48 | 26.7042 | 23.0424 | 22.8541 | 23.1385 | 22.4371 | 0.1069 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.381 | 0.0074 | 0.003686 | 0.6614 | 0.4731 | 0.7575 | 0.0561
2012-10-30 21:48 | 26.7049 | 23.007 | 22.8628 | 23.1432 | 22.4218 | 0.1034 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3616 | 0.0071 | 0.003688 | 0.6454 | 0.5012 | 0.7816 | 0.0602
2012-10-30 21:48 | 26.6989 | 23.005 | 22.8207 | 23.1092 | 22.4519 | 0.1046 | 0.006744 | 0.32653 | 0.002202 | 48.4177 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 22.3757 | 0.0072 | 0.003692 | 0.6293 | 0.445 | 0.7335 | 0.0762
2012-10-30 21:48 | 26.6952 | 22.9933 | 22.8691 | 23.1495 | 22.4402 | 0.1127 | 0.006722 | 0.32656 | 0.002195 | 48.5803 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.3158 | 0.0078 | 0.003684 | 0.6775 | 0.5533 | 0.8337 | 0.1244
2012-10-30 21:48 | 26.7034 | 22.9895 | 22.8493 | 23.1697 | 22.4444 | 0.0959 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006921 | 0.2148 | 0.001487 | 31.0361 | 22.3481 | 0.0066 | 0.003684 | 0.6414 | 0.5012 | 0.8216 | 0.0963
2012-10-30 21:48 | 26.7027 | 22.9928 | 22.8686 | 23.157 | 22.4316 | 0.1052 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.3675 | 0.0073 | 0.003686 | 0.6253 | 0.5011 | 0.7895 | 0.0641
2012-10-30 21:48 | 26.7079 | 23.026 | 22.9219 | 23.1542 | 22.3887 | 0.1029 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.21487 | 0.001491 | 30.966 | 22.3606 | 0.0071 | 0.003688 | 0.6654 | 0.5613 | 0.7936 | 0.0281
2012-10-30 21:48 | 26.7109 | 22.997 | 22.8568 | 23.1852 | 22.4559 | 0.0988 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.3716 | 0.0068 | 0.003691 | 0.6254 | 0.4852 | 0.8136 | 0.0843
2012-10-30 21:48 | 26.6989 | 23.0251 | 22.8768 | 23.1292 | 22.3797 | 0.1023 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.3155 | 0.0071 | 0.003686 | 0.7096 | 0.5613 | 0.8137 | 0.0642
2012-10-30 21:48 | 26.7072 | 22.9933 | 22.8491 | 23.1215 | 22.4361 | 0.1063 | 0.006754 | 0.32646 | 0.002205 | 48.3352 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.3479 | 0.0073 | 0.003695 | 0.6454 | 0.5012 | 0.7736 | 0.0882
2012-10-30 21:48 | 26.7019 | 23.004 | 22.8438 | 23.1442 | 22.4509 | 0.1046 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3386 | 0.0072 | 0.003688 | 0.6654 | 0.5052 | 0.8056 | 0.1123
2012-10-30 21:48 | 26.7072 | 23.0334 | 22.8731 | 23.1255 | 22.4161 | 0.0936 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3399 | 0.0065 | 0.003683 | 0.6935 | 0.5332 | 0.7856 | 0.0762
2012-10-30 21:49 | 26.7012 | 23.0113 | 22.8391 | 23.1315 | 22.4301 | 0.1046 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 22.3579 | 0.0072 | 0.003688 | 0.6534 | 0.4812 | 0.7736 | 0.0722
2012-10-30 21:49 | 26.6989 | 23.005 | 22.8688 | 23.1332 | 22.4238 | 0.1017 | 0.006713 | 0.32651 | 0.002192 | 48.6391 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.2995 | 0.007 | 0.003682 | 0.7055 | 0.5693 | 0.8337 | 0.1243
2012-10-30 21:49 | 26.7139 | 23.02 | 22.8438 | 23.1362 | 22.4308 | 0.093 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3466 | 0.0064 | 0.003682 | 0.6734 | 0.4972 | 0.7896 | 0.0842
2012-10-30 21:49 | 26.7057 | 23.0078 | 22.8516 | 23.136 | 22.3745 | 0.1087 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3424 | 0.0075 | 0.003687 | 0.6654 | 0.5092 | 0.7936 | 0.0321
2012-10-30 21:49 | 26.7034 | 22.9935 | 22.8252 | 23.1297 | 22.4323 | 0.115 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.3682 | 0.0079 | 0.003689 | 0.6253 | 0.457 | 0.7615 | 0.0641
2012-10-30 21:49 | 26.7064 | 23.0205 | 22.8242 | 23.1006 | 22.4434 | 0.0971 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 22.3511 | 0.0067 | 0.003682 | 0.6694 | 0.4731 | 0.7495 | 0.0923
2012-10-30 21:49 | 26.7079 | 23.014 | 22.8378 | 23.1101 | 22.4208 | 0.1046 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3285 | 0.0072 | 0.003685 | 0.6855 | 0.5093 | 0.7816 | 0.0923
2012-10-30 21:49 | 26.7117 | 23.0419 | 22.8496 | 23.15 | 22.4366 | 0.111 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.3524 | 0.0077 | 0.003686 | 0.6895 | 0.4972 | 0.7976 | 0.0842
2012-10-30 21:49 | 26.6974 | 23.0236 | 22.8593 | 23.0996 | 22.3983 | 0.1081 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.318 | 0.0075 | 0.003684 | 0.7056 | 0.5413 | 0.7816 | 0.0803
2012-10-30 21:49 | 26.6974 | 23.0276 | 22.8353 | 23.1437 | 22.4063 | 0.1075 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.306 | 0.0074 | 0.003684 | 0.7216 | 0.5293 | 0.8377 | 0.1003
2012-10-30 21:49 | 26.6915 | 23.0257 | 22.8374 | 23.1138 | 22.4244 | 0.1058 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 22.3482 | 0.0073 | 0.003692 | 0.6775 | 0.4892 | 0.7656 | 0.0762
2012-10-30 21:49 | 26.693 | 23.0432 | 22.8389 | 23.1473 | 22.446 | 0.1063 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 22.3257 | 0.0073 | 0.003686 | 0.7175 | 0.5132 | 0.8216 | 0.1203
2012-10-30 21:49 | 26.693 | 22.9991 | 22.8589 | 23.1073 | 22.434 | 0.0988 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 22.3497 | 0.0068 | 0.003684 | 0.6494 | 0.5092 | 0.7576 | 0.0843
2012-10-30 21:49 | 26.7012 | 23.0394 | 22.8791 | 23.1195 | 22.4141 | 0.1081 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3539 | 0.0075 | 0.003682 | 0.6855 | 0.5252 | 0.7656 | 0.0602
2012-10-30 21:49 | 26.6907 | 23.0249 | 22.8686 | 23.137 | 22.4397 | 0.1087 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3354 | 0.0075 | 0.003688 | 0.6895 | 0.5332 | 0.8016 | 0.1043
2012-10-30 21:49 | 26.6937 | 23.0279 | 22.8316 | 23.148 | 22.4226 | 0.1179 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006917 | 0.21477 | 0.001486 | 31.0499 | 22.3504 | 0.0081 | 0.003683 | 0.6775 | 0.4812 | 0.7976 | 0.0722
2012-10-30 21:49 | 26.6945 | 23.0607 | 22.8724 | 23.1088 | 22.4475 | 0.0988 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3071 | 0.0068 | 0.003683 | 0.7536 | 0.5653 | 0.8017 | 0.1404
2012-10-30 21:49 | 26.6922 | 23.0624 | 22.8301 | 23.0985 | 22.4251 | 0.1144 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.00693 | 0.21487 | 0.001489 | 31.0062 | 22.3369 | 0.0079 | 0.003686 | 0.7255 | 0.4932 | 0.7616 | 0.0882
2012-10-30 21:49 | 26.6885 | 23.0427 | 22.8344 | 23.0867 | 22.4094 | 0.1011 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.006924 | 0.21479 | 0.001487 | 31.0205 | 22.3292 | 0.007 | 0.003682 | 0.7135 | 0.5052 | 0.7575 | 0.0802
2012-10-30 21:49 | 26.6989 | 23.001 | 22.8408 | 23.1052 | 22.4479 | 0.1115 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 22.3757 | 0.0077 | 0.003683 | 0.6253 | 0.4651 | 0.7295 | 0.0722
2012-10-30 21:49 | 26.6855 | 23.0477 | 22.8073 | 23.0998 | 22.4064 | 0.1058 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 22.3864 | 0.0073 | 0.003682 | 0.6613 | 0.4209 | 0.7134 | 0.02
2012-10-30 21:49 | 26.6892 | 23.0234 | 22.771 | 23.0834 | 22.4061 | 0.1087 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.2978 | 0.0075 | 0.003684 | 0.7256 | 0.4732 | 0.7856 | 0.1083
2012-10-30 21:49 | 26.6862 | 23.0404 | 22.804 | 23.1085 | 22.4151 | 0.1075 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 22.3349 | 0.0074 | 0.003686 | 0.7055 | 0.4691 | 0.7736 | 0.0802
2012-10-30 21:49 | 26.6847 | 23.0069 | 22.8306 | 23.111 | 22.4016 | 0.1133 | 0.006707 | 0.32643 | 0.002189 | 48.6696 | 0.006917 | 0.21479 | 0.001486 | 31.0519 | 22.3695 | 0.0078 | 0.003675 | 0.6374 | 0.4611 | 0.7415 | 0.0321
2012-10-30 21:49 | 26.6802 | 23.0464 | 22.786 | 23.1065 | 22.3891 | 0.1092 | 0.006751 | 0.32643 | 0.002204 | 48.3524 | 0.006911 | 0.21476 | 0.001484 | 31.0748 | 22.3169 | 0.0075 | 0.003688 | 0.7295 | 0.4691 | 0.7896 | 0.0722
2012-10-30 21:49 | 26.684 | 23.0462 | 22.7778 | 23.1103 | 22.441 | 0.1104 | 0.006744 | 0.32646 | 0.002202 | 48.4069 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3207 | 0.0076 | 0.003693 | 0.7255 | 0.4571 | 0.7896 | 0.1203
2012-10-30 21:49 | 26.693 | 23.0912 | 22.8188 | 23.0912 | 22.4259 | 0.1058 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.3377 | 0.0073 | 0.003687 | 0.7535 | 0.4811 | 0.7535 | 0.0882
2012-10-30 21:49 | 26.6787 | 23.0129 | 22.7765 | 23.105 | 22.4116 | 0.1063 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 22.3074 | 0.0073 | 0.003687 | 0.7055 | 0.4691 | 0.7976 | 0.1042
2012-10-30 21:50 | 26.6885 | 23.0307 | 22.7943 | 23.0747 | 22.3813 | 0.1121 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.2931 | 0.0077 | 0.003681 | 0.7376 | 0.5012 | 0.7816 | 0.0882
2012-10-30 21:50 | 26.6825 | 23.0367 | 22.7843 | 23.0727 | 22.3874 | 0.1034 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3553 | 0.0071 | 0.003684 | 0.6814 | 0.429 | 0.7174 | 0.0321
2012-10-30 21:50 | 26.6877 | 23.0419 | 22.7695 | 23.0739 | 22.4046 | 0.1219 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.3645 | 0.0084 | 0.003685 | 0.6774 | 0.405 | 0.7094 | 0.0401
2012-10-30 21:50 | 26.6787 | 23.0289 | 22.7765 | 23.0729 | 22.4116 | 0.1075 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.006927 | 0.21471 | 0.001487 | 30.9967 | 22.3354 | 0.0074 | 0.003685 | 0.6935 | 0.4411 | 0.7375 | 0.0762
2012-10-30 21:50 | 26.6742 | 23.0444 | 22.78 | 23.0925 | 22.3751 | 0.111 | 0.006707 | 0.32638 | 0.002189 | 48.6632 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3189 | 0.0077 | 0.003677 | 0.7255 | 0.4611 | 0.7736 | 0.0562
2012-10-30 21:50 | 26.6847 | 23.0349 | 22.8186 | 23.0829 | 22.3815 | 0.1104 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3294 | 0.0076 | 0.003682 | 0.7055 | 0.4892 | 0.7535 | 0.0521
2012-10-30 21:50 | 26.6742 | 23.0604 | 22.8 | 23.0684 | 22.4152 | 0.1006 | 0.006741 | 0.32643 | 0.0022 | 48.4241 | 0.006924 | 0.21471 | 0.001487 | 31.0101 | 22.3029 | 0.0069 | 0.003687 | 0.7575 | 0.4971 | 0.7655 | 0.1123
2012-10-30 21:50 | 26.6802 | 23.0464 | 22.774 | 23.0945 | 22.4212 | 0.1087 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.2968 | 0.0075 | 0.003691 | 0.7496 | 0.4772 | 0.7977 | 0.1244
2012-10-30 21:50 | 26.666 | 23.0282 | 22.8039 | 23.0763 | 22.3709 | 0.1069 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2786 | 0.0074 | 0.003682 | 0.7496 | 0.5253 | 0.7977 | 0.0923
2012-10-30 21:50 | 26.6675 | 23.0257 | 22.7533 | 23.0377 | 22.3844 | 0.1011 | 0.006707 | 0.32647 | 0.00219 | 48.6762 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.3242 | 0.007 | 0.00368 | 0.7015 | 0.4291 | 0.7135 | 0.0602
2012-10-30 21:50 | 26.669 | 23.0112 | 22.7427 | 23.0552 | 22.3899 | 0.1029 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 22.3458 | 0.0071 | 0.003681 | 0.6654 | 0.3969 | 0.7094 | 0.0441
2012-10-30 21:50 | 26.6653 | 23.0155 | 22.7711 | 23.0435 | 22.3662 | 0.0994 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006917 | 0.21477 | 0.001486 | 31.0499 | 22.3421 | 0.0069 | 0.003683 | 0.6734 | 0.429 | 0.7014 | 0.0241
2012-10-30 21:50 | 26.6653 | 23.0435 | 22.7471 | 23.0635 | 22.4143 | 0.1063 | 0.0067 | 0.32649 | 0.002187 | 48.7291 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.306 | 0.0073 | 0.003676 | 0.7375 | 0.4411 | 0.7575 | 0.1083
2012-10-30 21:50 | 26.6735 | 23.0277 | 22.8034 | 23.0878 | 22.4105 | 0.1115 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3102 | 0.0077 | 0.003684 | 0.7175 | 0.4932 | 0.7776 | 0.1003
2012-10-30 21:50 | 26.6802 | 23.0544 | 22.806 | 23.0744 | 22.369 | 0.1121 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 22.2968 | 0.0077 | 0.003682 | 0.7576 | 0.5092 | 0.7776 | 0.0722
2012-10-30 21:50 | 26.663 | 23.0012 | 22.7408 | 23.0773 | 22.3959 | 0.1029 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 22.2917 | 0.0071 | 0.003685 | 0.7095 | 0.4491 | 0.7856 | 0.1042
2012-10-30 21:50 | 26.6727 | 23.0629 | 22.7505 | 23.083 | 22.3334 | 0.1081 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006921 | 0.21476 | 0.001486 | 31.0299 | 22.2813 | 0.0075 | 0.003683 | 0.7816 | 0.4692 | 0.8017 | 0.0521
2012-10-30 21:50 | 26.6765 | 23.0828 | 22.7863 | 23.0707 | 22.3693 | 0.1069 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.265 | 0.0074 | 0.003689 | 0.8178 | 0.5213 | 0.8057 | 0.1043
2012-10-30 21:50 | 26.6765 | 23.0507 | 22.7663 | 23.0627 | 22.3332 | 0.1017 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006917 | 0.21482 | 0.001486 | 31.0561 | 22.257 | 0.007 | 0.003684 | 0.7937 | 0.5093 | 0.8057 | 0.0762
2012-10-30 21:50 | 26.6757 | 23.0179 | 22.7735 | 23.0419 | 22.3846 | 0.1034 | 0.006703 | 0.32643 | 0.002188 | 48.6986 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.2562 | 0.0071 | 0.003679 | 0.7617 | 0.5173 | 0.7857 | 0.1284
2012-10-30 21:50 | 26.675 | 23.0412 | 22.7528 | 23.0252 | 22.4079 | 0.1058 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.006936 | 0.21476 | 0.00149 | 30.9628 | 22.2675 | 0.0073 | 0.003686 | 0.7737 | 0.4853 | 0.7577 | 0.1404
2012-10-30 21:50 | 26.6735 | 23.0878 | 22.7553 | 23.0717 | 22.3784 | 0.1069 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.238 | 0.0074 | 0.003683 | 0.8498 | 0.5173 | 0.8337 | 0.1404
2012-10-30 21:50 | 26.6765 | 23.0067 | 22.7583 | 23.0507 | 22.3934 | 0.115 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.2209 | 0.0079 | 0.003688 | 0.7858 | 0.5374 | 0.8298 | 0.1725
2012-10-30 21:50 | 26.6787 | 23.0449 | 22.7765 | 23.0289 | 22.3876 | 0.1133 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.2913 | 0.0078 | 0.00369 | 0.7536 | 0.4852 | 0.7376 | 0.0963
2012-10-30 21:50 | 26.678 | 23.0162 | 22.7678 | 23.0722 | 22.3829 | 0.1139 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.2705 | 0.0079 | 0.003681 | 0.7457 | 0.4973 | 0.8017 | 0.1124
2012-10-30 21:50 | 26.675 | 23.0412 | 22.7407 | 23.0372 | 22.3598 | 0.1069 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.2996 | 0.0074 | 0.003682 | 0.7416 | 0.4411 | 0.7376 | 0.0602
2012-10-30 21:50 | 26.6742 | 23.0124 | 22.752 | 23.0204 | 22.3831 | 0.1121 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.2547 | 0.0077 | 0.003686 | 0.7577 | 0.4973 | 0.7657 | 0.1284
2012-10-30 21:50 | 26.6735 | 23.0037 | 22.7272 | 23.0798 | 22.3623 | 0.1058 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3102 | 0.0073 | 0.003688 | 0.6935 | 0.417 | 0.7696 | 0.0521
2012-10-30 21:50 | 26.6705 | 23.0527 | 22.7322 | 23.0327 | 22.3754 | 0.1029 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006911 | 0.21477 | 0.001484 | 31.0768 | 22.2831 | 0.0071 | 0.003682 | 0.7696 | 0.4491 | 0.7496 | 0.0923
2012-10-30 21:50 | 26.669 | 23.0072 | 22.7267 | 23.0352 | 22.3378 | 0.0994 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.2575 | 0.0069 | 0.003683 | 0.7497 | 0.4692 | 0.7777 | 0.0803
2012-10-30 21:51 | 26.6727 | 23.0069 | 22.7384 | 23.0469 | 22.3816 | 0.1058 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.2893 | 0.0073 | 0.003687 | 0.7176 | 0.4491 | 0.7576 | 0.0923
2012-10-30 21:51 | 26.675 | 22.9691 | 22.7447 | 23.0132 | 22.3678 | 0.1029 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.2756 | 0.0071 | 0.003688 | 0.6935 | 0.4691 | 0.7376 | 0.0922
2012-10-30 21:51 | 26.6757 | 23.0099 | 22.7494 | 23.0499 | 22.3485 | 0.111 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006921 | 0.21486 | 0.001487 | 31.0445 | 22.3084 | 0.0077 | 0.003686 | 0.7015 | 0.441 | 0.7415 | 0.0401
2012-10-30 21:51 | 26.6735 | 22.9957 | 22.7513 | 23.0197 | 22.3864 | 0.1167 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.006949 | 0.21486 | 0.001493 | 30.9194 | 22.3383 | 0.008 | 0.003686 | 0.6574 | 0.413 | 0.6814 | 0.0481
2012-10-30 21:51 | 26.6765 | 22.9987 | 22.7743 | 23.0587 | 22.3413 | 0.1 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.2971 | 0.0069 | 0.00369 | 0.7016 | 0.4772 | 0.7616 | 0.0442
2012-10-30 21:51 | 26.6712 | 22.9893 | 22.753 | 23.0254 | 22.364 | 0.1081 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.2999 | 0.0075 | 0.003688 | 0.6894 | 0.4531 | 0.7255 | 0.0641
2012-10-30 21:51 | 26.6877 | 22.9698 | 22.7574 | 23.0579 | 22.3805 | 0.1034 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3404 | 0.0071 | 0.003684 | 0.6294 | 0.417 | 0.7175 | 0.0401
2012-10-30 21:51 | 26.6742 | 22.9964 | 22.756 | 23.0324 | 22.371 | 0.1029 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.2988 | 0.0071 | 0.00369 | 0.6976 | 0.4572 | 0.7336 | 0.0722
2012-10-30 21:51 | 26.6765 | 22.9746 | 22.7583 | 23.0307 | 22.3413 | 0.111 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3172 | 0.0077 | 0.003689 | 0.6574 | 0.4411 | 0.7135 | 0.0241
2012-10-30 21:51 | 26.6742 | 23.0284 | 22.7399 | 23.0324 | 22.3229 | 0.1023 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 22.3069 | 0.0071 | 0.00369 | 0.7215 | 0.433 | 0.7255 | 0.016
2012-10-30 21:51 | 26.6817 | 22.9598 | 22.7114 | 23.0239 | 22.3384 | 0.115 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006917 | 0.21487 | 0.001486 | 31.0645 | 22.2863 | 0.0079 | 0.003678 | 0.6735 | 0.4251 | 0.7376 | 0.0521
2012-10-30 21:51 | 26.672 | 22.9982 | 22.7538 | 22.9942 | 22.3488 | 0.0977 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006917 | 0.21492 | 0.001487 | 31.0707 | 22.2645 | 0.0067 | 0.003685 | 0.7337 | 0.4893 | 0.7297 | 0.0843
2012-10-30 21:51 | 26.6772 | 22.9873 | 22.755 | 23.0434 | 22.354 | 0.0925 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 22.2376 | 0.0064 | 0.003689 | 0.7497 | 0.5174 | 0.8058 | 0.1164
2012-10-30 21:51 | 26.6862 | 22.9763 | 22.7199 | 23.0204 | 22.3951 | 0.1046 | 0.006725 | 0.32653 | 0.002196 | 48.5545 | 0.006943 | 0.21487 | 0.001492 | 30.9481 | 22.2868 | 0.0072 | 0.003688 | 0.6895 | 0.4331 | 0.7336 | 0.1083
2012-10-30 21:51 | 26.684 | 22.9941 | 22.7377 | 22.9981 | 22.3608 | 0.1098 | 0.006741 | 0.32654 | 0.002201 | 48.4413 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3046 | 0.0076 | 0.00369 | 0.6895 | 0.4331 | 0.6935 | 0.0562
2012-10-30 21:51 | 26.6885 | 23.0387 | 22.7182 | 23.0187 | 22.3613 | 0.1017 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.2489 | 0.007 | 0.003687 | 0.7898 | 0.4693 | 0.7698 | 0.1124
2012-10-30 21:51 | 26.6772 | 23.0394 | 22.7189 | 22.9953 | 22.3339 | 0.1087 | 0.006729 | 0.32653 | 0.002197 | 48.5256 | 0.006933 | 0.21489 | 0.00149 | 30.995 | 22.2818 | 0.0075 | 0.003687 | 0.7576 | 0.4371 | 0.7135 | 0.0521
2012-10-30 21:51 | 26.6795 | 23.0017 | 22.7973 | 23.0217 | 22.3683 | 0.1029 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.3001 | 0.0071 | 0.003688 | 0.7016 | 0.4972 | 0.7216 | 0.0682
2012-10-30 21:51 | 26.6675 | 23.0257 | 22.7172 | 23.0297 | 22.3323 | 0.1115 | 0.006722 | 0.32659 | 0.002195 | 48.5846 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.3042 | 0.0077 | 0.003685 | 0.7215 | 0.413 | 0.7255 | 0.0281
2012-10-30 21:51 | 26.6787 | 22.9808 | 22.7364 | 23.0369 | 22.3595 | 0.1173 | 0.006725 | 0.32654 | 0.002196 | 48.5566 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.2632 | 0.0081 | 0.003685 | 0.7176 | 0.4732 | 0.7737 | 0.0963
2012-10-30 21:51 | 26.6787 | 22.9968 | 22.7565 | 23.0249 | 22.3836 | 0.1104 | 0.006738 | 0.3265 | 0.0022 | 48.4565 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.2913 | 0.0076 | 0.003691 | 0.7055 | 0.4652 | 0.7336 | 0.0923
2012-10-30 21:51 | 26.6847 | 22.9868 | 22.7224 | 23.0309 | 22.3374 | 0.1098 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.2572 | 0.0076 | 0.003688 | 0.7296 | 0.4652 | 0.7737 | 0.0802
2012-10-30 21:51 | 26.6727 | 23.0109 | 22.7344 | 23.0069 | 22.3655 | 0.0994 | 0.006732 | 0.32656 | 0.002198 | 48.5082 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.2773 | 0.0069 | 0.003688 | 0.7336 | 0.4571 | 0.7296 | 0.0882
2012-10-30 21:51 | 26.6742 | 22.9283 | 22.7079 | 23.0324 | 22.3349 | 0.0936 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.2908 | 0.0065 | 0.003685 | 0.6375 | 0.4171 | 0.7416 | 0.0441
2012-10-30 21:51 | 26.6772 | 23.0074 | 22.6908 | 23.0474 | 22.374 | 0.1034 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006917 | 0.21479 | 0.001486 | 31.0519 | 22.3139 | 0.0071 | 0.003685 | 0.6935 | 0.3769 | 0.7335 | 0.0601
2012-10-30 21:51 | 26.6757 | 22.9778 | 22.6853 | 23.0459 | 22.3445 | 0.1052 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.2723 | 0.0073 | 0.003685 | 0.7055 | 0.413 | 0.7736 | 0.0722
2012-10-30 21:51 | 26.6772 | 23.0194 | 22.767 | 23.0074 | 22.366 | 0.0959 | 0.006747 | 0.32647 | 0.002203 | 48.3876 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.2818 | 0.0066 | 0.00369 | 0.7376 | 0.4852 | 0.7256 | 0.0842
2012-10-30 21:51 | 26.672 | 23.0102 | 22.7417 | 23.0502 | 22.3408 | 0.1069 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.2485 | 0.0074 | 0.00368 | 0.7617 | 0.4932 | 0.8017 | 0.0923
2012-10-30 21:52 | 26.6697 | 22.9838 | 22.7394 | 23.0199 | 22.3585 | 0.1104 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3264 | 0.0076 | 0.003681 | 0.6574 | 0.413 | 0.6935 | 0.0321
2012-10-30 21:52 | 26.678 | 23.0202 | 22.7197 | 23.0122 | 22.3748 | 0.1029 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 22.3147 | 0.0071 | 0.00369 | 0.7055 | 0.405 | 0.6975 | 0.0601
2012-10-30 21:52 | 26.669 | 22.9711 | 22.7227 | 22.9912 | 22.3538 | 0.1075 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.006943 | 0.21484 | 0.001492 | 30.944 | 22.2615 | 0.0074 | 0.003683 | 0.7096 | 0.4612 | 0.7297 | 0.0923
2012-10-30 21:52 | 26.6757 | 23.0219 | 22.7374 | 22.9818 | 22.3565 | 0.0988 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.2963 | 0.0068 | 0.003685 | 0.7256 | 0.4411 | 0.6855 | 0.0602
2012-10-30 21:52 | 26.6787 | 23.0009 | 22.7124 | 23.0329 | 22.4197 | 0.0942 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 22.3715 | 0.0065 | 0.00369 | 0.6294 | 0.3409 | 0.6614 | 0.0482
2012-10-30 21:52 | 26.6832 | 23.0094 | 22.7369 | 23.0174 | 22.352 | 0.1069 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.006949 | 0.21487 | 0.001493 | 30.9214 | 22.3239 | 0.0074 | 0.003688 | 0.6855 | 0.413 | 0.6935 | 0.0281
2012-10-30 21:52 | 26.6787 | 22.9528 | 22.7324 | 23.0129 | 22.3916 | 0.1202 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.3555 | 0.0083 | 0.003691 | 0.5973 | 0.3769 | 0.6574 | 0.0361
2012-10-30 21:52 | 26.678 | 22.9601 | 22.7477 | 22.9961 | 22.3508 | 0.1029 | 0.006719 | 0.32653 | 0.002194 | 48.5979 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3428 | 0.0071 | 0.003684 | 0.6173 | 0.4049 | 0.6533 | 0.008
2012-10-30 21:52 | 26.6832 | 23.0094 | 22.7529 | 23.0294 | 22.368 | 0.1104 | 0.006735 | 0.32653 | 0.002199 | 48.4824 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.364 | 0.0076 | 0.003689 | 0.6454 | 0.3889 | 0.6654 | 0.004
2012-10-30 21:52 | 26.6817 | 22.9798 | 22.7474 | 23.0439 | 22.3826 | 0.0982 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.3785 | 0.0068 | 0.003688 | 0.6013 | 0.3689 | 0.6654 | 0.0041
2012-10-30 21:52 | 26.675 | 22.9611 | 22.7087 | 22.9811 | 22.3638 | 0.0982 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3518 | 0.0068 | 0.003688 | 0.6093 | 0.3569 | 0.6293 | 0.012
2012-10-30 21:52 | 26.6727 | 22.9548 | 22.7104 | 23.0149 | 22.3495 | 0.0994 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006946 | 0.21489 | 0.001493 | 30.9369 | 22.3415 | 0.0069 | 0.003687 | 0.6133 | 0.3689 | 0.6734 | 0.008
2012-10-30 21:52 | 26.6787 | 23.0009 | 22.7805 | 23.0169 | 22.3394 | 0.1069 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.3234 | 0.0074 | 0.003689 | 0.6775 | 0.4571 | 0.6935 | 0.016
2012-10-30 21:52 | 26.6817 | 22.9558 | 22.7595 | 22.9918 | 22.3264 | 0.0948 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3505 | 0.0065 | 0.003688 | 0.6053 | 0.409 | 0.6413 | -0.0241
2012-10-30 21:52 | 26.678 | 23.0082 | 22.7678 | 22.9881 | 22.3668 | 0.1023 | 0.006741 | 0.32643 | 0.0022 | 48.4241 | 0.006921 | 0.21483 | 0.001487 | 31.0403 | 22.3307 | 0.0071 | 0.003687 | 0.6775 | 0.4371 | 0.6574 | 0.0361
2012-10-30 21:52 | 26.6757 | 22.9498 | 22.7615 | 23.0259 | 22.3405 | 0.1029 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3405 | 0.0071 | 0.003687 | 0.6093 | 0.421 | 0.6854 | 0
2012-10-30 21:52 | 26.675 | 22.9731 | 22.7487 | 22.9891 | 22.3759 | 0.1087 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 22.2916 | 0.0075 | 0.00369 | 0.6815 | 0.4571 | 0.6975 | 0.0843
2012-10-30 21:52 | 26.684 | 22.9981 | 22.7217 | 23.0021 | 22.3808 | 0.0994 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 22.3247 | 0.0069 | 0.003688 | 0.6734 | 0.397 | 0.6774 | 0.0561
2012-10-30 21:52 | 26.6757 | 23.0219 | 22.7334 | 23.0019 | 22.3364 | 0.1063 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.3164 | 0.0073 | 0.003682 | 0.7055 | 0.417 | 0.6855 | 0.02
2012-10-30 21:52 | 26.6795 | 22.9976 | 22.7532 | 22.9576 | 22.3603 | 0.0982 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.2921 | 0.0068 | 0.003691 | 0.7055 | 0.4611 | 0.6655 | 0.0682
2012-10-30 21:52 | 26.6757 | 22.9618 | 22.6813 | 22.9578 | 22.3605 | 0.0959 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006921 | 0.2148 | 0.001487 | 31.0361 | 22.3084 | 0.0066 | 0.003682 | 0.6534 | 0.3729 | 0.6494 | 0.0521
2012-10-30 21:53 | 26.6802 | 22.9663 | 22.7339 | 22.9903 | 22.357 | 0.1063 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3008 | 0.0073 | 0.003687 | 0.6655 | 0.4331 | 0.6895 | 0.0562
2012-10-30 21:53 | 26.6787 | 23.0329 | 22.6923 | 23.0049 | 22.3956 | 0.0965 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.199 | 0.0067 | 0.003684 | 0.8339 | 0.4933 | 0.8059 | 0.1966
2012-10-30 21:53 | 26.6795 | 22.9336 | 22.6971 | 23.0137 | 22.3563 | 0.1156 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.248 | 0.008 | 0.003686 | 0.6856 | 0.4491 | 0.7657 | 0.1083
2012-10-30 21:53 | 26.681 | 22.9351 | 22.7227 | 22.9871 | 22.3658 | 0.1017 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.1893 | 0.007 | 0.003687 | 0.7458 | 0.5334 | 0.7978 | 0.1765
2012-10-30 21:53 | 26.6787 | 22.9448 | 22.7084 | 22.9328 | 22.3114 | 0.0936 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.2231 | 0.0065 | 0.003682 | 0.7217 | 0.4853 | 0.7097 | 0.0883
2012-10-30 21:53 | 26.6847 | 22.9828 | 22.7384 | 22.9788 | 22.3374 | 0.0965 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.2692 | 0.0067 | 0.003684 | 0.7136 | 0.4692 | 0.7096 | 0.0682
2012-10-30 21:53 | 26.6742 | 22.9002 | 22.7319 | 22.9763 | 22.371 | 0.0988 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.2146 | 0.0068 | 0.003685 | 0.6856 | 0.5173 | 0.7617 | 0.1564
2012-10-30 21:53 | 26.6787 | 22.9688 | 22.7204 | 22.9368 | 22.3194 | 0.1075 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006927 | 0.21486 | 0.001488 | 31.0176 | 22.191 | 0.0074 | 0.003683 | 0.7778 | 0.5294 | 0.7458 | 0.1284
2012-10-30 21:53 | 26.6832 | 22.9853 | 22.7008 | 22.9533 | 22.3159 | 0.1029 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.2356 | 0.0071 | 0.003685 | 0.7497 | 0.4652 | 0.7177 | 0.0803
2012-10-30 21:53 | 26.6825 | 22.9526 | 22.7522 | 22.9846 | 22.3352 | 0.1052 | 0.006735 | 0.32651 | 0.002199 | 48.4802 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.263 | 0.0073 | 0.003688 | 0.6896 | 0.4892 | 0.7216 | 0.0722
2012-10-30 21:53 | 26.6795 | 22.9496 | 22.7332 | 23.0057 | 22.3483 | 0.0965 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 22.2399 | 0.0067 | 0.00369 | 0.7097 | 0.4933 | 0.7658 | 0.1084
2012-10-30 21:53 | 26.6802 | 22.9182 | 22.7499 | 22.9623 | 22.3249 | 0.0959 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.2246 | 0.0066 | 0.003686 | 0.6936 | 0.5253 | 0.7377 | 0.1003
2012-10-30 21:53 | 26.6795 | 22.9536 | 22.7532 | 22.9496 | 22.3322 | 0.0954 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006924 | 0.21484 | 0.001488 | 31.0289 | 22.2159 | 0.0066 | 0.003683 | 0.7377 | 0.5373 | 0.7337 | 0.1163
2012-10-30 21:53 | 26.6735 | 22.8995 | 22.6751 | 22.9276 | 22.3463 | 0.0942 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.242 | 0.0065 | 0.003688 | 0.6575 | 0.4331 | 0.6856 | 0.1043
2012-10-30 21:53 | 26.6802 | 22.9463 | 22.6978 | 22.9783 | 22.3008 | 0.1075 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 22.2527 | 0.0074 | 0.003687 | 0.6936 | 0.4451 | 0.7256 | 0.0481
2012-10-30 21:53 | 26.6727 | 22.9268 | 22.6823 | 22.9388 | 22.3094 | 0.1069 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006939 | 0.21487 | 0.001491 | 30.966 | 22.2331 | 0.0074 | 0.003685 | 0.6937 | 0.4492 | 0.7057 | 0.0763
2012-10-30 21:53 | 26.6653 | 22.9434 | 22.6789 | 22.9634 | 22.322 | 0.1087 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006921 | 0.21484 | 0.001487 | 31.0423 | 22.1936 | 0.0075 | 0.003681 | 0.7498 | 0.4853 | 0.7698 | 0.1284
2012-10-30 21:53 | 26.6757 | 22.9418 | 22.7134 | 22.9498 | 22.3645 | 0.1052 | 0.006744 | 0.32647 | 0.002202 | 48.4091 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 22.2161 | 0.0073 | 0.00369 | 0.7257 | 0.4973 | 0.7337 | 0.1484
2012-10-30 21:53 | 26.6675 | 22.9175 | 22.6931 | 22.9216 | 22.3403 | 0.1075 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.26 | 0.0074 | 0.003685 | 0.6575 | 0.4331 | 0.6616 | 0.0803
2012-10-30 21:53 | 26.6742 | 22.9563 | 22.7359 | 22.9563 | 22.351 | 0.1075 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.2387 | 0.0074 | 0.003683 | 0.7176 | 0.4972 | 0.7176 | 0.1123
2012-10-30 21:53 | 26.6757 | 22.9538 | 22.7374 | 22.9818 | 22.3204 | 0.1017 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.00693 | 0.21473 | 0.001488 | 30.9854 | 22.2482 | 0.007 | 0.003686 | 0.7056 | 0.4892 | 0.7336 | 0.0722
2012-10-30 21:53 | 26.6735 | 22.9676 | 22.7192 | 22.9596 | 22.3463 | 0.1058 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2701 | 0.0073 | 0.003686 | 0.6975 | 0.4491 | 0.6895 | 0.0762
2012-10-30 21:53 | 26.6593 | 22.9454 | 22.721 | 22.9254 | 22.3281 | 0.0873 | 0.00671 | 0.32643 | 0.00219 | 48.6478 | 0.006924 | 0.21486 | 0.001488 | 31.0311 | 22.2518 | 0.006 | 0.003678 | 0.6936 | 0.4692 | 0.6736 | 0.0763
2012-10-30 21:53 | 26.6712 | 22.9853 | 22.7049 | 22.9773 | 22.364 | 0.0994 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2597 | 0.0069 | 0.00369 | 0.7256 | 0.4452 | 0.7176 | 0.1043
2012-10-30 21:53 | 26.6705 | 22.8965 | 22.6921 | 22.9967 | 22.3754 | 0.1046 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.243 | 0.0072 | 0.003681 | 0.6535 | 0.4491 | 0.7537 | 0.1324
2012-10-30 21:54 | 26.6563 | 22.9224 | 22.6819 | 22.9945 | 22.3371 | 0.1011 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.2368 | 0.007 | 0.003688 | 0.6856 | 0.4451 | 0.7577 | 0.1003
2012-10-30 21:54 | 26.6608 | 23.015 | 22.7025 | 22.9269 | 22.3376 | 0.0913 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.2614 | 0.0063 | 0.003684 | 0.7536 | 0.4411 | 0.6655 | 0.0762
2012-10-30 21:54 | 26.6638 | 22.9219 | 22.7175 | 22.9579 | 22.3286 | 0.0948 | 0.006713 | 0.32651 | 0.002192 | 48.6391 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 22.2363 | 0.0065 | 0.003683 | 0.6856 | 0.4812 | 0.7216 | 0.0923
2012-10-30 21:54 | 26.657 | 22.9271 | 22.7147 | 22.9591 | 22.3017 | 0.111 | 0.006741 | 0.32653 | 0.002201 | 48.4393 | 0.006927 | 0.21487 | 0.001488 | 31.0196 | 22.1974 | 0.0077 | 0.003689 | 0.7297 | 0.5173 | 0.7617 | 0.1043
2012-10-30 21:54 | 26.6585 | 22.9326 | 22.6841 | 22.9526 | 22.3353 | 0.1023 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.2189 | 0.0071 | 0.003687 | 0.7137 | 0.4652 | 0.7337 | 0.1164
2012-10-30 21:54 | 26.6615 | 22.9396 | 22.6751 | 22.9636 | 22.3343 | 0.1075 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006943 | 0.21474 | 0.001491 | 30.9294 | 22.1939 | 0.0074 | 0.003685 | 0.7457 | 0.4812 | 0.7697 | 0.1404
2012-10-30 21:54 | 26.6578 | 22.8998 | 22.7195 | 22.9519 | 22.3185 | 0.1098 | 0.006744 | 0.32653 | 0.002202 | 48.4177 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 22.2263 | 0.0076 | 0.003691 | 0.6735 | 0.4932 | 0.7256 | 0.0922
2012-10-30 21:54 | 26.6668 | 22.9329 | 22.6965 | 22.9569 | 22.3436 | 0.104 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.2313 | 0.0072 | 0.003688 | 0.7016 | 0.4652 | 0.7256 | 0.1123
2012-10-30 21:54 | 26.657 | 22.9311 | 22.6706 | 22.907 | 22.3418 | 0.0977 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006939 | 0.21486 | 0.001491 | 30.964 | 22.2295 | 0.0067 | 0.00369 | 0.7016 | 0.4411 | 0.6775 | 0.1123
2012-10-30 21:54 | 26.6585 | 22.9366 | 22.7082 | 22.9646 | 22.3233 | 0.1 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006924 | 0.21486 | 0.001488 | 31.0311 | 22.2189 | 0.0069 | 0.003682 | 0.7177 | 0.4893 | 0.7457 | 0.1044
2012-10-30 21:54 | 26.6548 | 22.9329 | 22.7005 | 22.9329 | 22.3276 | 0.1058 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006933 | 0.21486 | 0.00149 | 30.9908 | 22.2313 | 0.0073 | 0.003686 | 0.7016 | 0.4692 | 0.7016 | 0.0963
2012-10-30 21:54 | 26.6548 | 22.8688 | 22.7245 | 22.9008 | 22.2915 | 0.1 | 0.006741 | 0.3265 | 0.002201 | 48.435 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.2433 | 0.0069 | 0.003691 | 0.6255 | 0.4812 | 0.6575 | 0.0482
2012-10-30 21:54 | 26.6503 | 22.9204 | 22.6639 | 22.9044 | 22.3271 | 0.1075 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006914 | 0.21484 | 0.001485 | 31.0738 | 22.2308 | 0.0074 | 0.003683 | 0.6896 | 0.4331 | 0.6736 | 0.0963
2012-10-30 21:54 | 26.6555 | 22.8975 | 22.7172 | 22.9176 | 22.3002 | 0.1 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 22.2119 | 0.0069 | 0.003686 | 0.6856 | 0.5053 | 0.7057 | 0.0883
2012-10-30 21:54 | 26.651 | 22.9331 | 22.7087 | 22.9331 | 22.3398 | 0.0954 | 0.006713 | 0.32643 | 0.002191 | 48.6261 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 22.2235 | 0.0066 | 0.003684 | 0.7096 | 0.4852 | 0.7096 | 0.1163
2012-10-30 21:54 | 26.645 | 22.9191 | 22.6907 | 22.9351 | 22.3298 | 0.1017 | 0.006738 | 0.32653 | 0.0022 | 48.4608 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.2175 | 0.007 | 0.003689 | 0.7016 | 0.4732 | 0.7176 | 0.1123
2012-10-30 21:54 | 26.6548 | 22.9089 | 22.7125 | 22.9289 | 22.3356 | 0.1104 | 0.006738 | 0.32651 | 0.0022 | 48.4586 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.1872 | 0.0076 | 0.00369 | 0.7217 | 0.5253 | 0.7417 | 0.1484
2012-10-30 21:54 | 26.6428 | 22.9209 | 22.6765 | 22.9049 | 22.3437 | 0.0977 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.2073 | 0.0067 | 0.003686 | 0.7136 | 0.4692 | 0.6976 | 0.1364
2012-10-30 21:54 | 26.6421 | 22.9122 | 22.6718 | 22.9482 | 22.3349 | 0.1144 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.2146 | 0.0079 | 0.003687 | 0.6976 | 0.4572 | 0.7336 | 0.1203
2012-10-30 21:54 | 26.6465 | 22.9046 | 22.7082 | 22.9446 | 22.3393 | 0.104 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006936 | 0.2149 | 0.001491 | 30.9836 | 22.1949 | 0.0072 | 0.003688 | 0.7097 | 0.5133 | 0.7497 | 0.1444
2012-10-30 21:54 | 26.6436 | 22.9177 | 22.7013 | 22.9617 | 22.3364 | 0.0994 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006946 | 0.21486 | 0.001492 | 30.9328 | 22.2121 | 0.0069 | 0.003688 | 0.7056 | 0.4892 | 0.7496 | 0.1243
2012-10-30 21:54 | 26.6473 | 22.9174 | 22.713 | 22.9134 | 22.3241 | 0.1023 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 22.2158 | 0.0071 | 0.003688 | 0.7016 | 0.4972 | 0.6976 | 0.1083
2012-10-30 21:54 | 26.6436 | 22.9417 | 22.6853 | 22.9537 | 22.3645 | 0.0861 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 22.196 | 0.0059 | 0.003689 | 0.7457 | 0.4893 | 0.7577 | 0.1685
2012-10-30 21:54 | 26.6338 | 22.9039 | 22.6675 | 22.9199 | 22.3106 | 0.104 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.2023 | 0.0072 | 0.003684 | 0.7016 | 0.4652 | 0.7176 | 0.1083
2012-10-30 21:54 | 26.6293 | 22.9154 | 22.687 | 22.9114 | 22.3261 | 0.1069 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006917 | 0.21483 | 0.001486 | 31.0583 | 22.2539 | 0.0074 | 0.003684 | 0.6615 | 0.4331 | 0.6575 | 0.0722
2012-10-30 21:54 | 26.6428 | 22.9489 | 22.7085 | 22.9369 | 22.3437 | 0.1034 | 0.006735 | 0.3265 | 0.002199 | 48.4781 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.2353 | 0.0071 | 0.003689 | 0.7136 | 0.4732 | 0.7016 | 0.1084
2012-10-30 21:54 | 26.6278 | 22.9259 | 22.7056 | 22.958 | 22.3447 | 0.1 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2043 | 0.0069 | 0.003687 | 0.7216 | 0.5013 | 0.7537 | 0.1404
2012-10-30 21:55 | 26.6203 | 22.8904 | 22.6459 | 22.8744 | 22.3372 | 0.0948 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.1928 | 0.0065 | 0.003684 | 0.6976 | 0.4531 | 0.6816 | 0.1444
2012-10-30 21:55 | 26.6376 | 22.9277 | 22.6953 | 22.9477 | 22.2983 | 0.1075 | 0.006744 | 0.32649 | 0.002202 | 48.4112 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.2021 | 0.0074 | 0.003692 | 0.7256 | 0.4932 | 0.7456 | 0.0962
2012-10-30 21:55 | 26.6301 | 22.9563 | 22.6798 | 22.9322 | 22.3269 | 0.1052 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 22.1865 | 0.0073 | 0.003687 | 0.7698 | 0.4933 | 0.7457 | 0.1404
2012-10-30 21:55 | 26.6346 | 22.9407 | 22.7043 | 22.9527 | 22.3314 | 0.0988 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006917 | 0.21486 | 0.001486 | 31.0625 | 22.2191 | 0.0068 | 0.003683 | 0.7216 | 0.4852 | 0.7336 | 0.1123
2012-10-30 21:55 | 26.6346 | 22.9247 | 22.7083 | 22.9527 | 22.2833 | 0.1063 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.195 | 0.0073 | 0.003684 | 0.7297 | 0.5133 | 0.7577 | 0.0883
2012-10-30 21:55 | 26.6361 | 22.9583 | 22.6938 | 22.9422 | 22.345 | 0.1063 | 0.006719 | 0.32651 | 0.002194 | 48.5956 | 0.006943 | 0.21484 | 0.001492 | 30.944 | 22.2166 | 0.0073 | 0.003686 | 0.7417 | 0.4772 | 0.7256 | 0.1284
2012-10-30 21:55 | 26.6346 | 22.9047 | 22.6562 | 22.9568 | 22.3274 | 0.0977 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.2151 | 0.0067 | 0.003683 | 0.6896 | 0.4411 | 0.7417 | 0.1123
2012-10-30 21:55 | 26.6286 | 22.8947 | 22.6462 | 22.9027 | 22.3375 | 0.0982 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.2171 | 0.0068 | 0.003688 | 0.6776 | 0.4291 | 0.6856 | 0.1204
2012-10-30 21:55 | 26.6278 | 22.954 | 22.6534 | 22.9139 | 22.3567 | 0.1144 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006933 | 0.21487 | 0.00149 | 30.9928 | 22.1682 | 0.0079 | 0.003686 | 0.7858 | 0.4852 | 0.7457 | 0.1885
2012-10-30 21:55 | 26.6323 | 22.8944 | 22.662 | 22.8904 | 22.293 | 0.0936 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.006955 | 0.2148 | 0.001494 | 30.8844 | 22.1727 | 0.0065 | 0.003687 | 0.7217 | 0.4893 | 0.7177 | 0.1203
2012-10-30 21:55 | 26.6211 | 22.9272 | 22.6467 | 22.8752 | 22.2899 | 0.1 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 22.1575 | 0.0069 | 0.003683 | 0.7697 | 0.4892 | 0.7177 | 0.1324
2012-10-30 21:55 | 26.6278 | 22.9099 | 22.6294 | 22.8979 | 22.3327 | 0.0965 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 22.1401 | 0.0067 | 0.003681 | 0.7698 | 0.4893 | 0.7578 | 0.1926
2012-10-30 21:55 | 26.6218 | 22.8919 | 22.6835 | 22.9399 | 22.3387 | 0.104 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.1943 | 0.0072 | 0.003684 | 0.6976 | 0.4892 | 0.7456 | 0.1444
2012-10-30 21:55 | 26.6241 | 22.9182 | 22.6658 | 22.9302 | 22.337 | 0.1069 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2327 | 0.0074 | 0.003686 | 0.6855 | 0.4331 | 0.6975 | 0.1043
2012-10-30 21:55 | 26.6248 | 22.9149 | 22.7146 | 22.955 | 22.2936 | 0.1017 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.1852 | 0.007 | 0.003683 | 0.7297 | 0.5294 | 0.7698 | 0.1084
2012-10-30 21:55 | 26.6226 | 22.9407 | 22.6803 | 22.9247 | 22.3194 | 0.1029 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.1951 | 0.0071 | 0.003686 | 0.7456 | 0.4852 | 0.7296 | 0.1243
2012-10-30 21:55 | 26.6211 | 22.9513 | 22.7189 | 22.9112 | 22.3059 | 0.0873 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.2658 | 0.006 | 0.003686 | 0.6855 | 0.4531 | 0.6454 | 0.0401
2012-10-30 21:55 | 26.6189 | 22.885 | 22.6886 | 22.913 | 22.3037 | 0.1052 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 22.2155 | 0.0073 | 0.003683 | 0.6695 | 0.4731 | 0.6975 | 0.0882
2012-10-30 21:55 | 26.6218 | 22.8718 | 22.6955 | 22.8799 | 22.2705 | 0.1058 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2063 | 0.0073 | 0.003679 | 0.6655 | 0.4892 | 0.6736 | 0.0642
2012-10-30 21:55 | 26.6218 | 22.9319 | 22.7036 | 22.948 | 22.3026 | 0.1017 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.2184 | 0.007 | 0.003688 | 0.7135 | 0.4852 | 0.7296 | 0.0842
2012-10-30 21:55 | 26.6226 | 22.8606 | 22.6843 | 22.8967 | 22.3034 | 0.1052 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.183 | 0.0073 | 0.003686 | 0.6776 | 0.5013 | 0.7137 | 0.1204
2012-10-30 21:55 | 26.6174 | 22.9155 | 22.6791 | 22.9355 | 22.3263 | 0.1156 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.234 | 0.008 | 0.003686 | 0.6815 | 0.4451 | 0.7015 | 0.0923
2012-10-30 21:55 | 26.6174 | 22.9075 | 22.6871 | 22.8835 | 22.3142 | 0.0948 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.222 | 0.0065 | 0.003686 | 0.6855 | 0.4651 | 0.6615 | 0.0922
2012-10-30 21:55 | 26.6114 | 22.8695 | 22.6851 | 22.8574 | 22.3243 | 0.1179 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 22.212 | 0.0081 | 0.003685 | 0.6575 | 0.4731 | 0.6454 | 0.1123
2012-10-30 21:55 | 26.6174 | 22.8955 | 22.6992 | 22.9075 | 22.3183 | 0.1034 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 22.238 | 0.0071 | 0.003684 | 0.6575 | 0.4612 | 0.6695 | 0.0803
2012-10-30 21:56 | 26.6129 | 22.903 | 22.6866 | 22.919 | 22.2977 | 0.115 | 0.006713 | 0.32644 | 0.002191 | 48.6283 | 0.006921 | 0.2148 | 0.001487 | 31.0361 | 22.1934 | 0.0079 | 0.003678 | 0.7096 | 0.4932 | 0.7256 | 0.1043
2012-10-30 21:56 | 26.6136 | 22.8837 | 22.6272 | 22.9037 | 22.3345 | 0.0977 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.2061 | 0.0067 | 0.00369 | 0.6776 | 0.4211 | 0.6976 | 0.1284
2012-10-30 21:56 | 26.6144 | 22.9125 | 22.6641 | 22.9165 | 22.2952 | 0.0994 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006946 | 0.21479 | 0.001492 | 30.9223 | 22.1949 | 0.0069 | 0.00369 | 0.7176 | 0.4692 | 0.7216 | 0.1003
2012-10-30 21:56 | 26.6226 | 22.8927 | 22.6563 | 22.9127 | 22.3395 | 0.1046 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 22.2272 | 0.0072 | 0.003682 | 0.6655 | 0.4291 | 0.6855 | 0.1123
2012-10-30 21:56 | 26.6203 | 22.8904 | 22.686 | 22.8503 | 22.3412 | 0.1156 | 0.006732 | 0.32653 | 0.002198 | 48.504 | 0.006921 | 0.2148 | 0.001487 | 31.0361 | 22.2329 | 0.008 | 0.003685 | 0.6575 | 0.4531 | 0.6174 | 0.1083
2012-10-30 21:56 | 26.6203 | 22.8784 | 22.6379 | 22.9384 | 22.281 | 0.1104 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.1888 | 0.0076 | 0.003687 | 0.6896 | 0.4491 | 0.7496 | 0.0922
2012-10-30 21:56 | 26.6106 | 22.8967 | 22.6924 | 22.8927 | 22.2834 | 0.1087 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.1991 | 0.0075 | 0.003687 | 0.6976 | 0.4933 | 0.6936 | 0.0843
2012-10-30 21:56 | 26.6203 | 22.8623 | 22.674 | 22.9144 | 22.2891 | 0.111 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006952 | 0.21484 | 0.001494 | 30.9039 | 22.2369 | 0.0077 | 0.003688 | 0.6254 | 0.4371 | 0.6775 | 0.0522
2012-10-30 21:56 | 26.6203 | 22.8944 | 22.686 | 22.9224 | 22.3212 | 0.1052 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.006921 | 0.21486 | 0.001487 | 31.0445 | 22.2329 | 0.0073 | 0.003679 | 0.6615 | 0.4531 | 0.6895 | 0.0883
2012-10-30 21:56 | 26.6174 | 22.8875 | 22.6671 | 22.8514 | 22.2902 | 0.1058 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.226 | 0.0073 | 0.003689 | 0.6615 | 0.4411 | 0.6254 | 0.0642
2012-10-30 21:56 | 26.6196 | 22.8696 | 22.7014 | 22.8897 | 22.3205 | 0.1052 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.2442 | 0.0073 | 0.003687 | 0.6254 | 0.4572 | 0.6455 | 0.0763
2012-10-30 21:56 | 26.6106 | 22.8967 | 22.6643 | 22.8767 | 22.3355 | 0.1006 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.00693 | 0.21471 | 0.001488 | 30.9833 | 22.2513 | 0.0069 | 0.003683 | 0.6454 | 0.413 | 0.6254 | 0.0842
2012-10-30 21:56 | 26.6099 | 22.904 | 22.6476 | 22.872 | 22.3268 | 0.0965 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.2867 | 0.0067 | 0.003686 | 0.6173 | 0.3609 | 0.5853 | 0.0401
2012-10-30 21:56 | 26.6076 | 22.8576 | 22.6934 | 22.9418 | 22.2964 | 0.1139 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.2443 | 0.0079 | 0.003687 | 0.6133 | 0.4491 | 0.6975 | 0.0521
2012-10-30 21:56 | 26.6099 | 22.864 | 22.6756 | 22.88 | 22.2867 | 0.104 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 22.2666 | 0.0072 | 0.003687 | 0.5974 | 0.409 | 0.6134 | 0.0201
2012-10-30 21:56 | 26.6166 | 22.8867 | 22.6703 | 22.9347 | 22.2894 | 0.1023 | 0.006716 | 0.32638 | 0.002192 | 48.598 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.2292 | 0.0071 | 0.003682 | 0.6575 | 0.4411 | 0.7055 | 0.0602
2012-10-30 21:56 | 26.6136 | 22.9117 | 22.6713 | 22.8797 | 22.3305 | 0.1029 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.2503 | 0.0071 | 0.003685 | 0.6614 | 0.421 | 0.6294 | 0.0802
2012-10-30 21:56 | 26.6121 | 22.8621 | 22.6618 | 22.9142 | 22.3089 | 0.0907 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006911 | 0.21479 | 0.001484 | 31.0789 | 22.2929 | 0.0063 | 0.003681 | 0.5692 | 0.3689 | 0.6213 | 0.016
2012-10-30 21:56 | 26.6091 | 22.8912 | 22.6909 | 22.8872 | 22.2979 | 0.1127 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.2658 | 0.0078 | 0.003687 | 0.6254 | 0.4251 | 0.6214 | 0.0321
2012-10-30 21:56 | 26.6039 | 22.89 | 22.6696 | 22.886 | 22.3208 | 0.1006 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.2847 | 0.0069 | 0.003688 | 0.6053 | 0.3849 | 0.6013 | 0.0361
2012-10-30 21:56 | 26.6084 | 22.8785 | 22.7022 | 22.8745 | 22.2772 | 0.1046 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006908 | 0.21476 | 0.001484 | 31.0883 | 22.2691 | 0.0072 | 0.003679 | 0.6094 | 0.4331 | 0.6054 | 0.0081
2012-10-30 21:56 | 26.6031 | 22.8692 | 22.6728 | 22.8772 | 22.328 | 0.0977 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.006921 | 0.21476 | 0.001486 | 31.0299 | 22.2959 | 0.0067 | 0.003683 | 0.5733 | 0.3769 | 0.5813 | 0.0321
2012-10-30 21:56 | 26.6069 | 22.877 | 22.6446 | 22.885 | 22.2997 | 0.1069 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.006914 | 0.2148 | 0.001485 | 31.0675 | 22.2636 | 0.0074 | 0.003678 | 0.6134 | 0.381 | 0.6214 | 0.0361
2012-10-30 21:56 | 26.6039 | 22.87 | 22.5935 | 22.8499 | 22.3048 | 0.1208 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 22.2526 | 0.0083 | 0.003685 | 0.6174 | 0.3409 | 0.5973 | 0.0522
2012-10-30 21:56 | 26.6016 | 22.8356 | 22.6032 | 22.8757 | 22.2864 | 0.0959 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.006952 | 0.21479 | 0.001493 | 30.8956 | 22.2383 | 0.0066 | 0.003689 | 0.5973 | 0.3649 | 0.6374 | 0.0481
2012-10-30 21:56 | 26.6061 | 22.8602 | 22.6277 | 22.8642 | 22.2949 | 0.1087 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.2267 | 0.0075 | 0.003687 | 0.6335 | 0.401 | 0.6375 | 0.0682
2012-10-30 21:57 | 26.6218 | 22.8839 | 22.6194 | 22.8919 | 22.3026 | 0.0965 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.006914 | 0.21479 | 0.001485 | 31.0654 | 22.2344 | 0.0067 | 0.003677 | 0.6495 | 0.385 | 0.6575 | 0.0682
2012-10-30 21:57 | 26.6091 | 22.9152 | 22.6468 | 22.8391 | 22.314 | 0.0959 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006946 | 0.21474 | 0.001492 | 30.9161 | 22.1856 | 0.0066 | 0.003687 | 0.7296 | 0.4612 | 0.6535 | 0.1284
2012-10-30 21:57 | 26.6196 | 22.8977 | 22.6533 | 22.8616 | 22.2884 | 0.0965 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.2242 | 0.0067 | 0.003682 | 0.6735 | 0.4291 | 0.6374 | 0.0642
2012-10-30 21:57 | 26.6196 | 22.8336 | 22.5891 | 22.8136 | 22.3164 | 0.0982 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.006914 | 0.21486 | 0.001486 | 31.0759 | 22.1841 | 0.0068 | 0.00368 | 0.6495 | 0.405 | 0.6295 | 0.1323
2012-10-30 21:57 | 26.6241 | 22.8661 | 22.5896 | 22.8421 | 22.2808 | 0.0942 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.1765 | 0.0065 | 0.00369 | 0.6896 | 0.4131 | 0.6656 | 0.1043
2012-10-30 21:57 | 26.6233 | 22.8533 | 22.5928 | 22.8493 | 22.284 | 0.0977 | 0.00671 | 0.32637 | 0.00219 | 48.6393 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 22.2078 | 0.0067 | 0.003678 | 0.6455 | 0.385 | 0.6415 | 0.0762
2012-10-30 21:57 | 26.6256 | 22.8797 | 22.5871 | 22.8556 | 22.2703 | 0.1029 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.2061 | 0.0071 | 0.003683 | 0.6736 | 0.381 | 0.6495 | 0.0642
2012-10-30 21:57 | 26.6241 | 22.8501 | 22.5856 | 22.8741 | 22.2929 | 0.1092 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.2086 | 0.0075 | 0.003685 | 0.6415 | 0.377 | 0.6655 | 0.0843
2012-10-30 21:57 | 26.6353 | 22.8333 | 22.5808 | 22.8413 | 22.3321 | 0.1046 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 22.2038 | 0.0072 | 0.003684 | 0.6295 | 0.377 | 0.6375 | 0.1283
2012-10-30 21:57 | 26.6383 | 22.8363 | 22.6319 | 22.8363 | 22.3111 | 0.1058 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.2028 | 0.0073 | 0.003688 | 0.6335 | 0.4291 | 0.6335 | 0.1083
2012-10-30 21:57 | 26.6383 | 22.8443 | 22.5838 | 22.9004 | 22.299 | 0.1006 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.1827 | 0.0069 | 0.003691 | 0.6616 | 0.4011 | 0.7177 | 0.1163
2012-10-30 21:57 | 26.6368 | 22.8268 | 22.5622 | 22.8628 | 22.2895 | 0.0988 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.1772 | 0.0068 | 0.003688 | 0.6496 | 0.385 | 0.6856 | 0.1123
2012-10-30 21:57 | 26.6465 | 22.8605 | 22.612 | 22.8565 | 22.2872 | 0.0902 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006927 | 0.21489 | 0.001489 | 31.0218 | 22.2591 | 0.0062 | 0.003684 | 0.6014 | 0.3529 | 0.5974 | 0.0281
2012-10-30 21:57 | 26.6443 | 22.8262 | 22.6459 | 22.8863 | 22.277 | 0.0942 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.277 | 0.0065 | 0.003687 | 0.5492 | 0.3689 | 0.6093 | 0
2012-10-30 21:57 | 26.6495 | 22.8395 | 22.619 | 22.8635 | 22.2661 | 0.1011 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.218 | 0.007 | 0.003684 | 0.6215 | 0.401 | 0.6455 | 0.0481
2012-10-30 21:57 | 26.6458 | 22.8478 | 22.6274 | 22.8598 | 22.2945 | 0.1023 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 22.1782 | 0.0071 | 0.003681 | 0.6696 | 0.4492 | 0.6816 | 0.1163
2012-10-30 21:57 | 26.654 | 22.8119 | 22.6195 | 22.8319 | 22.3188 | 0.1034 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.2546 | 0.0071 | 0.003686 | 0.5573 | 0.3649 | 0.5773 | 0.0642
2012-10-30 21:57 | 26.6615 | 22.8154 | 22.603 | 22.8114 | 22.2861 | 0.104 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.254 | 0.0072 | 0.003683 | 0.5614 | 0.349 | 0.5574 | 0.0321
2012-10-30 21:57 | 26.6638 | 22.8497 | 22.5772 | 22.8698 | 22.2644 | 0.1029 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.2082 | 0.0071 | 0.003687 | 0.6415 | 0.369 | 0.6616 | 0.0562
2012-10-30 21:57 | 26.6668 | 22.8407 | 22.6243 | 22.8407 | 22.2995 | 0.0959 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.006924 | 0.21474 | 0.001487 | 31.0143 | 22.2353 | 0.0066 | 0.003682 | 0.6054 | 0.389 | 0.6054 | 0.0642
2012-10-30 21:57 | 26.6727 | 22.8506 | 22.6262 | 22.8426 | 22.2853 | 0.1179 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 22.2973 | 0.0081 | 0.003686 | 0.5533 | 0.3289 | 0.5453 | -0.012
2012-10-30 21:57 | 26.675 | 22.8169 | 22.6485 | 22.8209 | 22.2916 | 0.0925 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.2716 | 0.0064 | 0.00368 | 0.5453 | 0.3769 | 0.5493 | 0.02
2012-10-30 21:57 | 26.6757 | 22.8416 | 22.6212 | 22.8616 | 22.3405 | 0.0988 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006936 | 0.21486 | 0.00149 | 30.9774 | 22.2803 | 0.0068 | 0.003688 | 0.5613 | 0.3409 | 0.5813 | 0.0602
2012-10-30 21:57 | 26.678 | 22.8359 | 22.6315 | 22.8439 | 22.3026 | 0.1173 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006914 | 0.21477 | 0.001485 | 31.0633 | 22.2906 | 0.0081 | 0.00368 | 0.5453 | 0.3409 | 0.5533 | 0.012
2012-10-30 21:57 | 26.6802 | 22.798 | 22.6257 | 22.8101 | 22.2366 | 0.0896 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.3048 | 0.0062 | 0.003689 | 0.4932 | 0.3209 | 0.5053 | -0.0682
2012-10-30 21:57 | 26.678 | 22.8159 | 22.6515 | 22.8038 | 22.2705 | 0.0902 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.2425 | 0.0062 | 0.003689 | 0.5734 | 0.409 | 0.5613 | 0.028
2012-10-30 21:58 | 26.6855 | 22.8154 | 22.639 | 22.8354 | 22.3142 | 0.1115 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.2981 | 0.0077 | 0.003685 | 0.5173 | 0.3409 | 0.5373 | 0.0161
2012-10-30 21:58 | 26.693 | 22.8389 | 22.6425 | 22.8269 | 22.2936 | 0.1115 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.2775 | 0.0077 | 0.003683 | 0.5614 | 0.365 | 0.5494 | 0.0161
2012-10-30 21:58 | 26.6945 | 22.8043 | 22.636 | 22.8404 | 22.267 | 0.1063 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.2911 | 0.0073 | 0.003691 | 0.5132 | 0.3449 | 0.5493 | -0.0241
2012-10-30 21:58 | 26.6952 | 22.8251 | 22.7088 | 22.821 | 22.2998 | 0.1058 | 0.006732 | 0.32651 | 0.002198 | 48.5018 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.2877 | 0.0073 | 0.003688 | 0.5374 | 0.4211 | 0.5333 | 0.0121
2012-10-30 21:58 | 26.7004 | 22.8062 | 22.6699 | 22.8423 | 22.2849 | 0.1006 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.309 | 0.0069 | 0.003689 | 0.4972 | 0.3609 | 0.5333 | -0.0241
2012-10-30 21:58 | 26.7004 | 22.7982 | 22.6739 | 22.8343 | 22.297 | 0.1046 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.2649 | 0.0072 | 0.003684 | 0.5333 | 0.409 | 0.5694 | 0.0321
2012-10-30 21:58 | 26.6997 | 22.8055 | 22.6612 | 22.8496 | 22.2963 | 0.093 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.3003 | 0.0064 | 0.003686 | 0.5052 | 0.3609 | 0.5493 | -0.004
2012-10-30 21:58 | 26.6997 | 22.7815 | 22.6091 | 22.8296 | 22.3203 | 0.1058 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.2922 | 0.0073 | 0.003683 | 0.4893 | 0.3169 | 0.5374 | 0.0281
2012-10-30 21:58 | 26.6945 | 22.8364 | 22.652 | 22.8003 | 22.3111 | 0.1069 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 22.2991 | 0.0074 | 0.003688 | 0.5373 | 0.3529 | 0.5012 | 0.012
2012-10-30 21:58 | 26.6974 | 22.8152 | 22.6589 | 22.8112 | 22.298 | 0.1069 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.2779 | 0.0074 | 0.003684 | 0.5373 | 0.381 | 0.5333 | 0.0201
2012-10-30 21:58 | 26.7012 | 22.8391 | 22.6587 | 22.8471 | 22.3098 | 0.1 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3138 | 0.0069 | 0.003685 | 0.5253 | 0.3449 | 0.5333 | -0.004
2012-10-30 21:58 | 26.7087 | 22.8426 | 22.6542 | 22.8386 | 22.3093 | 0.1029 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006939 | 0.21486 | 0.001491 | 30.964 | 22.3213 | 0.0071 | 0.003686 | 0.5213 | 0.3329 | 0.5173 | -0.012
2012-10-30 21:58 | 26.7042 | 22.806 | 22.6537 | 22.8541 | 22.2727 | 0.1058 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3369 | 0.0073 | 0.003683 | 0.4691 | 0.3168 | 0.5172 | -0.0642
2012-10-30 21:58 | 26.7057 | 22.7915 | 22.6592 | 22.8195 | 22.2822 | 0.0959 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.3544 | 0.0066 | 0.003683 | 0.4371 | 0.3048 | 0.4651 | -0.0722
2012-10-30 21:58 | 26.7117 | 22.8255 | 22.6492 | 22.8255 | 22.3123 | 0.0988 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.006914 | 0.21473 | 0.001485 | 31.0571 | 22.3444 | 0.0068 | 0.003683 | 0.4811 | 0.3048 | 0.4811 | -0.0321
2012-10-30 21:58 | 26.7079 | 22.8137 | 22.6093 | 22.8177 | 22.2764 | 0.0988 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 22.3366 | 0.0068 | 0.003684 | 0.4771 | 0.2727 | 0.4811 | -0.0602
2012-10-30 21:58 | 26.7132 | 22.835 | 22.6467 | 22.8511 | 22.3057 | 0.104 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.3459 | 0.0072 | 0.003687 | 0.4891 | 0.3008 | 0.5052 | -0.0402
2012-10-30 21:58 | 26.7042 | 22.826 | 22.6577 | 22.806 | 22.2767 | 0.1017 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.4051 | 0.007 | 0.003682 | 0.4209 | 0.2526 | 0.4009 | -0.1284
2012-10-30 21:58 | 26.7117 | 22.8135 | 22.6812 | 22.8696 | 22.2882 | 0.1029 | 0.0067 | 0.32636 | 0.002187 | 48.7097 | 0.006936 | 0.21473 | 0.001489 | 30.9586 | 22.4085 | 0.0071 | 0.003676 | 0.405 | 0.2727 | 0.4611 | -0.1203
2012-10-30 21:58 | 26.7117 | 22.7935 | 22.6572 | 22.8375 | 22.3002 | 0.0965 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.3885 | 0.0067 | 0.003687 | 0.405 | 0.2687 | 0.449 | -0.0883
2012-10-30 21:58 | 26.7019 | 22.7877 | 22.6754 | 22.8478 | 22.2824 | 0.1006 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.4228 | 0.0069 | 0.003684 | 0.3649 | 0.2526 | 0.425 | -0.1404
2012-10-30 21:58 | 26.7117 | 22.7854 | 22.6692 | 22.8055 | 22.3123 | 0.1017 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.4005 | 0.007 | 0.00369 | 0.3849 | 0.2687 | 0.405 | -0.0882
2012-10-30 21:58 | 26.7027 | 22.8245 | 22.6161 | 22.8446 | 22.2391 | 0.1191 | 0.006757 | 0.32646 | 0.002206 | 48.3137 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3875 | 0.0082 | 0.003693 | 0.437 | 0.2286 | 0.4571 | -0.1484
2012-10-30 21:58 | 26.7147 | 22.8005 | 22.6682 | 22.8566 | 22.3233 | 0.1034 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 22.4115 | 0.0071 | 0.003687 | 0.389 | 0.2567 | 0.4451 | -0.0882
2012-10-30 21:58 | 26.7154 | 22.7932 | 22.6889 | 22.8372 | 22.2999 | 0.1081 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.4122 | 0.0075 | 0.003688 | 0.381 | 0.2767 | 0.425 | -0.1123
2012-10-30 21:58 | 26.7042 | 22.794 | 22.6577 | 22.826 | 22.2847 | 0.1052 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.381 | 0.0073 | 0.003689 | 0.413 | 0.2767 | 0.445 | -0.0963
2012-10-30 21:58 | 26.7154 | 22.8412 | 22.697 | 22.8292 | 22.2999 | 0.1162 | 0.006713 | 0.32651 | 0.002192 | 48.6391 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.34 | 0.008 | 0.003682 | 0.5012 | 0.357 | 0.4892 | -0.0401
2012-10-30 21:58 | 26.7132 | 22.795 | 22.6827 | 22.839 | 22.3057 | 0.0971 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.386 | 0.0067 | 0.003688 | 0.409 | 0.2967 | 0.453 | -0.0803
2012-10-30 21:59 | 26.7124 | 22.7821 | 22.6298 | 22.8262 | 22.321 | 0.0994 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.006949 | 0.21479 | 0.001493 | 30.9089 | 22.3451 | 0.0069 | 0.003689 | 0.437 | 0.2847 | 0.4811 | -0.0241
2012-10-30 21:59 | 26.7169 | 22.7786 | 22.6544 | 22.8347 | 22.2412 | 0.0936 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3977 | 0.0065 | 0.003685 | 0.3809 | 0.2567 | 0.437 | -0.1565
2012-10-30 21:59 | 26.7229 | 22.7926 | 22.6804 | 22.8087 | 22.3114 | 0.0988 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3516 | 0.0068 | 0.003683 | 0.441 | 0.3288 | 0.4571 | -0.0402
2012-10-30 21:59 | 26.7177 | 22.8115 | 22.6832 | 22.8355 | 22.3303 | 0.0948 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006924 | 0.21486 | 0.001488 | 31.0311 | 22.3624 | 0.0065 | 0.003683 | 0.4491 | 0.3208 | 0.4731 | -0.0321
2012-10-30 21:59 | 26.7184 | 22.7721 | 22.6599 | 22.8242 | 22.339 | 0.1023 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.3791 | 0.0071 | 0.003689 | 0.393 | 0.2808 | 0.4451 | -0.0401
2012-10-30 21:59 | 26.7192 | 22.817 | 22.6326 | 22.825 | 22.3117 | 0.0971 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.388 | 0.0067 | 0.003686 | 0.429 | 0.2446 | 0.437 | -0.0763
2012-10-30 21:59 | 26.7221 | 22.7999 | 22.6475 | 22.7999 | 22.3106 | 0.0867 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 22.3828 | 0.006 | 0.003681 | 0.4171 | 0.2647 | 0.4171 | -0.0722
2012-10-30 21:59 | 26.7281 | 22.7658 | 22.6656 | 22.8299 | 22.3166 | 0.1115 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 22.3848 | 0.0077 | 0.003687 | 0.381 | 0.2808 | 0.4451 | -0.0682
2012-10-30 21:59 | 26.7236 | 22.7813 | 22.6129 | 22.7813 | 22.2439 | 0.1092 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006917 | 0.21483 | 0.001486 | 31.0583 | 22.3402 | 0.0075 | 0.00368 | 0.4411 | 0.2727 | 0.4411 | -0.0963
2012-10-30 21:59 | 26.7184 | 22.7801 | 22.6519 | 22.74 | 22.323 | 0.1 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3551 | 0.0069 | 0.003688 | 0.425 | 0.2968 | 0.3849 | -0.0321
2012-10-30 21:59 | 26.7334 | 22.747 | 22.6468 | 22.8352 | 22.3219 | 0.1121 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.3821 | 0.0077 | 0.003684 | 0.3649 | 0.2647 | 0.4531 | -0.0602
2012-10-30 21:59 | 26.7251 | 22.7507 | 22.6425 | 22.7748 | 22.2695 | 0.1081 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.3538 | 0.0075 | 0.003681 | 0.3969 | 0.2887 | 0.421 | -0.0843
2012-10-30 21:59 | 26.7236 | 22.8054 | 22.629 | 22.7653 | 22.28 | 0.1098 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006933 | 0.21489 | 0.00149 | 30.995 | 22.3924 | 0.0076 | 0.003689 | 0.413 | 0.2366 | 0.3729 | -0.1124
2012-10-30 21:59 | 26.7266 | 22.7442 | 22.6119 | 22.7683 | 22.3071 | 0.1081 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006911 | 0.21484 | 0.001485 | 31.0873 | 22.4154 | 0.0075 | 0.003684 | 0.3288 | 0.1965 | 0.3529 | -0.1083
2012-10-30 21:59 | 26.7319 | 22.7215 | 22.7255 | 22.7736 | 22.2883 | 0.1092 | 0.006703 | 0.32643 | 0.002188 | 48.6986 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.4207 | 0.0075 | 0.003679 | 0.3008 | 0.3048 | 0.3529 | -0.1324
2012-10-30 21:59 | 26.7341 | 22.8038 | 22.6876 | 22.8038 | 22.2504 | 0.1058 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 22.4069 | 0.0073 | 0.003688 | 0.3969 | 0.2807 | 0.3969 | -0.1565
2012-10-30 21:59 | 26.7401 | 22.7698 | 22.7056 | 22.8339 | 22.3005 | 0.1104 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.457 | 0.0076 | 0.003688 | 0.3128 | 0.2486 | 0.3769 | -0.1565
2012-10-30 21:59 | 26.7281 | 22.7217 | 22.7257 | 22.7738 | 22.3086 | 0.0878 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006939 | 0.21471 | 0.00149 | 30.9431 | 22.4089 | 0.0061 | 0.003685 | 0.3128 | 0.3168 | 0.3649 | -0.1003
2012-10-30 21:59 | 26.7349 | 22.7846 | 22.6684 | 22.8006 | 22.2994 | 0.1017 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.4077 | 0.007 | 0.003681 | 0.3769 | 0.2607 | 0.3929 | -0.1083
2012-10-30 21:59 | 26.7296 | 22.7552 | 22.651 | 22.7392 | 22.27 | 0.1087 | 0.006741 | 0.32636 | 0.0022 | 48.4134 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.4144 | 0.0075 | 0.003689 | 0.3408 | 0.2366 | 0.3248 | -0.1444
2012-10-30 21:59 | 26.7356 | 22.7292 | 22.661 | 22.7933 | 22.3121 | 0.0971 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.4405 | 0.0067 | 0.003686 | 0.2887 | 0.2205 | 0.3528 | -0.1284
2012-10-30 21:59 | 26.7461 | 22.7878 | 22.6515 | 22.8198 | 22.3386 | 0.1017 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006927 | 0.21477 | 0.001488 | 31.0051 | 22.467 | 0.007 | 0.003687 | 0.3208 | 0.1845 | 0.3528 | -0.1284
2012-10-30 21:59 | 26.7289 | 22.7746 | 22.6263 | 22.7946 | 22.3295 | 0.1 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 22.4257 | 0.0069 | 0.003679 | 0.3489 | 0.2006 | 0.3689 | -0.0962
2012-10-30 21:59 | 26.7259 | 22.7556 | 22.6553 | 22.8357 | 22.2944 | 0.0977 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.4027 | 0.0067 | 0.003686 | 0.3529 | 0.2526 | 0.433 | -0.1083
2012-10-30 21:59 | 26.7244 | 22.75 | 22.6458 | 22.7821 | 22.2889 | 0.1104 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006914 | 0.21479 | 0.001485 | 31.0654 | 22.4373 | 0.0076 | 0.003679 | 0.3127 | 0.2085 | 0.3448 | -0.1484
2012-10-30 21:59 | 26.7184 | 22.724 | 22.6519 | 22.8242 | 22.3069 | 0.1069 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.4393 | 0.0074 | 0.003681 | 0.2847 | 0.2126 | 0.3849 | -0.1324
2012-10-30 21:59 | 26.7162 | 22.7739 | 22.6857 | 22.7739 | 22.2807 | 0.0982 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 22.4291 | 0.0068 | 0.003683 | 0.3448 | 0.2566 | 0.3448 | -0.1484
2012-10-30 21:59 | 26.7192 | 22.7729 | 22.6527 | 22.841 | 22.3358 | 0.1156 | 0.00671 | 0.32643 | 0.00219 | 48.6478 | 0.006917 | 0.21482 | 0.001486 | 31.0561 | 22.4602 | 0.008 | 0.003676 | 0.3127 | 0.1925 | 0.3808 | -0.1244
2012-10-30 21:59 | 26.7214 | 22.7791 | 22.6629 | 22.8072 | 22.318 | 0.1271 | 0.00671 | 0.32641 | 0.00219 | 48.6458 | 0.006924 | 0.21484 | 0.001488 | 31.0289 | 22.4102 | 0.0088 | 0.003678 | 0.3689 | 0.2527 | 0.397 | -0.0922
2012-10-30 22:00 | 26.7311 | 22.7808 | 22.6806 | 22.8008 | 22.2835 | 0.1052 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.448 | 0.0073 | 0.003681 | 0.3328 | 0.2326 | 0.3528 | -0.1645
2012-10-30 22:00 | 26.7169 | 22.7425 | 22.6584 | 22.8067 | 22.3094 | 0.1156 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.4218 | 0.008 | 0.003688 | 0.3207 | 0.2366 | 0.3849 | -0.1124
2012-10-30 22:00 | 26.7251 | 22.7708 | 22.6425 | 22.8069 | 22.2936 | 0.1029 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.006936 | 0.21486 | 0.00149 | 30.9774 | 22.4621 | 0.0071 | 0.003684 | 0.3087 | 0.1804 | 0.3448 | -0.1685
2012-10-30 22:00 | 26.7206 | 22.7382 | 22.626 | 22.7903 | 22.3372 | 0.1058 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.4455 | 0.0073 | 0.003685 | 0.2927 | 0.1805 | 0.3448 | -0.1083
2012-10-30 22:00 | 26.7169 | 22.7666 | 22.6463 | 22.8027 | 22.2773 | 0.1023 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.00693 | 0.21487 | 0.001489 | 31.0062 | 22.4378 | 0.0071 | 0.003683 | 0.3288 | 0.2085 | 0.3649 | -0.1605
2012-10-30 22:00 | 26.7192 | 22.7248 | 22.6887 | 22.805 | 22.3559 | 0.1069 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.408 | 0.0074 | 0.003686 | 0.3168 | 0.2807 | 0.397 | -0.0521
2012-10-30 22:00 | 26.7259 | 22.7756 | 22.6553 | 22.7956 | 22.3586 | 0.1144 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.006917 | 0.21482 | 0.001486 | 31.0561 | 22.4508 | 0.0079 | 0.003678 | 0.3248 | 0.2045 | 0.3448 | -0.0922
2012-10-30 22:00 | 26.7221 | 22.7838 | 22.6074 | 22.8079 | 22.3347 | 0.111 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.4871 | 0.0077 | 0.003685 | 0.2967 | 0.1203 | 0.3208 | -0.1524
2012-10-30 22:00 | 26.7214 | 22.743 | 22.6027 | 22.7671 | 22.2899 | 0.1162 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.4904 | 0.008 | 0.003683 | 0.2526 | 0.1123 | 0.2767 | -0.2005
2012-10-30 22:00 | 26.7177 | 22.7714 | 22.6271 | 22.7794 | 22.3022 | 0.0994 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.4426 | 0.0069 | 0.003681 | 0.3288 | 0.1845 | 0.3368 | -0.1404
2012-10-30 22:00 | 26.7072 | 22.7489 | 22.6326 | 22.785 | 22.2957 | 0.1225 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.006921 | 0.21483 | 0.001487 | 31.0403 | 22.4642 | 0.0084 | 0.003685 | 0.2847 | 0.1684 | 0.3208 | -0.1685
2012-10-30 22:00 | 26.7124 | 22.7741 | 22.6459 | 22.7821 | 22.313 | 0.1104 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.4534 | 0.0076 | 0.003682 | 0.3207 | 0.1925 | 0.3287 | -0.1404
2012-10-30 22:00 | 26.7087 | 22.7504 | 22.594 | 22.7504 | 22.2852 | 0.1046 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.006908 | 0.21484 | 0.001484 | 31.1008 | 22.4256 | 0.0072 | 0.00368 | 0.3248 | 0.1684 | 0.3248 | -0.1404
2012-10-30 22:00 | 26.7027 | 22.7604 | 22.6241 | 22.7764 | 22.2872 | 0.1081 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.006946 | 0.2148 | 0.001492 | 30.9244 | 22.4036 | 0.0075 | 0.003684 | 0.3568 | 0.2205 | 0.3728 | -0.1164
2012-10-30 22:00 | 26.7004 | 22.706 | 22.6339 | 22.7341 | 22.301 | 0.1087 | 0.006713 | 0.32644 | 0.002191 | 48.6283 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 22.4213 | 0.0075 | 0.003679 | 0.2847 | 0.2126 | 0.3128 | -0.1203
2012-10-30 22:00 | 26.7072 | 22.7729 | 22.6126 | 22.7649 | 22.2877 | 0.115 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006946 | 0.21487 | 0.001493 | 30.9348 | 22.4361 | 0.0079 | 0.003686 | 0.3368 | 0.1765 | 0.3288 | -0.1484
2012-10-30 22:00 | 26.7034 | 22.7611 | 22.5807 | 22.7411 | 22.316 | 0.1139 | 0.006719 | 0.32628 | 0.002192 | 48.5612 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3922 | 0.0079 | 0.003679 | 0.3689 | 0.1885 | 0.3489 | -0.0762
2012-10-30 22:00 | 26.7012 | 22.7709 | 22.6427 | 22.7589 | 22.2777 | 0.1011 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.4181 | 0.007 | 0.003679 | 0.3528 | 0.2246 | 0.3408 | -0.1404
2012-10-30 22:00 | 26.7012 | 22.791 | 22.5865 | 22.791 | 22.2295 | 0.0954 | 0.006716 | 0.32638 | 0.002192 | 48.598 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3619 | 0.0066 | 0.003679 | 0.4291 | 0.2246 | 0.4291 | -0.1324
2012-10-30 22:00 | 26.7049 | 22.7786 | 22.6263 | 22.7506 | 22.2974 | 0.1011 | 0.006713 | 0.3264 | 0.002191 | 48.6219 | 0.006917 | 0.21483 | 0.001486 | 31.0583 | 22.3777 | 0.007 | 0.003677 | 0.4009 | 0.2486 | 0.3729 | -0.0803
2012-10-30 22:00 | 26.6974 | 22.7471 | 22.6028 | 22.7671 | 22.2859 | 0.0988 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.006921 | 0.21476 | 0.001486 | 31.0299 | 22.3541 | 0.0068 | 0.003679 | 0.393 | 0.2487 | 0.413 | -0.0682
2012-10-30 22:00 | 26.6922 | 22.7659 | 22.5775 | 22.7138 | 22.2888 | 0.104 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3249 | 0.0072 | 0.003683 | 0.441 | 0.2526 | 0.3889 | -0.0361
2012-10-30 22:00 | 26.7042 | 22.7699 | 22.5535 | 22.7579 | 22.3008 | 0.1 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 22.3168 | 0.0069 | 0.003679 | 0.4531 | 0.2367 | 0.4411 | -0.016
2012-10-30 22:00 | 26.7012 | 22.7509 | 22.5745 | 22.783 | 22.3098 | 0.0948 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.378 | 0.0065 | 0.003683 | 0.3729 | 0.1965 | 0.405 | -0.0682
2012-10-30 22:00 | 26.7064 | 22.7962 | 22.5717 | 22.724 | 22.311 | 0.1121 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.4193 | 0.0077 | 0.003684 | 0.3769 | 0.1524 | 0.3047 | -0.1083
2012-10-30 22:00 | 26.7177 | 22.7834 | 22.583 | 22.7754 | 22.3062 | 0.1081 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.4627 | 0.0075 | 0.003682 | 0.3207 | 0.1203 | 0.3127 | -0.1565
2012-10-30 22:00 | 26.7079 | 22.7416 | 22.6053 | 22.7736 | 22.2683 | 0.093 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.4408 | 0.0064 | 0.003687 | 0.3008 | 0.1645 | 0.3328 | -0.1725
2012-10-30 22:00 | 26.7162 | 22.7739 | 22.5975 | 22.7619 | 22.3087 | 0.1006 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.5013 | 0.0069 | 0.003688 | 0.2726 | 0.0962 | 0.2606 | -0.1926
2012-10-30 22:01 | 26.7147 | 22.8125 | 22.592 | 22.7444 | 22.2751 | 0.104 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.4396 | 0.0072 | 0.003691 | 0.3729 | 0.1524 | 0.3048 | -0.1645
2012-10-30 22:01 | 26.7072 | 22.7729 | 22.5885 | 22.7609 | 22.2877 | 0.1144 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.4522 | 0.0079 | 0.003679 | 0.3207 | 0.1363 | 0.3087 | -0.1645
2012-10-30 22:01 | 26.7094 | 22.719 | 22.5587 | 22.7551 | 22.2538 | 0.1006 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 22.3782 | 0.0069 | 0.003685 | 0.3408 | 0.1805 | 0.3769 | -0.1244
2012-10-30 22:01 | 26.7034 | 22.7451 | 22.5927 | 22.721 | 22.2638 | 0.1133 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.4163 | 0.0078 | 0.003684 | 0.3288 | 0.1764 | 0.3047 | -0.1525
2012-10-30 22:01 | 26.7034 | 22.7691 | 22.5927 | 22.721 | 22.2839 | 0.1063 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006914 | 0.21484 | 0.001485 | 31.0738 | 22.4123 | 0.0073 | 0.00368 | 0.3568 | 0.1804 | 0.3087 | -0.1284
2012-10-30 22:01 | 26.7132 | 22.7789 | 22.5985 | 22.7108 | 22.2937 | 0.0942 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3619 | 0.0065 | 0.003684 | 0.417 | 0.2366 | 0.3489 | -0.0682
2012-10-30 22:01 | 26.7094 | 22.7751 | 22.5907 | 22.735 | 22.2578 | 0.1017 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.4103 | 0.007 | 0.003683 | 0.3648 | 0.1804 | 0.3247 | -0.1525
2012-10-30 22:01 | 26.6989 | 22.7646 | 22.5281 | 22.7446 | 22.2634 | 0.1029 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.4359 | 0.0071 | 0.003687 | 0.3287 | 0.0922 | 0.3087 | -0.1725
2012-10-30 22:01 | 26.7132 | 22.791 | 22.5825 | 22.7228 | 22.3057 | 0.1092 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.4822 | 0.0075 | 0.003683 | 0.3088 | 0.1003 | 0.2406 | -0.1765
2012-10-30 22:01 | 26.7132 | 22.7549 | 22.5665 | 22.7148 | 22.2897 | 0.1011 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21487 | 0.00149 | 30.9928 | 22.4461 | 0.007 | 0.003685 | 0.3088 | 0.1204 | 0.2687 | -0.1564
2012-10-30 22:01 | 26.7184 | 22.7801 | 22.6198 | 22.7521 | 22.2949 | 0.0988 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.4193 | 0.0068 | 0.003689 | 0.3608 | 0.2005 | 0.3328 | -0.1244
2012-10-30 22:01 | 26.7229 | 22.7125 | 22.5641 | 22.7526 | 22.2673 | 0.1219 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.4318 | 0.0084 | 0.003686 | 0.2807 | 0.1323 | 0.3208 | -0.1645
2012-10-30 22:01 | 26.7177 | 22.7955 | 22.5429 | 22.7353 | 22.246 | 0.1115 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.4426 | 0.0077 | 0.003681 | 0.3529 | 0.1003 | 0.2927 | -0.1966
2012-10-30 22:01 | 26.7259 | 22.7836 | 22.5711 | 22.6954 | 22.2984 | 0.0942 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.4869 | 0.0065 | 0.003681 | 0.2967 | 0.0842 | 0.2085 | -0.1885
2012-10-30 22:01 | 26.7236 | 22.7813 | 22.6129 | 22.7452 | 22.28 | 0.1052 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006921 | 0.21483 | 0.001487 | 31.0403 | 22.5007 | 0.0073 | 0.003682 | 0.2806 | 0.1122 | 0.2445 | -0.2207
2012-10-30 22:01 | 26.7251 | 22.7628 | 22.5744 | 22.7307 | 22.2695 | 0.115 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006943 | 0.21479 | 0.001491 | 30.9356 | 22.434 | 0.0079 | 0.003688 | 0.3288 | 0.1404 | 0.2967 | -0.1645
2012-10-30 22:01 | 26.7274 | 22.745 | 22.5125 | 22.7009 | 22.2919 | 0.1029 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.3922 | 0.0071 | 0.003686 | 0.3528 | 0.1203 | 0.3087 | -0.1003
2012-10-30 22:01 | 26.7304 | 22.752 | 22.5075 | 22.7039 | 22.2708 | 0.0936 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3751 | 0.0065 | 0.003683 | 0.3769 | 0.1324 | 0.3288 | -0.1043
2012-10-30 22:01 | 26.7319 | 22.7575 | 22.5411 | 22.7335 | 22.2602 | 0.1127 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.4287 | 0.0078 | 0.003684 | 0.3288 | 0.1124 | 0.3048 | -0.1685
2012-10-30 22:01 | 26.7424 | 22.7801 | 22.5235 | 22.732 | 22.2587 | 0.1034 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.4553 | 0.0071 | 0.003689 | 0.3248 | 0.0682 | 0.2767 | -0.1966
2012-10-30 22:01 | 26.7476 | 22.7853 | 22.5287 | 22.7051 | 22.3 | 0.1121 | 0.006713 | 0.3265 | 0.002192 | 48.637 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.4284 | 0.0077 | 0.003682 | 0.3569 | 0.1003 | 0.2767 | -0.1284
2012-10-30 22:01 | 26.7446 | 22.7181 | 22.5337 | 22.6901 | 22.285 | 0.1052 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.4374 | 0.0073 | 0.003688 | 0.2807 | 0.0963 | 0.2527 | -0.1524
2012-10-30 22:01 | 26.7431 | 22.7287 | 22.5482 | 22.7327 | 22.2755 | 0.1098 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.46 | 0.0076 | 0.003684 | 0.2687 | 0.0882 | 0.2727 | -0.1845
2012-10-30 22:01 | 26.7409 | 22.7665 | 22.5902 | 22.6904 | 22.2572 | 0.1104 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.4578 | 0.0076 | 0.003688 | 0.3087 | 0.1324 | 0.2326 | -0.2006
2012-10-30 22:01 | 26.7491 | 22.7507 | 22.5141 | 22.7387 | 22.2614 | 0.115 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.4941 | 0.0079 | 0.003682 | 0.2566 | 0.02 | 0.2446 | -0.2327
2012-10-30 22:01 | 26.7409 | 22.7706 | 22.5621 | 22.6944 | 22.2733 | 0.1139 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.4257 | 0.0079 | 0.003684 | 0.3449 | 0.1364 | 0.2687 | -0.1524
2012-10-30 22:01 | 26.7483 | 22.7339 | 22.5615 | 22.7258 | 22.2967 | 0.1104 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.4772 | 0.0076 | 0.003686 | 0.2567 | 0.0843 | 0.2486 | -0.1805
2012-10-30 22:01 | 26.7521 | 22.7617 | 22.5613 | 22.7136 | 22.2524 | 0.1087 | 0.006744 | 0.32641 | 0.002201 | 48.4005 | 0.006936 | 0.21487 | 0.00149 | 30.9794 | 22.4329 | 0.0075 | 0.003691 | 0.3288 | 0.1284 | 0.2807 | -0.1805
2012-10-30 22:02 | 26.7439 | 22.7054 | 22.5771 | 22.7214 | 22.2602 | 0.1115 | 0.006744 | 0.32646 | 0.002202 | 48.4069 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.4488 | 0.0077 | 0.00369 | 0.2566 | 0.1283 | 0.2726 | -0.1886
2012-10-30 22:02 | 26.7596 | 22.7532 | 22.5808 | 22.689 | 22.2639 | 0.1081 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.4404 | 0.0075 | 0.003682 | 0.3128 | 0.1404 | 0.2486 | -0.1765
2012-10-30 22:02 | 26.7536 | 22.7752 | 22.5628 | 22.679 | 22.2659 | 0.1046 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.5186 | 0.0072 | 0.003687 | 0.2566 | 0.0442 | 0.1604 | -0.2527
2012-10-30 22:02 | 26.7611 | 22.7306 | 22.5582 | 22.7306 | 22.2734 | 0.1115 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.5502 | 0.0077 | 0.003688 | 0.1804 | 0.008 | 0.1804 | -0.2768
2012-10-30 22:02 | 26.7596 | 22.7171 | 22.5607 | 22.661 | 22.2639 | 0.1167 | 0.00671 | 0.32649 | 0.002191 | 48.6565 | 0.006924 | 0.2149 | 0.001488 | 31.0373 | 22.5327 | 0.008 | 0.003679 | 0.1844 | 0.028 | 0.1283 | -0.2688
2012-10-30 22:02 | 26.7671 | 22.6885 | 22.5642 | 22.6725 | 22.2995 | 0.1196 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.5803 | 0.0082 | 0.003684 | 0.1082 | -0.0161 | 0.0922 | -0.2808
2012-10-30 22:02 | 26.7603 | 22.7258 | 22.5293 | 22.6697 | 22.2726 | 0.1063 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 22.5775 | 0.0073 | 0.003684 | 0.1483 | -0.0482 | 0.0922 | -0.3049
2012-10-30 22:02 | 26.7715 | 22.7611 | 22.5446 | 22.6889 | 22.2838 | 0.1052 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.006914 | 0.21473 | 0.001485 | 31.0571 | 22.6087 | 0.0073 | 0.00368 | 0.1524 | -0.0641 | 0.0802 | -0.3249
2012-10-30 22:02 | 26.7745 | 22.744 | 22.5756 | 22.736 | 22.2868 | 0.0919 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.5796 | 0.0063 | 0.003682 | 0.1644 | -0.004 | 0.1564 | -0.2928
2012-10-30 22:02 | 26.7641 | 22.7256 | 22.5372 | 22.6895 | 22.2443 | 0.1046 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.5773 | 0.0072 | 0.003684 | 0.1483 | -0.0401 | 0.1122 | -0.333
2012-10-30 22:02 | 26.776 | 22.7135 | 22.5491 | 22.7054 | 22.2843 | 0.1156 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.006921 | 0.21477 | 0.001486 | 31.0319 | 22.5049 | 0.008 | 0.003683 | 0.2086 | 0.0442 | 0.2005 | -0.2206
2012-10-30 22:02 | 26.7753 | 22.7408 | 22.5604 | 22.7007 | 22.2796 | 0.1075 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 22.4842 | 0.0074 | 0.003686 | 0.2566 | 0.0762 | 0.2165 | -0.2046
2012-10-30 22:02 | 26.7738 | 22.7433 | 22.5469 | 22.6872 | 22.25 | 0.1092 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 22.4626 | 0.0075 | 0.003682 | 0.2807 | 0.0843 | 0.2246 | -0.2126
2012-10-30 22:02 | 26.7768 | 22.7664 | 22.5218 | 22.6862 | 22.261 | 0.1058 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.4977 | 0.0073 | 0.003687 | 0.2687 | 0.0241 | 0.1885 | -0.2367
2012-10-30 22:02 | 26.7843 | 22.7618 | 22.5413 | 22.7017 | 22.2765 | 0.1115 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.4972 | 0.0077 | 0.003682 | 0.2646 | 0.0441 | 0.2045 | -0.2207
2012-10-30 22:02 | 26.7775 | 22.711 | 22.5345 | 22.6628 | 22.2818 | 0.1121 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.4543 | 0.0077 | 0.003686 | 0.2567 | 0.0802 | 0.2085 | -0.1725
2012-10-30 22:02 | 26.7775 | 22.6989 | 22.5666 | 22.6668 | 22.2617 | 0.1185 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.4663 | 0.0082 | 0.003683 | 0.2326 | 0.1003 | 0.2005 | -0.2046
2012-10-30 22:02 | 26.785 | 22.7425 | 22.5621 | 22.6984 | 22.2732 | 0.13 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.4578 | 0.009 | 0.003683 | 0.2847 | 0.1043 | 0.2406 | -0.1846
2012-10-30 22:02 | 26.7813 | 22.7268 | 22.5664 | 22.7308 | 22.2655 | 0.1185 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006917 | 0.21479 | 0.001486 | 31.0519 | 22.434 | 0.0082 | 0.003685 | 0.2928 | 0.1324 | 0.2968 | -0.1685
2012-10-30 22:02 | 26.7903 | 22.7638 | 22.5353 | 22.6796 | 22.2464 | 0.1133 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.431 | 0.0078 | 0.00368 | 0.3328 | 0.1043 | 0.2486 | -0.1846
2012-10-30 22:02 | 26.7888 | 22.7543 | 22.5939 | 22.6581 | 22.269 | 0.1243 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.4014 | 0.0086 | 0.00369 | 0.3529 | 0.1925 | 0.2567 | -0.1324
2012-10-30 22:02 | 26.7805 | 22.758 | 22.5495 | 22.7059 | 22.2567 | 0.1214 | 0.006741 | 0.32649 | 0.002201 | 48.4327 | 0.006917 | 0.21484 | 0.001486 | 31.0603 | 22.3771 | 0.0084 | 0.003687 | 0.3809 | 0.1724 | 0.3288 | -0.1204
2012-10-30 22:02 | 26.779 | 22.7485 | 22.5761 | 22.6844 | 22.2351 | 0.1121 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.4558 | 0.0077 | 0.003689 | 0.2927 | 0.1203 | 0.2286 | -0.2207
2012-10-30 22:02 | 26.776 | 22.7175 | 22.5491 | 22.6613 | 22.2843 | 0.1052 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006917 | 0.21493 | 0.001487 | 31.0729 | 22.3846 | 0.0073 | 0.003685 | 0.3329 | 0.1645 | 0.2767 | -0.1003
2012-10-30 22:02 | 26.7843 | 22.7578 | 22.5493 | 22.6857 | 22.2725 | 0.1081 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.3969 | 0.0075 | 0.003682 | 0.3609 | 0.1524 | 0.2888 | -0.1244
2012-10-30 22:02 | 26.7895 | 22.723 | 22.5144 | 22.7029 | 22.2737 | 0.1098 | 0.006716 | 0.32643 | 0.002192 | 48.6044 | 0.006924 | 0.21486 | 0.001488 | 31.0311 | 22.4141 | 0.0076 | 0.00368 | 0.3089 | 0.1003 | 0.2888 | -0.1404
2012-10-30 22:02 | 26.782 | 22.7074 | 22.535 | 22.6754 | 22.2461 | 0.1121 | 0.006751 | 0.32638 | 0.002203 | 48.346 | 0.006921 | 0.2149 | 0.001487 | 31.0507 | 22.4066 | 0.0077 | 0.00369 | 0.3008 | 0.1284 | 0.2688 | -0.1605
2012-10-30 22:03 | 26.7768 | 22.6862 | 22.5539 | 22.6581 | 22.265 | 0.1081 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 22.4375 | 0.0075 | 0.003687 | 0.2487 | 0.1164 | 0.2206 | -0.1725
2012-10-30 22:03 | 26.7858 | 22.7193 | 22.5388 | 22.7072 | 22.3101 | 0.1075 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.4345 | 0.0074 | 0.003681 | 0.2848 | 0.1043 | 0.2727 | -0.1244
2012-10-30 22:03 | 26.782 | 22.7395 | 22.5711 | 22.7034 | 22.2301 | 0.1075 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 22.3986 | 0.0074 | 0.003679 | 0.3409 | 0.1725 | 0.3048 | -0.1685
2012-10-30 22:03 | 26.7835 | 22.749 | 22.5525 | 22.6849 | 22.2516 | 0.1185 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006917 | 0.21482 | 0.001486 | 31.0561 | 22.4041 | 0.0082 | 0.003679 | 0.3449 | 0.1484 | 0.2808 | -0.1525
2012-10-30 22:03 | 26.782 | 22.7475 | 22.551 | 22.6874 | 22.2461 | 0.1046 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006936 | 0.21482 | 0.00149 | 30.971 | 22.3946 | 0.0072 | 0.003688 | 0.3529 | 0.1564 | 0.2928 | -0.1485
2012-10-30 22:03 | 26.7813 | 22.7188 | 22.5784 | 22.6346 | 22.2735 | 0.1162 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3939 | 0.008 | 0.003679 | 0.3249 | 0.1845 | 0.2407 | -0.1204
2012-10-30 22:03 | 26.7858 | 22.7273 | 22.5789 | 22.6431 | 22.278 | 0.1098 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3743 | 0.0076 | 0.003685 | 0.353 | 0.2046 | 0.2688 | -0.0963
2012-10-30 22:03 | 26.7843 | 22.7097 | 22.5734 | 22.6977 | 22.2444 | 0.104 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3929 | 0.0072 | 0.00368 | 0.3168 | 0.1805 | 0.3048 | -0.1485
2012-10-30 22:03 | 26.7783 | 22.7639 | 22.5834 | 22.6636 | 22.2344 | 0.1011 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 22.419 | 0.007 | 0.003678 | 0.3449 | 0.1644 | 0.2446 | -0.1846
2012-10-30 22:03 | 26.7828 | 22.7603 | 22.5438 | 22.6842 | 22.263 | 0.111 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006927 | 0.21486 | 0.001488 | 31.0176 | 22.3593 | 0.0077 | 0.003686 | 0.401 | 0.1845 | 0.3249 | -0.0963
2012-10-30 22:03 | 26.7835 | 22.7129 | 22.5445 | 22.6488 | 22.2677 | 0.1092 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.4322 | 0.0075 | 0.003681 | 0.2807 | 0.1123 | 0.2166 | -0.1645
2012-10-30 22:03 | 26.7858 | 22.7633 | 22.5348 | 22.6591 | 22.258 | 0.1139 | 0.006713 | 0.32644 | 0.002191 | 48.6283 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3904 | 0.0079 | 0.00368 | 0.3729 | 0.1444 | 0.2687 | -0.1324
2012-10-30 22:03 | 26.7813 | 22.7308 | 22.5503 | 22.6626 | 22.2454 | 0.111 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3939 | 0.0077 | 0.003682 | 0.3369 | 0.1564 | 0.2687 | -0.1485
2012-10-30 22:03 | 26.7903 | 22.7678 | 22.5794 | 22.6676 | 22.2665 | 0.1191 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.447 | 0.0082 | 0.003682 | 0.3208 | 0.1324 | 0.2206 | -0.1805
2012-10-30 22:03 | 26.7925 | 22.734 | 22.5656 | 22.6538 | 22.2767 | 0.4386 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.3931 | 0.0302 | 0.003688 | 0.3409 | 0.1725 | 0.2607 | -0.1164
2012-10-30 22:03 | 26.7865 | 22.74 | 22.6678 | 22.6558 | 22.3148 | 6.4705 | 0.006738 | 0.32656 | 0.0022 | 48.465 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.4312 | 0.4461 | 0.003688 | 0.3088 | 0.2366 | 0.2246 | -0.1164
2012-10-30 22:03 | 26.7925 | 22.8542 | 22.726 | 22.6538 | 22.3168 | 15.2076 | 0.006716 | 0.3265 | 0.002193 | 48.6152 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.4332 | 1.0484 | 0.003683 | 0.421 | 0.2928 | 0.2206 | -0.1164
2012-10-30 22:03 | 26.7865 | 22.9324 | 22.7761 | 22.6478 | 22.3108 | 22.308 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006914 | 0.21477 | 0.001485 | 31.0633 | 22.4272 | 1.5379 | 0.003681 | 0.5052 | 0.3489 | 0.2206 | -0.1164
2012-10-30 22:03 | 26.7828 | 22.9687 | 22.7924 | 22.7122 | 22.3232 | 27.2891 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3994 | 1.8813 | 0.003685 | 0.5693 | 0.393 | 0.3128 | -0.0762
2012-10-30 22:03 | 26.7858 | 22.9717 | 22.7754 | 22.7152 | 22.3783 | 30.1793 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006952 | 0.21479 | 0.001493 | 30.8956 | 22.3623 | 2.0806 | 0.00369 | 0.6094 | 0.4131 | 0.3529 | 0.016
2012-10-30 22:03 | 26.7843 | 22.9983 | 22.7378 | 22.7538 | 22.4009 | 31.4103 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3969 | 2.1654 | 0.003686 | 0.6014 | 0.3409 | 0.3569 | 0.004
2012-10-30 22:03 | 26.7865 | 23.0285 | 22.744 | 22.8121 | 22.4352 | 32.0322 | 0.006713 | 0.32644 | 0.002191 | 48.6283 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3871 | 2.2083 | 0.003678 | 0.6414 | 0.3569 | 0.425 | 0.0481
2012-10-30 22:03 | 26.7903 | 23.0123 | 22.6997 | 22.852 | 22.435 | 32.3113 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.3628 | 2.2275 | 0.003685 | 0.6495 | 0.3369 | 0.4892 | 0.0722
2012-10-30 22:03 | 26.785 | 23.015 | 22.6743 | 22.8026 | 22.4498 | 32.4321 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006936 | 0.21489 | 0.00149 | 30.9815 | 22.3495 | 2.2359 | 0.003685 | 0.6655 | 0.3248 | 0.4531 | 0.1003
2012-10-30 22:03 | 26.7918 | 22.9777 | 22.7373 | 22.8615 | 22.4766 | 32.5332 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.4084 | 2.2428 | 0.003684 | 0.5693 | 0.3289 | 0.4531 | 0.0682
2012-10-30 22:03 | 26.7775 | 23.0075 | 22.6949 | 22.8272 | 22.4703 | 32.613 | 0.006722 | 0.32649 | 0.002195 | 48.5696 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.4142 | 2.2483 | 0.003684 | 0.5933 | 0.2807 | 0.413 | 0.0561
2012-10-30 22:03 | 26.788 | 22.986 | 22.7495 | 22.8297 | 22.4808 | 32.6638 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.4086 | 2.2518 | 0.003685 | 0.5774 | 0.3409 | 0.4211 | 0.0722
2012-10-30 22:04 | 26.779 | 22.993 | 22.7525 | 22.8367 | 22.4718 | 32.7008 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006908 | 0.21483 | 0.001484 | 31.0987 | 22.3756 | 2.2544 | 0.00368 | 0.6174 | 0.3769 | 0.4611 | 0.0962
2012-10-30 22:04 | 26.7925 | 22.9784 | 22.738 | 22.8662 | 22.4893 | 32.7465 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 22.373 | 2.2575 | 0.003684 | 0.6054 | 0.365 | 0.4932 | 0.1163
2012-10-30 22:04 | 26.7933 | 23.0193 | 22.7348 | 22.859 | 22.4701 | 32.7725 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.3417 | 2.2593 | 0.003685 | 0.6776 | 0.3931 | 0.5173 | 0.1284
2012-10-30 22:04 | 26.7933 | 22.9993 | 22.7588 | 22.851 | 22.5102 | 33.0684 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3497 | 2.2797 | 0.003683 | 0.6496 | 0.4091 | 0.5013 | 0.1605
2012-10-30 22:04 | 26.7843 | 22.9903 | 22.7418 | 22.858 | 22.4932 | 33.5625 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.006921 | 0.2148 | 0.001487 | 31.0361 | 22.3488 | 2.3138 | 0.00368 | 0.6415 | 0.393 | 0.5092 | 0.1444
2012-10-30 22:04 | 26.7888 | 22.9627 | 22.7663 | 22.8465 | 22.5338 | 33.7746 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3613 | 2.3284 | 0.003682 | 0.6014 | 0.405 | 0.4852 | 0.1725
2012-10-30 22:04 | 26.788 | 23.006 | 22.7335 | 22.8698 | 22.4969 | 33.9052 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.006911 | 0.21479 | 0.001484 | 31.0789 | 22.3123 | 2.3374 | 0.003682 | 0.6937 | 0.4212 | 0.5575 | 0.1846
2012-10-30 22:04 | 26.788 | 23.002 | 22.7816 | 22.8778 | 22.5209 | 33.989 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.4006 | 2.3432 | 0.003679 | 0.6014 | 0.381 | 0.4772 | 0.1203
2012-10-30 22:04 | 26.785 | 22.987 | 22.7545 | 22.8547 | 22.5099 | 34.0416 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.00693 | 0.21476 | 0.001488 | 30.9896 | 22.3495 | 2.3468 | 0.003683 | 0.6375 | 0.405 | 0.5052 | 0.1604
2012-10-30 22:04 | 26.782 | 23 | 22.7916 | 22.8477 | 22.4909 | 34.0821 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.4227 | 2.3496 | 0.003686 | 0.5773 | 0.3689 | 0.425 | 0.0682
2012-10-30 22:04 | 26.7783 | 22.9763 | 22.7999 | 22.84 | 22.5433 | 34.1127 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.407 | 2.3517 | 0.003681 | 0.5693 | 0.3929 | 0.433 | 0.1363
2012-10-30 22:04 | 26.7783 | 22.9923 | 22.7759 | 22.844 | 22.5193 | 34.1433 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.4471 | 2.3538 | 0.003683 | 0.5452 | 0.3288 | 0.3969 | 0.0722
2012-10-30 22:04 | 26.7768 | 22.9507 | 22.7624 | 22.8986 | 22.5579 | 34.1699 | 0.006703 | 0.32638 | 0.002188 | 48.6922 | 0.006921 | 0.21483 | 0.001487 | 31.0403 | 22.4335 | 2.3557 | 0.003675 | 0.5172 | 0.3289 | 0.4651 | 0.1244
2012-10-30 22:04 | 26.7753 | 22.9773 | 22.7769 | 22.829 | 22.5123 | 34.1948 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3879 | 2.3574 | 0.003684 | 0.5894 | 0.389 | 0.4411 | 0.1244
2012-10-30 22:04 | 26.7723 | 22.9302 | 22.7939 | 22.8741 | 22.5413 | 34.2057 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3809 | 2.3581 | 0.003682 | 0.5493 | 0.413 | 0.4932 | 0.1604
2012-10-30 22:04 | 26.7633 | 22.9693 | 22.7729 | 22.8651 | 22.5524 | 34.2231 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3759 | 2.3593 | 0.003684 | 0.5934 | 0.397 | 0.4892 | 0.1765
2012-10-30 22:04 | 26.7701 | 22.9881 | 22.8278 | 22.8198 | 22.5231 | 34.2254 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.3907 | 2.3595 | 0.003683 | 0.5974 | 0.4371 | 0.4291 | 0.1324
2012-10-30 22:04 | 26.7663 | 22.9603 | 22.836 | 22.8761 | 22.5755 | 34.2364 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.3749 | 2.3603 | 0.003685 | 0.5854 | 0.4611 | 0.5012 | 0.2006
2012-10-30 22:04 | 26.7641 | 22.9701 | 22.7577 | 22.8459 | 22.5532 | 34.245 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006936 | 0.21479 | 0.00149 | 30.9668 | 22.3566 | 2.3609 | 0.003687 | 0.6135 | 0.4011 | 0.4893 | 0.1966
2012-10-30 22:04 | 26.7581 | 22.936 | 22.7797 | 22.8359 | 22.5552 | 34.2624 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 22.3747 | 2.362 | 0.00368 | 0.5613 | 0.405 | 0.4612 | 0.1805
2012-10-30 22:04 | 26.7693 | 22.9593 | 22.7949 | 22.8551 | 22.5744 | 34.2745 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 22.3578 | 2.3629 | 0.00368 | 0.6015 | 0.4371 | 0.4973 | 0.2166
2012-10-30 22:04 | 26.7573 | 22.9072 | 22.7348 | 22.823 | 22.5223 | 34.2699 | 0.0067 | 0.32643 | 0.002187 | 48.7204 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3619 | 2.3626 | 0.003676 | 0.5453 | 0.3729 | 0.4611 | 0.1604
2012-10-30 22:04 | 26.7588 | 22.9327 | 22.7925 | 22.8446 | 22.5078 | 34.2895 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.3393 | 2.3639 | 0.003682 | 0.5934 | 0.4532 | 0.5053 | 0.1685
2012-10-30 22:04 | 26.7663 | 22.9242 | 22.8 | 22.8681 | 22.5313 | 34.3051 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006917 | 0.21484 | 0.001486 | 31.0603 | 22.3548 | 2.365 | 0.003681 | 0.5694 | 0.4452 | 0.5133 | 0.1765
2012-10-30 22:04 | 26.7641 | 22.946 | 22.8098 | 22.8338 | 22.5412 | 34.2924 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3928 | 2.3641 | 0.003683 | 0.5532 | 0.417 | 0.441 | 0.1484
2012-10-30 22:04 | 26.7513 | 22.9413 | 22.789 | 22.805 | 22.5244 | 34.3051 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.3639 | 2.365 | 0.003684 | 0.5774 | 0.4251 | 0.4411 | 0.1605
2012-10-30 22:05 | 26.7648 | 22.9467 | 22.8065 | 22.8305 | 22.5499 | 34.3479 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.3694 | 2.3679 | 0.003688 | 0.5773 | 0.4371 | 0.4611 | 0.1805
2012-10-30 22:05 | 26.7551 | 22.9491 | 22.8048 | 22.8208 | 22.5362 | 35.0628 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3597 | 2.4172 | 0.003686 | 0.5894 | 0.4451 | 0.4611 | 0.1765
2012-10-30 22:05 | 26.7498 | 22.9237 | 22.8035 | 22.8155 | 22.5549 | 35.5344 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.3584 | 2.4497 | 0.003687 | 0.5653 | 0.4451 | 0.4571 | 0.1965
2012-10-30 22:05 | 26.7626 | 22.9566 | 22.7602 | 22.8003 | 22.5276 | 35.7511 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3953 | 2.4647 | 0.003688 | 0.5613 | 0.3649 | 0.405 | 0.1323
2012-10-30 22:05 | 26.7588 | 22.9447 | 22.8085 | 22.8606 | 22.5198 | 35.8783 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.3794 | 2.4735 | 0.003683 | 0.5653 | 0.4291 | 0.4812 | 0.1404
2012-10-30 22:05 | 26.7476 | 22.9416 | 22.8093 | 22.7692 | 22.5888 | 35.947 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 22.3602 | 2.4782 | 0.003686 | 0.5814 | 0.4491 | 0.409 | 0.2286
2012-10-30 22:05 | 26.7566 | 22.9305 | 22.7983 | 22.8183 | 22.5818 | 36.0002 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006933 | 0.21486 | 0.00149 | 30.9908 | 22.3572 | 2.4819 | 0.003688 | 0.5733 | 0.4411 | 0.4611 | 0.2246
2012-10-30 22:05 | 26.7468 | 22.9087 | 22.7765 | 22.8005 | 22.564 | 36.0424 | 0.006732 | 0.3265 | 0.002198 | 48.4997 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.3474 | 2.4848 | 0.003686 | 0.5613 | 0.4291 | 0.4531 | 0.2166
2012-10-30 22:05 | 26.7461 | 22.9361 | 22.7958 | 22.7677 | 22.5472 | 36.0811 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006943 | 0.2148 | 0.001491 | 30.9378 | 22.3386 | 2.4874 | 0.003688 | 0.5975 | 0.4572 | 0.4291 | 0.2086
2012-10-30 22:05 | 26.7446 | 22.8905 | 22.7943 | 22.8224 | 22.5578 | 36.0956 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3853 | 2.4884 | 0.003685 | 0.5052 | 0.409 | 0.4371 | 0.1725
2012-10-30 22:05 | 26.7491 | 22.927 | 22.7948 | 22.8108 | 22.5623 | 36.1141 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.3737 | 2.4897 | 0.003683 | 0.5533 | 0.4211 | 0.4371 | 0.1886
2012-10-30 22:05 | 26.7476 | 22.9055 | 22.7813 | 22.7773 | 22.5728 | 36.1505 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3522 | 2.4922 | 0.003685 | 0.5533 | 0.4291 | 0.4251 | 0.2206
2012-10-30 22:05 | 26.7454 | 22.9153 | 22.8031 | 22.7871 | 22.5666 | 36.1522 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.4061 | 2.4923 | 0.003684 | 0.5092 | 0.397 | 0.381 | 0.1605
2012-10-30 22:05 | 26.7483 | 22.9342 | 22.786 | 22.782 | 22.5695 | 36.1695 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006933 | 0.21476 | 0.001489 | 30.9762 | 22.385 | 2.4935 | 0.003683 | 0.5492 | 0.401 | 0.397 | 0.1845
2012-10-30 22:05 | 26.7439 | 22.9298 | 22.8257 | 22.7896 | 22.5811 | 36.1701 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.3806 | 2.4936 | 0.003688 | 0.5492 | 0.4451 | 0.409 | 0.2005
2012-10-30 22:05 | 26.7461 | 22.916 | 22.7958 | 22.7798 | 22.6114 | 36.1938 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3948 | 2.4952 | 0.003686 | 0.5212 | 0.401 | 0.385 | 0.2166
2012-10-30 22:05 | 26.7409 | 22.9188 | 22.8026 | 22.7906 | 22.5661 | 36.1996 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.3856 | 2.4956 | 0.00369 | 0.5332 | 0.417 | 0.405 | 0.1805
2012-10-30 22:05 | 26.7364 | 22.9023 | 22.7981 | 22.7901 | 22.5696 | 36.2239 | 0.006725 | 0.3265 | 0.002196 | 48.5502 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.3771 | 2.4973 | 0.003687 | 0.5252 | 0.421 | 0.413 | 0.1925
2012-10-30 22:05 | 26.7401 | 22.8499 | 22.7818 | 22.8018 | 22.5533 | 36.2169 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006936 | 0.21486 | 0.00149 | 30.9774 | 22.3647 | 2.4968 | 0.003682 | 0.4852 | 0.4171 | 0.4371 | 0.1886
2012-10-30 22:05 | 26.7349 | 22.9008 | 22.8127 | 22.7686 | 22.5761 | 36.2314 | 0.006703 | 0.32643 | 0.002188 | 48.6986 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3876 | 2.4978 | 0.003676 | 0.5132 | 0.4251 | 0.381 | 0.1885
2012-10-30 22:05 | 26.7319 | 22.9219 | 22.7616 | 22.8097 | 22.5411 | 36.2325 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3926 | 2.4979 | 0.003685 | 0.5293 | 0.369 | 0.4171 | 0.1485
2012-10-30 22:05 | 26.7341 | 22.88 | 22.8279 | 22.7998 | 22.5352 | 36.2377 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006933 | 0.21473 | 0.001489 | 30.972 | 22.3628 | 2.4982 | 0.003691 | 0.5172 | 0.4651 | 0.437 | 0.1724
2012-10-30 22:05 | 26.7304 | 22.9043 | 22.744 | 22.748 | 22.5556 | 36.2649 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.3631 | 2.5001 | 0.003688 | 0.5412 | 0.3809 | 0.3849 | 0.1925
2012-10-30 22:05 | 26.7319 | 22.8698 | 22.8056 | 22.7535 | 22.5691 | 36.2499 | 0.006713 | 0.3264 | 0.002191 | 48.6219 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3565 | 2.4991 | 0.00368 | 0.5133 | 0.4491 | 0.397 | 0.2126
2012-10-30 22:05 | 26.7319 | 22.8698 | 22.8337 | 22.7976 | 22.5892 | 36.2539 | 0.006741 | 0.32633 | 0.0022 | 48.4091 | 0.006917 | 0.2148 | 0.001486 | 31.0541 | 22.3926 | 2.4993 | 0.003686 | 0.4772 | 0.4411 | 0.405 | 0.1966
2012-10-30 22:05 | 26.7281 | 22.882 | 22.8099 | 22.8139 | 22.5533 | 36.2695 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006943 | 0.21476 | 0.001491 | 30.9316 | 22.3327 | 2.5004 | 0.003686 | 0.5493 | 0.4772 | 0.4812 | 0.2206
2012-10-30 22:06 | 26.7311 | 22.9251 | 22.7768 | 22.7608 | 22.5723 | 36.277 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.006911 | 0.21474 | 0.001484 | 31.0726 | 22.2996 | 2.5009 | 0.00368 | 0.6255 | 0.4772 | 0.4612 | 0.2727
2012-10-30 22:06 | 26.7229 | 22.8928 | 22.7766 | 22.7886 | 22.5641 | 36.288 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.2994 | 2.5017 | 0.003684 | 0.5934 | 0.4772 | 0.4892 | 0.2647
2012-10-30 22:06 | 26.7124 | 22.8703 | 22.7701 | 22.7781 | 22.5576 | 36.2765 | 0.00671 | 0.32631 | 0.00219 | 48.6307 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3009 | 2.5009 | 0.003678 | 0.5694 | 0.4692 | 0.4772 | 0.2567
2012-10-30 22:06 | 26.7139 | 22.9159 | 22.8037 | 22.7476 | 22.5632 | 36.4423 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 22.3426 | 2.5123 | 0.00368 | 0.5733 | 0.4611 | 0.405 | 0.2206
2012-10-30 22:06 | 26.7177 | 22.9117 | 22.8155 | 22.7554 | 22.5269 | 37.4733 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.006917 | 0.21476 | 0.001485 | 31.0479 | 22.3303 | 2.5834 | 0.003677 | 0.5814 | 0.4852 | 0.4251 | 0.1966
2012-10-30 22:06 | 26.7094 | 22.8873 | 22.7912 | 22.7631 | 22.5667 | 37.8802 | 0.006719 | 0.32628 | 0.002192 | 48.5612 | 0.006939 | 0.21476 | 0.00149 | 30.9494 | 22.2899 | 2.6115 | 0.003682 | 0.5974 | 0.5013 | 0.4732 | 0.2768
2012-10-30 22:06 | 26.7132 | 22.9072 | 22.7789 | 22.7789 | 22.5544 | 38.0495 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 22.3017 | 2.6231 | 0.003681 | 0.6055 | 0.4772 | 0.4772 | 0.2527
2012-10-30 22:06 | 26.6997 | 22.9257 | 22.7975 | 22.7534 | 22.565 | 38.1628 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.006908 | 0.21476 | 0.001484 | 31.0883 | 22.3524 | 2.6309 | 0.00368 | 0.5733 | 0.4451 | 0.401 | 0.2126
2012-10-30 22:06 | 26.6989 | 22.9449 | 22.7847 | 22.7646 | 22.5321 | 38.2206 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.006921 | 0.21479 | 0.001487 | 31.034 | 22.3476 | 2.6349 | 0.003682 | 0.5973 | 0.4371 | 0.417 | 0.1845
2012-10-30 22:06 | 26.6937 | 22.9197 | 22.7514 | 22.7354 | 22.5349 | 38.2587 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006914 | 0.21476 | 0.001485 | 31.0613 | 22.3023 | 2.6376 | 0.003679 | 0.6174 | 0.4491 | 0.4331 | 0.2326
2012-10-30 22:06 | 26.7019 | 22.9159 | 22.7756 | 22.7716 | 22.5672 | 38.3726 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3506 | 2.6454 | 0.003683 | 0.5653 | 0.425 | 0.421 | 0.2166
2012-10-30 22:06 | 26.7012 | 22.9272 | 22.811 | 22.783 | 22.5745 | 39.9451 | 0.00671 | 0.32643 | 0.00219 | 48.6478 | 0.006924 | 0.21484 | 0.001488 | 31.0289 | 22.378 | 2.7538 | 0.003678 | 0.5492 | 0.433 | 0.405 | 0.1965
2012-10-30 22:06 | 26.6959 | 22.8979 | 22.8217 | 22.7135 | 22.5452 | 40.6791 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006952 | 0.2149 | 0.001494 | 30.9123 | 22.3446 | 2.8044 | 0.003688 | 0.5533 | 0.4771 | 0.3689 | 0.2006
2012-10-30 22:06 | 26.6967 | 22.8786 | 22.8065 | 22.7624 | 22.5259 | 40.9213 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006924 | 0.21483 | 0.001487 | 31.0269 | 22.3494 | 2.8211 | 0.003682 | 0.5292 | 0.4571 | 0.413 | 0.1765
2012-10-30 22:06 | 26.6877 | 22.9297 | 22.7855 | 22.7414 | 22.541 | 41.0287 | 0.006707 | 0.32637 | 0.002189 | 48.6611 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.3364 | 2.8285 | 0.003678 | 0.5933 | 0.4491 | 0.405 | 0.2046
2012-10-30 22:06 | 26.6982 | 22.9002 | 22.808 | 22.7519 | 22.5875 | 41.0987 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3349 | 2.8333 | 0.003687 | 0.5653 | 0.4731 | 0.417 | 0.2526
2012-10-30 22:06 | 26.6945 | 22.9285 | 22.8364 | 22.7362 | 22.5798 | 42.5464 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3392 | 2.9331 | 0.003687 | 0.5893 | 0.4972 | 0.397 | 0.2406
2012-10-30 22:06 | 26.6922 | 22.9342 | 22.8221 | 22.774 | 22.5815 | 43.7947 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.006939 | 0.21489 | 0.001491 | 30.9682 | 22.373 | 3.0192 | 0.003685 | 0.5612 | 0.4491 | 0.401 | 0.2085
2012-10-30 22:06 | 26.6885 | 22.9185 | 22.8264 | 22.7622 | 22.5698 | 44.3039 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006933 | 0.21487 | 0.00149 | 30.9928 | 22.3813 | 3.0543 | 0.003689 | 0.5372 | 0.4451 | 0.3809 | 0.1885
2012-10-30 22:06 | 26.6959 | 22.9379 | 22.8177 | 22.7777 | 22.5933 | 44.5466 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006936 | 0.21474 | 0.001489 | 30.9606 | 22.3727 | 3.071 | 0.003686 | 0.5652 | 0.445 | 0.405 | 0.2206
2012-10-30 22:06 | 26.6974 | 22.9074 | 22.8192 | 22.7952 | 22.5787 | 44.8616 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.006921 | 0.21482 | 0.001487 | 31.0381 | 22.3662 | 3.0928 | 0.003684 | 0.5412 | 0.453 | 0.429 | 0.2125
2012-10-30 22:06 | 26.6982 | 22.9322 | 22.812 | 22.7639 | 22.5996 | 45.814 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.4031 | 3.1584 | 0.003683 | 0.5291 | 0.4089 | 0.3608 | 0.1965
2012-10-30 22:06 | 26.7019 | 22.9359 | 22.8318 | 22.7756 | 22.6273 | 46.2584 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.006927 | 0.21486 | 0.001488 | 31.0176 | 22.3907 | 3.1891 | 0.003683 | 0.5452 | 0.4411 | 0.3849 | 0.2366
2012-10-30 22:07 | 26.6907 | 22.9247 | 22.8045 | 22.7685 | 22.58 | 46.5456 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.4036 | 3.2089 | 0.003687 | 0.5211 | 0.4009 | 0.3649 | 0.1764
2012-10-30 22:07 | 26.693 | 22.915 | 22.7908 | 22.7748 | 22.5823 | 46.8057 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.3818 | 3.2268 | 0.003684 | 0.5332 | 0.409 | 0.393 | 0.2005
2012-10-30 22:07 | 26.6937 | 22.9357 | 22.8796 | 22.7955 | 22.575 | 47.6732 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.3544 | 3.2866 | 0.003686 | 0.5813 | 0.5252 | 0.4411 | 0.2206
2012-10-30 22:07 | 26.6855 | 22.8995 | 22.8394 | 22.7913 | 22.5869 | 47.9373 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006939 | 0.21477 | 0.00149 | 30.9514 | 22.3623 | 3.3048 | 0.00369 | 0.5372 | 0.4771 | 0.429 | 0.2246
2012-10-30 22:07 | 26.6952 | 22.9573 | 22.8291 | 22.785 | 22.6166 | 48.0592 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006927 | 0.21487 | 0.001488 | 31.0196 | 22.4161 | 3.3132 | 0.003682 | 0.5412 | 0.413 | 0.3689 | 0.2005
2012-10-30 22:07 | 26.6862 | 22.9523 | 22.8281 | 22.78 | 22.5836 | 48.154 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.375 | 3.3197 | 0.003691 | 0.5773 | 0.4531 | 0.405 | 0.2086
2012-10-30 22:07 | 26.6907 | 22.9127 | 22.8206 | 22.7885 | 22.6161 | 48.2008 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.4076 | 3.323 | 0.003687 | 0.5051 | 0.413 | 0.3809 | 0.2085
2012-10-30 22:07 | 26.6855 | 22.9235 | 22.7833 | 22.7873 | 22.651 | 48.2413 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.3944 | 3.3258 | 0.003683 | 0.5291 | 0.3889 | 0.3929 | 0.2566
2012-10-30 22:07 | 26.6915 | 22.9255 | 22.8214 | 22.7933 | 22.629 | 48.28 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.4325 | 3.3284 | 0.003689 | 0.493 | 0.3889 | 0.3608 | 0.1965
2012-10-30 22:07 | 26.6855 | 22.9155 | 22.8113 | 22.7352 | 22.5468 | 48.31 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006949 | 0.21486 | 0.001493 | 30.9194 | 22.3703 | 3.3305 | 0.00369 | 0.5452 | 0.441 | 0.3649 | 0.1765
2012-10-30 22:07 | 26.6832 | 22.9092 | 22.8131 | 22.789 | 22.6086 | 48.3303 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.00693 | 0.21494 | 0.00149 | 31.0166 | 22.3199 | 3.3319 | 0.003684 | 0.5893 | 0.4932 | 0.4691 | 0.2887
2012-10-30 22:07 | 26.6765 | 22.9466 | 22.8023 | 22.7903 | 22.5739 | 48.3655 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3693 | 3.3343 | 0.003687 | 0.5773 | 0.433 | 0.421 | 0.2046
2012-10-30 22:07 | 26.6847 | 22.8987 | 22.8226 | 22.8386 | 22.6182 | 48.3788 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006943 | 0.21483 | 0.001492 | 30.942 | 22.3936 | 3.3352 | 0.003688 | 0.5051 | 0.429 | 0.445 | 0.2246
2012-10-30 22:07 | 26.675 | 22.905 | 22.8008 | 22.7808 | 22.6205 | 48.3863 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.006939 | 0.21486 | 0.001491 | 30.964 | 22.3839 | 3.3358 | 0.003689 | 0.5211 | 0.4169 | 0.3969 | 0.2366
2012-10-30 22:07 | 26.6825 | 22.8644 | 22.7963 | 22.7643 | 22.628 | 48.4054 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.3793 | 3.3371 | 0.00369 | 0.4851 | 0.417 | 0.385 | 0.2487
2012-10-30 22:07 | 26.6862 | 22.9202 | 22.8241 | 22.7399 | 22.6317 | 48.7435 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006924 | 0.21489 | 0.001488 | 31.0352 | 22.4151 | 3.3604 | 0.003686 | 0.5051 | 0.409 | 0.3248 | 0.2166
2012-10-30 22:07 | 26.678 | 22.908 | 22.7918 | 22.7678 | 22.6034 | 49.0845 | 0.006732 | 0.32654 | 0.002198 | 48.5061 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.3829 | 3.3839 | 0.003689 | 0.5251 | 0.4089 | 0.3849 | 0.2205
2012-10-30 22:07 | 26.6825 | 22.8885 | 22.8404 | 22.7763 | 22.6079 | 49.218 | 0.006732 | 0.32649 | 0.002198 | 48.4975 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.4034 | 3.3931 | 0.003686 | 0.4851 | 0.437 | 0.3729 | 0.2045
2012-10-30 22:07 | 26.6832 | 22.8932 | 22.8411 | 22.793 | 22.5846 | 49.3122 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.372 | 3.3996 | 0.003685 | 0.5212 | 0.4691 | 0.421 | 0.2126
2012-10-30 22:07 | 26.6825 | 22.8885 | 22.8204 | 22.7883 | 22.5959 | 49.3879 | 0.006716 | 0.32646 | 0.002192 | 48.6087 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.4154 | 3.4048 | 0.00368 | 0.4731 | 0.405 | 0.3729 | 0.1805
2012-10-30 22:07 | 26.684 | 22.898 | 22.8379 | 22.7898 | 22.5894 | 49.4318 | 0.006713 | 0.32641 | 0.002191 | 48.624 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3688 | 3.4078 | 0.00368 | 0.5292 | 0.4691 | 0.421 | 0.2206
2012-10-30 22:07 | 26.6862 | 22.8802 | 22.8201 | 22.8281 | 22.5796 | 49.4751 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3991 | 3.4108 | 0.003682 | 0.4811 | 0.421 | 0.429 | 0.1805
2012-10-30 22:07 | 26.687 | 22.9451 | 22.8529 | 22.8409 | 22.6525 | 49.5214 | 0.006719 | 0.32647 | 0.002194 | 48.5892 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 22.3758 | 3.414 | 0.003683 | 0.5693 | 0.4771 | 0.4651 | 0.2767
2012-10-30 22:08 | 26.6795 | 22.9175 | 22.8214 | 22.7773 | 22.625 | 49.526 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3563 | 3.4143 | 0.003679 | 0.5612 | 0.4651 | 0.421 | 0.2687
2012-10-30 22:08 | 26.6855 | 22.8875 | 22.8073 | 22.7753 | 22.6109 | 49.5676 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3823 | 3.4172 | 0.003686 | 0.5052 | 0.425 | 0.393 | 0.2286
2012-10-30 22:08 | 26.687 | 22.917 | 22.8128 | 22.7808 | 22.5884 | 49.5711 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 22.3999 | 3.4174 | 0.003683 | 0.5171 | 0.4129 | 0.3809 | 0.1885
2012-10-30 22:08 | 26.6817 | 22.8797 | 22.8035 | 22.7835 | 22.5911 | 49.6029 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3785 | 3.4196 | 0.003685 | 0.5012 | 0.425 | 0.405 | 0.2126
2012-10-30 22:08 | 26.6855 | 22.9195 | 22.7673 | 22.7753 | 22.6069 | 49.6254 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.00693 | 0.21483 | 0.001489 | 31 | 22.3864 | 3.4212 | 0.003688 | 0.5331 | 0.3809 | 0.3889 | 0.2205
2012-10-30 22:08 | 26.6795 | 22.9416 | 22.8174 | 22.7733 | 22.645 | 49.6387 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.006921 | 0.21486 | 0.001487 | 31.0445 | 22.3804 | 3.4221 | 0.003683 | 0.5612 | 0.437 | 0.3929 | 0.2646
2012-10-30 22:08 | 26.678 | 22.916 | 22.8359 | 22.7598 | 22.5834 | 49.6653 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006936 | 0.21492 | 0.001491 | 30.9856 | 22.3668 | 3.4239 | 0.003686 | 0.5492 | 0.4691 | 0.393 | 0.2166
2012-10-30 22:08 | 26.6832 | 22.8892 | 22.805 | 22.7489 | 22.6086 | 49.6815 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.006921 | 0.21486 | 0.001487 | 31.0445 | 22.3921 | 3.425 | 0.003683 | 0.4971 | 0.4129 | 0.3568 | 0.2165
2012-10-30 22:08 | 26.681 | 22.8589 | 22.8229 | 22.7467 | 22.5824 | 49.708 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.3217 | 3.4269 | 0.003681 | 0.5372 | 0.5012 | 0.425 | 0.2607
2012-10-30 22:08 | 26.687 | 22.889 | 22.8008 | 22.7487 | 22.6124 | 49.7069 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.006936 | 0.21486 | 0.00149 | 30.9774 | 22.3558 | 3.4268 | 0.003685 | 0.5332 | 0.445 | 0.3929 | 0.2566
2012-10-30 22:08 | 26.6922 | 22.9262 | 22.814 | 22.77 | 22.5896 | 49.7265 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.006927 | 0.2148 | 0.001488 | 31.0092 | 22.3449 | 3.4281 | 0.003681 | 0.5813 | 0.4691 | 0.4251 | 0.2447
2012-10-30 22:08 | 26.681 | 22.8589 | 22.8229 | 22.7948 | 22.5944 | 49.7364 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3698 | 3.4288 | 0.003689 | 0.4891 | 0.4531 | 0.425 | 0.2246
2012-10-30 22:08 | 26.6787 | 22.8727 | 22.7524 | 22.7524 | 22.6001 | 49.7514 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3475 | 3.4299 | 0.003682 | 0.5252 | 0.4049 | 0.4049 | 0.2526
2012-10-30 22:08 | 26.681 | 22.871 | 22.7507 | 22.7347 | 22.5784 | 49.7554 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006927 | 0.21489 | 0.001489 | 31.0218 | 22.3578 | 3.4301 | 0.003687 | 0.5132 | 0.3929 | 0.3769 | 0.2206
2012-10-30 22:08 | 26.675 | 22.901 | 22.8008 | 22.7568 | 22.5724 | 49.7861 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.006921 | 0.21483 | 0.001487 | 31.0403 | 22.3759 | 3.4323 | 0.003686 | 0.5251 | 0.4249 | 0.3809 | 0.1965
2012-10-30 22:08 | 26.6735 | 22.8955 | 22.7593 | 22.7513 | 22.5789 | 49.7959 | 0.006719 | 0.32649 | 0.002194 | 48.5913 | 0.006946 | 0.21486 | 0.001492 | 30.9328 | 22.3383 | 3.4329 | 0.003686 | 0.5572 | 0.421 | 0.413 | 0.2406
2012-10-30 22:08 | 26.6675 | 22.9015 | 22.7693 | 22.7653 | 22.5849 | 49.7982 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.006949 | 0.21486 | 0.001493 | 30.9194 | 22.3603 | 3.4331 | 0.003686 | 0.5412 | 0.409 | 0.405 | 0.2246
2012-10-30 22:08 | 26.6683 | 22.8583 | 22.7861 | 22.726 | 22.6138 | 49.8236 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.006939 | 0.2149 | 0.001491 | 30.9702 | 22.3611 | 3.4348 | 0.00369 | 0.4972 | 0.425 | 0.3649 | 0.2527
2012-10-30 22:08 | 26.6735 | 22.8755 | 22.7513 | 22.8074 | 22.6029 | 49.8265 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3383 | 3.435 | 0.003685 | 0.5372 | 0.413 | 0.4691 | 0.2646
2012-10-30 22:08 | 26.6683 | 22.8983 | 22.8062 | 22.7621 | 22.5857 | 49.8323 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.006936 | 0.21487 | 0.00149 | 30.9794 | 22.3692 | 3.4354 | 0.003688 | 0.5291 | 0.437 | 0.3929 | 0.2165
2012-10-30 22:08 | 26.6697 | 22.8556 | 22.7835 | 22.7595 | 22.5911 | 49.8595 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.21486 | 0.001488 | 31.0176 | 22.3264 | 3.4373 | 0.003683 | 0.5292 | 0.4571 | 0.4331 | 0.2647
2012-10-30 22:08 | 26.6645 | 22.8865 | 22.7743 | 22.7503 | 22.5819 | 49.8693 | 0.006741 | 0.32644 | 0.002201 | 48.4263 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.3333 | 3.438 | 0.00369 | 0.5532 | 0.441 | 0.417 | 0.2486
2012-10-30 22:08 | 26.666 | 22.856 | 22.8039 | 22.7598 | 22.6075 | 49.8785 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006917 | 0.21482 | 0.001486 | 31.0561 | 22.3027 | 3.4386 | 0.003684 | 0.5533 | 0.5012 | 0.4571 | 0.3048
2012-10-30 22:08 | 26.6727 | 22.8867 | 22.7945 | 22.6863 | 22.5901 | 49.8913 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 22.3254 | 3.4395 | 0.003686 | 0.5613 | 0.4691 | 0.3609 | 0.2647
2012-10-30 22:09 | 26.6772 | 22.8992 | 22.763 | 22.7509 | 22.6347 | 49.8976 | 0.006735 | 0.32631 | 0.002198 | 48.4502 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.3219 | 3.4399 | 0.003688 | 0.5773 | 0.4411 | 0.429 | 0.3128
2012-10-30 22:09 | 26.672 | 22.878 | 22.7778 | 22.7457 | 22.6014 | 49.9069 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.006924 | 0.21476 | 0.001487 | 31.0165 | 22.3247 | 3.4406 | 0.00368 | 0.5533 | 0.4531 | 0.421 | 0.2767
2012-10-30 22:09 | 26.672 | 22.906 | 22.7698 | 22.7097 | 22.5172 | 49.9028 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.006924 | 0.21484 | 0.001488 | 31.0289 | 22.3287 | 3.4403 | 0.003684 | 0.5773 | 0.4411 | 0.381 | 0.1885
2012-10-30 22:09 | 26.6653 | 22.9113 | 22.7871 | 22.7431 | 22.5827 | 49.9311 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.006933 | 0.21487 | 0.00149 | 30.9928 | 22.306 | 3.4423 | 0.003681 | 0.6053 | 0.4811 | 0.4371 | 0.2767
2012-10-30 22:09 | 26.6675 | 22.8855 | 22.7773 | 22.7132 | 22.5689 | 49.9433 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.006914 | 0.2148 | 0.001485 | 31.0675 | 22.3242 | 3.4431 | 0.003683 | 0.5613 | 0.4531 | 0.389 | 0.2447
2012-10-30 22:09 | 26.6697 | 22.8637 | 22.7996 | 22.7475 | 22.5631 | 49.9485 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.3465 | 3.4434 | 0.003684 | 0.5172 | 0.4531 | 0.401 | 0.2166
2012-10-30 22:09 | 26.6675 | 22.8855 | 22.7893 | 22.7292 | 22.5809 | 49.9577 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006924 | 0.21484 | 0.001488 | 31.0289 | 22.3523 | 3.4441 | 0.003684 | 0.5332 | 0.437 | 0.3769 | 0.2286
2012-10-30 22:09 | 26.6712 | 22.8331 | 22.769 | 22.7089 | 22.5846 | 49.9762 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3159 | 3.4454 | 0.003685 | 0.5172 | 0.4531 | 0.393 | 0.2687
2012-10-30 22:09 | 26.678 | 22.88 | 22.7638 | 22.7317 | 22.5754 | 49.9716 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.006927 | 0.21476 | 0.001488 | 31.003 | 22.3668 | 3.445 | 0.003685 | 0.5132 | 0.397 | 0.3649 | 0.2086
2012-10-30 22:09 | 26.672 | 22.866 | 22.7377 | 22.7337 | 22.5413 | 49.9964 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006924 | 0.21482 | 0.001487 | 31.0247 | 22.3608 | 3.4468 | 0.003684 | 0.5052 | 0.3769 | 0.3729 | 0.1805
2012-10-30 22:09 | 26.6683 | 22.8462 | 22.7741 | 22.718 | 22.5617 | 49.9935 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006924 | 0.21477 | 0.001487 | 31.0185 | 22.3611 | 3.4466 | 0.003682 | 0.4851 | 0.413 | 0.3569 | 0.2006
2012-10-30 22:09 | 26.6727 | 22.8386 | 22.7825 | 22.7264 | 22.554 | 50.0224 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.006927 | 0.21489 | 0.001489 | 31.0218 | 22.4137 | 3.4485 | 0.003689 | 0.4249 | 0.3688 | 0.3127 | 0.1403
2012-10-30 22:09 | 26.672 | 22.8499 | 22.7978 | 22.7017 | 22.6055 | 50.0317 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006943 | 0.21477 | 0.001491 | 30.9336 | 22.4009 | 3.4492 | 0.003689 | 0.449 | 0.3969 | 0.3008 | 0.2046
2012-10-30 22:09 | 26.6615 | 22.8635 | 22.7473 | 22.6511 | 22.595 | 50.0467 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006917 | 0.21483 | 0.001486 | 31.0583 | 22.4025 | 3.4502 | 0.003684 | 0.461 | 0.3448 | 0.2486 | 0.1925
2012-10-30 22:09 | 26.6705 | 22.8364 | 22.7843 | 22.7082 | 22.5518 | 50.0421 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.4155 | 3.4499 | 0.003687 | 0.4209 | 0.3688 | 0.2927 | 0.1363
2012-10-30 22:09 | 26.6742 | 22.8561 | 22.768 | 22.7159 | 22.5836 | 50.0594 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.006933 | 0.21492 | 0.00149 | 30.999 | 22.3951 | 3.4511 | 0.00369 | 0.461 | 0.3729 | 0.3208 | 0.1885
2012-10-30 22:09 | 26.6675 | 22.8695 | 22.7613 | 22.6972 | 22.5769 | 50.0721 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.3523 | 3.452 | 0.003688 | 0.5172 | 0.409 | 0.3449 | 0.2246
2012-10-30 22:09 | 26.669 | 22.8549 | 22.7948 | 22.7027 | 22.5583 | 50.0953 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 22.3458 | 3.4536 | 0.003692 | 0.5091 | 0.449 | 0.3569 | 0.2125
2012-10-30 22:09 | 26.6742 | 22.8161 | 22.752 | 22.7159 | 22.5676 | 50.1028 | 0.006744 | 0.32644 | 0.002202 | 48.4048 | 0.00693 | 0.21484 | 0.001489 | 31.002 | 22.347 | 3.4541 | 0.003691 | 0.4691 | 0.405 | 0.3689 | 0.2206
2012-10-30 22:09 | 26.675 | 22.8369 | 22.7167 | 22.6886 | 22.5764 | 50.0981 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 22.3398 | 3.4538 | 0.003687 | 0.4971 | 0.3769 | 0.3488 | 0.2366
2012-10-30 22:09 | 26.669 | 22.8029 | 22.7107 | 22.6987 | 22.5824 | 50.1207 | 0.006744 | 0.32638 | 0.002201 | 48.3962 | 0.006927 | 0.21479 | 0.001488 | 31.0071 | 22.3418 | 3.4553 | 0.003689 | 0.4611 | 0.3689 | 0.3569 | 0.2406
2012-10-30 22:09 | 26.6697 | 22.8356 | 22.7154 | 22.7034 | 22.555 | 50.1143 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.006927 | 0.21474 | 0.001488 | 31.0009 | 22.3986 | 3.4549 | 0.003686 | 0.437 | 0.3168 | 0.3048 | 0.1564
2012-10-30 22:09 | 26.6817 | 22.8196 | 22.7474 | 22.7474 | 22.5229 | 50.1409 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.006946 | 0.21484 | 0.001492 | 30.9306 | 22.3946 | 3.4567 | 0.003688 | 0.425 | 0.3528 | 0.3528 | 0.1283
2012-10-30 22:09 | 26.6705 | 22.8404 | 22.7883 | 22.6841 | 22.5438 | 50.149 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.00693 | 0.21474 | 0.001488 | 30.9874 | 22.3673 | 3.4573 | 0.003686 | 0.4731 | 0.421 | 0.3168 | 0.1765
2012-10-30 22:09 | 26.6735 | 22.8474 | 22.7513 | 22.7272 | 22.5508 | 50.1513 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3503 | 3.4574 | 0.003685 | 0.4971 | 0.401 | 0.3769 | 0.2005
2012-10-30 22:09 | 26.6727 | 22.8266 | 22.7344 | 22.6663 | 22.558 | 50.1617 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.006927 | 0.21484 | 0.001488 | 31.0154 | 22.3535 | 3.4581 | 0.003683 | 0.4731 | 0.3809 | 0.3128 | 0.2045
2012-10-30 22:10 | 26.6862 | 22.808 | 22.7519 | 22.7038 | 22.5755 | 50.1767 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.006936 | 0.21484 | 0.00149 | 30.9752 | 22.375 | 3.4592 | 0.003689 | 0.433 | 0.3769 | 0.3288 | 0.2005
2012-10-30 22:10 | 26.6802 | 22.806 | 22.7539 | 22.6577 | 22.5736 | 50.1877 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.006946 | 0.2149 | 0.001493 | 30.939 | 22.377 | 3.4599 | 0.003691 | 0.429 | 0.3769 | 0.2807 | 0.1966
2012-10-30 22:10 | 26.6787 | 22.8246 | 22.7484 | 22.6883 | 22.536 | 50.1773 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006936 | 0.21477 | 0.00149 | 30.9648 | 22.3715 | 3.4592 | 0.003687 | 0.4531 | 0.3769 | 0.3168 | 0.1645
2012-10-30 22:10 | 26.6795 | 22.8414 | 22.7492 | 22.641 | 22.5488 | 50.1843 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006936 | 0.21471 | 0.001489 | 30.9565 | 22.3563 | 3.4597 | 0.003688 | 0.4851 | 0.3929 | 0.2847 | 0.1925
2012-10-30 22:10 | 26.6892 | 22.8471 | 22.7589 | 22.6627 | 22.5385 | 50.1785 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.006939 | 0.21479 | 0.00149 | 30.9535 | 22.378 | 3.4593 | 0.003687 | 0.4691 | 0.3809 | 0.2847 | 0.1605
2012-10-30 22:10 | 26.6817 | 22.8316 | 22.7514 | 22.6793 | 22.571 | 50.1808 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.006936 | 0.21486 | 0.00149 | 30.9774 | 22.3264 | 3.4595 | 0.003689 | 0.5052 | 0.425 | 0.3529 | 0.2446
2012-10-30 22:10 | 26.6817 | 22.8396 | 22.7715 | 22.6913 | 22.531 | 50.1819 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.006939 | 0.21482 | 0.001491 | 30.9576 | 22.3585 | 3.4595 | 0.003684 | 0.4811 | 0.413 | 0.3328 | 0.1725
2012-10-30 22:10 | 26.678 | 22.8319 | 22.7277 | 22.6676 | 22.5433 | 50.1871 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3267 | 3.4599 | 0.003682 | 0.5052 | 0.401 | 0.3409 | 0.2166
2012-10-30 22:10 | 26.69 | 22.8359 | 22.7477 | 22.6916 | 22.5793 | 50.1831 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.006946 | 0.21482 | 0.001492 | 30.9264 | 22.3507 | 3.4596 | 0.003685 | 0.4852 | 0.397 | 0.3409 | 0.2286
2012-10-30 22:10 | 26.6802 | 22.8181 | 22.766 | 22.6658 | 22.5295 | 50.1814 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.006933 | 0.21477 | 0.001489 | 30.9782 | 22.4011 | 3.4595 | 0.003686 | 0.417 | 0.3649 | 0.2647 | 0.1284
2012-10-30 22:10 | 26.687 | 22.7728 | 22.7648 | 22.6726 | 22.5483 | 50.1866 | 0.006744 | 0.3265 | 0.002202 | 48.4134 | 0.006933 | 0.21482 | 0.001489 | 30.9844 | 22.3437 | 3.4599 | 0.003691 | 0.4291 | 0.4211 | 0.3289 | 0.2046
2012-10-30 22:10 | 26.6915 | 22.8534 | 22.7773 | 22.6771 | 22.5969 | 50.1871 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006943 | 0.21482 | 0.001491 | 30.9398 | 22.3482 | 3.4599 | 0.003685 | 0.5052 | 0.4291 | 0.3289 | 0.2487
2012-10-30 22:10 | 26.69 | 22.8158 | 22.7357 | 22.7076 | 22.5513 | 50.1779 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.00693 | 0.21487 | 0.001489 | 31.0062 | 22.3748 | 3.4593 | 0.003689 | 0.441 | 0.3609 | 0.3328 | 0.1765
2012-10-30 22:10 | 26.6907 | 22.7685 | 22.7564 | 22.6642 | 22.568 | 50.1848 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3514 | 3.4597 | 0.003688 | 0.4171 | 0.405 | 0.3128 | 0.2166
2012-10-30 22:10 | 26.6915 | 22.8254 | 22.7212 | 22.6891 | 22.5568 | 50.1773 | 0.006735 | 0.32644 | 0.002199 | 48.4695 | 0.006939 | 0.21483 | 0.001491 | 30.9598 | 22.3883 | 3.4592 | 0.00369 | 0.4371 | 0.3329 | 0.3008 | 0.1685
2012-10-30 22:10 | 26.6974 | 22.7912 | 22.7551 | 22.6549 | 22.5747 | 50.1762 | 0.006741 | 0.32641 | 0.0022 | 48.422 | 0.006933 | 0.21483 | 0.001489 | 30.9866 | 22.3983 | 3.4591 | 0.003689 | 0.3929 | 0.3568 | 0.2566 | 0.1764
2012-10-30 22:10 | 26.6952 | 22.821 | 22.7329 | 22.6808 | 22.5364 | 50.1773 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.4001 | 3.4592 | 0.003686 | 0.4209 | 0.3328 | 0.2807 | 0.1363
2012-10-30 22:10 | 26.6915 | 22.7733 | 22.7532 | 22.6731 | 22.5608 | 50.1785 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.3402 | 3.4593 | 0.003688 | 0.4331 | 0.413 | 0.3329 | 0.2206
2012-10-30 22:10 | 26.6937 | 22.8155 | 22.7153 | 22.6552 | 22.5069 | 50.1692 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3665 | 3.4587 | 0.003681 | 0.449 | 0.3488 | 0.2887 | 0.1404
2012-10-30 22:10 | 26.7049 | 22.8307 | 22.7666 | 22.6664 | 22.5582 | 50.1999 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.006936 | 0.21489 | 0.00149 | 30.9815 | 22.4459 | 3.4608 | 0.003685 | 0.3848 | 0.3207 | 0.2205 | 0.1123
2012-10-30 22:10 | 26.6982 | 22.784 | 22.7238 | 22.6437 | 22.5354 | 50.3114 | 0.006732 | 0.32647 | 0.002198 | 48.4954 | 0.006911 | 0.21479 | 0.001484 | 31.0789 | 22.4031 | 3.4685 | 0.003682 | 0.3809 | 0.3207 | 0.2406 | 0.1323
2012-10-30 22:10 | 26.6982 | 22.816 | 22.7238 | 22.6757 | 22.5074 | 50.3547 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.006917 | 0.21484 | 0.001486 | 31.0603 | 22.367 | 3.4715 | 0.003685 | 0.449 | 0.3568 | 0.3087 | 0.1404
2012-10-30 22:10 | 26.7019 | 22.7837 | 22.7436 | 22.6754 | 22.5231 | 50.3692 | 0.006729 | 0.32649 | 0.002197 | 48.5191 | 0.006933 | 0.21479 | 0.001489 | 30.9802 | 22.4308 | 3.4725 | 0.003686 | 0.3529 | 0.3128 | 0.2446 | 0.0923
2012-10-30 22:10 | 26.7012 | 22.8391 | 22.7148 | 22.6306 | 22.5304 | 50.379 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.37 | 3.4731 | 0.003688 | 0.4691 | 0.3448 | 0.2606 | 0.1604
2012-10-30 22:10 | 26.7004 | 22.8222 | 22.722 | 22.698 | 22.4855 | 50.3784 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006933 | 0.21489 | 0.00149 | 30.995 | 22.3451 | 3.4731 | 0.003685 | 0.4771 | 0.3769 | 0.3529 | 0.1404
2012-10-30 22:10 | 26.7004 | 22.8182 | 22.722 | 22.6459 | 22.4975 | 50.3969 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006924 | 0.2148 | 0.001487 | 31.0227 | 22.3571 | 3.4744 | 0.003683 | 0.4611 | 0.3649 | 0.2888 | 0.1404
2012-10-30 22:11 | 26.7004 | 22.8142 | 22.694 | 22.6739 | 22.5055 | 50.4125 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.006943 | 0.21484 | 0.001492 | 30.944 | 22.3732 | 3.4754 | 0.003687 | 0.441 | 0.3208 | 0.3007 | 0.1323
2012-10-30 22:11 | 26.6959 | 22.8258 | 22.7296 | 22.6654 | 22.5051 | 50.401 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3607 | 3.4746 | 0.003681 | 0.4651 | 0.3689 | 0.3047 | 0.1444
2012-10-30 22:11 | 26.6915 | 22.8013 | 22.7532 | 22.657 | 22.5047 | 50.4276 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.00693 | 0.21486 | 0.001489 | 31.0042 | 22.3322 | 3.4765 | 0.003686 | 0.4691 | 0.421 | 0.3248 | 0.1725
2012-10-30 22:11 | 26.7034 | 22.7771 | 22.689 | 22.697 | 22.5446 | 50.4218 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.006933 | 0.21484 | 0.00149 | 30.9886 | 22.3441 | 3.4761 | 0.003689 | 0.433 | 0.3449 | 0.3529 | 0.2005
2012-10-30 22:11 | 26.7012 | 22.799 | 22.7028 | 22.6507 | 22.4943 | 50.438 | 0.006741 | 0.32643 | 0.0022 | 48.4241 | 0.00693 | 0.21477 | 0.001488 | 30.9916 | 22.3619 | 3.4772 | 0.003688 | 0.4371 | 0.3409 | 0.2888 | 0.1324
2012-10-30 22:11 | 26.7079 | 22.8017 | 22.7135 | 22.6293 | 22.5572 | 50.4403 | 0.006713 | 0.32644 | 0.002191 | 48.6283 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3727 | 3.4774 | 0.003681 | 0.429 | 0.3408 | 0.2566 | 0.1845
2012-10-30 22:11 | 26.7162 | 22.7338 | 22.7459 | 22.6416 | 22.5534 | 50.427 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.401 | 3.4764 | 0.003689 | 0.3328 | 0.3449 | 0.2406 | 0.1524
2012-10-30 22:11 | 26.7087 | 22.7905 | 22.6943 | 22.6742 | 22.5219 | 50.4403 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3735 | 3.4774 | 0.003686 | 0.417 | 0.3208 | 0.3007 | 0.1484
2012-10-30 22:11 | 26.7154 | 22.7811 | 22.701 | 22.6288 | 22.5286 | 50.4293 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.006936 | 0.2148 | 0.00149 | 30.969 | 22.3601 | 3.4766 | 0.003684 | 0.421 | 0.3409 | 0.2687 | 0.1685
2012-10-30 22:11 | 26.7072 | 22.817 | 22.7288 | 22.6487 | 22.4923 | 50.4253 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.006924 | 0.21473 | 0.001487 | 31.0123 | 22.384 | 3.4763 | 0.003683 | 0.433 | 0.3448 | 0.2647 | 0.1083
2012-10-30 22:11 | 26.7094 | 22.7471 | 22.7431 | 22.6348 | 22.5226 | 50.4305 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.006933 | 0.21474 | 0.001489 | 30.974 | 22.3541 | 3.4767 | 0.003682 | 0.393 | 0.389 | 0.2807 | 0.1685
2012-10-30 22:11 | 26.7184 | 22.7561 | 22.724 | 22.6559 | 22.5436 | 50.4345 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.006949 | 0.21474 | 0.001492 | 30.9027 | 22.3711 | 3.477 | 0.003685 | 0.385 | 0.3529 | 0.2848 | 0.1725
2012-10-30 22:11 | 26.7124 | 22.8182 | 22.706 | 22.6298 | 22.4975 | 50.431 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3731 | 3.4767 | 0.003686 | 0.4451 | 0.3329 | 0.2567 | 0.1244
2012-10-30 22:11 | 26.7064 | 22.7922 | 22.6839 | 22.6719 | 22.5156 | 50.4426 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.006939 | 0.2148 | 0.001491 | 30.9556 | 22.3631 | 3.4775 | 0.003688 | 0.4291 | 0.3208 | 0.3088 | 0.1525
2012-10-30 22:11 | 26.7117 | 22.7774 | 22.7013 | 22.6411 | 22.5128 | 50.4357 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.3684 | 3.477 | 0.003682 | 0.409 | 0.3329 | 0.2727 | 0.1444
2012-10-30 22:11 | 26.7109 | 22.7686 | 22.7486 | 22.6283 | 22.5321 | 50.4281 | 0.006751 | 0.32637 | 0.002203 | 48.3439 | 0.00693 | 0.21482 | 0.001489 | 30.9978 | 22.3556 | 3.4765 | 0.003692 | 0.413 | 0.393 | 0.2727 | 0.1765
2012-10-30 22:11 | 26.7109 | 22.7606 | 22.7245 | 22.6203 | 22.512 | 50.4345 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.006946 | 0.21476 | 0.001492 | 30.9182 | 22.3556 | 3.477 | 0.003689 | 0.405 | 0.3689 | 0.2647 | 0.1564
2012-10-30 22:11 | 26.7102 | 22.7358 | 22.7038 | 22.6637 | 22.5474 | 50.4357 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.006927 | 0.21482 | 0.001488 | 31.0113 | 22.3509 | 3.477 | 0.003685 | 0.3849 | 0.3529 | 0.3128 | 0.1965
2012-10-30 22:11 | 26.7034 | 22.7852 | 22.697 | 22.6409 | 22.5126 | 50.4368 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.006936 | 0.21483 | 0.00149 | 30.9732 | 22.312 | 3.4771 | 0.003687 | 0.4732 | 0.385 | 0.3289 | 0.2006
2012-10-30 22:11 | 26.7057 | 22.7434 | 22.6993 | 22.6111 | 22.5149 | 50.4253 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.006917 | 0.21487 | 0.001486 | 31.0645 | 22.3424 | 3.4763 | 0.003684 | 0.401 | 0.3569 | 0.2687 | 0.1725
2012-10-30 22:11 | 26.7019 | 22.7797 | 22.6875 | 22.6073 | 22.5151 | 50.4299 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.006933 | 0.2148 | 0.001489 | 30.9824 | 22.3185 | 3.4766 | 0.003685 | 0.4612 | 0.369 | 0.2888 | 0.1966
2012-10-30 22:11 | 26.7049 | 22.7386 | 22.7145 | 22.5942 | 22.51 | 50.4339 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.00693 | 0.2148 | 0.001489 | 30.9958 | 22.3616 | 3.4769 | 0.003688 | 0.377 | 0.3529 | 0.2326 | 0.1484
2012-10-30 22:11 | 26.7019 | 22.7797 | 22.7195 | 22.6354 | 22.499 | 50.4195 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.00693 | 0.21479 | 0.001488 | 30.9937 | 22.3306 | 3.4759 | 0.003683 | 0.4491 | 0.3889 | 0.3048 | 0.1684
2012-10-30 22:11 | 26.7102 | 22.796 | 22.6797 | 22.6236 | 22.4833 | 50.4333 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.006927 | 0.21483 | 0.001488 | 31.0134 | 22.3108 | 3.4769 | 0.003683 | 0.4852 | 0.3689 | 0.3128 | 0.1725
2012-10-30 22:11 | 26.7034 | 22.7531 | 22.693 | 22.6328 | 22.4724 | 50.4391 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.006939 | 0.21484 | 0.001491 | 30.9618 | 22.328 | 3.4773 | 0.003688 | 0.4251 | 0.365 | 0.3048 | 0.1444
2012-10-30 22:12 | 26.7049 | 22.6865 | 22.7065 | 22.6143 | 22.482 | 50.4253 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.006946 | 0.21483 | 0.001492 | 30.9286 | 22.3255 | 3.4763 | 0.003686 | 0.361 | 0.381 | 0.2888 | 0.1565
2012-10-30 22:12 | 26.7012 | 22.7669 | 22.6787 | 22.6066 | 22.5023 | 50.427 | 0.006741 | 0.32647 | 0.002201 | 48.4306 | 0.00693 | 0.21487 | 0.001489 | 31.0062 | 22.3258 | 3.4764 | 0.00369 | 0.4411 | 0.3529 | 0.2808 | 0.1765
2012-10-30 22:12 | 26.7094 | 22.7751 | 22.727 | 22.6028 | 22.5065 | 50.4362 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.05403 | 0.96596 | 0.052191 | 17.8782 | 22.322 | 3.4771 | 0.054391 | 0.4531 | 0.405 | 0.2808 | 0.1845
2012-10-30 22:12 | 26.7094 | 22.7591 | 22.6589 | 22.6108 | 22.5266 | 50.4461 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.048906 | 0.88446 | 0.043255 | 18.0849 | 22.3381 | 3.4778 | 0.045449 | 0.421 | 0.3208 | 0.2727 | 0.1885
2012-10-30 22:12 | 26.7132 | 22.7749 | 22.7028 | 22.6266 | 22.4822 | 50.4328 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.048695 | 0.88112 | 0.042906 | 18.0947 | 22.3699 | 3.4768 | 0.045106 | 0.405 | 0.3329 | 0.2567 | 0.1123
2012-10-30 22:12 | 26.7034 | 22.7611 | 22.6809 | 22.6128 | 22.4765 | 50.4247 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048648 | 0.88017 | 0.042818 | 18.0926 | 22.324 | 3.4763 | 0.045012 | 0.4371 | 0.3569 | 0.2888 | 0.1525
2012-10-30 22:12 | 26.7102 | 22.7318 | 22.6877 | 22.6196 | 22.4993 | 50.4316 | 0.00671 | 0.3264 | 0.00219 | 48.6437 | 0.048695 | 0.88095 | 0.042898 | 18.0911 | 22.375 | 3.4768 | 0.045088 | 0.3568 | 0.3127 | 0.2446 | 0.1243
2012-10-30 22:12 | 26.7087 | 22.7504 | 22.6862 | 22.6462 | 22.4978 | 50.4206 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048629 | 0.87986 | 0.042787 | 18.0934 | 22.3815 | 3.476 | 0.044981 | 0.3689 | 0.3047 | 0.2647 | 0.1163
2012-10-30 22:12 | 26.7057 | 22.7714 | 22.6913 | 22.6311 | 22.5108 | 50.4229 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.048636 | 0.8803 | 0.042814 | 18.0997 | 22.3624 | 3.4762 | 0.045006 | 0.409 | 0.3289 | 0.2687 | 0.1484
2012-10-30 22:12 | 26.7027 | 22.7484 | 22.6802 | 22.6402 | 22.4557 | 50.427 | 0.006744 | 0.32637 | 0.002201 | 48.3941 | 0.048645 | 0.88023 | 0.042819 | 18.0949 | 22.3594 | 3.4764 | 0.04502 | 0.389 | 0.3208 | 0.2808 | 0.0963
2012-10-30 22:12 | 26.7094 | 22.7872 | 22.6749 | 22.6268 | 22.5346 | 50.4397 | 0.006703 | 0.32625 | 0.002187 | 48.6728 | 0.048617 | 0.87963 | 0.042765 | 18.0931 | 22.3541 | 3.4773 | 0.044952 | 0.4331 | 0.3208 | 0.2727 | 0.1805
2012-10-30 22:12 | 26.6989 | 22.7606 | 22.6684 | 22.5963 | 22.492 | 50.4333 | 0.006722 | 0.32628 | 0.002193 | 48.5396 | 0.048645 | 0.87991 | 0.042803 | 18.0884 | 22.3677 | 3.4769 | 0.044996 | 0.3929 | 0.3007 | 0.2286 | 0.1243
2012-10-30 22:12 | 26.7094 | 22.7471 | 22.691 | 22.6549 | 22.5266 | 50.4189 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.048645 | 0.88015 | 0.042815 | 18.0934 | 22.3541 | 3.4759 | 0.045014 | 0.393 | 0.3369 | 0.3008 | 0.1725
2012-10-30 22:12 | 26.7064 | 22.7842 | 22.6679 | 22.6118 | 22.4995 | 50.4218 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048639 | 0.88017 | 0.04281 | 18.0959 | 22.315 | 3.4761 | 0.045004 | 0.4692 | 0.3529 | 0.2968 | 0.1845
2012-10-30 22:12 | 26.7004 | 22.7741 | 22.6779 | 22.6178 | 22.4935 | 50.4125 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048607 | 0.87969 | 0.042759 | 18.098 | 22.321 | 3.4754 | 0.044953 | 0.4531 | 0.3569 | 0.2968 | 0.1725
2012-10-30 22:12 | 26.6974 | 22.7551 | 22.6509 | 22.6268 | 22.4945 | 50.4189 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048661 | 0.8802 | 0.042831 | 18.0883 | 22.2819 | 3.4759 | 0.04503 | 0.4732 | 0.369 | 0.3449 | 0.2126
2012-10-30 22:12 | 26.6952 | 22.7569 | 22.6527 | 22.6166 | 22.4963 | 50.4229 | 0.006744 | 0.32636 | 0.002201 | 48.3919 | 0.048661 | 0.8804 | 0.042841 | 18.0925 | 22.2918 | 3.4762 | 0.045042 | 0.4651 | 0.3609 | 0.3248 | 0.2045
2012-10-30 22:12 | 26.693 | 22.7627 | 22.6184 | 22.5984 | 22.4821 | 50.4235 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048629 | 0.87994 | 0.04279 | 18.0949 | 22.2735 | 3.4762 | 0.044983 | 0.4892 | 0.3449 | 0.3249 | 0.2086
2012-10-30 22:12 | 26.6974 | 22.7752 | 22.683 | 22.6068 | 22.4705 | 50.4264 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048683 | 0.88069 | 0.042875 | 18.0903 | 22.2699 | 3.4764 | 0.045071 | 0.5053 | 0.4131 | 0.3369 | 0.2006
2012-10-30 22:12 | 26.6915 | 22.7813 | 22.6811 | 22.6049 | 22.4806 | 50.4287 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048604 | 0.87968 | 0.042756 | 18.0989 | 22.3041 | 3.4766 | 0.044955 | 0.4772 | 0.377 | 0.3008 | 0.1765
2012-10-30 22:12 | 26.7034 | 22.7651 | 22.6489 | 22.5968 | 22.4323 | 50.4143 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048683 | 0.88075 | 0.042877 | 18.0914 | 22.308 | 3.4756 | 0.045071 | 0.4571 | 0.3409 | 0.2888 | 0.1243
2012-10-30 22:13 | 26.6922 | 22.7539 | 22.6617 | 22.5976 | 22.4773 | 50.4166 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048639 | 0.88004 | 0.042804 | 18.0933 | 22.2647 | 3.4757 | 0.045001 | 0.4892 | 0.397 | 0.3329 | 0.2126
2012-10-30 22:13 | 26.687 | 22.7648 | 22.6365 | 22.6084 | 22.452 | 50.4079 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048645 | 0.88021 | 0.042818 | 18.0946 | 22.2394 | 3.4751 | 0.045013 | 0.5254 | 0.3971 | 0.369 | 0.2126
2012-10-30 22:13 | 26.693 | 22.7828 | 22.6425 | 22.5743 | 22.458 | 50.4218 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048648 | 0.88046 | 0.042832 | 18.0985 | 22.1852 | 3.4761 | 0.045029 | 0.5976 | 0.4573 | 0.3891 | 0.2728
2012-10-30 22:13 | 26.6885 | 22.7342 | 22.658 | 22.5939 | 22.4335 | 50.4183 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.048661 | 0.88028 | 0.042835 | 18.0901 | 22.1888 | 3.4758 | 0.045033 | 0.5454 | 0.4692 | 0.4051 | 0.2447
2012-10-30 22:13 | 26.6937 | 22.7354 | 22.6512 | 22.5991 | 22.4587 | 50.4172 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048639 | 0.88007 | 0.042806 | 18.0939 | 22.2582 | 3.4758 | 0.045002 | 0.4772 | 0.393 | 0.3409 | 0.2005
2012-10-30 22:13 | 26.6825 | 22.7643 | 22.6961 | 22.6079 | 22.4355 | 50.4235 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048673 | 0.88044 | 0.042854 | 18.0889 | 22.1787 | 3.4762 | 0.045049 | 0.5856 | 0.5174 | 0.4292 | 0.2568
2012-10-30 22:13 | 26.6892 | 22.7549 | 22.6667 | 22.5986 | 22.4542 | 50.4229 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048648 | 0.88014 | 0.042817 | 18.092 | 22.2095 | 3.4762 | 0.045009 | 0.5454 | 0.4572 | 0.3891 | 0.2447
2012-10-30 22:13 | 26.6892 | 22.7469 | 22.6788 | 22.6186 | 22.4623 | 50.4206 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048683 | 0.88062 | 0.042871 | 18.0888 | 22.2095 | 3.476 | 0.045068 | 0.5374 | 0.4693 | 0.4091 | 0.2528
2012-10-30 22:13 | 26.6847 | 22.7705 | 22.6663 | 22.5781 | 22.4738 | 50.4137 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048667 | 0.8805 | 0.042851 | 18.0923 | 22.2492 | 3.4755 | 0.045048 | 0.5213 | 0.4171 | 0.3289 | 0.2246
2012-10-30 22:13 | 26.6922 | 22.7539 | 22.6577 | 22.5695 | 22.4653 | 50.4189 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048661 | 0.88034 | 0.042838 | 18.0913 | 22.2567 | 3.4759 | 0.045033 | 0.4972 | 0.401 | 0.3128 | 0.2086
2012-10-30 22:13 | 26.6892 | 22.7509 | 22.6748 | 22.5866 | 22.4623 | 50.4039 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048673 | 0.88046 | 0.042854 | 18.0892 | 22.2456 | 3.4748 | 0.045052 | 0.5053 | 0.4292 | 0.341 | 0.2167
2012-10-30 22:13 | 26.6847 | 22.7504 | 22.6582 | 22.6101 | 22.4818 | 50.401 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048642 | 0.88004 | 0.042807 | 18.0921 | 22.2371 | 3.4746 | 0.045005 | 0.5133 | 0.4211 | 0.373 | 0.2447
2012-10-30 22:13 | 26.6757 | 22.7454 | 22.6933 | 22.6172 | 22.4808 | 50.4085 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048673 | 0.88072 | 0.042867 | 18.0946 | 22.2642 | 3.4752 | 0.045066 | 0.4812 | 0.4291 | 0.353 | 0.2166
2012-10-30 22:13 | 26.6817 | 22.7354 | 22.6673 | 22.6111 | 22.4668 | 50.4044 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048642 | 0.88014 | 0.042812 | 18.0942 | 22.2502 | 3.4749 | 0.045008 | 0.4852 | 0.4171 | 0.3609 | 0.2166
2012-10-30 22:13 | 26.6795 | 22.7372 | 22.7011 | 22.5929 | 22.5047 | 50.4218 | 0.006741 | 0.32636 | 0.0022 | 48.4134 | 0.048639 | 0.88002 | 0.042803 | 18.0929 | 22.244 | 3.4761 | 0.045003 | 0.4932 | 0.4571 | 0.3489 | 0.2607
2012-10-30 22:13 | 26.693 | 22.7788 | 22.6906 | 22.5944 | 22.454 | 50.4085 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048686 | 0.88075 | 0.04288 | 18.0903 | 22.2775 | 3.4752 | 0.045074 | 0.5013 | 0.4131 | 0.3169 | 0.1765
2012-10-30 22:13 | 26.6877 | 22.7735 | 22.7174 | 22.5851 | 22.4728 | 50.412 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048642 | 0.88001 | 0.042805 | 18.0915 | 22.2802 | 3.4754 | 0.045003 | 0.4933 | 0.4372 | 0.3049 | 0.1926
2012-10-30 22:13 | 26.6885 | 22.7342 | 22.6821 | 22.5778 | 22.4455 | 50.4044 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.04867 | 0.88038 | 0.042848 | 18.0889 | 22.281 | 3.4749 | 0.045043 | 0.4532 | 0.4011 | 0.2968 | 0.1645
2012-10-30 22:13 | 26.6832 | 22.7409 | 22.6808 | 22.5966 | 22.4522 | 50.4108 | 0.006741 | 0.32634 | 0.0022 | 48.4114 | 0.048648 | 0.88024 | 0.042822 | 18.0941 | 22.3279 | 3.4753 | 0.045022 | 0.413 | 0.3529 | 0.2687 | 0.1243
2012-10-30 22:13 | 26.6735 | 22.7352 | 22.6991 | 22.5749 | 22.4626 | 50.4177 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048642 | 0.87991 | 0.0428 | 18.0895 | 22.246 | 3.4758 | 0.044997 | 0.4892 | 0.4531 | 0.3289 | 0.2166
2012-10-30 22:14 | 26.6847 | 22.7184 | 22.6903 | 22.5861 | 22.4778 | 50.416 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.048642 | 0.88034 | 0.042822 | 18.0984 | 22.2933 | 3.4757 | 0.045019 | 0.4251 | 0.397 | 0.2928 | 0.1845
2012-10-30 22:14 | 26.6817 | 22.7635 | 22.6793 | 22.5831 | 22.5029 | 50.405 | 0.006722 | 0.32647 | 0.002195 | 48.5675 | 0.048651 | 0.88028 | 0.042827 | 18.0938 | 22.3023 | 3.4749 | 0.045022 | 0.4612 | 0.377 | 0.2808 | 0.2006
2012-10-30 22:14 | 26.678 | 22.7397 | 22.6676 | 22.6395 | 22.4912 | 50.4021 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.048664 | 0.8804 | 0.042844 | 18.0914 | 22.2705 | 3.4747 | 0.045037 | 0.4692 | 0.3971 | 0.369 | 0.2207
2012-10-30 22:14 | 26.6802 | 22.7299 | 22.6457 | 22.5896 | 22.4894 | 50.4125 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048645 | 0.88005 | 0.04281 | 18.0913 | 22.2768 | 3.4754 | 0.045006 | 0.4531 | 0.3689 | 0.3128 | 0.2126
2012-10-30 22:14 | 26.6787 | 22.7444 | 22.7164 | 22.5921 | 22.4999 | 50.4137 | 0.00671 | 0.32638 | 0.00219 | 48.6414 | 0.048661 | 0.8803 | 0.042836 | 18.0904 | 22.2712 | 3.4755 | 0.045026 | 0.4732 | 0.4452 | 0.3209 | 0.2287
2012-10-30 22:14 | 26.6765 | 22.7543 | 22.6621 | 22.618 | 22.4776 | 50.4224 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.048661 | 0.88044 | 0.042843 | 18.0934 | 22.253 | 3.4761 | 0.045038 | 0.5013 | 0.4091 | 0.365 | 0.2246
2012-10-30 22:14 | 26.675 | 22.7447 | 22.6646 | 22.6125 | 22.4721 | 50.4166 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048667 | 0.88049 | 0.042851 | 18.0921 | 22.2595 | 3.4757 | 0.045048 | 0.4852 | 0.4051 | 0.353 | 0.2126
2012-10-30 22:14 | 26.681 | 22.7307 | 22.6946 | 22.6024 | 22.4581 | 50.412 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.048661 | 0.88052 | 0.042847 | 18.0949 | 22.3177 | 3.4754 | 0.045045 | 0.413 | 0.3769 | 0.2847 | 0.1404
2012-10-30 22:14 | 26.6705 | 22.7603 | 22.648 | 22.604 | 22.5037 | 50.4143 | 0.006744 | 0.32638 | 0.002201 | 48.3962 | 0.048639 | 0.8801 | 0.042807 | 18.0945 | 22.263 | 3.4756 | 0.045008 | 0.4973 | 0.385 | 0.341 | 0.2407
2012-10-30 22:14 | 26.6638 | 22.7456 | 22.6694 | 22.5732 | 22.489 | 50.4206 | 0.006729 | 0.3265 | 0.002197 | 48.5213 | 0.048667 | 0.88033 | 0.042843 | 18.0888 | 22.2443 | 3.476 | 0.04504 | 0.5013 | 0.4251 | 0.3289 | 0.2447
2012-10-30 22:14 | 26.6638 | 22.7375 | 22.6534 | 22.5932 | 22.4609 | 50.4328 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.048661 | 0.88049 | 0.042845 | 18.0943 | 22.2363 | 3.4768 | 0.045039 | 0.5012 | 0.4171 | 0.3569 | 0.2246
2012-10-30 22:14 | 26.6705 | 22.7162 | 22.6681 | 22.5919 | 22.4756 | 50.4062 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048711 | 0.88111 | 0.04292 | 18.0885 | 22.2189 | 3.475 | 0.045118 | 0.4973 | 0.4492 | 0.373 | 0.2567
2012-10-30 22:14 | 26.6683 | 22.73 | 22.6499 | 22.6258 | 22.4253 | 50.4102 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048642 | 0.87992 | 0.042801 | 18.0898 | 22.2287 | 3.4753 | 0.044994 | 0.5013 | 0.4212 | 0.3971 | 0.1966
2012-10-30 22:14 | 26.6623 | 22.736 | 22.692 | 22.6078 | 22.3952 | 50.4102 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.048711 | 0.88145 | 0.042937 | 18.0956 | 22.2428 | 3.4753 | 0.045135 | 0.4932 | 0.4492 | 0.365 | 0.1524
2012-10-30 22:14 | 26.6712 | 22.749 | 22.6768 | 22.6167 | 22.4643 | 50.4102 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.048655 | 0.88024 | 0.042828 | 18.0915 | 22.2236 | 3.4753 | 0.045023 | 0.5254 | 0.4532 | 0.3931 | 0.2407
2012-10-30 22:14 | 26.6638 | 22.7456 | 22.6814 | 22.5973 | 22.4168 | 50.4073 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.04867 | 0.8805 | 0.042854 | 18.0912 | 22.2563 | 3.4751 | 0.045048 | 0.4893 | 0.4251 | 0.341 | 0.1605
2012-10-30 22:14 | 26.666 | 22.7397 | 22.6716 | 22.6075 | 22.4711 | 50.4131 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048717 | 0.88111 | 0.042925 | 18.0862 | 22.2425 | 3.4755 | 0.04512 | 0.4972 | 0.4291 | 0.365 | 0.2286
2012-10-30 22:14 | 26.6593 | 22.709 | 22.6849 | 22.5847 | 22.4484 | 50.405 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.048617 | 0.87997 | 0.042781 | 18.1 | 22.2719 | 3.4749 | 0.044979 | 0.4371 | 0.413 | 0.3128 | 0.1765
2012-10-30 22:14 | 26.657 | 22.7668 | 22.6947 | 22.6025 | 22.4621 | 50.4068 | 0.006741 | 0.32637 | 0.0022 | 48.4157 | 0.048673 | 0.88043 | 0.042853 | 18.0886 | 22.2255 | 3.475 | 0.045053 | 0.5413 | 0.4692 | 0.377 | 0.2366
2012-10-30 22:14 | 26.657 | 22.7388 | 22.6706 | 22.5864 | 22.4341 | 50.401 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.048623 | 0.87997 | 0.042787 | 18.0977 | 22.2295 | 3.4746 | 0.044985 | 0.5093 | 0.4411 | 0.3569 | 0.2046
2012-10-30 22:14 | 26.6555 | 22.7212 | 22.6972 | 22.609 | 22.4366 | 50.4108 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.048655 | 0.88018 | 0.042825 | 18.0903 | 22.24 | 3.4753 | 0.045024 | 0.4812 | 0.4572 | 0.369 | 0.1966
2012-10-30 22:14 | 26.648 | 22.7418 | 22.6857 | 22.5574 | 22.4411 | 50.4102 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.048633 | 0.88004 | 0.042799 | 18.0955 | 22.2125 | 3.4753 | 0.044997 | 0.5293 | 0.4732 | 0.3449 | 0.2286
2012-10-30 22:14 | 26.6518 | 22.7255 | 22.7095 | 22.5772 | 22.477 | 50.4004 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.048655 | 0.88028 | 0.04283 | 18.0923 | 22.2564 | 3.4746 | 0.045027 | 0.4691 | 0.4531 | 0.3208 | 0.2206
2012-10-30 22:14 | 26.648 | 22.7097 | 22.6656 | 22.6135 | 22.4531 | 50.412 | 0.006735 | 0.32649 | 0.002199 | 48.4759 | 0.048645 | 0.88004 | 0.042809 | 18.091 | 22.2125 | 3.4754 | 0.045008 | 0.4972 | 0.4531 | 0.401 | 0.2406
2012-10-30 22:14 | 26.6391 | 22.7249 | 22.6527 | 22.6166 | 22.4483 | 50.4039 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048673 | 0.88067 | 0.042865 | 18.0937 | 22.2557 | 3.4748 | 0.04506 | 0.4692 | 0.397 | 0.3609 | 0.1926
2012-10-30 22:14 | 26.6495 | 22.7433 | 22.7232 | 22.615 | 22.4827 | 50.401 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048642 | 0.88014 | 0.042812 | 18.0942 | 22.2862 | 3.4746 | 0.045008 | 0.4571 | 0.437 | 0.3288 | 0.1965
2012-10-30 22:15 | 26.645 | 22.7789 | 22.6987 | 22.6025 | 22.5143 | 50.4102 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048699 | 0.88111 | 0.042909 | 18.0929 | 22.2456 | 3.4753 | 0.045106 | 0.5333 | 0.4531 | 0.3569 | 0.2687
2012-10-30 22:15 | 26.6413 | 22.7351 | 22.711 | 22.5707 | 22.4585 | 50.4044 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048645 | 0.88012 | 0.042814 | 18.0928 | 22.2499 | 3.4749 | 0.045007 | 0.4852 | 0.4611 | 0.3208 | 0.2086
2012-10-30 22:15 | 26.6488 | 22.7145 | 22.7506 | 22.5983 | 22.4379 | 50.4131 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.04868 | 0.88069 | 0.042872 | 18.0914 | 22.2935 | 3.4755 | 0.045072 | 0.421 | 0.4571 | 0.3048 | 0.1444
2012-10-30 22:15 | 26.6421 | 22.7559 | 22.6758 | 22.5836 | 22.4593 | 50.4033 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048658 | 0.88011 | 0.042824 | 18.0877 | 22.2868 | 3.4748 | 0.045022 | 0.4691 | 0.389 | 0.2968 | 0.1725
2012-10-30 22:15 | 26.6488 | 22.7706 | 22.7145 | 22.6143 | 22.486 | 50.4033 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048724 | 0.88132 | 0.042942 | 18.0881 | 22.2855 | 3.4748 | 0.045141 | 0.4851 | 0.429 | 0.3288 | 0.2005
2012-10-30 22:15 | 26.6413 | 22.7231 | 22.7271 | 22.5868 | 22.4625 | 50.4039 | 0.006744 | 0.32636 | 0.002201 | 48.3919 | 0.048645 | 0.88012 | 0.042814 | 18.0928 | 22.278 | 3.4748 | 0.045015 | 0.4451 | 0.4491 | 0.3088 | 0.1845
2012-10-30 22:15 | 26.6353 | 22.7211 | 22.7291 | 22.5848 | 22.4324 | 50.4079 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048686 | 0.88086 | 0.042886 | 18.0927 | 22.3281 | 3.4751 | 0.045081 | 0.393 | 0.401 | 0.2567 | 0.1043
2012-10-30 22:15 | 26.6368 | 22.6985 | 22.6785 | 22.5943 | 22.446 | 50.4004 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048639 | 0.88005 | 0.042805 | 18.0935 | 22.2935 | 3.4746 | 0.045002 | 0.405 | 0.385 | 0.3008 | 0.1525
2012-10-30 22:15 | 26.6376 | 22.7033 | 22.6913 | 22.567 | 22.5149 | 50.3952 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048677 | 0.88073 | 0.042871 | 18.0934 | 22.2783 | 3.4742 | 0.045069 | 0.425 | 0.413 | 0.2887 | 0.2366
2012-10-30 22:15 | 26.6413 | 22.675 | 22.7391 | 22.5948 | 22.4745 | 50.4027 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048651 | 0.88017 | 0.042821 | 18.0915 | 22.29 | 3.4748 | 0.045019 | 0.385 | 0.4491 | 0.3048 | 0.1845
2012-10-30 22:15 | 26.6376 | 22.7354 | 22.6512 | 22.539 | 22.4668 | 50.4102 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.048655 | 0.88007 | 0.04282 | 18.0879 | 22.2903 | 3.4753 | 0.045019 | 0.4451 | 0.3609 | 0.2487 | 0.1765
2012-10-30 22:15 | 26.6428 | 22.7085 | 22.6925 | 22.5682 | 22.4479 | 50.4097 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.048651 | 0.88021 | 0.042823 | 18.0924 | 22.2554 | 3.4752 | 0.045023 | 0.4531 | 0.4371 | 0.3128 | 0.1925
2012-10-30 22:15 | 26.6406 | 22.6983 | 22.6983 | 22.5861 | 22.4578 | 50.4004 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048661 | 0.88027 | 0.042835 | 18.0898 | 22.2973 | 3.4746 | 0.045033 | 0.401 | 0.401 | 0.2888 | 0.1605
2012-10-30 22:15 | 26.6421 | 22.6838 | 22.6557 | 22.6076 | 22.4793 | 50.4091 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048765 | 0.88226 | 0.043024 | 18.0921 | 22.3149 | 3.4752 | 0.045219 | 0.3689 | 0.3408 | 0.2927 | 0.1644
2012-10-30 22:15 | 26.6338 | 22.7156 | 22.6915 | 22.5913 | 22.459 | 50.4044 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048639 | 0.88011 | 0.042808 | 18.0947 | 22.2745 | 3.4749 | 0.045003 | 0.4411 | 0.417 | 0.3168 | 0.1845
2012-10-30 22:15 | 26.6331 | 22.7149 | 22.6948 | 22.5746 | 22.4743 | 50.4027 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048658 | 0.8803 | 0.042834 | 18.0915 | 22.2738 | 3.4748 | 0.045028 | 0.4411 | 0.421 | 0.3008 | 0.2005
2012-10-30 22:15 | 26.6383 | 22.7201 | 22.7521 | 22.5838 | 22.4755 | 50.4097 | 0.00671 | 0.32634 | 0.00219 | 48.635 | 0.048683 | 0.88092 | 0.042886 | 18.095 | 22.275 | 3.4752 | 0.045076 | 0.4451 | 0.4771 | 0.3088 | 0.2005
2012-10-30 22:15 | 26.6413 | 22.703 | 22.683 | 22.6028 | 22.4825 | 50.3929 | 0.006741 | 0.32641 | 0.0022 | 48.422 | 0.048651 | 0.88002 | 0.042814 | 18.0885 | 22.3141 | 3.4741 | 0.045014 | 0.3889 | 0.3689 | 0.2887 | 0.1684
2012-10-30 22:15 | 26.6421 | 22.6998 | 22.6758 | 22.6156 | 22.4673 | 50.4206 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048661 | 0.88018 | 0.042831 | 18.088 | 22.2988 | 3.476 | 0.045029 | 0.401 | 0.377 | 0.3168 | 0.1685
2012-10-30 22:15 | 26.6406 | 22.6903 | 22.6703 | 22.6181 | 22.4417 | 50.3935 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048692 | 0.88096 | 0.042896 | 18.0925 | 22.3174 | 3.4741 | 0.045094 | 0.3729 | 0.3529 | 0.3007 | 0.1243
2012-10-30 22:15 | 26.6301 | 22.7319 | 22.6918 | 22.6157 | 22.4473 | 50.3865 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048645 | 0.88026 | 0.04282 | 18.0955 | 22.2828 | 3.4736 | 0.045016 | 0.4491 | 0.409 | 0.3329 | 0.1645
2012-10-30 22:15 | 26.6263 | 22.696 | 22.7 | 22.5678 | 22.4475 | 50.3952 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048648 | 0.88021 | 0.042821 | 18.0935 | 22.3071 | 3.4742 | 0.045014 | 0.3889 | 0.3929 | 0.2607 | 0.1404
2012-10-30 22:15 | 26.6271 | 22.6768 | 22.6808 | 22.5846 | 22.4483 | 50.3946 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.048655 | 0.8803 | 0.042831 | 18.0927 | 22.2878 | 3.4742 | 0.045028 | 0.389 | 0.393 | 0.2968 | 0.1605
2012-10-30 22:15 | 26.6233 | 22.7091 | 22.697 | 22.6008 | 22.4164 | 50.3969 | 0.006713 | 0.32643 | 0.002191 | 48.6261 | 0.048683 | 0.88089 | 0.042884 | 18.0944 | 22.2881 | 3.4744 | 0.045075 | 0.421 | 0.4089 | 0.3127 | 0.1283
2012-10-30 22:15 | 26.6278 | 22.6855 | 22.7015 | 22.5653 | 22.445 | 50.3952 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048629 | 0.88004 | 0.042795 | 18.097 | 22.3006 | 3.4742 | 0.044991 | 0.3849 | 0.4009 | 0.2647 | 0.1444
2012-10-30 22:15 | 26.6278 | 22.6775 | 22.6655 | 22.5572 | 22.469 | 50.3975 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048702 | 0.88092 | 0.042903 | 18.0879 | 22.3126 | 3.4744 | 0.045101 | 0.3649 | 0.3529 | 0.2446 | 0.1564
2012-10-30 22:15 | 26.6174 | 22.6631 | 22.6711 | 22.5428 | 22.4426 | 50.3964 | 0.006732 | 0.32646 | 0.002198 | 48.4932 | 0.048639 | 0.88023 | 0.042813 | 18.0971 | 22.2822 | 3.4743 | 0.045011 | 0.3809 | 0.3889 | 0.2606 | 0.1604
2012-10-30 22:16 | 26.6256 | 22.7034 | 22.6913 | 22.5631 | 22.4628 | 50.394 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048642 | 0.88008 | 0.042809 | 18.093 | 22.2823 | 3.4742 | 0.045006 | 0.4211 | 0.409 | 0.2808 | 0.1805
2012-10-30 22:16 | 26.6196 | 22.6533 | 22.6412 | 22.6332 | 22.4648 | 50.3917 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.048655 | 0.88033 | 0.042832 | 18.0932 | 22.2523 | 3.474 | 0.045027 | 0.401 | 0.3889 | 0.3809 | 0.2125
2012-10-30 22:16 | 26.6226 | 22.7004 | 22.6803 | 22.5601 | 22.4598 | 50.3981 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048645 | 0.88036 | 0.042825 | 18.0976 | 22.2513 | 3.4744 | 0.045022 | 0.4491 | 0.429 | 0.3088 | 0.2085
2012-10-30 22:16 | 26.6233 | 22.7011 | 22.6289 | 22.5648 | 22.4645 | 50.3923 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.04867 | 0.88069 | 0.042863 | 18.0951 | 22.272 | 3.474 | 0.045056 | 0.4291 | 0.3569 | 0.2928 | 0.1925
2012-10-30 22:16 | 26.6151 | 22.7089 | 22.6528 | 22.5405 | 22.4844 | 50.3842 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.048655 | 0.88043 | 0.042837 | 18.0953 | 22.2558 | 3.4735 | 0.045029 | 0.4531 | 0.397 | 0.2847 | 0.2286
2012-10-30 22:16 | 26.6144 | 22.7202 | 22.6641 | 22.5719 | 22.4556 | 50.3958 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.04867 | 0.88056 | 0.042857 | 18.0924 | 22.2511 | 3.4743 | 0.045053 | 0.4691 | 0.413 | 0.3208 | 0.2045
2012-10-30 22:16 | 26.6121 | 22.6498 | 22.6177 | 22.6057 | 22.4373 | 50.4016 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048636 | 0.88018 | 0.042809 | 18.0973 | 22.2167 | 3.4747 | 0.045003 | 0.4331 | 0.401 | 0.389 | 0.2206
2012-10-30 22:16 | 26.6031 | 22.6768 | 22.6408 | 22.5366 | 22.4564 | 50.3865 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048651 | 0.88036 | 0.04283 | 18.0953 | 22.2277 | 3.4736 | 0.045025 | 0.4491 | 0.4131 | 0.3089 | 0.2287
2012-10-30 22:16 | 26.6054 | 22.6872 | 22.6832 | 22.5509 | 22.4587 | 50.3854 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048661 | 0.8804 | 0.042841 | 18.0925 | 22.2661 | 3.4736 | 0.045036 | 0.4211 | 0.4171 | 0.2848 | 0.1926
2012-10-30 22:16 | 26.6076 | 22.7094 | 22.6854 | 22.5892 | 22.4609 | 50.3923 | 0.006707 | 0.32638 | 0.002189 | 48.6632 | 0.048645 | 0.88038 | 0.042826 | 18.0982 | 22.2523 | 3.474 | 0.045015 | 0.4571 | 0.4331 | 0.3369 | 0.2086
2012-10-30 22:16 | 26.6061 | 22.7159 | 22.6277 | 22.6077 | 22.4353 | 50.3912 | 0.006744 | 0.32636 | 0.002201 | 48.3919 | 0.04867 | 0.88044 | 0.042851 | 18.09 | 22.2588 | 3.474 | 0.045052 | 0.4571 | 0.3689 | 0.3489 | 0.1765
2012-10-30 22:16 | 26.5979 | 22.6917 | 22.6556 | 22.5835 | 22.4431 | 50.4016 | 0.006751 | 0.32631 | 0.002203 | 48.3354 | 0.048655 | 0.8803 | 0.042831 | 18.0927 | 22.2627 | 3.4747 | 0.045034 | 0.429 | 0.3929 | 0.3208 | 0.1804
2012-10-30 22:16 | 26.5986 | 22.6643 | 22.6323 | 22.5641 | 22.4278 | 50.3987 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048651 | 0.88037 | 0.042831 | 18.0956 | 22.2353 | 3.4745 | 0.045026 | 0.429 | 0.397 | 0.3288 | 0.1925
2012-10-30 22:16 | 26.5927 | 22.6705 | 22.6705 | 22.5662 | 22.4379 | 50.3912 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.048655 | 0.8803 | 0.042831 | 18.0927 | 22.2374 | 3.474 | 0.045025 | 0.4331 | 0.4331 | 0.3288 | 0.2005
2012-10-30 22:16 | 26.5912 | 22.6569 | 22.669 | 22.5647 | 22.4485 | 50.3848 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.048661 | 0.88037 | 0.04284 | 18.0919 | 22.2479 | 3.4735 | 0.045038 | 0.409 | 0.4211 | 0.3168 | 0.2006
2012-10-30 22:16 | 26.5912 | 22.713 | 22.669 | 22.5527 | 22.4244 | 50.3992 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048658 | 0.88028 | 0.042833 | 18.0912 | 22.2519 | 3.4745 | 0.045027 | 0.4611 | 0.4171 | 0.3008 | 0.1725
2012-10-30 22:16 | 26.5889 | 22.7107 | 22.6626 | 22.5424 | 22.4341 | 50.4004 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.048645 | 0.88024 | 0.042819 | 18.0952 | 22.1935 | 3.4746 | 0.045012 | 0.5172 | 0.4691 | 0.3489 | 0.2406
2012-10-30 22:16 | 26.5904 | 22.6922 | 22.6601 | 22.5599 | 22.4557 | 50.3796 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048651 | 0.88033 | 0.042829 | 18.0947 | 22.2672 | 3.4732 | 0.045023 | 0.425 | 0.3929 | 0.2927 | 0.1885
2012-10-30 22:16 | 26.5986 | 22.6884 | 22.6844 | 22.5721 | 22.4479 | 50.3802 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048667 | 0.88021 | 0.042837 | 18.0864 | 22.2634 | 3.4732 | 0.045034 | 0.425 | 0.421 | 0.3087 | 0.1845
2012-10-30 22:16 | 26.5807 | 22.6865 | 22.6424 | 22.5863 | 22.45 | 50.3813 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048689 | 0.88079 | 0.042885 | 18.0901 | 22.2695 | 3.4733 | 0.045084 | 0.417 | 0.3729 | 0.3168 | 0.1805
2012-10-30 22:16 | 26.5784 | 22.7002 | 22.6401 | 22.5479 | 22.4758 | 50.3848 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048705 | 0.88112 | 0.042915 | 18.091 | 22.2552 | 3.4735 | 0.045113 | 0.445 | 0.3849 | 0.2927 | 0.2206
2012-10-30 22:16 | 26.5837 | 22.6574 | 22.6655 | 22.6374 | 22.4249 | 50.3808 | 0.006741 | 0.32636 | 0.0022 | 48.4134 | 0.04868 | 0.88067 | 0.042871 | 18.0911 | 22.1963 | 3.4733 | 0.045071 | 0.4611 | 0.4692 | 0.4411 | 0.2286
2012-10-30 22:16 | 26.5837 | 22.6895 | 22.6735 | 22.5853 | 22.457 | 50.3877 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048667 | 0.88073 | 0.042863 | 18.0971 | 22.2444 | 3.4737 | 0.045058 | 0.4451 | 0.4291 | 0.3409 | 0.2126
2012-10-30 22:16 | 26.5792 | 22.685 | 22.677 | 22.5808 | 22.4605 | 50.3854 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048692 | 0.88091 | 0.042893 | 18.0914 | 22.2921 | 3.4736 | 0.045089 | 0.3929 | 0.3849 | 0.2887 | 0.1684
2012-10-30 22:16 | 26.5762 | 22.6139 | 22.674 | 22.5578 | 22.4255 | 50.3906 | 0.00671 | 0.3264 | 0.00219 | 48.6437 | 0.04868 | 0.88073 | 0.042874 | 18.0923 | 22.2329 | 3.4739 | 0.045064 | 0.381 | 0.4411 | 0.3249 | 0.1926
2012-10-30 22:16 | 26.5807 | 22.6985 | 22.6985 | 22.5983 | 22.438 | 50.3704 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.048705 | 0.88105 | 0.042911 | 18.0895 | 22.2936 | 3.4725 | 0.045111 | 0.4049 | 0.4049 | 0.3047 | 0.1444
2012-10-30 22:16 | 26.5792 | 22.701 | 22.685 | 22.6008 | 22.4485 | 50.3848 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048711 | 0.88122 | 0.042925 | 18.0908 | 22.2841 | 3.4735 | 0.045121 | 0.4169 | 0.4009 | 0.3167 | 0.1644
2012-10-30 22:16 | 26.5724 | 22.6582 | 22.6902 | 22.554 | 22.4136 | 50.3958 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048711 | 0.88135 | 0.042932 | 18.0935 | 22.2733 | 3.4743 | 0.045131 | 0.3849 | 0.4169 | 0.2807 | 0.1403
2012-10-30 22:17 | 26.568 | 22.6538 | 22.6698 | 22.5536 | 22.4453 | 50.3894 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048721 | 0.88108 | 0.042927 | 18.0842 | 22.2849 | 3.4738 | 0.045124 | 0.3689 | 0.3849 | 0.2687 | 0.1604
2012-10-30 22:17 | 26.5687 | 22.6304 | 22.6785 | 22.5703 | 22.418 | 50.3906 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048689 | 0.88093 | 0.042892 | 18.0931 | 22.2615 | 3.4739 | 0.045089 | 0.3689 | 0.417 | 0.3088 | 0.1565
2012-10-30 22:17 | 26.5657 | 22.6795 | 22.6956 | 22.5673 | 22.455 | 50.3969 | 0.006747 | 0.32644 | 0.002203 | 48.3833 | 0.048705 | 0.88106 | 0.042912 | 18.0898 | 22.2545 | 3.4744 | 0.045115 | 0.425 | 0.4411 | 0.3128 | 0.2005
2012-10-30 22:17 | 26.565 | 22.6628 | 22.6868 | 22.5866 | 22.4263 | 50.3779 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048677 | 0.88062 | 0.042866 | 18.091 | 22.2779 | 3.4731 | 0.045061 | 0.3849 | 0.4089 | 0.3087 | 0.1484
2012-10-30 22:17 | 26.5724 | 22.7023 | 22.6742 | 22.5299 | 22.4257 | 50.3958 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048673 | 0.88075 | 0.042869 | 18.0952 | 22.2933 | 3.4743 | 0.045065 | 0.409 | 0.3809 | 0.2366 | 0.1324
2012-10-30 22:17 | 26.5702 | 22.656 | 22.684 | 22.5678 | 22.4395 | 50.3894 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048673 | 0.88067 | 0.042865 | 18.0937 | 22.243 | 3.4738 | 0.045063 | 0.413 | 0.441 | 0.3248 | 0.1965
2012-10-30 22:17 | 26.568 | 22.6818 | 22.6778 | 22.5696 | 22.4453 | 50.3854 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048749 | 0.88155 | 0.042975 | 18.0835 | 22.2368 | 3.4736 | 0.045172 | 0.445 | 0.441 | 0.3328 | 0.2085
2012-10-30 22:17 | 26.5665 | 22.6723 | 22.6763 | 22.5801 | 22.3997 | 50.3912 | 0.006747 | 0.32647 | 0.002203 | 48.3876 | 0.048774 | 0.88226 | 0.043031 | 18.0888 | 22.2593 | 3.474 | 0.045234 | 0.413 | 0.417 | 0.3208 | 0.1404
2012-10-30 22:17 | 26.5732 | 22.695 | 22.651 | 22.5588 | 22.4585 | 50.3773 | 0.006744 | 0.32636 | 0.002201 | 48.3919 | 0.048752 | 0.88189 | 0.042994 | 18.0892 | 22.2781 | 3.473 | 0.045195 | 0.4169 | 0.3729 | 0.2807 | 0.1804
2012-10-30 22:17 | 26.5635 | 22.6773 | 22.6733 | 22.5731 | 22.4488 | 50.3825 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.048714 | 0.88125 | 0.042929 | 18.0903 | 22.2884 | 3.4734 | 0.045122 | 0.3889 | 0.3849 | 0.2847 | 0.1604
2012-10-30 22:17 | 26.5552 | 22.641 | 22.665 | 22.5608 | 22.4405 | 50.3871 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.048727 | 0.88154 | 0.042955 | 18.0914 | 22.2561 | 3.4737 | 0.045149 | 0.3849 | 0.4089 | 0.3047 | 0.1844
2012-10-30 22:17 | 26.562 | 22.6357 | 22.6478 | 22.5836 | 22.4393 | 50.3865 | 0.006747 | 0.32641 | 0.002202 | 48.379 | 0.048736 | 0.8816 | 0.042966 | 18.0893 | 22.319 | 3.4736 | 0.045168 | 0.3167 | 0.3288 | 0.2646 | 0.1203
2012-10-30 22:17 | 26.5492 | 22.635 | 22.663 | 22.5428 | 22.4426 | 50.3831 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048777 | 0.88212 | 0.043027 | 18.0847 | 22.2902 | 3.4734 | 0.045223 | 0.3448 | 0.3728 | 0.2526 | 0.1524
2012-10-30 22:17 | 26.5507 | 22.6204 | 22.6765 | 22.5483 | 22.44 | 50.3888 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048695 | 0.88119 | 0.04291 | 18.0962 | 22.2676 | 3.4738 | 0.045109 | 0.3528 | 0.4089 | 0.2807 | 0.1724
2012-10-30 22:17 | 26.556 | 22.6738 | 22.7059 | 22.5255 | 22.4373 | 50.3894 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.048777 | 0.88236 | 0.043039 | 18.0898 | 22.2809 | 3.4738 | 0.045232 | 0.3929 | 0.425 | 0.2446 | 0.1564
2012-10-30 22:17 | 26.5575 | 22.6513 | 22.6753 | 22.5591 | 22.4308 | 50.3819 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048711 | 0.8814 | 0.042934 | 18.0944 | 22.2744 | 3.4733 | 0.045133 | 0.3769 | 0.4009 | 0.2847 | 0.1564
2012-10-30 22:17 | 26.5642 | 22.6379 | 22.662 | 22.5458 | 22.4616 | 50.3808 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048903 | 0.88433 | 0.043246 | 18.0833 | 22.249 | 3.4733 | 0.045442 | 0.3889 | 0.413 | 0.2968 | 0.2126
2012-10-30 22:17 | 26.5545 | 22.6683 | 22.6683 | 22.5441 | 22.4238 | 50.3784 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.04878 | 0.88219 | 0.043033 | 18.0851 | 22.2955 | 3.4731 | 0.04523 | 0.3728 | 0.3728 | 0.2486 | 0.1283
2012-10-30 22:17 | 26.556 | 22.6418 | 22.6778 | 22.5736 | 22.4494 | 50.386 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.04878 | 0.88239 | 0.043043 | 18.0892 | 22.309 | 3.4736 | 0.045237 | 0.3328 | 0.3688 | 0.2646 | 0.1404
2012-10-30 22:17 | 26.5575 | 22.6633 | 22.6553 | 22.5751 | 22.4709 | 50.3779 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.04885 | 0.88338 | 0.043153 | 18.0834 | 22.2864 | 3.4731 | 0.045346 | 0.3769 | 0.3689 | 0.2887 | 0.1845
2012-10-30 22:17 | 26.5605 | 22.6583 | 22.6783 | 22.5501 | 22.4418 | 50.3813 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048847 | 0.88335 | 0.043149 | 18.0839 | 22.2453 | 3.4733 | 0.045346 | 0.413 | 0.433 | 0.3048 | 0.1965
2012-10-30 22:17 | 26.553 | 22.6468 | 22.6869 | 22.5586 | 22.4464 | 50.3894 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048787 | 0.88249 | 0.043054 | 18.0887 | 22.2378 | 3.4738 | 0.045249 | 0.409 | 0.4491 | 0.3208 | 0.2086
2012-10-30 22:17 | 26.559 | 22.6608 | 22.7049 | 22.5686 | 22.4002 | 50.379 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048768 | 0.88225 | 0.043025 | 18.0907 | 22.2599 | 3.4731 | 0.045223 | 0.4009 | 0.445 | 0.3087 | 0.1403
2012-10-30 22:17 | 26.5597 | 22.6535 | 22.5974 | 22.5332 | 22.429 | 50.379 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.04879 | 0.88231 | 0.043048 | 18.0837 | 22.2565 | 3.4731 | 0.045243 | 0.397 | 0.3409 | 0.2767 | 0.1725
2012-10-30 22:17 | 26.5612 | 22.639 | 22.6349 | 22.5468 | 22.4947 | 50.3727 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.048755 | 0.88199 | 0.043001 | 18.0902 | 22.254 | 3.4727 | 0.045194 | 0.385 | 0.3809 | 0.2928 | 0.2407
2012-10-30 22:17 | 26.556 | 22.6338 | 22.6498 | 22.5496 | 22.4213 | 50.379 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048812 | 0.88288 | 0.043095 | 18.0874 | 22.2208 | 3.4731 | 0.045292 | 0.413 | 0.429 | 0.3288 | 0.2005
2012-10-30 22:17 | 26.5605 | 22.6423 | 22.6623 | 22.5781 | 22.4418 | 50.3825 | 0.006744 | 0.32641 | 0.002201 | 48.4005 | 0.048806 | 0.88291 | 0.043091 | 18.0903 | 22.2212 | 3.4734 | 0.045292 | 0.4211 | 0.4411 | 0.3569 | 0.2206
2012-10-30 22:18 | 26.5567 | 22.6224 | 22.6304 | 22.5022 | 22.406 | 50.3836 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048774 | 0.88218 | 0.043027 | 18.087 | 22.2335 | 3.4734 | 0.045224 | 0.3889 | 0.3969 | 0.2687 | 0.1725
2012-10-30 22:18 | 26.5605 | 22.6302 | 22.6463 | 22.5661 | 22.4619 | 50.3917 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.048831 | 0.88307 | 0.043121 | 18.0843 | 22.2293 | 3.474 | 0.045314 | 0.4009 | 0.417 | 0.3368 | 0.2326
2012-10-30 22:18 | 26.553 | 22.6548 | 22.6708 | 22.5746 | 22.4023 | 50.3721 | 0.006719 | 0.32646 | 0.002193 | 48.587 | 0.048799 | 0.88274 | 0.043077 | 18.0893 | 22.2178 | 3.4727 | 0.04527 | 0.437 | 0.453 | 0.3568 | 0.1845
2012-10-30 22:18 | 26.547 | 22.6328 | 22.6528 | 22.5205 | 22.4083 | 50.3779 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048812 | 0.88285 | 0.043094 | 18.0868 | 22.2438 | 3.4731 | 0.045289 | 0.389 | 0.409 | 0.2767 | 0.1645
2012-10-30 22:18 | 26.55 | 22.6718 | 22.6638 | 22.5636 | 22.4594 | 50.386 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048793 | 0.8827 | 0.043069 | 18.0906 | 22.2027 | 3.4736 | 0.045266 | 0.4691 | 0.4611 | 0.3609 | 0.2567
2012-10-30 22:18 | 26.5492 | 22.663 | 22.6109 | 22.5468 | 22.4025 | 50.3779 | 0.006744 | 0.32637 | 0.002201 | 48.3941 | 0.04879 | 0.88254 | 0.043059 | 18.0885 | 22.238 | 3.4731 | 0.04526 | 0.425 | 0.3729 | 0.3088 | 0.1645
2012-10-30 22:18 | 26.5575 | 22.6753 | 22.6152 | 22.5471 | 22.4428 | 50.3842 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.048806 | 0.88272 | 0.043082 | 18.0864 | 22.2584 | 3.4735 | 0.045277 | 0.4169 | 0.3568 | 0.2887 | 0.1844
2012-10-30 22:18 | 26.553 | 22.6267 | 22.6147 | 22.5506 | 22.4063 | 50.3761 | 0.006735 | 0.32647 | 0.002199 | 48.4738 | 0.048809 | 0.88265 | 0.043081 | 18.0838 | 22.2458 | 3.4729 | 0.04528 | 0.3809 | 0.3689 | 0.3048 | 0.1605
2012-10-30 22:18 | 26.553 | 22.6308 | 22.6267 | 22.5426 | 22.4263 | 50.3675 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.048815 | 0.88294 | 0.043101 | 18.0875 | 22.2458 | 3.4723 | 0.045295 | 0.385 | 0.3809 | 0.2968 | 0.1805
2012-10-30 22:18 | 26.5462 | 22.648 | 22.5919 | 22.5478 | 22.3714 | 50.3727 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048834 | 0.88317 | 0.043129 | 18.0852 | 22.231 | 3.4727 | 0.045325 | 0.417 | 0.3609 | 0.3168 | 0.1404
2012-10-30 22:18 | 26.5522 | 22.6219 | 22.6259 | 22.5498 | 22.4456 | 50.3744 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.048887 | 0.88394 | 0.043213 | 18.0813 | 22.2531 | 3.4728 | 0.045409 | 0.3688 | 0.3728 | 0.2967 | 0.1925
2012-10-30 22:18 | 26.553 | 22.6308 | 22.6227 | 22.5786 | 22.4103 | 50.3761 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048869 | 0.88381 | 0.043191 | 18.0853 | 22.2458 | 3.4729 | 0.045387 | 0.385 | 0.3769 | 0.3328 | 0.1645
2012-10-30 22:18 | 26.5462 | 22.652 | 22.6199 | 22.5358 | 22.4035 | 50.3721 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.0489 | 0.88429 | 0.043242 | 18.0836 | 22.219 | 3.4727 | 0.045439 | 0.433 | 0.4009 | 0.3168 | 0.1845
2012-10-30 22:18 | 26.5552 | 22.677 | 22.641 | 22.5007 | 22.4245 | 50.386 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.0489 | 0.88423 | 0.043239 | 18.0824 | 22.2601 | 3.4736 | 0.045436 | 0.4169 | 0.3809 | 0.2406 | 0.1644
2012-10-30 22:18 | 26.5462 | 22.664 | 22.644 | 22.5398 | 22.3834 | 50.3686 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048774 | 0.88235 | 0.043036 | 18.0906 | 22.215 | 3.4724 | 0.04523 | 0.449 | 0.429 | 0.3248 | 0.1684
2012-10-30 22:18 | 26.5485 | 22.6022 | 22.6182 | 22.5461 | 22.3937 | 50.3709 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048954 | 0.88509 | 0.043329 | 18.0801 | 22.2052 | 3.4726 | 0.045528 | 0.397 | 0.413 | 0.3409 | 0.1885
2012-10-30 22:18 | 26.5522 | 22.5979 | 22.642 | 22.5297 | 22.4095 | 50.3802 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048935 | 0.88478 | 0.043297 | 18.0807 | 22.225 | 3.4732 | 0.045491 | 0.3729 | 0.417 | 0.3047 | 0.1845
2012-10-30 22:18 | 26.553 | 22.6227 | 22.6588 | 22.5586 | 22.3982 | 50.3784 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048976 | 0.88569 | 0.043377 | 18.0841 | 22.2539 | 3.4731 | 0.045573 | 0.3688 | 0.4049 | 0.3047 | 0.1443
2012-10-30 22:18 | 26.5515 | 22.6533 | 22.6293 | 22.5331 | 22.4449 | 50.3732 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048979 | 0.88538 | 0.043365 | 18.0768 | 22.2363 | 3.4727 | 0.045559 | 0.417 | 0.393 | 0.2968 | 0.2086
2012-10-30 22:18 | 26.5575 | 22.6393 | 22.6793 | 22.5791 | 22.4468 | 50.3646 | 0.006716 | 0.32647 | 0.002193 | 48.6109 | 0.048897 | 0.88413 | 0.043231 | 18.0814 | 22.2624 | 3.4721 | 0.045424 | 0.3769 | 0.4169 | 0.3167 | 0.1844
2012-10-30 22:18 | 26.559 | 22.6528 | 22.6368 | 22.6007 | 22.3882 | 50.3617 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048919 | 0.88458 | 0.043273 | 18.0824 | 22.2478 | 3.4719 | 0.045467 | 0.405 | 0.389 | 0.3529 | 0.1404
2012-10-30 22:18 | 26.5545 | 22.6563 | 22.6443 | 22.512 | 22.4038 | 50.3692 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048969 | 0.8854 | 0.043357 | 18.0808 | 22.2513 | 3.4725 | 0.045554 | 0.405 | 0.393 | 0.2607 | 0.1525
2012-10-30 22:18 | 26.5635 | 22.6653 | 22.6533 | 22.5731 | 22.4368 | 50.3732 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048928 | 0.88447 | 0.043275 | 18.077 | 22.2403 | 3.4727 | 0.045472 | 0.425 | 0.413 | 0.3328 | 0.1965
2012-10-30 22:18 | 26.5537 | 22.6555 | 22.6234 | 22.5874 | 22.4671 | 50.3513 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.049042 | 0.88648 | 0.043475 | 18.076 | 22.2024 | 3.4712 | 0.045674 | 0.4531 | 0.421 | 0.385 | 0.2647
2012-10-30 22:18 | 26.5552 | 22.6129 | 22.6049 | 22.5047 | 22.4325 | 50.3761 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.049023 | 0.88615 | 0.043442 | 18.0762 | 22.2159 | 3.4729 | 0.045637 | 0.397 | 0.389 | 0.2888 | 0.2166
2012-10-30 22:18 | 26.553 | 22.6708 | 22.6388 | 22.5426 | 22.3982 | 50.3767 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048966 | 0.88551 | 0.04336 | 18.0843 | 22.2378 | 3.473 | 0.045557 | 0.433 | 0.401 | 0.3048 | 0.1604
2012-10-30 22:18 | 26.556 | 22.6057 | 22.5897 | 22.5295 | 22.4133 | 50.3646 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.04879 | 0.88264 | 0.043064 | 18.0906 | 22.2488 | 3.4721 | 0.045262 | 0.3569 | 0.3409 | 0.2807 | 0.1645
2012-10-30 22:19 | 26.565 | 22.6187 | 22.6387 | 22.5506 | 22.4223 | 50.3669 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.04885 | 0.88374 | 0.043171 | 18.0908 | 22.2498 | 3.4723 | 0.045366 | 0.3689 | 0.3889 | 0.3008 | 0.1725
2012-10-30 22:19 | 26.5515 | 22.6573 | 22.6052 | 22.513 | 22.3526 | 50.3738 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.048878 | 0.88391 | 0.043204 | 18.084 | 22.2403 | 3.4728 | 0.045403 | 0.417 | 0.3649 | 0.2727 | 0.1123
2012-10-30 22:19 | 26.562 | 22.6558 | 22.6478 | 22.5155 | 22.4273 | 50.3686 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.048859 | 0.88356 | 0.04317 | 18.0839 | 22.2789 | 3.4724 | 0.045363 | 0.3769 | 0.3689 | 0.2366 | 0.1484
2012-10-30 22:19 | 26.5507 | 22.6565 | 22.6365 | 22.5122 | 22.408 | 50.3582 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.048982 | 0.88547 | 0.043372 | 18.0775 | 22.2235 | 3.4717 | 0.04557 | 0.433 | 0.413 | 0.2887 | 0.1845
2012-10-30 22:19 | 26.5373 | 22.603 | 22.571 | 22.5469 | 22.3866 | 50.3565 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048894 | 0.88394 | 0.043219 | 18.0787 | 22.2061 | 3.4716 | 0.045414 | 0.3969 | 0.3649 | 0.3408 | 0.1805
2012-10-30 22:19 | 26.553 | 22.6708 | 22.6027 | 22.5466 | 22.4383 | 50.3611 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.048932 | 0.88475 | 0.043292 | 18.0812 | 22.2258 | 3.4719 | 0.045488 | 0.445 | 0.3769 | 0.3208 | 0.2125
2012-10-30 22:19 | 26.55 | 22.6518 | 22.6077 | 22.5235 | 22.3912 | 50.3605 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.049004 | 0.88579 | 0.043407 | 18.0758 | 22.2428 | 3.4719 | 0.045603 | 0.409 | 0.3649 | 0.2807 | 0.1484
2012-10-30 22:19 | 26.5395 | 22.6052 | 22.6213 | 22.5451 | 22.4048 | 50.3727 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.049101 | 0.88733 | 0.043569 | 18.0716 | 22.2283 | 3.4727 | 0.045769 | 0.3769 | 0.393 | 0.3168 | 0.1765
2012-10-30 22:19 | 26.55 | 22.6358 | 22.6157 | 22.5596 | 22.4153 | 50.3553 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.049026 | 0.88611 | 0.043442 | 18.0742 | 22.2388 | 3.4715 | 0.045639 | 0.397 | 0.3769 | 0.3208 | 0.1765
2012-10-30 22:19 | 26.5425 | 22.6924 | 22.6483 | 22.5561 | 22.4479 | 50.3646 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.049013 | 0.88606 | 0.043429 | 18.0781 | 22.1872 | 3.4721 | 0.045626 | 0.5052 | 0.4611 | 0.3689 | 0.2607
2012-10-30 22:19 | 26.5388 | 22.6727 | 22.6406 | 22.5364 | 22.4241 | 50.3599 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.048809 | 0.8827 | 0.043084 | 18.0847 | 22.2236 | 3.4718 | 0.045276 | 0.4491 | 0.417 | 0.3128 | 0.2005
2012-10-30 22:19 | 26.5418 | 22.6396 | 22.6596 | 22.5715 | 22.4472 | 50.3495 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048928 | 0.88455 | 0.043279 | 18.0785 | 22.2747 | 3.4711 | 0.045473 | 0.3649 | 0.3849 | 0.2968 | 0.1725
2012-10-30 22:19 | 26.526 | 22.6839 | 22.6318 | 22.5396 | 22.3712 | 50.3599 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.048944 | 0.88481 | 0.043306 | 18.0779 | 22.2549 | 3.4718 | 0.045505 | 0.429 | 0.3769 | 0.2847 | 0.1163
2012-10-30 22:19 | 26.526 | 22.6478 | 22.6398 | 22.5236 | 22.4033 | 50.3553 | 0.00671 | 0.32638 | 0.00219 | 48.6414 | 0.048784 | 0.88244 | 0.043049 | 18.0886 | 22.2509 | 3.4715 | 0.045239 | 0.3969 | 0.3889 | 0.2727 | 0.1524
2012-10-30 22:19 | 26.5283 | 22.6461 | 22.6501 | 22.5459 | 22.3896 | 50.3669 | 0.006741 | 0.32643 | 0.0022 | 48.4241 | 0.048859 | 0.88345 | 0.043164 | 18.0816 | 22.2653 | 3.4723 | 0.045364 | 0.3808 | 0.3848 | 0.2806 | 0.1243
2012-10-30 22:19 | 26.532 | 22.6378 | 22.6418 | 22.5216 | 22.4254 | 50.3657 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048796 | 0.88264 | 0.043069 | 18.0883 | 22.2569 | 3.4722 | 0.045269 | 0.3809 | 0.3849 | 0.2647 | 0.1685
2012-10-30 22:19 | 26.5268 | 22.6406 | 22.6567 | 22.5284 | 22.3801 | 50.3628 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.049161 | 0.88821 | 0.043666 | 18.0675 | 22.3159 | 3.472 | 0.045862 | 0.3247 | 0.3408 | 0.2125 | 0.0642
2012-10-30 22:19 | 26.523 | 22.6208 | 22.5967 | 22.5166 | 22.4524 | 50.3519 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.048843 | 0.88338 | 0.043147 | 18.086 | 22.2359 | 3.4713 | 0.045347 | 0.3849 | 0.3608 | 0.2807 | 0.2165
2012-10-30 22:19 | 26.532 | 22.6538 | 22.6619 | 22.5296 | 22.3973 | 50.3594 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048998 | 0.88576 | 0.0434 | 18.0775 | 22.2128 | 3.4718 | 0.045597 | 0.441 | 0.4491 | 0.3168 | 0.1845
2012-10-30 22:19 | 26.5253 | 22.6191 | 22.583 | 22.4988 | 22.4026 | 50.3634 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048988 | 0.88575 | 0.043391 | 18.0809 | 22.2262 | 3.4721 | 0.045587 | 0.3929 | 0.3568 | 0.2726 | 0.1764
2012-10-30 22:19 | 26.523 | 22.6008 | 22.6048 | 22.4765 | 22.3843 | 50.3576 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.048749 | 0.8818 | 0.042987 | 18.0886 | 22.2519 | 3.4717 | 0.045185 | 0.3489 | 0.3529 | 0.2246 | 0.1324
2012-10-30 22:19 | 26.5275 | 22.5932 | 22.6253 | 22.5411 | 22.4128 | 50.3599 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.049114 | 0.88751 | 0.043589 | 18.0703 | 22.2685 | 3.4718 | 0.045783 | 0.3247 | 0.3568 | 0.2726 | 0.1443
2012-10-30 22:19 | 26.5238 | 22.6537 | 22.6216 | 22.5174 | 22.3971 | 50.3536 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048865 | 0.88397 | 0.043195 | 18.09 | 22.2527 | 3.4714 | 0.045392 | 0.401 | 0.3689 | 0.2647 | 0.1444
2012-10-30 22:19 | 26.5208 | 22.6466 | 22.6106 | 22.5344 | 22.4342 | 50.3455 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.048988 | 0.88546 | 0.043377 | 18.075 | 22.2537 | 3.4708 | 0.045576 | 0.3929 | 0.3569 | 0.2807 | 0.1805
2012-10-30 22:19 | 26.5223 | 22.6201 | 22.6201 | 22.5199 | 22.3836 | 50.3617 | 0.006747 | 0.32644 | 0.002203 | 48.3833 | 0.04885 | 0.88338 | 0.043153 | 18.0834 | 22.2392 | 3.4719 | 0.045356 | 0.3809 | 0.3809 | 0.2807 | 0.1444
2012-10-30 22:19 | 26.526 | 22.6559 | 22.6198 | 22.5236 | 22.4073 | 50.3386 | 0.006754 | 0.32638 | 0.002204 | 48.3245 | 0.04895 | 0.88508 | 0.043325 | 18.0813 | 22.271 | 3.4703 | 0.045529 | 0.3849 | 0.3488 | 0.2526 | 0.1363
2012-10-30 22:20 | 26.526 | 22.6278 | 22.6118 | 22.5396 | 22.3953 | 50.353 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048683 | 0.88086 | 0.042883 | 18.0938 | 22.2228 | 3.4713 | 0.04508 | 0.405 | 0.389 | 0.3168 | 0.1725
2012-10-30 22:20 | 26.523 | 22.6248 | 22.6168 | 22.5406 | 22.4123 | 50.3438 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.049142 | 0.88807 | 0.043642 | 18.0715 | 22.2198 | 3.4707 | 0.045835 | 0.405 | 0.397 | 0.3208 | 0.1925
2012-10-30 22:20 | 26.5148 | 22.6366 | 22.6086 | 22.4883 | 22.4162 | 50.3652 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.048812 | 0.88285 | 0.043094 | 18.0868 | 22.2598 | 3.4722 | 0.045293 | 0.3768 | 0.3488 | 0.2285 | 0.1564
2012-10-30 22:20 | 26.5193 | 22.6171 | 22.6211 | 22.5289 | 22.3846 | 50.3617 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048711 | 0.88128 | 0.042928 | 18.092 | 22.2643 | 3.4719 | 0.045125 | 0.3528 | 0.3568 | 0.2646 | 0.1203
2012-10-30 22:20 | 26.5275 | 22.6493 | 22.6413 | 22.5531 | 22.4329 | 50.3611 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.048721 | 0.88122 | 0.042934 | 18.0871 | 22.2524 | 3.4719 | 0.04513 | 0.3969 | 0.3889 | 0.3007 | 0.1805
2012-10-30 22:20 | 26.5328 | 22.6627 | 22.6426 | 22.5705 | 22.3821 | 50.3582 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048686 | 0.88082 | 0.042884 | 18.0918 | 22.2698 | 3.4717 | 0.045079 | 0.3929 | 0.3728 | 0.3007 | 0.1123
2012-10-30 22:20 | 26.5245 | 22.6463 | 22.6503 | 22.5181 | 22.4138 | 50.3547 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048717 | 0.88137 | 0.042938 | 18.0916 | 22.2374 | 3.4715 | 0.045133 | 0.4089 | 0.4129 | 0.2807 | 0.1764
2012-10-30 22:20 | 26.5335 | 22.6593 | 22.6353 | 22.5712 | 22.4269 | 50.3524 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.04868 | 0.88063 | 0.042869 | 18.0902 | 22.2664 | 3.4713 | 0.045061 | 0.3929 | 0.3689 | 0.3048 | 0.1605
2012-10-30 22:20 | 26.5223 | 22.6441 | 22.6241 | 22.5159 | 22.3996 | 50.3495 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048705 | 0.88112 | 0.042915 | 18.091 | 22.2151 | 3.4711 | 0.04511 | 0.429 | 0.409 | 0.3008 | 0.1845
2012-10-30 22:20 | 26.5283 | 22.6541 | 22.6221 | 22.5219 | 22.3816 | 50.3542 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048686 | 0.88106 | 0.042895 | 18.0969 | 22.2492 | 3.4714 | 0.045094 | 0.4049 | 0.3729 | 0.2727 | 0.1324
2012-10-30 22:20 | 26.5268 | 22.6366 | 22.6366 | 22.5204 | 22.4041 | 50.3571 | 0.006725 | 0.32647 | 0.002196 | 48.5459 | 0.048651 | 0.88033 | 0.042829 | 18.0947 | 22.2397 | 3.4716 | 0.045025 | 0.3969 | 0.3969 | 0.2807 | 0.1644
2012-10-30 22:20 | 26.5298 | 22.6356 | 22.6637 | 22.5154 | 22.4071 | 50.3443 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048667 | 0.88053 | 0.042853 | 18.0929 | 22.2627 | 3.4707 | 0.045048 | 0.3729 | 0.401 | 0.2527 | 0.1444
2012-10-30 22:20 | 26.5305 | 22.6483 | 22.6203 | 22.504 | 22.4239 | 50.3495 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.04868 | 0.88076 | 0.042875 | 18.0929 | 22.2514 | 3.4711 | 0.045068 | 0.3969 | 0.3689 | 0.2526 | 0.1725
2012-10-30 22:20 | 26.5223 | 22.6401 | 22.6401 | 22.5359 | 22.3675 | 50.3513 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048708 | 0.88109 | 0.042916 | 18.0893 | 22.2392 | 3.4712 | 0.045114 | 0.4009 | 0.4009 | 0.2967 | 0.1283
2012-10-30 22:20 | 26.526 | 22.6038 | 22.6278 | 22.5356 | 22.3993 | 50.3478 | 0.006716 | 0.32631 | 0.002192 | 48.5873 | 0.048667 | 0.8805 | 0.042851 | 18.0923 | 22.2509 | 3.471 | 0.045043 | 0.3529 | 0.3769 | 0.2847 | 0.1484
2012-10-30 22:20 | 26.5238 | 22.6336 | 22.6056 | 22.5254 | 22.4091 | 50.3495 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048677 | 0.8808 | 0.042875 | 18.0949 | 22.2527 | 3.4711 | 0.045072 | 0.3809 | 0.3529 | 0.2727 | 0.1564
2012-10-30 22:20 | 26.5208 | 22.6667 | 22.6026 | 22.5424 | 22.4142 | 50.3553 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.048689 | 0.88076 | 0.042883 | 18.0895 | 22.2337 | 3.4715 | 0.045081 | 0.433 | 0.3689 | 0.3087 | 0.1805
2012-10-30 22:20 | 26.5171 | 22.6149 | 22.6309 | 22.5307 | 22.3904 | 50.349 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048695 | 0.88091 | 0.042896 | 18.0903 | 22.25 | 3.4711 | 0.04509 | 0.3649 | 0.3809 | 0.2807 | 0.1404
2012-10-30 22:20 | 26.523 | 22.6368 | 22.5967 | 22.5206 | 22.4083 | 50.3565 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.048699 | 0.88088 | 0.042898 | 18.0882 | 22.2519 | 3.4716 | 0.045093 | 0.3849 | 0.3448 | 0.2687 | 0.1564
2012-10-30 22:20 | 26.5245 | 22.6223 | 22.6063 | 22.5221 | 22.3938 | 50.3588 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048658 | 0.88036 | 0.042836 | 18.0927 | 22.2775 | 3.4717 | 0.04503 | 0.3448 | 0.3288 | 0.2446 | 0.1163
2012-10-30 22:20 | 26.5171 | 22.6269 | 22.6269 | 22.5427 | 22.4024 | 50.3467 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048686 | 0.88086 | 0.042886 | 18.0927 | 22.242 | 3.4709 | 0.045082 | 0.3849 | 0.3849 | 0.3007 | 0.1604
2012-10-30 22:20 | 26.5223 | 22.6001 | 22.6281 | 22.5239 | 22.4237 | 50.3536 | 0.006741 | 0.32637 | 0.0022 | 48.4157 | 0.048636 | 0.8802 | 0.042809 | 18.0976 | 22.2272 | 3.4714 | 0.045009 | 0.3729 | 0.4009 | 0.2967 | 0.1965
2012-10-30 22:20 | 26.5215 | 22.6073 | 22.6433 | 22.5552 | 22.4229 | 50.3524 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.04867 | 0.88065 | 0.042861 | 18.0942 | 22.2504 | 3.4713 | 0.045056 | 0.3569 | 0.3929 | 0.3048 | 0.1725
2012-10-30 22:20 | 26.5208 | 22.6106 | 22.6226 | 22.5464 | 22.4262 | 50.3605 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048639 | 0.88027 | 0.042815 | 18.098 | 22.2457 | 3.4719 | 0.045009 | 0.3649 | 0.3769 | 0.3007 | 0.1805
2012-10-30 22:20 | 26.5193 | 22.6291 | 22.6051 | 22.5369 | 22.4006 | 50.3351 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048633 | 0.88031 | 0.042812 | 18.1011 | 22.2362 | 3.4701 | 0.045009 | 0.3929 | 0.3689 | 0.3007 | 0.1644
2012-10-30 22:21 | 26.5171 | 22.655 | 22.6029 | 22.4987 | 22.4385 | 50.3426 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.048655 | 0.88034 | 0.042833 | 18.0935 | 22.246 | 3.4706 | 0.045031 | 0.409 | 0.3569 | 0.2527 | 0.1925
2012-10-30 22:21 | 26.5208 | 22.6266 | 22.6426 | 22.5224 | 22.4302 | 50.3588 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048727 | 0.8814 | 0.042948 | 18.0885 | 22.2618 | 3.4717 | 0.045144 | 0.3648 | 0.3808 | 0.2606 | 0.1684
2012-10-30 22:21 | 26.5156 | 22.6575 | 22.6214 | 22.5132 | 22.3849 | 50.3472 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048661 | 0.88046 | 0.042844 | 18.0937 | 22.2566 | 3.4709 | 0.045039 | 0.4009 | 0.3648 | 0.2566 | 0.1283
2012-10-30 22:21 | 26.5133 | 22.6351 | 22.6231 | 22.4668 | 22.3906 | 50.3449 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.04867 | 0.88067 | 0.042862 | 18.0948 | 22.3144 | 3.4708 | 0.04506 | 0.3207 | 0.3087 | 0.1524 | 0.0762
2012-10-30 22:21 | 26.5073 | 22.5931 | 22.6091 | 22.5089 | 22.4087 | 50.3438 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048636 | 0.88037 | 0.042818 | 18.1012 | 22.2563 | 3.4707 | 0.045015 | 0.3368 | 0.3528 | 0.2526 | 0.1524
2012-10-30 22:21 | 26.5036 | 22.6254 | 22.6054 | 22.5333 | 22.433 | 50.3455 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048661 | 0.88057 | 0.04285 | 18.0961 | 22.2085 | 3.4708 | 0.045048 | 0.4169 | 0.3969 | 0.3248 | 0.2245
2012-10-30 22:21 | 26.4991 | 22.629 | 22.6209 | 22.5408 | 22.3644 | 50.3461 | 0.006725 | 0.32625 | 0.002194 | 48.5136 | 0.048667 | 0.88049 | 0.042851 | 18.0921 | 22.2481 | 3.4709 | 0.045045 | 0.3809 | 0.3728 | 0.2927 | 0.1163
2012-10-30 22:21 | 26.5051 | 22.6229 | 22.6309 | 22.5147 | 22.4145 | 50.3449 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048642 | 0.88011 | 0.04281 | 18.0936 | 22.21 | 3.4708 | 0.045009 | 0.4129 | 0.4209 | 0.3047 | 0.2045
2012-10-30 22:21 | 26.5013 | 22.6111 | 22.6312 | 22.5229 | 22.4147 | 50.3449 | 0.00671 | 0.32625 | 0.002189 | 48.6221 | 0.048677 | 0.88066 | 0.042868 | 18.0919 | 22.2383 | 3.4708 | 0.045057 | 0.3728 | 0.3929 | 0.2846 | 0.1764
2012-10-30 22:21 | 26.4991 | 22.6209 | 22.5568 | 22.5087 | 22.3564 | 50.3519 | 0.006719 | 0.32627 | 0.002192 | 48.559 | 0.048648 | 0.8802 | 0.04282 | 18.0932 | 22.1799 | 3.4713 | 0.045012 | 0.441 | 0.3769 | 0.3288 | 0.1765
2012-10-30 22:21 | 26.4871 | 22.6089 | 22.6089 | 22.4847 | 22.3885 | 50.3449 | 0.006722 | 0.32625 | 0.002193 | 48.5353 | 0.048664 | 0.88056 | 0.042851 | 18.0946 | 22.2 | 3.4708 | 0.045044 | 0.4089 | 0.4089 | 0.2847 | 0.1885
2012-10-30 22:21 | 26.4916 | 22.6455 | 22.5894 | 22.4932 | 22.417 | 50.3484 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.048667 | 0.88052 | 0.042852 | 18.0926 | 22.2125 | 3.471 | 0.045047 | 0.433 | 0.3769 | 0.2807 | 0.2045
2012-10-30 22:21 | 26.4849 | 22.6067 | 22.5627 | 22.4785 | 22.4304 | 50.3438 | 0.006722 | 0.32631 | 0.002193 | 48.5439 | 0.048658 | 0.88057 | 0.042847 | 18.0972 | 22.2178 | 3.4707 | 0.04504 | 0.3889 | 0.3449 | 0.2607 | 0.2126
2012-10-30 22:21 | 26.4841 | 22.622 | 22.5578 | 22.4737 | 22.3574 | 50.3415 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.04867 | 0.88072 | 0.042864 | 18.0957 | 22.213 | 3.4705 | 0.045056 | 0.409 | 0.3448 | 0.2607 | 0.1444
2012-10-30 22:21 | 26.4871 | 22.6009 | 22.5889 | 22.4887 | 22.3764 | 50.3357 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048651 | 0.88043 | 0.042834 | 18.0968 | 22.212 | 3.4701 | 0.045031 | 0.3889 | 0.3769 | 0.2767 | 0.1644
2012-10-30 22:21 | 26.4856 | 22.6355 | 22.5634 | 22.5233 | 22.379 | 50.3472 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048711 | 0.88116 | 0.042922 | 18.0897 | 22.2667 | 3.4709 | 0.045117 | 0.3688 | 0.2967 | 0.2566 | 0.1123
2012-10-30 22:21 | 26.4849 | 22.6388 | 22.6067 | 22.5105 | 22.3823 | 50.3484 | 0.006722 | 0.32628 | 0.002193 | 48.5396 | 0.048683 | 0.88062 | 0.042871 | 18.0888 | 22.27 | 3.471 | 0.045064 | 0.3688 | 0.3367 | 0.2405 | 0.1123
2012-10-30 22:21 | 26.4856 | 22.6315 | 22.5994 | 22.4832 | 22.407 | 50.3432 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.04867 | 0.88046 | 0.042852 | 18.0903 | 22.2747 | 3.4707 | 0.045052 | 0.3568 | 0.3247 | 0.2085 | 0.1323
2012-10-30 22:21 | 26.4901 | 22.628 | 22.5959 | 22.5117 | 22.4035 | 50.3426 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048692 | 0.88076 | 0.042886 | 18.0884 | 22.2511 | 3.4706 | 0.045079 | 0.3769 | 0.3448 | 0.2606 | 0.1524
2012-10-30 22:21 | 26.4826 | 22.6285 | 22.6125 | 22.4962 | 22.392 | 50.3484 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048658 | 0.88053 | 0.042845 | 18.0963 | 22.2356 | 3.471 | 0.045039 | 0.3929 | 0.3769 | 0.2606 | 0.1564
2012-10-30 22:21 | 26.4826 | 22.6565 | 22.5724 | 22.5002 | 22.376 | 50.3478 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048765 | 0.882 | 0.043011 | 18.0868 | 22.2476 | 3.471 | 0.045208 | 0.4089 | 0.3248 | 0.2526 | 0.1284
2012-10-30 22:21 | 26.4849 | 22.6107 | 22.5867 | 22.4905 | 22.3742 | 50.3282 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048683 | 0.88076 | 0.042878 | 18.0917 | 22.2379 | 3.4696 | 0.045072 | 0.3728 | 0.3488 | 0.2526 | 0.1363
2012-10-30 22:21 | 26.4916 | 22.6174 | 22.5774 | 22.5213 | 22.401 | 50.3443 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048727 | 0.88119 | 0.042938 | 18.0843 | 22.2687 | 3.4707 | 0.045135 | 0.3487 | 0.3087 | 0.2526 | 0.1323
2012-10-30 22:21 | 26.4811 | 22.639 | 22.5468 | 22.4907 | 22.3985 | 50.3472 | 0.006725 | 0.32649 | 0.002196 | 48.548 | 0.04873 | 0.88134 | 0.042948 | 18.0861 | 22.2261 | 3.4709 | 0.045144 | 0.4129 | 0.3207 | 0.2646 | 0.1724
2012-10-30 22:21 | 26.4796 | 22.6175 | 22.5894 | 22.4972 | 22.3369 | 50.3339 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048689 | 0.88086 | 0.042888 | 18.0916 | 22.2406 | 3.47 | 0.045087 | 0.3769 | 0.3488 | 0.2566 | 0.0963
2012-10-30 22:22 | 26.4841 | 22.65 | 22.5859 | 22.4817 | 22.4176 | 50.3455 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.04868 | 0.88069 | 0.042872 | 18.0914 | 22.2531 | 3.4708 | 0.045064 | 0.3969 | 0.3328 | 0.2286 | 0.1645
2012-10-30 22:22 | 26.4789 | 22.6168 | 22.5887 | 22.4765 | 22.3682 | 50.3519 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.048664 | 0.88031 | 0.04284 | 18.0896 | 22.1918 | 3.4713 | 0.045039 | 0.425 | 0.3969 | 0.2847 | 0.1764
2012-10-30 22:22 | 26.4864 | 22.6643 | 22.5521 | 22.504 | 22.3677 | 50.3455 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048667 | 0.88056 | 0.042854 | 18.0935 | 22.1953 | 3.4708 | 0.045047 | 0.469 | 0.3568 | 0.3087 | 0.1724
2012-10-30 22:22 | 26.4826 | 22.5924 | 22.5604 | 22.4922 | 22.3399 | 50.3542 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048651 | 0.88002 | 0.042814 | 18.0885 | 22.2035 | 3.4714 | 0.045011 | 0.3889 | 0.3569 | 0.2887 | 0.1364
2012-10-30 22:22 | 26.4849 | 22.6308 | 22.5827 | 22.5146 | 22.3662 | 50.3374 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048689 | 0.88083 | 0.042887 | 18.091 | 22.2138 | 3.4703 | 0.045086 | 0.417 | 0.3689 | 0.3008 | 0.1524
2012-10-30 22:22 | 26.4849 | 22.5907 | 22.5747 | 22.5065 | 22.3662 | 50.3472 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.048692 | 0.88096 | 0.042896 | 18.0925 | 22.1737 | 3.4709 | 0.045095 | 0.417 | 0.401 | 0.3328 | 0.1925
2012-10-30 22:22 | 26.4856 | 22.5954 | 22.5313 | 22.5233 | 22.391 | 50.3455 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.04867 | 0.88047 | 0.042853 | 18.0906 | 22.2226 | 3.4708 | 0.045051 | 0.3728 | 0.3087 | 0.3007 | 0.1684
2012-10-30 22:22 | 26.4901 | 22.6079 | 22.5879 | 22.5037 | 22.3794 | 50.3334 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048683 | 0.88072 | 0.042876 | 18.0909 | 22.211 | 3.47 | 0.045073 | 0.3969 | 0.3769 | 0.2927 | 0.1684
2012-10-30 22:22 | 26.4871 | 22.6089 | 22.5769 | 22.5328 | 22.3684 | 50.3322 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.04868 | 0.88088 | 0.042881 | 18.0952 | 22.2401 | 3.4699 | 0.045076 | 0.3688 | 0.3368 | 0.2927 | 0.1283
2012-10-30 22:22 | 26.4804 | 22.6303 | 22.5902 | 22.494 | 22.3577 | 50.3264 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.048664 | 0.8805 | 0.042849 | 18.0935 | 22.2053 | 3.4695 | 0.04504 | 0.425 | 0.3849 | 0.2887 | 0.1524
2012-10-30 22:22 | 26.4774 | 22.6273 | 22.5592 | 22.491 | 22.3708 | 50.3576 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048673 | 0.88066 | 0.042864 | 18.0934 | 22.2224 | 3.4717 | 0.045062 | 0.4049 | 0.3368 | 0.2686 | 0.1484
2012-10-30 22:22 | 26.4849 | 22.6188 | 22.5627 | 22.5186 | 22.3943 | 50.3386 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048708 | 0.88114 | 0.042918 | 18.0902 | 22.2299 | 3.4703 | 0.045113 | 0.3889 | 0.3328 | 0.2887 | 0.1644
2012-10-30 22:22 | 26.4856 | 22.6195 | 22.6155 | 22.4832 | 22.3749 | 50.3432 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.04868 | 0.88083 | 0.042879 | 18.0943 | 22.2185 | 3.4707 | 0.045073 | 0.401 | 0.397 | 0.2647 | 0.1564
2012-10-30 22:22 | 26.4804 | 22.6223 | 22.5782 | 22.482 | 22.3697 | 50.327 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048664 | 0.88054 | 0.042851 | 18.0943 | 22.2414 | 3.4695 | 0.045048 | 0.3809 | 0.3368 | 0.2406 | 0.1283
2012-10-30 22:22 | 26.4841 | 22.626 | 22.5939 | 22.5138 | 22.3855 | 50.3328 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048677 | 0.88046 | 0.042858 | 18.0877 | 22.217 | 3.4699 | 0.045054 | 0.409 | 0.3769 | 0.2968 | 0.1685
2012-10-30 22:22 | 26.4766 | 22.5904 | 22.5984 | 22.5063 | 22.3619 | 50.327 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.048708 | 0.88115 | 0.042919 | 18.0905 | 22.1975 | 3.4695 | 0.045115 | 0.3929 | 0.4009 | 0.3088 | 0.1644
2012-10-30 22:22 | 26.4781 | 22.656 | 22.608 | 22.4757 | 22.3715 | 50.3374 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048758 | 0.88174 | 0.042992 | 18.0841 | 22.2471 | 3.4703 | 0.045189 | 0.4089 | 0.3609 | 0.2286 | 0.1244
2012-10-30 22:22 | 26.4796 | 22.6375 | 22.5694 | 22.4972 | 22.401 | 50.3299 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048727 | 0.88153 | 0.042954 | 18.0911 | 22.1564 | 3.4697 | 0.045151 | 0.4811 | 0.413 | 0.3408 | 0.2446
2012-10-30 22:22 | 26.4804 | 22.6423 | 22.6062 | 22.482 | 22.3297 | 50.3368 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048721 | 0.88125 | 0.042935 | 18.0877 | 22.2174 | 3.4702 | 0.045132 | 0.4249 | 0.3888 | 0.2646 | 0.1123
2012-10-30 22:22 | 26.4804 | 22.6623 | 22.5942 | 22.5261 | 22.3537 | 50.3386 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048736 | 0.88173 | 0.042972 | 18.0919 | 22.2494 | 3.4703 | 0.045167 | 0.4129 | 0.3448 | 0.2767 | 0.1043
2012-10-30 22:22 | 26.4826 | 22.6445 | 22.5964 | 22.4882 | 22.384 | 50.3357 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048815 | 0.88309 | 0.043108 | 18.0905 | 22.2075 | 3.4701 | 0.045305 | 0.437 | 0.3889 | 0.2807 | 0.1765
2012-10-30 22:22 | 26.4692 | 22.5991 | 22.551 | 22.5149 | 22.3826 | 50.3259 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048711 | 0.88141 | 0.042934 | 18.0947 | 22.2062 | 3.4695 | 0.045131 | 0.3929 | 0.3448 | 0.3087 | 0.1764
2012-10-30 22:22 | 26.4796 | 22.5854 | 22.5934 | 22.4892 | 22.377 | 50.3316 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.048768 | 0.88208 | 0.043017 | 18.0872 | 22.2486 | 3.4699 | 0.045209 | 0.3368 | 0.3448 | 0.2406 | 0.1284
2012-10-30 22:22 | 26.4834 | 22.6173 | 22.5852 | 22.5131 | 22.3246 | 50.3311 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048812 | 0.8829 | 0.043096 | 18.0877 | 22.2204 | 3.4698 | 0.045294 | 0.3969 | 0.3648 | 0.2927 | 0.1042
2012-10-30 22:22 | 26.4729 | 22.5507 | 22.5707 | 22.4985 | 22.3863 | 50.3287 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048847 | 0.88333 | 0.043148 | 18.0836 | 22.246 | 3.4697 | 0.045348 | 0.3047 | 0.3247 | 0.2525 | 0.1403
2012-10-30 22:23 | 26.4856 | 22.6034 | 22.5794 | 22.4631 | 22.395 | 50.3195 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048821 | 0.8828 | 0.043099 | 18.0823 | 22.2627 | 3.469 | 0.045296 | 0.3407 | 0.3167 | 0.2004 | 0.1323
2012-10-30 22:23 | 26.4856 | 22.6114 | 22.5914 | 22.4912 | 22.3629 | 50.3368 | 0.006738 | 0.32649 | 0.0022 | 48.4543 | 0.048743 | 0.8816 | 0.042972 | 18.0867 | 22.2506 | 3.4702 | 0.045172 | 0.3608 | 0.3408 | 0.2406 | 0.1123
2012-10-30 22:23 | 26.4834 | 22.6213 | 22.5932 | 22.489 | 22.3888 | 50.3397 | 0.006747 | 0.32637 | 0.002202 | 48.3726 | 0.04878 | 0.88229 | 0.043038 | 18.0872 | 22.2163 | 3.4704 | 0.04524 | 0.405 | 0.3769 | 0.2727 | 0.1725
2012-10-30 22:23 | 26.4856 | 22.5994 | 22.5834 | 22.4832 | 22.3589 | 50.3351 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048714 | 0.88121 | 0.042927 | 18.0894 | 22.2386 | 3.4701 | 0.045121 | 0.3608 | 0.3448 | 0.2446 | 0.1203
2012-10-30 22:23 | 26.4819 | 22.6037 | 22.5837 | 22.4514 | 22.3753 | 50.3259 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048771 | 0.88216 | 0.043024 | 18.0878 | 22.2911 | 3.4695 | 0.045221 | 0.3126 | 0.2926 | 0.1603 | 0.0842
2012-10-30 22:23 | 26.4931 | 22.6189 | 22.6109 | 22.4827 | 22.3744 | 50.3351 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048837 | 0.88307 | 0.043127 | 18.082 | 22.2581 | 3.4701 | 0.045324 | 0.3608 | 0.3528 | 0.2246 | 0.1163
2012-10-30 22:23 | 26.4766 | 22.5864 | 22.5704 | 22.5022 | 22.394 | 50.3391 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.048733 | 0.88171 | 0.042969 | 18.0927 | 22.2537 | 3.4704 | 0.045167 | 0.3327 | 0.3167 | 0.2485 | 0.1403
2012-10-30 22:23 | 26.4796 | 22.6095 | 22.6095 | 22.4531 | 22.3329 | 50.3276 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048793 | 0.88232 | 0.043051 | 18.0829 | 22.2647 | 3.4696 | 0.045248 | 0.3448 | 0.3448 | 0.1884 | 0.0682
2012-10-30 22:23 | 26.4894 | 22.5792 | 22.6032 | 22.487 | 22.3627 | 50.3345 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048765 | 0.88203 | 0.043012 | 18.0874 | 22.2665 | 3.4701 | 0.045204 | 0.3127 | 0.3367 | 0.2205 | 0.0962
2012-10-30 22:23 | 26.4916 | 22.5854 | 22.5894 | 22.4772 | 22.3529 | 50.342 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048799 | 0.8827 | 0.043075 | 18.0884 | 22.2446 | 3.4706 | 0.04527 | 0.3408 | 0.3448 | 0.2326 | 0.1083
2012-10-30 22:23 | 26.5006 | 22.6144 | 22.6024 | 22.5062 | 22.3859 | 50.3334 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048727 | 0.88147 | 0.042951 | 18.0899 | 22.2656 | 3.47 | 0.045146 | 0.3488 | 0.3368 | 0.2406 | 0.1203
2012-10-30 22:23 | 26.4976 | 22.5794 | 22.5633 | 22.4711 | 22.403 | 50.3345 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048749 | 0.88192 | 0.042993 | 18.091 | 22.2787 | 3.4701 | 0.045187 | 0.3007 | 0.2846 | 0.1924 | 0.1243
2012-10-30 22:23 | 26.5006 | 22.6184 | 22.5984 | 22.4982 | 22.3779 | 50.3368 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048809 | 0.88281 | 0.043089 | 18.0871 | 22.2456 | 3.4702 | 0.045287 | 0.3728 | 0.3528 | 0.2526 | 0.1323
2012-10-30 22:23 | 26.4976 | 22.5874 | 22.5794 | 22.4751 | 22.3549 | 50.3363 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048837 | 0.8831 | 0.043128 | 18.0826 | 22.2426 | 3.4702 | 0.045324 | 0.3448 | 0.3368 | 0.2325 | 0.1123
2012-10-30 22:23 | 26.4998 | 22.6417 | 22.6256 | 22.4934 | 22.3571 | 50.3339 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048777 | 0.88235 | 0.043038 | 18.0894 | 22.2327 | 3.47 | 0.045236 | 0.409 | 0.3929 | 0.2607 | 0.1244
2012-10-30 22:23 | 26.4968 | 22.6026 | 22.6026 | 22.4784 | 22.33 | 50.3432 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048837 | 0.88313 | 0.043129 | 18.0832 | 22.2017 | 3.4707 | 0.045329 | 0.4009 | 0.4009 | 0.2767 | 0.1283
2012-10-30 22:23 | 26.5021 | 22.6079 | 22.5598 | 22.4636 | 22.3674 | 50.3235 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048765 | 0.88199 | 0.04301 | 18.0865 | 22.2631 | 3.4693 | 0.045209 | 0.3448 | 0.2967 | 0.2005 | 0.1043
2012-10-30 22:23 | 26.5066 | 22.6124 | 22.5803 | 22.4841 | 22.3959 | 50.338 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.04879 | 0.88252 | 0.043058 | 18.0882 | 22.2436 | 3.4703 | 0.045254 | 0.3688 | 0.3367 | 0.2405 | 0.1523
2012-10-30 22:23 | 26.5096 | 22.5633 | 22.5833 | 22.4871 | 22.423 | 50.3287 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048771 | 0.88226 | 0.043029 | 18.0899 | 22.2466 | 3.4697 | 0.045224 | 0.3167 | 0.3367 | 0.2405 | 0.1764
2012-10-30 22:23 | 26.5156 | 22.5613 | 22.5573 | 22.4571 | 22.421 | 50.3328 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048799 | 0.88248 | 0.043064 | 18.084 | 22.2806 | 3.4699 | 0.045263 | 0.2807 | 0.2767 | 0.1765 | 0.1404
2012-10-30 22:23 | 26.5036 | 22.5373 | 22.5773 | 22.4811 | 22.3729 | 50.3218 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048708 | 0.88132 | 0.042928 | 18.094 | 22.2245 | 3.4692 | 0.045123 | 0.3128 | 0.3528 | 0.2566 | 0.1484
2012-10-30 22:23 | 26.5163 | 22.5299 | 22.566 | 22.4898 | 22.3816 | 50.3299 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048721 | 0.88166 | 0.042955 | 18.096 | 22.2172 | 3.4697 | 0.045152 | 0.3127 | 0.3488 | 0.2726 | 0.1644
2012-10-30 22:23 | 26.5141 | 22.5718 | 22.5558 | 22.4556 | 22.3593 | 50.3224 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048768 | 0.88218 | 0.043022 | 18.0892 | 22.243 | 3.4692 | 0.045219 | 0.3288 | 0.3128 | 0.2126 | 0.1163
2012-10-30 22:23 | 26.5193 | 22.565 | 22.5971 | 22.4888 | 22.3846 | 50.3183 | 0.006741 | 0.32643 | 0.0022 | 48.4241 | 0.048815 | 0.88307 | 0.043107 | 18.0902 | 22.2081 | 3.4689 | 0.045307 | 0.3569 | 0.389 | 0.2807 | 0.1765
2012-10-30 22:23 | 26.5156 | 22.5813 | 22.6054 | 22.4931 | 22.3448 | 50.3212 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.048825 | 0.88309 | 0.043117 | 18.0868 | 22.2084 | 3.4691 | 0.045312 | 0.3729 | 0.397 | 0.2847 | 0.1364
2012-10-30 22:23 | 26.5208 | 22.5304 | 22.5785 | 22.5144 | 22.3941 | 50.3259 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048771 | 0.88216 | 0.043024 | 18.0878 | 22.2096 | 3.4695 | 0.045217 | 0.3208 | 0.3689 | 0.3048 | 0.1845
2012-10-30 22:23 | 26.5215 | 22.5872 | 22.5632 | 22.475 | 22.3908 | 50.327 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.048768 | 0.88229 | 0.043028 | 18.0916 | 22.2224 | 3.4695 | 0.045226 | 0.3648 | 0.3408 | 0.2526 | 0.1684
2012-10-30 22:24 | 26.5223 | 22.576 | 22.56 | 22.4638 | 22.3716 | 50.3374 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048812 | 0.88261 | 0.043082 | 18.0818 | 22.175 | 3.4703 | 0.045278 | 0.401 | 0.385 | 0.2888 | 0.1966
2012-10-30 22:24 | 26.529 | 22.5747 | 22.5747 | 22.4945 | 22.3662 | 50.3235 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.048818 | 0.88299 | 0.043106 | 18.0873 | 22.1978 | 3.4693 | 0.045307 | 0.3769 | 0.3769 | 0.2967 | 0.1684
2012-10-30 22:24 | 26.532 | 22.5857 | 22.5737 | 22.4855 | 22.3492 | 50.3207 | 0.006741 | 0.32637 | 0.0022 | 48.4157 | 0.048796 | 0.88248 | 0.043061 | 18.0851 | 22.2248 | 3.4691 | 0.045261 | 0.3609 | 0.3489 | 0.2607 | 0.1244
2012-10-30 22:24 | 26.52 | 22.5176 | 22.5497 | 22.4735 | 22.3612 | 50.3322 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.04884 | 0.88314 | 0.043133 | 18.0824 | 22.2209 | 3.4699 | 0.045328 | 0.2967 | 0.3288 | 0.2526 | 0.1403
2012-10-30 22:24 | 26.5358 | 22.6095 | 22.5735 | 22.5093 | 22.3409 | 50.3287 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.048787 | 0.88235 | 0.043047 | 18.0857 | 22.1564 | 3.4697 | 0.045243 | 0.4531 | 0.4171 | 0.3529 | 0.1845
2012-10-30 22:24 | 26.5268 | 22.5725 | 22.5524 | 22.4683 | 22.4041 | 50.3149 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048787 | 0.88241 | 0.04305 | 18.0869 | 22.1875 | 3.4687 | 0.045245 | 0.385 | 0.3649 | 0.2808 | 0.2166
2012-10-30 22:24 | 26.526 | 22.6118 | 22.5997 | 22.4915 | 22.3793 | 50.3287 | 0.006741 | 0.32633 | 0.0022 | 48.4091 | 0.04885 | 0.88335 | 0.043151 | 18.0828 | 22.1707 | 3.4697 | 0.045351 | 0.4411 | 0.429 | 0.3208 | 0.2086
2012-10-30 22:24 | 26.5245 | 22.5942 | 22.5421 | 22.474 | 22.3858 | 50.3218 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.048793 | 0.88247 | 0.043058 | 18.0859 | 22.2173 | 3.4692 | 0.045252 | 0.3769 | 0.3248 | 0.2567 | 0.1685
2012-10-30 22:24 | 26.5305 | 22.5842 | 22.5922 | 22.496 | 22.3677 | 50.327 | 0.006744 | 0.32636 | 0.002201 | 48.3919 | 0.048818 | 0.88288 | 0.043101 | 18.0852 | 22.2073 | 3.4695 | 0.045302 | 0.3769 | 0.3849 | 0.2887 | 0.1604
2012-10-30 22:24 | 26.529 | 22.5747 | 22.5587 | 22.5106 | 22.3702 | 50.3247 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048758 | 0.88203 | 0.043006 | 18.09 | 22.2178 | 3.4694 | 0.045201 | 0.3569 | 0.3409 | 0.2928 | 0.1524
2012-10-30 22:24 | 26.5171 | 22.5908 | 22.5708 | 22.4706 | 22.3463 | 50.323 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048853 | 0.88355 | 0.043164 | 18.0859 | 22.2139 | 3.4693 | 0.045359 | 0.3769 | 0.3569 | 0.2567 | 0.1324
2012-10-30 22:24 | 26.5223 | 22.584 | 22.56 | 22.4918 | 22.3635 | 50.3282 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048796 | 0.88254 | 0.043064 | 18.0863 | 22.2071 | 3.4696 | 0.045263 | 0.3769 | 0.3529 | 0.2847 | 0.1564
2012-10-30 22:24 | 26.5245 | 22.5622 | 22.5582 | 22.486 | 22.3617 | 50.3299 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048906 | 0.88436 | 0.04325 | 18.0828 | 22.1491 | 3.4697 | 0.04545 | 0.4131 | 0.4091 | 0.3369 | 0.2126
2012-10-30 22:24 | 26.5298 | 22.5875 | 22.5514 | 22.4713 | 22.355 | 50.3247 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048843 | 0.88333 | 0.043145 | 18.0851 | 22.2066 | 3.4694 | 0.045344 | 0.3809 | 0.3448 | 0.2647 | 0.1484
2012-10-30 22:24 | 26.5223 | 22.576 | 22.552 | 22.4638 | 22.3595 | 50.3224 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048856 | 0.88353 | 0.043166 | 18.0845 | 22.2111 | 3.4692 | 0.045362 | 0.3649 | 0.3409 | 0.2527 | 0.1484
2012-10-30 22:24 | 26.523 | 22.5607 | 22.5446 | 22.4645 | 22.3562 | 50.3149 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.04885 | 0.8832 | 0.043144 | 18.0799 | 22.1958 | 3.4687 | 0.04534 | 0.3649 | 0.3488 | 0.2687 | 0.1604
2012-10-30 22:24 | 26.5245 | 22.5542 | 22.5782 | 22.466 | 22.3497 | 50.323 | 0.006751 | 0.32641 | 0.002204 | 48.3503 | 0.048837 | 0.88323 | 0.043134 | 18.0853 | 22.1973 | 3.4693 | 0.045338 | 0.3569 | 0.3809 | 0.2687 | 0.1524
2012-10-30 22:24 | 26.5238 | 22.5735 | 22.5735 | 22.5174 | 22.333 | 50.3183 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.04885 | 0.88342 | 0.043155 | 18.0843 | 22.1845 | 3.4689 | 0.045352 | 0.389 | 0.389 | 0.3329 | 0.1485
2012-10-30 22:24 | 26.5223 | 22.58 | 22.56 | 22.4598 | 22.3355 | 50.3143 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.04884 | 0.88348 | 0.043149 | 18.0892 | 22.2272 | 3.4687 | 0.045346 | 0.3528 | 0.3328 | 0.2326 | 0.1083
2012-10-30 22:24 | 26.5298 | 22.5675 | 22.5755 | 22.4913 | 22.347 | 50.3143 | 0.006738 | 0.32631 | 0.002199 | 48.4286 | 0.04879 | 0.88249 | 0.043057 | 18.0876 | 22.2427 | 3.4687 | 0.045256 | 0.3248 | 0.3328 | 0.2486 | 0.1043
2012-10-30 22:24 | 26.5298 | 22.5955 | 22.5715 | 22.4592 | 22.3791 | 50.3212 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.048771 | 0.88255 | 0.043043 | 18.0958 | 22.1946 | 3.4691 | 0.045237 | 0.4009 | 0.3769 | 0.2646 | 0.1845
2012-10-30 22:24 | 26.532 | 22.5777 | 22.5817 | 22.4815 | 22.3532 | 50.3085 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.04878 | 0.88223 | 0.043035 | 18.086 | 22.2288 | 3.4683 | 0.045231 | 0.3489 | 0.3529 | 0.2527 | 0.1244
2012-10-30 22:24 | 26.5223 | 22.588 | 22.5319 | 22.4558 | 22.3475 | 50.3328 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.048812 | 0.8826 | 0.043081 | 18.0815 | 22.2071 | 3.4699 | 0.045275 | 0.3809 | 0.3248 | 0.2487 | 0.1404
2012-10-30 22:24 | 26.5193 | 22.6171 | 22.573 | 22.4528 | 22.3686 | 50.3207 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048891 | 0.88407 | 0.043223 | 18.0824 | 22.2121 | 3.4691 | 0.045418 | 0.405 | 0.3609 | 0.2407 | 0.1565
2012-10-30 22:25 | 26.5208 | 22.5505 | 22.5945 | 22.4943 | 22.358 | 50.3178 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.048802 | 0.88275 | 0.04308 | 18.0885 | 22.2217 | 3.4689 | 0.045277 | 0.3288 | 0.3728 | 0.2726 | 0.1363
2012-10-30 22:25 | 26.5245 | 22.5902 | 22.5862 | 22.5341 | 22.3697 | 50.3178 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048856 | 0.88348 | 0.043163 | 18.0833 | 22.1852 | 3.4689 | 0.045358 | 0.405 | 0.401 | 0.3489 | 0.1845
2012-10-30 22:25 | 26.5335 | 22.5712 | 22.5792 | 22.483 | 22.3587 | 50.316 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048831 | 0.88329 | 0.043132 | 18.0887 | 22.1983 | 3.4688 | 0.045331 | 0.3729 | 0.3809 | 0.2847 | 0.1604
2012-10-30 22:25 | 26.523 | 22.5847 | 22.5527 | 22.4845 | 22.3803 | 50.3253 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048847 | 0.88342 | 0.043152 | 18.0854 | 22.2118 | 3.4694 | 0.045347 | 0.3729 | 0.3409 | 0.2727 | 0.1685
2012-10-30 22:25 | 26.5283 | 22.586 | 22.562 | 22.4577 | 22.4016 | 50.3074 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048806 | 0.88258 | 0.043075 | 18.0835 | 22.2131 | 3.4682 | 0.045269 | 0.3729 | 0.3489 | 0.2446 | 0.1885
2012-10-30 22:25 | 26.5275 | 22.6213 | 22.5331 | 22.4569 | 22.3326 | 50.3062 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048862 | 0.88352 | 0.043171 | 18.0819 | 22.1802 | 3.4681 | 0.045369 | 0.4411 | 0.3529 | 0.2767 | 0.1524
2012-10-30 22:25 | 26.5343 | 22.592 | 22.5599 | 22.4998 | 22.3715 | 50.3154 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048834 | 0.88339 | 0.043139 | 18.0897 | 22.183 | 3.4687 | 0.045336 | 0.409 | 0.3769 | 0.3168 | 0.1885
2012-10-30 22:25 | 26.5305 | 22.5561 | 22.5682 | 22.464 | 22.3557 | 50.3253 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.04884 | 0.88316 | 0.043133 | 18.0827 | 22.1872 | 3.4694 | 0.045326 | 0.3689 | 0.381 | 0.2768 | 0.1685
2012-10-30 22:25 | 26.5208 | 22.5986 | 22.5585 | 22.4663 | 22.358 | 50.3149 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048812 | 0.88285 | 0.043094 | 18.0868 | 22.2257 | 3.4687 | 0.045289 | 0.3729 | 0.3328 | 0.2406 | 0.1323
2012-10-30 22:25 | 26.5238 | 22.6056 | 22.5494 | 22.4853 | 22.341 | 50.3178 | 0.006744 | 0.32638 | 0.002201 | 48.3962 | 0.048847 | 0.88336 | 0.04315 | 18.0842 | 22.1685 | 3.4689 | 0.045351 | 0.4371 | 0.3809 | 0.3168 | 0.1725
2012-10-30 22:25 | 26.529 | 22.6068 | 22.5827 | 22.4825 | 22.3221 | 50.3102 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.048919 | 0.88444 | 0.043266 | 18.0798 | 22.2178 | 3.4684 | 0.04546 | 0.389 | 0.3649 | 0.2647 | 0.1043
2012-10-30 22:25 | 26.526 | 22.5717 | 22.5396 | 22.4595 | 22.3552 | 50.3299 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048991 | 0.88572 | 0.043392 | 18.0792 | 22.2228 | 3.4697 | 0.045587 | 0.3489 | 0.3168 | 0.2367 | 0.1324
2012-10-30 22:25 | 26.5245 | 22.5822 | 22.5662 | 22.49 | 22.3337 | 50.3097 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.048938 | 0.88475 | 0.043298 | 18.079 | 22.1973 | 3.4684 | 0.045495 | 0.3849 | 0.3689 | 0.2927 | 0.1364
2012-10-30 22:25 | 26.5268 | 22.6206 | 22.5444 | 22.5043 | 22.356 | 50.3241 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.048988 | 0.88563 | 0.043385 | 18.0785 | 22.2036 | 3.4693 | 0.045584 | 0.417 | 0.3408 | 0.3007 | 0.1524
2012-10-30 22:25 | 26.523 | 22.6008 | 22.5767 | 22.4845 | 22.3161 | 50.3143 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.049016 | 0.88624 | 0.04344 | 18.0805 | 22.2279 | 3.4687 | 0.045637 | 0.3729 | 0.3488 | 0.2566 | 0.0882
2012-10-30 22:25 | 26.5268 | 22.5725 | 22.5364 | 22.4482 | 22.3801 | 50.3224 | 0.006747 | 0.32638 | 0.002202 | 48.3747 | 0.049001 | 0.88608 | 0.043419 | 18.0828 | 22.2277 | 3.4692 | 0.045621 | 0.3448 | 0.3087 | 0.2205 | 0.1524
2012-10-30 22:25 | 26.5178 | 22.5875 | 22.5675 | 22.4873 | 22.331 | 50.3235 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048887 | 0.8841 | 0.043221 | 18.0845 | 22.2467 | 3.4693 | 0.045416 | 0.3408 | 0.3208 | 0.2406 | 0.0843
2012-10-30 22:25 | 26.5245 | 22.5782 | 22.5582 | 22.474 | 22.3457 | 50.3126 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.049061 | 0.88696 | 0.043515 | 18.0787 | 22.2133 | 3.4686 | 0.045709 | 0.3649 | 0.3449 | 0.2607 | 0.1324
2012-10-30 22:25 | 26.5208 | 22.5865 | 22.5745 | 22.4543 | 22.334 | 50.3212 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.048988 | 0.88577 | 0.043392 | 18.0814 | 22.2096 | 3.4691 | 0.045587 | 0.3769 | 0.3649 | 0.2447 | 0.1244
2012-10-30 22:25 | 26.5268 | 22.5805 | 22.5845 | 22.4843 | 22.372 | 50.3143 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048847 | 0.88319 | 0.043141 | 18.0807 | 22.2116 | 3.4687 | 0.045337 | 0.3689 | 0.3729 | 0.2727 | 0.1604
2012-10-30 22:25 | 26.523 | 22.5967 | 22.5326 | 22.4965 | 22.3682 | 50.316 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.049054 | 0.88658 | 0.04349 | 18.0736 | 22.272 | 3.4688 | 0.045688 | 0.3247 | 0.2606 | 0.2245 | 0.0962
2012-10-30 22:25 | 26.5253 | 22.575 | 22.5469 | 22.4788 | 22.3986 | 50.3218 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.049061 | 0.88677 | 0.043506 | 18.0749 | 22.2502 | 3.4692 | 0.045701 | 0.3248 | 0.2967 | 0.2286 | 0.1484
2012-10-30 22:25 | 26.5208 | 22.5585 | 22.5344 | 22.4583 | 22.33 | 50.3091 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.049032 | 0.88647 | 0.043465 | 18.0794 | 22.2257 | 3.4683 | 0.045661 | 0.3328 | 0.3087 | 0.2326 | 0.1043
2012-10-30 22:25 | 26.5238 | 22.6056 | 22.5695 | 22.4492 | 22.365 | 50.3154 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.04907 | 0.887 | 0.043525 | 18.0763 | 22.2527 | 3.4687 | 0.045722 | 0.3529 | 0.3168 | 0.1965 | 0.1123
2012-10-30 22:26 | 26.5223 | 22.588 | 22.556 | 22.4517 | 22.3595 | 50.3114 | 0.006741 | 0.32641 | 0.0022 | 48.422 | 0.049054 | 0.88657 | 0.04349 | 18.0733 | 22.2151 | 3.4685 | 0.04569 | 0.3729 | 0.3409 | 0.2366 | 0.1444
2012-10-30 22:26 | 26.5171 | 22.5868 | 22.5548 | 22.4666 | 22.3583 | 50.297 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.04878 | 0.88244 | 0.043045 | 18.0901 | 22.246 | 3.4675 | 0.045244 | 0.3408 | 0.3088 | 0.2206 | 0.1123
2012-10-30 22:26 | 26.5171 | 22.5788 | 22.5748 | 22.5027 | 22.3784 | 50.3149 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.048793 | 0.8826 | 0.043064 | 18.0886 | 22.2059 | 3.4687 | 0.045256 | 0.3729 | 0.3689 | 0.2968 | 0.1725
2012-10-30 22:26 | 26.5193 | 22.6451 | 22.561 | 22.4968 | 22.3285 | 50.3108 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048894 | 0.88429 | 0.043236 | 18.0858 | 22.2081 | 3.4684 | 0.04543 | 0.437 | 0.3529 | 0.2887 | 0.1204
2012-10-30 22:26 | 26.5208 | 22.5745 | 22.5464 | 22.4783 | 22.35 | 50.3039 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.049057 | 0.88681 | 0.043504 | 18.0772 | 22.2417 | 3.468 | 0.045698 | 0.3328 | 0.3047 | 0.2366 | 0.1083
2012-10-30 22:26 | 26.5163 | 22.562 | 22.554 | 22.4538 | 22.3375 | 50.3149 | 0.006716 | 0.32643 | 0.002192 | 48.6044 | 0.048853 | 0.8834 | 0.043157 | 18.0829 | 22.2412 | 3.4687 | 0.045349 | 0.3208 | 0.3128 | 0.2126 | 0.0963
2012-10-30 22:26 | 26.5238 | 22.5655 | 22.5975 | 22.4613 | 22.345 | 50.2987 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048944 | 0.88459 | 0.043295 | 18.0735 | 22.2287 | 3.4676 | 0.045491 | 0.3368 | 0.3688 | 0.2326 | 0.1163
2012-10-30 22:26 | 26.5171 | 22.5748 | 22.5227 | 22.4626 | 22.3142 | 50.3062 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.048831 | 0.88301 | 0.043118 | 18.0831 | 22.1778 | 3.4681 | 0.045314 | 0.397 | 0.3449 | 0.2848 | 0.1364
2012-10-30 22:26 | 26.5215 | 22.5872 | 22.5391 | 22.463 | 22.3186 | 50.3062 | 0.006747 | 0.32638 | 0.002202 | 48.3747 | 0.048667 | 0.88044 | 0.042848 | 18.0911 | 22.2424 | 3.4681 | 0.04505 | 0.3448 | 0.2967 | 0.2206 | 0.0762
2012-10-30 22:26 | 26.5245 | 22.5381 | 22.5381 | 22.486 | 22.3417 | 50.3114 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.049039 | 0.88629 | 0.043463 | 18.0732 | 22.2334 | 3.4685 | 0.045662 | 0.3047 | 0.3047 | 0.2526 | 0.1083
2012-10-30 22:26 | 26.5186 | 22.5723 | 22.5643 | 22.4641 | 22.3679 | 50.3131 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048944 | 0.88472 | 0.043302 | 18.0761 | 22.2195 | 3.4686 | 0.045497 | 0.3528 | 0.3448 | 0.2446 | 0.1484
2012-10-30 22:26 | 26.5208 | 22.5705 | 22.5545 | 22.4663 | 22.3099 | 50.3114 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048869 | 0.88366 | 0.043184 | 18.0823 | 22.1896 | 3.4685 | 0.045383 | 0.3809 | 0.3649 | 0.2767 | 0.1203
2012-10-30 22:26 | 26.5268 | 22.5645 | 22.5565 | 22.4482 | 22.348 | 50.3068 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048944 | 0.88484 | 0.043307 | 18.0785 | 22.2156 | 3.4682 | 0.045501 | 0.3489 | 0.3409 | 0.2326 | 0.1324
2012-10-30 22:26 | 26.5186 | 22.5763 | 22.5803 | 22.4521 | 22.3318 | 50.3079 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048755 | 0.88212 | 0.043008 | 18.0929 | 22.1914 | 3.4682 | 0.045201 | 0.3849 | 0.3889 | 0.2607 | 0.1404
2012-10-30 22:26 | 26.5245 | 22.6263 | 22.5301 | 22.462 | 22.3296 | 50.2993 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048982 | 0.88566 | 0.043381 | 18.0813 | 22.2013 | 3.4676 | 0.04558 | 0.425 | 0.3288 | 0.2607 | 0.1283
2012-10-30 22:26 | 26.5171 | 22.5387 | 22.5588 | 22.4506 | 22.3343 | 50.2987 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048784 | 0.8821 | 0.043033 | 18.0818 | 22.1979 | 3.4676 | 0.045232 | 0.3408 | 0.3609 | 0.2527 | 0.1364
2012-10-30 22:26 | 26.5178 | 22.6316 | 22.5354 | 22.4673 | 22.339 | 50.2929 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048821 | 0.88297 | 0.043108 | 18.0859 | 22.1866 | 3.4672 | 0.045304 | 0.445 | 0.3488 | 0.2807 | 0.1524
2012-10-30 22:26 | 26.5118 | 22.5495 | 22.5495 | 22.4292 | 22.321 | 50.2935 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048938 | 0.88466 | 0.043294 | 18.0772 | 22.1926 | 3.4672 | 0.045491 | 0.3569 | 0.3569 | 0.2366 | 0.1284
2012-10-30 22:26 | 26.5186 | 22.6004 | 22.5483 | 22.4881 | 22.3077 | 50.3085 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.049061 | 0.88668 | 0.043502 | 18.0731 | 22.1793 | 3.4683 | 0.045699 | 0.4211 | 0.369 | 0.3088 | 0.1284
2012-10-30 22:26 | 26.5118 | 22.5455 | 22.5415 | 22.4493 | 22.349 | 50.2981 | 0.006719 | 0.32644 | 0.002193 | 48.5849 | 0.04885 | 0.88316 | 0.043142 | 18.079 | 22.1926 | 3.4676 | 0.045335 | 0.3529 | 0.3489 | 0.2567 | 0.1564
2012-10-30 22:26 | 26.5215 | 22.5832 | 22.5471 | 22.467 | 22.3547 | 50.305 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048784 | 0.88226 | 0.04304 | 18.0851 | 22.1822 | 3.468 | 0.045238 | 0.401 | 0.3649 | 0.2848 | 0.1725
2012-10-30 22:26 | 26.5171 | 22.5868 | 22.5588 | 22.4506 | 22.3383 | 50.3045 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.04901 | 0.88612 | 0.043429 | 18.0804 | 22.1819 | 3.468 | 0.045627 | 0.4049 | 0.3769 | 0.2687 | 0.1564
2012-10-30 22:27 | 26.5111 | 22.5688 | 22.5448 | 22.4686 | 22.3363 | 50.3027 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048932 | 0.88481 | 0.043295 | 18.0824 | 22.2079 | 3.4679 | 0.04549 | 0.3609 | 0.3369 | 0.2607 | 0.1284
2012-10-30 22:27 | 26.5118 | 22.5334 | 22.5254 | 22.4693 | 22.345 | 50.2987 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.049174 | 0.88858 | 0.043695 | 18.0701 | 22.2006 | 3.4676 | 0.045887 | 0.3328 | 0.3248 | 0.2687 | 0.1444
2012-10-30 22:27 | 26.5141 | 22.5798 | 22.5478 | 22.4676 | 22.3834 | 50.3022 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.048922 | 0.88463 | 0.043278 | 18.0825 | 22.215 | 3.4678 | 0.045476 | 0.3648 | 0.3328 | 0.2526 | 0.1684
2012-10-30 22:27 | 26.5126 | 22.5944 | 22.5823 | 22.43 | 22.3458 | 50.2993 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048881 | 0.88371 | 0.043197 | 18.0788 | 22.1814 | 3.4676 | 0.045391 | 0.413 | 0.4009 | 0.2486 | 0.1644
2012-10-30 22:27 | 26.5133 | 22.587 | 22.551 | 22.4468 | 22.3946 | 50.301 | 0.00671 | 0.32637 | 0.00219 | 48.6393 | 0.049123 | 0.88791 | 0.043617 | 18.0753 | 22.2142 | 3.4678 | 0.045807 | 0.3728 | 0.3368 | 0.2326 | 0.1804
2012-10-30 22:27 | 26.5156 | 22.5573 | 22.5573 | 22.413 | 22.3408 | 50.2993 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048988 | 0.88538 | 0.043373 | 18.0735 | 22.1763 | 3.4676 | 0.045568 | 0.381 | 0.381 | 0.2367 | 0.1645
2012-10-30 22:27 | 26.5126 | 22.5262 | 22.5262 | 22.4701 | 22.3137 | 50.2975 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048865 | 0.88369 | 0.043182 | 18.0844 | 22.1653 | 3.4675 | 0.045375 | 0.3609 | 0.3609 | 0.3048 | 0.1484
2012-10-30 22:27 | 26.5111 | 22.5728 | 22.5367 | 22.4846 | 22.3483 | 50.3039 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.048891 | 0.88379 | 0.04321 | 18.0768 | 22.216 | 3.468 | 0.045402 | 0.3568 | 0.3207 | 0.2686 | 0.1323
2012-10-30 22:27 | 26.5163 | 22.57 | 22.562 | 22.4297 | 22.3134 | 50.3074 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.049101 | 0.88755 | 0.04358 | 18.076 | 22.2292 | 3.4682 | 0.045777 | 0.3408 | 0.3328 | 0.2005 | 0.0842
2012-10-30 22:27 | 26.5133 | 22.5389 | 22.571 | 22.4668 | 22.3345 | 50.3114 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048903 | 0.88423 | 0.043241 | 18.0813 | 22.1821 | 3.4685 | 0.045436 | 0.3568 | 0.3889 | 0.2847 | 0.1524
2012-10-30 22:27 | 26.5141 | 22.5718 | 22.5438 | 22.4676 | 22.3192 | 50.297 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048928 | 0.88484 | 0.043293 | 18.0844 | 22.215 | 3.4675 | 0.045491 | 0.3568 | 0.3288 | 0.2526 | 0.1042
2012-10-30 22:27 | 26.5111 | 22.5327 | 22.5488 | 22.4365 | 22.2882 | 50.3016 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048771 | 0.88219 | 0.043025 | 18.0884 | 22.228 | 3.4678 | 0.045221 | 0.3047 | 0.3208 | 0.2085 | 0.0602
2012-10-30 22:27 | 26.5171 | 22.5949 | 22.5427 | 22.4506 | 22.3583 | 50.2941 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048994 | 0.88567 | 0.043393 | 18.0772 | 22.226 | 3.4673 | 0.04559 | 0.3689 | 0.3167 | 0.2246 | 0.1323
2012-10-30 22:27 | 26.5133 | 22.6191 | 22.4989 | 22.4668 | 22.3425 | 50.2923 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.049042 | 0.88664 | 0.043483 | 18.0792 | 22.2222 | 3.4672 | 0.045677 | 0.3969 | 0.2767 | 0.2446 | 0.1203
2012-10-30 22:27 | 26.5126 | 22.5543 | 22.5382 | 22.434 | 22.3378 | 50.3108 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.049029 | 0.88642 | 0.04346 | 18.0796 | 22.2496 | 3.4684 | 0.045652 | 0.3047 | 0.2886 | 0.1844 | 0.0882
2012-10-30 22:27 | 26.5178 | 22.5515 | 22.5274 | 22.4432 | 22.339 | 50.2975 | 0.006722 | 0.32624 | 0.002193 | 48.5332 | 0.04895 | 0.88502 | 0.043322 | 18.0801 | 22.2066 | 3.4675 | 0.045515 | 0.3449 | 0.3208 | 0.2366 | 0.1324
2012-10-30 22:27 | 26.5141 | 22.5959 | 22.5237 | 22.4355 | 22.2791 | 50.2946 | 0.006722 | 0.3265 | 0.002195 | 48.5719 | 0.049061 | 0.88674 | 0.043504 | 18.0743 | 22.223 | 3.4673 | 0.045699 | 0.3729 | 0.3007 | 0.2125 | 0.0561
2012-10-30 22:27 | 26.5126 | 22.5663 | 22.5463 | 22.43 | 22.3659 | 50.2889 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048828 | 0.88338 | 0.043133 | 18.0916 | 22.1974 | 3.4669 | 0.045328 | 0.3689 | 0.3489 | 0.2326 | 0.1685
2012-10-30 22:27 | 26.5171 | 22.5868 | 22.5508 | 22.4706 | 22.3182 | 50.2912 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048994 | 0.88567 | 0.043393 | 18.0772 | 22.2019 | 3.4671 | 0.045588 | 0.3849 | 0.3489 | 0.2687 | 0.1163
2012-10-30 22:27 | 26.5081 | 22.6019 | 22.5217 | 22.4416 | 22.3413 | 50.2866 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048802 | 0.88275 | 0.04308 | 18.0885 | 22.1969 | 3.4668 | 0.045277 | 0.405 | 0.3248 | 0.2447 | 0.1444
2012-10-30 22:28 | 26.5178 | 22.5956 | 22.5555 | 22.4352 | 22.331 | 50.2935 | 0.006744 | 0.32634 | 0.002201 | 48.3898 | 0.048812 | 0.88287 | 0.043095 | 18.0872 | 22.2026 | 3.4672 | 0.045296 | 0.393 | 0.3529 | 0.2326 | 0.1284
2012-10-30 22:28 | 26.5111 | 22.5969 | 22.5488 | 22.4606 | 22.3082 | 50.2946 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.049054 | 0.88661 | 0.043492 | 18.0742 | 22.1999 | 3.4673 | 0.045687 | 0.397 | 0.3489 | 0.2607 | 0.1083
2012-10-30 22:28 | 26.4991 | 22.5167 | 22.5608 | 22.4165 | 22.2922 | 50.2918 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048825 | 0.88291 | 0.043108 | 18.0832 | 22.204 | 3.4671 | 0.045306 | 0.3127 | 0.3568 | 0.2125 | 0.0882
2012-10-30 22:28 | 26.5148 | 22.5805 | 22.5244 | 22.3921 | 22.324 | 50.2946 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.0489 | 0.8843 | 0.043242 | 18.0838 | 22.2397 | 3.4673 | 0.045441 | 0.3408 | 0.2847 | 0.1524 | 0.0843
2012-10-30 22:28 | 26.5081 | 22.5738 | 22.5658 | 22.4776 | 22.3213 | 50.2923 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048878 | 0.88375 | 0.043196 | 18.0808 | 22.241 | 3.4672 | 0.045393 | 0.3328 | 0.3248 | 0.2366 | 0.0803
2012-10-30 22:28 | 26.5148 | 22.5445 | 22.5404 | 22.4402 | 22.34 | 50.2894 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048806 | 0.88291 | 0.043091 | 18.0903 | 22.2157 | 3.467 | 0.045284 | 0.3288 | 0.3247 | 0.2245 | 0.1243
2012-10-30 22:28 | 26.5066 | 22.5563 | 22.5282 | 22.4481 | 22.3198 | 50.2987 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.049029 | 0.88616 | 0.043448 | 18.0743 | 22.1994 | 3.4676 | 0.045641 | 0.3569 | 0.3288 | 0.2487 | 0.1204
2012-10-30 22:28 | 26.5118 | 22.5815 | 22.5254 | 22.4653 | 22.3771 | 50.2883 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048806 | 0.88299 | 0.043095 | 18.0917 | 22.2046 | 3.4669 | 0.045288 | 0.3769 | 0.3208 | 0.2607 | 0.1725
2012-10-30 22:28 | 26.5066 | 22.5563 | 22.5443 | 22.4441 | 22.3478 | 50.297 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.048793 | 0.8826 | 0.043064 | 18.0886 | 22.1954 | 3.4675 | 0.045256 | 0.3609 | 0.3489 | 0.2487 | 0.1524
2012-10-30 22:28 | 26.5073 | 22.5851 | 22.557 | 22.4367 | 22.3165 | 50.2958 | 0.006722 | 0.32624 | 0.002193 | 48.5332 | 0.048909 | 0.88436 | 0.043253 | 18.0817 | 22.2162 | 3.4674 | 0.045446 | 0.3689 | 0.3408 | 0.2205 | 0.1003
2012-10-30 22:28 | 26.5006 | 22.5503 | 22.5423 | 22.4421 | 22.3258 | 50.2964 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048752 | 0.88182 | 0.04299 | 18.0878 | 22.1974 | 3.4674 | 0.045183 | 0.3529 | 0.3449 | 0.2447 | 0.1284
2012-10-30 22:28 | 26.5013 | 22.6031 | 22.547 | 22.4548 | 22.3546 | 50.3039 | 0.006707 | 0.32633 | 0.002189 | 48.6545 | 0.048938 | 0.88499 | 0.04331 | 18.084 | 22.1861 | 3.468 | 0.045499 | 0.417 | 0.3609 | 0.2687 | 0.1685
2012-10-30 22:28 | 26.4916 | 22.5653 | 22.4972 | 22.4691 | 22.3248 | 50.301 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.04896 | 0.88508 | 0.043334 | 18.0776 | 22.1884 | 3.4678 | 0.045527 | 0.3769 | 0.3088 | 0.2807 | 0.1364
2012-10-30 22:28 | 26.4946 | 22.5563 | 22.5122 | 22.4361 | 22.3038 | 50.2993 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.049051 | 0.8866 | 0.043488 | 18.075 | 22.2275 | 3.4676 | 0.045682 | 0.3288 | 0.2847 | 0.2086 | 0.0763
2012-10-30 22:28 | 26.4864 | 22.5561 | 22.5161 | 22.4439 | 22.3196 | 50.2918 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.048862 | 0.88366 | 0.043178 | 18.0849 | 22.2113 | 3.4671 | 0.045369 | 0.3448 | 0.3048 | 0.2326 | 0.1083
2012-10-30 22:28 | 26.4849 | 22.5827 | 22.5186 | 22.4504 | 22.3141 | 50.2877 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048755 | 0.88192 | 0.042998 | 18.0887 | 22.2018 | 3.4668 | 0.045194 | 0.3809 | 0.3168 | 0.2486 | 0.1123
2012-10-30 22:28 | 26.4826 | 22.5684 | 22.5283 | 22.408 | 22.3078 | 50.279 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048834 | 0.88326 | 0.043133 | 18.087 | 22.2115 | 3.4662 | 0.045328 | 0.3569 | 0.3168 | 0.1965 | 0.0963
2012-10-30 22:28 | 26.4766 | 22.5423 | 22.5463 | 22.4341 | 22.3379 | 50.2883 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048906 | 0.88429 | 0.043247 | 18.0813 | 22.1975 | 3.4669 | 0.045444 | 0.3448 | 0.3488 | 0.2366 | 0.1404
2012-10-30 22:28 | 26.4811 | 22.5629 | 22.5268 | 22.4506 | 22.3344 | 50.2894 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.049079 | 0.88728 | 0.043547 | 18.0785 | 22.194 | 3.467 | 0.045744 | 0.3689 | 0.3328 | 0.2566 | 0.1404
2012-10-30 22:28 | 26.4841 | 22.5458 | 22.5298 | 22.4576 | 22.3574 | 50.2785 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048796 | 0.88271 | 0.043073 | 18.0898 | 22.205 | 3.4662 | 0.045266 | 0.3408 | 0.3248 | 0.2526 | 0.1524
2012-10-30 22:28 | 26.4834 | 22.5692 | 22.5571 | 22.4209 | 22.3246 | 50.2814 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048998 | 0.88592 | 0.043408 | 18.0807 | 22.2284 | 3.4664 | 0.045605 | 0.3408 | 0.3287 | 0.1925 | 0.0962
2012-10-30 22:28 | 26.4856 | 22.5634 | 22.5273 | 22.4591 | 22.3429 | 50.2918 | 0.006713 | 0.32637 | 0.002191 | 48.6176 | 0.049051 | 0.88645 | 0.043481 | 18.0721 | 22.1744 | 3.4671 | 0.045672 | 0.389 | 0.3529 | 0.2847 | 0.1685
2012-10-30 22:28 | 26.4796 | 22.5373 | 22.5413 | 22.4772 | 22.3449 | 50.2819 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.049101 | 0.88746 | 0.043575 | 18.0743 | 22.2166 | 3.4664 | 0.045772 | 0.3207 | 0.3247 | 0.2606 | 0.1283
2012-10-30 22:29 | 26.4781 | 22.5599 | 22.5318 | 22.4316 | 22.3153 | 50.2802 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048925 | 0.88456 | 0.043277 | 18.0799 | 22.179 | 3.4663 | 0.045473 | 0.3809 | 0.3528 | 0.2526 | 0.1363
2012-10-30 22:29 | 26.4804 | 22.5381 | 22.5622 | 22.4219 | 22.3497 | 50.2871 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048869 | 0.88366 | 0.043184 | 18.0823 | 22.1933 | 3.4668 | 0.04538 | 0.3448 | 0.3689 | 0.2286 | 0.1564
2012-10-30 22:29 | 26.4804 | 22.5541 | 22.5141 | 22.3858 | 22.3056 | 50.2923 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.048963 | 0.88527 | 0.043345 | 18.0803 | 22.1772 | 3.4672 | 0.045545 | 0.3769 | 0.3369 | 0.2086 | 0.1284
2012-10-30 22:29 | 26.4819 | 22.5356 | 22.4875 | 22.4635 | 22.2911 | 50.2808 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048932 | 0.88458 | 0.043284 | 18.0776 | 22.1948 | 3.4664 | 0.045482 | 0.3408 | 0.2927 | 0.2687 | 0.0963
2012-10-30 22:29 | 26.4841 | 22.5538 | 22.5258 | 22.4456 | 22.3053 | 50.2808 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048859 | 0.88365 | 0.043174 | 18.0857 | 22.205 | 3.4664 | 0.045369 | 0.3488 | 0.3208 | 0.2406 | 0.1003
2012-10-30 22:29 | 26.4841 | 22.5178 | 22.5138 | 22.4095 | 22.3414 | 50.2894 | 0.006725 | 0.32646 | 0.002195 | 48.5436 | 0.048954 | 0.88524 | 0.043336 | 18.0831 | 22.2211 | 3.467 | 0.045531 | 0.2967 | 0.2927 | 0.1884 | 0.1203
2012-10-30 22:29 | 26.4796 | 22.5694 | 22.4972 | 22.4211 | 22.3529 | 50.2871 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048994 | 0.88543 | 0.043381 | 18.0722 | 22.2045 | 3.4668 | 0.045578 | 0.3649 | 0.2927 | 0.2166 | 0.1484
2012-10-30 22:29 | 26.4789 | 22.5767 | 22.5446 | 22.4645 | 22.2961 | 50.2738 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048935 | 0.88468 | 0.043292 | 18.0786 | 22.2118 | 3.4659 | 0.045489 | 0.3649 | 0.3328 | 0.2527 | 0.0843
2012-10-30 22:29 | 26.4736 | 22.5393 | 22.5353 | 22.4471 | 22.3229 | 50.2848 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.04879 | 0.88244 | 0.043054 | 18.0864 | 22.2146 | 3.4666 | 0.04525 | 0.3247 | 0.3207 | 0.2325 | 0.1083
2012-10-30 22:29 | 26.4871 | 22.5087 | 22.5528 | 22.4326 | 22.3003 | 50.2866 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048944 | 0.88512 | 0.043322 | 18.0844 | 22.212 | 3.4668 | 0.04552 | 0.2967 | 0.3408 | 0.2206 | 0.0883
2012-10-30 22:29 | 26.4819 | 22.5156 | 22.5236 | 22.4755 | 22.3472 | 50.2779 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048881 | 0.88413 | 0.043217 | 18.0873 | 22.2189 | 3.4662 | 0.045413 | 0.2967 | 0.3047 | 0.2566 | 0.1283
2012-10-30 22:29 | 26.4736 | 22.5393 | 22.5353 | 22.4271 | 22.3269 | 50.2866 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048875 | 0.88391 | 0.043201 | 18.0851 | 22.1905 | 3.4668 | 0.045396 | 0.3488 | 0.3448 | 0.2366 | 0.1364
2012-10-30 22:29 | 26.4684 | 22.5542 | 22.5221 | 22.4139 | 22.3337 | 50.2842 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048947 | 0.88476 | 0.043306 | 18.0759 | 22.2214 | 3.4666 | 0.045501 | 0.3328 | 0.3007 | 0.1925 | 0.1123
2012-10-30 22:29 | 26.4699 | 22.5717 | 22.5196 | 22.4354 | 22.3552 | 50.2854 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048727 | 0.88137 | 0.042946 | 18.0879 | 22.1948 | 3.4667 | 0.045139 | 0.3769 | 0.3248 | 0.2406 | 0.1604
2012-10-30 22:29 | 26.4729 | 22.5867 | 22.5266 | 22.4464 | 22.3141 | 50.2842 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048928 | 0.88478 | 0.04329 | 18.0832 | 22.2139 | 3.4666 | 0.045482 | 0.3728 | 0.3127 | 0.2325 | 0.1002
2012-10-30 22:29 | 26.4721 | 22.5218 | 22.5218 | 22.4336 | 22.3414 | 50.297 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048919 | 0.88456 | 0.043272 | 18.0821 | 22.2131 | 3.4675 | 0.045468 | 0.3087 | 0.3087 | 0.2205 | 0.1283
2012-10-30 22:29 | 26.4811 | 22.5228 | 22.5388 | 22.4346 | 22.3825 | 50.2814 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.049001 | 0.88567 | 0.043399 | 18.0746 | 22.202 | 3.4664 | 0.045592 | 0.3208 | 0.3368 | 0.2326 | 0.1805
2012-10-30 22:29 | 26.4677 | 22.5174 | 22.4974 | 22.4292 | 22.321 | 50.2866 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048985 | 0.88575 | 0.043388 | 18.082 | 22.2247 | 3.4668 | 0.045584 | 0.2927 | 0.2727 | 0.2045 | 0.0963
2012-10-30 22:29 | 26.4781 | 22.5438 | 22.4837 | 22.4075 | 22.3113 | 50.2819 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048717 | 0.88122 | 0.042931 | 18.0886 | 22.199 | 3.4664 | 0.045124 | 0.3448 | 0.2847 | 0.2085 | 0.1123
2012-10-30 22:29 | 26.4729 | 22.5306 | 22.5386 | 22.4184 | 22.3302 | 50.286 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048872 | 0.88361 | 0.043184 | 18.08 | 22.1978 | 3.4667 | 0.045377 | 0.3328 | 0.3408 | 0.2206 | 0.1324
2012-10-30 22:29 | 26.4766 | 22.5223 | 22.5183 | 22.4141 | 22.2858 | 50.2814 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048944 | 0.88511 | 0.043321 | 18.0841 | 22.2456 | 3.4664 | 0.045516 | 0.2767 | 0.2727 | 0.1685 | 0.0402
2012-10-30 22:29 | 26.4729 | 22.5266 | 22.5346 | 22.4304 | 22.3342 | 50.2819 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048736 | 0.88155 | 0.042963 | 18.0884 | 22.2179 | 3.4664 | 0.04516 | 0.3087 | 0.3167 | 0.2125 | 0.1163
2012-10-30 22:29 | 26.4856 | 22.4992 | 22.5353 | 22.4792 | 22.2988 | 50.279 | 0.006741 | 0.32636 | 0.0022 | 48.4134 | 0.049064 | 0.88664 | 0.043502 | 18.0711 | 22.2506 | 3.4662 | 0.045702 | 0.2486 | 0.2847 | 0.2286 | 0.0482
2012-10-30 22:30 | 26.4684 | 22.5101 | 22.5101 | 22.3978 | 22.2896 | 50.279 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.04901 | 0.88587 | 0.043417 | 18.0754 | 22.2294 | 3.4662 | 0.045613 | 0.2807 | 0.2807 | 0.1684 | 0.0602
2012-10-30 22:30 | 26.4706 | 22.5443 | 22.5283 | 22.4201 | 22.3279 | 50.2761 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048717 | 0.8813 | 0.042934 | 18.0901 | 22.2396 | 3.466 | 0.04513 | 0.3047 | 0.2887 | 0.1805 | 0.0883
2012-10-30 22:30 | 26.4856 | 22.5193 | 22.5153 | 22.4431 | 22.3148 | 50.2819 | 0.006741 | 0.32636 | 0.0022 | 48.4134 | 0.048815 | 0.88268 | 0.043088 | 18.0822 | 22.2546 | 3.4664 | 0.045288 | 0.2647 | 0.2607 | 0.1885 | 0.0602
2012-10-30 22:30 | 26.4819 | 22.5356 | 22.4875 | 22.4394 | 22.3231 | 50.2779 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.049001 | 0.88569 | 0.0434 | 18.0749 | 22.255 | 3.4662 | 0.045596 | 0.2806 | 0.2325 | 0.1844 | 0.0681
2012-10-30 22:30 | 26.4789 | 22.4845 | 22.5045 | 22.4244 | 22.3081 | 50.286 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.049164 | 0.88836 | 0.043675 | 18.0693 | 22.2159 | 3.4667 | 0.045873 | 0.2686 | 0.2886 | 0.2085 | 0.0922
2012-10-30 22:30 | 26.4781 | 22.5238 | 22.5438 | 22.3995 | 22.3073 | 50.2906 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048947 | 0.88505 | 0.043321 | 18.0818 | 22.2271 | 3.467 | 0.045515 | 0.2967 | 0.3167 | 0.1724 | 0.0802
2012-10-30 22:30 | 26.4871 | 22.5168 | 22.5047 | 22.4406 | 22.3323 | 50.2767 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048922 | 0.88481 | 0.043286 | 18.0861 | 22.2241 | 3.4661 | 0.045479 | 0.2927 | 0.2806 | 0.2165 | 0.1082
2012-10-30 22:30 | 26.4766 | 22.5143 | 22.4822 | 22.4101 | 22.2978 | 50.2831 | 0.006716 | 0.32638 | 0.002192 | 48.598 | 0.048875 | 0.88378 | 0.043195 | 18.0825 | 22.2055 | 3.4665 | 0.045387 | 0.3088 | 0.2767 | 0.2046 | 0.0923
2012-10-30 22:30 | 26.4841 | 22.4777 | 22.5338 | 22.4336 | 22.2973 | 50.2738 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048972 | 0.88531 | 0.043355 | 18.0779 | 22.213 | 3.4659 | 0.045554 | 0.2647 | 0.3208 | 0.2206 | 0.0843
2012-10-30 22:30 | 26.4841 | 22.5178 | 22.5619 | 22.4055 | 22.3093 | 50.2819 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048932 | 0.88492 | 0.043301 | 18.0847 | 22.189 | 3.4664 | 0.045494 | 0.3288 | 0.3729 | 0.2165 | 0.1203
2012-10-30 22:30 | 26.4871 | 22.5127 | 22.5368 | 22.4246 | 22.2922 | 50.2819 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048853 | 0.88332 | 0.043153 | 18.0811 | 22.1799 | 3.4664 | 0.04535 | 0.3328 | 0.3569 | 0.2447 | 0.1123
2012-10-30 22:30 | 26.4879 | 22.5336 | 22.5176 | 22.4334 | 22.277 | 50.2663 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048834 | 0.8828 | 0.043111 | 18.0775 | 22.1687 | 3.4654 | 0.045311 | 0.3649 | 0.3489 | 0.2647 | 0.1083
2012-10-30 22:30 | 26.4879 | 22.5256 | 22.5055 | 22.4254 | 22.285 | 50.2733 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048938 | 0.88472 | 0.043296 | 18.0784 | 22.2008 | 3.4658 | 0.045496 | 0.3248 | 0.3047 | 0.2246 | 0.0842
2012-10-30 22:30 | 26.4968 | 22.5345 | 22.5144 | 22.4503 | 22.322 | 50.2854 | 0.006716 | 0.32644 | 0.002192 | 48.6066 | 0.049023 | 0.88614 | 0.043441 | 18.0759 | 22.2338 | 3.4667 | 0.045633 | 0.3007 | 0.2806 | 0.2165 | 0.0882
2012-10-30 22:30 | 26.4879 | 22.5015 | 22.5176 | 22.4053 | 22.3011 | 50.2738 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048935 | 0.88458 | 0.043287 | 18.0765 | 22.261 | 3.4659 | 0.045483 | 0.2405 | 0.2566 | 0.1443 | 0.0401
2012-10-30 22:30 | 26.4953 | 22.5169 | 22.5811 | 22.4408 | 22.3125 | 50.2663 | 0.006738 | 0.32638 | 0.002199 | 48.4393 | 0.049117 | 0.88785 | 0.043609 | 18.0763 | 22.2082 | 3.4654 | 0.045808 | 0.3087 | 0.3729 | 0.2326 | 0.1043
2012-10-30 22:30 | 26.4931 | 22.4987 | 22.5428 | 22.4145 | 22.2982 | 50.2709 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.049083 | 0.88715 | 0.043544 | 18.0744 | 22.2381 | 3.4657 | 0.045743 | 0.2606 | 0.3047 | 0.1764 | 0.0601
2012-10-30 22:30 | 26.4894 | 22.515 | 22.5351 | 22.4349 | 22.3146 | 50.279 | 0.006735 | 0.32646 | 0.002199 | 48.4716 | 0.048768 | 0.88212 | 0.043019 | 18.088 | 22.2384 | 3.4662 | 0.045218 | 0.2766 | 0.2967 | 0.1965 | 0.0762
2012-10-30 22:30 | 26.4953 | 22.529 | 22.525 | 22.3806 | 22.3125 | 50.2761 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048919 | 0.8845 | 0.043269 | 18.081 | 22.2603 | 3.466 | 0.045462 | 0.2687 | 0.2647 | 0.1203 | 0.0522
2012-10-30 22:30 | 26.5043 | 22.5299 | 22.562 | 22.4378 | 22.3295 | 50.2681 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048947 | 0.88512 | 0.043324 | 18.0833 | 22.2533 | 3.4655 | 0.045518 | 0.2766 | 0.3087 | 0.1845 | 0.0762
2012-10-30 22:30 | 26.5028 | 22.5004 | 22.5325 | 22.4242 | 22.328 | 50.2808 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048831 | 0.88322 | 0.043128 | 18.0872 | 22.2598 | 3.4664 | 0.045325 | 0.2406 | 0.2727 | 0.1644 | 0.0682
2012-10-30 22:30 | 26.4909 | 22.5125 | 22.5326 | 22.3762 | 22.3121 | 50.275 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.049039 | 0.8865 | 0.043473 | 18.0774 | 22.2359 | 3.466 | 0.045669 | 0.2766 | 0.2967 | 0.1403 | 0.0762
2012-10-30 22:30 | 26.4931 | 22.5147 | 22.5228 | 22.4105 | 22.3103 | 50.2605 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.049079 | 0.88693 | 0.04353 | 18.0715 | 22.226 | 3.465 | 0.045727 | 0.2887 | 0.2968 | 0.1845 | 0.0843
2012-10-30 22:30 | 26.4983 | 22.4598 | 22.556 | 22.4077 | 22.3034 | 50.2825 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048935 | 0.88497 | 0.043306 | 18.0845 | 22.2192 | 3.4665 | 0.0455 | 0.2406 | 0.3368 | 0.1885 | 0.0842
2012-10-30 22:31 | 26.4961 | 22.5338 | 22.5498 | 22.4175 | 22.3253 | 50.2756 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.049073 | 0.88667 | 0.043512 | 18.0684 | 22.201 | 3.466 | 0.04571 | 0.3328 | 0.3488 | 0.2165 | 0.1243
2012-10-30 22:31 | 26.5043 | 22.4979 | 22.554 | 22.4137 | 22.2934 | 50.26 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.049026 | 0.88642 | 0.043458 | 18.0807 | 22.2212 | 3.4649 | 0.045653 | 0.2767 | 0.3328 | 0.1925 | 0.0722
2012-10-30 22:31 | 26.4976 | 22.4832 | 22.5353 | 22.415 | 22.2787 | 50.2733 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.04885 | 0.88342 | 0.043155 | 18.0843 | 22.2025 | 3.4658 | 0.045353 | 0.2807 | 0.3328 | 0.2125 | 0.0762
2012-10-30 22:31 | 26.5043 | 22.5179 | 22.5259 | 22.3656 | 22.3014 | 50.2652 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.049004 | 0.88583 | 0.043409 | 18.0767 | 22.2372 | 3.4653 | 0.045607 | 0.2807 | 0.2887 | 0.1284 | 0.0642
2012-10-30 22:31 | 26.5073 | 22.4969 | 22.5289 | 22.4247 | 22.3004 | 50.2715 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.04878 | 0.88232 | 0.04304 | 18.0878 | 22.2322 | 3.4657 | 0.045234 | 0.2647 | 0.2967 | 0.1925 | 0.0682
2012-10-30 22:31 | 26.5118 | 22.5615 | 22.5455 | 22.4292 | 22.2969 | 50.2692 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.04907 | 0.88687 | 0.043519 | 18.0736 | 22.2488 | 3.4656 | 0.045715 | 0.3127 | 0.2967 | 0.1804 | 0.0481
2012-10-30 22:31 | 26.5051 | 22.5067 | 22.5147 | 22.4265 | 22.3062 | 50.2704 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048878 | 0.88377 | 0.043197 | 18.081 | 22.21 | 3.4656 | 0.045394 | 0.2967 | 0.3047 | 0.2165 | 0.0962
2012-10-30 22:31 | 26.5073 | 22.4969 | 22.5329 | 22.4648 | 22.3245 | 50.2646 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.048771 | 0.88206 | 0.043019 | 18.0858 | 22.2804 | 3.4652 | 0.04521 | 0.2165 | 0.2525 | 0.1844 | 0.0441
2012-10-30 22:31 | 26.5118 | 22.5334 | 22.5294 | 22.4092 | 22.2969 | 50.26 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.048891 | 0.88416 | 0.043227 | 18.0842 | 22.2488 | 3.4649 | 0.045421 | 0.2846 | 0.2806 | 0.1604 | 0.0481
2012-10-30 22:31 | 26.5066 | 22.4841 | 22.4721 | 22.408 | 22.2917 | 50.2663 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.049117 | 0.88788 | 0.04361 | 18.0769 | 22.2516 | 3.4654 | 0.045805 | 0.2325 | 0.2205 | 0.1564 | 0.0401
2012-10-30 22:31 | 26.5081 | 22.5097 | 22.4816 | 22.4375 | 22.3213 | 50.2727 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.049051 | 0.88683 | 0.0435 | 18.0797 | 22.1929 | 3.4658 | 0.045695 | 0.3168 | 0.2887 | 0.2446 | 0.1284
2012-10-30 22:31 | 26.5081 | 22.4937 | 22.5418 | 22.4335 | 22.3132 | 50.2663 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.04912 | 0.88796 | 0.043616 | 18.0773 | 22.229 | 3.4654 | 0.045812 | 0.2647 | 0.3128 | 0.2045 | 0.0842
2012-10-30 22:31 | 26.5088 | 22.5344 | 22.4904 | 22.4102 | 22.3059 | 50.2773 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.049196 | 0.88894 | 0.043732 | 18.0693 | 22.2297 | 3.4661 | 0.045926 | 0.3047 | 0.2607 | 0.1805 | 0.0762
2012-10-30 22:31 | 26.5036 | 22.5092 | 22.5172 | 22.405 | 22.3007 | 50.2623 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.048891 | 0.88414 | 0.043227 | 18.0839 | 22.2004 | 3.4651 | 0.04542 | 0.3088 | 0.3168 | 0.2046 | 0.1003
2012-10-30 22:31 | 26.5013 | 22.4949 | 22.5229 | 22.4548 | 22.3024 | 50.2698 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048935 | 0.88508 | 0.043311 | 18.0868 | 22.2102 | 3.4656 | 0.045506 | 0.2847 | 0.3127 | 0.2446 | 0.0922
2012-10-30 22:31 | 26.4961 | 22.4696 | 22.5458 | 22.4255 | 22.3293 | 50.264 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048762 | 0.88208 | 0.043012 | 18.0894 | 22.1929 | 3.4652 | 0.045212 | 0.2767 | 0.3529 | 0.2326 | 0.1364
2012-10-30 22:31 | 26.4998 | 22.5415 | 22.5335 | 22.4052 | 22.333 | 50.2808 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.048913 | 0.8844 | 0.043259 | 18.0811 | 22.2368 | 3.4664 | 0.045457 | 0.3047 | 0.2967 | 0.1684 | 0.0962
2012-10-30 22:31 | 26.4939 | 22.4915 | 22.5276 | 22.4113 | 22.291 | 50.2605 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048834 | 0.88346 | 0.043143 | 18.0911 | 22.1867 | 3.465 | 0.045335 | 0.3048 | 0.3409 | 0.2246 | 0.1043
2012-10-30 22:31 | 26.4976 | 22.5353 | 22.5353 | 22.3829 | 22.3428 | 50.2744 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.048771 | 0.88221 | 0.043026 | 18.0887 | 22.2265 | 3.4659 | 0.045218 | 0.3088 | 0.3088 | 0.1564 | 0.1163
2012-10-30 22:31 | 26.4961 | 22.4937 | 22.5137 | 22.4255 | 22.3454 | 50.2715 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048963 | 0.88531 | 0.043347 | 18.0812 | 22.2331 | 3.4657 | 0.045542 | 0.2606 | 0.2806 | 0.1924 | 0.1123
2012-10-30 22:31 | 26.4939 | 22.5236 | 22.5195 | 22.3592 | 22.3231 | 50.2669 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048963 | 0.88534 | 0.043349 | 18.0818 | 22.2188 | 3.4654 | 0.045545 | 0.3048 | 0.3007 | 0.1404 | 0.1043
2012-10-30 22:31 | 26.4924 | 22.498 | 22.5501 | 22.4619 | 22.2815 | 50.275 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.048994 | 0.88577 | 0.043398 | 18.0792 | 22.2494 | 3.466 | 0.045597 | 0.2486 | 0.3007 | 0.2125 | 0.0321
2012-10-30 22:31 | 26.4953 | 22.5129 | 22.557 | 22.3967 | 22.2844 | 50.2698 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048932 | 0.8847 | 0.04329 | 18.0803 | 22.1841 | 3.4656 | 0.045487 | 0.3288 | 0.3729 | 0.2126 | 0.1003
2012-10-30 22:31 | 26.4894 | 22.483 | 22.511 | 22.4229 | 22.2665 | 50.275 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.048774 | 0.88203 | 0.04302 | 18.0841 | 22.2103 | 3.466 | 0.045211 | 0.2727 | 0.3007 | 0.2126 | 0.0562
2012-10-30 22:32 | 26.4939 | 22.4915 | 22.5155 | 22.4233 | 22.2589 | 50.2814 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048938 | 0.88466 | 0.043294 | 18.0772 | 22.2188 | 3.4664 | 0.045491 | 0.2727 | 0.2967 | 0.2045 | 0.0401
2012-10-30 22:32 | 26.5028 | 22.5124 | 22.5365 | 22.4363 | 22.2718 | 50.2709 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048847 | 0.88359 | 0.043161 | 18.089 | 22.2237 | 3.4657 | 0.045357 | 0.2887 | 0.3128 | 0.2126 | 0.0481
2012-10-30 22:32 | 26.5006 | 22.5022 | 22.4982 | 22.414 | 22.2817 | 50.2663 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048947 | 0.88482 | 0.043309 | 18.0771 | 22.1894 | 3.4654 | 0.045504 | 0.3128 | 0.3088 | 0.2246 | 0.0923
2012-10-30 22:32 | 26.4916 | 22.5333 | 22.5052 | 22.3769 | 22.2526 | 50.2646 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048897 | 0.88408 | 0.043229 | 18.0805 | 22.1925 | 3.4652 | 0.045426 | 0.3408 | 0.3127 | 0.1844 | 0.0601
2012-10-30 22:32 | 26.4916 | 22.5172 | 22.4611 | 22.417 | 22.3048 | 50.2704 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048894 | 0.8843 | 0.043237 | 18.0861 | 22.2286 | 3.4656 | 0.045433 | 0.2886 | 0.2325 | 0.1884 | 0.0762
2012-10-30 22:32 | 26.5066 | 22.4841 | 22.5403 | 22.412 | 22.3238 | 50.2623 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.048865 | 0.88366 | 0.04318 | 18.0838 | 22.2395 | 3.4651 | 0.045374 | 0.2446 | 0.3008 | 0.1725 | 0.0843
2012-10-30 22:32 | 26.5028 | 22.5605 | 22.4884 | 22.3921 | 22.316 | 50.2669 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.049067 | 0.88677 | 0.043511 | 18.0727 | 22.2277 | 3.4654 | 0.045705 | 0.3328 | 0.2607 | 0.1644 | 0.0883
2012-10-30 22:32 | 26.5081 | 22.4977 | 22.5057 | 22.4055 | 22.3173 | 50.2681 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048897 | 0.88426 | 0.043238 | 18.0841 | 22.213 | 3.4655 | 0.045434 | 0.2847 | 0.2927 | 0.1925 | 0.1043
2012-10-30 22:32 | 26.5036 | 22.4571 | 22.5092 | 22.3849 | 22.2646 | 50.2594 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.049016 | 0.88632 | 0.043444 | 18.0823 | 22.2245 | 3.4649 | 0.045638 | 0.2326 | 0.2847 | 0.1604 | 0.0401
2012-10-30 22:32 | 26.5066 | 22.4721 | 22.4962 | 22.4 | 22.2676 | 50.2756 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048976 | 0.88554 | 0.04337 | 18.0812 | 22.2275 | 3.466 | 0.045563 | 0.2446 | 0.2687 | 0.1725 | 0.0401
2012-10-30 22:32 | 26.5006 | 22.4781 | 22.4982 | 22.394 | 22.2696 | 50.2478 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048979 | 0.88548 | 0.04337 | 18.0789 | 22.1694 | 3.4641 | 0.045565 | 0.3087 | 0.3288 | 0.2246 | 0.1002
2012-10-30 22:32 | 26.5111 | 22.4927 | 22.5047 | 22.3884 | 22.3002 | 50.2577 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048966 | 0.88528 | 0.043349 | 18.0795 | 22.1518 | 3.4648 | 0.045544 | 0.3409 | 0.3529 | 0.2366 | 0.1484
2012-10-30 22:32 | 26.5058 | 22.4954 | 22.5154 | 22.4192 | 22.2789 | 50.2657 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048758 | 0.88187 | 0.042998 | 18.0867 | 22.1385 | 3.4653 | 0.045197 | 0.3569 | 0.3769 | 0.2807 | 0.1404
2012-10-30 22:32 | 26.5088 | 22.5224 | 22.5304 | 22.3981 | 22.2658 | 50.2582 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048897 | 0.8841 | 0.04323 | 18.0808 | 22.1816 | 3.4648 | 0.045425 | 0.3408 | 0.3488 | 0.2165 | 0.0842
2012-10-30 22:32 | 26.5096 | 22.4912 | 22.5192 | 22.415 | 22.2867 | 50.2617 | 0.006713 | 0.32646 | 0.002191 | 48.6304 | 0.048828 | 0.88324 | 0.043127 | 18.0889 | 22.1703 | 3.465 | 0.045318 | 0.3209 | 0.3489 | 0.2447 | 0.1164
2012-10-30 22:32 | 26.5096 | 22.5032 | 22.4671 | 22.407 | 22.2786 | 50.2704 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.04885 | 0.88349 | 0.043159 | 18.0858 | 22.1744 | 3.4656 | 0.045354 | 0.3288 | 0.2927 | 0.2326 | 0.1042
2012-10-30 22:32 | 26.5193 | 22.5209 | 22.5169 | 22.4046 | 22.3325 | 50.2652 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.049016 | 0.88605 | 0.043431 | 18.0767 | 22.2121 | 3.4653 | 0.045623 | 0.3088 | 0.3048 | 0.1925 | 0.1204
2012-10-30 22:32 | 26.5073 | 22.4688 | 22.4768 | 22.3966 | 22.2763 | 50.2559 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.049048 | 0.88664 | 0.043488 | 18.077 | 22.156 | 3.4646 | 0.045687 | 0.3128 | 0.3208 | 0.2406 | 0.1203
2012-10-30 22:32 | 26.5073 | 22.5129 | 22.5129 | 22.3686 | 22.2924 | 50.2652 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048743 | 0.88189 | 0.042986 | 18.0926 | 22.1801 | 3.4653 | 0.045182 | 0.3328 | 0.3328 | 0.1885 | 0.1123
2012-10-30 22:32 | 26.5141 | 22.4756 | 22.5117 | 22.3994 | 22.3192 | 50.2634 | 0.006744 | 0.32631 | 0.002201 | 48.3855 | 0.049029 | 0.88634 | 0.043456 | 18.0778 | 22.1869 | 3.4652 | 0.045657 | 0.2887 | 0.3248 | 0.2125 | 0.1323
2012-10-30 22:32 | 26.5148 | 22.5004 | 22.4763 | 22.4001 | 22.3039 | 50.2657 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048875 | 0.88398 | 0.043205 | 18.0866 | 22.2116 | 3.4653 | 0.045401 | 0.2888 | 0.2647 | 0.1885 | 0.0923
2012-10-30 22:32 | 26.5171 | 22.5227 | 22.4946 | 22.3944 | 22.2581 | 50.2692 | 0.006713 | 0.32637 | 0.002191 | 48.6176 | 0.048749 | 0.88192 | 0.042993 | 18.091 | 22.1979 | 3.4656 | 0.045184 | 0.3248 | 0.2967 | 0.1965 | 0.0602
2012-10-30 22:32 | 26.5156 | 22.4731 | 22.4691 | 22.3809 | 22.2887 | 50.2559 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048809 | 0.88285 | 0.043091 | 18.088 | 22.2205 | 3.4646 | 0.045287 | 0.2526 | 0.2486 | 0.1604 | 0.0682
2012-10-30 22:33 | 26.5186 | 22.4961 | 22.4961 | 22.4039 | 22.2836 | 50.2623 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048928 | 0.88468 | 0.043285 | 18.0812 | 22.2235 | 3.4651 | 0.04548 | 0.2726 | 0.2726 | 0.1804 | 0.0601
2012-10-30 22:33 | 26.5238 | 22.5254 | 22.5013 | 22.3771 | 22.2608 | 50.2634 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.04885 | 0.88342 | 0.043155 | 18.0843 | 22.2367 | 3.4652 | 0.045352 | 0.2887 | 0.2646 | 0.1404 | 0.0241
2012-10-30 22:33 | 26.5163 | 22.4858 | 22.4979 | 22.3936 | 22.2853 | 50.2623 | 0.006716 | 0.32638 | 0.002192 | 48.598 | 0.048859 | 0.88385 | 0.043184 | 18.0898 | 22.2172 | 3.4651 | 0.045376 | 0.2686 | 0.2807 | 0.1764 | 0.0681
2012-10-30 22:33 | 26.52 | 22.4935 | 22.5336 | 22.4013 | 22.2971 | 50.2605 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.049123 | 0.88794 | 0.043618 | 18.0759 | 22.2489 | 3.465 | 0.045816 | 0.2446 | 0.2847 | 0.1524 | 0.0482
2012-10-30 22:33 | 26.5238 | 22.4933 | 22.4813 | 22.3731 | 22.2808 | 50.2559 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.049101 | 0.88759 | 0.043582 | 18.0769 | 22.2808 | 3.4646 | 0.045777 | 0.2125 | 0.2005 | 0.0923 | 0
2012-10-30 22:33 | 26.5215 | 22.475 | 22.5231 | 22.3988 | 22.2865 | 50.2501 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.049139 | 0.88829 | 0.04365 | 18.077 | 22.2183 | 3.4642 | 0.045849 | 0.2567 | 0.3048 | 0.1805 | 0.0682
2012-10-30 22:33 | 26.52 | 22.4535 | 22.5176 | 22.4174 | 22.277 | 50.264 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.049114 | 0.88775 | 0.043601 | 18.0754 | 22.2048 | 3.4652 | 0.045797 | 0.2487 | 0.3128 | 0.2126 | 0.0722
2012-10-30 22:33 | 26.5186 | 22.5002 | 22.5202 | 22.3999 | 22.2676 | 50.2605 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048887 | 0.8842 | 0.043226 | 18.0866 | 22.2796 | 3.465 | 0.045421 | 0.2206 | 0.2406 | 0.1203 | -0.012
2012-10-30 22:33 | 26.5193 | 22.4968 | 22.4688 | 22.4006 | 22.3244 | 50.2461 | 0.006747 | 0.32634 | 0.002202 | 48.3683 | 0.048815 | 0.88285 | 0.043097 | 18.0857 | 22.2603 | 3.464 | 0.045299 | 0.2365 | 0.2085 | 0.1403 | 0.0641
2012-10-30 22:33 | 26.5238 | 22.5174 | 22.5013 | 22.4212 | 22.3009 | 50.2536 | 0.006735 | 0.3263 | 0.002198 | 48.448 | 0.048796 | 0.88241 | 0.043058 | 18.0836 | 22.2608 | 3.4645 | 0.045256 | 0.2566 | 0.2405 | 0.1604 | 0.0401
2012-10-30 22:33 | 26.5245 | 22.502 | 22.5261 | 22.3978 | 22.3056 | 50.2467 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048859 | 0.8833 | 0.043157 | 18.0786 | 22.2775 | 3.464 | 0.045352 | 0.2245 | 0.2486 | 0.1203 | 0.0281
2012-10-30 22:33 | 26.5186 | 22.4561 | 22.4761 | 22.3959 | 22.2756 | 50.2559 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048859 | 0.88362 | 0.043173 | 18.0851 | 22.2636 | 3.4646 | 0.045366 | 0.1925 | 0.2125 | 0.1323 | 0.012
2012-10-30 22:33 | 26.5208 | 22.4903 | 22.5184 | 22.3781 | 22.2898 | 50.2473 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048991 | 0.88576 | 0.043394 | 18.08 | 22.2537 | 3.464 | 0.045589 | 0.2366 | 0.2647 | 0.1244 | 0.0361
2012-10-30 22:33 | 26.5208 | 22.5304 | 22.4983 | 22.3901 | 22.2698 | 50.2478 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048825 | 0.88291 | 0.043108 | 18.0832 | 22.2417 | 3.4641 | 0.045302 | 0.2887 | 0.2566 | 0.1484 | 0.0281
2012-10-30 22:33 | 26.5215 | 22.475 | 22.495 | 22.3748 | 22.3146 | 50.2571 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.049051 | 0.88677 | 0.043497 | 18.0786 | 22.2825 | 3.4647 | 0.045691 | 0.1925 | 0.2125 | 0.0923 | 0.0321
2012-10-30 22:33 | 26.5208 | 22.4943 | 22.5184 | 22.3861 | 22.2898 | 50.2559 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.049007 | 0.88596 | 0.043418 | 18.0783 | 22.2618 | 3.4646 | 0.045612 | 0.2325 | 0.2566 | 0.1243 | 0.028
2012-10-30 22:33 | 26.5186 | 22.4961 | 22.5442 | 22.3799 | 22.3077 | 50.2605 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048991 | 0.88572 | 0.043392 | 18.0792 | 22.1994 | 3.465 | 0.045591 | 0.2967 | 0.3448 | 0.1805 | 0.1083
2012-10-30 22:33 | 26.5171 | 22.4906 | 22.5227 | 22.3864 | 22.2701 | 50.2501 | 0.006722 | 0.32631 | 0.002193 | 48.5439 | 0.048847 | 0.88329 | 0.043146 | 18.0828 | 22.2821 | 3.4642 | 0.045339 | 0.2085 | 0.2406 | 0.1043 | -0.012
2012-10-30 22:33 | 26.5238 | 22.4412 | 22.4853 | 22.4332 | 22.337 | 50.2559 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.049013 | 0.88618 | 0.043434 | 18.0805 | 22.2567 | 3.4646 | 0.045632 | 0.1845 | 0.2286 | 0.1765 | 0.0803
2012-10-30 22:34 | 26.5163 | 22.5219 | 22.5099 | 22.3936 | 22.2974 | 50.2605 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048884 | 0.88395 | 0.043211 | 18.0827 | 22.2653 | 3.465 | 0.045407 | 0.2566 | 0.2446 | 0.1283 | 0.0321
2012-10-30 22:34 | 26.5171 | 22.4746 | 22.4906 | 22.4265 | 22.2701 | 50.2513 | 0.00671 | 0.32633 | 0.00219 | 48.6328 | 0.049054 | 0.88665 | 0.043494 | 18.0751 | 22.2701 | 3.4643 | 0.045684 | 0.2045 | 0.2205 | 0.1564 | 0
2012-10-30 22:34 | 26.5156 | 22.4811 | 22.5332 | 22.3849 | 22.2887 | 50.2507 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.049042 | 0.8865 | 0.043476 | 18.0763 | 22.2726 | 3.4643 | 0.04567 | 0.2085 | 0.2606 | 0.1123 | 0.0161
2012-10-30 22:34 | 26.5163 | 22.4538 | 22.4698 | 22.4257 | 22.3014 | 50.2571 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.049039 | 0.88641 | 0.043469 | 18.0756 | 22.2613 | 3.4647 | 0.045661 | 0.1925 | 0.2085 | 0.1644 | 0.0401
2012-10-30 22:34 | 26.5103 | 22.5119 | 22.4638 | 22.4117 | 22.2874 | 50.2501 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.048768 | 0.88218 | 0.043022 | 18.0892 | 22.2633 | 3.4642 | 0.045219 | 0.2486 | 0.2005 | 0.1484 | 0.0241
2012-10-30 22:34 | 26.5088 | 22.4984 | 22.5064 | 22.4302 | 22.2899 | 50.2507 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048944 | 0.88489 | 0.04331 | 18.0797 | 22.2377 | 3.4643 | 0.045504 | 0.2607 | 0.2687 | 0.1925 | 0.0522
2012-10-30 22:34 | 26.5073 | 22.4969 | 22.4768 | 22.4047 | 22.2563 | 50.26 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048825 | 0.88299 | 0.043112 | 18.0847 | 22.2362 | 3.4649 | 0.04531 | 0.2607 | 0.2406 | 0.1685 | 0.0201
2012-10-30 22:34 | 26.5096 | 22.4992 | 22.4671 | 22.431 | 22.2867 | 50.253 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.049123 | 0.88757 | 0.0436 | 18.0682 | 22.2466 | 3.4644 | 0.045794 | 0.2526 | 0.2205 | 0.1844 | 0.0401
2012-10-30 22:34 | 26.5006 | 22.4581 | 22.4822 | 22.398 | 22.2656 | 50.2519 | 0.006747 | 0.32637 | 0.002202 | 48.3726 | 0.048944 | 0.88521 | 0.043326 | 18.0862 | 22.2255 | 3.4644 | 0.045528 | 0.2326 | 0.2567 | 0.1725 | 0.0401
2012-10-30 22:34 | 26.5058 | 22.4874 | 22.4914 | 22.4032 | 22.2829 | 50.2536 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048806 | 0.88275 | 0.043084 | 18.087 | 22.2147 | 3.4645 | 0.045276 | 0.2727 | 0.2767 | 0.1885 | 0.0682
2012-10-30 22:34 | 26.5066 | 22.5122 | 22.4801 | 22.3559 | 22.2917 | 50.253 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048972 | 0.88543 | 0.043361 | 18.0803 | 22.2355 | 3.4644 | 0.045554 | 0.2767 | 0.2446 | 0.1204 | 0.0562
2012-10-30 22:34 | 26.5036 | 22.4531 | 22.4811 | 22.3889 | 22.3087 | 50.2461 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.048906 | 0.88433 | 0.043249 | 18.0822 | 22.1844 | 3.464 | 0.045442 | 0.2687 | 0.2967 | 0.2045 | 0.1243
2012-10-30 22:34 | 26.5066 | 22.5042 | 22.5242 | 22.3799 | 22.2957 | 50.2611 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048859 | 0.88352 | 0.043168 | 18.0831 | 22.2195 | 3.465 | 0.045363 | 0.2847 | 0.3047 | 0.1604 | 0.0762
2012-10-30 22:34 | 26.5028 | 22.4523 | 22.5204 | 22.4282 | 22.2959 | 50.2634 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.0489 | 0.88433 | 0.043244 | 18.0844 | 22.2237 | 3.4652 | 0.045442 | 0.2286 | 0.2967 | 0.2045 | 0.0722
2012-10-30 22:34 | 26.4939 | 22.5075 | 22.4314 | 22.3752 | 22.275 | 50.2571 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048708 | 0.88121 | 0.042922 | 18.0916 | 22.2188 | 3.4647 | 0.045115 | 0.2887 | 0.2126 | 0.1564 | 0.0562
2012-10-30 22:34 | 26.4953 | 22.4889 | 22.5009 | 22.3766 | 22.2603 | 50.2588 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.048954 | 0.88498 | 0.043323 | 18.0778 | 22.2042 | 3.4648 | 0.045517 | 0.2847 | 0.2967 | 0.1724 | 0.0561
2012-10-30 22:34 | 26.4909 | 22.4725 | 22.5045 | 22.3722 | 22.268 | 50.2513 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.048847 | 0.88365 | 0.043164 | 18.0902 | 22.1797 | 3.4643 | 0.045362 | 0.2928 | 0.3248 | 0.1925 | 0.0883
2012-10-30 22:34 | 26.4976 | 22.5353 | 22.4511 | 22.3428 | 22.3148 | 50.2525 | 0.006741 | 0.3264 | 0.0022 | 48.42 | 0.048878 | 0.88388 | 0.043202 | 18.0834 | 22.2185 | 3.4644 | 0.045402 | 0.3168 | 0.2326 | 0.1243 | 0.0963
2012-10-30 22:34 | 26.4871 | 22.4967 | 22.4526 | 22.3724 | 22.2963 | 50.2559 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048869 | 0.88375 | 0.043188 | 18.0841 | 22.1679 | 3.4646 | 0.045381 | 0.3288 | 0.2847 | 0.2045 | 0.1284
2012-10-30 22:34 | 26.5028 | 22.5325 | 22.4763 | 22.4002 | 22.2919 | 50.2588 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048957 | 0.88499 | 0.043327 | 18.077 | 22.1635 | 3.4648 | 0.045522 | 0.369 | 0.3128 | 0.2367 | 0.1284
2012-10-30 22:34 | 26.4856 | 22.5032 | 22.4631 | 22.391 | 22.2907 | 50.26 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048925 | 0.88466 | 0.043282 | 18.082 | 22.1985 | 3.4649 | 0.045475 | 0.3047 | 0.2646 | 0.1925 | 0.0922
2012-10-30 22:34 | 26.4871 | 22.4887 | 22.4687 | 22.3805 | 22.2561 | 50.2542 | 0.006716 | 0.32627 | 0.002191 | 48.5807 | 0.048758 | 0.88199 | 0.043004 | 18.0891 | 22.1759 | 3.4645 | 0.045195 | 0.3128 | 0.2928 | 0.2046 | 0.0802
2012-10-30 22:34 | 26.4909 | 22.4965 | 22.4484 | 22.3802 | 22.296 | 50.2496 | 0.006735 | 0.32624 | 0.002197 | 48.4395 | 0.048847 | 0.88329 | 0.043146 | 18.0828 | 22.1958 | 3.4642 | 0.045343 | 0.3007 | 0.2526 | 0.1844 | 0.1002
2012-10-30 22:34 | 26.4804 | 22.47 | 22.486 | 22.4179 | 22.2615 | 50.2559 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048928 | 0.8847 | 0.043287 | 18.0818 | 22.1692 | 3.4646 | 0.04548 | 0.3008 | 0.3168 | 0.2487 | 0.0923
2012-10-30 22:34 | 26.4931 | 22.4827 | 22.4586 | 22.4266 | 22.2902 | 50.2519 | 0.006713 | 0.32637 | 0.002191 | 48.6176 | 0.048878 | 0.88389 | 0.043203 | 18.0837 | 22.202 | 3.4644 | 0.045394 | 0.2807 | 0.2566 | 0.2246 | 0.0882
2012-10-30 22:34 | 26.4781 | 22.4356 | 22.4837 | 22.3674 | 22.2712 | 50.2501 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048821 | 0.88333 | 0.043125 | 18.0933 | 22.1228 | 3.4642 | 0.045317 | 0.3128 | 0.3609 | 0.2446 | 0.1484
2012-10-30 22:34 | 26.4841 | 22.4937 | 22.4737 | 22.3574 | 22.3053 | 50.2444 | 0.00671 | 0.32633 | 0.00219 | 48.6328 | 0.048875 | 0.88414 | 0.043212 | 18.0898 | 22.1368 | 3.4638 | 0.045402 | 0.3569 | 0.3369 | 0.2206 | 0.1685
2012-10-30 22:35 | 26.4781 | 22.5118 | 22.4637 | 22.3394 | 22.3073 | 50.2415 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048985 | 0.88572 | 0.043387 | 18.0814 | 22.183 | 3.4636 | 0.045581 | 0.3288 | 0.2807 | 0.1564 | 0.1243
2012-10-30 22:35 | 26.4804 | 22.506 | 22.4539 | 22.3858 | 22.2655 | 50.2484 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048825 | 0.88303 | 0.043114 | 18.0856 | 22.101 | 3.4641 | 0.045307 | 0.405 | 0.3529 | 0.2848 | 0.1645
2012-10-30 22:35 | 26.4759 | 22.4935 | 22.4655 | 22.3813 | 22.3171 | 50.2449 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.049117 | 0.88757 | 0.043595 | 18.0704 | 22.1086 | 3.4639 | 0.045788 | 0.3849 | 0.3569 | 0.2727 | 0.2085
2012-10-30 22:35 | 26.4699 | 22.4755 | 22.4274 | 22.3272 | 22.2349 | 50.2432 | 0.006738 | 0.3263 | 0.002199 | 48.4264 | 0.049205 | 0.88908 | 0.043747 | 18.0689 | 22.0544 | 3.4638 | 0.045946 | 0.4211 | 0.373 | 0.2728 | 0.1805
2012-10-30 22:35 | 26.4834 | 22.489 | 22.4409 | 22.3888 | 22.2565 | 50.2559 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.048695 | 0.88101 | 0.042901 | 18.0923 | 22.1281 | 3.4646 | 0.045099 | 0.3609 | 0.3128 | 0.2607 | 0.1284
2012-10-30 22:35 | 26.4811 | 22.5348 | 22.4667 | 22.3464 | 22.21 | 50.2455 | 0.006716 | 0.32631 | 0.002192 | 48.5873 | 0.0489 | 0.88416 | 0.043235 | 18.0809 | 22.1539 | 3.4639 | 0.045427 | 0.3809 | 0.3128 | 0.1925 | 0.0561
2012-10-30 22:35 | 26.4736 | 22.5033 | 22.4231 | 22.3309 | 22.2467 | 50.2582 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.048743 | 0.88158 | 0.042971 | 18.0864 | 22.1063 | 3.4648 | 0.045165 | 0.397 | 0.3168 | 0.2246 | 0.1404
2012-10-30 22:35 | 26.4699 | 22.4835 | 22.4515 | 22.3512 | 22.2309 | 50.2513 | 0.006707 | 0.32636 | 0.002189 | 48.6589 | 0.048916 | 0.88465 | 0.043273 | 18.085 | 22.1186 | 3.4643 | 0.045462 | 0.3649 | 0.3329 | 0.2326 | 0.1123
2012-10-30 22:35 | 26.4706 | 22.5323 | 22.4842 | 22.38 | 22.1995 | 50.2403 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048859 | 0.88382 | 0.043183 | 18.0893 | 22.1554 | 3.4636 | 0.045377 | 0.3769 | 0.3288 | 0.2246 | 0.0441
2012-10-30 22:35 | 26.4766 | 22.4902 | 22.4221 | 22.3539 | 22.2256 | 50.2467 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048925 | 0.88462 | 0.04328 | 18.0811 | 22.1012 | 3.464 | 0.045477 | 0.389 | 0.3209 | 0.2527 | 0.1244
2012-10-30 22:35 | 26.4706 | 22.5123 | 22.4441 | 22.4081 | 22.2076 | 50.2403 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.049127 | 0.88775 | 0.043613 | 18.0706 | 22.1474 | 3.4636 | 0.045805 | 0.3649 | 0.2967 | 0.2607 | 0.0602
2012-10-30 22:35 | 26.4811 | 22.4987 | 22.4867 | 22.3985 | 22.2181 | 50.238 | 0.006747 | 0.32636 | 0.002202 | 48.3704 | 0.04885 | 0.88336 | 0.043152 | 18.0831 | 22.1138 | 3.4634 | 0.045354 | 0.3849 | 0.3729 | 0.2847 | 0.1043
2012-10-30 22:35 | 26.4811 | 22.5108 | 22.4627 | 22.3785 | 22.2943 | 50.2519 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048695 | 0.88096 | 0.042898 | 18.0914 | 22.1699 | 3.4644 | 0.045094 | 0.3409 | 0.2928 | 0.2086 | 0.1244
2012-10-30 22:35 | 26.4774 | 22.5191 | 22.483 | 22.3427 | 22.2505 | 50.2455 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048837 | 0.88333 | 0.043139 | 18.0874 | 22.1462 | 3.4639 | 0.045333 | 0.3729 | 0.3368 | 0.1965 | 0.1043
2012-10-30 22:35 | 26.4766 | 22.4702 | 22.4541 | 22.394 | 22.2537 | 50.2548 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.04879 | 0.88248 | 0.043056 | 18.0873 | 22.1574 | 3.4646 | 0.045251 | 0.3128 | 0.2967 | 0.2366 | 0.0963
2012-10-30 22:35 | 26.4736 | 22.4992 | 22.4592 | 22.3589 | 22.2747 | 50.2496 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048944 | 0.88505 | 0.043318 | 18.0829 | 22.1143 | 3.4642 | 0.045514 | 0.3849 | 0.3449 | 0.2446 | 0.1604
2012-10-30 22:35 | 26.4789 | 22.5005 | 22.4524 | 22.3843 | 22.2479 | 50.2438 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048828 | 0.88306 | 0.043118 | 18.0851 | 22.1356 | 3.4638 | 0.045313 | 0.3649 | 0.3168 | 0.2487 | 0.1123
2012-10-30 22:35 | 26.4781 | 22.5118 | 22.4797 | 22.3795 | 22.2993 | 50.2438 | 0.006713 | 0.3264 | 0.002191 | 48.6219 | 0.048802 | 0.88299 | 0.043091 | 18.0932 | 22.1469 | 3.4638 | 0.045282 | 0.3649 | 0.3328 | 0.2326 | 0.1524
2012-10-30 22:35 | 26.4744 | 22.472 | 22.4399 | 22.3718 | 22.2555 | 50.2386 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048922 | 0.88476 | 0.043284 | 18.0852 | 22.1833 | 3.4634 | 0.045479 | 0.2887 | 0.2566 | 0.1885 | 0.0722
2012-10-30 22:35 | 26.4699 | 22.4955 | 22.4515 | 22.4154 | 22.243 | 50.2351 | 0.006747 | 0.32638 | 0.002202 | 48.3747 | 0.048793 | 0.88264 | 0.043067 | 18.0894 | 22.1266 | 3.4632 | 0.045269 | 0.3689 | 0.3249 | 0.2888 | 0.1164
2012-10-30 22:35 | 26.4766 | 22.4822 | 22.4541 | 22.39 | 22.2376 | 50.2438 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.048727 | 0.88157 | 0.042956 | 18.092 | 22.1414 | 3.4638 | 0.045155 | 0.3408 | 0.3127 | 0.2486 | 0.0962
2012-10-30 22:35 | 26.4766 | 22.5143 | 22.4421 | 22.3499 | 22.2456 | 50.2513 | 0.006751 | 0.32638 | 0.002203 | 48.346 | 0.049061 | 0.88677 | 0.043506 | 18.0749 | 22.1173 | 3.4643 | 0.045709 | 0.397 | 0.3248 | 0.2326 | 0.1283
2012-10-30 22:35 | 26.4759 | 22.4935 | 22.4655 | 22.3853 | 22.2449 | 50.2444 | 0.006741 | 0.32633 | 0.0022 | 48.4091 | 0.048938 | 0.88484 | 0.043302 | 18.0807 | 22.1487 | 3.4638 | 0.045502 | 0.3448 | 0.3168 | 0.2366 | 0.0962
2012-10-30 22:35 | 26.4654 | 22.5151 | 22.483 | 22.3347 | 22.2385 | 50.2444 | 0.006722 | 0.32628 | 0.002193 | 48.5396 | 0.048771 | 0.88193 | 0.043013 | 18.0831 | 22.1221 | 3.4638 | 0.045206 | 0.393 | 0.3609 | 0.2126 | 0.1164
2012-10-30 22:35 | 26.4677 | 22.4853 | 22.4412 | 22.345 | 22.2528 | 50.2392 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048935 | 0.88492 | 0.043304 | 18.0836 | 22.1565 | 3.4635 | 0.045502 | 0.3288 | 0.2847 | 0.1885 | 0.0963
2012-10-30 22:35 | 26.4669 | 22.4364 | 22.4404 | 22.3683 | 22.272 | 50.2455 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.048743 | 0.88167 | 0.042975 | 18.0881 | 22.1397 | 3.4639 | 0.045166 | 0.2967 | 0.3007 | 0.2286 | 0.1323
2012-10-30 22:36 | 26.4692 | 22.4628 | 22.4508 | 22.3706 | 22.2583 | 50.2542 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048903 | 0.88437 | 0.043248 | 18.0842 | 22.158 | 3.4645 | 0.045445 | 0.3048 | 0.2928 | 0.2126 | 0.1003
2012-10-30 22:36 | 26.4699 | 22.4755 | 22.4675 | 22.3753 | 22.243 | 50.2426 | 0.006741 | 0.32637 | 0.0022 | 48.4157 | 0.049262 | 0.88969 | 0.043828 | 18.0604 | 22.1066 | 3.4637 | 0.046028 | 0.3689 | 0.3609 | 0.2687 | 0.1364
2012-10-30 22:36 | 26.4692 | 22.4788 | 22.4467 | 22.3425 | 22.2543 | 50.2577 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048771 | 0.88225 | 0.043028 | 18.0896 | 22.15 | 3.4648 | 0.045225 | 0.3288 | 0.2967 | 0.1925 | 0.1043
2012-10-30 22:36 | 26.4729 | 22.4785 | 22.4865 | 22.3502 | 22.262 | 50.2438 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.048837 | 0.8832 | 0.043133 | 18.0847 | 22.1697 | 3.4638 | 0.045331 | 0.3088 | 0.3168 | 0.1805 | 0.0923
2012-10-30 22:36 | 26.4684 | 22.4299 | 22.4019 | 22.3778 | 22.2535 | 50.2484 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048683 | 0.88083 | 0.042882 | 18.0932 | 22.1572 | 3.4641 | 0.04508 | 0.2727 | 0.2447 | 0.2206 | 0.0963
2012-10-30 22:36 | 26.4669 | 22.4004 | 22.4685 | 22.3603 | 22.1958 | 50.2397 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048818 | 0.88287 | 0.0431 | 18.0849 | 22.1718 | 3.4635 | 0.045295 | 0.2286 | 0.2967 | 0.1885 | 0.024
2012-10-30 22:36 | 26.4692 | 22.4668 | 22.4948 | 22.3626 | 22.2583 | 50.2397 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.048724 | 0.8813 | 0.04294 | 18.0875 | 22.158 | 3.4635 | 0.045134 | 0.3088 | 0.3368 | 0.2046 | 0.1003
2012-10-30 22:36 | 26.4692 | 22.4548 | 22.4427 | 22.3706 | 22.2342 | 50.2438 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.048787 | 0.88248 | 0.043054 | 18.0884 | 22.162 | 3.4638 | 0.045253 | 0.2928 | 0.2807 | 0.2086 | 0.0722
2012-10-30 22:36 | 26.4699 | 22.4394 | 22.4715 | 22.3713 | 22.275 | 50.2345 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048752 | 0.88196 | 0.042997 | 18.0907 | 22.1868 | 3.4632 | 0.04519 | 0.2526 | 0.2847 | 0.1845 | 0.0882
2012-10-30 22:36 | 26.4721 | 22.4496 | 22.4617 | 22.4056 | 22.2251 | 50.2282 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048938 | 0.88491 | 0.043306 | 18.0822 | 22.185 | 3.4627 | 0.045504 | 0.2646 | 0.2767 | 0.2206 | 0.0401
2012-10-30 22:36 | 26.4744 | 22.472 | 22.5321 | 22.3998 | 22.2635 | 50.2357 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048727 | 0.88137 | 0.042946 | 18.0879 | 22.1833 | 3.4632 | 0.045144 | 0.2887 | 0.3488 | 0.2165 | 0.0802
2012-10-30 22:36 | 26.4699 | 22.4515 | 22.4835 | 22.3793 | 22.259 | 50.249 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048809 | 0.88287 | 0.043092 | 18.0883 | 22.1908 | 3.4642 | 0.045288 | 0.2607 | 0.2927 | 0.1885 | 0.0682
2012-10-30 22:36 | 26.4609 | 22.4785 | 22.5026 | 22.3703 | 22.254 | 50.227 | 0.006729 | 0.32647 | 0.002197 | 48.517 | 0.048708 | 0.88116 | 0.04292 | 18.0908 | 22.1698 | 3.4626 | 0.045117 | 0.3087 | 0.3328 | 0.2005 | 0.0842
2012-10-30 22:36 | 26.4639 | 22.4455 | 22.4775 | 22.3613 | 22.2731 | 50.2351 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048862 | 0.88372 | 0.04318 | 18.0861 | 22.2089 | 3.4632 | 0.045373 | 0.2366 | 0.2686 | 0.1524 | 0.0642
2012-10-30 22:36 | 26.4609 | 22.4625 | 22.4745 | 22.3623 | 22.258 | 50.2473 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048705 | 0.88106 | 0.042912 | 18.0898 | 22.2019 | 3.464 | 0.04511 | 0.2606 | 0.2726 | 0.1604 | 0.0561
2012-10-30 22:36 | 26.4639 | 22.4655 | 22.5376 | 22.3292 | 22.257 | 50.2438 | 0.006738 | 0.32638 | 0.002199 | 48.4393 | 0.048721 | 0.88153 | 0.042949 | 18.0933 | 22.1527 | 3.4638 | 0.045148 | 0.3128 | 0.3849 | 0.1765 | 0.1043
2012-10-30 22:36 | 26.4617 | 22.4753 | 22.4633 | 22.3791 | 22.2428 | 50.2484 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048881 | 0.88394 | 0.043208 | 18.0835 | 22.1746 | 3.4641 | 0.0454 | 0.3007 | 0.2887 | 0.2045 | 0.0682
2012-10-30 22:36 | 26.4639 | 22.4976 | 22.4655 | 22.3653 | 22.265 | 50.2444 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048717 | 0.88157 | 0.042947 | 18.0957 | 22.1567 | 3.4638 | 0.04514 | 0.3409 | 0.3088 | 0.2086 | 0.1083
2012-10-30 22:36 | 26.4684 | 22.486 | 22.478 | 22.3818 | 22.2615 | 50.2432 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.048777 | 0.88231 | 0.043036 | 18.0886 | 22.1773 | 3.4638 | 0.045228 | 0.3087 | 0.3007 | 0.2045 | 0.0842
2012-10-30 22:36 | 26.4677 | 22.4653 | 22.5014 | 22.3611 | 22.2488 | 50.2374 | 0.00671 | 0.32634 | 0.00219 | 48.635 | 0.048847 | 0.8834 | 0.043152 | 18.0851 | 22.1565 | 3.4634 | 0.045342 | 0.3088 | 0.3449 | 0.2046 | 0.0923
2012-10-30 22:36 | 26.4706 | 22.4842 | 22.5123 | 22.4121 | 22.2477 | 50.2328 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048853 | 0.88343 | 0.043158 | 18.0835 | 22.1795 | 3.463 | 0.045356 | 0.3047 | 0.3328 | 0.2326 | 0.0682
2012-10-30 22:36 | 26.4639 | 22.4655 | 22.4735 | 22.3733 | 22.241 | 50.2328 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048774 | 0.88221 | 0.043029 | 18.0876 | 22.1046 | 3.463 | 0.045225 | 0.3609 | 0.3689 | 0.2687 | 0.1364
2012-10-30 22:36 | 26.4564 | 22.45 | 22.4861 | 22.3618 | 22.2776 | 50.2293 | 0.006713 | 0.32633 | 0.002191 | 48.6111 | 0.048733 | 0.8817 | 0.042968 | 18.0924 | 22.1452 | 3.4628 | 0.045159 | 0.3048 | 0.3409 | 0.2166 | 0.1324
2012-10-30 22:36 | 26.4579 | 22.4234 | 22.4755 | 22.3633 | 22.235 | 50.2444 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.049139 | 0.88811 | 0.043641 | 18.0735 | 22.1387 | 3.4638 | 0.045838 | 0.2847 | 0.3368 | 0.2246 | 0.0963
2012-10-30 22:37 | 26.4692 | 22.4548 | 22.4668 | 22.3826 | 22.2743 | 50.2282 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.04878 | 0.88232 | 0.04304 | 18.0878 | 22.15 | 3.4627 | 0.045233 | 0.3048 | 0.3168 | 0.2326 | 0.1243
2012-10-30 22:37 | 26.4654 | 22.4429 | 22.483 | 22.3668 | 22.2705 | 50.2311 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048916 | 0.88447 | 0.043265 | 18.0815 | 22.1582 | 3.4629 | 0.045459 | 0.2847 | 0.3248 | 0.2086 | 0.1123
2012-10-30 22:37 | 26.4594 | 22.445 | 22.477 | 22.3568 | 22.2445 | 50.2288 | 0.006719 | 0.32643 | 0.002193 | 48.5827 | 0.048878 | 0.88397 | 0.043207 | 18.0852 | 22.1242 | 3.4628 | 0.0454 | 0.3208 | 0.3528 | 0.2326 | 0.1203
2012-10-30 22:37 | 26.4662 | 22.4678 | 22.4558 | 22.3275 | 22.2473 | 50.2397 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048746 | 0.88167 | 0.042978 | 18.087 | 22.131 | 3.4635 | 0.045175 | 0.3368 | 0.3248 | 0.1965 | 0.1163
2012-10-30 22:37 | 26.4602 | 22.4818 | 22.4337 | 22.3335 | 22.2493 | 50.2322 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048749 | 0.88189 | 0.042991 | 18.0904 | 22.149 | 3.463 | 0.045187 | 0.3328 | 0.2847 | 0.1845 | 0.1003
2012-10-30 22:37 | 26.4639 | 22.4294 | 22.4374 | 22.3372 | 22.237 | 50.2253 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048831 | 0.88316 | 0.043126 | 18.086 | 22.1327 | 3.4625 | 0.045322 | 0.2967 | 0.3047 | 0.2045 | 0.1043
2012-10-30 22:37 | 26.4684 | 22.474 | 22.466 | 22.3577 | 22.2174 | 50.2357 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048897 | 0.88427 | 0.043238 | 18.0844 | 22.085 | 3.4632 | 0.045431 | 0.389 | 0.381 | 0.2727 | 0.1324
2012-10-30 22:37 | 26.4617 | 22.4954 | 22.4473 | 22.3791 | 22.2428 | 50.2149 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.048919 | 0.8845 | 0.043269 | 18.081 | 22.1024 | 3.4618 | 0.045466 | 0.393 | 0.3449 | 0.2767 | 0.1404
2012-10-30 22:37 | 26.4647 | 22.4863 | 22.4783 | 22.3661 | 22.2337 | 50.2247 | 0.006719 | 0.32628 | 0.002192 | 48.5612 | 0.048724 | 0.88151 | 0.042951 | 18.0919 | 22.0773 | 3.4625 | 0.045143 | 0.409 | 0.401 | 0.2888 | 0.1564
2012-10-30 22:37 | 26.4594 | 22.4971 | 22.4249 | 22.3528 | 22.2686 | 50.2357 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.048746 | 0.88171 | 0.04298 | 18.0879 | 22.0961 | 3.4632 | 0.045173 | 0.401 | 0.3288 | 0.2567 | 0.1725
2012-10-30 22:37 | 26.4512 | 22.4929 | 22.4328 | 22.3686 | 22.2202 | 50.2264 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.049048 | 0.88671 | 0.043491 | 18.0785 | 22.1039 | 3.4626 | 0.04569 | 0.389 | 0.3289 | 0.2647 | 0.1163
2012-10-30 22:37 | 26.4572 | 22.4828 | 22.4147 | 22.3305 | 22.2142 | 50.2247 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048774 | 0.88225 | 0.043031 | 18.0885 | 22.1059 | 3.4625 | 0.045224 | 0.3769 | 0.3088 | 0.2246 | 0.1083
2012-10-30 22:37 | 26.4542 | 22.4839 | 22.4237 | 22.3476 | 22.2273 | 50.2218 | 0.006713 | 0.32631 | 0.002191 | 48.609 | 0.048828 | 0.88301 | 0.043116 | 18.0842 | 22.1029 | 3.4623 | 0.045307 | 0.381 | 0.3208 | 0.2447 | 0.1244
2012-10-30 22:37 | 26.4482 | 22.4779 | 22.3977 | 22.3135 | 22.2333 | 50.2311 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.048802 | 0.88264 | 0.043074 | 18.0861 | 22.0929 | 3.4629 | 0.045267 | 0.385 | 0.3048 | 0.2206 | 0.1404
2012-10-30 22:37 | 26.4512 | 22.4969 | 22.4007 | 22.3405 | 22.2042 | 50.2305 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048831 | 0.88314 | 0.043125 | 18.0857 | 22.0959 | 3.4629 | 0.045317 | 0.401 | 0.3048 | 0.2446 | 0.1083
2012-10-30 22:37 | 26.4557 | 22.4894 | 22.4292 | 22.3771 | 22.1967 | 50.2103 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048755 | 0.8819 | 0.042997 | 18.0884 | 22.0763 | 3.4615 | 0.045192 | 0.4131 | 0.3529 | 0.3008 | 0.1204
2012-10-30 22:37 | 26.4482 | 22.5099 | 22.4217 | 22.3656 | 22.2614 | 50.2345 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048777 | 0.88222 | 0.043032 | 18.0868 | 22.125 | 3.4632 | 0.045225 | 0.3849 | 0.2967 | 0.2406 | 0.1364
2012-10-30 22:37 | 26.4422 | 22.4839 | 22.4478 | 22.3155 | 22.1952 | 50.2328 | 0.00671 | 0.32634 | 0.00219 | 48.635 | 0.048859 | 0.88355 | 0.043169 | 18.0836 | 22.0147 | 3.463 | 0.045359 | 0.4692 | 0.4331 | 0.3008 | 0.1805
2012-10-30 22:37 | 26.4527 | 22.5144 | 22.4062 | 22.342 | 22.2258 | 50.2316 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048699 | 0.88114 | 0.04291 | 18.0935 | 22.0894 | 3.463 | 0.045107 | 0.425 | 0.3168 | 0.2526 | 0.1364
2012-10-30 22:37 | 26.4534 | 22.5031 | 22.3949 | 22.3909 | 22.2265 | 50.2345 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048891 | 0.88417 | 0.043228 | 18.0845 | 22.066 | 3.4632 | 0.045421 | 0.4371 | 0.3289 | 0.3249 | 0.1605
2012-10-30 22:37 | 26.4445 | 22.4782 | 22.378 | 22.3659 | 22.1895 | 50.2218 | 0.006741 | 0.32634 | 0.0022 | 48.4114 | 0.048938 | 0.88488 | 0.043304 | 18.0816 | 22.0491 | 3.4623 | 0.045504 | 0.4291 | 0.3289 | 0.3168 | 0.1404
2012-10-30 22:37 | 26.4482 | 22.4939 | 22.4017 | 22.3536 | 22.2092 | 50.2201 | 0.0067 | 0.32634 | 0.002186 | 48.7076 | 0.048721 | 0.88145 | 0.042945 | 18.0919 | 22.0929 | 3.4622 | 0.045131 | 0.401 | 0.3088 | 0.2607 | 0.1163
2012-10-30 22:37 | 26.4422 | 22.4678 | 22.4398 | 22.3115 | 22.2153 | 50.2247 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048916 | 0.88459 | 0.043271 | 18.0838 | 22.0548 | 3.4625 | 0.045469 | 0.413 | 0.385 | 0.2567 | 0.1605
2012-10-30 22:37 | 26.4407 | 22.4623 | 22.4102 | 22.3341 | 22.2258 | 50.2201 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.048736 | 0.88141 | 0.042956 | 18.0854 | 22.0934 | 3.4622 | 0.045154 | 0.3689 | 0.3168 | 0.2407 | 0.1324
2012-10-30 22:37 | 26.4377 | 22.4193 | 22.4112 | 22.3511 | 22.2268 | 50.2282 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048752 | 0.88182 | 0.04299 | 18.0878 | 22.1145 | 3.4627 | 0.045186 | 0.3048 | 0.2967 | 0.2366 | 0.1123
2012-10-30 22:37 | 26.437 | 22.4747 | 22.4105 | 22.3344 | 22.19 | 50.2253 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.04878 | 0.88232 | 0.04304 | 18.0878 | 22.0777 | 3.4625 | 0.045236 | 0.397 | 0.3328 | 0.2567 | 0.1123
2012-10-30 22:38 | 26.4474 | 22.4971 | 22.437 | 22.3167 | 22.2726 | 50.2253 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048812 | 0.88278 | 0.04309 | 18.0854 | 22.0961 | 3.4625 | 0.045287 | 0.401 | 0.3409 | 0.2206 | 0.1765
2012-10-30 22:38 | 26.4519 | 22.4896 | 22.4094 | 22.3613 | 22.241 | 50.2264 | 0.006744 | 0.32634 | 0.002201 | 48.3898 | 0.048982 | 0.88541 | 0.043369 | 18.0763 | 22.1046 | 3.4626 | 0.04557 | 0.385 | 0.3048 | 0.2567 | 0.1364
2012-10-30 22:38 | 26.4474 | 22.441 | 22.4049 | 22.3728 | 22.2205 | 50.2345 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.048762 | 0.88209 | 0.043012 | 18.0897 | 22.0961 | 3.4632 | 0.045207 | 0.3449 | 0.3088 | 0.2767 | 0.1244
2012-10-30 22:38 | 26.4482 | 22.4658 | 22.4137 | 22.3456 | 22.2213 | 50.2114 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048743 | 0.88166 | 0.042975 | 18.0878 | 22.113 | 3.4616 | 0.045172 | 0.3528 | 0.3007 | 0.2326 | 0.1083
2012-10-30 22:38 | 26.4422 | 22.4638 | 22.3957 | 22.3837 | 22.2153 | 50.2218 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048809 | 0.88252 | 0.043075 | 18.0812 | 22.1511 | 3.4623 | 0.045269 | 0.3127 | 0.2446 | 0.2326 | 0.0642
2012-10-30 22:38 | 26.4497 | 22.4473 | 22.4232 | 22.3431 | 22.2067 | 50.2247 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048743 | 0.88182 | 0.042982 | 18.0911 | 22.0944 | 3.4625 | 0.045179 | 0.3529 | 0.3288 | 0.2487 | 0.1123
2012-10-30 22:38 | 26.4504 | 22.4199 | 22.4199 | 22.3839 | 22.2315 | 50.2224 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048894 | 0.8844 | 0.043242 | 18.0881 | 22.1031 | 3.4623 | 0.045436 | 0.3168 | 0.3168 | 0.2808 | 0.1284
2012-10-30 22:38 | 26.4512 | 22.4368 | 22.4328 | 22.3606 | 22.2162 | 50.2311 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.048749 | 0.88186 | 0.04299 | 18.0898 | 22.1561 | 3.4629 | 0.045181 | 0.2807 | 0.2767 | 0.2045 | 0.0601
2012-10-30 22:38 | 26.4467 | 22.4483 | 22.4242 | 22.3481 | 22.2559 | 50.2311 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048865 | 0.88379 | 0.043187 | 18.0864 | 22.1476 | 3.4629 | 0.045384 | 0.3007 | 0.2766 | 0.2005 | 0.1083
2012-10-30 22:38 | 26.4572 | 22.4468 | 22.4067 | 22.3546 | 22.2664 | 50.2264 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048812 | 0.8828 | 0.043091 | 18.0857 | 22.142 | 3.4626 | 0.045287 | 0.3048 | 0.2647 | 0.2126 | 0.1244
2012-10-30 22:38 | 26.4504 | 22.468 | 22.4159 | 22.3478 | 22.2315 | 50.2155 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048802 | 0.88283 | 0.043084 | 18.09 | 22.1312 | 3.4619 | 0.045281 | 0.3368 | 0.2847 | 0.2166 | 0.1003
2012-10-30 22:38 | 26.4489 | 22.4104 | 22.4745 | 22.3262 | 22.2621 | 50.2351 | 0.006735 | 0.32643 | 0.002198 | 48.4673 | 0.048796 | 0.88245 | 0.04306 | 18.0845 | 22.1457 | 3.4632 | 0.045258 | 0.2647 | 0.3288 | 0.1805 | 0.1164
2012-10-30 22:38 | 26.4557 | 22.4453 | 22.4493 | 22.3611 | 22.2528 | 50.2201 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048994 | 0.8858 | 0.043399 | 18.0798 | 22.1886 | 3.4622 | 0.045596 | 0.2567 | 0.2607 | 0.1725 | 0.0642
2012-10-30 22:38 | 26.4512 | 22.4768 | 22.4448 | 22.3325 | 22.2443 | 50.2322 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048806 | 0.88254 | 0.043073 | 18.0826 | 22.0959 | 3.463 | 0.045268 | 0.3809 | 0.3489 | 0.2366 | 0.1484
2012-10-30 22:38 | 26.4549 | 22.4685 | 22.4685 | 22.3723 | 22.232 | 50.2224 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048749 | 0.88173 | 0.042983 | 18.0871 | 22.1237 | 3.4623 | 0.045178 | 0.3448 | 0.3448 | 0.2486 | 0.1083
2012-10-30 22:38 | 26.4497 | 22.4353 | 22.4553 | 22.3431 | 22.1987 | 50.2201 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048733 | 0.88161 | 0.042964 | 18.0907 | 22.1145 | 3.4622 | 0.045162 | 0.3208 | 0.3408 | 0.2286 | 0.0842
2012-10-30 22:38 | 26.4467 | 22.4603 | 22.4403 | 22.328 | 22.2198 | 50.2212 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048809 | 0.88268 | 0.043083 | 18.0844 | 22.0994 | 3.4622 | 0.045282 | 0.3609 | 0.3409 | 0.2286 | 0.1204
2012-10-30 22:38 | 26.4519 | 22.4134 | 22.4335 | 22.3252 | 22.229 | 50.2247 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048724 | 0.88138 | 0.042944 | 18.0893 | 22.1487 | 3.4625 | 0.045141 | 0.2647 | 0.2848 | 0.1765 | 0.0803
2012-10-30 22:38 | 26.4549 | 22.4405 | 22.4525 | 22.3442 | 22.236 | 50.2195 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.04878 | 0.88232 | 0.04304 | 18.0878 | 22.1197 | 3.4621 | 0.045236 | 0.3208 | 0.3328 | 0.2245 | 0.1163
2012-10-30 22:38 | 26.4594 | 22.5011 | 22.445 | 22.3487 | 22.2244 | 50.2212 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.048673 | 0.88054 | 0.042859 | 18.091 | 22.1522 | 3.4622 | 0.045059 | 0.3489 | 0.2928 | 0.1965 | 0.0722
2012-10-30 22:38 | 26.4549 | 22.4445 | 22.4605 | 22.3362 | 22.1959 | 50.2264 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048812 | 0.88301 | 0.043102 | 18.0901 | 22.1197 | 3.4626 | 0.045299 | 0.3248 | 0.3408 | 0.2165 | 0.0762
2012-10-30 22:38 | 26.4512 | 22.4488 | 22.4608 | 22.3325 | 22.2243 | 50.2201 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048853 | 0.88349 | 0.043161 | 18.0847 | 22.1521 | 3.4622 | 0.045359 | 0.2967 | 0.3087 | 0.1804 | 0.0722
2012-10-30 22:38 | 26.4617 | 22.4833 | 22.4673 | 22.3591 | 22.2107 | 50.2201 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.048736 | 0.88171 | 0.042971 | 18.0916 | 22.1305 | 3.4622 | 0.045169 | 0.3528 | 0.3368 | 0.2286 | 0.0802
2012-10-30 22:38 | 26.4489 | 22.4585 | 22.4305 | 22.3222 | 22.1939 | 50.2201 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048815 | 0.88294 | 0.043101 | 18.0875 | 22.1056 | 3.4622 | 0.045297 | 0.3529 | 0.3249 | 0.2166 | 0.0883
2012-10-30 22:38 | 26.4549 | 22.4565 | 22.4365 | 22.3362 | 22.248 | 50.223 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.048909 | 0.88437 | 0.043254 | 18.082 | 22.1036 | 3.4624 | 0.045452 | 0.3529 | 0.3329 | 0.2326 | 0.1444
2012-10-30 22:39 | 26.4474 | 22.465 | 22.429 | 22.2967 | 22.2044 | 50.2293 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048856 | 0.88349 | 0.043164 | 18.0836 | 22.1202 | 3.4628 | 0.045361 | 0.3448 | 0.3088 | 0.1765 | 0.0842
2012-10-30 22:39 | 26.4579 | 22.4555 | 22.4034 | 22.3513 | 22.2269 | 50.2345 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.048749 | 0.88177 | 0.042986 | 18.088 | 22.1307 | 3.4632 | 0.045184 | 0.3248 | 0.2727 | 0.2206 | 0.0962
2012-10-30 22:39 | 26.4497 | 22.4272 | 22.4513 | 22.3511 | 22.2468 | 50.227 | 0.006725 | 0.32641 | 0.002195 | 48.5372 | 0.04874 | 0.88173 | 0.042975 | 18.0904 | 22.0824 | 3.4626 | 0.04517 | 0.3448 | 0.3689 | 0.2687 | 0.1644
2012-10-30 22:39 | 26.4534 | 22.451 | 22.4309 | 22.3427 | 22.2465 | 50.223 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.048793 | 0.88265 | 0.043067 | 18.0897 | 22.1262 | 3.4624 | 0.045267 | 0.3248 | 0.3047 | 0.2165 | 0.1203
2012-10-30 22:39 | 26.4504 | 22.46 | 22.3959 | 22.3438 | 22.2515 | 50.2172 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.049042 | 0.88664 | 0.043483 | 18.0792 | 22.1071 | 3.462 | 0.04568 | 0.3529 | 0.2888 | 0.2367 | 0.1444
2012-10-30 22:39 | 26.4459 | 22.4595 | 22.4515 | 22.3072 | 22.239 | 50.2241 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048972 | 0.88515 | 0.043348 | 18.0747 | 22.0946 | 3.4624 | 0.045542 | 0.3649 | 0.3569 | 0.2126 | 0.1444
2012-10-30 22:39 | 26.4445 | 22.4421 | 22.414 | 22.3379 | 22.2256 | 50.2201 | 0.006732 | 0.32644 | 0.002198 | 48.4911 | 0.04879 | 0.88228 | 0.043046 | 18.0832 | 22.1253 | 3.4622 | 0.045244 | 0.3168 | 0.2887 | 0.2126 | 0.1003
2012-10-30 22:39 | 26.4422 | 22.4839 | 22.4438 | 22.3275 | 22.1992 | 50.2293 | 0.006738 | 0.32647 | 0.0022 | 48.4522 | 0.048869 | 0.88398 | 0.043199 | 18.0888 | 22.1431 | 3.4628 | 0.045399 | 0.3408 | 0.3007 | 0.1844 | 0.0561
2012-10-30 22:39 | 26.4437 | 22.4613 | 22.3852 | 22.313 | 22.2007 | 50.2264 | 0.006744 | 0.3264 | 0.002201 | 48.3984 | 0.048853 | 0.88352 | 0.043163 | 18.0853 | 22.0884 | 3.4626 | 0.045364 | 0.3729 | 0.2968 | 0.2246 | 0.1123
2012-10-30 22:39 | 26.4437 | 22.4293 | 22.4172 | 22.321 | 22.1726 | 50.2247 | 0.006713 | 0.32638 | 0.002191 | 48.6197 | 0.048843 | 0.88335 | 0.043145 | 18.0854 | 22.0884 | 3.4625 | 0.045336 | 0.3409 | 0.3288 | 0.2326 | 0.0842
2012-10-30 22:39 | 26.4527 | 22.4623 | 22.4543 | 22.326 | 22.2378 | 50.2212 | 0.006741 | 0.32643 | 0.0022 | 48.4241 | 0.048862 | 0.88382 | 0.043185 | 18.0881 | 22.0934 | 3.4622 | 0.045385 | 0.3689 | 0.3609 | 0.2326 | 0.1444
2012-10-30 22:39 | 26.4497 | 22.4633 | 22.4393 | 22.3591 | 22.2468 | 50.2218 | 0.006729 | 0.32646 | 0.002197 | 48.5148 | 0.048796 | 0.88249 | 0.043062 | 18.0854 | 22.0703 | 3.4623 | 0.045259 | 0.393 | 0.369 | 0.2888 | 0.1765
2012-10-30 22:39 | 26.4482 | 22.4698 | 22.4498 | 22.3416 | 22.2453 | 50.2149 | 0.006725 | 0.32644 | 0.002195 | 48.5416 | 0.048746 | 0.88164 | 0.042977 | 18.0864 | 22.0769 | 3.4618 | 0.045172 | 0.3929 | 0.3729 | 0.2647 | 0.1684
2012-10-30 22:39 | 26.4564 | 22.47 | 22.4299 | 22.3498 | 22.2295 | 50.2103 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.04885 | 0.88355 | 0.043161 | 18.087 | 22.1292 | 3.4615 | 0.045355 | 0.3408 | 0.3007 | 0.2206 | 0.1003
2012-10-30 22:39 | 26.4557 | 22.4773 | 22.4493 | 22.317 | 22.2007 | 50.2166 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.048727 | 0.88151 | 0.042953 | 18.0908 | 22.1205 | 3.4619 | 0.045151 | 0.3568 | 0.3288 | 0.1965 | 0.0802
2012-10-30 22:39 | 26.4542 | 22.4438 | 22.4678 | 22.3235 | 22.2072 | 50.2149 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048743 | 0.88176 | 0.042979 | 18.0899 | 22.1069 | 3.4618 | 0.045176 | 0.3369 | 0.3609 | 0.2166 | 0.1003
2012-10-30 22:39 | 26.4602 | 22.4778 | 22.4578 | 22.3175 | 22.2292 | 50.2155 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048746 | 0.8818 | 0.042984 | 18.0897 | 22.0889 | 3.4619 | 0.045183 | 0.3889 | 0.3689 | 0.2286 | 0.1403
2012-10-30 22:39 | 26.4534 | 22.443 | 22.463 | 22.3548 | 22.2064 | 50.227 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048887 | 0.88387 | 0.04321 | 18.0798 | 22.0981 | 3.4626 | 0.045404 | 0.3449 | 0.3649 | 0.2567 | 0.1083
2012-10-30 22:39 | 26.4549 | 22.4365 | 22.3964 | 22.3242 | 22.2199 | 50.2253 | 0.006713 | 0.32637 | 0.002191 | 48.6176 | 0.048699 | 0.88095 | 0.042901 | 18.0897 | 22.1277 | 3.4625 | 0.045092 | 0.3088 | 0.2687 | 0.1965 | 0.0922
2012-10-30 22:39 | 26.4527 | 22.4503 | 22.4142 | 22.2659 | 22.1736 | 50.2236 | 0.006722 | 0.32631 | 0.002193 | 48.5439 | 0.048881 | 0.88379 | 0.043201 | 18.0805 | 22.1175 | 3.4624 | 0.045394 | 0.3328 | 0.2967 | 0.1484 | 0.0561
2012-10-30 22:39 | 26.4579 | 22.4274 | 22.4154 | 22.3553 | 22.243 | 50.2207 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048831 | 0.88316 | 0.043126 | 18.086 | 22.1267 | 3.4622 | 0.045321 | 0.3007 | 0.2887 | 0.2286 | 0.1163
2012-10-30 22:39 | 26.4572 | 22.4628 | 22.4428 | 22.3826 | 22.2423 | 50.216 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.048815 | 0.88283 | 0.043095 | 18.0851 | 22.126 | 3.4619 | 0.045289 | 0.3368 | 0.3168 | 0.2566 | 0.1163
2012-10-30 22:39 | 26.4549 | 22.4204 | 22.4124 | 22.2961 | 22.256 | 50.212 | 0.006738 | 0.32646 | 0.0022 | 48.45 | 0.048944 | 0.88491 | 0.043311 | 18.08 | 22.1558 | 3.4616 | 0.045511 | 0.2646 | 0.2566 | 0.1403 | 0.1002
2012-10-30 22:39 | 26.4579 | 22.4274 | 22.4395 | 22.3593 | 22.2189 | 50.2085 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048743 | 0.88163 | 0.042973 | 18.0873 | 22.1788 | 3.4614 | 0.045169 | 0.2486 | 0.2607 | 0.1805 | 0.0401
2012-10-30 22:39 | 26.4602 | 22.4097 | 22.4418 | 22.3576 | 22.2292 | 50.2247 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048762 | 0.88225 | 0.04302 | 18.0929 | 22.1651 | 3.4625 | 0.045216 | 0.2446 | 0.2767 | 0.1925 | 0.0641
2012-10-30 22:39 | 26.4654 | 22.447 | 22.4269 | 22.3307 | 22.2826 | 50.2091 | 0.006741 | 0.32636 | 0.0022 | 48.4134 | 0.049164 | 0.8885 | 0.043682 | 18.0722 | 22.1783 | 3.4614 | 0.045882 | 0.2687 | 0.2486 | 0.1524 | 0.1043
2012-10-30 22:39 | 26.4609 | 22.4144 | 22.4625 | 22.3543 | 22.1858 | 50.2074 | 0.006729 | 0.32641 | 0.002196 | 48.5084 | 0.048758 | 0.88194 | 0.043002 | 18.0882 | 22.1377 | 3.4613 | 0.045198 | 0.2767 | 0.3248 | 0.2166 | 0.0481
2012-10-30 22:40 | 26.4609 | 22.4344 | 22.4384 | 22.3302 | 22.242 | 50.2033 | 0.00671 | 0.32638 | 0.00219 | 48.6414 | 0.0489 | 0.88417 | 0.043236 | 18.0812 | 22.1056 | 3.461 | 0.045426 | 0.3288 | 0.3328 | 0.2246 | 0.1364
2012-10-30 22:40 | 26.4594 | 22.3848 | 22.457 | 22.3167 | 22.2124 | 50.2218 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.049098 | 0.88745 | 0.043572 | 18.0751 | 22.1603 | 3.4623 | 0.045766 | 0.2245 | 0.2967 | 0.1564 | 0.0521
2012-10-30 22:40 | 26.4602 | 22.4377 | 22.4297 | 22.3415 | 22.2172 | 50.2051 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.048856 | 0.88345 | 0.043162 | 18.0827 | 22.1169 | 3.4611 | 0.045363 | 0.3208 | 0.3128 | 0.2246 | 0.1003
2012-10-30 22:40 | 26.4602 | 22.4538 | 22.4458 | 22.3896 | 22.2172 | 50.2103 | 0.006713 | 0.3264 | 0.002191 | 48.6219 | 0.048809 | 0.88272 | 0.043085 | 18.0853 | 22.129 | 3.4615 | 0.045276 | 0.3248 | 0.3168 | 0.2606 | 0.0882
2012-10-30 22:40 | 26.4519 | 22.4575 | 22.4615 | 22.3092 | 22.2651 | 50.2172 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.04885 | 0.8834 | 0.043154 | 18.084 | 22.0806 | 3.462 | 0.04535 | 0.3769 | 0.3809 | 0.2286 | 0.1845
2012-10-30 22:40 | 26.4632 | 22.4407 | 22.4608 | 22.3004 | 22.2162 | 50.2091 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048812 | 0.88303 | 0.043102 | 18.0904 | 22.1199 | 3.4614 | 0.045297 | 0.3208 | 0.3409 | 0.1805 | 0.0963
2012-10-30 22:40 | 26.4587 | 22.4523 | 22.4483 | 22.332 | 22.2398 | 50.2212 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048821 | 0.8829 | 0.043104 | 18.0844 | 22.1154 | 3.4622 | 0.0453 | 0.3369 | 0.3329 | 0.2166 | 0.1244
2012-10-30 22:40 | 26.4557 | 22.4373 | 22.4894 | 22.3851 | 22.2408 | 50.212 | 0.006741 | 0.32641 | 0.0022 | 48.422 | 0.04868 | 0.88075 | 0.042875 | 18.0926 | 22.0563 | 3.4616 | 0.045075 | 0.381 | 0.4331 | 0.3288 | 0.1845
2012-10-30 22:40 | 26.4624 | 22.46 | 22.4279 | 22.3678 | 22.2114 | 50.2126 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048812 | 0.88278 | 0.04309 | 18.0854 | 22.1191 | 3.4617 | 0.045284 | 0.3409 | 0.3088 | 0.2487 | 0.0923
2012-10-30 22:40 | 26.4594 | 22.4329 | 22.453 | 22.3287 | 22.2044 | 50.2097 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048793 | 0.88255 | 0.043062 | 18.0877 | 22.1442 | 3.4615 | 0.045259 | 0.2887 | 0.3088 | 0.1845 | 0.0602
2012-10-30 22:40 | 26.4564 | 22.438 | 22.462 | 22.3377 | 22.2656 | 50.2172 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.04874 | 0.8818 | 0.042979 | 18.0919 | 22.1452 | 3.462 | 0.045172 | 0.2928 | 0.3168 | 0.1925 | 0.1204
2012-10-30 22:40 | 26.4534 | 22.451 | 22.459 | 22.3427 | 22.2224 | 50.2149 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048865 | 0.88384 | 0.043189 | 18.0873 | 22.1342 | 3.4618 | 0.045383 | 0.3168 | 0.3248 | 0.2085 | 0.0882
2012-10-30 22:40 | 26.4557 | 22.4413 | 22.4653 | 22.333 | 22.2328 | 50.1987 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048869 | 0.88374 | 0.043187 | 18.0838 | 22.1285 | 3.4607 | 0.045381 | 0.3128 | 0.3368 | 0.2045 | 0.1043
2012-10-30 22:40 | 26.4602 | 22.4538 | 22.4658 | 22.3616 | 22.2453 | 50.2166 | 0.006738 | 0.32643 | 0.002199 | 48.4457 | 0.048711 | 0.88118 | 0.042923 | 18.0899 | 22.125 | 3.4619 | 0.045122 | 0.3288 | 0.3408 | 0.2366 | 0.1203
2012-10-30 22:40 | 26.4557 | 22.4252 | 22.4052 | 22.337 | 22.2247 | 50.2045 | 0.006707 | 0.32627 | 0.002188 | 48.6459 | 0.048878 | 0.88363 | 0.04319 | 18.0784 | 22.1245 | 3.4611 | 0.045378 | 0.3007 | 0.2807 | 0.2125 | 0.1002
2012-10-30 22:40 | 26.4467 | 22.4443 | 22.4443 | 22.308 | 22.2518 | 50.2103 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048853 | 0.88355 | 0.043164 | 18.0859 | 22.1074 | 3.4615 | 0.045362 | 0.3369 | 0.3369 | 0.2006 | 0.1444
2012-10-30 22:40 | 26.4459 | 22.4675 | 22.4756 | 22.3112 | 22.219 | 50.2028 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048837 | 0.88338 | 0.043141 | 18.0882 | 22.0946 | 3.461 | 0.045338 | 0.3729 | 0.381 | 0.2166 | 0.1244
2012-10-30 22:40 | 26.4452 | 22.4308 | 22.4508 | 22.3386 | 22.2343 | 50.2218 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048711 | 0.88127 | 0.042927 | 18.0917 | 22.13 | 3.4623 | 0.045122 | 0.3008 | 0.3208 | 0.2086 | 0.1043
2012-10-30 22:40 | 26.4459 | 22.4595 | 22.4635 | 22.3874 | 22.235 | 50.2091 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048796 | 0.88241 | 0.043058 | 18.0836 | 22.1307 | 3.4614 | 0.045256 | 0.3288 | 0.3328 | 0.2567 | 0.1043
2012-10-30 22:40 | 26.4347 | 22.4042 | 22.4764 | 22.3802 | 22.2238 | 50.2166 | 0.006738 | 0.32641 | 0.002199 | 48.4436 | 0.048758 | 0.8821 | 0.04301 | 18.0915 | 22.1476 | 3.4619 | 0.045209 | 0.2566 | 0.3288 | 0.2326 | 0.0762
2012-10-30 22:40 | 26.4362 | 22.4178 | 22.4699 | 22.3215 | 22.2454 | 50.212 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048891 | 0.8842 | 0.043229 | 18.0851 | 22.121 | 3.4616 | 0.045424 | 0.2968 | 0.3489 | 0.2005 | 0.1244
2012-10-30 22:40 | 26.4317 | 22.4213 | 22.4012 | 22.3451 | 22.2328 | 50.2137 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048878 | 0.88378 | 0.043197 | 18.0813 | 22.1285 | 3.4617 | 0.045392 | 0.2928 | 0.2727 | 0.2166 | 0.1043
2012-10-30 22:40 | 26.434 | 22.4156 | 22.4677 | 22.3514 | 22.2351 | 50.2126 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.04895 | 0.88514 | 0.043328 | 18.0825 | 22.1148 | 3.4617 | 0.045527 | 0.3008 | 0.3529 | 0.2366 | 0.1203
2012-10-30 22:40 | 26.4332 | 22.4067 | 22.4749 | 22.3827 | 22.2343 | 50.2103 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048784 | 0.88258 | 0.043056 | 18.0916 | 22.0699 | 3.4615 | 0.045251 | 0.3368 | 0.405 | 0.3128 | 0.1644
2012-10-30 22:40 | 26.4377 | 22.4152 | 22.4954 | 22.3471 | 22.2148 | 50.2189 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048843 | 0.88336 | 0.043146 | 18.0857 | 22.1345 | 3.4621 | 0.045343 | 0.2807 | 0.3609 | 0.2126 | 0.0803
2012-10-30 22:40 | 26.4302 | 22.4198 | 22.4639 | 22.3957 | 22.2634 | 50.2033 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048724 | 0.88115 | 0.042933 | 18.0845 | 22.1431 | 3.461 | 0.045132 | 0.2767 | 0.3208 | 0.2526 | 0.1203
2012-10-30 22:41 | 26.4355 | 22.4331 | 22.409 | 22.373 | 22.2086 | 50.2114 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.049023 | 0.88618 | 0.043443 | 18.0768 | 22.0962 | 3.4616 | 0.04564 | 0.3369 | 0.3128 | 0.2768 | 0.1124
2012-10-30 22:41 | 26.4175 | 22.4391 | 22.4592 | 22.3389 | 22.2507 | 50.2033 | 0.006719 | 0.32624 | 0.002192 | 48.5548 | 0.048777 | 0.88235 | 0.043038 | 18.0894 | 22.0823 | 3.461 | 0.04523 | 0.3568 | 0.3769 | 0.2566 | 0.1684
2012-10-30 22:41 | 26.4265 | 22.4521 | 22.4441 | 22.3319 | 22.2517 | 50.2016 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.0489 | 0.88444 | 0.043249 | 18.0868 | 22.1193 | 3.4609 | 0.045444 | 0.3328 | 0.3248 | 0.2126 | 0.1324
2012-10-30 22:41 | 26.4257 | 22.4113 | 22.4393 | 22.3632 | 22.2028 | 50.2143 | 0.006716 | 0.32631 | 0.002192 | 48.5873 | 0.048812 | 0.88281 | 0.043092 | 18.086 | 22.1025 | 3.4618 | 0.045284 | 0.3088 | 0.3368 | 0.2607 | 0.1003
2012-10-30 22:41 | 26.407 | 22.4286 | 22.4407 | 22.3084 | 22.2322 | 50.212 | 0.006713 | 0.32636 | 0.002191 | 48.6154 | 0.049004 | 0.88579 | 0.043407 | 18.0758 | 22.1399 | 3.4616 | 0.045598 | 0.2887 | 0.3008 | 0.1685 | 0.0923
2012-10-30 22:41 | 26.416 | 22.4537 | 22.4376 | 22.3334 | 22.2532 | 50.2178 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048925 | 0.88499 | 0.043298 | 18.0888 | 22.0808 | 3.462 | 0.045495 | 0.3729 | 0.3568 | 0.2526 | 0.1724
2012-10-30 22:41 | 26.4153 | 22.4369 | 22.449 | 22.3247 | 22.2445 | 50.2201 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.048837 | 0.88327 | 0.043136 | 18.0862 | 22.1643 | 3.4622 | 0.045328 | 0.2726 | 0.2847 | 0.1604 | 0.0802
2012-10-30 22:41 | 26.4085 | 22.4702 | 22.4863 | 22.346 | 22.2297 | 50.2033 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048831 | 0.88332 | 0.043133 | 18.0893 | 22.1053 | 3.461 | 0.04533 | 0.3649 | 0.381 | 0.2407 | 0.1244
2012-10-30 22:41 | 26.413 | 22.4346 | 22.4587 | 22.3825 | 22.2302 | 50.1952 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048966 | 0.8853 | 0.043349 | 18.0798 | 22.1259 | 3.4605 | 0.045545 | 0.3087 | 0.3328 | 0.2566 | 0.1043
2012-10-30 22:41 | 26.416 | 22.4176 | 22.4777 | 22.3374 | 22.2292 | 50.2056 | 0.006735 | 0.3264 | 0.002198 | 48.4631 | 0.04873 | 0.88145 | 0.042953 | 18.0885 | 22.1088 | 3.4612 | 0.045151 | 0.3088 | 0.3689 | 0.2286 | 0.1204
2012-10-30 22:41 | 26.4175 | 22.4191 | 22.4431 | 22.2988 | 22.2347 | 50.2004 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.048743 | 0.88186 | 0.042984 | 18.092 | 22.1063 | 3.4608 | 0.045176 | 0.3128 | 0.3368 | 0.1925 | 0.1284
2012-10-30 22:41 | 26.4145 | 22.384 | 22.4321 | 22.3159 | 22.2116 | 50.2051 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048913 | 0.88468 | 0.043272 | 18.0867 | 22.1314 | 3.4611 | 0.045468 | 0.2526 | 0.3007 | 0.1845 | 0.0802
2012-10-30 22:41 | 26.4115 | 22.3971 | 22.4612 | 22.3971 | 22.2848 | 50.2108 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048743 | 0.88182 | 0.042982 | 18.0911 | 22.1124 | 3.4615 | 0.045177 | 0.2847 | 0.3488 | 0.2847 | 0.1724
2012-10-30 22:41 | 26.4108 | 22.4284 | 22.4525 | 22.3362 | 22.2159 | 50.2189 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048903 | 0.88423 | 0.043241 | 18.0813 | 22.0996 | 3.4621 | 0.045438 | 0.3288 | 0.3529 | 0.2366 | 0.1163
2012-10-30 22:41 | 26.413 | 22.4186 | 22.4467 | 22.3465 | 22.1981 | 50.2091 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.04873 | 0.88164 | 0.042962 | 18.0924 | 22.0978 | 3.4614 | 0.045156 | 0.3208 | 0.3489 | 0.2487 | 0.1003
2012-10-30 22:41 | 26.4145 | 22.4401 | 22.4281 | 22.3359 | 22.2076 | 50.2108 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048831 | 0.8833 | 0.043133 | 18.089 | 22.1073 | 3.4615 | 0.045329 | 0.3328 | 0.3208 | 0.2286 | 0.1003
2012-10-30 22:41 | 26.4153 | 22.4409 | 22.449 | 22.3287 | 22.2405 | 50.1964 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048922 | 0.88456 | 0.043274 | 18.081 | 22.1001 | 3.4605 | 0.045473 | 0.3408 | 0.3489 | 0.2286 | 0.1404
2012-10-30 22:41 | 26.4153 | 22.4209 | 22.4329 | 22.3247 | 22.2365 | 50.2103 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048762 | 0.8821 | 0.043013 | 18.09 | 22.1162 | 3.4615 | 0.045209 | 0.3047 | 0.3167 | 0.2085 | 0.1203
2012-10-30 22:41 | 26.401 | 22.4507 | 22.4347 | 22.3304 | 22.2021 | 50.2062 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048799 | 0.88222 | 0.043051 | 18.0786 | 22.0698 | 3.4612 | 0.045244 | 0.3809 | 0.3649 | 0.2606 | 0.1323
2012-10-30 22:41 | 26.404 | 22.4176 | 22.4296 | 22.3375 | 22.2412 | 50.2056 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048721 | 0.88127 | 0.042936 | 18.088 | 22.0808 | 3.4612 | 0.045129 | 0.3368 | 0.3488 | 0.2567 | 0.1604
2012-10-30 22:41 | 26.4085 | 22.4181 | 22.4782 | 22.3299 | 22.2217 | 50.2056 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.048714 | 0.88119 | 0.042926 | 18.0891 | 22.0893 | 3.4612 | 0.045124 | 0.3288 | 0.3889 | 0.2406 | 0.1324
2012-10-30 22:41 | 26.4033 | 22.4049 | 22.3768 | 22.3167 | 22.1924 | 50.208 | 0.00671 | 0.32633 | 0.00219 | 48.6328 | 0.048865 | 0.88372 | 0.043183 | 18.085 | 22.064 | 3.4613 | 0.045373 | 0.3409 | 0.3128 | 0.2527 | 0.1284
2012-10-30 22:41 | 26.4025 | 22.4081 | 22.4442 | 22.3159 | 22.1956 | 50.2016 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048982 | 0.88567 | 0.043382 | 18.0816 | 22.0231 | 3.4609 | 0.045579 | 0.385 | 0.4211 | 0.2928 | 0.1725
2012-10-30 22:41 | 26.4063 | 22.4119 | 22.436 | 22.3478 | 22.2315 | 50.197 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048966 | 0.88537 | 0.043353 | 18.0813 | 22.0069 | 3.4606 | 0.045549 | 0.405 | 0.4291 | 0.3409 | 0.2246
2012-10-30 22:41 | 26.4003 | 22.4139 | 22.4219 | 22.3538 | 22.2095 | 50.2091 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.048872 | 0.88388 | 0.043197 | 18.0856 | 22.0891 | 3.4614 | 0.04539 | 0.3248 | 0.3328 | 0.2647 | 0.1204
2012-10-30 22:41 | 26.3913 | 22.4129 | 22.4169 | 22.3328 | 22.1884 | 50.1987 | 0.006713 | 0.32614 | 0.002189 | 48.583 | 0.048887 | 0.88423 | 0.043227 | 18.0872 | 22.044 | 3.4607 | 0.045416 | 0.3689 | 0.3729 | 0.2888 | 0.1444
2012-10-30 22:41 | 26.4063 | 22.444 | 22.4079 | 22.3197 | 22.2596 | 50.1947 | 0.006716 | 0.32628 | 0.002191 | 48.5829 | 0.04873 | 0.8816 | 0.04296 | 18.0915 | 22.067 | 3.4604 | 0.045151 | 0.377 | 0.3409 | 0.2527 | 0.1926
2012-10-30 22:41 | 26.4003 | 22.4099 | 22.43 | 22.3338 | 22.2215 | 50.2016 | 0.006735 | 0.32624 | 0.002197 | 48.4395 | 0.049039 | 0.88676 | 0.043486 | 18.0827 | 22.049 | 3.4609 | 0.045683 | 0.3609 | 0.381 | 0.2848 | 0.1725
2012-10-30 22:42 | 26.3951 | 22.4328 | 22.4167 | 22.2644 | 22.2283 | 50.2062 | 0.006716 | 0.32621 | 0.002191 | 48.5722 | 0.048774 | 0.88231 | 0.043034 | 18.0897 | 22.0799 | 3.4612 | 0.045225 | 0.3529 | 0.3368 | 0.1845 | 0.1484
2012-10-30 22:42 | 26.3958 | 22.4214 | 22.4295 | 22.2851 | 22.213 | 50.2068 | 0.006722 | 0.32621 | 0.002193 | 48.5289 | 0.048699 | 0.88102 | 0.042905 | 18.0911 | 22.0245 | 3.4613 | 0.045098 | 0.3969 | 0.405 | 0.2606 | 0.1885
2012-10-30 22:42 | 26.3906 | 22.4323 | 22.3681 | 22.2759 | 22.2399 | 50.201 | 0.006732 | 0.32625 | 0.002196 | 48.4632 | 0.048796 | 0.88238 | 0.043057 | 18.083 | 22.0393 | 3.4609 | 0.045253 | 0.393 | 0.3288 | 0.2366 | 0.2006
2012-10-30 22:42 | 26.3913 | 22.429 | 22.4129 | 22.3328 | 22.2245 | 50.2022 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.048702 | 0.88132 | 0.042922 | 18.0963 | 22.0561 | 3.4609 | 0.045114 | 0.3729 | 0.3568 | 0.2767 | 0.1684
2012-10-30 22:42 | 26.3928 | 22.4265 | 22.4064 | 22.3222 | 22.2501 | 50.1987 | 0.006725 | 0.32624 | 0.002194 | 48.5115 | 0.048919 | 0.88478 | 0.043282 | 18.0866 | 22.0295 | 3.4607 | 0.045476 | 0.397 | 0.3769 | 0.2927 | 0.2206
2012-10-30 22:42 | 26.3861 | 22.4478 | 22.3757 | 22.3155 | 22.1953 | 50.1993 | 0.00671 | 0.32622 | 0.002189 | 48.6177 | 0.048673 | 0.88065 | 0.042864 | 18.0931 | 22.0268 | 3.4607 | 0.045053 | 0.421 | 0.3489 | 0.2887 | 0.1685
2012-10-30 22:42 | 26.3921 | 22.4538 | 22.4017 | 22.3336 | 22.2213 | 50.2022 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048909 | 0.88437 | 0.043254 | 18.082 | 22.0528 | 3.4609 | 0.045448 | 0.401 | 0.3489 | 0.2808 | 0.1685
2012-10-30 22:42 | 26.3808 | 22.4145 | 22.3704 | 22.3102 | 22.21 | 50.1976 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.048793 | 0.88293 | 0.043081 | 18.0954 | 22.0095 | 3.4606 | 0.045276 | 0.405 | 0.3609 | 0.3007 | 0.2005
2012-10-30 22:42 | 26.3906 | 22.4323 | 22.3762 | 22.3321 | 22.1998 | 50.1999 | 0.00671 | 0.32628 | 0.002189 | 48.6264 | 0.04874 | 0.88158 | 0.042968 | 18.0875 | 22.0473 | 3.4608 | 0.045157 | 0.385 | 0.3289 | 0.2848 | 0.1525
2012-10-30 22:42 | 26.3995 | 22.4773 | 22.3811 | 22.3009 | 22.2127 | 50.2004 | 0.006735 | 0.32627 | 0.002197 | 48.4437 | 0.048692 | 0.88102 | 0.042899 | 18.0937 | 22.0723 | 3.4608 | 0.045096 | 0.405 | 0.3088 | 0.2286 | 0.1404
2012-10-30 22:42 | 26.3876 | 22.4533 | 22.4052 | 22.309 | 22.1847 | 50.19 | 0.006707 | 0.32627 | 0.002188 | 48.6459 | 0.048717 | 0.88135 | 0.042937 | 18.0913 | 22.0483 | 3.4601 | 0.045125 | 0.405 | 0.3569 | 0.2607 | 0.1364
2012-10-30 22:42 | 26.3786 | 22.4443 | 22.3642 | 22.3161 | 22.1998 | 50.2028 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.048878 | 0.88392 | 0.043204 | 18.0843 | 22.0674 | 3.461 | 0.045401 | 0.3769 | 0.2968 | 0.2487 | 0.1324
2012-10-30 22:42 | 26.3801 | 22.4298 | 22.4017 | 22.3015 | 22.2133 | 50.1947 | 0.006722 | 0.32628 | 0.002193 | 48.5396 | 0.048758 | 0.88192 | 0.043 | 18.0876 | 22.0729 | 3.4604 | 0.045193 | 0.3569 | 0.3288 | 0.2286 | 0.1404
2012-10-30 22:42 | 26.3726 | 22.4303 | 22.3501 | 22.29 | 22.1617 | 50.2022 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.048979 | 0.88525 | 0.043359 | 18.0742 | 22.0494 | 3.4609 | 0.045555 | 0.3809 | 0.3007 | 0.2406 | 0.1123
2012-10-30 22:42 | 26.3868 | 22.4485 | 22.4124 | 22.2561 | 22.232 | 50.1935 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048755 | 0.88186 | 0.042995 | 18.0875 | 22.0636 | 3.4603 | 0.045189 | 0.3849 | 0.3488 | 0.1925 | 0.1684
2012-10-30 22:42 | 26.3808 | 22.4345 | 22.3664 | 22.3022 | 22.1739 | 50.1964 | 0.006713 | 0.32631 | 0.002191 | 48.609 | 0.048928 | 0.88463 | 0.043283 | 18.0803 | 22.0857 | 3.4605 | 0.045474 | 0.3488 | 0.2807 | 0.2165 | 0.0882
2012-10-30 22:42 | 26.3816 | 22.4674 | 22.3912 | 22.3351 | 22.2068 | 50.1941 | 0.00671 | 0.32631 | 0.00219 | 48.6307 | 0.04874 | 0.88177 | 0.042978 | 18.0913 | 22.0504 | 3.4604 | 0.045168 | 0.417 | 0.3408 | 0.2847 | 0.1564
2012-10-30 22:42 | 26.3823 | 22.42 | 22.3919 | 22.2797 | 22.1915 | 50.1952 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048837 | 0.88324 | 0.043135 | 18.0856 | 22.0591 | 3.4605 | 0.045327 | 0.3609 | 0.3328 | 0.2206 | 0.1324
2012-10-30 22:42 | 26.3719 | 22.4136 | 22.3815 | 22.2853 | 22.1851 | 50.1883 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048815 | 0.88303 | 0.043105 | 18.0893 | 22.0728 | 3.46 | 0.0453 | 0.3408 | 0.3087 | 0.2125 | 0.1123
2012-10-30 22:42 | 26.3778 | 22.4395 | 22.3674 | 22.3193 | 22.1749 | 50.1918 | 0.006735 | 0.3263 | 0.002198 | 48.448 | 0.048711 | 0.88138 | 0.042933 | 18.0941 | 22.0506 | 3.4602 | 0.045131 | 0.3889 | 0.3168 | 0.2687 | 0.1243
2012-10-30 22:42 | 26.3801 | 22.4739 | 22.3777 | 22.3015 | 22.1652 | 50.1976 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048985 | 0.88583 | 0.043392 | 18.0837 | 22.105 | 3.4606 | 0.045587 | 0.3689 | 0.2727 | 0.1965 | 0.0602
2012-10-30 22:42 | 26.3733 | 22.415 | 22.3749 | 22.2987 | 22.2025 | 50.1924 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048755 | 0.88192 | 0.042998 | 18.0887 | 22.0701 | 3.4603 | 0.045191 | 0.3449 | 0.3048 | 0.2286 | 0.1324
2012-10-30 22:42 | 26.3816 | 22.3992 | 22.3992 | 22.287 | 22.2068 | 50.1947 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048869 | 0.88382 | 0.043192 | 18.0856 | 22.0664 | 3.4604 | 0.045386 | 0.3328 | 0.3328 | 0.2206 | 0.1404
2012-10-30 22:42 | 26.3823 | 22.416 | 22.3879 | 22.2837 | 22.1594 | 50.1964 | 0.006747 | 0.32634 | 0.002202 | 48.3683 | 0.048765 | 0.88205 | 0.043013 | 18.0877 | 22.0671 | 3.4605 | 0.045215 | 0.3489 | 0.3208 | 0.2166 | 0.0923
2012-10-30 22:42 | 26.3831 | 22.4649 | 22.4087 | 22.2925 | 22.2083 | 50.2074 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.048887 | 0.88388 | 0.04321 | 18.0801 | 22.0358 | 3.4613 | 0.045404 | 0.4291 | 0.3729 | 0.2567 | 0.1725
2012-10-30 22:42 | 26.3801 | 22.3897 | 22.3937 | 22.2935 | 22.1973 | 50.1999 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048711 | 0.88134 | 0.042931 | 18.0932 | 22.0489 | 3.4608 | 0.045125 | 0.3408 | 0.3448 | 0.2446 | 0.1484
2012-10-30 22:43 | 26.3838 | 22.3894 | 22.3894 | 22.3413 | 22.1889 | 50.1941 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048865 | 0.88382 | 0.043188 | 18.087 | 22.0325 | 3.4604 | 0.045383 | 0.3569 | 0.3569 | 0.3088 | 0.1564
2012-10-30 22:43 | 26.3861 | 22.4238 | 22.4077 | 22.3075 | 22.1912 | 50.1889 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048774 | 0.88244 | 0.04304 | 18.0923 | 22.0388 | 3.46 | 0.045237 | 0.385 | 0.3689 | 0.2687 | 0.1524
2012-10-30 22:43 | 26.3936 | 22.4192 | 22.4112 | 22.315 | 22.1707 | 50.1871 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048909 | 0.8844 | 0.043255 | 18.0826 | 22.0624 | 3.4599 | 0.045448 | 0.3568 | 0.3488 | 0.2526 | 0.1083
2012-10-30 22:43 | 26.3876 | 22.4092 | 22.3812 | 22.2689 | 22.1968 | 50.1906 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048724 | 0.88131 | 0.042941 | 18.0878 | 22.0885 | 3.4601 | 0.045139 | 0.3207 | 0.2927 | 0.1804 | 0.1083
2012-10-30 22:43 | 26.3891 | 22.3546 | 22.3907 | 22.2865 | 22.2023 | 50.1918 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048802 | 0.88284 | 0.043084 | 18.0903 | 22.09 | 3.4602 | 0.045277 | 0.2646 | 0.3007 | 0.1965 | 0.1123
2012-10-30 22:43 | 26.3823 | 22.3959 | 22.4079 | 22.3358 | 22.1955 | 50.1993 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.048746 | 0.88187 | 0.042988 | 18.0912 | 22.0832 | 3.4607 | 0.04518 | 0.3127 | 0.3247 | 0.2526 | 0.1123
2012-10-30 22:43 | 26.3831 | 22.3887 | 22.4047 | 22.3125 | 22.1521 | 50.1883 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.04885 | 0.88336 | 0.043152 | 18.0831 | 22.0519 | 3.46 | 0.045348 | 0.3368 | 0.3528 | 0.2606 | 0.1002
2012-10-30 22:43 | 26.3861 | 22.3837 | 22.4518 | 22.2795 | 22.1993 | 50.1837 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048743 | 0.88171 | 0.042977 | 18.089 | 22.0789 | 3.4597 | 0.045171 | 0.3048 | 0.3729 | 0.2006 | 0.1204
2012-10-30 22:43 | 26.3808 | 22.3624 | 22.4024 | 22.2902 | 22.1699 | 50.1993 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.048758 | 0.8821 | 0.04301 | 18.0915 | 22.0656 | 3.4607 | 0.045204 | 0.2968 | 0.3368 | 0.2246 | 0.1043
2012-10-30 22:43 | 26.3868 | 22.4004 | 22.4405 | 22.3363 | 22.212 | 50.1929 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.048784 | 0.88216 | 0.043035 | 18.083 | 22.0756 | 3.4603 | 0.045234 | 0.3248 | 0.3649 | 0.2607 | 0.1364
2012-10-30 22:43 | 26.3868 | 22.4124 | 22.3724 | 22.3002 | 22.1719 | 50.2016 | 0.006735 | 0.3263 | 0.002198 | 48.448 | 0.048837 | 0.88316 | 0.043131 | 18.0838 | 22.0917 | 3.4609 | 0.045329 | 0.3207 | 0.2807 | 0.2085 | 0.0802
2012-10-30 22:43 | 26.3868 | 22.3724 | 22.4285 | 22.3082 | 22.212 | 50.2033 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048818 | 0.88303 | 0.043108 | 18.0882 | 22.1077 | 3.461 | 0.045303 | 0.2647 | 0.3208 | 0.2005 | 0.1043
2012-10-30 22:43 | 26.3771 | 22.3747 | 22.4268 | 22.3186 | 22.1542 | 50.1877 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.04874 | 0.88169 | 0.042973 | 18.0896 | 22.094 | 3.4599 | 0.045172 | 0.2807 | 0.3328 | 0.2246 | 0.0602
2012-10-30 22:43 | 26.3906 | 22.3441 | 22.4523 | 22.32 | 22.2038 | 50.1924 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048856 | 0.88371 | 0.043174 | 18.088 | 22.1035 | 3.4603 | 0.04537 | 0.2406 | 0.3488 | 0.2165 | 0.1003
2012-10-30 22:43 | 26.3823 | 22.3959 | 22.452 | 22.3077 | 22.1714 | 50.1958 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048828 | 0.88307 | 0.043119 | 18.0854 | 22.1032 | 3.4605 | 0.045316 | 0.2927 | 0.3488 | 0.2045 | 0.0682
2012-10-30 22:43 | 26.3876 | 22.4052 | 22.4333 | 22.3251 | 22.1767 | 50.1993 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048862 | 0.88381 | 0.043185 | 18.0879 | 22.1085 | 3.4607 | 0.045382 | 0.2967 | 0.3248 | 0.2166 | 0.0682
2012-10-30 22:43 | 26.3831 | 22.4488 | 22.4047 | 22.3286 | 22.2283 | 50.1791 | 0.006735 | 0.32627 | 0.002197 | 48.4437 | 0.048727 | 0.88132 | 0.042944 | 18.087 | 22.0519 | 3.4593 | 0.045141 | 0.3969 | 0.3528 | 0.2767 | 0.1764
2012-10-30 22:43 | 26.3756 | 22.4093 | 22.4293 | 22.3211 | 22.1807 | 50.1999 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048802 | 0.88268 | 0.043077 | 18.087 | 22.0644 | 3.4608 | 0.045271 | 0.3449 | 0.3649 | 0.2567 | 0.1163
2012-10-30 22:43 | 26.3666 | 22.3522 | 22.4604 | 22.3281 | 22.1798 | 50.1958 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048812 | 0.88299 | 0.0431 | 18.0895 | 22.0434 | 3.4605 | 0.045297 | 0.3088 | 0.417 | 0.2847 | 0.1364
2012-10-30 22:43 | 26.3621 | 22.3998 | 22.4198 | 22.3036 | 22.1953 | 50.1877 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048736 | 0.8815 | 0.042961 | 18.0872 | 22.0389 | 3.4599 | 0.045156 | 0.3609 | 0.3809 | 0.2647 | 0.1564
2012-10-30 22:43 | 26.3659 | 22.3835 | 22.4356 | 22.3034 | 22.1951 | 50.1912 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048894 | 0.88418 | 0.043231 | 18.0837 | 22.0828 | 3.4602 | 0.045428 | 0.3007 | 0.3528 | 0.2206 | 0.1123
2012-10-30 22:43 | 26.3771 | 22.4268 | 22.3987 | 22.2785 | 22.2504 | 50.1877 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048859 | 0.88359 | 0.043171 | 18.0845 | 22.086 | 3.4599 | 0.045365 | 0.3408 | 0.3127 | 0.1925 | 0.1644
2012-10-30 22:43 | 26.3726 | 22.4423 | 22.4423 | 22.3381 | 22.1858 | 50.1918 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048865 | 0.88377 | 0.043185 | 18.0858 | 22.1055 | 3.4602 | 0.045378 | 0.3368 | 0.3368 | 0.2326 | 0.0803
2012-10-30 22:43 | 26.3748 | 22.4445 | 22.4405 | 22.3083 | 22.192 | 50.2016 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048847 | 0.8833 | 0.043147 | 18.0831 | 22.0516 | 3.4609 | 0.045345 | 0.3929 | 0.3889 | 0.2567 | 0.1404
2012-10-30 22:43 | 26.3748 | 22.4004 | 22.4045 | 22.3042 | 22.208 | 50.1993 | 0.00671 | 0.32628 | 0.002189 | 48.6264 | 0.048837 | 0.88322 | 0.043134 | 18.085 | 22.0877 | 3.4607 | 0.045323 | 0.3127 | 0.3168 | 0.2165 | 0.1203
2012-10-30 22:43 | 26.3689 | 22.3705 | 22.4066 | 22.2783 | 22.2101 | 50.2016 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048887 | 0.88398 | 0.043215 | 18.0821 | 22.0818 | 3.4609 | 0.045414 | 0.2887 | 0.3248 | 0.1965 | 0.1283
2012-10-30 22:43 | 26.3659 | 22.3595 | 22.4276 | 22.3034 | 22.2031 | 50.1762 | 0.00671 | 0.32631 | 0.00219 | 48.6307 | 0.048695 | 0.88112 | 0.042906 | 18.0947 | 22.0387 | 3.4591 | 0.045096 | 0.3208 | 0.3889 | 0.2647 | 0.1644
2012-10-30 22:43 | 26.3651 | 22.4268 | 22.3667 | 22.2905 | 22.1903 | 50.1999 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.048809 | 0.88265 | 0.043081 | 18.0838 | 22.0499 | 3.4608 | 0.045279 | 0.3769 | 0.3168 | 0.2406 | 0.1404
2012-10-30 22:44 | 26.3674 | 22.381 | 22.4171 | 22.2888 | 22.1685 | 50.1952 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048935 | 0.88476 | 0.043296 | 18.0804 | 22.0081 | 3.4605 | 0.045492 | 0.3729 | 0.409 | 0.2807 | 0.1604
2012-10-30 22:44 | 26.3666 | 22.3602 | 22.4003 | 22.3201 | 22.1677 | 50.1877 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.048686 | 0.88089 | 0.042887 | 18.0933 | 22.0474 | 3.4599 | 0.045081 | 0.3128 | 0.3529 | 0.2727 | 0.1203
2012-10-30 22:44 | 26.3651 | 22.3948 | 22.4549 | 22.2986 | 22.2063 | 50.1941 | 0.00671 | 0.32636 | 0.00219 | 48.6371 | 0.048717 | 0.88167 | 0.042952 | 18.0978 | 22.0459 | 3.4604 | 0.045142 | 0.3489 | 0.409 | 0.2527 | 0.1604
2012-10-30 22:44 | 26.3666 | 22.4083 | 22.3842 | 22.3081 | 22.2159 | 50.1837 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.048689 | 0.88096 | 0.042893 | 18.0937 | 22.0675 | 3.4597 | 0.04509 | 0.3408 | 0.3167 | 0.2406 | 0.1484
2012-10-30 22:44 | 26.3711 | 22.4008 | 22.4168 | 22.3326 | 22.1963 | 50.1779 | 0.006719 | 0.32628 | 0.002192 | 48.5612 | 0.048818 | 0.88287 | 0.0431 | 18.0849 | 22.0158 | 3.4593 | 0.045292 | 0.385 | 0.401 | 0.3168 | 0.1805
2012-10-30 22:44 | 26.3636 | 22.3973 | 22.3652 | 22.3251 | 22.1768 | 50.186 | 0.006719 | 0.32628 | 0.002192 | 48.5612 | 0.048746 | 0.88187 | 0.042988 | 18.0912 | 22.0404 | 3.4598 | 0.04518 | 0.3569 | 0.3248 | 0.2847 | 0.1364
2012-10-30 22:44 | 26.3733 | 22.3789 | 22.4591 | 22.3228 | 22.2105 | 50.1935 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.048733 | 0.8817 | 0.042968 | 18.0924 | 22.0381 | 3.4603 | 0.045166 | 0.3408 | 0.421 | 0.2847 | 0.1724
2012-10-30 22:44 | 26.3659 | 22.4116 | 22.4116 | 22.3154 | 22.2232 | 50.1947 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048777 | 0.88249 | 0.043045 | 18.0924 | 22.0547 | 3.4604 | 0.045242 | 0.3569 | 0.3569 | 0.2607 | 0.1685
2012-10-30 22:44 | 26.3681 | 22.4178 | 22.4218 | 22.3136 | 22.2013 | 50.1958 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048909 | 0.88452 | 0.043261 | 18.085 | 22.073 | 3.4605 | 0.045457 | 0.3448 | 0.3488 | 0.2406 | 0.1283
2012-10-30 22:44 | 26.3644 | 22.4021 | 22.386 | 22.3099 | 22.2056 | 50.1929 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048837 | 0.88335 | 0.04314 | 18.0876 | 22.0492 | 3.4603 | 0.045337 | 0.3529 | 0.3368 | 0.2607 | 0.1564
2012-10-30 22:44 | 26.3704 | 22.376 | 22.4201 | 22.3079 | 22.2076 | 50.186 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048762 | 0.88226 | 0.043021 | 18.0932 | 22.0913 | 3.4598 | 0.045217 | 0.2847 | 0.3288 | 0.2166 | 0.1163
2012-10-30 22:44 | 26.3704 | 22.4041 | 22.4281 | 22.2998 | 22.1675 | 50.1843 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048774 | 0.88229 | 0.043033 | 18.0894 | 22.0672 | 3.4597 | 0.045226 | 0.3369 | 0.3609 | 0.2326 | 0.1003
2012-10-30 22:44 | 26.3771 | 22.3907 | 22.4508 | 22.2945 | 22.2103 | 50.197 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048881 | 0.88395 | 0.043209 | 18.0838 | 22.086 | 3.4606 | 0.045402 | 0.3047 | 0.3648 | 0.2085 | 0.1243
2012-10-30 22:44 | 26.3651 | 22.3306 | 22.4188 | 22.2785 | 22.1743 | 50.1831 | 0.006744 | 0.32631 | 0.002201 | 48.3855 | 0.048771 | 0.88218 | 0.043025 | 18.0881 | 22.094 | 3.4596 | 0.045226 | 0.2366 | 0.3248 | 0.1845 | 0.0803
2012-10-30 22:44 | 26.3629 | 22.3725 | 22.3966 | 22.2923 | 22.152 | 50.186 | 0.006716 | 0.32628 | 0.002191 | 48.5829 | 0.04874 | 0.88155 | 0.042967 | 18.0869 | 22.0999 | 3.4598 | 0.045158 | 0.2726 | 0.2967 | 0.1924 | 0.0521
2012-10-30 22:44 | 26.3711 | 22.3847 | 22.4288 | 22.2805 | 22.2324 | 50.1889 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048721 | 0.88143 | 0.042944 | 18.0913 | 22.08 | 3.46 | 0.045139 | 0.3047 | 0.3488 | 0.2005 | 0.1524
2012-10-30 22:44 | 26.3644 | 22.3339 | 22.4341 | 22.2698 | 22.1856 | 50.1854 | 0.006738 | 0.3264 | 0.002199 | 48.4415 | 0.048743 | 0.88176 | 0.042979 | 18.0899 | 22.0492 | 3.4598 | 0.045178 | 0.2847 | 0.3849 | 0.2206 | 0.1364
2012-10-30 22:44 | 26.3681 | 22.3817 | 22.4138 | 22.2815 | 22.1532 | 50.1837 | 0.006747 | 0.32636 | 0.002202 | 48.3704 | 0.048799 | 0.88283 | 0.043081 | 18.0911 | 22.085 | 3.4597 | 0.045283 | 0.2967 | 0.3288 | 0.1965 | 0.0682
2012-10-30 22:44 | 26.3591 | 22.3326 | 22.4168 | 22.2885 | 22.1723 | 50.1814 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.048777 | 0.88232 | 0.043037 | 18.0889 | 22.08 | 3.4595 | 0.045234 | 0.2526 | 0.3368 | 0.2085 | 0.0923
2012-10-30 22:44 | 26.3584 | 22.376 | 22.4081 | 22.2838 | 22.1756 | 50.1785 | 0.006713 | 0.32643 | 0.002191 | 48.6261 | 0.048932 | 0.88469 | 0.04329 | 18.08 | 22.0552 | 3.4593 | 0.045481 | 0.3208 | 0.3529 | 0.2286 | 0.1204
2012-10-30 22:44 | 26.3659 | 22.3715 | 22.4116 | 22.3034 | 22.1991 | 50.1808 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.048708 | 0.88119 | 0.042921 | 18.0914 | 22.1029 | 3.4595 | 0.045118 | 0.2686 | 0.3087 | 0.2005 | 0.0962
2012-10-30 22:44 | 26.3621 | 22.3437 | 22.4358 | 22.2996 | 22.1793 | 50.1767 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.048806 | 0.88261 | 0.043077 | 18.084 | 22.1111 | 3.4592 | 0.045271 | 0.2326 | 0.3247 | 0.1885 | 0.0682
2012-10-30 22:44 | 26.3711 | 22.4008 | 22.4168 | 22.3166 | 22.2003 | 50.186 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048708 | 0.88145 | 0.042934 | 18.0967 | 22.096 | 3.4598 | 0.04513 | 0.3048 | 0.3208 | 0.2206 | 0.1043
2012-10-30 22:44 | 26.3644 | 22.354 | 22.4221 | 22.3099 | 22.1575 | 50.1964 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048724 | 0.88132 | 0.042942 | 18.0881 | 22.0933 | 3.4605 | 0.045137 | 0.2607 | 0.3288 | 0.2166 | 0.0642
2012-10-30 22:44 | 26.3599 | 22.3775 | 22.4256 | 22.3174 | 22.1931 | 50.1791 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048762 | 0.88197 | 0.043007 | 18.0873 | 22.0688 | 3.4593 | 0.045203 | 0.3087 | 0.3568 | 0.2486 | 0.1243
2012-10-30 22:44 | 26.3599 | 22.4056 | 22.3615 | 22.2573 | 22.2011 | 50.1733 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.048831 | 0.88293 | 0.043114 | 18.0813 | 22.0608 | 3.4589 | 0.045308 | 0.3448 | 0.3007 | 0.1965 | 0.1403
2012-10-30 22:44 | 26.3644 | 22.35 | 22.4582 | 22.3099 | 22.1776 | 50.1837 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048793 | 0.88272 | 0.043071 | 18.0912 | 22.0893 | 3.4597 | 0.045268 | 0.2607 | 0.3689 | 0.2206 | 0.0883
2012-10-30 22:45 | 26.3644 | 22.37 | 22.4622 | 22.3259 | 22.2056 | 50.1796 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048758 | 0.88209 | 0.043009 | 18.0912 | 22.0973 | 3.4594 | 0.045202 | 0.2727 | 0.3649 | 0.2286 | 0.1083
2012-10-30 22:45 | 26.3606 | 22.3943 | 22.4223 | 22.262 | 22.2018 | 50.1796 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048733 | 0.88205 | 0.042985 | 18.0996 | 22.0895 | 3.4594 | 0.045178 | 0.3048 | 0.3328 | 0.1725 | 0.1123
2012-10-30 22:45 | 26.3636 | 22.3933 | 22.4494 | 22.273 | 22.2088 | 50.1825 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048768 | 0.88245 | 0.043035 | 18.0949 | 22.0765 | 3.4596 | 0.04523 | 0.3168 | 0.3729 | 0.1965 | 0.1323
2012-10-30 22:45 | 26.3584 | 22.3921 | 22.4362 | 22.2558 | 22.2117 | 50.1895 | 0.006744 | 0.32631 | 0.002201 | 48.3855 | 0.048749 | 0.88186 | 0.04299 | 18.0898 | 22.1034 | 3.4601 | 0.045191 | 0.2887 | 0.3328 | 0.1524 | 0.1083
2012-10-30 22:45 | 26.3681 | 22.4178 | 22.4499 | 22.3336 | 22.2454 | 50.1993 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048765 | 0.88222 | 0.043021 | 18.0912 | 22.0529 | 3.4607 | 0.045219 | 0.3649 | 0.397 | 0.2807 | 0.1925
2012-10-30 22:45 | 26.3584 | 22.364 | 22.4041 | 22.3319 | 22.1836 | 50.1895 | 0.00671 | 0.32636 | 0.00219 | 48.6371 | 0.048837 | 0.8831 | 0.043128 | 18.0826 | 22.0833 | 3.4601 | 0.045318 | 0.2807 | 0.3208 | 0.2486 | 0.1003
2012-10-30 22:45 | 26.3591 | 22.3888 | 22.3968 | 22.3246 | 22.2164 | 50.1866 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048815 | 0.88281 | 0.043094 | 18.0849 | 22.068 | 3.4599 | 0.045286 | 0.3208 | 0.3288 | 0.2566 | 0.1484
2012-10-30 22:45 | 26.3606 | 22.3862 | 22.3983 | 22.3021 | 22.2058 | 50.1848 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.048903 | 0.88436 | 0.043248 | 18.0839 | 22.0695 | 3.4597 | 0.04544 | 0.3167 | 0.3288 | 0.2326 | 0.1363
2012-10-30 22:45 | 26.3584 | 22.3881 | 22.4081 | 22.2678 | 22.1836 | 50.175 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048982 | 0.88556 | 0.043376 | 18.0792 | 22.0392 | 3.4591 | 0.045571 | 0.3489 | 0.3689 | 0.2286 | 0.1444
2012-10-30 22:45 | 26.3591 | 22.3928 | 22.4409 | 22.3206 | 22.1843 | 50.1906 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048777 | 0.88225 | 0.043033 | 18.0874 | 22.0559 | 3.4601 | 0.04523 | 0.3369 | 0.385 | 0.2647 | 0.1284
2012-10-30 22:45 | 26.3576 | 22.3873 | 22.3752 | 22.3071 | 22.1908 | 50.175 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048799 | 0.88261 | 0.04307 | 18.0866 | 22.0464 | 3.4591 | 0.045264 | 0.3409 | 0.3288 | 0.2607 | 0.1444
2012-10-30 22:45 | 26.3591 | 22.3647 | 22.4088 | 22.3366 | 22.1963 | 50.1895 | 0.006713 | 0.32637 | 0.002191 | 48.6176 | 0.048979 | 0.88554 | 0.043373 | 18.0801 | 22.06 | 3.4601 | 0.045564 | 0.3047 | 0.3488 | 0.2766 | 0.1363
2012-10-30 22:45 | 26.3599 | 22.4056 | 22.4336 | 22.2974 | 22.1931 | 50.1773 | 0.006735 | 0.32638 | 0.002198 | 48.4609 | 0.048809 | 0.88265 | 0.043081 | 18.0838 | 22.0648 | 3.4592 | 0.045279 | 0.3408 | 0.3688 | 0.2326 | 0.1283
2012-10-30 22:45 | 26.3599 | 22.3775 | 22.3815 | 22.3294 | 22.1891 | 50.1843 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048774 | 0.88206 | 0.043022 | 18.0847 | 22.0527 | 3.4597 | 0.045219 | 0.3248 | 0.3288 | 0.2767 | 0.1364
2012-10-30 22:45 | 26.3644 | 22.39 | 22.3981 | 22.2698 | 22.2016 | 50.1704 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048812 | 0.88272 | 0.043088 | 18.0842 | 22.0452 | 3.4587 | 0.045283 | 0.3448 | 0.3529 | 0.2246 | 0.1564
2012-10-30 22:45 | 26.3629 | 22.3966 | 22.4326 | 22.2803 | 22.1761 | 50.186 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048768 | 0.88206 | 0.043016 | 18.0869 | 22.0678 | 3.4598 | 0.045213 | 0.3288 | 0.3648 | 0.2125 | 0.1083
2012-10-30 22:45 | 26.3674 | 22.377 | 22.389 | 22.2848 | 22.2167 | 50.1877 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.04879 | 0.88238 | 0.043051 | 18.0852 | 22.0923 | 3.4599 | 0.045248 | 0.2847 | 0.2967 | 0.1925 | 0.1244
2012-10-30 22:45 | 26.3599 | 22.3775 | 22.3936 | 22.3094 | 22.157 | 50.1877 | 0.006725 | 0.32643 | 0.002195 | 48.5393 | 0.048717 | 0.88132 | 0.042935 | 18.0907 | 22.0728 | 3.4599 | 0.04513 | 0.3047 | 0.3208 | 0.2366 | 0.0842
2012-10-30 22:45 | 26.3606 | 22.3782 | 22.4303 | 22.2981 | 22.1938 | 50.1843 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048724 | 0.8813 | 0.04294 | 18.0875 | 22.0334 | 3.4597 | 0.045137 | 0.3448 | 0.3969 | 0.2647 | 0.1604
2012-10-30 22:45 | 26.3644 | 22.354 | 22.4181 | 22.3099 | 22.1695 | 50.171 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.048708 | 0.88111 | 0.042917 | 18.0896 | 22.0532 | 3.4588 | 0.045116 | 0.3008 | 0.3649 | 0.2567 | 0.1163
2012-10-30 22:45 | 26.3689 | 22.3464 | 22.3986 | 22.2903 | 22.1821 | 50.175 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048692 | 0.88085 | 0.04289 | 18.0902 | 22.0698 | 3.4591 | 0.045083 | 0.2766 | 0.3288 | 0.2205 | 0.1123
2012-10-30 22:45 | 26.3696 | 22.4113 | 22.4273 | 22.299 | 22.2028 | 50.1739 | 0.006722 | 0.32641 | 0.002194 | 48.5589 | 0.048702 | 0.88115 | 0.042914 | 18.0927 | 22.0825 | 3.459 | 0.045108 | 0.3288 | 0.3448 | 0.2165 | 0.1203
2012-10-30 22:45 | 26.3681 | 22.3857 | 22.4018 | 22.3056 | 22.2093 | 50.1796 | 0.006713 | 0.32633 | 0.002191 | 48.6111 | 0.04867 | 0.88052 | 0.042855 | 18.0915 | 22.089 | 3.4594 | 0.045046 | 0.2967 | 0.3128 | 0.2166 | 0.1203
2012-10-30 22:45 | 26.3651 | 22.3547 | 22.4549 | 22.3106 | 22.1943 | 50.1727 | 0.006707 | 0.32634 | 0.002189 | 48.6568 | 0.048695 | 0.88096 | 0.042898 | 18.0914 | 22.0619 | 3.4589 | 0.045087 | 0.2928 | 0.393 | 0.2487 | 0.1324
2012-10-30 22:45 | 26.3666 | 22.3642 | 22.4323 | 22.3361 | 22.1878 | 50.1756 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048689 | 0.8807 | 0.04288 | 18.0883 | 22.1196 | 3.4591 | 0.045076 | 0.2446 | 0.3127 | 0.2165 | 0.0682
2012-10-30 22:45 | 26.3719 | 22.3575 | 22.4096 | 22.3094 | 22.173 | 50.1785 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.048692 | 0.88102 | 0.042899 | 18.0937 | 22.0848 | 3.4593 | 0.045092 | 0.2727 | 0.3248 | 0.2246 | 0.0882
2012-10-30 22:45 | 26.3666 | 22.4083 | 22.4564 | 22.272 | 22.2038 | 50.1727 | 0.006722 | 0.32646 | 0.002194 | 48.5653 | 0.048733 | 0.8815 | 0.042958 | 18.0883 | 22.1116 | 3.4589 | 0.045152 | 0.2967 | 0.3448 | 0.1604 | 0.0922
2012-10-30 22:46 | 26.3733 | 22.3709 | 22.435 | 22.2667 | 22.1825 | 50.1762 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048802 | 0.88268 | 0.043077 | 18.087 | 22.0742 | 3.4591 | 0.045273 | 0.2967 | 0.3608 | 0.1925 | 0.1083
2012-10-30 22:46 | 26.3681 | 22.3456 | 22.4178 | 22.2975 | 22.1933 | 50.1687 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048733 | 0.88145 | 0.042956 | 18.0874 | 22.089 | 3.4586 | 0.045153 | 0.2566 | 0.3288 | 0.2085 | 0.1043
2012-10-30 22:46 | 26.3606 | 22.3782 | 22.4063 | 22.266 | 22.1337 | 50.1744 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.04867 | 0.88057 | 0.042857 | 18.0927 | 22.0815 | 3.459 | 0.04505 | 0.2967 | 0.3248 | 0.1845 | 0.0522
2012-10-30 22:46 | 26.3621 | 22.3877 | 22.4118 | 22.2875 | 22.1753 | 50.1721 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.04878 | 0.88238 | 0.043042 | 18.0889 | 22.083 | 3.4589 | 0.04524 | 0.3047 | 0.3288 | 0.2045 | 0.0923
2012-10-30 22:46 | 26.3644 | 22.362 | 22.4261 | 22.3019 | 22.2137 | 50.1721 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.048733 | 0.88143 | 0.042954 | 18.0868 | 22.0452 | 3.4589 | 0.045146 | 0.3168 | 0.3809 | 0.2567 | 0.1685
2012-10-30 22:46 | 26.3606 | 22.3462 | 22.4143 | 22.258 | 22.1858 | 50.1756 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.04867 | 0.88063 | 0.04286 | 18.0939 | 22.0534 | 3.4591 | 0.045056 | 0.2928 | 0.3609 | 0.2046 | 0.1324
2012-10-30 22:46 | 26.3599 | 22.3695 | 22.3855 | 22.2452 | 22.1891 | 50.1704 | 0.006722 | 0.32643 | 0.002194 | 48.561 | 0.048667 | 0.88049 | 0.042851 | 18.0921 | 22.0367 | 3.4587 | 0.045045 | 0.3328 | 0.3488 | 0.2085 | 0.1524
2012-10-30 22:46 | 26.3621 | 22.3757 | 22.3797 | 22.2755 | 22.1993 | 50.1762 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048677 | 0.88066 | 0.042868 | 18.0919 | 22.087 | 3.4591 | 0.04506 | 0.2887 | 0.2927 | 0.1885 | 0.1123
2012-10-30 22:46 | 26.3614 | 22.343 | 22.4271 | 22.2828 | 22.1746 | 50.1675 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048758 | 0.88184 | 0.042997 | 18.0861 | 22.0582 | 3.4585 | 0.04519 | 0.2848 | 0.3689 | 0.2246 | 0.1164
2012-10-30 22:46 | 26.3741 | 22.4318 | 22.4198 | 22.3156 | 22.2073 | 50.1785 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048762 | 0.88215 | 0.043015 | 18.0909 | 22.0509 | 3.4593 | 0.045208 | 0.3809 | 0.3689 | 0.2647 | 0.1564
2012-10-30 22:46 | 26.3569 | 22.3745 | 22.3745 | 22.2182 | 22.1861 | 50.1704 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.04868 | 0.88043 | 0.042859 | 18.086 | 22.0858 | 3.4587 | 0.045055 | 0.2887 | 0.2887 | 0.1324 | 0.1003
2012-10-30 22:46 | 26.3696 | 22.3351 | 22.3792 | 22.271 | 22.1707 | 50.1704 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.048869 | 0.88355 | 0.043178 | 18.0799 | 22.0945 | 3.4587 | 0.04537 | 0.2406 | 0.2847 | 0.1765 | 0.0762
2012-10-30 22:46 | 26.3704 | 22.372 | 22.388 | 22.3079 | 22.2036 | 50.1646 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048771 | 0.88221 | 0.043026 | 18.0887 | 22.0753 | 3.4583 | 0.045221 | 0.2967 | 0.3127 | 0.2326 | 0.1283
2012-10-30 22:46 | 26.3636 | 22.3331 | 22.4414 | 22.257 | 22.1687 | 50.1646 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048799 | 0.88235 | 0.043058 | 18.0813 | 22.0805 | 3.4583 | 0.045253 | 0.2526 | 0.3609 | 0.1765 | 0.0882
2012-10-30 22:46 | 26.3674 | 22.381 | 22.4091 | 22.3209 | 22.1926 | 50.1721 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048802 | 0.88262 | 0.043074 | 18.0858 | 22.0883 | 3.4589 | 0.045271 | 0.2927 | 0.3208 | 0.2326 | 0.1043
2012-10-30 22:46 | 26.3704 | 22.356 | 22.384 | 22.2878 | 22.1796 | 50.1773 | 0.00671 | 0.32634 | 0.00219 | 48.635 | 0.048736 | 0.8817 | 0.04297 | 18.0913 | 22.0833 | 3.4592 | 0.04516 | 0.2727 | 0.3007 | 0.2045 | 0.0963
2012-10-30 22:46 | 26.3719 | 22.3414 | 22.4176 | 22.2773 | 22.169 | 50.1727 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048705 | 0.88118 | 0.042918 | 18.0922 | 22.0687 | 3.4589 | 0.045113 | 0.2727 | 0.3489 | 0.2086 | 0.1003
2012-10-30 22:46 | 26.3771 | 22.4308 | 22.3827 | 22.3065 | 22.1943 | 50.1744 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.048762 | 0.88197 | 0.043007 | 18.0873 | 22.082 | 3.459 | 0.0452 | 0.3488 | 0.3007 | 0.2245 | 0.1123
2012-10-30 22:46 | 26.3629 | 22.3445 | 22.4246 | 22.2763 | 22.16 | 50.175 | 0.006738 | 0.32638 | 0.002199 | 48.4393 | 0.048667 | 0.88073 | 0.042863 | 18.0971 | 22.0678 | 3.4591 | 0.045062 | 0.2767 | 0.3568 | 0.2085 | 0.0922
2012-10-30 22:46 | 26.3696 | 22.3632 | 22.4153 | 22.295 | 22.1908 | 50.1617 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.048843 | 0.8833 | 0.043143 | 18.0845 | 22.0905 | 3.4581 | 0.045334 | 0.2727 | 0.3248 | 0.2045 | 0.1003
2012-10-30 22:46 | 26.3726 | 22.3421 | 22.3942 | 22.298 | 22.1697 | 50.1727 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048755 | 0.88212 | 0.043008 | 18.0929 | 22.0614 | 3.4589 | 0.045202 | 0.2807 | 0.3328 | 0.2366 | 0.1083
2012-10-30 22:46 | 26.3674 | 22.3329 | 22.4131 | 22.2728 | 22.1846 | 50.1629 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048724 | 0.88122 | 0.042937 | 18.086 | 22.0602 | 3.4582 | 0.045133 | 0.2727 | 0.3529 | 0.2126 | 0.1244
2012-10-30 22:46 | 26.3778 | 22.3553 | 22.3914 | 22.2752 | 22.207 | 50.1733 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048683 | 0.88082 | 0.042881 | 18.0929 | 22.0827 | 3.4589 | 0.045076 | 0.2726 | 0.3087 | 0.1925 | 0.1243
2012-10-30 22:46 | 26.3689 | 22.3545 | 22.3945 | 22.3184 | 22.1981 | 50.1715 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048708 | 0.88132 | 0.042928 | 18.094 | 22.0617 | 3.4588 | 0.045127 | 0.2928 | 0.3328 | 0.2567 | 0.1364
2012-10-30 22:46 | 26.3696 | 22.3552 | 22.3872 | 22.3111 | 22.1747 | 50.16 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048872 | 0.88388 | 0.043197 | 18.0856 | 22.0263 | 3.458 | 0.045396 | 0.3289 | 0.3609 | 0.2848 | 0.1484
2012-10-30 22:46 | 26.3689 | 22.3464 | 22.4106 | 22.3064 | 22.142 | 50.1715 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048802 | 0.88281 | 0.043083 | 18.0897 | 22.0698 | 3.4588 | 0.045275 | 0.2766 | 0.3408 | 0.2366 | 0.0722
2012-10-30 22:46 | 26.3704 | 22.352 | 22.4441 | 22.3239 | 22.1635 | 50.1692 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048699 | 0.88111 | 0.042909 | 18.0929 | 22.0552 | 3.4587 | 0.045105 | 0.2968 | 0.3889 | 0.2687 | 0.1083
2012-10-30 22:47 | 26.3674 | 22.357 | 22.393 | 22.2848 | 22.1725 | 50.1675 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048755 | 0.882 | 0.043002 | 18.0905 | 22.0402 | 3.4585 | 0.045198 | 0.3168 | 0.3528 | 0.2446 | 0.1323
2012-10-30 22:47 | 26.3599 | 22.3535 | 22.3575 | 22.2974 | 22.157 | 50.1791 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048752 | 0.88182 | 0.04299 | 18.0878 | 22.0407 | 3.4593 | 0.045187 | 0.3128 | 0.3168 | 0.2567 | 0.1163
2012-10-30 22:47 | 26.3569 | 22.3825 | 22.3745 | 22.2863 | 22.138 | 50.1727 | 0.00671 | 0.32636 | 0.00219 | 48.6371 | 0.048705 | 0.88122 | 0.04292 | 18.0931 | 22.0096 | 3.4589 | 0.04511 | 0.3729 | 0.3649 | 0.2767 | 0.1284
2012-10-30 22:47 | 26.3599 | 22.3855 | 22.3374 | 22.2693 | 22.2011 | 50.1681 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048727 | 0.88154 | 0.042955 | 18.0914 | 22.0447 | 3.4586 | 0.045147 | 0.3408 | 0.2927 | 0.2246 | 0.1564
2012-10-30 22:47 | 26.3689 | 22.3945 | 22.3464 | 22.2823 | 22.1981 | 50.1681 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048793 | 0.88261 | 0.043065 | 18.0889 | 22.0858 | 3.4586 | 0.045262 | 0.3087 | 0.2606 | 0.1965 | 0.1123
2012-10-30 22:47 | 26.3561 | 22.3417 | 22.3817 | 22.2575 | 22.1492 | 50.1606 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.048664 | 0.88049 | 0.042848 | 18.0932 | 22.0209 | 3.4581 | 0.045045 | 0.3208 | 0.3608 | 0.2366 | 0.1283
2012-10-30 22:47 | 26.3629 | 22.4006 | 22.4206 | 22.2803 | 22.1721 | 50.1692 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048765 | 0.88203 | 0.043012 | 18.0874 | 22.0357 | 3.4587 | 0.045206 | 0.3649 | 0.3849 | 0.2446 | 0.1364
2012-10-30 22:47 | 26.3644 | 22.354 | 22.386 | 22.3139 | 22.1736 | 50.1681 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048642 | 0.88047 | 0.042828 | 18.101 | 22.0452 | 3.4586 | 0.045023 | 0.3088 | 0.3408 | 0.2687 | 0.1284
2012-10-30 22:47 | 26.3629 | 22.3525 | 22.3805 | 22.2763 | 22.1761 | 50.1739 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.048762 | 0.88213 | 0.043015 | 18.0906 | 22.0317 | 3.459 | 0.045211 | 0.3208 | 0.3488 | 0.2446 | 0.1444
2012-10-30 22:47 | 26.3621 | 22.3877 | 22.3958 | 22.3196 | 22.1472 | 50.1663 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.04884 | 0.88336 | 0.043143 | 18.0868 | 22.067 | 3.4585 | 0.045342 | 0.3207 | 0.3288 | 0.2526 | 0.0802
2012-10-30 22:47 | 26.3524 | 22.3179 | 22.366 | 22.2979 | 22.1535 | 50.1583 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048673 | 0.88067 | 0.042865 | 18.0937 | 22.0452 | 3.4579 | 0.045061 | 0.2727 | 0.3208 | 0.2527 | 0.1083
2012-10-30 22:47 | 26.3629 | 22.3885 | 22.4206 | 22.3044 | 22.1881 | 50.1681 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.048743 | 0.88203 | 0.042993 | 18.0956 | 22.0838 | 3.4586 | 0.04519 | 0.3047 | 0.3368 | 0.2206 | 0.1043
2012-10-30 22:47 | 26.3599 | 22.3495 | 22.3775 | 22.2813 | 22.145 | 50.1687 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048765 | 0.88215 | 0.043018 | 18.0898 | 22.0327 | 3.4586 | 0.045217 | 0.3168 | 0.3448 | 0.2486 | 0.1123
2012-10-30 22:47 | 26.3659 | 22.3515 | 22.4116 | 22.3114 | 22.163 | 50.1611 | 0.006725 | 0.3264 | 0.002195 | 48.5352 | 0.048806 | 0.88285 | 0.043089 | 18.0891 | 22.0547 | 3.4581 | 0.045284 | 0.2968 | 0.3569 | 0.2567 | 0.1083
2012-10-30 22:47 | 26.3644 | 22.346 | 22.39 | 22.2818 | 22.1294 | 50.164 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.048708 | 0.8815 | 0.042936 | 18.0976 | 22.0572 | 3.4583 | 0.045128 | 0.2888 | 0.3328 | 0.2246 | 0.0722
2012-10-30 22:47 | 26.3599 | 22.3575 | 22.3615 | 22.2813 | 22.1731 | 50.1617 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048673 | 0.88092 | 0.042877 | 18.0987 | 22.0327 | 3.4581 | 0.045071 | 0.3248 | 0.3288 | 0.2486 | 0.1404
2012-10-30 22:47 | 26.3651 | 22.3667 | 22.3787 | 22.3386 | 22.1422 | 50.175 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.04874 | 0.88176 | 0.042977 | 18.091 | 22.074 | 3.4591 | 0.045173 | 0.2927 | 0.3047 | 0.2646 | 0.0682
2012-10-30 22:47 | 26.3711 | 22.3406 | 22.4088 | 22.3246 | 22.1562 | 50.1721 | 0.006741 | 0.32638 | 0.0022 | 48.4177 | 0.048695 | 0.88115 | 0.042908 | 18.0953 | 22.0399 | 3.4589 | 0.045108 | 0.3007 | 0.3689 | 0.2847 | 0.1163
2012-10-30 22:47 | 26.3644 | 22.362 | 22.35 | 22.2738 | 22.1976 | 50.1646 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048664 | 0.88049 | 0.042848 | 18.0932 | 22.0332 | 3.4583 | 0.045042 | 0.3288 | 0.3168 | 0.2406 | 0.1644
2012-10-30 22:47 | 26.3681 | 22.3657 | 22.3256 | 22.2454 | 22.1652 | 50.164 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048806 | 0.8827 | 0.043081 | 18.0858 | 22.0409 | 3.4583 | 0.045275 | 0.3248 | 0.2847 | 0.2045 | 0.1243
2012-10-30 22:47 | 26.3681 | 22.3617 | 22.3497 | 22.2615 | 22.1612 | 50.1767 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048724 | 0.88155 | 0.042953 | 18.0928 | 22.0609 | 3.4592 | 0.045149 | 0.3008 | 0.2888 | 0.2006 | 0.1003
2012-10-30 22:47 | 26.3689 | 22.3304 | 22.3625 | 22.2262 | 22.174 | 50.171 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048702 | 0.88105 | 0.042909 | 18.0906 | 22.0377 | 3.4588 | 0.045106 | 0.2927 | 0.3248 | 0.1885 | 0.1363
2012-10-30 22:47 | 26.3733 | 22.3108 | 22.3749 | 22.2707 | 22.1544 | 50.164 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048692 | 0.88101 | 0.042898 | 18.0934 | 22.0661 | 3.4583 | 0.045093 | 0.2447 | 0.3088 | 0.2046 | 0.0883
2012-10-30 22:48 | 26.3741 | 22.3236 | 22.3757 | 22.3035 | 22.1792 | 50.1727 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048727 | 0.88153 | 0.042954 | 18.0911 | 22.0268 | 3.4589 | 0.045148 | 0.2968 | 0.3489 | 0.2767 | 0.1524
2012-10-30 22:48 | 26.3763 | 22.3739 | 22.414 | 22.2416 | 22.1413 | 50.1663 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048743 | 0.88177 | 0.04298 | 18.0902 | 22.017 | 3.4585 | 0.045173 | 0.3569 | 0.397 | 0.2246 | 0.1243
2012-10-30 22:48 | 26.3651 | 22.3707 | 22.3507 | 22.2464 | 22.1542 | 50.1559 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.048818 | 0.88283 | 0.043098 | 18.084 | 22.0419 | 3.4577 | 0.04529 | 0.3288 | 0.3088 | 0.2045 | 0.1123
2012-10-30 22:48 | 26.3704 | 22.364 | 22.364 | 22.2477 | 22.1435 | 50.1687 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048695 | 0.88093 | 0.042897 | 18.0909 | 22.0191 | 3.4586 | 0.045092 | 0.3449 | 0.3449 | 0.2286 | 0.1244
2012-10-30 22:48 | 26.3659 | 22.3515 | 22.3715 | 22.2512 | 22.151 | 50.1629 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.048686 | 0.88069 | 0.042877 | 18.0891 | 22.0467 | 3.4582 | 0.045076 | 0.3048 | 0.3248 | 0.2045 | 0.1043
2012-10-30 22:48 | 26.3689 | 22.3344 | 22.3865 | 22.2703 | 22.158 | 50.16 | 0.006729 | 0.32644 | 0.002197 | 48.5127 | 0.048692 | 0.88101 | 0.042898 | 18.0934 | 22.0337 | 3.458 | 0.045095 | 0.3007 | 0.3528 | 0.2366 | 0.1243
2012-10-30 22:48 | 26.3719 | 22.3655 | 22.3615 | 22.3054 | 22.1329 | 50.1606 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.048752 | 0.8819 | 0.042994 | 18.0896 | 22.0848 | 3.4581 | 0.045186 | 0.2807 | 0.2767 | 0.2206 | 0.0481
2012-10-30 22:48 | 26.3726 | 22.3582 | 22.3782 | 22.278 | 22.1376 | 50.1715 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048677 | 0.88066 | 0.042868 | 18.0919 | 22.0414 | 3.4588 | 0.045061 | 0.3168 | 0.3368 | 0.2366 | 0.0962
2012-10-30 22:48 | 26.3763 | 22.3458 | 22.418 | 22.2777 | 22.1213 | 50.1571 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048774 | 0.88231 | 0.043034 | 18.0897 | 22.029 | 3.4578 | 0.045229 | 0.3168 | 0.389 | 0.2487 | 0.0923
2012-10-30 22:48 | 26.3666 | 22.3762 | 22.3401 | 22.2519 | 22.1557 | 50.1669 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048743 | 0.88182 | 0.042982 | 18.0911 | 22.0354 | 3.4585 | 0.045177 | 0.3408 | 0.3047 | 0.2165 | 0.1203
2012-10-30 22:48 | 26.3636 | 22.3652 | 22.3772 | 22.265 | 22.1367 | 50.1542 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048831 | 0.88322 | 0.043128 | 18.0872 | 22.0284 | 3.4576 | 0.045321 | 0.3368 | 0.3488 | 0.2366 | 0.1083
2012-10-30 22:48 | 26.3681 | 22.3657 | 22.3777 | 22.2735 | 22.1692 | 50.1617 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048724 | 0.88157 | 0.042954 | 18.0931 | 22.0288 | 3.4581 | 0.045149 | 0.3369 | 0.3489 | 0.2447 | 0.1404
2012-10-30 22:48 | 26.3666 | 22.3842 | 22.3802 | 22.268 | 22.1517 | 50.16 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.048683 | 0.88067 | 0.042874 | 18.09 | 22.0634 | 3.458 | 0.045074 | 0.3208 | 0.3168 | 0.2046 | 0.0883
2012-10-30 22:48 | 26.3756 | 22.3251 | 22.3652 | 22.297 | 22.1807 | 50.1675 | 0.006716 | 0.32631 | 0.002192 | 48.5873 | 0.048651 | 0.8806 | 0.042842 | 18.1004 | 22.0684 | 3.4585 | 0.045034 | 0.2567 | 0.2968 | 0.2286 | 0.1123
2012-10-30 22:48 | 26.3711 | 22.3326 | 22.3607 | 22.2645 | 22.1682 | 50.1577 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048733 | 0.88147 | 0.042957 | 18.0877 | 22.0439 | 3.4579 | 0.04515 | 0.2887 | 0.3168 | 0.2206 | 0.1243
2012-10-30 22:48 | 26.3644 | 22.39 | 22.362 | 22.2578 | 22.1495 | 50.1663 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048752 | 0.88193 | 0.042996 | 18.0901 | 22.0091 | 3.4585 | 0.04519 | 0.3809 | 0.3529 | 0.2487 | 0.1404
2012-10-30 22:48 | 26.3704 | 22.38 | 22.376 | 22.2638 | 22.1916 | 50.1565 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.04867 | 0.88062 | 0.04286 | 18.0936 | 22.0592 | 3.4578 | 0.045057 | 0.3208 | 0.3168 | 0.2046 | 0.1324
2012-10-30 22:48 | 26.3726 | 22.3542 | 22.3942 | 22.282 | 22.1416 | 50.1617 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048762 | 0.88218 | 0.043017 | 18.0915 | 22.0414 | 3.4581 | 0.045214 | 0.3128 | 0.3528 | 0.2406 | 0.1002
2012-10-30 22:48 | 26.3666 | 22.3842 | 22.3562 | 22.288 | 22.1758 | 50.1652 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048673 | 0.88077 | 0.04287 | 18.0958 | 21.9993 | 3.4584 | 0.045066 | 0.3849 | 0.3569 | 0.2887 | 0.1765
2012-10-30 22:48 | 26.3666 | 22.3802 | 22.3802 | 22.284 | 22.1637 | 50.1577 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.04873 | 0.88169 | 0.042965 | 18.0933 | 22.0394 | 3.4579 | 0.045158 | 0.3408 | 0.3408 | 0.2446 | 0.1243
2012-10-30 22:48 | 26.3569 | 22.3465 | 22.3745 | 22.2823 | 22.1861 | 50.1646 | 0.006703 | 0.32636 | 0.002188 | 48.6879 | 0.048708 | 0.88119 | 0.042921 | 18.0914 | 22.0457 | 3.4583 | 0.045109 | 0.3008 | 0.3288 | 0.2366 | 0.1404
2012-10-30 22:48 | 26.3674 | 22.381 | 22.353 | 22.2487 | 22.1605 | 50.1611 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048749 | 0.88174 | 0.042984 | 18.0874 | 22.0362 | 3.4581 | 0.045177 | 0.3448 | 0.3168 | 0.2125 | 0.1243
2012-10-30 22:48 | 26.3659 | 22.3755 | 22.3675 | 22.2873 | 22.155 | 50.1681 | 0.006713 | 0.32628 | 0.00219 | 48.6046 | 0.048717 | 0.88124 | 0.042931 | 18.0889 | 22.0467 | 3.4586 | 0.045121 | 0.3288 | 0.3208 | 0.2406 | 0.1083
2012-10-30 22:48 | 26.3584 | 22.364 | 22.348 | 22.2598 | 22.1475 | 50.1635 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048683 | 0.8808 | 0.04288 | 18.0926 | 22.0713 | 3.4583 | 0.045076 | 0.2927 | 0.2767 | 0.1885 | 0.0762
2012-10-30 22:48 | 26.3464 | 22.332 | 22.34 | 22.2157 | 22.1195 | 50.1542 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048702 | 0.88112 | 0.042912 | 18.0921 | 22.0392 | 3.4576 | 0.045108 | 0.2928 | 0.3008 | 0.1765 | 0.0803
2012-10-30 22:48 | 26.3509 | 22.3765 | 22.3445 | 22.2162 | 22.128 | 50.1692 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.048762 | 0.88183 | 0.043 | 18.0843 | 22.0558 | 3.4587 | 0.045198 | 0.3207 | 0.2887 | 0.1604 | 0.0722
2012-10-30 22:49 | 26.3524 | 22.3299 | 22.366 | 22.2818 | 22.1696 | 50.1577 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048661 | 0.88067 | 0.042854 | 18.0981 | 22.0091 | 3.4579 | 0.045047 | 0.3208 | 0.3569 | 0.2727 | 0.1605
2012-10-30 22:49 | 26.3501 | 22.3517 | 22.3276 | 22.2635 | 22.1392 | 50.1554 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.04874 | 0.88167 | 0.042973 | 18.0892 | 21.9948 | 3.4577 | 0.045167 | 0.3569 | 0.3328 | 0.2687 | 0.1444
2012-10-30 22:49 | 26.3584 | 22.34 | 22.364 | 22.2678 | 22.1395 | 50.1652 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.04873 | 0.88171 | 0.042966 | 18.0939 | 22.0753 | 3.4584 | 0.045163 | 0.2647 | 0.2887 | 0.1925 | 0.0642
2012-10-30 22:49 | 26.3487 | 22.3904 | 22.3663 | 22.2621 | 22.1538 | 50.1559 | 0.006719 | 0.3262 | 0.002192 | 48.5483 | 0.048752 | 0.88225 | 0.043011 | 18.0967 | 22.0295 | 3.4577 | 0.045203 | 0.3609 | 0.3368 | 0.2326 | 0.1243
2012-10-30 22:49 | 26.3487 | 22.3904 | 22.3222 | 22.226 | 22.1338 | 50.1559 | 0.006738 | 0.32622 | 0.002198 | 48.4157 | 0.04868 | 0.88073 | 0.042874 | 18.0923 | 22.0375 | 3.4577 | 0.045072 | 0.3529 | 0.2847 | 0.1885 | 0.0963
2012-10-30 22:49 | 26.3427 | 22.3403 | 22.3403 | 22.2762 | 22.1438 | 50.1536 | 0.006735 | 0.32625 | 0.002197 | 48.4416 | 0.048777 | 0.88235 | 0.043038 | 18.0894 | 21.9954 | 3.4576 | 0.045235 | 0.3449 | 0.3449 | 0.2808 | 0.1484
2012-10-30 22:49 | 26.3501 | 22.3958 | 22.3557 | 22.2916 | 22.1232 | 50.1588 | 0.006719 | 0.32621 | 0.002192 | 48.5505 | 0.048771 | 0.88213 | 0.043023 | 18.0872 | 22.0149 | 3.4579 | 0.045215 | 0.3809 | 0.3408 | 0.2767 | 0.1083
2012-10-30 22:49 | 26.3501 | 22.3958 | 22.3557 | 22.2836 | 22.1633 | 50.1681 | 0.006735 | 0.32625 | 0.002197 | 48.4416 | 0.048711 | 0.88138 | 0.042933 | 18.0941 | 22.0149 | 3.4586 | 0.04513 | 0.3809 | 0.3408 | 0.2687 | 0.1484
2012-10-30 22:49 | 26.3516 | 22.3933 | 22.3692 | 22.269 | 22.1447 | 50.1525 | 0.006707 | 0.32628 | 0.002188 | 48.6481 | 0.048692 | 0.88098 | 0.042897 | 18.0928 | 22.0123 | 3.4575 | 0.045085 | 0.381 | 0.3569 | 0.2567 | 0.1324
2012-10-30 22:49 | 26.3479 | 22.3816 | 22.3816 | 22.2773 | 22.153 | 50.1525 | 0.006738 | 0.32631 | 0.002199 | 48.4286 | 0.04867 | 0.88053 | 0.042855 | 18.0918 | 22.0207 | 3.4575 | 0.045054 | 0.3609 | 0.3609 | 0.2566 | 0.1323
2012-10-30 22:49 | 26.3397 | 22.3533 | 22.3373 | 22.2812 | 22.189 | 50.1698 | 0.006722 | 0.32625 | 0.002193 | 48.5353 | 0.048705 | 0.88127 | 0.042922 | 18.094 | 22.0726 | 3.4587 | 0.045115 | 0.2807 | 0.2647 | 0.2086 | 0.1164
2012-10-30 22:49 | 26.3262 | 22.3919 | 22.3719 | 22.2677 | 22.1273 | 50.1525 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048755 | 0.88186 | 0.042995 | 18.0875 | 22.0591 | 3.4575 | 0.045187 | 0.3328 | 0.3128 | 0.2086 | 0.0682
2012-10-30 22:49 | 26.3284 | 22.35 | 22.334 | 22.2458 | 22.1576 | 50.16 | 0.006716 | 0.32622 | 0.002191 | 48.5743 | 0.048711 | 0.88111 | 0.04292 | 18.0885 | 22.0493 | 3.458 | 0.045111 | 0.3007 | 0.2847 | 0.1965 | 0.1083
2012-10-30 22:49 | 26.3337 | 22.3994 | 22.3674 | 22.2391 | 22.191 | 50.1629 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.048695 | 0.88091 | 0.042896 | 18.0903 | 22.0707 | 3.4582 | 0.045094 | 0.3287 | 0.2967 | 0.1684 | 0.1203
2012-10-30 22:49 | 26.3307 | 22.4044 | 22.3924 | 22.2481 | 22.1358 | 50.1669 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048702 | 0.88101 | 0.042907 | 18.0897 | 22.0877 | 3.4585 | 0.045103 | 0.3167 | 0.3047 | 0.1604 | 0.0481
2012-10-30 22:49 | 26.3389 | 22.3846 | 22.4086 | 22.2643 | 22.1481 | 50.1635 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048689 | 0.88102 | 0.042896 | 18.0948 | 22.0398 | 3.4583 | 0.045091 | 0.3448 | 0.3688 | 0.2245 | 0.1083
2012-10-30 22:49 | 26.3322 | 22.3699 | 22.3458 | 22.2777 | 22.1333 | 50.171 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048733 | 0.88153 | 0.042959 | 18.0889 | 22.0531 | 3.4588 | 0.045151 | 0.3168 | 0.2927 | 0.2246 | 0.0802
2012-10-30 22:49 | 26.3344 | 22.4122 | 22.3721 | 22.2518 | 22.0834 | 50.1461 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048746 | 0.88187 | 0.042988 | 18.0912 | 22.0272 | 3.4571 | 0.045185 | 0.385 | 0.3449 | 0.2246 | 0.0562
2012-10-30 22:49 | 26.3254 | 22.3751 | 22.347 | 22.2348 | 22.1426 | 50.1663 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048686 | 0.88093 | 0.042889 | 18.0942 | 22.0463 | 3.4585 | 0.045082 | 0.3288 | 0.3007 | 0.1885 | 0.0963
2012-10-30 22:49 | 26.3254 | 22.3591 | 22.343 | 22.2669 | 22.1265 | 50.1548 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.048702 | 0.88105 | 0.042909 | 18.0906 | 22.0543 | 3.4577 | 0.045101 | 0.3048 | 0.2887 | 0.2126 | 0.0722
2012-10-30 22:49 | 26.3367 | 22.3744 | 22.3864 | 22.2501 | 22.1178 | 50.1571 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.048705 | 0.88127 | 0.042922 | 18.094 | 22.0456 | 3.4578 | 0.045122 | 0.3288 | 0.3408 | 0.2045 | 0.0722
2012-10-30 22:49 | 26.324 | 22.3737 | 22.3296 | 22.2494 | 22.1372 | 50.1548 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048677 | 0.88079 | 0.042874 | 18.0946 | 22.0489 | 3.4577 | 0.045069 | 0.3248 | 0.2807 | 0.2005 | 0.0883
2012-10-30 22:49 | 26.3299 | 22.3716 | 22.3355 | 22.2553 | 22.135 | 50.1542 | 0.006719 | 0.3264 | 0.002193 | 48.5785 | 0.04873 | 0.88151 | 0.042956 | 18.0897 | 22.0267 | 3.4576 | 0.045149 | 0.3449 | 0.3088 | 0.2286 | 0.1083
2012-10-30 22:49 | 26.3232 | 22.3809 | 22.3609 | 22.2166 | 22.1324 | 50.149 | 0.006725 | 0.32627 | 0.002194 | 48.5157 | 0.048711 | 0.88125 | 0.042927 | 18.0914 | 22.0361 | 3.4573 | 0.045121 | 0.3448 | 0.3248 | 0.1805 | 0.0963
2012-10-30 22:49 | 26.3254 | 22.3591 | 22.339 | 22.2829 | 22.1346 | 50.1519 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048692 | 0.88076 | 0.042886 | 18.0884 | 22.0022 | 3.4575 | 0.045083 | 0.3569 | 0.3368 | 0.2807 | 0.1324
2012-10-30 22:49 | 26.3202 | 22.402 | 22.3659 | 22.2336 | 22.1534 | 50.1606 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.04878 | 0.88228 | 0.043037 | 18.0869 | 22.0251 | 3.4581 | 0.045233 | 0.3769 | 0.3408 | 0.2085 | 0.1283
2012-10-30 22:50 | 26.3277 | 22.3574 | 22.3413 | 22.2692 | 22.1449 | 50.145 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048733 | 0.88164 | 0.042965 | 18.0913 | 22.0486 | 3.457 | 0.045158 | 0.3088 | 0.2927 | 0.2206 | 0.0963
2012-10-30 22:50 | 26.3202 | 22.3458 | 22.3819 | 22.2336 | 22.1093 | 50.1536 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048717 | 0.88135 | 0.042937 | 18.0913 | 22.0451 | 3.4576 | 0.045131 | 0.3007 | 0.3368 | 0.1885 | 0.0642
2012-10-30 22:50 | 26.3299 | 22.3676 | 22.3515 | 22.2393 | 22.1591 | 50.1479 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048758 | 0.88196 | 0.043003 | 18.0885 | 22.0588 | 3.4572 | 0.0452 | 0.3088 | 0.2927 | 0.1805 | 0.1003
2012-10-30 22:50 | 26.3322 | 22.3739 | 22.3619 | 22.2336 | 22.1093 | 50.1507 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048825 | 0.88313 | 0.043119 | 18.0877 | 22.0732 | 3.4574 | 0.045312 | 0.3007 | 0.2887 | 0.1604 | 0.0361
2012-10-30 22:50 | 26.3254 | 22.3631 | 22.351 | 22.2147 | 22.1426 | 50.1658 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048683 | 0.8808 | 0.04288 | 18.0926 | 22.0704 | 3.4584 | 0.045075 | 0.2927 | 0.2806 | 0.1443 | 0.0722
2012-10-30 22:50 | 26.3232 | 22.3368 | 22.3609 | 22.2687 | 22.1845 | 50.1444 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048702 | 0.88108 | 0.04291 | 18.0912 | 22.0401 | 3.457 | 0.045106 | 0.2967 | 0.3208 | 0.2286 | 0.1444
2012-10-30 22:50 | 26.3262 | 22.3398 | 22.3478 | 22.2597 | 22.1313 | 50.1461 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048721 | 0.88132 | 0.042939 | 18.0892 | 22.0792 | 3.4571 | 0.045132 | 0.2606 | 0.2686 | 0.1805 | 0.0521
2012-10-30 22:50 | 26.324 | 22.3296 | 22.3897 | 22.2575 | 22.1291 | 50.1484 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048771 | 0.88215 | 0.043023 | 18.0875 | 22.081 | 3.4572 | 0.045222 | 0.2486 | 0.3087 | 0.1765 | 0.0481
2012-10-30 22:50 | 26.3254 | 22.319 | 22.347 | 22.2228 | 22.1185 | 50.1473 | 0.006732 | 0.32637 | 0.002197 | 48.4804 | 0.048714 | 0.88127 | 0.04293 | 18.0906 | 22.0864 | 3.4572 | 0.045127 | 0.2326 | 0.2606 | 0.1364 | 0.0321
2012-10-30 22:50 | 26.3232 | 22.3328 | 22.3729 | 22.2687 | 22.1283 | 50.1479 | 0.006732 | 0.3264 | 0.002197 | 48.4847 | 0.048692 | 0.88095 | 0.042895 | 18.0923 | 22.0521 | 3.4572 | 0.045092 | 0.2807 | 0.3208 | 0.2166 | 0.0762
2012-10-30 22:50 | 26.3269 | 22.3806 | 22.3566 | 22.2443 | 22.1561 | 50.1594 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048733 | 0.8815 | 0.042958 | 18.0883 | 22.0518 | 3.458 | 0.045152 | 0.3288 | 0.3048 | 0.1925 | 0.1043
2012-10-30 22:50 | 26.324 | 22.3296 | 22.3737 | 22.2334 | 22.1452 | 50.1646 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.048717 | 0.88114 | 0.042926 | 18.0868 | 22.0329 | 3.4583 | 0.045125 | 0.2967 | 0.3408 | 0.2005 | 0.1123
2012-10-30 22:50 | 26.3322 | 22.3178 | 22.3819 | 22.2777 | 22.1373 | 50.1461 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048667 | 0.88066 | 0.042859 | 18.0956 | 22.0451 | 3.4571 | 0.045054 | 0.2727 | 0.3368 | 0.2326 | 0.0922
2012-10-30 22:50 | 26.3269 | 22.3525 | 22.3566 | 22.2924 | 22.132 | 50.1583 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048812 | 0.88281 | 0.043092 | 18.086 | 22.0438 | 3.4579 | 0.045286 | 0.3087 | 0.3128 | 0.2486 | 0.0882
2012-10-30 22:50 | 26.3284 | 22.326 | 22.3581 | 22.2418 | 22.1937 | 50.1513 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048692 | 0.88086 | 0.042891 | 18.0905 | 21.9972 | 3.4574 | 0.045087 | 0.3288 | 0.3609 | 0.2446 | 0.1965
2012-10-30 22:50 | 26.3337 | 22.3473 | 22.3353 | 22.2391 | 22.1429 | 50.1519 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048749 | 0.88194 | 0.042994 | 18.0916 | 22.0666 | 3.4575 | 0.045189 | 0.2807 | 0.2687 | 0.1725 | 0.0763
2012-10-30 22:50 | 26.324 | 22.3136 | 22.3296 | 22.2414 | 22.1051 | 50.1502 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048695 | 0.88098 | 0.042899 | 18.0917 | 22.0489 | 3.4574 | 0.045093 | 0.2647 | 0.2807 | 0.1925 | 0.0562
2012-10-30 22:50 | 26.3322 | 22.3338 | 22.3498 | 22.2737 | 22.1574 | 50.1421 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048686 | 0.88105 | 0.042895 | 18.0966 | 21.9929 | 3.4568 | 0.045087 | 0.3409 | 0.3569 | 0.2808 | 0.1645
2012-10-30 22:50 | 26.3344 | 22.332 | 22.3681 | 22.2358 | 22.1436 | 50.1588 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.048749 | 0.88163 | 0.042978 | 18.085 | 22.0513 | 3.4579 | 0.045178 | 0.2807 | 0.3168 | 0.1845 | 0.0923
2012-10-30 22:50 | 26.3374 | 22.3711 | 22.347 | 22.2468 | 22.1305 | 50.149 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.04873 | 0.88141 | 0.042951 | 18.0876 | 22.0142 | 3.4573 | 0.045146 | 0.3569 | 0.3328 | 0.2326 | 0.1163
2012-10-30 22:50 | 26.3344 | 22.348 | 22.356 | 22.2438 | 22.1195 | 50.1467 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048705 | 0.88132 | 0.042925 | 18.0951 | 22.0393 | 3.4571 | 0.045118 | 0.3087 | 0.3167 | 0.2045 | 0.0802
2012-10-30 22:50 | 26.3344 | 22.36 | 22.3841 | 22.2278 | 22.2077 | 50.1525 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.04874 | 0.88167 | 0.042973 | 18.0892 | 22.0152 | 3.4575 | 0.045171 | 0.3448 | 0.3689 | 0.2126 | 0.1925
2012-10-30 22:50 | 26.3464 | 22.36 | 22.34 | 22.2879 | 22.0994 | 50.1652 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.048765 | 0.8821 | 0.043016 | 18.0889 | 22.0633 | 3.4584 | 0.045209 | 0.2967 | 0.2767 | 0.2246 | 0.0361
2012-10-30 22:50 | 26.3277 | 22.3413 | 22.3173 | 22.2612 | 22.1128 | 50.1525 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048777 | 0.88231 | 0.043036 | 18.0886 | 22.0245 | 3.4575 | 0.045233 | 0.3168 | 0.2928 | 0.2367 | 0.0883
2012-10-30 22:50 | 26.3299 | 22.3596 | 22.3515 | 22.2513 | 22.135 | 50.1519 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048705 | 0.88118 | 0.042918 | 18.0922 | 22.0508 | 3.4575 | 0.045113 | 0.3088 | 0.3007 | 0.2005 | 0.0842
2012-10-30 22:51 | 26.3307 | 22.3483 | 22.3443 | 22.2521 | 22.1238 | 50.1473 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048784 | 0.88232 | 0.043043 | 18.0863 | 22.0436 | 3.4572 | 0.045239 | 0.3047 | 0.3007 | 0.2085 | 0.0802
2012-10-30 22:51 | 26.3262 | 22.3478 | 22.3478 | 22.2356 | 22.1594 | 50.1461 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048762 | 0.88205 | 0.04301 | 18.0888 | 21.999 | 3.4571 | 0.045205 | 0.3488 | 0.3488 | 0.2366 | 0.1604
2012-10-30 22:51 | 26.3284 | 22.322 | 22.354 | 22.2538 | 22.1255 | 50.1479 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048717 | 0.8813 | 0.042934 | 18.0901 | 22.0573 | 3.4572 | 0.045129 | 0.2647 | 0.2967 | 0.1965 | 0.0682
2012-10-30 22:51 | 26.324 | 22.2975 | 22.3577 | 22.2133 | 22.1372 | 50.1583 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048736 | 0.88155 | 0.042963 | 18.0884 | 22.0128 | 3.4579 | 0.045159 | 0.2847 | 0.3449 | 0.2005 | 0.1244
2012-10-30 22:51 | 26.324 | 22.3296 | 22.3416 | 22.2254 | 22.1372 | 50.1531 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048736 | 0.88169 | 0.04297 | 18.091 | 22.0409 | 3.4576 | 0.045165 | 0.2887 | 0.3007 | 0.1845 | 0.0963
2012-10-30 22:51 | 26.3232 | 22.3529 | 22.3288 | 22.2366 | 22.1003 | 50.1525 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048664 | 0.8808 | 0.042863 | 18.0997 | 22.0441 | 3.4575 | 0.045057 | 0.3088 | 0.2847 | 0.1925 | 0.0562
2012-10-30 22:51 | 26.3277 | 22.2972 | 22.4135 | 22.2411 | 22.185 | 50.1479 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048677 | 0.8808 | 0.042875 | 18.0949 | 22.0125 | 3.4572 | 0.045067 | 0.2847 | 0.401 | 0.2286 | 0.1725
2012-10-30 22:51 | 26.321 | 22.3466 | 22.3587 | 22.2665 | 22.1101 | 50.1542 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048683 | 0.88089 | 0.042884 | 18.0944 | 21.9978 | 3.4576 | 0.045079 | 0.3488 | 0.3609 | 0.2687 | 0.1123
2012-10-30 22:51 | 26.3195 | 22.3171 | 22.3451 | 22.2169 | 22.1206 | 50.1502 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048705 | 0.88109 | 0.042914 | 18.0904 | 22.0645 | 3.4574 | 0.04511 | 0.2526 | 0.2806 | 0.1524 | 0.0561
2012-10-30 22:51 | 26.3165 | 22.3421 | 22.3061 | 22.274 | 22.1577 | 50.1432 | 0.006738 | 0.32631 | 0.002199 | 48.4286 | 0.048699 | 0.88109 | 0.042908 | 18.0926 | 22.0053 | 3.4569 | 0.045107 | 0.3368 | 0.3008 | 0.2687 | 0.1524
2012-10-30 22:51 | 26.3307 | 22.3724 | 22.3604 | 22.2601 | 22.188 | 50.1611 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048746 | 0.88197 | 0.042993 | 18.0933 | 21.9634 | 3.4581 | 0.045187 | 0.409 | 0.397 | 0.2967 | 0.2246
2012-10-30 22:51 | 26.315 | 22.3447 | 22.3567 | 22.2685 | 22.1121 | 50.1554 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048661 | 0.88073 | 0.042857 | 18.0993 | 21.9958 | 3.4577 | 0.045052 | 0.3489 | 0.3609 | 0.2727 | 0.1163
2012-10-30 22:51 | 26.3127 | 22.3063 | 22.3063 | 22.2261 | 22.1098 | 50.1473 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048699 | 0.88128 | 0.042917 | 18.0965 | 22.0296 | 3.4572 | 0.045111 | 0.2767 | 0.2767 | 0.1965 | 0.0802
2012-10-30 22:51 | 26.3075 | 22.3412 | 22.3331 | 22.2369 | 22.1207 | 50.1415 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.04873 | 0.88166 | 0.042963 | 18.0927 | 21.9883 | 3.4568 | 0.045158 | 0.3529 | 0.3448 | 0.2486 | 0.1324
2012-10-30 22:51 | 26.312 | 22.2895 | 22.3176 | 22.2374 | 22.1372 | 50.1374 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048702 | 0.88116 | 0.042914 | 18.093 | 22.053 | 3.4565 | 0.045108 | 0.2365 | 0.2646 | 0.1844 | 0.0842
2012-10-30 22:51 | 26.3127 | 22.3744 | 22.3784 | 22.2221 | 22.0898 | 50.138 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.048699 | 0.88118 | 0.042913 | 18.0944 | 21.9654 | 3.4565 | 0.045105 | 0.409 | 0.413 | 0.2567 | 0.1244
2012-10-30 22:51 | 26.3112 | 22.3288 | 22.3208 | 22.2767 | 22.1244 | 50.1438 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.048708 | 0.88099 | 0.042911 | 18.0872 | 21.988 | 3.4569 | 0.045104 | 0.3408 | 0.3328 | 0.2887 | 0.1364
2012-10-30 22:51 | 26.3112 | 22.3368 | 22.3328 | 22.1965 | 22.1163 | 50.1426 | 0.006707 | 0.32634 | 0.002189 | 48.6568 | 0.048683 | 0.88086 | 0.042883 | 18.0938 | 22.0201 | 3.4568 | 0.045072 | 0.3167 | 0.3127 | 0.1764 | 0.0962
2012-10-30 22:51 | 26.312 | 22.3657 | 22.3256 | 22.2455 | 22.1452 | 50.1525 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048736 | 0.88161 | 0.042966 | 18.0896 | 21.9808 | 3.4575 | 0.045159 | 0.3849 | 0.3448 | 0.2647 | 0.1644
2012-10-30 22:51 | 26.3037 | 22.3694 | 22.3013 | 22.2091 | 22.1289 | 50.1479 | 0.006713 | 0.32628 | 0.00219 | 48.6046 | 0.048705 | 0.88108 | 0.042913 | 18.0901 | 21.9604 | 3.4572 | 0.045103 | 0.409 | 0.3409 | 0.2487 | 0.1685
2012-10-30 22:51 | 26.3082 | 22.3739 | 22.3338 | 22.2376 | 22.1374 | 50.1444 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.048721 | 0.88131 | 0.042938 | 18.0889 | 21.977 | 3.457 | 0.045133 | 0.3969 | 0.3568 | 0.2606 | 0.1604
2012-10-30 22:51 | 26.3008 | 22.3385 | 22.3064 | 22.2222 | 22.1581 | 50.145 | 0.006738 | 0.32627 | 0.002198 | 48.4221 | 0.048699 | 0.88102 | 0.042905 | 18.0911 | 21.9575 | 3.457 | 0.045103 | 0.381 | 0.3489 | 0.2647 | 0.2006
2012-10-30 22:51 | 26.3075 | 22.3853 | 22.3211 | 22.2289 | 22.1487 | 50.1467 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048673 | 0.88069 | 0.042866 | 18.094 | 21.9843 | 3.4571 | 0.04506 | 0.401 | 0.3368 | 0.2446 | 0.1644
2012-10-30 22:51 | 26.3082 | 22.3699 | 22.3058 | 22.1975 | 22.1133 | 50.1484 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048743 | 0.88179 | 0.042981 | 18.0905 | 22.0091 | 3.4572 | 0.045175 | 0.3608 | 0.2967 | 0.1884 | 0.1042
2012-10-30 22:51 | 26.3067 | 22.3845 | 22.3163 | 22.2241 | 22.1199 | 50.1473 | 0.006713 | 0.32636 | 0.002191 | 48.6154 | 0.048768 | 0.88208 | 0.043017 | 18.0872 | 22.0076 | 3.4572 | 0.045208 | 0.3769 | 0.3087 | 0.2165 | 0.1123
2012-10-30 22:52 | 26.3037 | 22.3093 | 22.3494 | 22.2091 | 22.1169 | 50.1502 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.048673 | 0.88096 | 0.042879 | 18.0996 | 21.9604 | 3.4574 | 0.045077 | 0.3489 | 0.389 | 0.2487 | 0.1565
2012-10-30 22:52 | 26.3067 | 22.3283 | 22.3043 | 22.2522 | 22.1319 | 50.149 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048689 | 0.88105 | 0.042897 | 18.0954 | 22.0116 | 3.4573 | 0.045093 | 0.3167 | 0.2927 | 0.2406 | 0.1203
2012-10-30 22:52 | 26.3022 | 22.3559 | 22.3238 | 22.2076 | 22.0873 | 50.1432 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048699 | 0.88112 | 0.04291 | 18.0932 | 21.991 | 3.4569 | 0.045106 | 0.3649 | 0.3328 | 0.2166 | 0.0963
2012-10-30 22:52 | 26.3052 | 22.3669 | 22.3308 | 22.1905 | 22.0983 | 50.164 | 0.006729 | 0.3264 | 0.002196 | 48.5063 | 0.048677 | 0.88082 | 0.042876 | 18.0952 | 21.9579 | 3.4583 | 0.045072 | 0.409 | 0.3729 | 0.2326 | 0.1404
2012-10-30 22:52 | 26.3015 | 22.3392 | 22.2991 | 22.2229 | 22.1066 | 50.1559 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.048705 | 0.88124 | 0.042921 | 18.0934 | 21.9863 | 3.4577 | 0.045115 | 0.3529 | 0.3128 | 0.2366 | 0.1203
2012-10-30 22:52 | 26.306 | 22.2916 | 22.3116 | 22.2635 | 22.1272 | 50.1513 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.048743 | 0.88186 | 0.042984 | 18.092 | 21.9988 | 3.4574 | 0.045176 | 0.2928 | 0.3128 | 0.2647 | 0.1284
2012-10-30 22:52 | 26.315 | 22.3807 | 22.3086 | 22.2485 | 22.1201 | 50.1513 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048705 | 0.88122 | 0.04292 | 18.0931 | 22.0118 | 3.4574 | 0.045113 | 0.3689 | 0.2968 | 0.2367 | 0.1083
2012-10-30 22:52 | 26.312 | 22.3577 | 22.3216 | 22.2174 | 22.1171 | 50.1467 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.048689 | 0.88066 | 0.042878 | 18.0874 | 22.0129 | 3.4571 | 0.045071 | 0.3448 | 0.3087 | 0.2045 | 0.1042
2012-10-30 22:52 | 26.3082 | 22.3138 | 22.3258 | 22.2176 | 22.0772 | 50.1467 | 0.006722 | 0.3264 | 0.002194 | 48.5568 | 0.048677 | 0.88086 | 0.042878 | 18.096 | 21.989 | 3.4571 | 0.045072 | 0.3248 | 0.3368 | 0.2286 | 0.0882
2012-10-30 22:52 | 26.318 | 22.3356 | 22.3436 | 22.2434 | 22.083 | 50.1484 | 0.006716 | 0.3264 | 0.002192 | 48.6002 | 0.048711 | 0.88109 | 0.042919 | 18.0882 | 22.0068 | 3.4572 | 0.045111 | 0.3288 | 0.3368 | 0.2366 | 0.0762
2012-10-30 22:52 | 26.3082 | 22.3419 | 22.3419 | 22.1855 | 22.0933 | 50.1421 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048711 | 0.88153 | 0.04294 | 18.0971 | 22.0091 | 3.4568 | 0.045134 | 0.3328 | 0.3328 | 0.1764 | 0.0842
2012-10-30 22:52 | 26.312 | 22.3096 | 22.3056 | 22.2374 | 22.081 | 50.1513 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048724 | 0.88143 | 0.042947 | 18.0902 | 22.0048 | 3.4574 | 0.045145 | 0.3048 | 0.3008 | 0.2326 | 0.0762
2012-10-30 22:52 | 26.3082 | 22.3258 | 22.3218 | 22.2176 | 22.1214 | 50.1455 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.048752 | 0.88184 | 0.042992 | 18.0884 | 21.9609 | 3.457 | 0.045188 | 0.3649 | 0.3609 | 0.2567 | 0.1605
2012-10-30 22:52 | 26.3075 | 22.3532 | 22.2971 | 22.2089 | 22.1086 | 50.1583 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.048689 | 0.88092 | 0.042891 | 18.0928 | 21.9803 | 3.4579 | 0.045089 | 0.3729 | 0.3168 | 0.2286 | 0.1283
2012-10-30 22:52 | 26.3037 | 22.2772 | 22.3013 | 22.185 | 22.1048 | 50.1502 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048714 | 0.88138 | 0.042936 | 18.093 | 22.0005 | 3.4574 | 0.045132 | 0.2767 | 0.3008 | 0.1845 | 0.1043
2012-10-30 22:52 | 26.3105 | 22.3482 | 22.3121 | 22.2279 | 22.0996 | 50.1444 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048673 | 0.88067 | 0.042865 | 18.0937 | 21.9913 | 3.457 | 0.04506 | 0.3569 | 0.3208 | 0.2366 | 0.1083
2012-10-30 22:52 | 26.3082 | 22.3138 | 22.2817 | 22.2256 | 22.1133 | 50.1536 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048708 | 0.8813 | 0.042926 | 18.0934 | 21.985 | 3.4576 | 0.045118 | 0.3288 | 0.2967 | 0.2406 | 0.1283
2012-10-30 22:52 | 26.3142 | 22.3118 | 22.3278 | 22.2396 | 22.0993 | 50.1444 | 0.006725 | 0.32625 | 0.002194 | 48.5136 | 0.048727 | 0.88164 | 0.04296 | 18.0935 | 21.999 | 3.457 | 0.045154 | 0.3128 | 0.3288 | 0.2406 | 0.1003
2012-10-30 22:52 | 26.3187 | 22.3043 | 22.3043 | 22.196 | 22.1359 | 50.1403 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.048796 | 0.88275 | 0.043075 | 18.0907 | 22.0155 | 3.4567 | 0.045272 | 0.2888 | 0.2888 | 0.1805 | 0.1204
2012-10-30 22:52 | 26.3142 | 22.2677 | 22.3318 | 22.2436 | 22.1073 | 50.1363 | 0.006713 | 0.32628 | 0.00219 | 48.6046 | 0.04868 | 0.88091 | 0.042882 | 18.0958 | 21.9709 | 3.4564 | 0.045072 | 0.2968 | 0.3609 | 0.2727 | 0.1364
2012-10-30 22:52 | 26.318 | 22.3276 | 22.3116 | 22.2515 | 22.1432 | 50.1432 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.048733 | 0.88161 | 0.042964 | 18.0907 | 21.9868 | 3.4569 | 0.045158 | 0.3408 | 0.3248 | 0.2647 | 0.1564
2012-10-30 22:52 | 26.3172 | 22.3148 | 22.3228 | 22.1584 | 22.1384 | 50.1409 | 0.006722 | 0.32628 | 0.002193 | 48.5396 | 0.048758 | 0.88189 | 0.042999 | 18.087 | 21.9739 | 3.4567 | 0.045192 | 0.3409 | 0.3489 | 0.1845 | 0.1645
2012-10-30 22:52 | 26.3225 | 22.288 | 22.3121 | 22.2479 | 22.1317 | 50.1351 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.048777 | 0.88226 | 0.043034 | 18.0877 | 22.0274 | 3.4563 | 0.045231 | 0.2606 | 0.2847 | 0.2205 | 0.1043
2012-10-30 22:52 | 26.3142 | 22.2877 | 22.2998 | 22.2076 | 22.0913 | 50.1513 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.048695 | 0.88111 | 0.042906 | 18.0944 | 21.987 | 3.4574 | 0.045101 | 0.3007 | 0.3128 | 0.2206 | 0.1043
2012-10-30 22:52 | 26.3157 | 22.3173 | 22.3133 | 22.2331 | 22.0847 | 50.1588 | 0.00671 | 0.32631 | 0.00219 | 48.6307 | 0.048673 | 0.88066 | 0.042864 | 18.0934 | 21.9965 | 3.4579 | 0.045054 | 0.3208 | 0.3168 | 0.2366 | 0.0882
2012-10-30 22:53 | 26.3165 | 22.3141 | 22.3061 | 22.2219 | 22.1016 | 50.145 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048727 | 0.88144 | 0.04295 | 18.0893 | 22.0093 | 3.457 | 0.045146 | 0.3048 | 0.2968 | 0.2126 | 0.0923
2012-10-30 22:53 | 26.3172 | 22.3028 | 22.3108 | 22.2065 | 22.0983 | 50.1542 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048727 | 0.88148 | 0.042952 | 18.0902 | 21.99 | 3.4576 | 0.045147 | 0.3128 | 0.3208 | 0.2165 | 0.1083
2012-10-30 22:53 | 26.3082 | 22.2978 | 22.2978 | 22.2016 | 22.0933 | 50.1421 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048746 | 0.88186 | 0.042987 | 18.0909 | 21.993 | 3.4568 | 0.045181 | 0.3048 | 0.3048 | 0.2086 | 0.1003
2012-10-30 22:53 | 26.3135 | 22.3391 | 22.3351 | 22.1908 | 22.1307 | 50.1531 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048705 | 0.88121 | 0.042919 | 18.0928 | 22.0063 | 3.4576 | 0.045115 | 0.3328 | 0.3288 | 0.1845 | 0.1244
2012-10-30 22:53 | 26.3157 | 22.3253 | 22.3213 | 22.2331 | 22.165 | 50.1577 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.048702 | 0.88112 | 0.042912 | 18.0921 | 22.0767 | 3.4579 | 0.045104 | 0.2486 | 0.2446 | 0.1564 | 0.0883
2012-10-30 22:53 | 26.3202 | 22.3458 | 22.3258 | 22.2256 | 22.1053 | 50.1403 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.048711 | 0.88158 | 0.042943 | 18.0983 | 22.0411 | 3.4567 | 0.045136 | 0.3047 | 0.2847 | 0.1845 | 0.0642
2012-10-30 22:53 | 26.3165 | 22.3301 | 22.3181 | 22.1858 | 22.1056 | 50.1513 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.048777 | 0.88222 | 0.043032 | 18.0868 | 22.0133 | 3.4574 | 0.045229 | 0.3168 | 0.3048 | 0.1725 | 0.0923
2012-10-30 22:53 | 26.3172 | 22.2988 | 22.3188 | 22.2707 | 22.1063 | 50.138 | 0.006735 | 0.32633 | 0.002198 | 48.4523 | 0.048717 | 0.88138 | 0.042938 | 18.0919 | 22.0221 | 3.4565 | 0.045136 | 0.2767 | 0.2967 | 0.2486 | 0.0842
2012-10-30 22:53 | 26.3202 | 22.2977 | 22.3058 | 22.2697 | 22.0812 | 50.1583 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048651 | 0.88043 | 0.042834 | 18.0968 | 22.0331 | 3.4579 | 0.045029 | 0.2646 | 0.2727 | 0.2366 | 0.0481
2012-10-30 22:53 | 26.3135 | 22.3111 | 22.3231 | 22.2069 | 22.1227 | 50.1675 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048648 | 0.88046 | 0.042832 | 18.0985 | 22.0184 | 3.4585 | 0.045024 | 0.2927 | 0.3047 | 0.1885 | 0.1043
2012-10-30 22:53 | 26.3075 | 22.3091 | 22.2971 | 22.2249 | 22.1086 | 50.1658 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.048724 | 0.88155 | 0.042953 | 18.0928 | 21.9923 | 3.4584 | 0.04515 | 0.3168 | 0.3048 | 0.2326 | 0.1163
2012-10-30 22:53 | 26.3187 | 22.3123 | 22.3524 | 22.2522 | 22.0837 | 50.1565 | 0.006719 | 0.32631 | 0.002192 | 48.5656 | 0.04867 | 0.88066 | 0.042862 | 18.0945 | 21.9955 | 3.4578 | 0.045054 | 0.3168 | 0.3569 | 0.2567 | 0.0882
2012-10-30 22:53 | 26.315 | 22.3126 | 22.3447 | 22.2124 | 22.1081 | 50.1606 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048721 | 0.88153 | 0.042949 | 18.0933 | 21.9717 | 3.4581 | 0.045143 | 0.3409 | 0.373 | 0.2407 | 0.1364
2012-10-30 22:53 | 26.3187 | 22.2922 | 22.3604 | 22.208 | 22.1359 | 50.1733 | 0.006735 | 0.3263 | 0.002198 | 48.448 | 0.048664 | 0.88057 | 0.042852 | 18.0949 | 22.0196 | 3.4589 | 0.04505 | 0.2726 | 0.3408 | 0.1884 | 0.1163
2012-10-30 22:53 | 26.3127 | 22.2902 | 22.3023 | 22.19 | 22.0497 | 50.1536 | 0.006713 | 0.32631 | 0.002191 | 48.609 | 0.048692 | 0.8808 | 0.042888 | 18.0893 | 22.0176 | 3.4576 | 0.045079 | 0.2726 | 0.2847 | 0.1724 | 0.0321
2012-10-30 22:54 | 26.3225 | 22.276 | 22.3441 | 22.2199 | 22.0875 | 50.1646 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048683 | 0.88088 | 0.042884 | 18.0941 | 22.0033 | 3.4583 | 0.045076 | 0.2727 | 0.3408 | 0.2166 | 0.0842
2012-10-30 22:54 | 26.321 | 22.3026 | 22.3226 | 22.2264 | 22.0981 | 50.1531 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.048755 | 0.88186 | 0.042995 | 18.0875 | 22.0219 | 3.4576 | 0.045189 | 0.2807 | 0.3007 | 0.2045 | 0.0762
2012-10-30 22:54 | 26.3232 | 22.3007 | 22.3248 | 22.2246 | 22.1163 | 50.1681 | 0.006719 | 0.32627 | 0.002192 | 48.559 | 0.048733 | 0.8817 | 0.042968 | 18.0924 | 21.996 | 3.4586 | 0.04516 | 0.3047 | 0.3288 | 0.2286 | 0.1203
2012-10-30 22:54 | 26.3195 | 22.3091 | 22.3331 | 22.2008 | 22.0725 | 50.1519 | 0.006719 | 0.32628 | 0.002192 | 48.5612 | 0.04874 | 0.8818 | 0.042979 | 18.0919 | 22.0003 | 3.4575 | 0.045171 | 0.3088 | 0.3328 | 0.2005 | 0.0722
2012-10-30 22:54 | 26.3195 | 22.269 | 22.3211 | 22.1848 | 22.0966 | 50.1455 | 0.006719 | 0.32627 | 0.002192 | 48.559 | 0.048695 | 0.88114 | 0.042907 | 18.095 | 22.0123 | 3.457 | 0.045099 | 0.2567 | 0.3088 | 0.1725 | 0.0843
2012-10-30 22:54 | 26.318 | 22.2755 | 22.3477 | 22.1993 | 22.0911 | 50.1588 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048692 | 0.88089 | 0.042892 | 18.0911 | 21.9828 | 3.4579 | 0.045088 | 0.2927 | 0.3649 | 0.2165 | 0.1083
2012-10-30 22:54 | 26.324 | 22.3096 | 22.3216 | 22.2174 | 22.0971 | 50.1623 | 0.006713 | 0.32627 | 0.00219 | 48.6024 | 0.048711 | 0.88124 | 0.042926 | 18.0911 | 22.0128 | 3.4582 | 0.045116 | 0.2968 | 0.3088 | 0.2046 | 0.0843
2012-10-30 22:54 | 26.3232 | 22.2887 | 22.3088 | 22.2326 | 22.1243 | 50.1588 | 0.006735 | 0.32622 | 0.002197 | 48.4373 | 0.048749 | 0.88184 | 0.042989 | 18.0895 | 21.996 | 3.4579 | 0.045186 | 0.2927 | 0.3128 | 0.2366 | 0.1283
2012-10-30 22:54 | 26.3299 | 22.2874 | 22.2834 | 22.1952 | 22.1471 | 50.145 | 0.00671 | 0.32636 | 0.00219 | 48.6371 | 0.048752 | 0.88213 | 0.043006 | 18.0943 | 21.9947 | 3.457 | 0.045196 | 0.2927 | 0.2887 | 0.2005 | 0.1524
2012-10-30 22:54 | 26.3284 | 22.33 | 22.322 | 22.2057 | 22.0894 | 50.1513 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048686 | 0.88093 | 0.042889 | 18.0942 | 22.0212 | 3.4574 | 0.045083 | 0.3088 | 0.3008 | 0.1845 | 0.0682
2012-10-30 22:54 | 26.324 | 22.2775 | 22.2895 | 22.2093 | 22.0489 | 50.1525 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.04868 | 0.8806 | 0.042868 | 18.0896 | 22.0128 | 3.4575 | 0.045064 | 0.2647 | 0.2767 | 0.1965 | 0.0361
2012-10-30 22:54 | 26.3262 | 22.3238 | 22.2917 | 22.1875 | 22.1153 | 50.1513 | 0.006725 | 0.32627 | 0.002194 | 48.5157 | 0.048717 | 0.88148 | 0.042943 | 18.094 | 22.0271 | 3.4574 | 0.045137 | 0.2967 | 0.2646 | 0.1604 | 0.0882
2012-10-30 22:54 | 26.3202 | 22.2977 | 22.3058 | 22.1895 | 22.0892 | 50.1571 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048699 | 0.88104 | 0.042906 | 18.0914 | 22.013 | 3.4578 | 0.045102 | 0.2847 | 0.2928 | 0.1765 | 0.0762
2012-10-30 22:54 | 26.3307 | 22.3082 | 22.3002 | 22.216 | 22.1078 | 50.1507 | 0.006725 | 0.32628 | 0.002194 | 48.5179 | 0.048717 | 0.88128 | 0.042933 | 18.0898 | 21.9834 | 3.4574 | 0.045127 | 0.3248 | 0.3168 | 0.2326 | 0.1244
2012-10-30 22:54 | 26.324 | 22.2935 | 22.3056 | 22.1853 | 22.1011 | 50.1484 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048705 | 0.88109 | 0.042914 | 18.0904 | 21.9888 | 3.4572 | 0.04511 | 0.3047 | 0.3168 | 0.1965 | 0.1123
2012-10-30 22:54 | 26.3202 | 22.2617 | 22.2857 | 22.1815 | 22.0933 | 50.1496 | 0.006741 | 0.32627 | 0.002199 | 48.4005 | 0.048695 | 0.88102 | 0.042901 | 18.0926 | 22.005 | 3.4573 | 0.0451 | 0.2567 | 0.2807 | 0.1765 | 0.0883
2012-10-30 22:54 | 26.3232 | 22.2727 | 22.2847 | 22.2005 | 22.1283 | 50.1577 | 0.006713 | 0.32633 | 0.002191 | 48.6111 | 0.048714 | 0.88124 | 0.042929 | 18.09 | 22.0361 | 3.4579 | 0.04512 | 0.2366 | 0.2486 | 0.1644 | 0.0922
2012-10-30 22:54 | 26.3232 | 22.2767 | 22.3408 | 22.2166 | 22.1003 | 50.1525 | 0.006725 | 0.32624 | 0.002194 | 48.5115 | 0.048743 | 0.88173 | 0.042978 | 18.0893 | 22.0441 | 3.4575 | 0.045172 | 0.2326 | 0.2967 | 0.1725 | 0.0562
2012-10-30 22:54 | 26.3165 | 22.278 | 22.3221 | 22.2058 | 22.0855 | 50.1513 | 0.006738 | 0.32634 | 0.002199 | 48.4329 | 0.048711 | 0.88132 | 0.04293 | 18.0929 | 22.0013 | 3.4574 | 0.045129 | 0.2767 | 0.3208 | 0.2045 | 0.0842
2012-10-30 22:54 | 26.3307 | 22.3243 | 22.3363 | 22.212 | 22.1158 | 50.145 | 0.006722 | 0.32631 | 0.002193 | 48.5439 | 0.048673 | 0.88065 | 0.042864 | 18.0931 | 22.0115 | 3.457 | 0.045057 | 0.3128 | 0.3248 | 0.2005 | 0.1043
2012-10-30 22:54 | 26.3299 | 22.3195 | 22.3275 | 22.2112 | 22.119 | 50.1588 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.048667 | 0.88067 | 0.04286 | 18.0959 | 22.0588 | 3.4579 | 0.045051 | 0.2607 | 0.2687 | 0.1524 | 0.0602
2012-10-30 22:54 | 26.318 | 22.3036 | 22.3236 | 22.2354 | 22.1272 | 50.1484 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.048762 | 0.88193 | 0.043005 | 18.0864 | 22.0229 | 3.4572 | 0.045196 | 0.2807 | 0.3007 | 0.2125 | 0.1043
2012-10-30 22:54 | 26.3187 | 22.3003 | 22.3363 | 22.2121 | 22.0516 | 50.1409 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.048815 | 0.883 | 0.043104 | 18.0887 | 22.0155 | 3.4567 | 0.045302 | 0.2848 | 0.3208 | 0.1966 | 0.0361
2012-10-30 22:54 | 26.3202 | 22.3018 | 22.3178 | 22.1975 | 22.0892 | 50.1496 | 0.006732 | 0.32621 | 0.002196 | 48.4568 | 0.048711 | 0.88134 | 0.042931 | 18.0932 | 22.0371 | 3.4573 | 0.045127 | 0.2647 | 0.2807 | 0.1604 | 0.0521
2012-10-30 22:54 | 26.3202 | 22.2977 | 22.3298 | 22.1534 | 22.0692 | 50.1548 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048733 | 0.88143 | 0.042954 | 18.0868 | 22.0531 | 3.4577 | 0.04515 | 0.2446 | 0.2767 | 0.1003 | 0.0161
2012-10-30 22:55 | 26.3247 | 22.2822 | 22.3103 | 22.186 | 22.0536 | 50.1536 | 0.006716 | 0.32633 | 0.002192 | 48.5893 | 0.048714 | 0.88122 | 0.042928 | 18.0897 | 22.0456 | 3.4576 | 0.04512 | 0.2366 | 0.2647 | 0.1404 | 0.008
2012-10-30 22:55 | 26.3232 | 22.2927 | 22.3128 | 22.1765 | 22.0882 | 50.1473 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048708 | 0.88128 | 0.042925 | 18.0931 | 22.02 | 3.4572 | 0.045123 | 0.2727 | 0.2928 | 0.1565 | 0.0682
2012-10-30 22:55 | 26.3225 | 22.288 | 22.3241 | 22.1998 | 22.0956 | 50.1525 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048768 | 0.88194 | 0.043011 | 18.0845 | 22.0434 | 3.4575 | 0.045206 | 0.2446 | 0.2807 | 0.1564 | 0.0522
2012-10-30 22:55 | 26.3269 | 22.3004 | 22.3085 | 22.1882 | 22.108 | 50.1559 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.048743 | 0.88183 | 0.042983 | 18.0914 | 22.0358 | 3.4577 | 0.045182 | 0.2646 | 0.2727 | 0.1524 | 0.0722
2012-10-30 22:55 | 26.3292 | 22.2907 | 22.3188 | 22.2065 | 22.0902 | 50.1484 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048762 | 0.88197 | 0.043007 | 18.0873 | 22.014 | 3.4572 | 0.045202 | 0.2767 | 0.3048 | 0.1925 | 0.0762
2012-10-30 22:55 | 26.3262 | 22.2757 | 22.3358 | 22.2075 | 22.0993 | 50.1461 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.048749 | 0.88186 | 0.04299 | 18.0898 | 22.015 | 3.4571 | 0.045185 | 0.2607 | 0.3208 | 0.1925 | 0.0843
2012-10-30 22:55 | 26.3172 | 22.2466 | 22.3068 | 22.1985 | 22.1023 | 50.1363 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048683 | 0.8807 | 0.042875 | 18.0905 | 22.0301 | 3.4564 | 0.045069 | 0.2165 | 0.2767 | 0.1684 | 0.0722
2012-10-30 22:55 | 26.3195 | 22.285 | 22.3011 | 22.2129 | 22.0885 | 50.1403 | 0.006716 | 0.3263 | 0.002191 | 48.585 | 0.048692 | 0.88106 | 0.042901 | 18.0946 | 21.9762 | 3.4567 | 0.045092 | 0.3088 | 0.3249 | 0.2367 | 0.1123
2012-10-30 22:55 | 26.324 | 22.2615 | 22.2855 | 22.2053 | 22.061 | 50.1502 | 0.00671 | 0.32636 | 0.00219 | 48.6371 | 0.048837 | 0.88324 | 0.043135 | 18.0856 | 21.9928 | 3.4574 | 0.045325 | 0.2687 | 0.2927 | 0.2125 | 0.0682
2012-10-30 22:55 | 26.3202 | 22.2617 | 22.2937 | 22.2136 | 22.1173 | 50.149 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048714 | 0.88135 | 0.042934 | 18.0924 | 22.005 | 3.4573 | 0.04513 | 0.2567 | 0.2887 | 0.2086 | 0.1123
2012-10-30 22:55 | 26.3232 | 22.3168 | 22.2807 | 22.2085 | 22.0602 | 50.1455 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048705 | 0.88102 | 0.04291 | 18.0889 | 21.9559 | 3.457 | 0.045102 | 0.3609 | 0.3248 | 0.2526 | 0.1043
2012-10-30 22:55 | 26.3254 | 22.2589 | 22.311 | 22.2268 | 22.0583 | 50.1484 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.048686 | 0.88102 | 0.042893 | 18.096 | 22.0022 | 3.4572 | 0.045085 | 0.2567 | 0.3088 | 0.2246 | 0.0561
2012-10-30 22:55 | 26.3254 | 22.2789 | 22.2909 | 22.2067 | 22.0904 | 50.1484 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048695 | 0.88091 | 0.042896 | 18.0903 | 21.9902 | 3.4572 | 0.045092 | 0.2887 | 0.3007 | 0.2165 | 0.1002
2012-10-30 22:55 | 26.3262 | 22.2797 | 22.2717 | 22.1714 | 22.0832 | 50.1479 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048724 | 0.88144 | 0.042947 | 18.0904 | 22.015 | 3.4572 | 0.04514 | 0.2647 | 0.2567 | 0.1564 | 0.0682
2012-10-30 22:55 | 26.3292 | 22.2827 | 22.3508 | 22.1825 | 22.0742 | 50.1357 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.04868 | 0.88095 | 0.042885 | 18.0967 | 21.994 | 3.4564 | 0.045078 | 0.2887 | 0.3568 | 0.1885 | 0.0802
2012-10-30 22:55 | 26.3299 | 22.2674 | 22.3275 | 22.2072 | 22.0669 | 50.1334 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.048736 | 0.88174 | 0.042973 | 18.0922 | 22.0348 | 3.4562 | 0.045168 | 0.2326 | 0.2927 | 0.1724 | 0.0321
2012-10-30 22:55 | 26.3232 | 22.2526 | 22.2927 | 22.1644 | 22.0842 | 50.1357 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048721 | 0.88143 | 0.042944 | 18.0913 | 22.0281 | 3.4564 | 0.04514 | 0.2245 | 0.2646 | 0.1363 | 0.0561
2012-10-30 22:55 | 26.3232 | 22.2767 | 22.3128 | 22.1965 | 22.0842 | 50.1409 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048708 | 0.88141 | 0.042932 | 18.0958 | 22.016 | 3.4567 | 0.045128 | 0.2607 | 0.2968 | 0.1805 | 0.0682
2012-10-30 22:55 | 26.3187 | 22.2321 | 22.2962 | 22.18 | 22.0757 | 50.1334 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.048724 | 0.88131 | 0.042941 | 18.0878 | 22.0075 | 3.4562 | 0.045137 | 0.2246 | 0.2887 | 0.1725 | 0.0682
2012-10-30 22:55 | 26.3157 | 22.2652 | 22.2932 | 22.205 | 22.0888 | 50.1461 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.048828 | 0.88316 | 0.043123 | 18.0871 | 22.0005 | 3.4571 | 0.04532 | 0.2647 | 0.2927 | 0.2045 | 0.0883
2012-10-30 22:55 | 26.318 | 22.2675 | 22.2955 | 22.1953 | 22.071 | 50.1444 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.048743 | 0.88189 | 0.042986 | 18.0926 | 22.0309 | 3.457 | 0.045185 | 0.2366 | 0.2646 | 0.1644 | 0.0401
2012-10-30 22:55 | 26.3225 | 22.264 | 22.288 | 22.2038 | 22.1196 | 50.149 | 0.006713 | 0.32634 | 0.002191 | 48.6133 | 0.048733 | 0.88171 | 0.042969 | 18.0927 | 22.0394 | 3.4573 | 0.04516 | 0.2246 | 0.2486 | 0.1644 | 0.0802
2012-10-30 22:55 | 26.3232 | 22.2887 | 22.3448 | 22.2085 | 22.0802 | 50.134 | 0.006735 | 0.32625 | 0.002197 | 48.4416 | 0.048708 | 0.88127 | 0.042925 | 18.0928 | 22.0361 | 3.4562 | 0.045122 | 0.2526 | 0.3087 | 0.1724 | 0.0441
2012-10-30 22:55 | 26.3067 | 22.2722 | 22.3283 | 22.176 | 22.0396 | 50.145 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.048721 | 0.88128 | 0.042937 | 18.0883 | 22.0356 | 3.457 | 0.045133 | 0.2366 | 0.2927 | 0.1404 | 0.004
2012-10-30 22:55 | 26.3165 | 22.286 | 22.3462 | 22.2339 | 22.0855 | 50.1415 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.048689 | 0.88105 | 0.042897 | 18.0954 | 22.0174 | 3.4568 | 0.045094 | 0.2686 | 0.3288 | 0.2165 | 0.0681
2012-10-30 22:55 | 26.3142 | 22.2717 | 22.2877 | 22.2035 | 22.1033 | 50.1374 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048695 | 0.88115 | 0.042908 | 18.0953 | 22.011 | 3.4565 | 0.045101 | 0.2607 | 0.2767 | 0.1925 | 0.0923
2012-10-30 22:56 | 26.3082 | 22.2978 | 22.3058 | 22.2256 | 22.1053 | 50.1426 | 0.006722 | 0.32634 | 0.002194 | 48.5482 | 0.048746 | 0.88194 | 0.042991 | 18.0927 | 21.993 | 3.4568 | 0.045185 | 0.3048 | 0.3128 | 0.2326 | 0.1123
2012-10-30 22:56 | 26.3067 | 22.3283 | 22.2642 | 22.2121 | 22.0878 | 50.1346 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.048699 | 0.88122 | 0.042915 | 18.0953 | 21.9995 | 3.4563 | 0.045111 | 0.3288 | 0.2647 | 0.2126 | 0.0883
2012-10-30 22:56 | 26.3097 | 22.2712 | 22.3394 | 22.1509 | 22.0988 | 50.1421 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048755 | 0.8821 | 0.043007 | 18.0926 | 22.0186 | 3.4568 | 0.045203 | 0.2526 | 0.3208 | 0.1323 | 0.0802
2012-10-30 22:56 | 26.3045 | 22.3021 | 22.2981 | 22.1898 | 22.0575 | 50.1409 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.048664 | 0.88063 | 0.042855 | 18.0961 | 22.0094 | 3.4567 | 0.045048 | 0.2927 | 0.2887 | 0.1804 | 0.0481
2012-10-30 22:56 | 26.3008 | 22.2703 | 22.3385 | 22.2222 | 22.0899 | 50.1542 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.048708 | 0.88125 | 0.042924 | 18.0925 | 22.0137 | 3.4576 | 0.045119 | 0.2566 | 0.3248 | 0.2085 | 0.0762
2012-10-30 22:56 | 26.303 | 22.2605 | 22.3327 | 22.2084 | 22.0761 | 50.1461 | 0.006725 | 0.32625 | 0.002194 | 48.5136 | 0.048721 | 0.88144 | 0.042945 | 18.0916 | 22.0079 | 3.4571 | 0.045139 | 0.2526 | 0.3248 | 0.2005 | 0.0682
2012-10-30 22:56 | 26.2985 | 22.268 | 22.3282 | 22.1718 | 22.0996 | 50.1548 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.048699 | 0.88114 | 0.04291 | 18.0935 | 22.0114 | 3.4577 | 0.045106 | 0.2566 | 0.3168 | 0.1604 | 0.0882
2012-10-30 22:56 | 26.2978 | 22.2713 | 22.3515 | 22.1631 | 22.111 | 50.1513 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.048717 | 0.88144 | 0.042941 | 18.093 | 21.9987 | 3.4574 | 0.045138 | 0.2726 | 0.3528 | 0.1644 | 0.1123
2012-10-30 22:56 | 26.3022 | 22.2196 | 22.3278 | 22.2116 | 22.0873 | 50.1346 | 0.006716 | 0.32634 | 0.002192 | 48.5916 | 0.048686 | 0.88093 | 0.042889 | 18.0942 | 22.0071 | 3.4563 | 0.045081 | 0.2125 | 0.3207 | 0.2045 | 0.0802
2012-10-30 22:56 | 26.2985 | 22.272 | 22.2881 | 22.2119 | 22.1157 | 50.1369 | 0.006716 | 0.32631 | 0.002192 | 48.5873 | 0.048708 | 0.88134 | 0.042928 | 18.0943 | 22.0034 | 3.4564 | 0.04512 | 0.2686 | 0.2847 | 0.2085 | 0.1123
2012-10-30 22:56 | 26.2978 | 22.2232 | 22.2914 | 22.2272 | 22.0989 | 50.1432 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.048699 | 0.88132 | 0.04292 | 18.0974 | 21.9666 | 3.4569 | 0.045118 | 0.2566 | 0.3248 | 0.2606 | 0.1323
2012-10-30 22:56 | 26.2978 | 22.3114 | 22.3074 | 22.1952 | 22.0709 | 50.1432 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.048692 | 0.88093 | 0.042894 | 18.092 | 21.9786 | 3.4569 | 0.045089 | 0.3328 | 0.3288 | 0.2166 | 0.0923
2012-10-30 22:56 | 26.2805 | 22.234 | 22.3182 | 22.1859 | 22.0856 | 50.1351 | 0.006719 | 0.32636 | 0.002193 | 48.572 | 0.048673 | 0.88069 | 0.042866 | 18.094 | 21.9814 | 3.4563 | 0.045059 | 0.2526 | 0.3368 | 0.2045 | 0.1042
2012-10-30 22:56 | 26.285 | 22.2385 | 22.3347 | 22.1503 | 22.0821 | 50.1409 | 0.006729 | 0.32628 | 0.002196 | 48.4891 | 0.048755 | 0.88194 | 0.042999 | 18.0893 | 21.9939 | 3.4567 | 0.045195 | 0.2446 | 0.3408 | 0.1564 | 0.0882
2012-10-30 22:56 | 26.285 | 22.2625 | 22.2866 | 22.2064 | 22.0942 | 50.1322 | 0.006716 | 0.32641 | 0.002192 | 48.6023 | 0.04874 | 0.88184 | 0.042981 | 18.0928 | 21.9859 | 3.4561 | 0.045173 | 0.2766 | 0.3007 | 0.2205 | 0.1083
2012-10-30 22:56 | 26.2918 | 22.2894 | 22.2854 | 22.1852 | 22.0849 | 50.1369 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.048667 | 0.88076 | 0.042864 | 18.0977 | 22.0047 | 3.4564 | 0.045059 | 0.2847 | 0.2807 | 0.1805 | 0.0802
2012-10-30 22:56 | 26.2835 | 22.2771 | 22.3011 | 22.2129 | 22.1007 | 50.1421 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.04873 | 0.88158 | 0.04296 | 18.0912 | 21.9924 | 3.4568 | 0.045155 | 0.2847 | 0.3087 | 0.2205 | 0.1083
2012-10-30 22:56 | 26.2813 | 22.2909 | 22.2829 | 22.2067 | 22.0624 | 50.1294 | 0.006713 | 0.32637 | 0.002191 | 48.6176 | 0.048692 | 0.88075 | 0.042885 | 18.0881 | 21.9902 | 3.4559 | 0.045076 | 0.3007 | 0.2927 | 0.2165 | 0.0722
2012-10-30 22:56 | 26.2798 | 22.2894 | 22.3415 | 22.1972 | 22.0689 | 50.1409 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.048724 | 0.88145 | 0.042948 | 18.0908 | 22.0047 | 3.4567 | 0.045145 | 0.2847 | 0.3368 | 0.1925 | 0.0642
2012-10-30 22:56 | 26.2776 | 22.2672 | 22.3233 | 22.2111 | 22.1188 | 50.1473 | 0.006722 | 0.32636 | 0.002194 | 48.5503 | 0.048708 | 0.88116 | 0.04292 | 18.0908 | 21.9744 | 3.4572 | 0.045114 | 0.2928 | 0.3489 | 0.2367 | 0.1444
2012-10-30 22:56 | 26.2813 | 22.2829 | 22.327 | 22.2188 | 22.1025 | 50.1421 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.048692 | 0.88101 | 0.042898 | 18.0934 | 21.9822 | 3.4568 | 0.045092 | 0.3007 | 0.3448 | 0.2366 | 0.1203
2012-10-30 22:56 | 26.2805 | 22.3342 | 22.3061 | 22.1698 | 22.1017 | 50.1374 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.048714 | 0.88115 | 0.042924 | 18.0882 | 22.0054 | 3.4565 | 0.045117 | 0.3288 | 0.3007 | 0.1644 | 0.0963
2012-10-30 22:56 | 26.282 | 22.2876 | 22.2876 | 22.2034 | 22.0831 | 50.1398 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.04874 | 0.88177 | 0.042978 | 18.0913 | 22.0069 | 3.4566 | 0.045175 | 0.2807 | 0.2807 | 0.1965 | 0.0762
2012-10-30 22:56 | 26.2768 | 22.2664 | 22.2584 | 22.1942 | 22.0659 | 50.1357 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.04867 | 0.88067 | 0.042862 | 18.0948 | 21.9897 | 3.4564 | 0.045054 | 0.2767 | 0.2687 | 0.2045 | 0.0762
2012-10-30 22:56 | 26.2813 | 22.2949 | 22.3069 | 22.1747 | 22.0945 | 50.1357 | 0.006725 | 0.32625 | 0.002194 | 48.5136 | 0.048686 | 0.88101 | 0.042893 | 18.0957 | 21.9741 | 3.4564 | 0.045087 | 0.3208 | 0.3328 | 0.2006 | 0.1204
2012-10-30 22:57 | 26.2776 | 22.2351 | 22.3153 | 22.183 | 22.0507 | 50.1311 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.048689 | 0.88102 | 0.042896 | 18.0948 | 21.9865 | 3.456 | 0.04509 | 0.2486 | 0.3288 | 0.1965 | 0.0642
2012-10-30 22:57 | 26.2865 | 22.256 | 22.2801 | 22.1879 | 22.0916 | 50.134 | 0.006725 | 0.32636 | 0.002195 | 48.5286 | 0.11092 | 1.8737 | 0.20783 | 16.8921 | 21.9553 | 3.4562 | 0.21002 | 0.3007 | 0.3248 | 0.2326 | 0.1363
2012-10-30 22:57 | 26.2865 | 22.2961 | 22.2721 | 22.1718 | 22.0716 | 50.1565 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.14215 | 2.3718 | 0.33714 | 16.6858 | 21.9874 | 3.4578 | 0.33934 | 0.3087 | 0.2847 | 0.1844 | 0.0842
2012-10-30 22:57 | 26.2731 | 22.2827 | 22.2707 | 22.1745 | 22.0823 | 50.1507 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.14213 | 2.3716 | 0.33708 | 16.6862 | 21.9499 | 3.4574 | 0.33927 | 0.3328 | 0.3208 | 0.2246 | 0.1324
2012-10-30 22:57 | 26.2805 | 22.2821 | 22.3142 | 22.1698 | 22.0415 | 50.1479 | 0.006719 | 0.32634 | 0.002193 | 48.5699 | 0.14212 | 2.3713 | 0.337 | 16.6852 | 22.0094 | 3.4572 | 0.33919 | 0.2727 | 0.3048 | 0.1604 | 0.0321
2012-10-30 22:57 | 26.2828 | 22.3445 | 22.2884 | 22.1681 | 22.0679 | 50.1513 | 0.006741 | 0.32646 | 0.002201 | 48.4284 | 0.14207 | 2.3706 | 0.33679 | 16.6853 | 22.0117 | 3.4574 | 0.33899 | 0.3328 | 0.2767 | 0.1564 | 0.0562
2012-10-30 22:57 | 26.2805 | 22.3021 | 22.2741 | 22.1779 | 22.1217 | 50.1548 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.14211 | 2.3712 | 0.33699 | 16.6854 | 21.9934 | 3.4577 | 0.33919 | 0.3087 | 0.2807 | 0.1845 | 0.1283
2012-10-30 22:57 | 26.2895 | 22.3031 | 22.2871 | 22.2069 | 22.1027 | 50.1588 | 0.006744 | 0.32633 | 0.002201 | 48.3876 | 0.14215 | 2.3719 | 0.33716 | 16.6859 | 21.9703 | 3.4579 | 0.33936 | 0.3328 | 0.3168 | 0.2366 | 0.1324
2012-10-30 22:57 | 26.2828 | 22.3205 | 22.3084 | 22.1802 | 22.096 | 50.1479 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.14217 | 2.3719 | 0.33721 | 16.684 | 22.0278 | 3.4572 | 0.33941 | 0.2927 | 0.2806 | 0.1524 | 0.0682
2012-10-30 22:57 | 26.2948 | 22.3084 | 22.3084 | 22.2002 | 22.12 | 50.149 | 0.006719 | 0.32641 | 0.002193 | 48.5806 | 0.14216 | 2.372 | 0.3372 | 16.6855 | 22.0037 | 3.4573 | 0.33939 | 0.3047 | 0.3047 | 0.1965 | 0.1163
2012-10-30 22:57 | 26.279 | 22.3207 | 22.3207 | 22.1884 | 22.0641 | 50.164 | 0.006735 | 0.32641 | 0.002198 | 48.4652 | 0.14219 | 2.3724 | 0.33733 | 16.6845 | 22.0079 | 3.4583 | 0.33953 | 0.3128 | 0.3128 | 0.1805 | 0.0562
2012-10-30 22:57 | 26.2895 | 22.3071 | 22.3352 | 22.223 | 22.1147 | 50.1374 | 0.006747 | 0.32636 | 0.002202 | 48.3704 | 0.14223 | 2.373 | 0.3375 | 16.6841 | 21.9904 | 3.4565 | 0.3397 | 0.3167 | 0.3448 | 0.2326 | 0.1243
2012-10-30 22:57 | 26.2873 | 22.3089 | 22.3009 | 22.1726 | 22.0884 | 50.1467 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.14221 | 2.3726 | 0.3374 | 16.6845 | 21.9922 | 3.4571 | 0.3396 | 0.3167 | 0.3087 | 0.1804 | 0.0962
2012-10-30 22:57 | 26.291 | 22.3447 | 22.3126 | 22.2325 | 22.1002 | 50.16 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.14216 | 2.3718 | 0.33716 | 16.6843 | 21.9838 | 3.458 | 0.33936 | 0.3609 | 0.3288 | 0.2487 | 0.1164
2012-10-30 22:57 | 26.2873 | 22.329 | 22.3009 | 22.1686 | 22.0684 | 50.1617 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.14215 | 2.3719 | 0.33716 | 16.6858 | 21.9601 | 3.4581 | 0.33936 | 0.3689 | 0.3408 | 0.2085 | 0.1083
2012-10-30 22:57 | 26.2828 | 22.3686 | 22.2764 | 22.2082 | 22.0799 | 50.1548 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.14211 | 2.3712 | 0.33696 | 16.6861 | 21.9516 | 3.4577 | 0.33915 | 0.417 | 0.3248 | 0.2566 | 0.1283
2012-10-30 22:57 | 26.2865 | 22.3202 | 22.3282 | 22.2119 | 22.0916 | 50.1536 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.1421 | 2.3712 | 0.33694 | 16.6873 | 21.9352 | 3.4576 | 0.33913 | 0.385 | 0.393 | 0.2767 | 0.1564
2012-10-30 22:57 | 26.2873 | 22.3089 | 22.3691 | 22.2328 | 22.0804 | 50.1606 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.14211 | 2.3712 | 0.33697 | 16.6866 | 21.94 | 3.4581 | 0.33916 | 0.3689 | 0.4291 | 0.2928 | 0.1404
2012-10-30 22:57 | 26.282 | 22.3517 | 22.3277 | 22.2475 | 22.0831 | 50.1715 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.14213 | 2.3714 | 0.33706 | 16.6845 | 21.9668 | 3.4588 | 0.33926 | 0.3849 | 0.3609 | 0.2807 | 0.1163
2012-10-30 22:57 | 26.2835 | 22.3412 | 22.3212 | 22.229 | 22.1127 | 50.1554 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.14216 | 2.3721 | 0.33721 | 16.6862 | 21.9483 | 3.4577 | 0.33941 | 0.3929 | 0.3729 | 0.2807 | 0.1644
2012-10-30 22:57 | 26.285 | 22.3547 | 22.3467 | 22.2465 | 22.1022 | 50.1525 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.14218 | 2.3724 | 0.3373 | 16.6861 | 22.0179 | 3.4575 | 0.3395 | 0.3368 | 0.3288 | 0.2286 | 0.0843
2012-10-30 22:57 | 26.2843 | 22.354 | 22.334 | 22.2418 | 22.0654 | 50.1583 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.14216 | 2.3719 | 0.33718 | 16.6854 | 21.9771 | 3.4579 | 0.33938 | 0.3769 | 0.3569 | 0.2647 | 0.0883
2012-10-30 22:57 | 26.2925 | 22.3743 | 22.3422 | 22.254 | 22.0976 | 50.1571 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.14216 | 2.3721 | 0.3372 | 16.6863 | 22.0174 | 3.4578 | 0.3394 | 0.3569 | 0.3248 | 0.2366 | 0.0802
2012-10-30 22:57 | 26.2873 | 22.349 | 22.329 | 22.3009 | 22.1045 | 50.1588 | 0.006719 | 0.32627 | 0.002192 | 48.559 | 0.14218 | 2.3723 | 0.33729 | 16.6854 | 21.9841 | 3.4579 | 0.33948 | 0.3649 | 0.3449 | 0.3168 | 0.1204
2012-10-30 22:57 | 26.2858 | 22.4036 | 22.3636 | 22.2714 | 22.099 | 50.1531 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.14211 | 2.371 | 0.33693 | 16.685 | 21.9826 | 3.4576 | 0.33913 | 0.421 | 0.381 | 0.2888 | 0.1164
2012-10-30 22:57 | 26.285 | 22.3948 | 22.3267 | 22.3187 | 22.1302 | 50.1715 | 0.006725 | 0.32634 | 0.002195 | 48.5265 | 0.1421 | 2.3709 | 0.33691 | 16.6852 | 21.9417 | 3.4588 | 0.3391 | 0.4531 | 0.385 | 0.377 | 0.1885
2012-10-30 22:58 | 26.279 | 22.3968 | 22.3527 | 22.2886 | 22.1202 | 50.1513 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.14208 | 2.3707 | 0.33684 | 16.6856 | 21.9919 | 3.4574 | 0.33904 | 0.4049 | 0.3608 | 0.2967 | 0.1283
2012-10-30 22:58 | 26.2925 | 22.4304 | 22.3462 | 22.2861 | 22.1137 | 50.1525 | 0.006716 | 0.3263 | 0.002191 | 48.585 | 0.14211 | 2.371 | 0.33695 | 16.6843 | 21.9853 | 3.4575 | 0.33914 | 0.4451 | 0.3609 | 0.3008 | 0.1284
2012-10-30 22:58 | 26.2843 | 22.3901 | 22.354 | 22.2739 | 22.1255 | 50.1692 | 0.006716 | 0.3263 | 0.002191 | 48.585 | 0.14211 | 2.3712 | 0.33698 | 16.6851 | 21.9852 | 3.4587 | 0.33918 | 0.4049 | 0.3688 | 0.2887 | 0.1403
2012-10-30 22:58 | 26.291 | 22.4329 | 22.3367 | 22.3287 | 22.1242 | 50.1536 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.14213 | 2.3716 | 0.33708 | 16.6863 | 21.9798 | 3.4576 | 0.33927 | 0.4531 | 0.3569 | 0.3489 | 0.1444
2012-10-30 22:58 | 26.288 | 22.4339 | 22.3497 | 22.2776 | 22.1212 | 50.1548 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.14218 | 2.3724 | 0.3373 | 16.6859 | 21.9688 | 3.4577 | 0.33949 | 0.4651 | 0.3809 | 0.3088 | 0.1524
2012-10-30 22:58 | 26.285 | 22.4469 | 22.3307 | 22.2906 | 22.1102 | 50.1583 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.14217 | 2.3722 | 0.33726 | 16.6858 | 21.9698 | 3.4579 | 0.33946 | 0.4771 | 0.3609 | 0.3208 | 0.1404
2012-10-30 22:58 | 26.2865 | 22.4083 | 22.3322 | 22.3081 | 22.1317 | 50.1739 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.14216 | 2.3721 | 0.33723 | 16.6862 | 21.9753 | 3.459 | 0.33943 | 0.433 | 0.3569 | 0.3328 | 0.1564
2012-10-30 22:58 | 26.2948 | 22.4447 | 22.3525 | 22.3605 | 22.116 | 50.1611 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.14211 | 2.3714 | 0.33701 | 16.6863 | 21.9435 | 3.4581 | 0.3392 | 0.5012 | 0.409 | 0.417 | 0.1725
2012-10-30 22:58 | 26.294 | 22.4719 | 22.3437 | 22.3277 | 22.1392 | 50.1629 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.14215 | 2.3718 | 0.33714 | 16.685 | 21.9828 | 3.4582 | 0.33934 | 0.4891 | 0.3609 | 0.3449 | 0.1564
2012-10-30 22:58 | 26.2918 | 22.4337 | 22.3215 | 22.3094 | 22.133 | 50.1583 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.1421 | 2.3709 | 0.33689 | 16.685 | 21.9726 | 3.4579 | 0.33909 | 0.4611 | 0.3489 | 0.3368 | 0.1604
2012-10-30 22:58 | 26.2888 | 22.4507 | 22.3706 | 22.3305 | 22.118 | 50.1635 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.1421 | 2.3708 | 0.33688 | 16.6844 | 21.9656 | 3.4583 | 0.33908 | 0.4851 | 0.405 | 0.3649 | 0.1524
2012-10-30 22:58 | 26.2925 | 22.4063 | 22.3262 | 22.3141 | 22.1297 | 50.1617 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.14213 | 2.3715 | 0.33705 | 16.6856 | 21.9613 | 3.4581 | 0.33924 | 0.445 | 0.3649 | 0.3528 | 0.1684
2012-10-30 22:58 | 26.3022 | 22.4521 | 22.3679 | 22.3399 | 22.1314 | 50.1669 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.14209 | 2.3707 | 0.33686 | 16.6839 | 22.0231 | 3.4585 | 0.33906 | 0.429 | 0.3448 | 0.3168 | 0.1083
2012-10-30 22:58 | 26.288 | 22.4619 | 22.3577 | 22.3337 | 22.1172 | 50.164 | 0.006719 | 0.32637 | 0.002193 | 48.5742 | 0.14205 | 2.3703 | 0.33671 | 16.6864 | 21.9768 | 3.4583 | 0.3389 | 0.4851 | 0.3809 | 0.3569 | 0.1404
2012-10-30 22:58 | 26.2978 | 22.4477 | 22.3715 | 22.3916 | 22.1912 | 50.1773 | 0.006729 | 0.32627 | 0.002195 | 48.4868 | 0.14212 | 2.3713 | 0.33701 | 16.6845 | 21.9425 | 3.4592 | 0.33921 | 0.5052 | 0.429 | 0.4491 | 0.2487
2012-10-30 22:58 | 26.2933 | 22.4312 | 22.343 | 22.323 | 22.1025 | 50.1715 | 0.006735 | 0.32627 | 0.002197 | 48.4437 | 0.14215 | 2.3714 | 0.33709 | 16.6831 | 22.0062 | 3.4588 | 0.33929 | 0.425 | 0.3368 | 0.3168 | 0.0963
2012-10-30 22:58 | 26.291 | 22.4168 | 22.3688 | 22.3487 | 22.1242 | 50.1635 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.14214 | 2.3713 | 0.33705 | 16.6834 | 21.9718 | 3.4583 | 0.33925 | 0.445 | 0.397 | 0.3769 | 0.1524
2012-10-30 22:58 | 26.2948 | 22.4807 | 22.3685 | 22.3645 | 22.1521 | 50.171 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.14213 | 2.3715 | 0.33707 | 16.6847 | 21.9836 | 3.4588 | 0.33926 | 0.4971 | 0.3849 | 0.3809 | 0.1685
2012-10-30 22:58 | 26.291 | 22.4369 | 22.3888 | 22.3407 | 22.1362 | 50.1698 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.14211 | 2.3712 | 0.33697 | 16.6853 | 21.9638 | 3.4587 | 0.33917 | 0.4731 | 0.425 | 0.3769 | 0.1724
2012-10-30 22:58 | 26.2925 | 22.4704 | 22.3783 | 22.3903 | 22.1778 | 50.1715 | 0.006722 | 0.32644 | 0.002194 | 48.5632 | 0.14214 | 2.3718 | 0.33712 | 16.6868 | 21.9893 | 3.4588 | 0.33932 | 0.4811 | 0.389 | 0.401 | 0.1885
2012-10-30 22:58 | 26.2948 | 22.4647 | 22.3806 | 22.3766 | 22.1441 | 50.1652 | 0.006738 | 0.32633 | 0.002199 | 48.4307 | 0.14215 | 2.3718 | 0.33716 | 16.6847 | 21.9716 | 3.4584 | 0.33936 | 0.4931 | 0.409 | 0.405 | 0.1725
2012-10-30 22:58 | 26.3008 | 22.5028 | 22.3545 | 22.3745 | 22.1501 | 50.1756 | 0.006732 | 0.32643 | 0.002198 | 48.4889 | 0.14212 | 2.3714 | 0.33703 | 16.6861 | 21.9896 | 3.4591 | 0.33923 | 0.5132 | 0.3649 | 0.3849 | 0.1605
2012-10-30 22:58 | 26.2873 | 22.4773 | 22.3891 | 22.3731 | 22.1165 | 50.1629 | 0.006719 | 0.32633 | 0.002193 | 48.5676 | 0.14215 | 2.3719 | 0.33716 | 16.6867 | 21.9801 | 3.4582 | 0.33936 | 0.4972 | 0.409 | 0.393 | 0.1364
2012-10-30 22:58 | 26.2963 | 22.5023 | 22.3901 | 22.3861 | 22.1856 | 50.1606 | 0.006725 | 0.32627 | 0.002194 | 48.5157 | 0.14214 | 2.3716 | 0.33709 | 16.6847 | 21.9931 | 3.4581 | 0.33929 | 0.5092 | 0.397 | 0.393 | 0.1925
2012-10-30 22:58 | 26.2895 | 22.4995 | 22.3793 | 22.3592 | 22.1388 | 50.1658 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.14214 | 2.3716 | 0.33709 | 16.6855 | 21.9663 | 3.4584 | 0.33929 | 0.5332 | 0.413 | 0.3929 | 0.1725
2012-10-30 22:58 | 26.2925 | 22.5025 | 22.3983 | 22.3983 | 22.1137 | 50.1588 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.14216 | 2.372 | 0.3372 | 16.6855 | 21.9813 | 3.4579 | 0.3394 | 0.5212 | 0.417 | 0.417 | 0.1324
2012-10-30 22:59 | 26.2985 | 22.4804 | 22.3923 | 22.3803 | 22.1237 | 50.1617 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.1422 | 2.3723 | 0.33733 | 16.6829 | 21.9994 | 3.4581 | 0.33953 | 0.481 | 0.3929 | 0.3809 | 0.1243
2012-10-30 22:59 | 26.303 | 22.533 | 22.3928 | 22.4008 | 22.1563 | 50.1721 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14215 | 2.3719 | 0.33718 | 16.6857 | 21.9517 | 3.4589 | 0.33937 | 0.5813 | 0.4411 | 0.4491 | 0.2046
2012-10-30 22:59 | 26.3 | 22.514 | 22.4218 | 22.4098 | 22.1733 | 50.1773 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.14216 | 2.3717 | 0.33717 | 16.6832 | 21.9648 | 3.4592 | 0.33937 | 0.5492 | 0.457 | 0.445 | 0.2085
2012-10-30 22:59 | 26.2925 | 22.5345 | 22.4143 | 22.3943 | 22.1738 | 50.1611 | 0.006722 | 0.32633 | 0.002194 | 48.546 | 0.14215 | 2.3717 | 0.33712 | 16.6846 | 21.9653 | 3.4581 | 0.33931 | 0.5692 | 0.449 | 0.429 | 0.2085
2012-10-30 22:59 | 26.3015 | 22.5315 | 22.3873 | 22.3953 | 22.1548 | 50.1658 | 0.006747 | 0.32633 | 0.002202 | 48.3661 | 0.14215 | 2.3716 | 0.33712 | 16.6839 | 21.9622 | 3.4584 | 0.33932 | 0.5693 | 0.4251 | 0.4331 | 0.1926
2012-10-30 22:59 | 26.2918 | 22.5178 | 22.3696 | 22.4056 | 22.133 | 50.1681 | 0.006719 | 0.32638 | 0.002193 | 48.5763 | 0.14215 | 2.3716 | 0.33714 | 16.6837 | 21.9927 | 3.4586 | 0.33933 | 0.5251 | 0.3769 | 0.4129 | 0.1403
2012-10-30 22:59 | 26.2858 | 22.5559 | 22.3756 | 22.4317 | 22.1351 | 50.1727 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.14212 | 2.3713 | 0.33702 | 16.6847 | 21.9706 | 3.4589 | 0.33921 | 0.5853 | 0.405 | 0.4611 | 0.1645
2012-10-30 22:59 | 26.306 | 22.52 | 22.4118 | 22.4399 | 22.1913 | 50.1663 | 0.006735 | 0.3263 | 0.002198 | 48.448 | 0.14212 | 2.3713 | 0.33701 | 16.6852 | 21.9467 | 3.4585 | 0.33921 | 0.5733 | 0.4651 | 0.4932 | 0.2446
2012-10-30 22:59 | 26.2993 | 22.5453 | 22.3851 | 22.4091 | 22.1927 | 50.1594 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14219 | 2.3722 | 0.33732 | 16.6833 | 21.948 | 3.458 | 0.33951 | 0.5973 | 0.4371 | 0.4611 | 0.2447
2012-10-30 22:59 | 26.3037 | 22.5818 | 22.4055 | 22.4456 | 22.169 | 50.1739 | 0.006744 | 0.3264 | 0.002201 | 48.3984 | 0.1421 | 2.3712 | 0.33694 | 16.6868 | 21.9604 | 3.459 | 0.33914 | 0.6214 | 0.4451 | 0.4852 | 0.2086
2012-10-30 22:59 | 26.2895 | 22.5355 | 22.3953 | 22.4514 | 22.1347 | 50.1814 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.14213 | 2.3712 | 0.33702 | 16.6835 | 21.9703 | 3.4595 | 0.33922 | 0.5652 | 0.425 | 0.4811 | 0.1644
2012-10-30 22:59 | 26.3 | 22.546 | 22.3737 | 22.4459 | 22.2014 | 50.1594 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.1421 | 2.3711 | 0.33694 | 16.686 | 22.0209 | 3.458 | 0.33914 | 0.5251 | 0.3528 | 0.425 | 0.1805
2012-10-30 22:59 | 26.2985 | 22.5926 | 22.4163 | 22.4804 | 22.1999 | 50.1733 | 0.006738 | 0.32637 | 0.002199 | 48.4372 | 0.14212 | 2.3715 | 0.33704 | 16.6866 | 21.9713 | 3.4589 | 0.33924 | 0.6213 | 0.445 | 0.5091 | 0.2286
2012-10-30 22:59 | 26.2933 | 22.5794 | 22.4111 | 22.4512 | 22.1867 | 50.1785 | 0.006732 | 0.32634 | 0.002197 | 48.4761 | 0.14212 | 2.3713 | 0.33701 | 16.6852 | 21.9661 | 3.4593 | 0.33921 | 0.6133 | 0.445 | 0.4851 | 0.2206
2012-10-30 22:59 | 26.282 | 22.5881 | 22.4038 | 22.476 | 22.1473 | 50.16 | 0.006722 | 0.32631 | 0.002193 | 48.5439 | 0.14212 | 2.3712 | 0.33699 | 16.6841 | 21.9748 | 3.458 | 0.33919 | 0.6133 | 0.429 | 0.5012 | 0.1725
2012-10-30 22:59 | 26.2873 | 22.5333 | 22.4292 | 22.4532 | 22.1606 | 50.1658 | 0.006722 | 0.32638 | 0.002194 | 48.5546 | 0.14212 | 2.3712 | 0.33699 | 16.6846 | 21.9601 | 3.4584 | 0.33918 | 0.5732 | 0.4691 | 0.4931 | 0.2005
2012-10-30 22:59 | 26.2955 | 22.5175 | 22.4734 | 22.4774 | 22.1728 | 50.1767 | 0.006747 | 0.32636 | 0.002202 | 48.3704 | 0.14213 | 2.3713 | 0.33704 | 16.6836 | 21.9442 | 3.4592 | 0.33925 | 0.5733 | 0.5292 | 0.5332 | 0.2286
2012-10-30 22:59 | 26.2903 | 22.5884 | 22.4202 | 22.4161 | 22.2037 | 50.1819 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.14215 | 2.3719 | 0.33716 | 16.6864 | 21.9631 | 3.4595 | 0.33935 | 0.6253 | 0.4571 | 0.453 | 0.2406
2012-10-30 22:59 | 26.2828 | 22.5369 | 22.4247 | 22.5128 | 22.1842 | 50.1739 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.14211 | 2.3712 | 0.33699 | 16.6853 | 21.9315 | 3.459 | 0.33918 | 0.6054 | 0.4932 | 0.5813 | 0.2527
2012-10-30 22:59 | 26.2888 | 22.615 | 22.4347 | 22.4467 | 22.1461 | 50.1779 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.14213 | 2.3714 | 0.33703 | 16.6848 | 21.9576 | 3.4593 | 0.33923 | 0.6574 | 0.4771 | 0.4891 | 0.1885
2012-10-30 22:59 | 26.285 | 22.5551 | 22.4269 | 22.479 | 22.1623 | 50.1606 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.14211 | 2.3712 | 0.33697 | 16.6853 | 21.9337 | 3.4581 | 0.33917 | 0.6214 | 0.4932 | 0.5453 | 0.2286
2012-10-30 22:59 | 26.2903 | 22.6044 | 22.3961 | 22.4843 | 22.1997 | 50.1663 | 0.006722 | 0.32628 | 0.002193 | 48.5396 | 0.14212 | 2.3711 | 0.33698 | 16.6843 | 21.9751 | 3.4585 | 0.33918 | 0.6293 | 0.421 | 0.5092 | 0.2246
2012-10-30 22:59 | 26.2888 | 22.6069 | 22.4387 | 22.4667 | 22.1942 | 50.1733 | 0.006719 | 0.32622 | 0.002192 | 48.5526 | 0.14207 | 2.3704 | 0.33676 | 16.6852 | 21.9576 | 3.4589 | 0.33896 | 0.6493 | 0.4811 | 0.5091 | 0.2366
2012-10-30 22:59 | 26.285 | 22.5591 | 22.4669 | 22.495 | 22.1663 | 50.1687 | 0.006732 | 0.3263 | 0.002197 | 48.4695 | 0.14199 | 2.3691 | 0.33638 | 16.6857 | 21.9738 | 3.4586 | 0.33858 | 0.5853 | 0.4931 | 0.5212 | 0.1925
2012-10-30 22:59 | 26.2925 | 22.6106 | 22.4344 | 22.5265 | 22.2179 | 50.1635 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.14211 | 2.3712 | 0.33697 | 16.6852 | 21.9653 | 3.4583 | 0.33917 | 0.6453 | 0.4691 | 0.5612 | 0.2526
2012-10-30 22:59 | 26.288 | 22.6021 | 22.4259 | 22.522 | 22.2094 | 50.1721 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.14208 | 2.3705 | 0.3368 | 16.6848 | 21.9728 | 3.4589 | 0.33899 | 0.6293 | 0.4531 | 0.5492 | 0.2366
2012-10-30 22:59 | 26.2903 | 22.6125 | 22.4562 | 22.4883 | 22.2117 | 50.1681 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14201 | 2.3696 | 0.33651 | 16.6861 | 21.9831 | 3.4586 | 0.33871 | 0.6294 | 0.4731 | 0.5052 | 0.2286
2012-10-30 22:59 | 26.2918 | 22.5899 | 22.4417 | 22.4938 | 22.1932 | 50.1629 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.14201 | 2.3695 | 0.3365 | 16.6856 | 21.9525 | 3.4582 | 0.3387 | 0.6374 | 0.4892 | 0.5413 | 0.2407
2012-10-30 23:00 | 26.2925 | 22.6066 | 22.4784 | 22.4985 | 22.2219 | 50.1669 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14204 | 2.3701 | 0.33664 | 16.6867 | 21.9853 | 3.4585 | 0.33884 | 0.6213 | 0.4931 | 0.5132 | 0.2366
2012-10-30 23:00 | 26.291 | 22.6212 | 22.4449 | 22.501 | 22.2164 | 50.1727 | 0.006719 | 0.32624 | 0.002192 | 48.5548 | 0.14208 | 2.3707 | 0.33683 | 16.6863 | 21.9517 | 3.4589 | 0.33902 | 0.6695 | 0.4932 | 0.5493 | 0.2647
2012-10-30 23:00 | 26.2895 | 22.6517 | 22.4394 | 22.5075 | 22.2149 | 50.1658 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.14209 | 2.3707 | 0.33684 | 16.6846 | 21.9743 | 3.4584 | 0.33904 | 0.6774 | 0.4651 | 0.5332 | 0.2406
2012-10-30 23:00 | 26.2903 | 22.5844 | 22.4562 | 22.5283 | 22.2238 | 50.1687 | 0.006747 | 0.32636 | 0.002202 | 48.3704 | 0.1421 | 2.3711 | 0.33694 | 16.6859 | 21.9591 | 3.4586 | 0.33914 | 0.6253 | 0.4971 | 0.5692 | 0.2647
2012-10-30 23:00 | 26.2828 | 22.5929 | 22.4527 | 22.5529 | 22.1681 | 50.1687 | 0.006738 | 0.32636 | 0.002199 | 48.435 | 0.14211 | 2.3713 | 0.337 | 16.6858 | 21.9796 | 3.4586 | 0.3392 | 0.6133 | 0.4731 | 0.5733 | 0.1885
2012-10-30 23:00 | 26.2738 | 22.5919 | 22.4638 | 22.5359 | 22.2113 | 50.1779 | 0.006725 | 0.32637 | 0.002195 | 48.5309 | 0.14215 | 2.372 | 0.33717 | 16.6864 | 21.9787 | 3.4593 | 0.33937 | 0.6132 | 0.4851 | 0.5572 | 0.2326
2012-10-30 23:00 | 26.2768 | 22.651 | 22.4788 | 22.5669 | 22.2543 | 50.1704 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.14218 | 2.3724 | 0.33731 | 16.6855 | 21.9817 | 3.4587 | 0.3395 | 0.6693 | 0.4971 | 0.5852 | 0.2726
2012-10-30 23:00 | 26.2746 | 22.6008 | 22.4926 | 22.5327 | 22.2441 | 50.1687 | 0.006716 | 0.32636 | 0.002192 | 48.5937 | 0.14224 | 2.373 | 0.33753 | 16.6836 | 21.9353 | 3.4586 | 0.33972 | 0.6655 | 0.5573 | 0.5974 | 0.3088
2012-10-30 23:00 | 26.2716 | 22.6218 | 22.4736 | 22.5377 | 22.2491 | 50.1687 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.14219 | 2.3725 | 0.33736 | 16.6853 | 21.9484 | 3.4586 | 0.33955 | 0.6734 | 0.5252 | 0.5893 | 0.3007
2012-10-30 23:00 | 26.2716 | 22.6138 | 22.4856 | 22.5577 | 22.201 | 50.1744 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.1422 | 2.3725 | 0.33737 | 16.6844 | 21.9684 | 3.459 | 0.33957 | 0.6454 | 0.5172 | 0.5893 | 0.2326
2012-10-30 23:00 | 26.2686 | 22.6268 | 22.4986 | 22.5186 | 22.2421 | 50.1629 | 0.006732 | 0.32641 | 0.002197 | 48.4868 | 0.14223 | 2.3732 | 0.33754 | 16.6855 | 22.0216 | 3.4582 | 0.33974 | 0.6052 | 0.477 | 0.497 | 0.2205
2012-10-30 23:00 | 26.2686 | 22.6388 | 22.5066 | 22.5667 | 22.2461 | 50.1733 | 0.006729 | 0.32638 | 0.002196 | 48.5041 | 0.14219 | 2.3725 | 0.33734 | 16.6851 | 21.9614 | 3.4589 | 0.33954 | 0.6774 | 0.5452 | 0.6053 | 0.2847
2012-10-30 23:00 | 26.2753 | 22.6295 | 22.4693 | 22.5614 | 22.2528 | 50.1681 | 0.006725 | 0.32638 | 0.002195 | 48.5329 | 0.1422 | 2.3726 | 0.33739 | 16.6845 | 21.9882 | 3.4586 | 0.33959 | 0.6413 | 0.4811 | 0.5732 | 0.2646
2012-10-30 23:00 | 26.2708 | 22.661 | 22.4608 | 22.5849 | 22.2604 | 50.1692 | 0.006747 | 0.32637 | 0.002202 | 48.3726 | 0.14219 | 2.3724 | 0.33733 | 16.6844 | 21.9676 | 3.4587 | 0.33953 | 0.6934 | 0.4932 | 0.6173 | 0.2928
2012-10-30 23:00 | 26.2618 | 22.6881 | 22.4437 | 22.5719 | 22.2434 | 50.1802 | 0.006729 | 0.32636 | 0.002196 | 48.4998 | 0.14214 | 2.3718 | 0.33713 | 16.6857 | 21.9546 | 3.4594 | 0.33932 | 0.7335 | 0.4891 | 0.6173 | 0.2888
2012-10-30 23:00 | 26.2708 | 22.661 | 22.4728 | 22.5609 | 22.2644 | 50.1727 | 0.006732 | 0.32636 | 0.002197 | 48.4782 | 0.14219 | 2.3723 | 0.33733 | 16.6838 | 21.9957 | 3.4589 | 0.33952 | 0.6653 | 0.4771 | 0.5652 | 0.2687
2012-10-30 23:00 | 26.2686 | 22.6548 | 22.4425 | 22.5827 | 22.2502 | 50.1617 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.14213 | 2.3716 | 0.33707 | 16.6859 | 21.9374 | 3.4581 | 0.33927 | 0.7174 | 0.5051 | 0.6453 | 0.3128
2012-10-30 23:00 | 26.2701 | 22.6723 | 22.4841 | 22.6083 | 22.2637 | 50.1785 | 0.006741 | 0.3263 | 0.0022 | 48.4048 | 0.1421 | 2.3709 | 0.3369 | 16.6847 | 21.991 | 3.4593 | 0.3391 | 0.6813 | 0.4931 | 0.6173 | 0.2727
2012-10-30 23:00 | 26.2663 | 22.6966 | 22.4923 | 22.5444 | 22.2479 | 50.1652 | 0.006735 | 0.32636 | 0.002198 | 48.4566 | 0.14207 | 2.3704 | 0.33677 | 16.6849 | 21.9752 | 3.4584 | 0.33897 | 0.7214 | 0.5171 | 0.5692 | 0.2727
2012-10-30 23:00 | 26.2678 | 22.6861 | 22.5098 | 22.5739 | 22.3055 | 50.1756 | 0.006741 | 0.32627 | 0.002199 | 48.4005 | 0.14206 | 2.3703 | 0.33671 | 16.6852 | 21.9526 | 3.4591 | 0.33891 | 0.7335 | 0.5572 | 0.6213 | 0.3529
2012-10-30 23:00 | 26.2716 | 22.6819 | 22.4936 | 22.5176 | 22.3373 | 50.1681 | 0.006729 | 0.32643 | 0.002197 | 48.5105 | 0.14205 | 2.3699 | 0.33664 | 16.6845 | 21.9644 | 3.4586 | 0.33884 | 0.7175 | 0.5292 | 0.5532 | 0.3729
2012-10-30 23:00 | 26.2746 | 22.6728 | 22.5006 | 22.5927 | 22.3002 | 50.1779 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.14212 | 2.371 | 0.33697 | 16.683 | 22.0075 | 3.4593 | 0.33916 | 0.6653 | 0.4931 | 0.5852 | 0.2927
2012-10-30 23:00 | 26.2618 | 22.6921 | 22.5239 | 22.608 | 22.2393 | 50.1715 | 0.006735 | 0.32634 | 0.002198 | 48.4545 | 0.14211 | 2.371 | 0.33695 | 16.6836 | 21.9787 | 3.4588 | 0.33915 | 0.7134 | 0.5452 | 0.6293 | 0.2606
2012-10-30 23:00 | 26.2656 | 22.6959 | 22.5597 | 22.5757 | 22.2391 | 50.1791 | 0.006722 | 0.32637 | 0.002194 | 48.5525 | 0.14213 | 2.3713 | 0.33702 | 16.6843 | 21.9905 | 3.4593 | 0.33922 | 0.7054 | 0.5692 | 0.5852 | 0.2486
2012-10-30 23:00 | 26.2678 | 22.7021 | 22.5419 | 22.626 | 22.2734 | 50.1727 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.14213 | 2.3715 | 0.33706 | 16.6856 | 22.0328 | 3.4589 | 0.33926 | 0.6693 | 0.5091 | 0.5932 | 0.2406
2012-10-30 23:00 | 26.2648 | 22.6991 | 22.5549 | 22.599 | 22.3065 | 50.1791 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.14222 | 2.3729 | 0.33746 | 16.6852 | 22.0298 | 3.4593 | 0.33966 | 0.6693 | 0.5251 | 0.5692 | 0.2767
2012-10-30 23:00 | 26.2641 | 22.6663 | 22.4941 | 22.6423 | 22.2737 | 50.171 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.14218 | 2.3722 | 0.33727 | 16.6847 | 22.0171 | 3.4588 | 0.33947 | 0.6492 | 0.477 | 0.6252 | 0.2566
2012-10-30 23:01 | 26.2611 | 22.6874 | 22.5111 | 22.6153 | 22.2547 | 50.1669 | 0.006732 | 0.32621 | 0.002196 | 48.4568 | 0.14217 | 2.3722 | 0.33725 | 16.6859 | 21.9459 | 3.4585 | 0.33945 | 0.7415 | 0.5652 | 0.6694 | 0.3088
2012-10-30 23:01 | 26.2671 | 22.6934 | 22.5091 | 22.5812 | 22.2807 | 50.1663 | 0.006713 | 0.32625 | 0.00219 | 48.6003 | 0.14216 | 2.3719 | 0.33719 | 16.6849 | 21.9599 | 3.4585 | 0.33938 | 0.7335 | 0.5492 | 0.6213 | 0.3208
2012-10-30 23:01 | 26.2671 | 22.7134 | 22.5091 | 22.6453 | 22.3008 | 50.164 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.14218 | 2.3721 | 0.33725 | 16.6837 | 21.98 | 3.4583 | 0.33945 | 0.7334 | 0.5291 | 0.6653 | 0.3208
2012-10-30 23:01 | 26.2603 | 22.6746 | 22.5264 | 22.5985 | 22.2659 | 50.1675 | 0.006713 | 0.32628 | 0.00219 | 48.6046 | 0.14214 | 2.3717 | 0.33712 | 16.6858 | 21.9652 | 3.4585 | 0.33931 | 0.7094 | 0.5612 | 0.6333 | 0.3007
2012-10-30 23:01 | 26.2663 | 22.6926 | 22.5724 | 22.6205 | 22.304 | 50.1721 | 0.006744 | 0.32637 | 0.002201 | 48.3941 | 0.14217 | 2.3722 | 0.33725 | 16.6852 | 21.9832 | 3.4589 | 0.33945 | 0.7094 | 0.5892 | 0.6373 | 0.3208
2012-10-30 23:01 | 26.2626 | 22.7049 | 22.5247 | 22.5968 | 22.2923 | 50.171 | 0.006725 | 0.32633 | 0.002195 | 48.5243 | 0.14214 | 2.3717 | 0.33711 | 16.6863 | 21.9554 | 3.4588 | 0.33931 | 0.7495 | 0.5693 | 0.6414 | 0.3369
2012-10-30 23:01 | 26.2581 | 22.6804 | 22.5362 | 22.6043 | 22.2878 | 50.1791 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.14215 | 2.3717 | 0.33714 | 16.6848 | 21.975 | 3.4593 | 0.33934 | 0.7054 | 0.5612 | 0.6293 | 0.3128
2012-10-30 23:01 | 26.2603 | 22.7346 | 22.5464 | 22.6585 | 22.302 | 50.1796 | 0.006729 | 0.32622 | 0.002195 | 48.4805 | 0.14222 | 2.3728 | 0.33746 | 16.6841 | 21.9732 | 3.4594 | 0.33965 | 0.7614 | 0.5732 | 0.6853 | 0.3288
2012-10-30 23:01 | 26.2648 | 22.7551 | 22.5389 | 22.627 | 22.2985 | 50.1796 | 0.006713 | 0.3263 | 0.00219 | 48.6067 | 0.14228 | 2.3736 | 0.3377 | 16.6831 | 21.9616 | 3.4594 | 0.33989 | 0.7935 | 0.5773 | 0.6654 | 0.3369
2012-10-30 23:01 | 26.2671 | 22.7014 | 22.5171 | 22.6093 | 22.3128 | 50.1669 | 0.006722 | 0.32624 | 0.002193 | 48.5332 | 0.14226 | 2.3733 | 0.33761 | 16.6831 | 21.972 | 3.4585 | 0.33981 | 0.7294 | 0.5451 | 0.6373 | 0.3408
2012-10-30 23:01 | 26.2633 | 22.7657 | 22.5414 | 22.6335 | 22.325 | 50.1611 | 0.006716 | 0.32628 | 0.002191 | 48.5829 | 0.14221 | 2.3727 | 0.33743 | 16.685 | 21.9401 | 3.4581 | 0.33962 | 0.8256 | 0.6013 | 0.6934 | 0.3849
2012-10-30 23:01 | 26.2671 | 22.7374 | 22.6013 | 22.6533 | 22.3489 | 50.1687 | 0.006729 | 0.32624 | 0.002195 | 48.4827 | 0.14219 | 2.3724 | 0.33732 | 16.6849 | 21.9479 | 3.4586 | 0.33952 | 0.7895 | 0.6534 | 0.7054 | 0.401
2012-10-30 23:01 | 26.2588 | 22.7371 | 22.5529 | 22.657 | 22.3165 | 50.171 | 0.006738 | 0.3263 | 0.002199 | 48.4264 | 0.14222 | 2.3728 | 0.33746 | 16.6842 | 21.9637 | 3.4588 | 0.33966 | 0.7734 | 0.5892 | 0.6933 | 0.3528
2012-10-30 23:01 | 26.2656 | 22.7359 | 22.5797 | 22.6759 | 22.3073 | 50.16 | 0.006735 | 0.32628 | 0.002198 | 48.4459 | 0.14221 | 2.3728 | 0.33745 | 16.6852 | 21.9825 | 3.458 | 0.33964 | 0.7534 | 0.5972 | 0.6934 | 0.3248
2012-10-30 23:01 | 26.2566 | 22.7269 | 22.5667 | 22.6829 | 22.3544 | 50.1635 | 0.006741 | 0.32631 | 0.0022 | 48.4071 | 0.14222 | 2.3728 | 0.33744 | 16.6844 | 21.9655 | 3.4583 | 0.33964 | 0.7614 | 0.6012 | 0.7174 | 0.3889
2012-10-30 23:01 | 26.2566 | 22.7389 | 22.5427 | 22.6949 | 22.3223 | 50.1629 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.14219 | 2.3723 | 0.33731 | 16.6845 | 21.9976 | 3.4582 | 0.33951 | 0.7413 | 0.5451 | 0.6973 | 0.3247
2012-10-30 23:01 | 26.2521 | 22.7224 | 22.5382 | 22.6463 | 22.2898 | 50.1733 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14217 | 2.372 | 0.33724 | 16.6835 | 21.969 | 3.4589 | 0.33943 | 0.7534 | 0.5692 | 0.6773 | 0.3208
2012-10-30 23:01 | 26.2521 | 22.7905 | 22.6063 | 22.6824 | 22.3419 | 50.1629 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.14216 | 2.3719 | 0.33718 | 16.6846 | 21.965 | 3.4582 | 0.33938 | 0.8255 | 0.6413 | 0.7174 | 0.3769
2012-10-30 23:01 | 26.2454 | 22.7437 | 22.5796 | 22.6637 | 22.2991 | 50.1715 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.14212 | 2.3712 | 0.33699 | 16.6842 | 21.9904 | 3.4588 | 0.33919 | 0.7533 | 0.5892 | 0.6733 | 0.3087
2012-10-30 23:01 | 26.2446 | 22.747 | 22.5948 | 22.6589 | 22.3304 | 50.1756 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14211 | 2.3714 | 0.337 | 16.6867 | 21.9655 | 3.4591 | 0.3392 | 0.7815 | 0.6293 | 0.6934 | 0.3649
2012-10-30 23:01 | 26.2514 | 22.7778 | 22.6176 | 22.6817 | 22.3171 | 50.1808 | 0.006729 | 0.3263 | 0.002196 | 48.4912 | 0.14208 | 2.3706 | 0.33681 | 16.6854 | 21.9523 | 3.4595 | 0.33901 | 0.8255 | 0.6653 | 0.7294 | 0.3648
2012-10-30 23:01 | 26.2491 | 22.7555 | 22.5993 | 22.7034 | 22.3629 | 50.175 | 0.006732 | 0.32633 | 0.002197 | 48.4739 | 0.14207 | 2.3705 | 0.33679 | 16.6852 | 21.9219 | 3.4591 | 0.33899 | 0.8336 | 0.6774 | 0.7815 | 0.441
2012-10-30 23:01 | 26.2446 | 22.751 | 22.5668 | 22.7029 | 22.3344 | 50.1727 | 0.006735 | 0.32637 | 0.002198 | 48.4588 | 0.14206 | 2.3703 | 0.33672 | 16.6853 | 21.9374 | 3.4589 | 0.33892 | 0.8136 | 0.6294 | 0.7655 | 0.397
2012-10-30 23:01 | 26.2349 | 22.7893 | 22.5891 | 22.7132 | 22.3688 | 50.1669 | 0.006738 | 0.32631 | 0.002199 | 48.4286 | 0.14206 | 2.3703 | 0.33672 | 16.6853 | 21.9438 | 3.4585 | 0.33892 | 0.8455 | 0.6453 | 0.7694 | 0.425
2012-10-30 23:01 | 26.2304 | 22.7768 | 22.5886 | 22.6647 | 22.3122 | 50.1652 | 0.006732 | 0.32625 | 0.002196 | 48.4632 | 0.14206 | 2.3701 | 0.33669 | 16.6841 | 21.9513 | 3.4584 | 0.33889 | 0.8255 | 0.6373 | 0.7134 | 0.3609
2012-10-30 23:01 | 26.2319 | 22.7823 | 22.5741 | 22.7102 | 22.3577 | 50.1739 | 0.006732 | 0.32638 | 0.002197 | 48.4825 | 0.142 | 2.3692 | 0.33641 | 16.6851 | 21.9127 | 3.459 | 0.33861 | 0.8696 | 0.6614 | 0.7975 | 0.445
2012-10-30 23:01 | 26.2282 | 22.7626 | 22.6345 | 22.6905 | 22.3661 | 50.164 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.14199 | 2.3693 | 0.33642 | 16.6861 | 21.9692 | 3.4583 | 0.33862 | 0.7934 | 0.6653 | 0.7213 | 0.3969
2012-10-30 23:02 | 26.2371 | 22.7555 | 22.5953 | 22.6634 | 22.3429 | 50.1687 | 0.006729 | 0.32625 | 0.002195 | 48.4848 | 0.14193 | 2.3684 | 0.33615 | 16.6866 | 21.95 | 3.4586 | 0.33834 | 0.8055 | 0.6453 | 0.7134 | 0.3929
2012-10-30 23:02 | 26.2401 | 22.8185 | 22.6103 | 22.7104 | 22.3579 | 50.164 | 0.006729 | 0.32637 | 0.002196 | 48.502 | 0.14194 | 2.3683 | 0.33616 | 16.6857 | 21.953 | 3.4583 | 0.33835 | 0.8655 | 0.6573 | 0.7574 | 0.4049
2012-10-30 23:02 | 26.2386 | 22.765 | 22.5848 | 22.6729 | 22.3284 | 50.1669 | 0.006716 | 0.32628 | 0.002191 | 48.5829 | 0.14202 | 2.3695 | 0.33652 | 16.6846 | 21.9435 | 3.4585 | 0.33871 | 0.8215 | 0.6413 | 0.7294 | 0.3849
2012-10-30 23:02 | 26.2304 | 22.7888 | 22.6006 | 22.7007 | 22.3442 | 50.1583 | 0.006738 | 0.32638 | 0.002199 | 48.4393 | 0.14203 | 2.37 | 0.33661 | 16.6867 | 21.9834 | 3.4579 | 0.33881 | 0.8054 | 0.6172 | 0.7173 | 0.3608
2012-10-30 23:02 | 26.2379 | 22.8083 | 22.6041 | 22.7282 | 22.3397 | 50.1617 | 0.006716 | 0.3263 | 0.002191 | 48.585 | 0.14206 | 2.3702 | 0.33672 | 16.6845 | 21.9347 | 3.4581 | 0.33891 | 0.8736 | 0.6694 | 0.7935 | 0.405
2012-10-30 23:02 | 26.2364 | 22.7988 | 22.6026 | 22.6907 | 22.3743 | 50.1623 | 0.006725 | 0.32627 | 0.002194 | 48.5157 | 0.14206 | 2.3703 | 0.33673 | 16.685 | 21.9734 | 3.4582 | 0.33892 | 0.8254 | 0.6292 | 0.7173 | 0.4009
2012-10-30 23:02 | 26.2274 | 22.7898 | 22.5936 | 22.6977 | 22.3092 | 50.1646 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.14201 | 2.3696 | 0.33652 | 16.686 | 21.9563 | 3.4583 | 0.33872 | 0.8335 | 0.6373 | 0.7414 | 0.3529
2012-10-30 23:02 | 26.2282 | 22.7826 | 22.6345 | 22.6985 | 22.326 | 50.1727 | 0.006738 | 0.3263 | 0.002199 | 48.4264 | 0.14193 | 2.3685 | 0.33615 | 16.6882 | 21.9772 | 3.4589 | 0.33835 | 0.8054 | 0.6573 | 0.7213 | 0.3488
2012-10-30 23:02 | 26.2282 | 22.7906 | 22.5704 | 22.7386 | 22.3701 | 50.1629 | 0.006729 | 0.32631 | 0.002196 | 48.4934 | 0.14195 | 2.3684 | 0.33621 | 16.6845 | 21.9491 | 3.4582 | 0.3384 | 0.8415 | 0.6213 | 0.7895 | 0.421
2012-10-30 23:02 | 26.2319 | 22.8183 | 22.6542 | 22.7302 | 22.3738 | 50.1675 | 0.006729 | 0.32633 | 0.002196 | 48.4955 | 0.14198 | 2.3691 | 0.33637 | 16.6859 | 22.013 | 3.4585 | 0.33857 | 0.8053 | 0.6412 | 0.7172 | 0.3608
2012-10-30 23:02 | 26.2289 | 22.7833 | 22.6712 | 22.7793 | 22.4068 | 50.1496 | 0.006754 | 0.32638 | 0.002204 | 48.3245 | 0.14201 | 2.3695 | 0.3365 | 16.685 | 21.9979 | 3.4573 | 0.33871 | 0.7854 | 0.6733 | 0.7814 | 0.4089
2012-10-30 23:02 | 26.2252 | 22.8396 | 22.6274 | 22.7316 | 22.3911 | 50.1577 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.14196 | 2.3686 | 0.33625 | 16.6851 | 22.0263 | 3.4579 | 0.33845 | 0.8133 | 0.6011 | 0.7053 | 0.3648
2012-10-30 23:02 | 26.2244 | 22.8068 | 22.6547 | 22.7107 | 22.3783 | 50.1739 | 0.006725 | 0.32631 | 0.002194 | 48.5222 | 0.14199 | 2.3692 | 0.33639 | 16.6859 | 21.9975 | 3.459 | 0.33858 | 0.8093 | 0.6572 | 0.7132 | 0.3808
2012-10-30 23:02 | 26.2162 | 22.8066 | 22.6545 | 22.7105 | 22.3821 | 50.175 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.14204 | 2.3701 | 0.33665 | 16.6856 | 22.0093 | 3.4591 | 0.33884 | 0.7973 | 0.6452 | 0.7012 | 0.3728
2012-10-30 23:02 | 26.2199 | 22.8223 | 22.6262 | 22.7623 | 22.3938 | 50.1652 | 0.006735 | 0.32631 | 0.002198 | 48.4502 | 0.14208 | 2.3705 | 0.33681 | 16.684 | 21.993 | 3.4584 | 0.33901 | 0.8293 | 0.6332 | 0.7693 | 0.4008
2012-10-30 23:02 | 26.2177 | 22.8401 | 22.672 | 22.7361 | 22.4277 | 50.1704 | 0.006729 | 0.32634 | 0.002196 | 48.4977 | 0.14208 | 2.3707 | 0.33683 | 16.6857 | 22.0349 | 3.4587 | 0.33903 | 0.8052 | 0.6371 | 0.7012 | 0.3928
2012-10-30 23:02 | 26.2207 | 22.7991 | 22.659 | 22.7391 | 22.3706 | 50.1791 | 0.006716 | 0.32637 | 0.002192 | 48.5959 | 0.14205 | 2.3703 | 0.33672 | 16.686 | 21.9697 | 3.4593 | 0.33891 | 0.8294 | 0.6893 | 0.7694 | 0.4009
2012-10-30 23:02 | 26.2169 | 22.8473 | 22.6872 | 22.7072 | 22.3668 | 50.1687 | 0.006725 | 0.3262 | 0.002194 | 48.505 | 0.14208 | 2.3707 | 0.33683 | 16.685 | 21.9699 | 3.4586 | 0.33902 | 0.8774 | 0.7173 | 0.7373 | 0.3969
2012-10-30 23:02 | 26.2162 | 22.8106 | 22.6585 | 22.7386 | 22.3621 | 50.1669 | 0.006713 | 0.32628 | 0.00219 | 48.6046 | 0.14204 | 2.3701 | 0.33664 | 16.6867 | 21.9532 | 3.4585 | 0.33883 | 0.8574 | 0.7053 | 0.7854 | 0.4089
2012-10-30 23:02 | 26.2139 | 22.7963 | 22.6562 | 22.7283 | 22.4119 | 50.1658 | 0.006732 | 0.32628 | 0.002197 | 48.4675 | 0.14207 | 2.3706 | 0.3368 | 16.6856 | 22.003 | 3.4584 | 0.339 | 0.7933 | 0.6532 | 0.7253 | 0.4089
2012-10-30 23:02 | 26.2072 | 22.7896 | 22.6535 | 22.7376 | 22.4052 | 50.1629 | 0.00671 | 0.32625 | 0.002189 | 48.6221 | 0.14206 | 2.3704 | 0.33675 | 16.6856 | 21.9762 | 3.4582 | 0.33894 | 0.8134 | 0.6773 | 0.7614 | 0.429
2012-10-30 23:02 | 26.2065 | 22.8169 | 22.6368 | 22.7409 | 22.4205 | 50.1704 | 0.006722 | 0.3263 | 0.002193 | 48.5417 | 0.14202 | 2.3698 | 0.33655 | 16.6867 | 21.9595 | 3.4587 | 0.33874 | 0.8574 | 0.6773 | 0.7814 | 0.461
2012-10-30 23:02 | 26.2102 | 22.8446 | 22.6405 | 22.7326 | 22.3841 | 50.1698 | 0.006719 | 0.3263 | 0.002192 | 48.5633 | 0.14199 | 2.3693 | 0.33642 | 16.6862 | 21.9873 | 3.4587 | 0.33861 | 0.8573 | 0.6532 | 0.7453 | 0.3968
2012-10-30 23:02 | 26.2042 | 22.8266 | 22.6585 | 22.7426 | 22.3821 | 50.164 | 0.006722 | 0.32627 | 0.002193 | 48.5373 | 0.14204 | 2.37 | 0.33664 | 16.6858 | 21.9692 | 3.4583 | 0.33883 | 0.8574 | 0.6893 | 0.7734 | 0.4129
2012-10-30 23:02 | 26.205 | 22.8194 | 22.6593 | 22.7514 | 22.395 | 50.1773 | 0.006732 | 0.32631 | 0.002197 | 48.4718 | 0.1421 | 2.3708 | 0.33688 | 16.6839 | 21.97 | 3.4592 | 0.33908 | 0.8494 | 0.6893 | 0.7814 | 0.425
2012-10-30 23:02 | 26.2012 | 22.8637 | 22.5994 | 22.7476 | 22.4112 | 50.1767 | 0.006729 | 0.32622 | 0.002195 | 48.4805 | 0.14206 | 2.3706 | 0.33677 | 16.6871 | 21.9382 | 3.4592 | 0.33896 | 0.9255 | 0.6612 | 0.8094 | 0.473
2012-10-30 23:02 | 26.2079 | 22.8303 | 22.6982 | 22.7263 | 22.4019 | 50.1658 | 0.006735 | 0.32627 | 0.002197 | 48.4437 | 0.14203 | 2.3697 | 0.33657 | 16.6851 | 21.9569 | 3.4584 | 0.33876 | 0.8734 | 0.7413 | 0.7694 | 0.445
2012-10-30 23:02 | 26.2087 | 22.8391 | 22.655 | 22.7591 | 22.3746 | 50.1721 | 0.006735 | 0.32627 | 0.002197 | 48.4437 | 0.14199 | 2.369 | 0.33637 | 16.6845 | 21.9216 | 3.4589 | 0.33857 | 0.9175 | 0.7334 | 0.8375 | 0.453
2012-10-30 23:03 | 26.1975 | 22.856 | 22.6638 | 22.7479 | 22.3474 | 50.1715 | 0.006725 | 0.32627 | 0.002194 | 48.5157 | 0.14203 | 2.3698 | 0.33657 | 16.6853 | 21.9505 | 3.4588 | 0.33876 | 0.9055 | 0.7133 | 0.7974 | 0.3969
2012-10-30 23:03 | 26.1982 | 22.8487 | 22.6365 | 22.8006 | 22.3681 | 50.1721 | 0.006725 | 0.3263 | 0.002194 | 48.52 | 0.14204 | 2.3699 | 0.33661 | 16.6853 | 21.9151 | 3.4589 | 0.33881 | 0.9336 | 0.7214 | 0.8855 | 0.453
2012-10-30 23:03 | 26.2012 | 22.8717 | 22.6715 | 22.7396 | 22.3751 | 50.1617 | 0.006725 | 0.3262 | 0.002194 | 48.505 | 0.1421 | 2.3708 | 0.33689 | 16.6848 | 21.9542 | 3.4581 | 0.33908 | 0.9175 | 0.7173 | 0.7854 | 0.4209
2012-10-30 23:03 | 26.1945 | 22.861 | 22.6888 | 22.8049 | 22.4045 | 50.1652 | 0.006725 | 0.32625 | 0.002194 | 48.5136 | 0.14209 | 2.3708 | 0.33686 | 16.685 | 21.9716 | 3.4584 | 0.33906 | 0.8894 | 0.7172 | 0.8333 | 0.4329